## Sheet1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21811495&pc=FRN | 2016493-2025-72182 | 2025-03-13 04:00 | Death | CAREFUSION SD | 2025-04-10 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem | Low Blood Pressure/ Hypotension | K133532 |  | 1 | Event Description: IT WAS MENTIONED IN THE ISMP MEDICATION SAFETY ALERT ARTICLE (MARCH 13, 2025, VOLUME 30, ISSUE 5) THAT, "A HOSPITAL RECENTLY IMPLEMENTED INTEROPERABILITY BETWEEN THEIR ELECTRONIC HEALTH RECORD (EHR) AND SMART INFUSION PUMPS. FOR A PATIENT WITH ATRIAL FIBRILLATION, A PRESCRIBER ORDERED DEXTROSE 5% WATER (D5W) TO INFUSE AT 75 ML/HOUR AS A MAINTENANCE INTRAVENOUS (IV) FLUID, AND AMIODARONE (450 MG/250 ML IN D5W) TO INFUSE AT 0.5 MG/MIN (16.7 ML/HOUR). TO INITIATE THE INFUSIONS AND EHR DOCUMENTATION, THE NURSE SCANNED THE PATIENT¿S IDENTIFICATION (ID) BAND, INFUSION PUMP, AND BARCODE ON THE D5W BAG. SHE THEN SCANNED THE PATIENT¿S ID BAND, A SECOND INFUSION PUMP CHANNEL, AND THE PHARMACY LABEL BARCODE ON THE AMIODARONE BAG TO INITIATE THAT INFUSION AND DOCUMENT ADMINISTRATION. LATER, THE PRESCRIBER INCREASED THE D5W MAINTENANCE INFUSION RATE TO 100 ML/ HOUR. AROUND THE SAME TIME, THE PATIENT'S AMIODARONE INFUSION WAS NEARLY COMPLETED, SO THE NURSE OBTAINED A REPLACEMENT BAG FROM THE PHARMACY. DURING THE EHR BARCODE SCANNING STEPS TO ASSOCIATE THE NEW AMIODARONE BAG WITH THE INFUSION PUMP, THE NURSE SCANNED THE PATIENT¿S ID BAND, THE PUMP CHANNEL THAT WAS INFUSING AMIODARONE, AND A BARCODE ON THE AMIODARONE INFUSION BAG. THE PUMP GENERATED AN ALERT, ¿CHANNEL IS CURRENTLY INFUSING,¿ WITH ERROR DETAILS, ¿STOP THE CHANNEL OR USE A DIFFERENT CHANNEL AND RESEND THE ORDER DETAILS TO THE PUMP.¿ TO RESOLVE THIS, THE NURSE STOPPED THE CURRENT INFUSION OF AMIODARONE AND CLEARED THE PUMP SETTINGS. SHE THEN RESTARTED THE PROCESS BY SCANNING THE PATIENT¿S ID BAND, THE PUMP, AND A BARCODE ON THE AMIODARONE INFUSION BAG, AND DID NOT RECEIVE ANY ALERTS ON THE PUMP. HOWEVER, AN EHR ALERT WAS GENERATED STATING ¿THIS PUMP IS ASSOCIATED TO A DIFFERENT ORDER.¿ NOT UNDERSTANDING WHAT THE ALERT MEANT AND/ OR DUE TO ALERT FATIGUE, THE NURSE BYPASSED THE WARNING AND INITIATED THE INFUSION. IT TURNED OUT THAT THE NURSE HAD SCANNED THE MANUFACTURER¿S BARCODE ON THE DILUENT BAG (D5W) THAT THE PHARMACY USED TO COMPOUND THE AMIODARONE RATHER THAN THE PHARMACY GENERATED BARCODE IDENTIFYING THE COMPOUNDED AMIODARONE PRODUCT. THE EHR ASSOCIATED THE ¿D5W¿ SCAN (ACTUALLY THE AMIODARONE BAG) WITH THE D5W MAINTENANCE INFUSION ORDER AND DISASSOCIATED THE PRIOR D5W MAINTENANCE INFUSION. SINCE THE DISASSOCIATION WAS NOT A DISCONTINUATION OF THE INFUSION, THE D5W MAINTENANCE INFUSION CONTINUED TO INFUSE AT 75 ML/HOUR INSTEAD OF 100 ML/HOUR AS ORDERED. HOWEVER, THE ASSOCIATED AMIODARONE IN D5W WAS ADMINISTERED AT 100 ML/HOUR INSTEAD OF THE ORDERED RATE OF 16.7 ML/HOUR. APPROXIMATELY 2.5 HOURS LATER, THE PUMP ALARMED THAT THE INFUSION BAG HAD REACHED THE VOLUME TO BE INFUSED (250 ML), AND THE NURSE IDENTIFIED THE ERROR. THE PATIENT BECAME HYPOTENSIVE, A RAPID RESPONSE TEAM WAS ACTIVATED, AND THE PATIENT WAS TREATED WITH IV FLUIDS. HOWEVER, THE PATIENT DIED THE FOLLOWING DAY. ALTHOUGH THE ERROR WAS DETERMINED TO NOT BE THE IMMEDIATE CAUSE OF THE PATIENT'S DETERIORATION, IT MAY HAVE POTENTIALLY EXACERBATED THEIR UNDERLYING ILLNESS. THE ORGANIZATION IDENTIFIED SEVERAL CONTRIBUTING FACTORS: MEDICATION LABELING ISSUES RESULTING IN MULTIPLE SCANNABLE BARCODES ON THE INFUSION BAG AS WELL AS STAFF UNFAMILIARITY WITH THE NEW INTEROPERABILITY TECHNOLOGY AND HOW TO RESPOND TO ALERTS." BD CONTACTED ISMP REQUESTING FOR ADDITIONAL INFORMATION. ISMP RESPONDED; THEY ARE UNABLE TO SHARE ANY INFORMATION AS THEY RECEIVED THE REPORT ANONYMOUSLY. Manufacturer Narrative: THE ACTUAL DATE OF EVENT IS UNKNOWN. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. ROOT CAUSE : THE ROOT CAUSE FOR THE REPORTED ISSUE THAT DEVICE HAD A SAFETY ALERT: DEATH - PROGRAMMING ERROR NO DEVICES WERE AVAILABLE FOR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21811497&pc=FRN | 2016493-2025-72183 | 2025-03-13 04:00 | Death | CAREFUSION SD | 2025-04-10 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem | Low Blood Pressure/ Hypotension | K133532 |  | 1 | Event Description: IT WAS MENTIONED IN THE ISMP MEDICATION SAFETY ALERT ARTICLE (MARCH 13, 2025, VOLUME 30, ISSUE 5) THAT, "A HOSPITAL RECENTLY IMPLEMENTED INTEROPERABILITY BETWEEN THEIR ELECTRONIC HEALTH RECORD (EHR) AND SMART INFUSION PUMPS. FOR A PATIENT WITH ATRIAL FIBRILLATION, A PRESCRIBER ORDERED DEXTROSE 5% WATER (D5W) TO INFUSE AT 75 ML/HOUR AS A MAINTENANCE INTRAVENOUS (IV) FLUID, AND AMIODARONE (450 MG/250 ML IN D5W) TO INFUSE AT 0.5 MG/MIN (16.7 ML/HOUR). TO INITIATE THE INFUSIONS AND EHR DOCUMENTATION, THE NURSE SCANNED THE PATIENT¿S IDENTIFICATION (ID) BAND, INFUSION PUMP, AND BARCODE ON THE D5W BAG. SHE THEN SCANNED THE PATIENT¿S ID BAND, A SECOND INFUSION PUMP CHANNEL, AND THE PHARMACY LABEL BARCODE ON THE AMIODARONE BAG TO INITIATE THAT INFUSION AND DOCUMENT ADMINISTRATION. LATER, THE PRESCRIBER INCREASED THE D5W MAINTENANCE INFUSION RATE TO 100 ML/ HOUR. AROUND THE SAME TIME, THE PATIENT'S AMIODARONE INFUSION WAS NEARLY COMPLETED, SO THE NURSE OBTAINED A REPLACEMENT BAG FROM THE PHARMACY. DURING THE EHR BARCODE SCANNING STEPS TO ASSOCIATE THE NEW AMIODARONE BAG WITH THE INFUSION PUMP, THE NURSE SCANNED THE PATIENT¿S ID BAND, THE PUMP CHANNEL THAT WAS INFUSING AMIODARONE, AND A BARCODE ON THE AMIODARONE INFUSION BAG. THE PUMP GENERATED AN ALERT, ¿CHANNEL IS CURRENTLY INFUSING,¿ WITH ERROR DETAILS, ¿STOP THE CHANNEL OR USE A DIFFERENT CHANNEL AND RESEND THE ORDER DETAILS TO THE PUMP.¿ TO RESOLVE THIS, THE NURSE STOPPED THE CURRENT INFUSION OF AMIODARONE AND CLEARED THE PUMP SETTINGS. SHE THEN RESTARTED THE PROCESS BY SCANNING THE PATIENT¿S ID BAND, THE PUMP, AND A BARCODE ON THE AMIODARONE INFUSION BAG, AND DID NOT RECEIVE ANY ALERTS ON THE PUMP. HOWEVER, AN EHR ALERT WAS GENERATED STATING ¿THIS PUMP IS ASSOCIATED TO A DIFFERENT ORDER.¿ NOT UNDERSTANDING WHAT THE ALERT MEANT AND/ OR DUE TO ALERT FATIGUE, THE NURSE BYPASSED THE WARNING AND INITIATED THE INFUSION. IT TURNED OUT THAT THE NURSE HAD SCANNED THE MANUFACTURER¿S BARCODE ON THE DILUENT BAG (D5W) THAT THE PHARMACY USED TO COMPOUND THE AMIODARONE RATHER THAN THE PHARMACY GENERATED BARCODE IDENTIFYING THE COMPOUNDED AMIODARONE PRODUCT. THE EHR ASSOCIATED THE ¿D5W¿ SCAN (ACTUALLY THE AMIODARONE BAG) WITH THE D5W MAINTENANCE INFUSION ORDER AND DISASSOCIATED THE PRIOR D5W MAINTENANCE INFUSION. SINCE THE DISASSOCIATION WAS NOT A DISCONTINUATION OF THE INFUSION, THE D5W MAINTENANCE INFUSION CONTINUED TO INFUSE AT 75 ML/HOUR INSTEAD OF 100 ML/HOUR AS ORDERED. HOWEVER, THE ASSOCIATED AMIODARONE IN D5W WAS ADMINISTERED AT 100 ML/HOUR INSTEAD OF THE ORDERED RATE OF 16.7 ML/HOUR. APPROXIMATELY 2.5 HOURS LATER, THE PUMP ALARMED THAT THE INFUSION BAG HAD REACHED THE VOLUME TO BE INFUSED (250 ML), AND THE NURSE IDENTIFIED THE ERROR. THE PATIENT BECAME HYPOTENSIVE, A RAPID RESPONSE TEAM WAS ACTIVATED, AND THE PATIENT WAS TREATED WITH IV FLUIDS. HOWEVER, THE PATIENT DIED THE FOLLOWING DAY. ALTHOUGH THE ERROR WAS DETERMINED TO NOT BE THE IMMEDIATE CAUSE OF THE PATIENT'S DETERIORATION, IT MAY HAVE POTENTIALLY EXACERBATED THEIR UNDERLYING ILLNESS. THE ORGANIZATION IDENTIFIED SEVERAL CONTRIBUTING FACTORS: MEDICATION LABELING ISSUES RESULTING IN MULTIPLE SCANNABLE BARCODES ON THE INFUSION BAG AS WELL AS STAFF UNFAMILIARITY WITH THE NEW INTEROPERABILITY TECHNOLOGY AND HOW TO RESPOND TO ALERTS." BD CONTACTED ISMP REQUESTING FOR ADDITIONAL INFORMATION. ISMP RESPONDED; THEY ARE UNABLE TO SHARE ANY INFORMATION AS THEY RECEIVED THE REPORT ANONYMOUSLY. Manufacturer Narrative: THE ACTUAL DATE OF EVENT IS UNKNOWN. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. ROOT CAUSE: THE ROOT CAUSE FOR THE REPORTED ISSUE THAT DEVICE HAD A SAFETY ALERT: DEATH - PROGRAMMING ERROR NO DEVICES WERE AVAILABLE FOR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21812800&pc=BZD | 2518422-2025-043077 | 2022-03-19 04:00 | Death | RESPIRONICS, INC. | 2025-04-10 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGES VISUALIZATION OF PARTICLES IN DEVICE. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: AS PER ADDITIONAL INFORMATION RECEIVED ON 05/20/2025 : THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING VISUALIZATION OF PARTICLES AND PATIENT PASSED AWAY WITH DEVICE NOT BEING INVOLVED. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IN THIS REPORT, BOX B, D, H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21813307&pc=NIQ | 2024168-2025-03712 | 2025-03-03 05:00 | Death | ABBOTT VASCULAR INC. | 2025-04-10 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Obstruction/Occlusion | P110019 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PRESENTED WITH ST ELEVATION MYOCARDIAL INFARCTION (STEMI) AND WENT INTO CARDIOGENIC SHOCK WITH ENSUING CIRCULATORY COLLAPSE. THE PROCEDURE WAS TO TREAT A HEAVY OCCLUDED LESION IN THE RIGHT CORONARY ARTERY (RCA). THE LESION WAS PRE-DILATED WITH A 2.50X12MM AND A 3.50X15MM BALLOON CATHETERS FOLLOWED BY IMPLANTATION OF A 3.0X38MM XIENCE SKYPOINT STENT AND A 3.0X33MM XIENCE SKYPOINT STENT. THERE WAS SEVERE NO FLOW TREATED WITH MEDICATION. DURING INTERVENTIONAL EFFORTS THE PATIENT CODED ONCE AND ADVANCED CARDIOVASCULAR LIFE SUPPORT (ALCS) WAS PROVIDED WITH INTUBATION ON MECHANICAL VENTILATION AND RETURN OF SPONTANEOUS CIRCULATION (ROSC). THE PROCEDURE WAS COMPLETED HOWEVER SHORTLY AFTER SUCCESSFUL REVASCULARIZATION THE PATIENT WENT INTO PULSELESS ELECTRICAL ACTIVITY AND DESPITE RESUSCITATIVE EFFORTS HE WAS PRONOUNCED DECEASED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECTS OF OCCLUSION AND DEATH ARE LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (IFU) AS KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21813384&pc=OZD | 1220648-2025-27354 | 2025-03-16 04:00 | Death | ABIOMED INC. | 2025-04-10 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT HAD LIMB ISCHEMIA. THE STAFF WAS UNABLE TO DOPPLER DISTAL PULSES. THE PUMP WAS EXPLANTED AND AN IMPELLA 5.5 WAS IMPLANTED. THE PATIENT EXPIRED ON SUPPORT OF THE REPLACEMENT PUMP. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21813624&pc=DQY | 21813624 | 2025-03-21 04:00 | Death | PENUMBRA, INC. | 2025-04-10 04:00 | DQY | PENUMBRA NEURON MAX | Material Fragmentation | Intracranial Hemorrhage; Brain Injury; Foreign Body Embolism |  |  | 1 | Event Description: A 64-YEAR-OLD FEMALE TRANSFERRED (B)(6) 2025 FROM AN OUTLYING HOSPITAL TO (B)(6) HOSPITAL FOLLOWING AN UNWITNESSED FALL WITH SUBSEQUENT APHASIA AND RIGHT-SIDED WEAKNESS. PATIENT WAS RECENTLY DISCHARGED ((B)(6) 2025) FOLLOWING AN ADMISSION FOR NSTEMI AND GI BLEED. CTA HEAD AND NECK IDENTIFIED AN M1 OCCLUSION. ON INITIAL ASSESSMENT AT SMH, NIHSS OF 23 WITH FORCED GAZE DEVIATION AND FACIAL DROOP ON THE RIGHT, 0/5 ON RIGHT UPPER AND RIGHT LOWER EXTREMITY, HEMIANOPIA/NO BLINK ON THE RIGHT. SHE WAS NOT A CANDIDATE FOR TNK DUE TO RECENT GI BLEED AND PATIENT WAS TAKEN TO THE INTERVENTIONAL RADIOLOGY SUITE FOR MECHANICAL THROMBECTOMY. WITH SOME DIFFICULTY, THE STENT AT THE ORIGIN OF THE LEFT COMMON CAROTID ARTERY WAS SUCCESSFULLY NEGOTIATED. UPON REMOVAL OF THE INTERNAL CATHETER FROM THE SHEATH, THE TIP OF THE DEVICE WAS NOTED TO BE SHEARED OFF. CONTRAST INJECTION DEMONSTRATED MINIMAL FLOW INTO THE LEFT MIDDLE CEREBRAL ARTERY AS WELL AS THE LEFT ANTERIOR CEREBRAL ARTERY DUE TO THE FRAGMENT. ATTEMPTS TO RETRIEVE THE FRAGMENT WERE UNSUCCESSFUL FOLLOWED BY GROSS SUBARACHNOID HEMORRHAGE. SUBSEQUENT COIL OCCLUSION OF THE LEFT INTERNAL CAROTID ARTERY WAS PERFORMED TO CONTROL BLEEDING. AT THIS POINT, THE PROCEDURE WAS ABANDONED AND NO ATTEMPT TO PERFORM MECHANICAL THROMBECTOMY COULD BE MADE. POST-PROCEDURE PATIENT REMAINED INTUBATED, NOT SEDATED, REQUIRING 4 VASOPRESSORS. PT WITH NO COUGH/NO GAG/NO CORNEAL REFLEX. PUPILS 7MM AND NONREACTIVE BILATERALLY. DISCUSSION WITH SPOUSE/FAMILY GIVEN HEMODYNAMIC INSTABILITY AND UNLIKELY MEANINGFUL OUTCOME, FAMILY DECIDED TO TRANSITION TO COMFORT CARE AND PATIENT EXPIRED (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21813837&pc=DQY | 2124215-2025-22000 | 2025-04-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-10 04:00 | DQY | WATCHMAN TRUSEAL ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION AND PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSEAL ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN WAS USING TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) TO PERFORM THE TRANSEPTAL PUNCTURE. AFTER CROSSING THE FOSSA OVALIS THROUGH THE PATENT FORAMEN OVALE (PFO), THE PHYSICIAN OBSERVED THAT THE PUNCTURE SITE WAS TOO SUPERIOR, RESULTING IN NON-COAXIAL ALIGNMENT WITH THE OSTIUM OF THE LEFT ATRIAL APPENDAGE. THE PHYSICIAN ATTEMPTED A SECOND TRANSEPTAL PUNCTURE. HOWEVER, WHILE THE TRANSEPTAL SHEATH REMAINED IN THE RIGHT ATRIUM, THE PATIENT EXPERIENCED HYPOTENSION AND DEVELOPED A PERICARDIAL EFFUSION. THE PHYSICIANS COULD NOT DETERMINATE THE EXACT LOCATION OF THE PERFORATION, BUT THEY SUSPECTED TO BE IN THE RIGHT ATRIAL APPENDAGE. PERICARDIOCENTESIS WAS PERFORMED TO TREAT THE EFFUSION. THE PATIENT STABILIZED FOR APPROXIMATELY THIRTY MINUTES BEFORE SUFFERING A FATAL CARDIAC ARREST.Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION AND PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSEAL ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN WAS USING TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) TO PERFORM THE TRANSEPTAL PUNCTURE. AFTER CROSSING THE FOSSA OVALIS THROUGH THE PATENT FORAMEN OVALE (PFO), THE PHYSICIAN OBSERVED THAT THE PUNCTURE SITE WAS TOO SUPERIOR, RESULTING IN NON-COAXIAL ALIGNMENT WITH THE OSTIUM OF THE LEFT ATRIAL APPENDAGE. THE PHYSICIAN ATTEMPTED A SECOND TRANSEPTAL PUNCTURE. HOWEVER, WHILE THE TRANSEPTAL SHEATH REMAINED IN THE RIGHT ATRIUM, THE PATIENT EXPERIENCED HYPOTENSION AND DEVELOPED A PERICARDIAL EFFUSION. THE PHYSICIANS COULD NOT DETERMINATE THE EXACT LOCATION OF THE PERFORATION, BUT THEY SUSPECTED TO BE IN THE RIGHT ATRIAL APPENDAGE. PERICARDIOCENTESIS WAS PERFORMED TO TREAT THE EFFUSION. THE PATIENT STABILIZED FOR APPROXIMATELY THIRTY MINUTES BEFORE SUFFERING A FATAL CARDIAC ARREST. Manufacturer Narrative: H6: IMPACT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21813879&pc=GDW | 3005075853-2025-02632 | 2025-04-03 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-10 04:00 | GDW | PROXIMATE*RL VASCULAR LIN STAP | Failure to Fire; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K020779 |  | 1 | Event Description: IT WAS REPORTED THAT THE SURGEON WAS PERFORMING A LIVER TRANSPLANT USING THE TX30 AFTER USING THE STAPLER THE SURGEON THEN CUT THE VEIN WHICH LED TO A BLEED, MAJOR HEMORRHAGE PROTOCOL WAS ACTIVATED, HOWEVER AFTER THIS BLEED THERE WAS ANOTHER MAJOR BLEED NOT DIRECTLY LINKED TO THE STAPLER. THE PATIENT WAS ULTIMATELY CLOSED SENT TO ICU. THE SURGEON INSPECTED THE VEIN AND COULD NOT SEE ANY STAPLES. HE HAS ALSO CALLED PATHOLOGY TO SEE IF THEY COULD SEE ANY STAPLES ON THE LIVER THERE WAS ALSO NONE ON THE LIVER. PATIENT DEATH AFTER THE OPERATION WAS COMPLETED Manufacturer Narrative: (B)(4). DATE SENT 4/9/2025. D4 BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. NO LOT OR BATCH NUMBER WAS PROVIDED THEREFORE A DEVICE HISTORY COULD NOT BE DONE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: DID THE SURGEON RELEASE THE DEVICE AND INSPECT THE STAPLE LINE PRIOR TO DIVIDING THE VESSEL? DID THE SURGEON PLACE A VASCULAR CLAP ACROSS THE VESSEL PRIOR TO DIVIDING? WOULD THE SURGEON LIKE TO SPEAK WITH ETHICON? DOES THE SURGEON BELIEVE THE PATIENTS DEATH WAS RELATED TO AN ALLEGED DEFICIENCY OF THE DEVICE OR WERE THERE OTHER CONTRIBUTING PATIENT FACTORS? PLEASE EXPLAIN. WHAT IS THE SURGEONS EXPERIENCE WITH THE DEVICE? WHAT IS THE SURGEONS EXPERIENCE WITH THE COMPETITIVE DEVICE? WAS THE CLOSING TRIGGER CLOSED PRIOR TO TRANSECTION? WAS THE DEVICE TRIGGER CLOSED PRIOR TO TRANSECTION? WAS THIS THE FIRST FIRING OR SECOND FIRING? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? HOW WAS THE BLEEDING ADDRESSED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WAS THE BLEEDING INTRALUMINAL OR EXTRALUMINAL? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 5/12/2025. ADDITIONAL INFORMATION: DURING LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STAPLING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. Manufacturer Narrative: (B)(4). DATE SENT: 5/15/2025. ADDITIONAL INFORMATION: 1. WHAT IS THE SURGEONS EXPERIENCE WITH THE DEVICE? I AM VERY FAMILIAR WITH ETHICON TX LINEAR STAPLER AND I HAVE BEEN USING FOR MANY YEARS SINCE I JOINED EDINBURGH TRANSPLANT UNIT IN 2015. I CANNOT PRECISELY SPECIFY WHEN EXACTLY I STARTED USING IT OR HOW MANY TIMES BUT I WOULD CERTAINLY ESTIMATE I USED IT TENS OF TIMES, BECAUSE NOT ONLY I USE IT FOR CONTROLLING HEPATIC VEINS DURING LIVER TRANSPLANT HEPATECTOMY, BUT ALSO FOR THE LIVER GRAFT BACK TABLE PREPARATION (TO CLOSE IVC TOP AND OR BOTTOM ENDS). 2. WAS THE CLOSING TRIGGER CLOSED PRIOR TO TRANSECTION? YES. 3. WAS THE DEVICE TRIGGER CLOSED PRIOR TO TRANSECTION? YES. 4. WAS THIS THE FIRST FIRING OR SECOND FIRING? IT WAS THE SECOND FIRING. 5. WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. THE STAPLE FORMATION APPEARED AS IF THE STAPLE LINES WERE PRESENT LIKELY DUE TO THE MARKS FORMED BY THE CLOSING TRIGGER ON THE VEIN, IN ADDITION TO THE LIMITATION POSED BY THE ANGLE OF VIEW WHEN THE RIGHT HEPATIC VEIN WAS STAPLED. HENCE, THERE WERE NO ISSUES NOTED WITH THE STAPLE FORMATION AT THAT STAGE. THEN SUBSEQUENTLY, SIGNIFICANT BLEEDING OCCURRED FROM THE OPEN RIGHT HEPATIC VEIN WHEN IT WAS CUT DUE TO ABSENCE OF THE STAPLES. ABSENCE OF STAPLES ON THE VEIN WAS EVIDENCED AT LATER STAGES WITH NO STAPLES OBSERVED ON PATIENT SIDE (OPEN VEIN ON THE INFERIOR VENA CAVA) AND ON THE EXPLANTED LIVER SIDE AFTER IT WAS CROSS-EXAMINED AT THE PATHOLOGY DEPARTMENT. THE CUSTOMER IS NOT ALLOWED TO ANSWER ANY QUESTIONS ABOUT THE PATIENT, THEY CAN ONLY ANSWER QUESTIONS DIRECTLY LINKED TO THE STAPLER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21813912&pc=PYX | 1220648-2025-27363 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-04-10 04:00 | PYX | IMPELLA RP FLEX | Failure to Advance | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT A PATIENT WAS HAVING WATCHMAN PROCEDURE IN THE CATH LAB WHEN THE LEFT ATRIUM WAS PERFORATED. CPR WAS INITIATED AND THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM FOR EMERGENT STERNOTOMY/REPAIR. DECISION WAS MADE FOR BIPELLA. THERE WERE CONCERNS OF PATIENT HAVING BILATERAL UPPER LIMB ISCHEMIA. VASCULAR HAD BEEN CONSULTED AND DISCUSSION AMONGST THE HEART TEAM ON APPROACH DUE TO VESSEL SIZES. THE DECISION WAS MADE TO PLACE IMPELLA CP VIA RIGHT FEMORAL ARTERY AND PLACE RP FLEX VIA RIJ. IT WAS NOTED THAT RP FLEX WOULD NOT ADVANCE INTO PULMONIC VALVE. THE DECISION WAS MADE TO REMOVE THE RP FLEX AND PATIENT BEGAN BLEEDING WHERE THE PATIENT WAS PLACED ON CARDIOPULMONARY BYPASS SUPPORT AND EMERGENT STERNOTOMY. THE IMPELLA CP WAS ALREADY IN THE LEFT VENTRICLE. ULTIMATELY THE PATIENT EXPIRED IN THE PROCEDURAL AREA. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21814461&pc=BZD | 2518422-2025-103936 | 2025-02-24 05:00 | Death | RESPIRONICS, INC. | 2025-04-10 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER TO REQUEST THAT SHE NO LONGER RECEIVE LETTERS REGARDING THE REPLACEMENT OF THE CPAP DEVICE. SHE INDICATED THAT SHE HAS NOT HAD POSSESSION OF THE MACHINE FOR APPROXIMATELY TWO TO TWO AND A HALF YEARS. THERE HAS BEEN NO SUGGESTION THAT THE DEVICE PLAYED ANY ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21814537&pc=DXY | 2182208-2025-01814 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-10 04:00 | DXY | MEDTRONIC ICD | Appropriate Term/Code Not Available | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTATION IN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY (ICM) VERSUS NONISCHEMIC CARDIOMYOPATHY (NICM). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS CARDIOVASCULAR, SUDDEN CARDIAC DEATH, OR NON-CARDIOVASCULAR. THERE WAS PATIENT WHO EXPERIENCED A POSTOPERATIVE INFECTION AND HAD THEIR DEVICE SYSTEM REMOVED. THERE WERE OTHER PATIENTS WHO UNDERWENT SURGICAL TREATMENT FOR POCKETS BECAUSE OF LOCALIZED HEMATOMA OR POOR INCISION HEALING, ISCHEMIC STROKE OR ACUTE CORONARY SYNDROME THAT RECEIVED PERCUTANEOUS CORONARY INTERVENTION TREATMENT, RE-HOSPITALIZATION DUE TO HEART FAILURE, OR INAPPROPRIATE THERAPY; HOWEVER, THE CAUSES OF THE INAPPROPRIATE THERAPY WERE UNKNOWN. THERE WERE LEADS WHICH WERE REPOSITIONED DUE TO DISLODGEMENT. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/62 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPLANTABLE CARDIOVERTER DEFIBRILLATOR IN NONISCHEMIC VERSUS ISCHEMIC CARDIOMYOPATHY: REAL-WORLD PRIMARY PREVENT ION STUDY. PACING AND CLINICAL ELECTROPHYSIOLOGY. 2025. 48:169¿179. DOI: 10.1111/PACE.15142. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21814582&pc=DXY | 2182208-2025-01815 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-10 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTATION IN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY (ICM) VERSUS NONISCHEMIC CARDIOMYOPATHY (NICM). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS CARDIOVASCULAR, SUDDEN CARDIAC DEATH, OR NON-CARDIOVASCULAR. THERE WAS PATIENT WHO EXPERIENCED A POSTOPERATIVE INFECTION AND HAD THEIR DEVICE SYSTEM REMOVED. THERE WERE OTHER PATIENTS WHO UNDERWENT SURGICAL TREATMENT FOR POCKETS BECAUSE OF LOCALIZED HEMATOMA OR POOR INCISION HEALING, ISCHEMIC STROKE OR ACUTE CORONARY SYNDROME THAT RECEIVED PERCUTANEOUS CORONARY INTERVENTION TREATMENT, RE-HOSPITALIZATION DUE TO HEART FAILURE, OR INAPPROPRIATE THERAPY; HOWEVER, THE CAUSES OF THE INAPPROPRIATE THERAPY WERE UNKNOWN. THERE WERE LEADS WHICH WERE REPOSITIONED DUE TO DISLODGEMENT. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/62 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPLANTABLE CARDIOVERTER DEFIBRILLATOR IN NONISCHEMIC VERSUS ISCHEMIC CARDIOMYOPATHY: REAL-WORLD PRIMARY PREVENT ION STUDY. PACING AND CLINICAL ELECTROPHYSIOLOGY. 2025. 48:169¿179. DOI: 10.1111/PACE.15142 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21814608&pc=MIH | 3007284313-2025-03898 | 2025-02-06 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-10 04:00 | MIH | GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS | Malposition of Device; Migration | Pneumonia; Renal Impairment | P200030 |  | 1 | Event Description: IT WAS REPORTED TO GORE, THAT A PATIENT UNDERWENT AN ENDOVASCULAR ANEURYSM REPAIR (EVAR) ON (B)(6) 2024, WITH A GORE® EXCLUDER® AAA ENDOPROSTHESIS. AN FINAL ANGIOGRAPHY AT THE END OF THE PROCEDURE INDICATED THAT THE GRAFT WAS POSITIONED ONE OR TWO MILLIMETERS TOO HIGH BUT WITH GOOD FLOW DOWN RENAL ARTERIES. IT WAS ADDITIONALLY REPORTED THAT A BLEEDING FROM GROIN REQUIRED PATCH ENDARTERECTOMY, WITH AN UNKNOWN DEVICE(S) INTRA-PROCEDURE EVENT. THE PATIENT WAS DISCHARGED FROM HOSPITAL ON (B)(6) 2024 WITH NORMAL RENAL FUNCTION, WITH PATIENT DOING WELL. IT WAS STATED THAT ANGIOGRAPHY IMAGES FROM INDEX PROCEDURE WERE REEXAMINED AND IT APPEARED THAT THE TRUNK MIGRATED NORTHWARDS/PROXIMALLY DURING THE PROCEDURE. REPORTEDLY, THE COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) ON (B)(6) 2025 WAS PERFORMED AND PATIENT CONDITION REMAINED STABLE. BUT ON (B)(6) 2025, AT THE EVENING, THE PATIENT ATTENDED EMERGENCY DEPARTMENT (ED) IN THE HOSPITAL ON THE ISLAND (B)(6) VERY UNWELL. DURING THE HOSPITALIZATION IN THAT EMERGENCY, THE PHYSICIAN STATED TO GORE: CTA (B)(6) 2025 SHOWED RIGHT RENAL INFARCTION, PATIENT STABLE BUT EGFR 22, CRP <5, PLANNED TO DISCUSS AT MDT (MULTIDISCIPLINARY TEAM APPROACH). CREATININE 700+. ALSO HAD CONSOLIDATION ON CHEST X-RAY WHICH WAS TO DIAGNOSE THE PNEUMONIA. DECEASED ON (B)(6) 2025 WITH UNKNOWN CAUSE OF DEATH. REPORTEDLY, THE PATIENT HAD AN ARREST, CARDIOPULMONARY RESUSCITATION (CPR) WITH RETURN OF SPONTANEOUS CIRCULATION (ROSC) BUT ARRESTED AGAIN, TIME OF DEATH 7:07 P.M. THE PATIENT'S DIAGNOSIS FROM (B)(6) 2025 SHOWED A SUSPECTED EUGLYCAEMIC DKA SECONDARY TO DAPAGLIFLOZIN, OR SEVERE ACUTE KIDNEY INJURY (AKI).Event Description: IT WAS REPORTED TO GORE, THAT A 78-YEAR-OLD PATIENT UNDERWENT AN ENDOVASCULAR ANEURYSM REPAIR ON (B)(6) 2024, WITH GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS DEVICES. A FINAL INTRA-PROCEDURE IMAGING INDICATED THAT THE GORE® EXCLUDER® CONFORMABLE AAA TRUNK-IPSILATERAL LEG ENDOPROSTHESIS WAS POSITIONED ONE OR TWO MILLIMETERS TOO HIGH BUT WITH GOOD FLOW DOWN INTO RENAL ARTERIES. ON (B)(6) 2024, THE PATIENT WAS DISCHARGED FROM THE HOSPITAL WITH NORMAL RENAL FUNCTION, WITH THE PATIENT DOING WELL. POST-PROCEDURE (AT UNKNOWN DATE), ANGIOGRAPHY IMAGES FROM THIS INDEX PROCEDURE WERE RE-EXAMINED AND IT APPEARED THAT THE TRUNK MIGRATED PROXIMALLY DURING THE PROCEDURE. ON (B)(6) 2025, A COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) IMAGING WAS PERFORMED SHOWING RIGHT RENAL INFARCTION AND STABLE PATIENT. HOWEVER, LABORATORY VALUES ESTIMATED GLOMERULAR FILTRATION RATE (EGFR) 22, C-REACTIVE PROTEIN (CRP) <5, AND CREATININE 700+. DURING AN MDT (MULTIDISCIPLINARY TEAM APPROACH) THEY HAD CONSOLIDATION ON A CHEST X-RAY, WHICH WAS TO DIAGNOSE THE PNEUMONIA. FURTHER, IT HAS WAS DIAGNOSED A SUSPECTED EUGLYCEMIC DIABETIC KETOACIDOSIS SECONDARY TO DAPAGLIFLOZIN, OR SEVERE ACUTE KIDNEY INJURY (AKI). ON (B)(6) 2025, AT THE EVENING, THE PATIENT ATTENDED EMERGENCY DEPARTMENT (ED) IN THE HOSPITAL ON THE (B)(6) VERY UNWELL. REPORTEDLY, THE PATIENT HAD EXPERIENCED A CARDIAC ARREST AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, RESULTING IN RETURN OF SPONTANEOUS CIRCULATION (ROSC), BUT ARRESTED AGAIN. THE PHYSICIAN REPORTED THAT THE PATIENT DECEASED ON (B)(6) 2025 AT 7:07 P.M. AND THE CAUSE OF DEATH WAS THE CARDIAC ARREST. Manufacturer Narrative: A1: PATIENT IDENTIFIER IS TRUNCATED BY THE SYSTEM, THE FULL IDENTIFIER REPORTED TO GORE WAS (B)(6). H3: THE IMPLANTED GORE DEVICE IS NOT RETURNED BECAUSE IT REMAINS IMPLANTED, A PRODUCT EVALUATION IS NOT POSSIBLE. ALSO, NO OTHER GORE DEVICES WERE IMPLANTED DURING THE INDEX PROCEDURE THAT NEEDS TO REPORT ON. H6: CODE B13, FURTHER INFORMATION REQUESTS ARE MADE AND MAKE INTERVIEWS. CODE B15, THERE ARE IMAGES REQUESTED AND AVAILABLE FOR EVALUATION BY GORE. CODE B14: ANALYSIS OF PRODUCTION RECORDS WAS PERFORMED BY GORE WITH THE FOLLOWING RESULT, A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET PRE-RELEASE SPECIFICATIONS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: UPDATED SECTION B5: DESCRIBE EVENT OR PROBLEM HAS ADDITIONAL INFORMATION. H6, TYPE OF INVESTIGATION: CODE B11 ADDED. MEDICAL DEVICE PROBLEM: CODE A050408 ADDED. COMPONENT CODES, IMDRF ANNEX G: CODE G04122 ADDED. INVESTIGATION FINDINGS: CODE C19 ADDED. CODE C21 IS NO LONGER APPLICABLE. INVESTIGATION CONCLUSIONS: CODE D15 AND D12 ADDED. CODE D16 IS NO LONGER APPLICABLE. HEALTH EFFECT, IMPACT CODES: CODE F0801 AND F15 ADDED. HEALTH EFFECT, CLINICAL CODES: CODE E0733 AND CODE E2121 ADDED. FOR H6, CODE B15: THE MANUFACTURING RECORDS WERE REVIEWED AND DOCUMENTED IN THE PRODUCT HISTORY TASK. THE DEVICE LOT MET ALL PRE-RELEASE SPECIFICATIONS. INVESTIGATION SUMMARY: THE PRIMARY REPORTED DEVICE FAILURE MODE, PROXIMAL DEVICE MIGRATION DURING THE INDEX PROCEDURE, WAS CONFIRMED DURING EVALUATION OF THE PROVIDED CLINICAL IMAGES. THE CLINICAL IMAGES FROM THE INDEX PROCEDURE SHOW THE PROXIMAL END OF THE DEVICE LOCATED JUST DISTAL TO THE RENAL ARTERIES AFTER RE-CONSTRAINING THE DEVICE, AND THE DEVICE WAS OBSERVED PARTIALLY COVERING THE RIGHT RENAL ARTERY (RRA) AFTER A SUBSEQUENT BALLOONING OPERATION. FLOW WAS MAINTAINED IN BOTH RENAL ARTERIES DESPITE THE PARTIAL RRA COVERAGE. APPROXIMATELY 50 DAYS AFTER THE INDEX PROCEDURE FOLLOW-UP IMAGING SHOWS THE RRA IS OCCLUDED. A RELATIONSHIP BETWEEN THE INTRA-PROCEDURAL BALLOON OPERATION AND THE PROXIMAL MIGRATION OF THE DEPLOYED ENDOPROSTHESIS, WHICH WAS OBSERVED AFTER THE BALLOON SEQUENCE, COULD BE NEITHER CONFIRMED NOR DISMISSED. THE CAUSE FOR THE CONFIRMED PROXIMAL DEVICE MIGRATION COULD NOT BE ESTABLISHED WITH THE AVAILABLE INFORMATION. THE TYPE II ENDOLEAK OBSERVED IN THE EVALUATION OF PROVIDED CLINICAL IMAGES IS NOT REASONABLY SUGGESTIVE OF A DEVICE MALFUNCTION. THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS DEVICE INSTRUCTIONS FOR USE (IFU) WERE REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD, AND THE FOLLOWING IFU STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. POTENTIAL DEVICE OR PROCEDURE RELATED ADVERSE EVENTS. ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE BUT ARE NOT LIMITED TO: ENDOLEAK / THIS WAS ADDED FOR CODE 'D12' AS FOR ANNEX D - INVESTIGATION CONCLUSIONS. GORE IFU CONTAINS MORE POTENTIAL ADVERSE EVENTS RELATED TO PROCEDURE AND DEATH IS INCLUDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21815141&pc=GDW | 3005075853-2025-02636 | 2025-03-13 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-10 04:00 | GDW | ENDOPATH ECHELON VASCULAR WHITE RELOAD FOR ADVANCED PLACEMENT TIP (35MM, 4 ROW) | Failure to Form Staple; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K141952 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A SEGMENTECTOMIA , STAPLING USING VASCULAR CHARGE IN THE ARTERY BECAME ´DROOLING´ FOR A WHILE. THEY MANAGED TO STALL BUT AFTER, THE STAPLES FELL BY OPENING THE STAPLING AND THERE WAS LEAKAGE. THE SURGERY WAS DELAYED BY 40 MINUTES. THE PATIENT EXPIRED. Manufacturer Narrative: (B)(4). DATE SENT: 5/13/2025. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. THE LOT#326D66 HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE EQUIPMENT. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE ISSUE WAS WITH THE ENSEAL THAT WAS THE DEVICE WITH THE ONE HE DID HEMOSTASIS ONCE THE STAPLER DIDN¿T WORK, THE PATIENT BLED 24 HOURS POST-OP IN THE PLACE HE DID THE HEMOSTASIS WITH THE ENSEAL. THE COAGULATION TEST OF THE PATIENT WERE NORMAL BUT ONE RELATIVE ACCEPTED THAT THE PATIENT HAVE HAD A COAGULOPATHY ONCE IN HER LIFE, WHICH WAS NO NOTED IN HER MEDICAL HISTORY. HIS QUESTION WAS RELATED TO BLEEDING CASES HOURS AFTER USING THE ENSEAL AND IF THERE IS A RESTRICTION TO USE ENSEAL IN PATIENTS WITH COAGULOPATHY. AN EXTERNAL RAPID RESPONSE CALL WILL BE SCHEDULED. Manufacturer Narrative: (B)(4). DATE SENT: 6/11/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: AN EXTERNAL RAPID RESPONSE CALL WAS HELD ON (B)(6) 2025 TO INCLUDE POST MARKET SURVEILLANCE, CUSTOMER QUALITY, MEDICAL SAFETY, RESEARCH AND DEVELOPMENT, QUALITY, THE LOCAL TEAM AND THE OPERATING SURGEON TO DISCUSS THE HEMOSTASIS ISSUE AND THE PATIENT¿S UNFORTUNATE PASSING. THE SURGEON PROVIDED A SUMMARY OF THE EVENTS. WHEN STAPLING AND CUTTING THE BRONCHI DURING THE RIGHT LOBECTOMY, IT STARTED TO BLEED. HE HAD TO OPEN THE SURGERY TO PUT SOME CLIPS IN TO STOP THE BLEEDING. AFTER COMPLETING THE PROCEDURE, THE PATIENT WENT TO THE ICU. THE INTUBATION TUBE WAS REMOVED THE NEXT MORNING. IN THE AFTERNOON, THE SURGEON TOLD THE PATIENT ABOUT THE PROCEDURE. AT ABOUT 8PM THE PATIENT HAD COUGH AND SEIZERS AND BLED 1.5 LITERS OF BLOOD. HE TOOK THE PATIENT BACK INTO SURGERY. HE SAW THAT THE ARTERY WHERE THE ENSEAL WAS USED IN THE ORIGINAL PROCEDURE (NOT THE STAPLER) WAS BLEEDING. IT WAS A DIFFERENT VESSEL THAN WAS BLEEDING IN THE ORIGINAL PROCEDURE. THE SURGEON ATTEMPTED TO SUTURE THE BLEED, BUT THE PATIENT ULTIMATELY PASSED. PATIENT HAD COAGULOPATHY ISSUES IN THE PAST, BUT THIS WAS NOT KNOWN DURING EITHER PROCEDURE. Manufacturer Narrative: (B)(4). DATE SENT: 4/10/2025. D4: BATCH # UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: ADDITIONAL INFORMATION RECEIVED FROM SALES REP. WHAT WERE THE INDICATIONS FOR SURGERY? LUNG NODULES. WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? DOCTORS HAVE BEEN USING OUR MATERIAL FOR OVER FOUR YEARS. WHAT IS THE COMPRESSED WIDTH AND THICKNESS OF THE VESSEL? 2MM. WHAT IS THE ESTIMATED COMPRESSED WIDTH OF THE TARGET VESSEL? DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? YES. DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? YES. DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? YES. WERE THE HEMOSTASIS ISSUES IDENTIFIED POST OPERATIVELY OR INTRA OPERATIVELY? POST-OPERATIVE WERE ANY SURGICAL CLIPS USED IN THE AREA OF THE DEVICE FIRING? AFTER OPENING THE SHEARS, IT WAS NECESSARY TO PLACE A CLIP WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? NO. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? BLEEDING AFTER OPENING THE SEAL IS ESTIMATED AT 2 LITERS. WAS A TRANSFUSION REQUIRED? YES. WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? NO INFORMATION. WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? NO INFORMATION WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? COAGULOPATHY PLEASE PROVIDE A TIMELINE OF EVENTS. AFTER THE CLAMPING OF THE VESSELS AND ARTERY, AFTER 1 HOUR OF THE PROCEDURE THE ARTERY CLAMPING OPENED. TWO CLIPS WERE PLACED IN THE SITE AND THERE WAS HEMOSTASIS. IN THE CASE OF THE SEAL OF THE ENSEAL, AFTER 24 HOURS BLEEDING AND OPENING OF THE SEAL WERE IDENTIFIED. IS THERE AN AUTOPSY REPORT AVAILABLE? NO. IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? YES. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21815282&pc=DSQ | 2916596-2025-01973 | 2024-11-01 04:00 | Death | THORATEC CORPORATION | 2025-04-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT WAS ON HOSPICE, AND THEIR DEVICE WAS TURNED OFF. THEY HAD AN INCURABLE DRIVELINE INFECTION. NO EQUIPMENT WAS RETURNING, AND THE DEVICE WAS FUNCTIONING AS INTENDED. A CULTURE FROM (B)(6) 2024 GREW STAPHYLOCOCCUS AUREUS. THE ONSET OF THE INFECTION WAS JUN2020. THE PATIENT WAS SAID TO BE VERY NON-COMPLIANT, WAS IN AND OUT OF JAIL, AND DID NOT TAKE THEIR MEDICATIONS ROUTINELY. DURING THEIR LAST ADMISSION IN (B)(6) 2025, THEY WERE POSITIVE FOR COCAINE AND METHAMPHETAMINES. ULTIMATELY, THEY DECIDED TO ENTER HOSPICE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE HEARTMATE 3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION. THE IFU ALSO PROVIDES SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21815651&pc=LWR | 3005687633-2025-00106 | 2025-02-17 05:00 | Death | CORCYM S.R.L. | 2025-04-10 04:00 | LWR | PERCEVAL PLUS SUTURELESS AORTIC HEART VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding; Rupture | P150011S013 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED THAT ON (B)(6) 2025, A PATIENT UNDERWENT AORTIC VALVE REPLACEMENT WITH A PERCIVAL PLUS SUTURELESS AORTIC BIOPROSTHESIS SIZE SMALL. THE VALVE WAS NEEDED TO BE EXPLANTED DUE TO LVOT RUPTURE AND EXCESSIVE BLEEDING. BASED ON FURTHER INFORMATION RECEIVED, THE PATIENT REMAINED STABLE THROUGHOUT THE SURGERY. HOWEVER, THE PATIENT'S CONDITION DETERIORATED DURING THE RE-EXPLORATION PERIOD, ULTIMATELY LEADING TO THEIR PASSING. ACCORDING TO MEDICAL JUDGEMENT, LVOT RUPTURE MIGHT HAVE OCCURRED WHILE TAKING THE NADIR SUTURES AND DURING DECALCIFICATION. THE MANUFACTURER HAS BEEN INFORMED THAT THE INVOLVED DEVICE WILL NOT BE RETURNED FOR ANALYSIS AS IT WAS DISCARDED AT THE HOSPITAL FACILITY.Event Description: THE MANUFACTURER WAS INFORMED THAT ON (B)(6) 2025, A PATIENT UNDERWENT AORTIC VALVE REPLACEMENT WITH A PERCEVAL PLUS SUTURELESS AORTIC BIOPROSTHESIS SIZE SMALL. THE VALVE WAS NEEDED TO BE EXPLANTED DUE TO LVOT RUPTURE AND EXCESSIVE BLEEDING. BASED ON FURTHER INFORMATION RECEIVED, THE PATIENT REMAINED STABLE THROUGHOUT THE SURGERY. HOWEVER, THE PATIENT'S CONDITION DETERIORATED DURING THE RE-EXPLORATION PERIOD, ULTIMATELY LEADING TO THEIR PASSING. ACCORDING TO MEDICAL JUDGEMENT, LVOT RUPTURE MIGHT HAVE OCCURRED WHILE TAKING THE NADIR SUTURES AND DURING DECALCIFICATION. THE MANUFACTURER HAS BEEN INFORMED THAT THE INVOLVED DEVICE WILL NOT BE RETURNED FOR ANALYSIS AS IT WAS DISCARDED AT THE HOSPITAL FACILITY. ADDITIONALLY, IT HAS BEEN CONFIRMED THAT AN AUTOPSY WAS NOT PERFORMED.Event Description: THE MANUFACTURER WAS INFORMED THAT ON (B)(6) 2025; A PATIENT UNDERWENT AORTIC VALVE REPLACEMENT WITH A PERCEVAL PLUS SUTURELESS AORTIC BIOPROSTHESIS SIZE SMALL. THE VALVE WAS NEEDED TO BE EXPLANTED DUE TO LVOT RUPTURE AND EXCESSIVE BLEEDING. BASED ON FURTHER INFORMATION RECEIVED, THE PATIENT REMAINED STABLE THROUGHOUT THE SURGERY. HOWEVER, THE PATIENT'S CONDITION DETERIORATED DURING THE RE-EXPLORATION PERIOD, ULTIMATELY LEADING TO THEIR PASSING. THE MANUFACTURER WAS INFORMED THAT AN AUTOPSY WAS NOT PERFORMED. ACCORDING TO MEDICAL JUDGEMENT, LVOT RUPTURE MIGHT HAVE OCCURRED WHILE TAKING THE NADIR SUTURES AND DURING DECALCIFICATION. ADDITIONALLY, THE SURGEON DETERMINED THAT THE EVENT WAS DUE TO SURGICAL COMPLICATIONS UNRELATED TO THE PERCEVAL VALVE. THE MANUFACTURER HAS BEEN INFORMED THAT THE INVOLVED DEVICE WILL NOT BE RETURNED FOR ANALYSIS AS IT WAS DISCARDED AT THE HOSPITAL FACILITY. Manufacturer Narrative: CORRECTED FIELDS: H6. SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS, THE MANUFACTURER COULD NOT PERFORM FURTHER INVESTIGATION ON THE INVOLVED PERCEVAL PROSTHESIS. ALTHOUGH ONLY LIMITED INVESTIGATIONS COULD BE CONDUCTED, BASED ON THE MEDICAL JUDGMENT PROVIDED, IT IS REASONABLE TO CONCLUDE THAT THE EVENT WAS DUE TO SURGICAL COMPLICATIONS RELATED TO THE AVR PROCEDURE AND NOT ASSOCIATED WITH THE PERCEVAL VALVE. Manufacturer Narrative: NOTE: THE DEVICE IS DISTRIBUTED OUTSIDE THE US AND NO GUDID INFORMATION EXISTS. FDA PREMARKET SUBMISSION NUMBER OF SIMILAR DEVICE: P150011 (SUPPLEMENT NUMBER 013); PRIMARY DI NUMBER FOR SIMILAR DEVICE: (B)(4). THE MANUFACTURING AND MATERIAL RECORDS FOR THE DEVICE INVOLVED IN THIS EVENT, AS THEY PERTAIN TO THE REPORTED ISSUE, WERE RETRIEVED AND REVIEWED BY THE MANUFACTURER¿S QUALITY ENGINEERING. THE RESULTS CONFIRMED THAT THIS PROSTHESIS SATISFIED ALL REQUIRED MATERIAL, VISUAL, AND PERFORMANCE STANDARDS AT THE TIME OF MANUFACTURE AND RELEASE. THE MANUFACTURER IS FOLLOWING UP WITH THE HEALTHCARE FACILITY TO RETRIEVE ADDITIONAL INFORMATION ON THE REPORTED EVENT AND THE DEVICE INVOLVED. A FOLLOW UP REPORT WILL BE SUBMITTED UPON RECEIPT OF ANY NEW INFORMATION OR AT THE COMPLETION OF THE INVESTIGATION. Manufacturer Narrative: THE MANUFACTURER IS CONTINUING EFFORTS TO FOLLOW UP ON THE EVENT, THE DEVICE INVOLVED, AND THE PATIENT'S RISK FACTORS AND MEDICAL CONDITIONS. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON RECEIPT OF ANY ADDITIONAL INFORMATION OR UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21815777&pc=DSQ | 3007042319-2025-00748 | 2022-10-11 04:00 | Death | HEARTWARE, INC. | 2025-04-10 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus; Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A META-ANALYSIS REVIEW ON THE EFFECT OF POST-IMPLANT PHOSPHODIESTERASE-5 INHIBITORS ON HEMOCOMPATIBILITY-RELATED EVENTS, SUCH AS PUMP THROMBOSIS, STROKE, GASTROINTESTINAL BLEEDS (GIBS), AND MORTALITY IN PATIENTS WITH CONTINUOUS-FLOW LEFT VENTRICULAR ASSIST DEVICES (VADS). FIFTEEN JOURNAL ARTICLES WITH VARIOUS STUDIES WERE USED IN EXTRACTING THE DATA. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED PUMP THROMBOSIS, ISCHEMIC STROKE, AND GIB. PUMP THROMBOSIS WAS DEFINED AS PRESENCE OF HEMOLYSIS, ABNORMAL PUMP PARAMETERS, OR ONE OR MORE OF THE FOLLOWING: STROKE OR TRANSIENT ISCHEMIC ATTACK, ARTERIAL NON-CENTRAL NERVOUS SYSTEM THROMBOEMBOLISM, DE NOVO NEED FOR INOTROPE THERAPY, TREATMENT WITH INTRAVENOUS ANTI-COAGULANTS, INTRAVENOUS THROMBOLYTICS, OR INTRAVENOUS ANTIPLATELET THERAPY, PUMP REPLACEMENT, PUMP EXPLANTATION WITH OR WITHOUT EXCHANGE, PUMP DEACTIVATION WITHOUT PUMP REMOVAL, OPERATION TO REPAIR OR REPLACE ANY INTERNAL COMPONENT OF THE CIRCULATORY SUPPORT SYSTEM, OR WITHIN THE BLOOD-CONTACTING SURFACES OF DEVICE INFLOW CANNULA OR OUTFLOW GRAFTS. ISCHEMIC STROKE WAS DIAGNOSED WITH NEUROIMAGING EVIDENCE OF CENTRAL NERVOUS SYSTEM INFARCTION IN THE CORRESPONDING VASCULAR TERRITORY. GIB WAS DEFINED AS BLEEDING FROM THE UPPER GASTROINTESTINAL TRACT, BLEEDING FROM THE LOWER GASTROINTESTINAL TRACT, OR BLEEDING OF UNKNOWN ORIGIN BUT WITH GUAIAC-POSITIVE STOOLS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS UNKNOWN. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: POST-IMPLANT PHOSPHODIESTERASE-5 INHIBITORS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICE: A SYSTEMATIC REVIEW AND META-ANALYSIS. JOURNAL OF CLINICAL MEDICINE. 2022. 11, 5988. DOI: 10.3390/JCM11205988, D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21815994&pc=LGW | 3008514029-2025-00288 | 2025-03-02 05:00 | Death | NEVRO CORP. | 2025-04-10 04:00 | LGW | SENZA | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Sepsis; Insufficient Information | P130022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. AN HCP ASSESSMENT REGARDING THE CAUSE OF DEATH WAS NOT AVAILABLE. THERE WERE NO REPORTS OF DEVICE-RELATED ISSUES FROM THE PATIENT PRIOR TO THE PASSING. Manufacturer Narrative: THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELEVANT NONCONFORMITIES WERE FOUND. NEVRO SUBMITS THIS REPORT IN COMPLIANCE WITH FDA'S MEDICAL DEVICE REPORTING REGULATIONS UNDER 21 CFR PART 803. NEVRO HAS COMPLIED WITH REGULATORY INVESTIGATION REQUIREMENTS AND IS SUBMITTING ALL INFORMATION THAT IS REASONABLY KNOWN TO US AT THIS TIME. HOWEVER, WE MAY NOT HAVE BEEN ABLE TO CONFIRM THIS INFORMATION OR COMPLETE THE INVESTIGATION WITHIN THE TIMEFRAME FOR FILING THIS REPORT. WE MAY HAVE GIVEN NO RESPONSE OR AN INCOMPLETE RESPONSE TO CERTAIN QUESTIONS BECAUSE WE DO NOT CURRENTLY HAVE INFORMATION AVAILABLE TO PROVIDE A COMPLETE RESPONSE. IF WE LATER OBTAIN ANY REQUIRED INFORMATION THAT WAS NOT AVAILABLE AT THE TIME OF THIS INITIAL REPORT, WE WILL SUBMIT A SUPPLEMENTAL REPORT. THIS REPORT IS NOT AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT WAS DEFECTIVE, THAT IT MALFUNCTIONED, OR THAT IT CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THIS REPORT. WE MAY CONCLUDE THAT THE DEVICE HAD NO DEFECT, DID NOT MALFUNCTION, OR DID NOT CAUSE OR CONTRIBUTE TO A REPORTABLE EVENT. SOME OF THE ITEMS ON THIS FORM INCLUDE FORCED-CHOICE TERMS USED BY THE FDA FOR REPORTING PURPOSES THAT DO NOT NECESSARILY REFLECT NEVRO¿S CONCLUSIONS ABOUT THE CAUSES OR NATURE OF THE EVENT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY FREEDOM OF INFORMATION ACT RESPONSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816045&pc=OWB | 3003768277-2025-003382 | 2025-04-07 04:00 | Death | PHILIPS MEDICAL SYSTEMS NEDERLAND B.V. | 2025-04-10 04:00 | OWB | ALLURA XPER FD | Poor Quality Image | Insufficient Information | K162859 |  | 1 | Event Description: IT WAS REPORTED TO PHILIPS THAT DURING AN INTRACRANIAL ARTERY EMBOLIZATION PROCEDURE, THE GUIDEWIRE WAS NOT VISIBLE. THE REPORT INDICATED THAT THE GUIDEWIRE WAS PUSHED TOO DEEP INTO THE VESSEL RESULTING IN A CEREBROVASCULAR INJURY. THE PROCEDURE WAS ABORTED, AND THE PATIENT WAS MOVED TO THE INTENSIVE CARE UNIT (ICU) WHERE THEY LATER PASSED AWAY. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816098&pc=LGW | 3008514029-2025-00290 | 2024-11-07 05:00 | Death | NEVRO CORP. | 2025-04-10 04:00 | LGW | SENZA | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Sepsis; Insufficient Information | P130022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. AN HCP ASSESSMENT REGARDING THE CAUSE OF DEATH WAS NOT AVAILABLE. THERE WERE NO REPORTS OF DEVICE-RELATED ISSUES FROM THE PATIENT PRIOR TO THE PASSING. Manufacturer Narrative: THE MANUFACTURING RECORDS WERE REVIEWED AND NO RELEVANT NONCONFORMITIES WERE FOUND. NEVRO SUBMITS THIS REPORT IN COMPLIANCE WITH FDA'S MEDICAL DEVICE REPORTING REGULATIONS UNDER 21 CFR PART 803. NEVRO HAS COMPLIED WITH REGULATORY INVESTIGATION REQUIREMENTS AND IS SUBMITTING ALL INFORMATION THAT IS REASONABLY KNOWN TO US AT THIS TIME. HOWEVER, WE MAY NOT HAVE BEEN ABLE TO CONFIRM THIS INFORMATION OR COMPLETE THE INVESTIGATION WITHIN THE TIMEFRAME FOR FILING THIS REPORT. WE MAY HAVE GIVEN NO RESPONSE OR AN INCOMPLETE RESPONSE TO CERTAIN QUESTIONS BECAUSE WE DO NOT CURRENTLY HAVE INFORMATION AVAILABLE TO PROVIDE A COMPLETE RESPONSE. IF WE LATER OBTAIN ANY REQUIRED INFORMATION THAT WAS NOT AVAILABLE AT THE TIME OF THIS INITIAL REPORT, WE WILL SUBMIT A SUPPLEMENTAL REPORT. THIS REPORT IS NOT AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT WAS DEFECTIVE, THAT IT MALFUNCTIONED, OR THAT IT CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THIS REPORT. WE MAY CONCLUDE THAT THE DEVICE HAD NO DEFECT, DID NOT MALFUNCTION, OR DID NOT CAUSE OR CONTRIBUTE TO A REPORTABLE EVENT. SOME OF THE ITEMS ON THIS FORM INCLUDE FORCED-CHOICE TERMS USED BY THE FDA FOR REPORTING PURPOSES THAT DO NOT NECESSARILY REFLECT NEVRO¿S CONCLUSIONS ABOUT THE CAUSES OR NATURE OF THE EVENT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY FREEDOM OF INFORMATION ACT RESPONSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816473&pc=DQK | MW5168817 | 2025-02-04 05:00 | Death | BIOSENSE WEBSTER, A JOHNSON & JOHNSON COMPANY | 2025-04-10 04:00 | DQK | CARTO 3 SYSTEM AND STSF DF CURVE CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation |  |  | 1 | Event Description: IT WAS REPORTED THAT WHILE PERFORMING A PROCEDURE IN THE ANIMAL LAB THEY PERFORATED THE AORTA WITH THE SL1 SHEATH AND THE ANIMAL PASSED AWAY. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816564&pc=OZD | 1220648-2025-27380 | 2025-03-16 04:00 | Death | ABIOMED, INC. | 2025-04-10 04:00 | OZD | IMPELLA CP | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WAS A PUMP STOP. STAFF ATTEMPTED TO RESTART THE PUMP TO NO AVAIL. THE PUMP WAS EXPLANTED AND REPLACED. THE PATIENT REMAINED STABLE AND THE PROCEDURAL OUTCOME WAS THE PATIENT SURVIVED SURGERY.Event Description: PER MEDICAL SAFETY OFFICER REVIEW; THE PUMP STOPPED DURING IMPELLA CP PUMP AFTER 12 DAYS START OF SUPPORT REQUIRING REPLACEMENT AND WAS THE CASCADING EVENT WHICH CANNOT BE DISSOCIATED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME OF DEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK CPI SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ Manufacturer Narrative: THE INVESTIGATION OF PUMP STOP HAS BEEN COMPLETED. THE IMPELLA CP RETURNED FOR EVALUATION. UPON VISUAL INSPECTION, LARGE PURGE GAP WAS FOUND. DATA LOGS WERE REVIEWED AND A HIGH MOTOR CURRENT PUMP STOP WITH ALARM OCCURRED ON 11TH DAY. MULTIPLE ATTEMPTS WERE MADE TO RESTART PUMP WITHOUT SUCCESS. THE ROOT CAUSE OF THE PUMP STOP WAS DUE TO BEARING WEAR AS PUMP WAS RUN PAST INDICATED DESIGN LIFE. Manufacturer Narrative: PER THE MEDICAL SAFETY OFFICER REVIEW; THE PUMP STOPPED DURING IMPELLA CP PUMP 1 MCS AFTER 12 DAYS START OF SUPPORT REQUIRING REPLACEMENT AND WAS THE CASCADING EVENT WHICH CANNOT BE DISSOCIATED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. BASED ON THE MEDICAL SAFETY OFFICER (MSO) REVIEW THE FOLLOWING FIELDS HAVE BEEN UPDATED: B1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27380 TO REFLECT BOTH PRODUCT PROBLEM AND ADVERSE EVENT. B2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27380 TO REFLECT DEATH AND DATE OF DEATH B5 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27380 TO REFLECT PATIENT OUTCOME AND MSO STATEMENT H1 TYPE OF REPORTABLE EVENT REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27380 TO REFLECT DEATH H6 HEALTH EFFECT - IMPACT CODE REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27380 TO REFLECT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816567&pc=NPT | 2015691-2025-02785 | 2024-04-08 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-04-10 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, DURING TAVR PROCEDURE OF 23MM SAPIEN 3 VALVE, CEREBRAL INFARCTION OCCURRED. THE PATIENT PASSED AWAY AS SEVERE NEUROLOGICAL IMPAIRMENT DUE TO INTRAOPERATIVE CEREBRAL INFARCTION AND PROGRESSION OF DISUSE SYNDROME DUE TO REPEATED PNEUMONIA. THE PATIENT PASSED AWAY AFTER BEING TRANSFERRED TO A LONG-TERM CARE HOSPITAL. THE EXACT DATE OF DEATH WAS UNKNOWN AND NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. LIMITED INFORMATION HAS BEEN RECEIVED FROM THE TAVI REGISTRY DATA AND THE EXACT CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN BUT IS LIKELY RELATED TO ONE OR MORE OF THE FACTORS AND/OR MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816587&pc=NPT | 2015691-2025-02786 | 2023-12-22 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-04-10 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, DURING TAVR PROCEDURE OF 29MM SAPIEN 3 ULTRA RESILIA VALVE, EXTENSIVE EMBOLIC CEREBRAL INFARCTION OCCURRED AND THE PATIENT EXPERIENCED PERSISTENT IMPAIRED CONSCIOUSNESS. EVENTUALLY, THE PATIENT PASSED AWAY AS SEVERE NEUROLOGICAL IMPAIRMENT DUE TO INTRAOPERATIVE CEREBRAL INFARCTION AND PROGRESSION OF DISUSE SYNDROME DUE TO REPEATED PNEUMONIA AT OTHER HOSPITAL. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. DESPITE REQUESTS FOR MORE INFORMATION, LIMITED INFORMATION WAS RECEIVED; THEREFORE, THE CAUSE OF THE EVENT REMAINS UNKNOWN, BUT MAY BE RELATED TO THE FACTORS AND/OR MECHANISMS MENTIONED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816754&pc=OZD | 1220648-2025-27382 | 2025-03-16 04:00 | Death | ABIOMED, INC. | 2025-04-10 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT EXPERIENCED LIMB ISCHEMIA. THE PATIENT'S IMPELLA LEG WAS LIVID AND HAD POOR CIRCULATION. CATECHOLAMINE COULD NOT BE REDUCED AND THERE WAS NO VASCULAR SURGEON ON SITE. THE PHYSICIAN DID NOT WANT TO DO A CROSSOVER BY PASS BECAUSE THE PATIENT WAS NOT ABLE TO BE TRANSFERRED. AS THE PATIENT WAS NOT WEANABLE FROM THE IMPELLA, IT WAS RECOMMENDED TO REMOVE THE PUMP AND IMPLANT A NEW ONE ON THE OPPOSITE SIDE. CARE WAS WITHDRAW AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21816757&pc=OZD | 1220648-2025-27383 | 2025-03-15 04:00 | Death | ABIOMED, INC. | 2025-04-10 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT HAD SKIN TRACT OOZING AT THEIR IMPELLA CP ACCESS SITE DURING SUPPORT. A PURSE STRING SUTURE WAS PLACED AND ONE UNIT OF PACKED RED BLOOD CELLS WAS GIVEN TO TREAT THE CONDITION. ADDITIONALLY, HEPARIN WAS TEMPORARILY PAUSED AND THE SITE WAS REDRESSED AND OOZING WAS RESOLVED. THE FOLLOWING DAY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE HEMOSTASIS WAS ACHIEVED BY ADDING ADDITIONAL SUTURE. H.6 INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS WERE REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27383 AS THE INVESTIGATION WAS COMPLETE. H.10 INVESTIGATION RESULTS WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27383. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21817489&pc=MNS | 2518422-2025-043477 | 2021-08-09 04:00 | Death | RESPIRONICS, INC. | 2025-04-10 04:00 | MNS | DREAMSTATION ST30 | Material Integrity Problem | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE THIRD-PARTY SERVICE CENTER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. REPORTED THE PATIENT HAS PASSED AWAY AND NO ALLEGATION THE DEVICE CONTRIBUTED TO THE DEATH. THE THIRD-PARTY SERVICE CENTER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21817697&pc=OZD | 1220648-2025-27386 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-10 04:00 | OZD | IMPELLA CP-JP | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Bronchial Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE 80-YEAR-OLD MALE PATIENT HAD A CP PUMP SUPPORT PLACED DURING CPR FOR SUPPORT DUE TO MYOCARDITIS OR PERICARDITIS. THE PUMP WAS PLACED WHILE UNDERGOING CPR BUT, THE PATIENT EXPIRED FROM PULMONARY HEMORRHAGE AND CARDIAC TAMPONADE. IT WAS THOUGHT THAT CPR CAUSED PULMONARY HEMORRHAGE AND CARDIAC TAMPONADE, AND THE DEATH HAD NO RELATIONSHIP TO IMPELLA PUMP USE. HOWEVER, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL AND VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VASCULAR DAMAGE - PERIPHERAL: CLINICAL DETAILS REPORT THAT PULMONARY HEMORRHAGE OCCURRED AS A RESULT OF CARDIOPULMONARY RESUSCITATION (CPR). PRODUCT WAS NOT RETURNED FOR INVESTIGATION AND DATA LOGS ARE NOT RELEVANT TO THE COMPLICATION. SINCE CLINICAL DETAILS REPORT THAT THE INJURY OCCURRED DUE TO CPR, THE ROOT CAUSE OF THE VASCULAR DAMAGE WAS DETERMINED TO MOST LIKELY BE A USE ISSUE. VENTRICLE PERFORATION: CLINICAL DETAILS REPORT THAT IMPELLA WAS INSERTED WHILE CPR WAS BEING DONE. AS A RESULT OF CPR, A CARDIAC TAMPONADE OCCURRED. DATA LOGS WERE REVIEWED, AND THERE WERE NO POSITIONING ALARMS DURING THE CASE. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. BASED ON THE CLINICAL DETAILS PROVIDED THAT THE CARDIAC TAMPONADE WAS A RESULT OF CPR, THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS DETERMINED TO MOST LIKELY BE A USE ISSUE. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER: 1220648-2025-27386.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21817709&pc=QSK | 1063481-2025-00031 |  | Death | ARTIVION, INC. ¿ KENNESA | 2025-04-10 04:00 | QSK | PHOTOFIX UNKNOWN CONFIGURATION | Insufficient Information | Wound Dehiscence; Endocarditis; Valvular Insufficiency/ Regurgitation | K172085 |  | 1 | Event Description: ACCORDING TO INITIAL REPORTS REGARDING THE RECENT PRE-PROOF OF THE ARTICLE TITLED: ¿MID-TERM RESULTS OF OZAKI TECHNIQUE FOR AORTIC VALVE RECONSTRUCTION COMPARING AUTOLOGOUS AND BOVINE PERICARDIUM IN CONGENITAL AORTIC VALVE DISEASE¿ BY CHAVEZ, M, ET AL., WAS RECENTLY RELEASED IN JTCVS STRUCTURAL AND ENDOVASCULAR; "IT REPORTS ON AORTIC VALVE RECONSTRUCTION COMPARING AUTOLOGOUS PERICARDIUM TO PHOTOFIX®. IT IS A RETROSPECTIVE REVIEW OF PATIENTS WITH THREE-LEAFLET RECONSTRUCTION FROM 8/2015 TO 6/2023. THE SERIES REPORTS ON 74 PATIENTS (31 AUTOLOGOUS AND 43 PHOTOFIX). TWENTY-FIVE OF 43 (58%) OF PHOTOFIX PATIENTS REQUIRED AORTIC VALVE REOPERATION. FAILURE MODE FOR PHOTOFIX WAS REPORTED AS DEHISCENCE AT THE ANNULAR SUTURE LINE. AUTHORS CONCLUDE POOR MID-TERM RESULTS WITH PHOTOFIX BOVINE PERICARDIUM. AS NOTED IN OUR LABELING, THE USE OF PHOTOFIX FOR LEAFLET RECONSTRUCTION IS A CONTRAINDICATION." THIS INVESTIGATION IS RELEGATED TO PHOTOFIX UNKNOWN CONFIGURATION FOR THE FOLLOWING 5 EVENTS. DEATH, REGURGITATION, REOPERATION, DEHISCENCE, ENDOCARDITIS. Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC. Manufacturer Narrative: ACCORDING TO INITIAL REPORTS REGARDING THE RECENT PRE-PROOF OF THE ARTICLE TITLED: ¿MID-TERM RESULTS OF OZAKI TECHNIQUE FOR AORTIC VALVE RECONSTRUCTION COMPARING AUTOLOGOUS AND BOVINE PERICARDIUM IN CONGENITAL AORTIC VALVE DISEASE¿ BY CHAVEZ, M, ET AL., WAS RECENTLY RELEASED IN JTCVS STRUCTURAL AND ENDOVASCULAR; "IT REPORTS ON AORTIC VALVE RECONSTRUCTION COMPARING AUTOLOGOUS PERICARDIUM TO PHOTOFIX. IT IS A RETROSPECTIVE REVIEW OF PATIENTS WITH THREE-LEAFLET RECONSTRUCTION FROM 8/2015 TO 6/2023. THE SERIES REPORTS ON 74 PATIENTS (31 AUTOLOGOUS AND 43 PHOTOFIX). TWENTY-FIVE OF 43 (58%) OF PHOTOFIX PATIENTS REQUIRED AORTIC VALVE REOPERATION. FAILURE MODE FOR PHOTOFIX WAS REPORTED AS DEHISCENCE AT THE ANNULAR SUTURE LINE. AUTHORS CONCLUDE POOR MID-TERM RESULTS WITH PHOTOFIX BOVINE PERICARDIUM. AS NOTED IN OUR LABELING, THE USE OF PHOTOFIX FOR LEAFLET RECONSTRUCTION IS A CONTRAINDICATION." THIS INVESTIGATION IS RELEGATED TO PHOTOFIX UNKNOWN CONFIGURATION FOR THE FOLLOWING 5 EVENTS.
DEATH,
REGURGITATION,
REOPERATION,
DEHISCENCE,
ENDOCARDITIS.
MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION HAVE GONE UNMET. NO ADDITIONAL INFORMATION IS FORTHCOMING.
A REVIEW OF MANUFACTURING RECORDS COULD NOT BE PERFORMED AS A DEFINITIVE LOT/SERIAL NUMBER AND DATE OF SURGERY WERE NOT PROVIDED BY THE COMPLAINANT.
A REVIEW OF THE AVAILABLE INFORMATION WAS PERFORMED. PER THE MANUSCRIPT, 2 PHOTOFIX PATIENTS RETURNED WITH VALVE ENDOCARDITIS AND THERE WERE 6 DEATHS (1 AUTOLOGOUS AND 5 PHOTOFIX RECIPIENTS) REPORTED AFTER DISCHARGE. THE CAUSE OF DEATH(S) LISTED IN THE REPORT ARE NOT SPECIFIC TO TISSUE RECIPIENT, NOR ARE THEY STATED TO BE DIRECTLY ATTRIBUTED TO EITHER PATCH USE. AORTIC REGURGITATION WAS NOTED IN BOTH TISSUE TYPES AND PATHOLOGIC EVALUATION WAS PERFORMED ON EXPLANTED TISSUE. FAILURE MODE FOR AUTOLOGOUS PERICARDIUM WAS RELATED TO EXTENSIVE FIBROUS GROWTH AND CALCIFICATION THROUGHOUT THE LEAFLETS, WITH THE PHOTOFIX SAMPLES SHOWING FIBROSIS WITH MYXOID DEGENERATION, PRESENTING WITH DEHISCENCE AT THE SUTURE LINE. THE PRIMARY ENDPOINT OF THE EVALUATION WAS SURGICAL RE-INTERVENTION. BOTH TISSUE TYPES HAD A 97% FREEDOM FROM RE-OPERATION AT 1-YEAR FOLLOW-UP, BUT THERE WAS A STATISTICALLY SIGNIFICANT DIFFERENCE AT FIVE YEARS FOLLOW-UP IN FAVOR OF AUTOLOGOUS PERICARDIUM (76% VS 30%, P=0.006). PATIENTS OF BOTH TISSUE TYPES UNDERWENT CONCOMITANT AORTIC VALVE/CARDIAC PROCEDURES DURING THEIR INCIDENT OPERATION; THE REPORT DOES NOT DELINEATE ANY POTENTIAL CORRELATION OF THE IMPACT OF THESE ON THE REPORTED OUTCOMES WITH EITHER TISSUE TYPE. ADHESION, DEATH, INFECTION, REJECTION AND INFLAMMATORY REACTIONS ARE LISTED AS KNOWN ADVERSE EVENTS FOR PERICARDIAL PATCHES AND ARE LISTED IN THE PHOTOFIX INSTRUCTIONS FOR USE (IFU). AS NOTED IN THE TITLE, THE PATIENT COHORT IN THIS EVALUATION ALL PRESENT WITH CONGENITAL AORTIC VALVE DISEASE; REOPERATION IS NOT AN UNEXPECTED EVENT FOR THIS PATIENT POPULATION. PHOTOFIX MANUFACTURING PROCESS INCLUDES 100% VISUAL INSPECTION PRIOR TO PACKAGING. THE REPORTED EVENTS ARE NOT UNEXPECTED OUTCOMES FOR THIS PATIENT POPULATION. THE RELATIONSHIP BETWEEN THE REPORTED EVENTS AND PHOTOFIX CANNOT BE DETERMINED WITHOUT ADDITIONAL INFORMATION. ADEQUATE PRECAUTIONS ARE LISTING IN OUR LABELING, THE PHOTOFIX IFU SPECIFIES THAT PHOTOFIX IS CONTRAINDICATED FOR VALVE LEAFLET REPAIR.
THE PHOTOFIX DECELLULARIZED BOVINE PERICARDIUM A/DFMEA WAS REVIEWED. THE REPORTED EVENT IS ADDRESSED. THIS COMPLAINT REPORTS ADVERSE EVENTS (AE) IDENTIFIED IN LITERATURE AND DOES NOT SPECIFY POTENTIAL CAUSES OF FAILURE. DUE TO THE LIMITED INFORMATION AND UNKNOWN TIME TO EVENTS, IT CANNOT BE DETERMINED IF THE REPORTED EVENTS ARE DIRECTLY RELATED TO THE PHOTOFIX. AS SUCH, NO RISK OCCURRENCE ANALYSIS WAS PERFORMED IN THIS REPORT. THE REPORTED EVENTS WILL CONTINUE TO BE MONITORED FOR TRENDS. NO UPDATES TO THE RMF ARE NEEDED IN RESPONSE TO THIS COMPLAINT. RISK HAS BEEN REDUCED AS LOW AS POSSIBLE AND OVERALL RESIDUAL RISK IS ACCEPTABLE.
A COMPLAINT AND RECALL QUERY COULD NOT BE PERFORMED AS THE LOT NUMBER IS UNKNOWN.
BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE FOR THIS EVENT CANNOT BE DETERMINED. THIS EVENT DOES NOT IDENTIFY ADDITIONAL HAZARDS OR MODIFY THE PROBABILITY AND SEVERITY OF EXISTING HAZARDS. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC AND THE IFU ADEQUATELY COMMUNICATES RISK. THIS COMPLAINT WAS REVIEWED FOR A CAPA EVALUATION AND A CAPA IS NOT WARRANTED AT THIS TIME. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY.
THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819185&pc=NPT | 2025587-2025-02550 | 2024-03-01 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-10 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 MONTHS AND 24 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A CEREBRAL INFARCTION OCCURRED. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL; HOWEVER, APPROXIMATELY 4 MONTHS LATER THE PATIENT EXPERIENCED ANOTHER CEREBRAL INFARCTION AND PASSED AWAY. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819228&pc=DSQ | 2916596-2025-02445 | 2024-12-28 05:00 | Death | THORATEC CORPORATION | 2025-04-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE WITH INABILITY TO WEAN THE VENT. THE FAMILY DECIDED TO PLACE THE PATIENT ON HOSPICE. CARE WAS WITHDRAWN ON (B)(6) 2024. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) FUNCTIONED WITHOUT ANY ISSUES. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819251&pc=QFG | 3013756811-2025-67680 | 2024-08-31 04:00 | Death | TANDEM DIABETES CARE | 2025-04-10 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: VOLUNTARY MEDWATCH RECEIVED ON 3/13/2025 REPORTED: PATIENT HAD REPORTEDLY COMPLAINED OF MALFUNCTIONING INSULIN PUMP WITH RECURRENT DKA (DIABETIC KETOACIDOSIS). EXPIRED WITH DKA. CALLED TANDEM MULTIPLE TIMES FOR PUMP INTERROGATION AND THEY HAVE NO FOLLOWED UP. A MEDICAL EXAMINER CONTACTED TANDEM ON (B)(6) 2025 TO REPORT THE DEATH EVENT AND REQUEST INTERROGATIONS TO DETERMINE THE CAUSE OF DEATH. NO ALLEGATIONS WERE MADE AGAINST THE PUMP AND THE TIME OF THE REPORT. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT (CTS) TO OBTAIN MORE INFORMATION. ON (B)(6) 2025, CTS SPOKE TO THE EXAMINER. ON (B)(6) 2025, A FOLLOW UP EMAIL WAS SENT TO THE EXAMINER. A THIRD CONTACT ATTEMPT WAS MADE ON (B)(6) 2025 AND CTS DISCUSSED THE PUMP RETURN PROCESS WITH THE MEDICAL EXAMINER. THE MEDICAL EXAMINER DID NOT ALLEGE A DEVICE DEFICIENCY IN ANY OF THESE INTERACTIONS WITH CTS. ON 3/13/2025, TANDEM RECEIVED NOTICE OF THE VOLUNTARY MEDWATCH #MW5167559. THROUGH ITS OWN INTERNAL RECORDS, TANDEM WAS ABLE TO VERIFY THAT THE MEDWATCH WAS REPORTED BY THE SAME MEDICAL EXAMINER WHO CONTACTED TANDEM ON (B)(6) 2025, HOWEVER, BECAUSE THE MEDWATCH DID NOT MENTION THE NAME OF THE PATIENT OR THE PUMP SERIAL NUMBER, IT WAS NOT CLEAR WHETHER IT CONCERNED THE SAME DEATH EVENT. AFTER SEVERAL ATTEMPTS TO REACH OUT TO THE MEDICAL EXAMINER, TANDEM WAS EVENTUALLY ABLE TO CONFIRM THAT THIS DEATH WAS THE SAME INDIVIDUAL AS THE ONE PREVIOUSLY REPORTED BY THEM ON (B)(6) 2025. VOLUNTARY MEDWATCH ALLEGES NEW INFORMATION THAT HAD NOT PREVIOUSLY BEEN REPORTED TO TANDEM. CTS FOLLOWED UP WITH THE MEDICAL EXAMINER WHO CONFIRMED THE PATIENT INFORMATION AND REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024 DUE TO DIABETIC KETOACIDOSIS. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED AND EVALUATION IS PERFORMED. RELATED MEDWATCH REPORT #MW5167559.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819334&pc=DSQ | 2916596-2025-02139 | 2025-03-24 04:00 | Death | THORATEC CORPORATION | 2025-04-10 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY, DUE TO AN UNKNOWN CAUSE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION E3 - OCCUPATION: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819356&pc=NPT | 2025587-2025-02551 | 2024-03-13 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-10 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 MONTHS AND 11 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819428&pc=NPT | 2025587-2025-02555 | 2024-01-17 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-10 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS CARDIAC RELATED BUT WAS NOT SPECIFIED. IT WAS NOTED THAT THE PATIENT HAD HEART FAILURE. PER THE PHYSICIAN, THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE.Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS CARDIAC RELATED, BUT WAS NOT SPECIFIED. IT WAS NOTED THAT THE PATIENT HAD HEART FAILURE. PER THE PHYSICIAN, THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. INFORMATION WAS INITIATED PROVIDED, BUT NOT PREVIOUSLY CAPTURED IN THE EVENT NARRATIVE WHICH REPORTED THE PATIENT'S DEATH OCCURRED MORE THAN SEVEN MONTHS FOLLOWING THE VALVE IMPLANT PROCEDURE. Manufacturer Narrative: UPDATED DATA: D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE, H4. DEVICE MFG DATE CORRECTED DATA: B5. SECOND PARAGRAPH E. FACILITY ADDRESS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819503&pc=NPT | 2025587-2025-02557 | 2024-09-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-10 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 11 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DEVELOPED SUDDEN CARDIOPULMONARY ARREST AND SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE IMPLANT PROCEDURE CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819612&pc=DRA | 9612164-2025-01873 | 2025-04-07 04:00 | Death | MEDTRONIC IRELAND | 2025-04-10 04:00 | DRA | FLEXCATH CONTOUR¿ STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K232321 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A CARDIAC ABLATION PROCEDURE, IT WAS DIFFICULT TO INSERT ANOTHER MANUFACTURER'S SHEATH INTO THE GROIN; THE TIP OF THE SHEATH WOULD BECOME JAMMED UPON INSERTION. THE SHEATH WAS REPLACED AND THE PROCEDURE CONTINUED AS EXPECTED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND A HEART CATHETER WAS USED, BUT NO FINDINGS WERE REVEALED. A COMPLICATION OCCURRED IN THE VASCULATURE IN THE GROIN AREA WHILE ATTEMPTING VASCULAR ACCESS. THIS LED TO A SIGNIFICANT VENOUS BLEED. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED, AND BLOOD TRANSFUSIONS WERE GIVEN, BUT THE BLEEDING WAS UNABLE TO BE STOPPED. THE PROCEDURE WAS ABORTED WHILE THE PATIENT WAS UNDER GENERAL ANESTHESIA. THE PATIENT WAS THEN ADMITTED TO THE INTENSIVE CARE UNIT AND LATER EXPIRED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEFT SUPERIOR PULMONARY VEIN (LSPV), LEFT INFERIOR PULMONARY VEIN (LIPV) AND RIGHT SUPERIOR PULMONARY VEIN (RSPV) WERE ABLATED AND THEN THE PATIENT'S BLOOD PRESSURE DROPPED. Manufacturer Narrative: UPDATED B5.DESC EVT PROBLEM UPDATED H6.PATIENT CODES (IME/ANNEX E). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 900310; PRODUCT TYPE: INTEGRATED DILATOR/NEEDLE PRODUCT ID NON-MEDTRONIC SL0; PRODUCT TYPE: SHEATH PRODUCT ID PSCC100; PRODUCT TYPE: LOOP CATHETER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819650&pc=MIH | 2017233-2025-06061 | 2025-03-18 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-10 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Insufficient Information | Stroke/CVA; Paralysis | P040043 |  | 1 | Event Description: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT A TYPE B AORTIC DISSECTION UTILIZING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL (CTAGAC). THERE WERE NO ISSUES DURING THE PROCEDURE, THE PATIENT TOLERATED THE PROCEDURE. ON 3/18/2024, FSA WAS NOTIFIED THAT THE PATIENT SUFFERED A STROKE ON (B)(6) 2025. PARALYSIS WAS ALSO OBSERVED, AND EVENTUALLY PATIENT PASSED AWAY. PHYSICIAN DOES NOT THINK ISSUES ARE RELATED TO THE PROCEDURE OR CTAGAC DEVICE, BUT CAUSE REMAINS UNKNOWN. NO FURTHER PATIENT INFORMATION IS AVAILABLE. Manufacturer Narrative: H6: CODE C19: A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE MANUFACTURING SPECIFICATIONS. AS THE DEVICE WAS NOT ACCESSIBLE, THE PRODUCT HISTORY REVIEW (PHR) REVIEW WAS THE EXTEND OF THE INVESTIGATION. NO DEVICE PROBLEM WAS FOUND PER REVIEW OF THESE RECORDS. H6: CODE C20: THE DEVICE REMAINS IMPLANTED AND, THEREFORE, WAS NOT AVAILABLE FOR DIRECT ANALYSIS BY GORE. IT SHOULD BE NOTED THAT, PER THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, DEATH AND NEUROLOGIC DAMAGE, LOCAL OR SYSTEMIC (E.G., STROKE, PARAPLEGIA, PARAPARESIS). W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819973&pc=QZI | 3002648230-2025-00302 | 2025-04-07 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-04-10 04:00 | QZI | PULSESELECT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | P230017 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A CARDIAC ABLATION PROCEDURE, IT WAS DIFFICULT TO INSERT ANOTHER MANUFACTURER'S SHEATH INTO THE GROIN; THE TIP OF THE SHEATH WOULD BECOME JAMMED UPON INSERTION. THE SHEATH WAS REPLACED AND THE PROCEDURE CONTINUED AS EXPECTED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND A HEART CATHETER WAS USED, BUT NO FINDINGS WERE REVEALED. A COMPLICATION OCCURRED IN THE VASCULATURE IN THE GROIN AREA WHILE ATTEMPTING VASCULAR ACCESS. THIS LED TO A SIGNIFICANT VENOUS BLEED. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED, AND BLOOD TRANSFUSIONS WERE GIVEN, BUT THE BLEEDING WAS UNABLE TO BE STOPPED. THE PROCEDURE WAS ABORTED WHILE THE PATIENT WAS UNDER GENERAL ANESTHESIA. THE PATIENT WAS THEN ADMITTED TO THE INTENSIVE CARE UNIT AND LATER EXPIRED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEFT SUPERIOR PULMONARY VEIN (LSPV), LEFT INFERIOR PULMONARY VEIN (LIPV) AND RIGHT SUPERIOR PULMONARY VEIN (RSPV) WERE ABLATED AND THEN THE PATIENT'S BLOOD PRESSURE DROPPED. Manufacturer Narrative: UPDATED B5.DESC EVT PROBLEM UPDATED H6.PATIENT CODES (IME/ANNEX E). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 900310; PRODUCT TYPE: INTEGRATED DILATOR/NEEDLE PRODUCT ID NON-MEDTRONIC SL0; PRODUCT TYPE: SHEATH PRODUCT ID 10FCC13; PRODUCT TYPE: SHEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21819979&pc=NPT | 2025587-2025-02565 | 2025-03-17 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-10 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) USING AN 18 MILLIMETER (MM) NON-MEDTRONIC BALLOON WAS PERFORMED DUE TO CALCIFICATION. FOLLOWING THE IMPLANT OF THE VALVE, A TEAR IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WAS OBSERVED. SUBSEQUENTLY, THE PATIENT'S BLOOD PRESSURE DROPPED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. ADDITIONALLY, A CORONARY OCCLUSION WAS OBSERVED AND THE PROCEDURE WAS SUBSEQUENTLY CONVERTED TO AN EMERGENCY CORONARY ARTERY BYPASS GRAFT (CABG); HOWEVER, THE PATIENT DIED LATER THAT DAY. IT WAS REPORTED THAT PRE-IMPLANT BAV MAY HAVE CONTRIBUTED TO THE TORN LAD, BUT THE TEAR WAS NOT DETECTED PRIOR TO THE IMPLANT OF THE VALVE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM DID NOT CAUSE OR CONTRIBUTE TO THE TORN LAD. THE GUIDEWIRE USED WAS A NON-MEDTRONIC PRODUCT. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME DELIV SYS D-EVOLUTFX-2329 PRODUCT ID D-EVOLUTFX-2329 SERIAL/LOT UNKNOWN USE BY DATE UNKNOWN UDI UNKNOWN. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21820754&pc=GDW | 2955842-2025-14012 | 2025-03-13 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-10 04:00 | GDW | SUREFORM | Insufficient Information | Hemorrhage/Bleeding; Perforation of Vessels | K190999 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED PULMONARY LOBECTOMY, BLEEDING OCCURRED AFTER STAPLING. THE PROCEDURE WAS CONVERTED TO OPEN WHERE FURTHER COMPLICATIONS OCCURRED AND THE PATIENT EXPIRED. THE SURGEON STATED THAT AFTER FIRING A GRAY SUREFORM 45 RELOAD ON THE TRUNCUS ANTERIOR ARTERY, BLEEDING OCCURRED IMMEDIATELY AT THE STAPLE LINE, REPORTEDLY DUE TO WEAKNESS IN THE TISSUE STRUCTURE FROM SURROUNDING LYMPH NODE CALCIFICATION. COMPRESSION WAS APPLIED, A CONTROLLED EMERGENCY UNDOCKING OF THE DA VINCI SYSTEM WAS COMPLETED, AND THE PROCEDURE WAS CONVERTED TO OPEN THORACOTOMY AFTER 50ML OF BLOOD LOSS. THE STAPLE LINE WAS REPORTED TO BE COMPLETELY APPROXIMATED AND WELL-FORMED STAPLES WERE IN PLACE. ONCE OPEN, NEW BLEEDING BEGAN AFTER A TEAR OF THE INTERLOBAR ARTERY OCCURRED WHILE BYPASSING THE UPPER LOBE BRONCHUS DUE TO THE TISSUE WEAKNESS. EFFORTS MADE TO ACHIEVE HEMOSTASIS INCLUDED COMPRESSION AND ATTEMPTS TO LOOP THE PULMONARY ARTERY THAT WERE UNSUCCESSFUL. THE ESTIMATED BLOOD LOSS DURING THE OPEN PORTION OF THE PROCEDURE WAS 4 LITERS. THE THREE THORACIC SURGEONS PRESENT FOR THE PROCEDURE EXCLUDED ANY INVOLVEMENT OF THE DA VINCI SYSTEM, OR INSTRUMENTS IN THE PATIENT'S OUTCOME AS THE LIFE-THREATENING EVENTS AROSE DURING THE OPEN PORTION OF THE PROCEDURE. Manufacturer Narrative: A REVIEW OF THE DA VINCI SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND NO ERRORS OCCURRED THAT WOULD BE RELATED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES ALSO FOUND NO SYSTEM ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. NO COMPLAINTS WERE RECEIVED FOR THE DA VINCI SYSTEM. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, PROGRASP FORCEPS, CADIERE FORCEPS, TIP-UP FENESTRATED GRASPER, LARGE NEEDLE DRIVER, SYNCHROSEAL, SUREFORM 45 CURVED TIP STAPLER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE REPORTED FOR ANY OF THE PRODUCTS. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE PROCEDURE CONFIRMED THERE WERE NO RELATED NON-CONFORMANCES. AN ADVANCED STAPLER LOG REVIEW FOUND THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 3 TIMES AND FIRED 3 GRAY RELOADS. ON EACH INSTALL, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH NO PAUSES FOR COMPRESSION. THE INSTRUMENT WAS REMOVED AND NOT USED AGAIN IN THE PROCEDURE. THE EMERGENCY STOP BUTTON APPEARS TO HAVE BEEN PRESSED ON SURGEON SIDE CONSOLE 2 (SSC) APPROXIMATELY 20 MINUTES AFTER THE FINAL COMPLETED UNCLAMPING. THERE WERE NO STAPLER RELATED ERRORS IN THE SYSTEM LOGS. A VIDEO CLIP FROM THE PROCEDURE WAS PROVIDED FOR EVALUATION. REVIEW FOUND THAT THE STAPLER FIRE APPEARS TO BE NORMAL, BUT FOLLOWING UNCLAMPING THERE IS BLEEDING FROM THE STAPLED ARTERY. NO MALFUNCTION OR UNEXPECTED BEHAVIOR FROM THE SYSTEM OR INSTRUMENTS IN VIEW WERE OBSERVED. BEFORE THE VIDEO ENDS, THE USER PLACES GAUZE ON THE BLEEDING AREA AS AN ATTEMPT FOR HEMOSTASIS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT BASED ON THE AVAILABLE INFORMATION, THE PATIENT IN THIS REPORT DIED FOLLOWING AN ATTEMPTED PULMONARY LOBECTOMY. AN INADEQUATE SEAL FROM A STAPLER ON A PULMONARY VESSEL CAUSED THE SURGEONS TO CONVERT TO AN OPEN OPERATION. ALTHOUGH THE ORIGINAL BLEEDING WAS REPORTED AS A SMALL VOLUME BLEED AND CONTROL WAS ESTABLISHED, A SECOND BLEEDING EVENT OCCURRED LATER IN THE PROCEDURE WITH A DIFFERENT VESSEL DURING THE OPEN PORTION OF THE OPERATION. THIS SECOND BLEED EVENTUALLY LED TO THIS PATIENT¿S DEATH. WHILE THE SURGEON SUGGESTED THE PATIENT¿S VESSEL ANATOMY WAS A POSSIBLE CONTRIBUTING FACTOR IN BOTH BLEEDING EVENTS, THE DETAILS OF HOW THEY GOT INTO THE SECOND BLEEDING AND EXACT CAUSE IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21820781&pc=QZI | 2124215-2025-22796 | 2025-03-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-10 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Respiratory Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT AT THE END OF A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. DURING THE PROCEDURE IT WAS IDENTIFIED THE PATIENT'S HEART WAS ENLARGED AND DILATED. "DISEASED SUBSTRATE" WAS FOUND DURING VOLTAGE MAPPING. ISOLATION OF THE PULMONARY VEINS WERE SUCCESSFULLY COMPLETED WITH THE FARAWAVE CATHETER. WHEN THE PATIENT WAS EXTUBATED, THEY "DESATURATED INSTANTLY" AND STAFF COULD NOT FIND A PULSE. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR APPROXIMATELY 45 MINUTES BEFORE THE PATIENT WAS DECLARED DECEASED. THE CAUSE OF DEATH WAS STATED TO BE RESPIRATORY ARREST WITH CARDIAC ARREST ALSO OCCURRING. IN THE PHYSICIAN'S OPINION NEITHER THE FARAWAVE NOR THE TREATMENT PROVIDED BY THE CATHETER CONTRIBUTED TO THE DEATH, BUT WITHOUT FURTHER INFORMATION THE DEVICE OR TREATMENT CANNOT BE FULLY RULED OUT. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. BSC WAS INFORMED THAT MORE SPECIFIC PRODUCT INFORMATION IS NOT AVAILABLE FOR THIS EVENT. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21821737&pc=LYJ | 1644487-2025-00392 | 2025-03-10 04:00 | Death | CYBERONICS - HOUSTON | 2025-04-10 04:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A SEIZURE. REPONSE WAS RECEIVED FROM THE PHYSICIAN THAT THEY ARE UNABLE TO PROVIDE FURTHER INFORMATION ON THE EVENT. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21822661&pc=NHL | 3006630150-2025-02381 | 2025-01-01 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-04-10 04:00 | NHL | VERCISE GENUS? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT PASSED AWAY OF AN UNKNOWN REASON. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD FAITH EFFORTS.Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT PASSED AWAY OF AN UNKNOWN REASON. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD FAITH EFFORTS. Manufacturer Narrative: BLOCK B3: APPROXIMATED BASED ON THE DATE THE MANUFACTURER BECAME AWARE OF THE EVENT. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: BLOCK B3: APPROXIMATED BASED ON THE DATE THE MANUFACTURER BECAME AWARE OF THE EVENT. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21803409&pc=DTZ | 9611109-2025-90049 | 2025-03-10 04:00 | Death | LIVANOVA DEUTSCHLAND GMBH | 2025-04-09 04:00 | DTZ | HEATER-COOLER SYSTEM | Microbial Contamination of Device | Bacterial Infection |  |  | 1 | Event Description: LIVANOVA DEUTSCHLAND RECEIVED A REPORT RELATED TO A DUST INFECTION OF TRICHOSPORON AUSTROAMERICANUM IN AN OPERATING ROOM THROUGH A HEATER-COOLER SYSTEM 3T DEVICE. FOLLOWING A 3RD CASE OF TRICHOSPORON AUSTROAMERICANUM SURGICAL SITE INFECTION, THAT HAS ALREADY LED TO TWO DEATHS IN THE HOSPITAL SINCE (B)(6) 2024, CUSTOMER MOVED TO OPERATING BLOCK 6 ON (B)(6). THE CUSTOMER THEN OBSERVED DUST BUILD-UP IN THE 3T DEVICE GRIDS IN BLOCK 6, ASKING LIVANOVA TO CARRY OUT A BIOCLEANING OF THE GRIDS FOR THIS DEVICE AND FOR THE ENTIRE FLEET URGENTLY.Event Description: SEE INITIAL REPORT. Manufacturer Narrative: A.1.-A.5. PATIENT INFORMATION WAS NOT PROVIDED. G.5. THE HEATER-COOLER 16-02-80 IS NOT DISTRIBUTED IN THE USA, AND IT IS SIMILAR TO HEATER-COOLER 16-02-85, WHICH IS DISTRIBUTED IN THE USA (510(K) NUMBER: K220635). H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE HEATER-COOLER SYSTEM 3T. THE 3T HEATER COOLER UNIT IS SIMPLY INVOLVED IN THE WHOLE PROCESS. THE USERS DISINFECT THE TWO BLACK GRILLES BY REGULAR CLEANING AS PER IFU REQUIREMENT. SAMPLES WERE TAKEN FROM THE GRID OF THE THERMAL CEC GENERATOR N° 16S30285: THE RESULT WAS NEGATIVE FOR TRICHOSPORON BUT OTHER SPECIES OF FUNGUS WERE FOUND. THE HC3TS ARE NORMALLY CLEANED EVERY WEEK. ACCORDING TO CUSTOMER, 3T HEATER COOLER UNIT, AS WELL AS EVERYTHING ELSE IN THE OPERATING ROOM, CONTRIBUTED TO THE CAUSE OF BOTH DEATHS. THE CUSTOMER ASKS TO STATE MORE CLEARLY THE MAINTENANCE TO REMOVE THE DUST AND A BIOCLEANING ACTION TO FIGHT AGAINST THE DUST (KNOWING THAT IT IS DUST FROM THE OPERATING FIELD). FROM A PRELIMINARY CLINICAL INVESTIGATION WITH LIVANOVA MEDICAL TEAM IT IS POSSIBLE TO STATE THAT THE FUNGUS MENTIONED BY THE CUSTOMER IN THE EVENT DOES NOT GROW IN DUST ACCUMULATIONS. ITS GROWTH ACTUALLY OCCURS INSTEAD IN A HUMID ENVIRONMENT. THEREFORE, THE INVESTIGATION IS ON-GOING TO VERIFY WHETHER THE ACCUMULATED DUST CAN BE A CONTRIBUTOR FACTOR OF THE EVENT OR NOT. IN ADDITION, FOLLOW-UP COMMUNICATION HAS BEEN STARTED WITH THE CUSTOMER TO ASK IF MICROBIAL TEST RESULT ON DUST HAVE BEEN CARRIED OUT AND IF RESULTS ARE AVAILABLE. IF ANY ADDITIONAL INFORMATION PERTINENT TO THE REPORTED EVENT IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: ON MARCH 10, 2025, THE CUSTOMER REPORTED TO LIVANOVA THREE CASES OF SURGICAL INFECTION WITH FUNGUS TRICHOSPORON AUSTROAMERICANUM. BASED ON INFORMATION PROVIDED, TWO DEATHS HAD ALREADY OCCURRED AT THE HOSPITAL SINCE (B)(6) 2024 (THESE EVENTS WERE NOT REPORTED TO LIVANOVA AT THE TIME), WHILE THE THIRD DEATH WAS REPORTED TO HAVE OCCURRED ON (B)(6) 2025. IN THE CONTEXT OF THE THIRD INFECTION, THE CUSTOMER ALLEGED ABOUT FUNGUS TRICHOSPORON AUSTROAMERICANUM GROWING IN DUST ENVIRONMENT AND THAT THIS MIGHT HAVE ORIGINATED FROM ANYTHING PRESENT IN THE OPERATING ROOM (ANY SURFACE AND/OR EQUIPMENT USED INCLUDING ALSO THE HEATER COOLER DEVICE). THROUGH FOLLOW UP COMMUNICATION WITH CUSTOMER LIVANOVA BECAME AWARE THAT THE HEATER COOLER 3T (B)(6) WAS USED FOR THE MENTIONED SURGERIES AND THAT IT IS NOT THE CAUSE OF THE THREE SURGICAL INFECTIONS. THIS WAS CONFIRMED BY THE FACT THAT SAMPLES OF DUST WERE TAKEN FROM THE GRID OF THE HEATER COOLER 3T (B)(6) AND ANALYZED. THE TEST RESULTED NEGATIVE TO FUNGUS TRICHOSPORON AUSTROAMERICANUM EXCLUDING THE HEATER COOLER DEVICE FROM BEING THE CAUSE OF THE SURGICAL INFECTIONS. CURRENTLY LIVANOVA IS NOT AWARE OF ANY OTHER ACTION UNDERTAKEN BY THE CUSTOMER ON THE DEVICE. IN CONCLUSION, BASED ON INVESTIGATION RESULTS, ANY RELATIONSHIP BETWEEN THE REPORTED INFECTION AND THE HEATER COOLER DEVICE IS EXCLUDED. NO SPECIFIC ACTION WAS CURRENTLY DEEMED NECESSARY, LIVANOVA MAINTAINS AND DOCUMENTS PERIODIC CUSTOMER EVENTS MONITORING PROCESS IN ORDER TO EVALUATE ACTIONS FOR PRODUCTS IMPROVEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21803577&pc=QSY | 3011137372-2025-00105 | 2024-09-01 04:00 | Death | Z-MEDICA, LLC | 2025-04-09 04:00 | QSY | QC HEMOSTATIC DRESSING, 2X2, 10X | Improper or Incorrect Procedure or Method | Insufficient Information | K123387 |  | 1 | Event Description: IT WAS REPORTED THAT: "THE PATIENT UNDERWENT SURGERY IN (B)(6) 2024 TO REMOVE A HEMOSTATIC DRESSING THAT HAD BEEN LEFT IN THE BODY DURING ANOTHER SURGERY ON (B)(6) 2022. THE HEMOSTATIC DRESSING HAD CAUSED AN ABSCESS TO FORM AND LED TO A SEVERE INFECTION, AND UNFORTUNATELY, THE PATIENT DIED." ADDITIONAL INFORMATION HAS BEEN REQUESTED AND A FOLLOW UP REPORT WILL BE SUBMITTED IF IT IS RECEIVED. Manufacturer Narrative: (B)(4) NO ITEM RETURNED FOR EVALUATION. IFU WARNINGS SECTION CLEARLY STATES:1) PRODUCT IS NOT ABSORBABLE AND MUST BE REMOVED FROM WOUND PRIOR TO WOUND CLOSURE.2) THE QUIKCLOT HEMOSTATIC DRESSING MUST NOT REMAIN IN THE WOUND FOR LONGER THAN 24 HOURS. THIS COMPLAINT IS THE RESULT OF USER ERROR, NOT ADHEARING TO IFU INSTRUCTIONS AND WARNINGS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21803651&pc=DTZ | 2184009-2025-00502 | 2025-03-31 04:00 | Death | PERFUSION SYSTEMS | 2025-04-09 04:00 | DTZ | FUSION HOLLOW FIBER OXYGENATOR | Gas Output Problem; Obstruction of Flow | Low Blood Pressure/ Hypotension; Ventilator Dependent; Low Oxygen Saturation | K230640 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF A FUSION OXYGENATOR, IT WAS REPORTED THAT THE PARTIAL PRESSURE OF OXYGEN WAS RECORDED AT 8 WITH 100% FRACTION OF INSPIRED OXYGEN (FI02) BEING DELIVERED THROUGH THE OXYGENATOR. THE CUSTOMER ALSO OBSERVED CLOTS IN THE OXYGENATOR. SEE ATTACHED PHOTOS. THE OXYGENATOR SWAPPED OUT WAS REPLACED WITH ANOTHER OXYGENATOR WHICH WORKED SUFFICIENTLY FOR THE REMAINING BYPASS TIME. THE PATIENT PASSED AWAY. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT WAS AT HIGH RISK WITH AN AORTIC DISSECTION. THE PATIENT HAD BEEN IN THE HOSPITAL FOR A FEW DAYS. THE PROCEDURE WAS NOT AN EMERGENCY/CRASH BYPASS. THE PATIENT DID RECEIVE TRASYLOL PRIOR TO THE CASE (9 ANTITHROMBOTIC). THE CUSTOMER MANAGED HIGH AND ACCORDINGLY ACTIVATED CLOTTING TIMES (ACT). THEIR HEMOCHRON ACT DEVICE WAS SUFFICIENT FOR TRASYLOL USE. THE OXYGENATOR REPLACEMENT TOOK 3-4 MINUTES BUT THE CHANGE OUT WAS REQUIRED AT A TIME WHEN THE PATIENT WAS STRUGGLING ON THE TABLE AND VENTILATION HAD TO COMMENCE DURING THE OXYGENATOR SWAP OUT PERIOD. HYPOTENSION WAS EXPERIENCED. THE FIRST BYPASS PERIOD WAS FOR APPROXIMATELY 173 MINUTES. IT WAS UNEVENTFUL FROM AN OXYGENATOR PERFORMANCE PERSPECTIVE. THE PATIENT'S BLOOD GASES, LACTATES, PARTIAL PRESSURE OF OXYGEN (PO2), PARTIAL PRESSURE OF CARBON DIOXIDE (PCO2) AND ACTS WERE ALL SUFFICIENT. AT 15.39PM, PO2 WAS 27MMHG AND ACT WAS 999S. THE FIRST BYPASS ENDED AT 15.44PM. PROTAMINE WAS ADMINISTERED TO THE PATIENT, HOWEVER, ALL SUCTION DEVICES WERE REMOVED FROM THE PATIENT. THERE WAS NO SUCTION BLOOD. THE ARTERIAL AND VENOUS CANNULA REMAINED IN SITU. THE PATIENT WAS TRAN SFUSION WITH VARIOUS DOSES OF 100ML FROM THE PERFUSION CIRCUIT DURING THIS OFF BYPASS STAGE, AVOIDING STASIS IN THE CIRCUIT. THESE 100ML TOP UPS OF PATIENT VOLUME OCCURRED APPROXIMATELY EVERY 30 SECONDS. UPON INSTIGATING A SECOND BYPASS AT 16.12PM, THE CUSTOMER NOTED THAT THE OXYGENATOR PERFORMANCE STARTED TO DETERIORATE. PO2 HAD DROPPED AND INCREASING FIO2 DID NOT RESOLVE THE ISSUE. THE FIO2 VALUES WERE FROM 60 TO 70, TO 80, TO 90, TO 100% WITH PO2¿S FALLING FROM 22.4, 15, 7.9, 7.6MMHG. THE CUSTOMER REPORTED OBSERVING CLOTTING IN THE OXYGENATOR, ALONG WITH PO2 CHALLENGES AND MAX FIO2. THE OXYGENATOR WAS SWAPPED OUT FOR ANOTHER FUSION OXYGENATOR AT APPROXIMATELY AT 17.20PM. SHORTLY AFTER THE OXYGENATOR WAS CHANGED OUT, THE OPERATION WAS STOPPED AS THE PATIENT PASSED AWAY. THE CUSTOMER NOTED THAT THE NEW OXYGENATOR WAS WORKING SUFFICIENTLY. THE CUSTOMER DOES NOT UNDERSTAND WHY THE OXYGENATOR WOULD STRUGGLE ON THE SECOND BYPASS PERIOD AS COMING OFF AND ON BYPASS WAS A COMMON OCCURRENCE. THE CUSTOMER STATED THAT THE OXYGENATOR FAILED AT A T IME IN THE OPERATION WHERE THE PATIENT WAS STRUGGLING DUE TO OPERATION RELATED ISSUES. THE PATIENT'S DEATH MAY BE MULTIFACTORIAL. PRE AND POST OXYGENATOR PRESSURES WERE NOT MEASURED. THE ACT¿S WERE KEPT HIGH THROUGHOUT THE CASE AND THEY WERE ALL ABOVE PROTOCOL. THE CUSTOMER NOTICED CLOTS IN THE OXYGENATOR ON THEIR SECOND BYPASS RUN ONLY. A NORMAL ROLLER PUMP CONFIGURATION WAS USED. Manufacturer Narrative: DEVICE EVALUATION SUMMARY: VISUAL INSPECTION SHOWS NO OUTWARD SIGNS OF PHYSICAL DAMAGE OR ABNORMALITIES. UNIT APPEARS TO HAVE BEEN USED. PRESSURE INTEGRITY TESTING SHOWS NO INTERNAL OR EXTERNAL LEAKS WHEN RUN AT 3 LPM WITH 23 PSI, (1189 MMHG) OF BACK PRESSURE FOR 10 MINUTES. THE DEVICE WAS SENT TO THE BLOOD LAB FOR PERFORMANCE TESTING. REASON FOR RETURN WAS UNDETERMINED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CUSTOMER STATED THAT MEDTRONIC ARE NOT PRIVILEGED TO THE FINAL INFORMATION REGARDING THE CAUSE OF DEATH, HOWEVER, THEY MENTIONED A RIGHT-SIDE HEART FAILURE BUT THERE WAS NO FORMAL NOTE. THE PRESSURE WAS MEASURED IN THE CIRCUIT POST OXYGENATOR. HEPARIN DOSING WAS MONITORED WITH AN ACT INSTRUMENT TO ACHIEVE OPTIMAL THERAPEUTIC RESPONSE. THE CARDIOTOMY FLOW RATES WERE STANDARD FOR THE SHUNT LINES (E.G. 500ML/MINUTE). THE CUSTOMER USED A FLOW PROBE TO MEASURE THE ACTUAL FLOW TO THE PATIENT. THE SUCKER BLOOD WAS NOT SEQUESTERED TO THE CELL SAVER BEFORE RETURNING TO CARDIOTOMY VENOUS RESERVOIR (CVR). THE LEVEL IN THE CVR WAS NOT UNUSUALLY LOW OR HIGH FOR CVR LEVELS. THERE WAS 500ML OF RINGER¿S SOLUTION, 1000ML OF HARTMANN'S SOLUTION AND 10,000 UNITS OF HEPARIN ADMINISTRATED FOR PRIMING. DURING CARDIOPULMONARY BYPASS (CPB), GELOFUSINE, RINGER'S SOLUTION, BLOOD, CELL SAVER BLOOD, APROTININ (TRASYLOL), MANNITOL (10%), METARAMINOL AND SODIUM BICARBONATE WERE ADMINISTRATED. POST OXYGENATOR PRESSURES WERE 193MMHG, 158MMHG AND 73MMHG. THE CARDIOPLEGIA TEMPERATURE WAS BETWEEN 6.1° TO 7°. THE PATIENT'S CORE TEMPERATURE WAS BETWEEN 36.5° TO 36.8°. THE ARTERIAL TEMPERATURE WAS BETWEEN 22.4° TO 32.7°. AT 17.32PM, IT WAS NOTED THAT THE PULMONARY ARTERY (PA) PRESSURE WAS TOO HIGH. THE PATIENT WAS GIVEN PROSTACYCLIN BEFORE COMING OFF BYPASS. THE PATIENT PASSED AWAY AT 17 .38PM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804493&pc=OZD | 1220648-2025-27313 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-09 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED LIMB ISCHEMIA, ATRIAL FIBRILLATION AND SUBSEQUENTLY PASSED AWAY. LEFT HEART CATHETERIZATION SHOWED CHRONIC TOTAL OCCLUSION TO THE RIGHT CORONARY ARTERY (RCA) GRAFT. AN INTRA-AORTIC BALLOON PUMP (IABP) WAS PLACED AT THE CARDIOTHORACIC SURGEONS REQUEST. THE PATIENT CONTINUED TO DECLINE REQUIRING INTUBATION, LEVOPHED/VASOPRESSIN/AMIODARONE TO BE ADDED. CARDIOTHORACIC SURGERY TEAM TURNED THE PATIENT DOWN FOR SURGERY DUE TO THE PATIENT'S INSTABILITY. THE TEAM ELECTED TO EXCHANGE THE IABP FOR AN IMPELLA CP TO OPTIMIZE. THE PATIENT REQUIRED TWO ROUNDS OF CARDIOPULMONARY RESUSCITATION BETWEEN IABP DISCONTINUATION AND THE IMPELLA CP PLACEMENT. TEMPORARY PACER WAS PLACED. THE PATIENT WAS SENT TO THE UNIT ON SUPPORT. THE SAME DAY, IT WAS REPORTED THAT THE PATIENT HAD NO PULSES LOWER EXTREMITY BILATERALLY. TWO DAYS LATER, VASOPRESSIN WAS REMOVED. NO ESCALATION TO AN IMPELLA 5.5 WAS DETERMINED. HOWEVER, CORONARY ARTERY BYPASS GRAPH (RCA) REDO IS NEEDED WHEN THE PATIENT IS MORE STABLE. LATER IN THE EVENING, IT WAS REPORTED THE NURSE WAS UNABLE TO OBTAIN PULSES. THE PATIENT EXPERIENCED PAROXYSMAL ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. THE FOLLOWING MORNING, THE PATIENT WAS MADE DO NOT RESUSCITATE AND THE PATIENT WOULD EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804603&pc=OEV | 9610691-2025-00007 | 2025-03-07 05:00 | Death | AMBU A/S | 2025-04-09 04:00 | OEV | SPUR II PEDIATRIC RESUSCITATOR | Component Missing | Low Oxygen Saturation |  |  | 1 | Event Description: AS PER THE CUSTOMER, DURING INTUBATION IN AN EMERGENCY ROOM, PEDIATRIC PATIENT WAS NOT GETTING CO2. PEDIATRIC BAG WAS USED ON A PATIENT BUT THE INLET VALVE AT THE BASE WAS MISSING AND THEY COULD NOT PRODUCE ANY CO2. LATER WHEN INSPECTING THE BAG, THEY NOTICED THAT THE INLET WAS FOUND AND INSIDE BY THE MASK. Manufacturer Narrative: WE VERIFIED THE REPORTED FAILURE OF A MISSING PATIENT VALVE DISC INSIDE PATIENT VALVE HOUSING FROM A PHOTO PROVIDED BY THE CUSTOMER. THE SAMPLE WAS NOT RETURNED FOR INVESTIGATION. DURING MANUFACTURING, PATIENT VALVE DISC IS INSTALLED INTO PATIENT VALVE HOUSING AND A FUNCTION TEST PERFORMED ON 100% OF ASSEMBLED PATIENT VALVES. THEN THE VALVE IS INSTALLED ONTO SPUR II RESUSCITATOR AND ANOTHER FUNCTION TEST IS PERFORMED ON ALL DEVICES. MISSING PATIENT VALVE DISC WOULD BE REVEALED BY BOTH FUNCTION TESTS; HOWEVER PRODUCTION RECORDS DO NOT INDICATE ANY ABNORMALITY. THEREFORE, THE DEVICE SHOULD BE WAS WITHIN SPECIFICATIONS. NEXT, PEEP VALVE, FACE MASK AND MANOMETER WERE PRE-ATTACHED, WHICH WOULD NOT IMPACT INSTALLED PATIENT VALVE DISC. NO ABNORMALITIES WERE FOUND IN THE PRODUCTION RECORDS; THUS, IT IS CONCLUDED THAT DURING MANUFACTURING, THE DISC WAS INSTALLED CORRECTLY. THE CUSTOMER REPORTED THAT THE PATIENT VALVE DISC WAS FOUND INSIDE THE BAG NEXT TO A MASK. POSSIBLE CAUSES FOR THIS ISSUE TO OCCUR ARE: DETACHMENT OF THE EXPIRATORY CONNECTOR AND PATIENT VALVE DISC DURING TRANSPORTATION; OR DISASSEMBLING OF THE EXPIRATORY CONNECTOR AND PATIENT VALVE DISC BY THE USER, WITH SUBSEQUENT WRONG ASSEMBLY. AS PER THE TRANSPORTATION TESTING, THE EXPIRATORY CONNECTOR AND PATIENT VALVE DISC WOULD NOT DETACH UNDER NORMAL TRANSPORTATION CONDITIONS. SEVERAL ATTEMPTS WERE MADE TO RETRIEVE MORE INFORMATION FORM THE CUSTOMER; HOWEVER, THEY WERE UNSUCCESSFUL. DUE TO LIMITED INFORMATION, THE ROOT CAUSE REMAINS UNKNOWN. THE REPORTED FAILURE IS DETECTABLE DURING A PRESCRIBED PRE-CHECK. AS PER THE INSTRUCTIONS FOR USE: 'ALWAYS INSPECT THE RESUSCITATOR AND PERFORM FUNCTIONAL TEST AFTER UNPACKING, CLEANING, ASSEMBLY AND PRIOR TO USE.' THIS ISSUE IS CONSIDERED TO BE AN ISOLATED CASE, SINCE THIS WAS THE FIRST TIME THAT A DETACHMENT OF PATIENT VALVE DISC WAS REPORTED IN THE PAST YEAR. PRODUCTION HAS BEEN INFORMED ABOUT THE FAILURE AND WE WILL KEEP MONITORING THE MARKET FOR SIMILAR INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804668&pc=DSQ | 2916596-2025-02164 | 2025-02-01 05:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fatigue; Muscle Weakness; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF PSEUDOMONAS DRIVELINE INFECTIONS AFTER DRIVELINE RELOCATION WITH OMENTAL FLAP IN (B)(6) 2025. THE PATIENT RECEIVED A 6-WEEK COURSE OF ANTIBIOTICS AND HAD THEIR PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) LINE REMOVED AFTER BEING SEEN AT THE CLINIC. THE PATIENT WAS SEEN IN THE CLINIC ON (B)(6) 2025 IN A NORMAL STATE. LATER, IN THE EARLY HOURS OF(B)(6) 2025, THE PATIENT PASSED AWAY UNEXPECTEDLY AT HOME AND NO OTHER DETAILS LEADING UP TO THE DEATH WERE NOTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED ON 23APR2025 THAT THE PATIENT WAS FOUND DECEASED IN BED AROUND 03:30 AFTER THE PATIENT'S WIFE WAS AWOKEN BY AN ALARM ON THE SYSTEM CONTROLLER. THE PATIENT'S WIFE STATED THAT THE PATIENT WAS NOT BREATHING AT THE TIME AND IMMEDIATELY CALLED 911. THE PATIENT'S WIFE DID A SHORT ROUND OF CHEST COMPRESSIONS, BUT THE PATIENT HAD ALREADY PASSED AWAY. THE SYSTEM CONTROLLER WAS DISCONNECTED BY THE FAMILY AND RETURNED TO THE CUSTOMER FOR INVESTIGATION. OF NOTE, THE PATIENT HAS NOT BEEN FEELING WELL FOR 2 DAYS LEADING UP TO THE DEATH. VAGUE COMPLAINTS OF NOT FEELING WELL AND WEAK/FATIGUE. LOG FILES SUBMITTED FOR REVIEW REVEALED CONSTANT LOW FLOW EVENTS STARTING ON (B)(6) 2025 AT 0214 THAT HAD CONTINUED UNTIL THE SYSTEM CONTROLLER WAS DISCONNECTED AT 0420. ESTIMATE FLOWS WERE IN THE 4 LITERS PER MINUTE (LPM) RANGE PRIOR TO THE LOW FLOW EVENTS AND THEN DROPPED IN THE 1 LPM RANGE AT 0200, DURING THE EVENTS, THE PULSATILITY INDEX (PI) VALUES WERE ON THE LOW END BETWEEN 1 AND 4. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED THE LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE DETERMINED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED SYSTEM CONTROLLER EVENT LOG FILES CAPTURED RELEVANT EVENTS FROM 21MAR2025 THROUGH 23MAR2025. A PERSISTENT LOW FLOW HAZARD ALARM WAS CAPTURED ON 23MAR2025 AND REMAINED ACTIVE UNTIL THE DRIVELINE WAS DISCONNECTED THAT SAME DAY. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINE SYSTEM CONTROLLER ALARMS, INCLUDING THE LOW FLOW HAZARD ALARM, AND PROVIDE INFORMATION REGARDING HOW TO RESPOND TO AND TROUBLESHOOT THE ALARMS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B3: EVENT DATE IS ESTIMATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804793&pc=NIK | 9614453-2025-01394 | 2025-03-17 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-09 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2021, 419488 LEAD IMPLANTED (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804794&pc=DTB | 2649622-2025-09512 | 2025-03-17 04:00 | Death | MPRI | 2025-04-09 04:00 | DTB | PIROUET+ | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2021, 419488 LEAD IMPLANTED (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804800&pc=DTB | 2649622-2025-09513 | 2025-03-17 04:00 | Death | MPRI | 2025-04-09 04:00 | DTB | CRYSTALLINE ACTFIX | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6)2021, 419488 LEAD IMPLANTED (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804806&pc=DTB | 2649622-2025-09514 | 2025-03-17 04:00 | Death | MPRI | 2025-04-09 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: CONTINUATION OF D10: 419488 LEAD IMPLANTED (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804808&pc=LWS | 2649622-2025-09515 | 2025-03-17 04:00 | Death | MPRI | 2025-04-09 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2021; 419488 LEAD IMPLANTED (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804811&pc=FNL | 1525712-2025-00003 | 2024-03-27 04:00 | Death | UNKNOWN | 2025-04-09 04:00 | FNL | RMA-BEDS | Insufficient Information | Physical Entrapment |  |  | 1 | Event Description: HEALTH CANADA INCIDENT # (B)(4): IT WAS REPORTED TO HEALTH CANADA THAT ON (B)(6) 2024, AN INCIDENT HAPPENED AT A LONG-TERM FACILITY WHICH LED TO A PATIENT'S DEATH. THE PATIENT, WHO WAS NOT INDEPENDENTLY MOBILE, GOT HER LOWER BODY OFF THE EDGE OF THE BED, AND WAS FOUND WITH THE LOWER PART OF HER BODY ON THE FLOOR AND HER HEAD CAUGHT BETWEEN THE BED AND THE ASSIST RAIL. THE ASSIST RAIL WAS IN THE UPPER POSITION AND OUT OF USE AT THE TIME OF THE EVENT. Manufacturer Narrative: A FACILITY, (B)(6) CANADA, REPORTED DIRECTLY TO HEALTH CANADA ON AN ENTRAPMENT INCIDENT THAT LED TO A PATIENT'S DEATH. LIMITED INFORMATION WAS PROVIDED. THE EVENT INVOLVES THE IHCSRLAS, ASSIST RAIL FOR CS BED, MANUFACTURED BY FLEXTRONICS (MEXICO) ON APRIL 30, 2015. THE MODEL AND SERIAL NUMBER OF THE ACTUAL BED THE RAILS WERE ATTACHED TO IS UNKNOWN AS WELL AS ANY INFORMATION RELATING TO THE MATTRESS BEING UTILIZED ON THE BED. NO INFORMATION WAS PROVIDED RELATING TO A SPECIFIC MALFUNCTION WITH THE ASSIST RAIL. THIS REPORT IS BEING FILED IN AN ABUNDANCE OF CAUTION. INVACARE CANADA IS ATTEMPTING TO OBTAIN ADDITIONAL INFORMATION FROM THE FACILITY. SHOULD ANY ADDITIONAL INFORMATION BECOME AVAILABLE, A FOLLOW UP REPORT WILL BE FILED. NOTE: THE MANUFACTURER OF THE SUBJECT BED IS UNKNOWN, THEREFORE, TAYLOR STREET MANUFACTURING ((B)(6)) CFN NUMBER IS BEING UTILIZED FOR THE MDR FILING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804814&pc=LWP | 2649622-2025-09516 | 2025-03-17 04:00 | Death | MPRI | 2025-04-09 04:00 | LWP | ATTAIN BIPOLAR OVER THE WIRE LEAD | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P010015 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2021. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21804817&pc=LWS | 2649622-2025-09517 | 2025-03-17 04:00 | Death | MPRI | 2025-04-09 04:00 | LWS | SPRINT FIDELIS | Adverse Event Without Identified Device or Use Problem | Pocket Erosion; Cardiac Perforation; Pericardial Effusion | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS POST GENERATOR AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT, THE PATIENT EXPERIENCED AN INFECTION AND EROSION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS EXPLANTED, AND A LASER LEAD EXTRACTION WAS PERFORMED ON 3 ACTIVE AND 3 CAPPED LEADS. A NEW RIGHT ATRIAL (RA) LEAD, RV LEAD, AND CRT-D WERE IMPLANTED. SHORTLY AFTER ANESTHESIOLOGIST NOTED A SIZABLE PERICARDIAL EFFUSION AND A PERICARDIAL DRAIN WAS PERFORMED. THE PATIENT¿S VITALS WERE STABILIZED. AFTER THE PROCEDURE, THE PATIENT CODED TWICE AND PASSED AWAY. IT WAS BELIEVED THE CAUSE OF DEATH HAD TO DO WITH PERFORATION AS A RESULT OF THE LASER LEAD EXTRACTION. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2021, 419488 LEAD IMPLANTED (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805010&pc=OZD | 1220648-2025-27319 | 2025-03-13 04:00 | Death | ABIOMED, INC. | 2025-04-09 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Insufficient Information | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND AFTER THE IMPLANT, THE PUMP SHIFTED TOWARDS SEPTAL WALL LIMITING ITS FLOW. KINK OCCURRED ON CATHETER THAT DID NOT ALLOW FORWARD MANEUVER. FINAL POSITION WAS AT 3.4 CM TOWARD MITRAL VALVE. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. ACCORDING TO THE CLINICAL NOTES, THE DOCTOR WAS NOT ABLE TO REPOSITION THE PUMP AFTER IT BECAME SHALLOW DURING THE ADVANCEMENT OF THE BLUE HUB. DOCTOR WAS ABLE TO REPOSITION THE PUMP. THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USE RELATED SINCE THE PUMP POSITION WAS COMPROMISED DURING ADVANCING THE BLUE CATH-ANCHOR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805101&pc=OZD | 1220648-2025-27320 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-09 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY PLACED ON BIPELLA SUPPORT (IMPELLA CP AND IMPELLA RP FLEX) FOR ONE DAY, THEN LEFT SIDED SUPPORT ONLY FOR THREE DAYS BEFORE BEING ESCALATED TO AN IMPELLA 5.5 DEVICE FOR CONTINUED MECHANICAL CIRCULATORY SUPPORT. APPROXIMATELY FOUR DAYS AFTER THE START OF THE IMPELLA 5.5 SUPPORT, THE PATIENT'S URINE OUTPUT WAS PINK-TINGED. THE FOLLOWING DAY, THE URINE WAS STILL PINK/RED TINGED WITH SMALL CLOTS. DIURETICS WERE DISCONTINUED. ALBUMIN WAS GIVEN. THE PATIENT CONTINUED ON SUPPORT. TWO DAYS LATER, THE PATIENT WAS REINTUBATED FOR AIRWAY CONCERNS AND FALLING OXYGEN SATURATIONS. CONTINUOUS VENO-VENOUS HEMOFILTRATION WAS STARTED. HOWEVER, THE PATIENT WOULD EXPIRE A COUPLE OF DAYS LATER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805178&pc=GBX | 1820334-2025-00365 | 2025-03-15 04:00 | Death | COOK INC | 2025-04-09 04:00 | GBX | FUHRMAN PLEURAL/PNEUMOPERICARDIAL DRAINAGE SET | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  |  | 1 | Event Description: IT WAS REPORTED THAT A 12FR FUHRMAN PLEURAL/PNEUMOPERICARDIAL DRAINAGE SET WAS REQUIRED TO DRAIN PLEURAL FLUID FROM A 2-YEAR-OLD PEDIATRIC PATIENT. PRIOR TO DEVICE PLACEMENT, THE PATIENT WAS IN CRITICAL CONDITION, INTUBATED, AND BEING TREATED FOR SEPSIS AND PNEUMONIA. DURING THE PROCEDURE, THE DEVICE WAS INSERTED INTO THE THIRD INTERCOSTAL SPACE, AND FLUID WAS SUCCESSFULLY ASPIRATED. THE PROCEDURE WENT SMOOTHLY WITH NO COMPLICATIONS, AND AN ADDITIONAL 150 ML OF FLUID WERE DRAINED AFTER THE CATHETER WAS PLACED. NO BLOOD IN THE FLUID, THE DRAIN, OR THE ENDOTRACHEAL TUBE WAS SEEN AT THIS TIME. A POST-PROCEDURE X-RAY WAS PERFORMED TO VERIFY CATHETER PLACEMENT, AND THE IMAGING CONFIRMED NORMAL PLACEMENT OF THE DEVICE. THREE HOURS AFTER THE DRAIN WAS PLACED, THE PATIENT BEGAN BLEEDING PROFUSELY FROM THEIR ENDOTRACHEAL TUBE AND FROM THE FUHRMAN DRAIN. THE PATIENT EXPERIENCED CARDIAC AND/OR RESPIRATORY ARREST AS RESULT AND LATER EXPIRED FROM THESE COMPLICATIONS. NO OTHER ADVERSE EFFECTS WERE REPORTED FOR THIS INCIDENT. ADDITIONAL INFORMATION REGARDING EVENT DETAILS HAS BEEN REQUESTED BUT IS NOT AVAILABLE AT THIS TIME.Event Description: ADDITIONAL INFORMATION WAS PROVIDED ON 10APR2025. THE COOK REPRESENTATIVE IS UNAWARE OF ANY ADDITIONAL PROCEDURES THAT WERE COMPLETED PRIOR TO THE PLACEMENT OF THE CHEST TUBE. THE FAMILY OF THE PATIENT DECLINED AN AUTOPSY. Manufacturer Narrative: D4 - CATALOG #: C-PPD-1200-CHILDRENS-082694-IMH. G4 ¿ PMA/510(K) #: EXEMPT. H3 - DEVICE EVALUATED BY MFG? DEVICE NOT RETURNED TO MANUFACTURER. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED OR THAT A DEATH OR SERIOUS INJURY OCCURRED; NOR IS IT ADMISSION THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805267&pc=QSK | 9612515-2025-00034 | 2025-04-01 04:00 | Death | VASCUTEK LTD | 2025-04-09 04:00 | QSK | THORAFLEX HYBRID | Insufficient Information; Appropriate Term/Code Not Available |  | P210006 |  | 1 | Event Description: EVENT TYPE: ACUTE DISSECTION. THE ISSUE WAS IDENTIFIED AS IMPROPER PLACEMENT OF THE DEVICE RATHER THAN A DEVICE FAILURE. SURGICAL INTERVENTION: ENDOVASCULAR REPAIR WAS RECOMMENDED AND COMPLETED. OUTCOME: PATIENT IS DOING WELL.Event Description: EVENT TYPE: ACUTE DISSECTION. THE ISSUE WAS IDENTIFIED AS IMPROPER PLACEMENT OF THE DEVICE RATHER THAN A DEVICE FAILURE. SURGICAL INTERVENTION: ENDOVASCULAR REPAIR WAS RECOMMENDED AND COMPLETED. OUTCOME: PATIENT IS DOING WELL. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP #1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00034. THE MANUFACTURER'S ARE AWAITING CT SCANS FROM THE COMPLAINANT. THE NEXT FOLLOW -UP REPORT WILL BE SENT BY 30 JUL 25.Event Description: EVENT TYPE: ACUTE DISSECTION. THE ISSUE WAS IDENTIFIED AS IMPROPER PLACEMENT OF THE DEVICE RATHER THAN A DEVICE FAILURE. SURGICAL INTERVENTION: ENDOVASCULAR REPAIR WAS RECOMMENDED AND COMPLETED. OUTCOME: PATIENT IS DOING WELL. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP #2 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00034. NO COMPUTED TOMOGRAPHY SCANS OR ADDITIONAL DETAILS WERE PROVIDED REGARDING THE EVENT. SHOULD ANY FURTHER INFORMATION BECOME AVAILABLE, A FOLLOW-UP WILL BE ISSUED. Manufacturer Narrative: CLINICAL CODE: 4870 - AORTIC DISSECTION- ACUTE DISSECTION IMPACT CODE: 4624- SURGICAL INTERVENTION- ENDOVASCULAR REPAIR WAS RECOMMENDED AND COMPLETED. 1802- DEATH- THE PATIENT STARTED TO IMPROVE BUT THEM QUICKLY WENT DOWNHILL AND EXPIRED WHILE IN RECOVERY. MEDICAL DEVICE PROBLEM: 3191- APPROPRIATE TERM/CODE NOT AVAILABLE- NO CAUSAL LINK BETWEEN THE EVENT AND DEVICE COULD BE ESTABLISHED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE TYPE OF INVESTIGATION: 4111 - COMMUNICATION/INTERVIEW: ADDITIONAL INFORMATION WAS RECEIVED ON 23 JUN 25 CONFIRMING THE LOCATION OF EVENT WAS TYPE B DISSECTION OF THE DESCENDING AORTA. THE DISSECTION ENTRY TEAR DID NOT CORRELATE WITH THE PROXIMAL OR DISTAL LANDING ZONE OF A PREVIOUS ENDOVASCULAR/ HYBRID DEVICE. THE SHEATH COULD BE RETRACTED. THE CORRECT DELIVERY SYSTEM DEPLOYMENT STEP WERE FOLLOWED. NO HIGH FORCE EXPERIENCED IN RETRACTING THE SHEATH. THE DELIVERY SYSTEM DID NOT MOVE WITHIN THE PATIENT DURING DEPLOYMENT. THE DELIVERY SYSTEM DID NOT MOVE DISTALLY OR PROXIMALLY. THE DELIVERY SYSTEM DID NOT BUCKLE WITHIN THE PATIENT DURING DEPLOYMENT. THE DELIVERY SYSTEM DID NOT BREAK OR BECOME DISSEMBLED DURING DEPLOYMENT. THE DEVICE DID NOT MOVE DISTALLY WITH THE SHEATH UPON DEPLOYMENT. THE DEPLOYMENT SITE WAS NOT A SMALL DIAMETER, STENOSED, CALCIFIED, HIGHLY ANGULATED OR OTHERWISE CHALLENGING FOR DEPLOYMENT. NO GUIDE WIRE WAS USED. IT IS NOT MANDATORY FOR A GUIDEWIRE TO BE USED HOWEVER IT IS RECOMMEND THE USE OF A GUIDEWIRE DURING DEPLOYMENT OF THE THORAFLEX HYBRID DEVICE AS PER THE INSTRUCTIONS FOR USE. POST THE FOLLOW UP CT SCANS, A THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) GRAFT WAS PLACED FROM THE TRUE LUMEN INTO THE LUMEN OF THE THORAFLEX HYBRID GRAFT. THE PATIENT STARTED TO IMPROVE BUT THEM QUICKLY WENT DOWNHILL AND EXPIRED WHILE IN RECOVERY. 3331- ANALYSIS OF PRODUCTION RECORDS: A FULL BATCH REVIEW WAS PERFORMED FOR (B)(4), AND NO ISSUES WERE IDENTIFIED DURING MANUFACTURING PROCESS FROM RAW MATERIALS TO FINISHED PRODUCTS. 4117- DEVICE NOT ACCESSIBLE FOR TESTING: DEVICE REMAINS IMPLANTED INVESTIGATION FINDING: 3221 NO FINDINGS AVAILABLE- NO CAUSAL LINK BETWEEN THE EVENT AND DEVICE COULD BE ESTABLISHED. INVESTIGATION CONCLUSION: 4315- - CAUSE NOT ESTABLISHED- A COMPLETE ROOT CAUSE COULD NOT BE COMPLETED DUE TO THE MINIMAL INFORMATION AND CT SCANS NOT BEING PROVIDED Manufacturer Narrative: CLINICAL CODE: 4870 - AORTIC DISSECTION- ACUTE DISSECTION. IMPACT CODE: 4624- SURGICAL INTERVENTION- ENDOVASCULAR REPAIR WAS RECOMMENDED AND COMPLETED. MEDICAL DEVICE PROBLEM: 3190- INSUFFICIENT INFORMATION- ADDITIONAL QUESTIONS SENT TO THE SITE. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4111 - COMMUNICATION/INTERVIEW: ADDITIONAL QUESTIONS SENT TO THE SITE. 3331- ANALYSIS OF PRODUCTION RECORDS: A REVIEW OF THE MANUFACTURING AND QUALITY CONTROL RECORDS FOR THIS BATCH CONFIRMED THAT THE PRODUCT WAS MANUFACTURED TO SPECIFICATION. 4117- DEVICE NOT ACCESSIBLE FOR TESTING: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDING: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE. Manufacturer Narrative: CLINICAL CODE: 4870: AORTIC DISSECTION: ACUTE DISSECTION. IMPACT CODE: 4624: SURGICAL INTERVENTION: ENDOVASCULAR REPAIR WAS RECOMMENDED AND COMPLETED. 1802: DEATH: THE PATIENT STARTED TO IMPROVE BUT THEM QUICKLY WENT DOWNHILL AND EXPIRED WHILE IN RECOVERY. MEDICAL DEVICE PROBLEM: 3190: INSUFFICIENT INFORMATION- ADDITIONAL QUESTIONS SENT TO THE SITE. 3009: POSITIONING PROBLEM- IMPROPER PLACEMENT OF THE DEVICE. COMPONENT CODE: 4755: PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4111: COMMUNICATION/INTERVIEW: ADDITIONAL QUESTIONS SENT TO THE SITE. 3331: ANALYSIS OF PRODUCTION RECORDS: A FULL BATCH REVIEW WAS PERFORMED FOR TAG0211, AND NO ISSUES WERE IDENTIFIED DURING MANUFACTURING PROCESS FROM RAW MATERIALS TO FINISHED PRODUCTS. 4117: DEVICE NOT ACCESSIBLE FOR TESTING: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDING: 3233: RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: 11: CONCLUSION NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805436&pc=NGV | 2124215-2025-21590 | 2025-04-04 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | NGV | WATCHMAN FLX? PRO | Activation Failure | Cardiac Arrest; Hemorrhage/Bleeding; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Asystole | P130013 |  | 1 | Event Description: IT WAS NOTED A PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, CARDIAC PERFORATION AND DEATH OCCURRED. DURING A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, PRE-PROCEDURAL IMAGING CONFIRMED THERE WAS NO PRE-EXISTING PE. AFTER THE FEMORAL VEIN ACCESS, THE TRANSSEPTAL PUNCTURE (TSP) WAS COMPLETED USING VERSACROSS CONNECT (VCC) TRANSSEPTAL ACCESS SYSTEM THROUGH THE WATCHMAN ACCESS SYSTEM (WAS) UNDER INTRACARDIAC ECHOCARDIOGRAM (ICE) GUIDANCE. THE WAS ADVANCED INTO THE LEFT ATRIUM AND A 31MM WATCHMAN FLX PRO DELIVERY SYSTEM AND CLOSURE DEVICE (WDS) WAS ADVANCED TO THE LAA. UPON DEPLOYMENT OF THE WDS, THE DEVICE REMAINED CINCHED IN THE MIDDLE, SO THE PHYSICIAN RECAPTURED AND REDEPLOYED THE WDS. AFTER INITIAL DEPLOYMENT AND SECONDS BEFORE THE FINAL DEPLOYMENT OF THE WDS, CARDIAC TAMPONADE, AND A PE WAS NOTED AT THE COUMADIN RIDGE. THE WDS WAS RELEASED AND THE CLOSURE DEVICE IMPLANTED. A PERICARDIOCENTESIS WAS PERFORMED AND AN UNKNOWN AMOUNT OF FLUID WAS REMOVED. CARDIAC SURGERY WAS PERFORMED, AND THE CARDIAC PERFORATION WAS CLOSED. THE PATIENT WAS IN ASYSTOLE AND EXPERIENCED CARDIAC ARREST. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY ONE (1) HOUR AFTER THE CLOSURE DEVICE WAS IMPLANTED. IN THE PHYSICIAN'S OPINION, IMPROPER LOCATION OF THE INITIAL DEPLOYMENT OF THE WDS CAUSED THE PE.Event Description: IT WAS NOTED A PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, CARDIAC PERFORATION AND DEATH OCCURRED. DURING A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, PRE-PROCEDURAL IMAGING CONFIRMED THERE WAS NO PRE-EXISTING PE. AFTER THE FEMORAL VEIN ACCESS, THE TRANSSEPTAL PUNCTURE (TSP) WAS COMPLETED USING VERSACROSS CONNECT (VCC) TRANSSEPTAL ACCESS SYSTEM THROUGH THE WATCHMAN ACCESS SYSTEM (WAS) UNDER INTRACARDIAC ECHOCARDIOGRAM (ICE) GUIDANCE. THE WAS ADVANCED INTO THE LEFT ATRIUM AND A 31MM WATCHMAN FLX PRO DELIVERY SYSTEM AND CLOSURE DEVICE (WDS) WAS ADVANCED TO THE LAA. UPON DEPLOYMENT OF THE WDS, THE DEVICE REMAINED CINCHED IN THE MIDDLE, SO THE PHYSICIAN RECAPTURED AND REDEPLOYED THE WDS. AFTER INITIAL DEPLOYMENT AND SECONDS BEFORE THE FINAL DEPLOYMENT OF THE WDS, CARDIAC TAMPONADE, AND A PE WAS NOTED AT THE COUMADIN RIDGE. THE WDS WAS RELEASED AND THE CLOSURE DEVICE IMPLANTED. A PERICARDIOCENTESIS WAS PERFORMED AND AN UNKNOWN AMOUNT OF FLUID WAS REMOVED. CARDIAC SURGERY WAS PERFORMED, AND THE CARDIAC PERFORATION WAS CLOSED. THE PATIENT WAS IN ASYSTOLE AND EXPERIENCED CARDIAC ARREST. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY ONE (1) HOUR AFTER THE CLOSURE DEVICE WAS IMPLANTED. IN THE PHYSICIAN'S OPINION, IMPROPER LOCATION OF THE INITIAL DEPLOYMENT OF THE WDS CAUSED THE PE. IT WAS FURTHER REPORTED THAT 200ML OF FLUID WAS REMOVED WITH THE PERICARDIOCENTESIS, AND THE PATIENT RECEIVED TWO (2) UNITES OF RED BLOOD CELLS AND ONE (1) UNIT OF PLASMA. THE WDS WAS INITIALLY DEPLOYED IN THE COUMADIN RIDGE. Manufacturer Narrative: B5: INFORMATION ADDED. B7: INFORMATION ADDED. H6 IMPACT CODE ADDED: BLOOD TRANSFUSION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805442&pc=NKE | 2124215-2025-21933 | 2025-03-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | NKE | VALITUDE? X4 CRT-P | Incorrect, Inadequate or Imprecise Result or Readings | No Clinical Signs, Symptoms or Conditions | P030005 |  | 1 | Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) WAS FOUND TO BE IN SAFETY MODE. THIS CRT-P WAS EXPLANTED AND REPLACED. THIS CRT-P CAN BE EXPECTED TO BE RETURNED. NO ADDITIONAL ADVERSE PATIENT EFFECTS.Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) WAS FOUND TO BE IN SAFETY MODE. THIS CRT-P WAS EXPLANTED AND REPLACED. WITHIN 24 HOURS OF THIS SURGICAL REPLACEMENT PROCEDURE, THE PATIENT PASSED AWAY. THIS CRT-P HAS BEEN RECEIVED FOR ANALYSIS AND IS PENDING THE INVESTIGATION CONCLUSION. ONCE THE ANALYSIS IS COMPLETED, THIS RECORD WILL BE UPDATED WITH RESULTS.Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) WAS FOUND TO BE IN SAFETY MODE. THIS CRT-P WAS EXPLANTED AND REPLACED. WITHIN 24 HOURS OF THIS SURGICAL REPLACEMENT PROCEDURE, THE PATIENT PASSED AWAY. THIS CRT-P HAS BEEN RECEIVED FOR ANALYSIS. Manufacturer Narrative: THE PRODUCT IS EXPECTED TO BE RETURNED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON RETURN AND COMPLETION OF ANALYSIS. Manufacturer Narrative: THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS B2 (OUTCOMES ATTRIB TO ADV EVENT), B5 (EVENT DESCRIPTION), H1 (TYPE OF REPORTABLE EVENT), AND H6 (IMPACT CODES). THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS B5 (EVENT DESCRIPTION), D9 (DEVICE AVAIL FOR EVALUATION AND RETURNED TO MANUFACTURER DATE), H3 (DEVICE EVAL BY MANUFACTURER), H6 (COMPONENT CODES, EVALUATION METHOD CODES, EVALUATION RESULT CODES, AND EVALUATION CONCLUSION CODES), H7 (IF REMEDIAL TYPE INIT, TYPE), AND H11 (ADDITIONAL MFR NARRATIVE). UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. ANALYSIS OF THE DEVICE CONFIRMED IT WAS OPERATING IN SAFETY MODE AND THAT CRITICAL THERAPY REMAINED AVAILABLE. ENGINEERS DETERMINED THAT THIS DEVICE WAS DEMONSTRATING BEHAVIOR CONSISTENT WITH A HIGH IMPEDANCE WITHIN THE BATTERY. THE HIGH BATTERY IMPEDANCE PUT THIS DEVICE AT RISK OF EXPERIENCING TRANSIENT VOLTAGE DECREASES (AND ASSOCIATED RESETS) DURING TELEMETRY OR OTHER NORMAL, HIGHER-POWER OPERATIONS. WHEN BATTERY VOLTAGE DROPS BELOW A MINIMUM THRESHOLD, A SYSTEM RESET IS PERFORMED. IF THREE SYSTEM RESETS OCCUR WITHIN A 48-HOUR PERIOD, THE DEVICE IS DESIGNED TO ENTER SAFETY MODE OPERATION TO MAINTAIN BACK-UP PACING WITH PRE-DEFINED SETTINGS. BOSTON SCIENTIFIC HAS ISSUED A FIELD SAFETY NOTICE TO NOTIFY PHYSICIANS OF THE POTENTIAL FOR ACCOLADE PACEMAKER AND CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) DEVICES TO EXHIBIT THIS BEHAVIOR. Manufacturer Narrative: THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS B2 (OUTCOMES ATTRIB TO ADV EVENT), B5 (EVENT DESCRIPTION), H1 (TYPE OF REPORTABLE EVENT), AND H6 (IMPACT CODES).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805473&pc=FOZ | 3012307300-2025-03992 | 2025-03-01 05:00 | Death | SMITHS MEDICAL ITALIA SRL | 2025-04-09 04:00 | FOZ | JELCO STRAIGHT RADIOPAQUE INTRAVENOUS CATHETER | Difficult to Remove | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT FOR THE FIRST STICK ATTEMPT, THE DEVICE ADVANCED EASILY WITH NO ISSUE. WHEN PLACED ON BUNG AND TAPED IN, SL WENT TO TOP UP SEDATION DOSE VIA IV AND THERE WAS BLOOD COMING FROM OUTSIDE OF THE BUNG. AFTER NUMEROUS ATTEMPTS TO RETRIEVE THE DEVICE, INCLUDING CUT DOWN, THE CATHETER WAS REMOVED POST MORTEM. WOULD HAVE REQUIRED SURGICAL REMOVAL AS WAS UNABLE TO REMOVE IT FROM VEIN. THERE WAS PATIENT INVOLVEMENT. Manufacturer Narrative: B3: EVENT DATE UNKNOWN. H3: THE REPORTED UNIT WAS NOT RETURNED, THEREFORE NO FAILURE INVESTIGATION COULD BE PERFORMED ON THE PRODUCT. THE VISUAL ANALYSIS PERFORMED IS LIMITED TO THE ANALYSIS OF THE MEDIA PROVIDED. FOUR PHOTOS WERE PROVIDED. IN THE PHOTOS, THE TUBE APPEARED TO BE SEVERED INTO TWO PARTS. THE PORTION REMAINING INSIDE THE HUB APPEARED TO BE CORRECTLY ASSEMBLED WITH THE EYELET AND WITH NO VISIBLE DAMAGES. THE TUBE SHOWS A HORIZONTAL CLEAN CUT IN PROXIMITY OF THE HUB. IT IS IMPROBABLE THAT THIS TYPE OF DAMAGE OCCURRED DURING THE MANUFACTURING PROCESS AS THE MACHINE WHERE THE CATHETER TUBE IS ASSEMBLED WITH THE HUB AND THE EYELET IS EQUIPPED BY A PULL TEST STATION THAT PERFORMS 100% CHECK OF THE ASSEMBLED UNITS TO VERIFY THAT THE CATHETER TUBE IS PROPERLY SECURED TO THE HUB. THE DEFECTIVE UNITS ARE AUTOMATICALLY DISCARDED. THE GOOD PRODUCT UNDERGOES THE ROUTINE STATISTICAL QC TESTS. THE DHR REVIEW SHOWS THAT THE AFOREMENTIONED TESTS HAVE BEEN PERFORMED WITH NO ANOMALIES NOTED; ALL THE SAMPLES PASSED THE TESTS, CONFIRMING COMPLIANCE OF THE PRODUCT WITH THE SPECIFICATIONS. A REVIEW OF THE COMPLAINTS DATABASE HAS BEEN PERFORMED AND NO OTHER COMPLAINT WAS FOUND ON THE LOT INVOLVED. ADDITIONAL INFORMATION ABOUT THE EVENT WAS REQUESTED BUT NOT PROVIDED. THE DESCRIPTION OF THE EVENT REPORTED BY THE CUSTOMER AND THE STATUS OF THE SAMPLE OBSERVED ARE COMPATIBLE WITH A POTENTIAL MISHANDLING OF THE PRODUCT DURING USE/REMOVAL. BASED ON THE EVIDENCE CURRENTLY AVAILABLE, THE INVESTIGATION DID NOT FIND ANY QUALITY RELATED ISSUE THAT COULD HAVE RESULTED IN THE CUSTOMER REPORTED PROBLEM. PROBABLE CAUSE COULD NOT BE DETERMINED WITH CONFIDENCE DUE TO NO PRODUCT RETURNED FOR EVALUATION. HOWEVER, BASED ON THE EVENT DESCRIPTION, MISHANDLING LIKELY OCCURRED AND MAY HAVE BEEN DIRECTLY RELATED TO THE TUBE SEVERANCE. BASED ON THE EVIDENCE CURRENTLY AVAILABLE, THE REPORTED ALLEGATION COULD NOT BE CONFIRMED, THEREFORE NO ACTION WILL BE IMPLEMENTED AT THIS TIME. IF THE PRODUCT WILL BECOME AVAILABLE, WE WILL REOPEN THE INVESTIGATION. COMPLAINTS DATA WILL CONTINUE TO BE MONITORED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805504&pc=DSQ | 2916596-2025-01914 | 2019-12-01 05:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿ECHOCARDIOGRAPHIC PREDICTORS OF VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES AND IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH ARRHYTHMIA, RIGHT VENTRICULAR DYSFUNCTION, INTERVENTRICULAR SEPTAL SHIFT, SUCTION EVENTS, AND DEATH. THIS RETROSPECTIVE STUDY EVALUATED 264 PATIENTS WITH A PRE-EXISTING IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD), INCLUDING SINGLE CHAMBER ICD (N = 101), DUAL CHAMBER ICD (N = 39), AND CARDIAC RESYNCHRONIZATION THERAPY WITH DEFIBRILLATOR (CRT-D) (N = 124), WHO UNDERWENT LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT BETWEEN MAY2011 AND DEC2019 TO EVALUATE THE PREDICTIVE VALUES OF PREOPERATIVE ECHOCARDIOGRAPHIC (ECHO) PARAMETERS FOR OCCURRENCE OF VENTRICULAR ARRHYTHMIAS (VA) IN THESE PATIENTS. OF THE 264 PATIENTS WHO UNDERWENT LVAD IMPLANT, 154 HAD A HEARTWARE HVAD, 21 HAD A HEARTMATE II, AND 89 HAD A HM3. RESULTS WERE NOT SPECIFIED BY LVAD TYPE. THE PATIENTS WERE PREDOMINANTLY MALE (89%) WITH NONISCHEMIC CARDIOMYOPATHY (NICM) AND A MEAN AGE OF 59 YEARS. ONLY EPISODES OF MONOMORPHIC VENTRICULAR TACHYCARDIA (MVT), POLYMORPHIC VENTRICULAR TACHYCARDIA (PVT), AND/OR VENTRICULAR FIBRILLATION (VF) TERMINATED BY APPROPRIATE ICD THERAPY VIA ANTI-TACHYCARDIA PACING OR SHOCK WERE ANALYZED. ECHOCARDIOGRAPHIC PARAMETERS WERE COMPARED IN PATIENTS WITH AND WITHOUT VAS PRIOR TO LVAD IMPLANT, AT 1-MONTH POST-IMPLANT, AND AT 1 YEAR POST-IMPLANT. TO CLARIFY THE POSSIBLE RELATIONSHIP WITH THE EXTENT OF LEFT VENTRICULAR (LV) UNLOADING AND THE OCCURRENCE OF VA, TRANSTHORACIC ECHOCARDIOGRAPHY (TTE) OF THE POSITION OF THE INTERVENTRICULAR SEPTUM (IVS) WAS ASSESSED AT 1 MONTH PRE-IMPLANT, 1-MONTH POST-IMPLANT, AND AT 3-MONTH INTERVALS THEREAFTER. A TOTAL OF 102 PATIENTS (39%) WITH AND WITHOUT VAS HAD APPROPRIATE ICD INTERVENTIONS IN THE FIRST-YEAR POST-IMPLANT, WITH INTERVENTIONS OCCURRING FOR MVT IN 100 PATIENTS, PVT AND VF IN 27 PATIENTS, AND BOTH VAS IN 19 PATIENTS; OF NOTE, THESE CATEGORIES WERE NOT MUTUALLY EXCLUSIVE. MEAN TIME FROM LVAD IMPLANT TO FIRST VA EPISODE IN THE FIRST-YEAR POST-IMPLANT WAS 83 DAYS. ATRIAL FIBRILLATION WAS ALSO NOTED TO HAVE OCCURRED IN A TOTAL OF 194 PATIENTS, OCCURRING IN 80 PATIENTS WITH POST-LVAD VAS AND 114 WITHOUT POST-LVAD VAS (P = 0.15). A TOTAL OF 55 PATIENTS PASSED AWAY WITHIN A YEAR OF LVAD IMPLANT, 18 PATIENTS WITH POST-LVAD VAS AND 37 WITHOUT POST-LVAD VAS: NO IMPACT OF VAS ON MORTALITY IN THE FIRST YEAR AFTER LVAD IMPLANT WAS FOUND (P = 0.31). ECHO RESULTS FOUND THAT AN INCREASED PRE-IMPLANT LV END-DIASTOLIC DIAMETER (LVEDD) >= 72MM PREDICTED THE OCCURRENCE OF VAS AFTER LVAD IMPLANT IN PATIENTS WITH ISCHEMIC CARDIOMYOPATHY (P = 0.012), WHILE A LARGER PRE-IMPLANT RIGHT VENTRICULAR END-DIASTOLIC DIAMETER (RVEDD) >= 46MM WAS PREDICTIVE IN NICM PATIENTS (P = 0.004). PRE-LVAD VAS WERE THE MOST SIGNIFICANT PREDICTOR OF POST-LVAD VAS IN NICM PATIENTS. A LARGER RVEDD AT 1-YEAR POST-IMPLANT WAS HIGHLY ASSOCIATED WITH VA IN THE FIRST POST-IMPLANT YEAR (50MM IN THOSE WITH POST-LVAD VAS VS. 45MM IN THOSE WITHOUT POST-LVAD VAS, P = 0.001). THE REDUCTION OF RV DIAMETER >= 2MM AT THE 1-YEAR FOLLOW-UP WAS SIGNIFICANTLY ASSOCIATED WITH FEWER VAS IN THE FIRST POST-IMPLANT YEAR (P = 0.032). ALL PATIENTS DEMONSTRATED A SIGNIFICANT DECREASED IN LVEDD AS WELL AS A REDUCTION OF SEVERITY OF MITRAL REGURGITATION AND TRICUSPID REGURGITATION DURING THE 1-YEAR FOLLOW-UP PERIOD, REFLECTING LV UNLOADING THROUGH THE LVAD. A SIGNIFICANT DECLINE OF RIGHT VENTRICULAR-RIGHT ATRIAL (RV-RA) PRESSURE GRADIENT WAS OBSERVED ON THE TOTAL PATIENT POPULATION ON THE FIRST POST-IMPLANT ECHO (P < 0.001) THAT REMAINED STATIC OVER THE REMAINING FOLLOW-UP PERIOD (P = 0. 08). DESPITE THIS RV-RA GRADIENT PRESSURE REDUCTION, RVEDD OVERALL DID NOT CHANGE SIGNIFICANTLY IN THE FIRST POST-IMPLANT YEAR. NONETHELESS, RVEDD OVERALL SHOWED A TREND TO REDUCTION AT FIRST POST-IMPLANT CONTROL (P = 0.05) AS WELL AS AT 1-YEAR (P = 0.06). A SMALL BUT STATISTICALLY SIGNIFICANT REDUCTION OF TRICUSPID ANNULAR PLANE SYSTOLIC EXCURSION (TAPSE) AFTER LVAD-IMPLANT WAS NOTED IN COMPARISON TO PRE-LVAD EXAMINATION (P < 0.001). BY TAKING RV-RA GRADIENT IMPROVEMENT INTO ACCOUNT, IT WAS POSTULATED THAT TAPSE REDUCTION WAS MORE LIKELY CAUSED BY THE LVAD SURGERY ITSELF RATHER THAN A DETERIORATION OF GLOBAL RV FUNCTION. ALL OTHER OBTAINED ECHO PARAMETERS WERE SIMILAR IN PATIENTS WITH AND WITHOUT VAS. TTE ASSESSMENT RESULTS FOUND THAT 10 PATIENTS WITH POST-LVAD VAS AND 19 WITHOUT POST-LVAD VAS DEMONSTRATED A LEFTWARD SHIFT OF THE IVS. AT 1-YEAR FOLLOW-UP, 12 PATIENTS FROM EACH GROUP DEMONSTRATED A LEFTWARD SHIFT OF THE IVS. NINE SUCTION EVENTS WERE ALSO DETECTED DURING FOLLOW-UP VISITS, BUT ONLY ONE PATIENT WITH AN HVAD HAD A SUCTION EVENT IDENTIFIED AS CHRONOLOGICALLY RELATED TO VA. THE STUDY CONCLUDED THAT LARGE LV AND RV DIAMETERS BEFORE LVAD IMPLANT PREDICTED THE OCCURRENCE OF VAS AFTER IMPLANT IN BOTH ICM AND NICM PATIENTS. PERSISTENT RV REMODELING POST-LVAD WAS ALSO ASSOCIATED WITH VAS. Manufacturer Narrative: SECTION B: DATE OF DEATH ESTIMATED AS EVENT DATE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT THE TIME OF THE EVENT. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. SECTION B3: DATE OF EVENT HAS BEEN LISTED AS 01DEC2019 AS PATIENTS WERE PLANTED BETWEEN MAY2011 AND DEC2019. AUTHOR INFORMATION: EFIMOVA, E., ZEYNALOVA, S., EIFERT, S., DASHKEVICH, A., BORGER, M. A., MEYER, A. L., GARBADE, J., DARMA, A., BODE, K., & ARYA, A. (2024). ECHOCARDIOGRAPHIC PREDICTORS OF VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES AND IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR. JOURNAL OF CARDIOVASCULAR ELECTROPHYSIOLOGY, 36(2), 387¿395. HTTPS://DOI.ORG/10.1111/JCE.16539. DEPARTMENT OF ELECTROPHYSIOLOGY, LEIPZIG HEART CENTER, LEIPZIG, GERMANY; INSTITUTE OF MEDICAL INFORMATICS, STATISTICS AND EPIDEMIOLOGY, UNIVERSITY OF LEIPZIG, LEIPZIG, GERMANY; DEPARTMENT OF CARDIAC SURGERY, UNIVERSITY CLINIC OF CARDIAC SURGERY, LEIPZIG HEART CENTER, LEIPZIG, GERMANY; DEPARTMENT OF CARDIAC SURGERY, HEIDELBERG UNIVERSITY HOSPITAL, HEIDELBERG, GERMANY; DEPARTMENT OF CARDIAC SURGERY, KLINIKUM LINKS DER WESER, BREMEN, GERMANY; DEPARTMENT OF CARDIOLOGY, UNIVERSITY HOSPITAL HALLE, MARTIN-LUTHER UNIVERSITY HALLE-WITTENBERG, HALLE (SAALE), GERMANY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805588&pc=NPT | 2025587-2025-02488 | 2024-01-17 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, THE PATIENT DIED AT ANOTHER HOSPITAL DUE TO HEART FAILURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805592&pc=DSQ | 2916596-2025-02108 | 2025-03-15 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUBARACHNOID HEMORRHAGE. THE DEATH WAS NOT DEVICE RELATED, AND THE PUMP WAS WORKING AS EXPECTED FOR THE WHOLE TIME OF SUPPORT. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 8.5¿ FROM THE PUMP HEADER. A DISTAL PORTION OF THE PUMP CABLE, MEASURING APPROXIMATELY 17¿ FROM THE INLINE CONNECTOR, WAS RETURNED ATTACHED TO THE MODULAR CABLE. THE SEALED OUTFLOW GRAFT WAS SECURED TO THE PUMP COVER OUTLET PORT. THE OUTFLOW GRAFT BEND RELIEF WAS ATTACHED AT THE GRAFT HARDWARE. THE APICAL CUFF WAS NOT RETURNED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE RETRIEVED LVAD EVENT LOG FILE CONTAINED RELEVANT DATA FROM (B)(6) 2025. THE DATA CAPTURED AFTER (B)(6) 2025 OCCURRED DURING THE TESTING OF (B)(6) AT ABBOTT. MEAN FLOW TYPICALLY RANGED FROM 3.7-4.8 LITERS PER MINUTE (LPM). MEAN FLOW DECREASED IN THE FINAL 2 MINUTES OF RELEVANT DATA, WHICH APPEARED CONSISTENT WITH THE REPORTED PATIENT OUTCOME. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, BLEEDING, AND DEATH. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805882&pc=NIK | 9614453-2025-01397 | 2023-04-15 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-09 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED TWO DAYS POST PROCEDURE AND UPGRADE TO CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD, IMPLANTED: (B)(6) 2022. PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805908&pc=NPT | 2025587-2025-02490 | 2024-11-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, THE PATIENT DIED AT ANOTHER HOSPITAL DUE TO WORSENING HEART FAILURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21805953&pc=BZD | 2518422-2025-103692 | 2025-03-31 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Fibrosis | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY DUE TO PULMONARY FIBROSIS. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806009&pc=PCU | 3005099803-2025-01540 | 2014-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE: "CLINICAL COURSE FOR PANCREATIC NECROSIS AND PANCREATIC PSEUDOCYSTS DUE TO SEVERE ACUTE OR CHRONIC PANCREATITIS," BY STEFANO FUSCO, ET AL. PER THE ARTICLE, A RETROSPECTIVE OBSERVATIONAL SINGLE-CENTER STUDY EXAMINED PATIENTS WITH SEVERE PANCREATITIS ON 46 PATIENTS FROM JANUARY 2014 TO DECEMBER 2020. OF THE PATIENTS, 27 DEVELOPED WALLED-OFF NECROSIS (WON) AND 19 DEVELOPED PANCREATIC PSEUDOCYST (PPC). SEVEN PATIENTS RECEIVED TREATMENT WITH LAMS, INCLUDING TWO FROM THE WON GROUP AND FIVE FROM THE PPC GROUP. ONE PATIENT SUCCUMBED TO INTRAPERITONEAL BLEEDING FOLLOWING A NECROSECTOMY PERFORMED TWO WEEKS EARLIER AT AN EARLY STAGE OF ACUTE PANCREATITIS (AP). Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED DATE OF PATIENT'S DEATH OF JANUARY 31, 2014, WAS CHOSEN AS THE BEST ESTIMATE BASED ON 30 DAYS FROM THE RETROSPECTIVE STUDY START DATE OF JANUARY 2014. BLOCK B3: THE EXACT EVENT ONSET DATE IS UNKNOWN. THE PROVIDED EVENT DATE OF JANUARY 1, 2014, WAS CHOSEN AS THE BEST ESTIMATE BASED ON THE RETROSPECTIVE STUDY START DATE OF JANUARY 2014. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: STEFANO FUSCO; GRETA M. HANKE; KARSTEN BURINGER; LISA MINN; GUNNAR BLUMENSTOCK; ULRIKE SCHEMPF; MARTIN GOTZ; NISAR P. MALEK; DORTE WICHMANN; CHRISTOPH R. WERNER. "CLINICAL COURSE FOR PANCREATIC NECROSIS AND PANCREATIC PSEUDOCYSTS DUE TO SEVERE ACUTE OR CHRONIC PANCREATITIS." THERAPEUTIC ADVANCES IN GASTROENTEROLOGY 2024; VOL. 17: 1-14. BLOCK H6: IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF INTRAPERITONEAL BLEEDING. IMDRF IMPACT CODE F02 CAPTURES THE PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806036&pc=DSQ | 3007042319-2025-00736 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-09 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Renal Failure; Sepsis; Tachycardia; Respiratory Failure; Multiple Organ Failure; Ischemia Stroke; Heart Failure/Congestive Heart Failure; Liver Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION PERFORMED VIA MEDIAN STERNOTOMY (MS) VERSUS A LESS-INVASIVE APPROACH RELATED TO RIGHT VENTRICULAR FAILURE (RVF). THE AUTHORS DESCRIBED SOME PATIENT DEATHS WITHIN FOURTEEN DAYS OF VAD IMPLANT DUE TO SEVERE RVF. THERE WERE OTHER DEATHS WITHIN THIRTY DAYS OF THE VAD IMPLANT WITH A HIGHER NUMBER OF DEATHS IN THE MEDIAN STERNOTOMY GROUP VERSUS THE LESS-INVASIVE GROUP. THE CAUSES OF DEATH CONSISTED OF SEPSIS, RECURRENT VENTRICULAR TACHYCARDIA (VT) WHICH REQUIRED DEFIBRILLATION, RENAL FAILURE WHICH REQUIRED RENAL REPLACEMENT THERAPY, RV DYSFUNCTION, HYPOXEMIC RESPIRATORY FAILURE (PULMONARY HEMORRHAGE, OR ACUTE RESPIRATORY DISTRESS SYNDROME), INTRACRANIAL HEMORRHAGE WITH HERNIATION, DIFFUSE BRAIN INJURY, ISCHEMIC CEREBROVASCULAR ACCIDENT (CVA), BOWEL ISCHEMIA, VASOPLEGIA, BOWEL PSEUDO-OBSTRUCTION, PROGRESSIVE MULTIORGAN FAILURE, RVF, AND LIVER/KIDNEY FAILURE. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED SEVERE RVF DEFINED AS ELEVATED POSTOPERATIVE CENTRAL VENOUS PRESSURE (CVP) COUPLED WITH CLINICAL MANIFESTATIONS OF PERIPHERAL EDEMA, ASCITES, HEPATOMEGALY, OR LABORATORY EVIDENCE OF WORSENING HEPATIC OR RENAL DYSFUNCTION AND THE NEED FOR INOTROPIC SUPPORT, INHALED NITRIC OXIDE, OR INTRAVENOUS VASODILATORS. OTHER PATIENTS EXPERIENCED STROKES, MAJOR BLEEDING IN WHICH PATIENTS WERE TRANSFUSED WITH PACKED RED BLOOD CELLS (PRBC), FRESH FROZEN PLASMA, OR PLATELETS, READMISSION FOR RIGHT HEART FAILURE, EXTENDED STAYS IN THE INTENSIVE CARE UNIT (ICU), AND PATIENTS WHO REQUIRED MECHANICAL VE NTILATION. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: RIGHT VENTRICULAR FUNCTION FOLLOWING STERNOTOMY VERSUS A LESS-INVASIVE APPROACH FOR LEFT VENTRICULAR ASSIST DEVICE IMPLANT: RETROSPECTIVE COHORT STUDY. JOURNAL OF CARDIOTHORACIC AND VASCULAR ANESTHESIA. 2025. 39; 79-87. DOI: 10.1053/J.JVCA.2024.04.044 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806058&pc=BZD | 2518422-2025-103816 | 2025-03-12 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT WAS SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION IT WAS NOTED THERE WERE NO ERROR CODES. THE DEVICE WAS TESTED AND PASSED. THE DEVICE WAS RETURNED TO STOCK. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806227&pc=FDS | 2518897-2025-00004 | 2025-04-02 04:00 | Death | HOYA CORPORATION PENTAX TOKYO OFFICE | 2025-04-09 04:00 | FDS | PENTAX MEDICAL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Respiratory Arrest |  |  | 1 | Event Description: ON 02-APR-2025, PENTAX MEDICAL BECAME AWARE OF A PATIENT DEATH INVOLVING A PENTAX MEDICAL VIDEO GASTROSCOPE, MODEL EG34-I10, SERIAL NUMBER (B)(6) ACCORDING TO THE INITIAL REPORT, THE PHYSICIAN PERFORMED A THERAPEUTIC ENDOSCOPY TO IDENTIFY A GASTROINTESTINAL (GI) BLEEDING SOURCE. THE BLEEDING SOURCE WAS FOUND, HEMOSTASIS WAS ACHIEVED USING ELECTROCOAGULATION, AND CLIPS WERE DEPLOYED TO PREVENT FURTHER BLEEDING. DURING THE PROCEDURE, THE PATIENT'S PHYSICAL CONDITION DETERIORATED. THE ENDOSCOPE WAS REMOVED, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND A CODE BLUE TEAM WAS ACTIVATED. THE PATIENT WAS PRONOUNCED DEAD ON THE SAME DAY, (B)(6) 2025. THE PENTAX MEDICAL VIDEO GASTROSCOPE REPORTEDLY FUNCTIONED NORMALLY DURING THE PROCEDURE. THERE WAS NO INDICATION OF DEVICE MALFUNCTION, AND THE DEATH WAS NOT ATTRIBUTED TO THE ENDOSCOPE. THE PROCEDURE WAS SCHEDULED IN ADVANCE AND WAS NOT AN EMERGENCY CASE. THE PATIENT HAD A HISTORY OF HEART SURGERY; HOWEVER, DETAILED MEDICAL HISTORY WAS NOT AVAILABLE. THIS EVENT MEETS THE REQUIREMENTS FOR FDA REPORTABILITY; HOWEVER, SUBMISSION OF THIS REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, MANUFACTURER OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT.Event Description: REFER TO H11. Manufacturer Narrative: F10 CONTINUED: INTERNATIONAL MEDICAL DEVICE REGULATORS FORUM (IMDRF) ADVERSE EVENT REPORTING HEALTH EFFECT - CLINICAL CODE: 1762 CARDIAC ARREST, 4461 RESPIRATORY ARREST HEALTH EFFECT - IMPACT CODE: 1802 DEATH, 4635 UNEXPECTED DETERIORATIONS MEDICAL DEVICE PROBLEM CODE: 2993 ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM COMPONENT CODE: 4776 INSUFFICIENT INFORMATION. ADDITIONAL EQUIPMENT USED: A PENTAX MEDICAL VIDEO PROCESSOR (MODEL EPK-I8020C, SERIAL NUMBER (B)(6), FIRMWARE VERSION 0107C-1.0055) WAS USED IN CONJUNCTION WITH THE VIDEO GASTROSCOPE. THE VIDEO PROCESSOR FUNCTIONED NORMALLY AND NO MALFUNCTION WAS NOTED. INVESTIGATION IS IN-PROCESS. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE FILED WITH THE NEW INFORMATION. Manufacturer Narrative: CORRECTION INFORMATION: B4: DATE OF THIS REPORT, B5: REFER TO H11, F7: FOLLOW UP #01, F11: UPDATED DATES, F13: UPDATED DATES. ADDITIONAL INFORMATION: D4: D4: PRIMARY UNIQUE DEVICE IDENTIFIER (UDI) CORRECTED, F9: AGE OF DEVICE, H11: EVALUATION SUMMARY. F10 CONTINUED: INTERNATIONAL MEDICAL DEVICE REGULATORS FORUM (IMDRF) ADVERSE EVENT REPORTING. COMPONENT CODE: ADDED 4755, REMOVED - 4776 INSUFFICIENT INFORMATION. EVALUATION SUMMARY: PENTAX MEDICAL AMERICA AND THE JAPANESE INVESTIGATION TEAM CONDUCTED A GOOD FAITH EFFORT (GFE) TO OBTAIN ADDITIONAL INFORMATION REGARDING THIS COMPLAINT. THE FOLLOWING TWO QUESTIONS WERE SUBMITTED TO THE SITE REPRESENTATIVE MULTIPLE TIMES BETWEEN 11-APR-2025 AND 15-MAY-2025: HAS THE SUBJECT ENDOSCOPE (EG34-I10_Q00Z7AB016) CONTINUED TO BE USED AT THE FACILITY AFTER THE COMPLAINT? WHO DETERMINED THAT THERE WAS NOTHING WRONG WITH THE SCOPE (A PHYSICIAN OR A FACILITY STAFF MEMBER)? AFTER A DELAY, A RESPONSE WAS PROVIDED BY THE SITE REPRESENTATIVE ON 15-MAY-2025. THE RESPONSE INDICATED THAT THE ENDOSCOPE IS STILL IN USE AT THE FACILITY AND NO COMPLAINTS OR ISSUES HAVE BEEN REPORTED. IT WAS CONFIRMED BY A FACILITY STAFF MEMBER (ENDOSCOPY SUPERVISOR) THAT THE DEVICE FUNCTIONED NORMALLY DURING THE PROCEDURE. AFTER THE PROCEDURE WAS ABORTED DUE TO THE PATIENT'S CONDITION, THE ENDOSCOPE WAS LEAK TESTED AND REPROCESSED, AND NO ABNORMALITIES WERE FOUND. UPON RECEIVING THIS INFORMATION, THE JAPANESE INVESTIGATION REPRESENTATIVE CONFIRMED ON 16-MAY-2025 THAT THERE WERE NO PROBLEMS WITH THE ENDOSCOPE AND THAT IT COULD CONTINUE TO BE USED. SINCE THE DEVICE HAD BEEN FUNCTIONING NORMALLY AND HAD UNDERGONE STANDARD REPROCESSING WITHOUT ANY ISSUES, IT WAS DETERMINED THAT FURTHER INSPECTION OF THE DEVICE WAS UNNECESSARY. THE EVENT THAT OCCURRED WAS A DETERIORATION IN THE PATIENT'S CONDITION. BASED ON THE INVESTIGATION AND GFE, THE GASTROINTESTINAL BLEEDING HAD OCCURRED BEFORE THE ENDOSCOPE WAS INSERTED, AND IT WAS DETERMINED THAT IT WAS NOT DUE TO EQUIPMENT FAILURE. FURTHERMORE, BASED ON THE TREND ANALYSIS, THERE IS NO SIGNIFICANT INCREASE OR UPWARD TREND IN REPAIR COMPLAINTS RELATED TO THIS FAILURE MODE/HAZARD. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED BY THE MANUFACTURER. THE REVIEW BY DHR CONFIRMED THAT THE INSTRUMENT WAS MANUFACTURED UNDER NORMAL CONDITIONS ON 05-JUL-2024, PASSED ALL REQUIRED INSPECTIONS, AND WAS RELEASED ACCORDINGLY. ALSO, THERE WERE NO REWORKS OR CONCESSIONS, AND THE DATES OF APPROVAL FOR SHIPMENT AND ACTUAL DATE SHIPPED WERE CONFIRMED ON 16-JUL-2024. PENTAX MEDICAL HAS NOT RECEIVED ANY FURTHER INFORMATION FOR THIS EVENT AND THEREFORE, CONSIDERS THIS MEDWATCH REPORT CLOSED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806234&pc=FRN | MW5168746 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-04-09 04:00 | FRN | NOVUM IQ LARGE VOLUME INFUSION PUMP (LVP) | Backflow; Excess Flow or Over-Infusion; Insufficient Flow or Under Infusion; Fail-Safe Did Not Operate | Cardiac Arrest; Phototoxicity; Respiratory Arrest; Reaction to Medicinal Component of Device |  |  | 1 | Event Description: I DO NOT HAVE THE EXACT DATES- BUT THESE EVENTS WERE DOCUMENTED AT (B)(6). NEW BAXTER PUMPS WERE IMPLEMENTED TO (B)(6) HOSPITALS IN 2024/2025 (NOVUM IQ LARGE VOLUME PUMPS AND NOVUM IQ SYRINGE PUMPS). ISSUES INCLUDE BUT ARE NOT LIMITED TO BLOOD BACKING UP DURING INFUSIONS, MEDICATIONS INFUSING TOO SLOWLY OR RAPIDLY, AND SAFETY MECHANISMS FAILING. THE TWO MOST CONCERNING EVENTS HAPPENED AT (B)(6) HOSPITAL AND (B)(6) HOSPITAL. A DEFECT WITH THE PUMPS THAT WAS ABLE TO BE REPLICATED BY RISK MANAGEMENT- IF THE TUBING TERMINATING AT THE PT RECEIVES JUST THE RIGHT AMOUNT OF TENSION IT CAUSES A SIPHONING EFFECT WHERE THE ENTIRE BAG OF MEDICATION WILL RAPIDLY INFUSE INTO THE PATIENT, WHILE IT APPEARS THAT THE MEDICATION IS INFUSING APPROPRIATELY ON THE PUMP SCREEN. TO MY KNOWLEDGE THE FIRST EVENT HAPPENED IN FEB 2025 AT (B)(6) HOSPITAL WHERE A BAG OF LIDOCAINE INFUSED INTO A PATIENT. THE PATIENT SHOWED SIGNS OF LIDOCAINE TOXICITY AND REQUIRED LIPIDS BUT SURVIVED. THE SECOND EVENT HAPPENED AT (B)(6) HOSPITAL, WHERE AN ENTIRE BAG OF LEVOPHED INFUSED INTO THE PATIENT RESULTING IN THE PATIENT'S DEATH. THERE ARE NO SAFETY MEASURES IN EFFECT AT THIS TIME TO PREVENT THIS FROM HAPPENING AGAIN. NURSES AND DOCTORS ARE REPORTING ALL EVENTS INTERNALLY. REFERENCE REPORT: MW5168747.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806235&pc=FRN | MW5168747 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-04-09 04:00 | FRN | NOVUM IQ LARGE VOLUME INFUSION PUMP (LVP) | Backflow; Inaccurate Flow Rate; Incorrect, Inadequate or Imprecise Result or Readings; Defective Device; Connection Problem; Free or Unrestricted Flow; No Fail-Safe Mechanism | Insufficient Information |  |  | 1 | Event Description: I DO NOT HAVE THE EXACT DATES- BUT THESE EVENTS WERE DOCUMENTED AT (B)(6). NEW BAXTER PUMPS WERE IMPLEMENTED TO (B)(6) HOSPITALS IN 2024/2025 (NOVUM IQ LARGE VOLUME PUMPS AND NOVUM IQ SYRINGE PUMPS). ISSUES INCLUDE BUT ARE NOT LIMITED TO BLOOD BACKING UP DURING INFUSIONS, MEDICATIONS INFUSING TOO SLOWLY OR RAPIDLY, AND SAFETY MECHANISMS FAILING. THE TWO MOST CONCERNING EVENTS HAPPENED AT (B)(6) HOSPITAL AND (B)(6) HOSPITAL. A DEFECT WITH THE PUMPS THAT WAS ABLE TO BE REPLICATED BY RISK MANAGEMENT- IF THE TUBING TERMINATING AT THE PT RECEIVES JUST THE RIGHT AMOUNT OF TENSION IT CAUSES A SIPHONING EFFECT WHERE THE ENTIRE BAG OF MEDICATION WILL RAPIDLY INFUSE INTO THE PATIENT, WHILE IT APPEARS THAT THE MEDICATION IS INFUSING APPROPRIATELY ON THE PUMP SCREEN. TO MY KNOWLEDGE THE FIRST EVENT HAPPENED IN FEB 2025 AT (B)(6) HOSPITAL WHERE A BAG OF LIDOCAINE INFUSED INTO A PATIENT. THE PATIENT SHOWED SIGNS OF LIDOCAINE TOXICITY AND REQUIRED LIPIDS BUT SURVIVED. THE SECOND EVENT HAPPENED AT (B)(6) HOSPITAL, WHERE AN ENTIRE BAG OF LEVOPHED INFUSED INTO THE PATIENT RESULTING IN THE PATIENT'S DEATH. THERE ARE NO SAFETY MEASURES IN EFFECT AT THIS TIME TO PREVENT THIS FROM HAPPENING AGAIN. NURSES AND DOCTORS ARE REPORTING ALL EVENTS INTERNALLY. REFERENCE REPORT: MW5168746.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806588&pc=OZD | 1220648-2025-27330 |  | Death | ABIOMED, INC. | 2025-04-09 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: ABIOMED JAPAN FORWARDED A PUBLICATION TITLED "ECHOCARDIOGRAPHY-GUIDED PERCUTANEOUS VENTRICULAR ASSIST DEVICE FOR POSTINFARCT VENTRICULAR SEPTAL RUPTURE," WHICH NOTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA 5.5 AND PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT. IN SUMMARY, THE PATIENT EXPERIENCED "UNCONTROLLED BLEEDING IN SETTING OF SYSTEMIC ANTICOAGULATION, LEADING TO HEMORRHAGIC SHOCK." CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE ACTUAL OR BEST ESTIMATE DATE OF THE EVENT IS UNKNOWN/COULD NOT BE DETERMINED. CONSEQUENTLY, SECTION B5 WAS LEFT BLANK. PUBLICATION REFERENCE: AVINASH D¿SOUZA, DO ET. AL (2021). "ECHOCARDIOGRAPHY-GUIDED PERCUTANEOUS VENTRICULAR ASSIST DEVICE FOR POSTINFARCT VENTRICULAR SEPTAL RUPTURE." CARDIOVASCULAR IMAGING CASE REPORTS. DECEMBER 2024 VOLUME 8 NUMBER 12.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806695&pc=NPT | 2025587-2025-02503 | 2024-07-29 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Septic Shock | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 4 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. PER THE PATIENT'S FAMILY, THE CAUSE OF DEATH WAS SEPTIC SHOCK. THE DETAILS REGARDING HOW THE PATIENT DEVELOPED SEPTIC SHOCK ARE UNSPECIFIED. IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS NOT BELIEVED TO HAVE BEEN THE DIRECT CAUSE OF DEATH. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806708&pc=DTB | 2649622-2025-09559 | 2025-03-23 04:00 | Death | MPRI | 2025-04-09 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING THAT RESULTED IN INAPPROPRIATE ATRIAL PACING AND TRIGGERED DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS AT TIMES. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING THAT RESULTED IN INAPPROPRIATE ATRIAL PACING AND TRIGGERED DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS AT TIMES. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT PASSED TWO WEEKS AFTER THE RA UNDERSENSING WAS OBSERVED. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED A PROXIMAL SEGMENT OF THE LEAD 7.2 CM LONG WAS RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 620RG33 HEART RING, IMPLANTED: (B)(6) 2019. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806712&pc=DSQ | 3007042319-2025-00744 | 2025-03-13 04:00 | Death | HEARTWARE, INC. | 2025-04-09 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY EXPERIENCED A STROKE, WHICH WAS IDENTIFIED AS THE CAUSE OF DEATH, WHILE USING A VENTRICULAR ASSIST DEVICE. THE PATIENT TRANSITIONED TO HOSPICE CARE AND SUBSEQUENTLY PASSED AWAY. A COMPUTED TOMOGRAPHY (CT) SCAN OF THE BRAIN/HEAD WAS CONDUCTED AS PART OF THE DIAGNOSTIC PROCESS. IT WAS NOTED THAT THE PATIENT WAS COMPLIANT WITH THEIR THERAPEUTIC ANTICOAGULATION THERAPY, INCLUDING ANTICOAGULANT AND ASPIRIN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF NEUROLOGICAL DYSFUNCTION. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806789&pc=DSQ | 2916596-2025-02131 | 2025-03-17 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21806916&pc=BZD | 2518422-2025-103765 | 2025-03-20 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE SPOUSE OF THE PATIENT REPORTED THE PATIENT WAS DECEASED. THE REPORTER INQUIRED ABOUT THE RETURN AND/OR DISPOSAL OF THE DEVICE. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE REPORTER. THE REPORTER WAS ADVISED IT WAS SAFE TO DISPOSE OF THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807098&pc=BZD | 3009096682-2025-00003 | 2025-02-22 05:00 | Death | DEVILBISS HEALTHCARE, LLC | 2025-04-09 04:00 | BZD | INTELLIPAP AUTO BILEVEL W/HEATED HUMIDIFIER | Detachment of Device or Device Component | Unspecified Respiratory Problem | K091919 |  | 1 | Event Description: REACT HEALTH RECEIVED A COMPLAINT FROM DEVILBISS HEALTHCARE, LLC BY EMAIL ON (B)(6) 2025 STATING THAT A PATIENT WAS IN HOSPICE CARE AND PASSED AWAY WHILE USING THE DEVILBISS INTELLIPAP DEVICE. THE PATIENT'S DAUGHTER ALLEGES THAT THE CPAP TUBING BECAME DISCONNECTED FROM THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO REACT HEALTH AS OF THIS REPORT. REACT HEALTH IS FILING AN INITIAL REPORT. IF FURTHER INFORMATION BECOMES AVAILABLE, A FOLLOW UP REPORT WILL BE FILED.Event Description: FROM DEVILBISS COMPLAINT RECEIVED BY EMAIL: "I WOULD LIKE TO FILE A COMPLAINT CONCERNING A MEDICAL DEVICE. ON (B)(6) 2025, MY MOTHER WAS PUT ON HOSPICE. SHE WAS DISCHARGED FROM THE HOSPITAL AND WENT HOME AS BIPAP DEPENDENT. MY MOTHER HAD COPD AND THOUGH SHE WAS PUT ON HOSPICE SHE WAS NO WHERE NEAR READY TO PASS. MY MOTHER ATE DINNER ON FRIDAY NIGHT, SHE GOT UP BY HERSELF, USED THE RESTROOM AND WAS FULLY ALLERT. SHE WENT TO SLEEP THAT NIGHT AT 1045PM AND WOKE UP SOMETIME BETWEEN 530 AND 6AM. AT THIS TIME MY MOTHER SAT UP IN HER BED AND WHEN DOING SO HER BIPAP TUBE CAME COMPLETELY OFF OF HER MACHINE. ONCE THIS HAPPENED SHE WAS UNABLE TO REATTACH IT AND I WAS TOO LATE GETTING TO HER. THE MASK WAS STILL FULLY ATTACHED TO HER FACE BUT THE END THAT ATTACHES TO THE BIPAP MACHINE CAME LOOSE. I HAD NO IDEA THAT COULD HAPPEN OR I WOULD HAVE STAYED WHERE I COULD WATCH HER EVERY SECOND OF THE DAY. THE MACHINE HAS NO EJECTION PORT, NO ALARM AND JUST SLIDES RIGHT OFF THE MACHINE. I WAS TOO LATE, BY THE TIME I GOT TO MY MOTHER SHE COULD NOT CATCH HER BREATH AND DIED AT 642AM. MY MOTHER DIED BECAUSE OF THIS MACHINES FAILURE MY MOTHER SUFFERED AND DIED A HORRIBLE DEATH. SOMETHING NEEDS TO BE DONE TO MAKE THESE MACHINES SAFE FOR PEOPLE THAT ARE DEPENDENT UPON THEM. THE SIMPLE SLIDING ON AND OFF IS NOT SAFE AND KILLS PEOPLE WHO ARE DEPENDENT UPON THESE MACHINES. THE MACHINE SHE WAS USING: DEVILBISS INTELLIPAP." SEE DEVILBISS COMPLAINT# (B)(4). MW5168016. Manufacturer Narrative: REACT HEALTH HAD PREVIOUSLY SUBMITTED A REPORT STATING THAT A PATIENT HAD DIED WHILE USING A DEVILBISS INTELLIPAP BIPAP. ADDITIONAL INFORMATION WAS RECEIVED ON 4/15/2025 FROM THE HOSPICE PROVIDER ON A PATIENT COMPLAINT FORM. THE HOSPICE PROVIDER ALSO FORWARDED A SERVICE REPORT FROM A THIRD PARTY SERVICE PROVIDER. PER THEIR PAPERWORK, IT WAS NOTED THE STREETS DME LLC MEDICAL DELIVERY TECHNICIAN WHO HAD DELIVERED THE BIPAP ON 2/20/2025 WITH ALL SUPPLIES (MASK, PRESSURE LINE ADAPTOR AND GREY CORRUGATED TUBING) NOTICED THE PATIENT HAD THEIR OWN 5 LITER AND 10 LITER CONCENTRATOR AND ALSO A PHILIPS RESPIRONICS TRILOGY WITH BLUE TUBING. ON 2/24/2025 THE DELIVERY TECHNICIAN PICKED UP ALL EQUIPMENT DUE TO PATIENT PASSING AND NOTICED THEIR BLUE TUBING FROM THE PHILIPS RESPIRONICS TRILOGY WAS ON THE DEVILBISS BIPAP, # (B)(6). THE BIPAP, (B)(6) WAS TESTED BY THE EQUIPMENT PROVIDER ON THE SAME DAY OF THE COMPLAINT WITH NO ISSUES FOUND. THE AIR PRESSURE WAS NORMAL. THE DEVICE WAS THEN SENT TO MEDICAL TECTRONICS, INC. IN HOUSTON, TEXAS FOR A PERFORMANCE EVALUATION. THE TECHNICIAN RAN FULL DIAGNOSTICS ON UNIT UPON INTAKE AND INDICATED NO ISSUES. THE UNIT WAS STATED AS OPERATING ACCORDING TO MANUFACTURER'S SPECIFICATIONS. MULTIPLE ATTEMPTS HAVE BEEN MADE TO HAVE THE DEVICE RETURNED TO REACT HEALTH FOR EVALUATION AND THE DEVICE HAS NOT BEEN RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807156&pc=NPT | 2025587-2025-02505 | 2024-10-29 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807222&pc=NPT | 2025587-2025-02508 | 2024-05-27 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY AT HOME. THE CAUSE OF DEATH WAS REPORTED TO BE "ANOTHER DISEASE." Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807411&pc=DSQ | 2916596-2025-02383 | 2025-01-12 05:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO UNKNOWN CAUSE. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807765&pc=NPT | 2025587-2025-02513 | 2025-03-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX DCS | Material Invagination; Device Difficult to Setup or Prepare; Adverse Event Without Identified Device or Use Problem; Insufficient Information; Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A ARCH-ASCENDING-AORTA-ANNULUS THAT EXHIBITED A SHORT AND STEEP CURVE WITH A 60-DEGREE ANGLE, MILD ANNULUS CALCIFICATION, AND AN ANNULUS PERIMETER OF 86 MM, ACCESS WAS OBTAINED ON THE RIGHT SIDE. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 22 MM NON-MEDTRONIC BALLOON. UPON FLUOROSCOPIC INSPECTION OF THE FIRST VALVE LOAD, A FRAME NODE OVERLAP THAT WAS ON THE BORDERLINE OF INVOLVING NODE 4 WAS NOTED. THE DECISION WAS MADE TO MOVE FORWARD WITH THE LOADED VALVE. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, ADVANCED TO THE ANNULUS, AND THE VALVE WAS DEPLOYED. AT 80% DEPLOYMENT, AND INFOLD WAS OBSERVED IN THE VALVE FRAME. THE VALVE WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. THE INFOLD COULD BE VISUALIZED OUTSIDE THE PATIENT IN THE UNSTERILE FIELD. SUBSEQUENTLY, A NEW VALVE WAS LOADED INTO A NEW DCS AND INSERTED INTO THE PATIENT. THROUGHOUT THE PROCEDURE, THE SYSTEM TENDED TO DRIFT TOWARDS THE GREATER CURVATURE, RESULTING IN DEEP PLACEMENT AT THE NON-CORONARY CUSP (NCC) AND SHALLOW PLACEMENT ON THE LEFT CORONARY CUSP (LCC) DURING DEPLOYMENT. THE VALVE WAS RECAPTURED MULTIPLE TIMES AFTER THE AORTIC VALVE WAS CROSSED DUE TO THESE POSITIONING DIFFICULT ES. DURING REPOSITIONING, THE SYSTEM WOULD OCCASIONALLY ATTACH TO THE GREATER CURVATURE OF THE ASCENDING AORTA NEAR THE PADDLE ATTACHMENT AREA. EVENTUALLY, FOLLOWING AN AXIAL ADJUSTMENT IN THE LEFT ANTERIOR OBLIQUE ALIGNMENT, THE VALVE WAS FULLY RELEASED. A LOCALIZED DISSECTION OF THE ASCENDING AORTA WAS NOTED IN THE FINAL AORTOGRAM. NO IMPACT ON THE PATIENT'S VITALS WAS OBSERVED. AN ECHOCARDIOGRAM AND CONTRAST IMAGING REVEALED NO DISSECTION EXPANSION. THE PROCEDURE WAS CONCLUDED WITH A PLAN FOR ONGOING OBSERVATION OF THE DISSECTION. THE PATIENT PASSED AWAY 6 DAYS LATER. THE CAUSE OF DEATH WAS AORTIC DISSECTION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE WAS RECAPTURED ONCE DUE TO THE INFOLD AND SUBSEQUENTLY WITHDRAWN FROM THE PATIENT. THE SECOND VALVE WAS RECAPTURED THREE TIMES PRIOR TO FULL RELEASE. A PROCEDURAL DELAY OCCURRED IN RESULT OF THE INFOLD. PER THE PHYSICIAN, IT WAS POSSIBLE THAT INFLOW STRUT OF THE SECOND VALVE MADE CONTACT WITH THE ASCENDING AORTA AT THE TIME OF RECAPTURE, CAUSING THE AORTIC DISSECTION, OR THE PADDLE ATTACHMENT MADE CONTACT WITH THE ASCENDING AORTIC WALL, CAUSING THE AORTIC DISSECTION. IT WAS FURTHER REPORTED THAT THE VALVE AND DCS COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A ARCH-ASCENDING-AORTA-ANNULUS T HAT EXHIBITED A SHORT AND STEEP CURVE WITH A 60-DEGREE ANGLE, MILD ANNULUS CALCIFICATION, AND AN ANNULUS PERIMETER OF 86 MM, ACCESS WAS OBTAINED ON THE RIGHT SIDE. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 22 MM NON-MEDTRONIC BALLOON. UPON FLUOROSCOPIC INSPECTION OF THE FIRST VALVE LOAD, A FRAME NODE OVERLAP THAT WAS ON THE BORDERLINE OF INVOLVING NODE 4 WAS NOTED. THE DECISION WAS MADE TO MOVE FORWARD WITH THE LOADED VALVE. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, ADVANCED TO THE ANNULUS, AND THE VALVE WAS DEPLOYED. AT 80% DEPLOYMENT, AND INFOLD WAS OBSERVED IN THE VALVE FRAME. THE VALVE WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. THE INFOLD COULD BE VISUALIZED OUTSIDE THE PATIENT IN THE UNSTERILE FIELD. SUBSEQUENTLY, A NEW VALVE WAS LOADED INTO A NEW DCS AND INSERTED INTO THE PATIENT. THROUGHOUT THE PROCEDURE, THE SYSTEM TENDED TO DRIFT TOWARDS THE GREATER CURVATURE, RESULTING IN DEEP PLACEMENT AT THE NON-CORONARY CUSP (NCC) AND SHALLOW PLACEMENT ON THE LEFT CORONARY CUSP (LCC) DURING DEPLOYMENT. THE VALVE WAS RECAPTURED MULTIPLE TIMES AFTER THE AORTIC VALVE WAS CROSSED DUE TO THESE POSITIONING DIFFICULTI ES. DURING REPOSITIONING, THE SYSTEM WOULD OCCASIONALLY ATTACH TO THE GREATER CURVATURE OF THE ASCENDING AORTA NEAR THE PADDLE ATTACHMENT AREA. EVENTUALLY, FOLLOWING AN AXIAL ADJUSTMENT IN THE LEFT ANTERIOR OBLIQUE ALIGNMENT, THE VALVE WAS FULLY RELEASED. A LOCALIZED DISSECTION OF THE ASCENDING AORTA WAS NOTED IN THE FINAL AORTOGRAM. NO IMPACT ON THE PATIENT'S VITALS WAS OBSERVED. AN ECHOCARDIOGRAM AND CONTRAST IMAGING REVEALED NO DISSECTION EXPANSION. THE PROCEDURE WAS CONCLUDED WITH A PLAN FOR ONGOING OBSERVATION OF THE DISSECTION. THE PATIENT PASSED AWAY 6 DAYS LATER. THE CAUSE OF DEATH WAS AORTIC DISSECTION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE WAS RECAPTURED ONCE DUE TO THE INFOLD AND SUBSEQUENTLY WITHDRAWN FROM THE PATIENT. THE SECOND VALVE WAS RECAPTURED THREE TIMES PRIOR TO FULL RELEASE. A PROCEDURAL DELAY OCCURRED IN RESULT OF THE INFOLD. PER THE PHYSICIAN, IT WAS POSSIBLE THAT INFLOW STRUT OF THE SECOND VALVE MADE CONTACT WITH THE ASCENDING AORTA AT THE TIME OF RECAPTURE, CAUSING THE AORTIC DISSECTION, OR THE PADDLE ATTACHMENT MADE CONTACT WITH THE ASCENDING AORTIC WALL, CAUSING THE AORTIC DISSECTION. IT WAS FURTHER REPORTED THAT THE VALVE AND DCS COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT PER THE PHYSICIAN, THE FIRST DCS DID NOT CONTRIBUTE TO THE AORTIC DISSECTION. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED SHOWED THERE IS NO INFORMATION TO REASONABLY SUGGEST THE DEVICE IN THIS REPORT CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE; LOT #, AND FULL UDI # H4. DEVICE MFG DATE H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: EVOLUTFX-34, SERIAL/LOT #: UNKNOWN SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807766&pc=FTL | 9615742-2025-00324 |  | Death | SOFRADIM PRODUCTION SAS | 2025-04-09 04:00 | FTL | PROGRIP | Adverse Event Without Identified Device or Use Problem | Pain; Burning Sensation; Anxiety; Discomfort; Depression; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K140941 |  | 1 | Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS.Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS. IT WAS NOTED THAT THE PATIENT DIED IN (B)(6) 2025.Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN. THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG WITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS.Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS. THE PATIENT WAS REPORTED TO HAVE DIED BY SUICIDE IN (B)(6) 2025. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: PP1208DL PROGRIP LT PPL/PLA 12X8CM (LOT#SYB2020X) MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: E1 (PREFIX, LAST NAME, FIRST NAME REMOVED), G3 CORRECTION: E3 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B2 (DEATH, DATE OF DEATH), B5, G3, H1, H6 CORRECTION: D4 (UNIQUE IDENTIFIER (UDI) #) NEW INFORMATION HAS BEEN RECEIVED PERTAINING TO THE EVENT. THIS EVENT HAS BEEN REASSESSED AND THE REPORTABILITY HAS BEEN DETERMINED TO BE A DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807783&pc=DSQ | 2916596-2025-02271 | 2025-03-18 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Disconnection; Adverse Event Without Identified Device or Use Problem | Cyanosis; Convulsion/Seizure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD SEIZURES AND THEN CYANOSIS AT HOME, AND THE PATIENT'S FAMILY SENT THEM TO A HOSPITAL. THE PATIENT PASSED AWAY. THE SURGEON DID NOT THINK THAT THE EVENT WAS RELATED TO THE PRODUCT. LOG FILE REVIEW INDICATED THAT THE PATIENT WAS CONNECTED TO THE MOBILE POWER UNIT (MPU) ON (B)(6) 2025 AT 16:30:46. AT 16:47:03 THERE WAS A NO EXTERNAL POWER ALARM FLAG ACTIVE. THE DATA INDICATED THAT THE SYSTEM WAS BEING SUPPORTED BY THE EMERGENCY BACKUP BATTERY (EBB). PUMP SPEED BEING MAINTAINED AT 4800 REVOLUTION PER MINUTES (RPMS), THE PATIENT SET SPEED. SILENCE ALARM BUTTON PRESSED EVENTS WERE RECORDED DURING THIS TIME. AT 16:50:18, A LOW FLOW ALARM FLAG BECOMES ACTIVE AS THE RECORDED CALCULATED AVERAGE FLOW REDUCES TO 1.9 LITER PER MINUTES (LPMS). SILENCE ALARM BUTTON PRESSED EVENTS WERE RECORDED DURING THIS TIME. AT 17:10:57, A 14 VOLT BATTERY WAS CONNECTED TO THE WHITE POWER LEAD. LOW FLOW ALARM FLAG REMAINED ACTIVE. 1.1 LPMS RECORDED. AT 17:11:20, A 14 VOLT BATTERY WAS CONNECTED TO THE BLACK POWER LEAD. LOW FLOW ALARM FLAG REMAINS ACTIVE. 1.0 LPMS WAS RECORDED. SILENCE ALARM BUTTON PRESSED EVENTS WERE RECORDED DURING THIS TIME. AT 17:28:23, A 14 VOLT BATTERY WAS DISCONNECTED TO THE WHITE POWER LEAD. LOW FLOW ALARM FLAG REMAINED ACTIVE. 1.0 LPMS WERE RECORDED. AT 17:28:36, A 14 VOLT BATTERY WAS DISCONNECTED TO THE BLACK POWER LEAD. LOW FLOW ALARM FLAG REMAINS ACTIVE. 1.0 LPMS WAS RECORDED. THE DATA INDICATED THAT THE SYSTEM IS BEING SUPPORTED BY THE EBB. PUMP SPEED BEING MAINTAINED AT 4800 RPMS, THE PATIENT SET SPEED. SILENCE ALARM BUTTON PRESSED EVENTS WERE RECORDED DURING THIS TIME. THE DATA INDICATED THAT THE SYSTEM WAS ABLE TO MAINTAIN PUMP SPEED ON THE EBB UNTIL A DRIVELINE DISCONNECT ALARM FLAG WAS RECORDED AT 17:46:00. A SHUTDOWN_SEQUENCE_STARTED EVENT WAS RECORDED AT (B)(6) 2025 AT 18:39:46. THE DEVICE OPERATED AS EXPECTED. THE MPU HAD A V-LOCK CONNECTOR ON THE AC POWER CORD.Event Description: IT WAS FURTHER REPORTED THAT ACCORDING TO PATIENT'S RECORD, BEFORE THE PATIENT HAD SEIZURE ATTACK, ALL SYSTEM WORKED NORMALLY. THE SYSTEM STARTED TO HAPPEN ALARM AFTER THE PATIENT LOSS OF VITAL SIGNS. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED THE LOW FLOW ALARMS AND PUMP STOP. A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION. REVIEW OF THE SUBMITTED CONTROLLER EVENT LOG FILE CAPTURED A PERSISTENT LOW FLOW HAZARD ALARM BEGINNING AT (B)(6) 2025. THE LOW FLOW HAZARD ALARM REMAINED ACTIVE UNTIL THE DRIVELINE WAS DISCONNECTED FROM THE CONTROLLER ON (B)(6) 2025. NO OTHER NOTABLE EVENTS OR ALARM WERE CAPTURED. ALTHOUGH THE CAUSE OF THE PATIENT¿S OUTCOME WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. IT WAS COMMUNICATED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING NEUROLOGICAL EVENTS (NOT STROKE RELATED) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT" (UNDER "ONGOING PATIENT ASSESSMENT AND CARE"), INCLUDES A LIST OF HEARTMATE 3 PATIENT ASSESSMENTS AND LISTS NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 2 OF THE IFU AND SECTION 2 OF THE PATIENT HANDBOOK INSTRUCT THE USER TO CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK ALSO EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. ADDITIONALLY, SECTION 2 OF THE IFU, UNDER "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807821&pc=MNS | 2518422-2025-103858 | 2025-04-07 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807861&pc=OZD | 1220648-2025-27337 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-09 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK UNDERWENT PERCUTANEOUS CORONARY INTERVENTION AND WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT WAS TRANSFERRED TO A RECEIVING FACILITY FOR ESCALATION IN CARE TO AN IMPELLA 5.5 DEVICE FOR CONTINUED MCS. THE SAME DAY, POST IMPLANT OF THE IMPELLA 5.5, IT WAS INDICATED THAT THE PATIENT EXPERIENCED BLEEDING, AND ANTICOAGULATION WAS WITHHELD (NOTES REFLECTED "WILL START SYSTEMIC ANTICOAGULATION TOMORROW PENDING BLEEDING."). THE FOLLOWING DAY HEPARIN STILL WAS WITHHELD. THE PATIENT HAD CLEAR TO YELLOW URINE SINCE BEING PLACED ON THE IMPELLA 5.5, AND REMAINED ON THREE PRESSORS WITH PLANS TO WEAN OFF SEDATION IN HOPES TO EXTUBATE. THE PATIENT CONTINUED ON IMPELLA 5.5 MCS. THREE DAYS LATER, IT WAS REPORTED THAT THE PATIENT HAD BEEN GOING IN AND OUT OF VENTRICULAR TACHYCARDIA/TORSADES DE POINTES INTO THE FOLLOWING DAY. THE PATIENT WAS DEFIBRILLATED. TWO DAYS LATER, THE PATIENT WAS PLACED ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AND TEMPORARY PACING WIRES. THE PATIENT RECEIVED FOUR UNITS OF PACKED RED BLOOD CELLS, THREE UNITS OF FRESH FROZEN PLASMA, AND ALBUMIN WERE NOTED IN THE LAST 24 HOURS. AFTER APPROXIMATELY THREE DAYS ON VA ECMO, THE PATIENT WAS DECANNULATED AND WAS NOTED TO BE DOING WELL WITH A PLAN TO EVENTUALLY START WEANING THE IMPELLA 5.5 DEVICE. THE PATIENT CONTINUED ON IMPELLA MCS, AND APPROXIMATELY 14 DAYS LATER, THE PATIENT EXPERIENCED A SUDDEN MENTAL STATUS CHANGE. A HEAD COMPUTED TOMOGRAPHY SCAN CONFIRMED A "LARGE" CYSTIC LESION WITH HEMORRHAGE IN THE RIGHT FRONTAL REGION WITH LARGE VOLUME INTRAVENTRICULAR HEMORRHAGE IN LATERAL VENTRICLES. NEUROLOGICAL EXAM WAS CONSISTENT WITH STATED FINDINGS. THERE WERE NOTABLE ELECTROCARDIOGRAM CHANGES THE MORNING OF THIS DAY, SUDDENLY, WIDENING QRS, AND THE PATIENT WOULD EXPIRE THIS DAY. CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807923&pc=NPT | 2025587-2025-02514 | 2025-03-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A ARCH-ASCENDING-AORTA-ANNULUS T HAT EXHIBITED A SHORT AND STEEP CURVE WITH A 60-DEGREE ANGLE, MILD ANNULUS CALCIFICATION, AND AN ANNULUS PERIMETER OF 86 MM, ACCESS WAS OBTAINED ON THE RIGHT SIDE. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 22 MM NON-MEDTRONIC BALLOON. UPON FLUOROSCOPIC INSPECTION OF THE FIRST VALVE LOAD, A FRAME NODE OVERLAP THAT WAS ON THE BORDERLINE OF INVOLVING NODE 4 WAS NOTED. THE DECISION WAS MADE TO MOVE FORWARD WITH THE LOADED VALVE. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, ADVANCED TO THE ANNULUS, AND THE VALVE WAS DEPLOYED. AT 80% DEPLOYMENT, AND INFOLD WAS OBSERVED IN THE VALVE FRAME. THE VALVE WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. THE INFOLD COULD BE VISUALIZED OUTSIDE THE PATIENT IN THE UNSTERILE FIELD. SUBSEQUENTLY, A NEW VALVE WAS LOADED INTO A NEW DCS AND INSERTED INTO THE PATIENT. THROUGHOUT THE PROCEDURE, THE SYSTEM TENDED TO DRIFT TOWARDS THE GREATER CURVATURE, RESULTING IN DEEP PLACEMENT AT THE NON-CORONARY CUSP (NCC) AND SHALLOW PLACEMENT ON THE LEFT CORONARY CUSP (LCC) DURING DEPLOYMENT. THE VALVE WAS RECAPTURED MULTIPLE TIMES AFTER THE AORTIC VALVE WAS CROSSED DUE TO THESE POSITIONING DIFFICULT ES. DURING REPOSITIONING, THE SYSTEM WOULD OCCASIONALLY ATTACH TO THE GREATER CURVATURE OF THE ASCENDING AORTA NEAR THE PADDLE ATTACHMENT AREA. EVENTUALLY, FOLLOWING AN AXIAL ADJUSTMENT IN THE LEFT ANTERIOR OBLIQUE ALIGNMENT, THE VALVE WAS FULLY RELEASED. A LOCALIZED DISSECTION OF THE ASCENDING AORTA WAS NOTED IN THE FINAL AORTOGRAM. NO IMPACT ON THE PATIENT'S VITALS WAS OBSERVED. AN ECHOCARDIOGRAM AND CONTRAST IMAGING REVEALED NO DISSECTION EXPANSION. THE PROCEDURE WAS CONCLUDED WITH A PLAN FOR ONGOING OBSERVATION OF THE DISSECTION. THE PATIENT PASSED AWAY 6 DAYS LATER. THE CAUSE OF DEATH WAS AORTIC DISSECTION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE WAS RECAPTURED ONCE DUE TO THE INFOLD AND SUBSEQUENTLY WITHDRAWN FROM THE PATIENT. THE SECOND VALVE WAS RECAPTURED THREE TIMES PRIOR TO FULL RELEASE. A PROCEDURAL DELAY OCCURRED IN RESULT OF THE INFOLD. PER THE PHYSICIAN, IT WAS POSSIBLE THAT INFLOW STRUT OF THE SECOND VALVE MADE CONTACT WITH THE ASCENDING AORTA AT THE TIME OF RECAPTURE, CAUSING THE AORTIC DISSECTION, OR THE PADDLE ATTACHMENT MADE CONTACT WITH THE ASCENDING AORTIC WALL, CAUSING THE AORTIC DISSECTION. IT WAS FURTHER REPORTED THAT THE VALVE AND DCS COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VOLUTFX-34, SERIAL/LOT #: UNKNOWN. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4 H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807940&pc=FTL | 9615742-2025-00325 | 2024-07-03 04:00 | Death | SOFRADIM PRODUCTION SAS | 2025-04-09 04:00 | FTL | PROGRIP | Adverse Event Without Identified Device or Use Problem | Pain; Burning Sensation; Anxiety; Discomfort; Depression; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K140941 |  | 1 | Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN. THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG WITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS.Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024, AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS.Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG WITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS. IT WAS NOTED THAT THE PATIENT DIED ON (B)(6) 2025.Event Description: THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS. THE PATIENT WAS REPORTED TO HAVE DIED BY SUICIDE IN (B)(6) 2025. Manufacturer Narrative: CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B2 (DEATH, DATE OF DEATH), B5, G3, H1, H6 CORRECTION: D4 (UNIQUE IDENTIFIER (UDI)#) NEW INFORMATION HAS BEEN RECEIVED PERTAINING TO THE EVENT. THIS EVENT HAS BEEN REASSESSED AND THE REPORTABILITY HAS BEEN DETERMINED TO BE A DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: E1 (FIRST NAME REMOVED). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B3, B5, G3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: PP1208DR, PP1208DR PROGRIP RT PPL/PLA 12X8CM (LOT#SYA1851X). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807984&pc=DSQ | 2916596-2025-02171 | 2024-08-19 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21807995&pc=BRY | 2016493-2025-72076 | 2025-03-13 04:00 | Death | CAREFUSION 303, INC. | 2025-04-09 04:00 | BRY | BD PYXIS¿ MEDBANK TOWER | Improper or Incorrect Procedure or Method; Application Program Problem | Insufficient Information | EXEMPT |  | 1 | Event Description: IT WAS REPORTED WHEN USING THE BD PYXIS¿ MEDBANK TOWER THE USER LOGGED IN TO OBTAIN NARCAN DURING AN "EMERGENCY SITUATION" IN A REHABILITATION FACILITY. HOWEVER, THE MEDICATION REPORTEDLY DID NOT SHOW UNDER THE PATIENT PROFILE. THE DELAY REPORTEDLY RESULTED IN "NOT BEING ABLE TO SAVE THE RESIDENT." THE CUSTOMER REPORTED THAT THE DEVICE DID HAVE 2 NARCAN VIALS. PRELIMINARY INVESTIGATION REVEALED THAT THE ISSUE WAS DUE TO USE ERROR. IT WAS REPORTED THAT SINCE THE MEDICATION WAS NOT ON THE PATIENT PROFILE, THE USER NEEDED TO CLICK THE ADD ITEM BUTTON TO ADD NARCAN ONTO THE PATIENT PROFILE. THE DEVICE KEYSTROKE LOGS SHOW THAT USER DID NOT CLICK THE ADD ITEM BUTTON AT THE TIME OF THE EVENT. IT WAS CONFIRMED THAT THE USER HAS DEVICE PERMISSION FOR GENERAL OVERRIDE, WHICH WOULD ALLOW THEM TO USE THE ADD ITEM OVERRIDE FEATURE AND ACCESS THE MEDICATION. MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN ADDITIONAL INFORMATION, NO RESPONSE HAS BEEN RECEIVED TO DATE.Event Description: IT WAS REPORTED WHEN USING THE BD PYXIS¿ MEDBANK TOWER THE USER LOGGED IN TO OBTAIN NARCAN DURING AN "EMERGENCY SITUATION" IN A REHABILITATION FACILITY. HOWEVER, THE MEDICATION REPORTEDLY DID NOT SHOW UNDER THE PATIENT PROFILE. THE DELAY REPORTEDLY RESULTED IN "NOT BEING ABLE TO SAVE THE RESIDENT." THE CUSTOMER REPORTED THAT THE DEVICE DID HAVE 2 NARCAN VIALS. PRELIMINARY INVESTIGATION REVEALED THAT THE ISSUE WAS DUE TO USE ERROR. IT WAS REPORTED THAT SINCE THE MEDICATION WAS NOT ON THE PATIENT PROFILE, THE USER NEEDED TO CLICK THE ADD ITEM BUTTON TO ADD NARCAN ONTO THE PATIENT PROFILE. THE DEVICE KEYSTROKE LOGS SHOW THAT USER DID NOT CLICK THE ADD ITEM BUTTON AT THE TIME OF THE EVENT. IT WAS CONFIRMED THAT THE USER HAS DEVICE PERMISSION FOR GENERAL OVERRIDE, WHICH WOULD ALLOW THEM TO USE THE ADD ITEM OVERRIDE FEATURE AND ACCESS THE MEDICATION. MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN ADDITIONAL INFORMATION, NO RESPONSE HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: COMPLAINT HISTORY REVIEW: A REVIEW OF THE COMPLAINT HISTORY FOR SN (B)(6) WAS PERFORMED IN SALESFORCE WHICH DID NOT LOCATE SIMILAR COMPLAINTS WITH THE SAME FAILURE MODE FOR THIS SERIAL NUMBER. DEVICE HISTORY REVIEW: A REVIEW OF THE DEVICE HISTORY RECORD FOR SN (B)(6) WAS PERFORMED FROM THE DATE OF MANUFACTURE, 30-NOV-2022 AND CONFIRMED THAT THIS DEVICE WAS NOT PREVIOUSLY RETURNED FOR SERVICING AND THERE WERE NO PRODUCTION FAILURES WHICH CORRELATES TO THE CUSTOMER REPORTED ISSUE. INVESTIGATION SUMMARY: UPON INVESTIGATING THE DEVICE INVOLVED IN THIS INCIDENT, IT WAS DETERMINED THAT THE USER DID NOT CLICK THE ADD ITEMS BUTTON TO ADD NARCAN TO THE PATIENT PROFILE, WHICH CAUSED THE MEDICATION TO BE UNAVAILABLE FOR DISPENSING AT THAT TIME. A TECHNICAL SUPPORT SPECIALIST ASSISTED THE USER IN CLICKING THE ADD ITEMS BUTTON ADDING NARCAN TO THE PATIENT PROFILE WITH GENERAL OVERRIDE PERMISSIONS, THE USER HAD ACCESS TO THE MEDICATION THROUGH THE ADD ITEMS FEATURE. THE SYSTEM FUNCTIONED AS INTENDED AFTER THE TECHNICAL SUPPORT SPECIALIST TROUBLESHOT THE DEVICE. Manufacturer Narrative: COMPLAINT HISTORY REVIEW: A REVIEW OF THE COMPLAINT HISTORY FOR SN (B)(6) WAS PERFORMED IN SALESFORCE WHICH DID NOT LOCATE SIMILAR COMPLAINTS WITH THE SAME FAILURE MODE FOR THIS SERIAL NUMBER. DEVICE HISTORY REVIEW: A REVIEW OF THE DEVICE HISTORY RECORD FOR SN (B)(6) WAS PERFORMED FROM THE DATE OF MANUFACTURE, 30-NOV-2022 AND CONFIRMED THAT THIS DEVICE WAS NOT PREVIOUSLY RETURNED FOR SERVICING AND THERE WERE NO PRODUCTION FAILURES WHICH CORRELATES TO THE CUSTOMER REPORTED ISSUE. INVESTIGATION SUMMARY: UPON INVESTIGATING THE DEVICE INVOLVED IN THIS INCIDENT BY REVIEWING LOG FILES, IT WAS DETERMINED THAT USER WAS NOT ABLE TO ACCESS THE NARCAN MEDICATION, WHICH WAS AVAILABLE FOR DISPENSING IN THE DEVICE, BECAUSE THE USER DID NOT ADD THIS DRUG TO THE PATIENT PROFILE. THIS IS DONE BY CLICKING THE ADD ITEMS BUTTON ONTO THE PROFILE. TSC CONFIRMED THE USER HAD GENERAL ACCESS, AND THE ITEM WAS GENERAL OVERRIDE. THIS INCIDENT IS DUE TO USE ERROR. IT IS NOT A MALFUNCTION OF THE DEVICE. THE DEVICE FUNCTIONED AS INTENDED DURING THE INCIDENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808124&pc=MKJ | 1220908-2025-01204 | 2025-03-18 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-04-09 04:00 | MKJ | X SERIES | Failure to Charge | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE WAS UNABLE TO CHARGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE DATA LOGS. HOWEVER, THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING INCLUDING BENCH TESTING WITHOUT DUPLICATING THE REPORT. THE DEVICE COULDN'T CONSISTENTLY RECOGNIZED A VALID CABLE ID, WHICH PREVENTED ENERGY DELIVERY AND CPR FEEDBACK. THIS SUGGESTS A POOR CONNECTION BETWEEN THE MULTIFUNCTION CABLE (MFC) AND THE DEVICE. THE MFC, MFC RECEPTACLE, AND CPRD ADAPTOR WERE REPLACED AS A PRECAUTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808250&pc=MNS | 2518422-2025-103891 | 2025-04-07 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808313&pc=OEV | 1220828-2025-00007 | 2025-03-07 05:00 | Death | AMBU A/S | 2025-04-09 04:00 | OEV | SPUR II PEDIATRIC RESUSCITATOR | Component Missing | Low Oxygen Saturation |  |  | 1 | Event Description: AS PER THE CUSTOMER, DURING INTUBATION IN AN EMERGENCY ROOM, PEDIATRIC PATIENT WAS NOT GETTING CO2. PEDIATRIC BAG WAS USED ON A PATIENT BUT THE INLET VALVE AT THE BASE WAS MISSING AND THEY COULD NOT PRODUCE ANY CO2. LATER WHEN INSPECTING THE BAG, THEY NOTICED THAT THE INLET WAS FOUND INSIDE BY THE MASK. Manufacturer Narrative: WE VERIFIED THE REPORTED FAILURE OF A MISSING PATIENT VALVE DISC INSIDE PATIENT VALVE HOUSING FROM A PHOTO PROVIDED BY THE CUSTOMER. THE SAMPLE WAS NOT RETURNED FOR INVESTIGATION. DURING MANUFACTURING, PATIENT VALVE DISC IS INSTALLED INTO PATIENT VALVE HOUSING AND A FUNCTION TEST PERFORMED ON 100% OF ASSEMBLED PATIENT VALVES. THEN THE VALVE IS INSTALLED ONTO SPUR II RESUSCITATOR AND ANOTHER FUNCTION TEST IS PERFORMED ON ALL DEVICES. MISSING PATIENT VALVE DISC WOULD BE REVEALED BY BOTH FUNCTION TESTS, HOWEVER PRODUCTION RECORDS DO NOT INDICATE ANY ABNORMALITY. THEREFORE, THE DEVICE SHOULD BE WITHIN SPECIFICATIONS. NEXT, PEEP VALVE, FACE MASK AND MANOMETER WERE PREATTACHED, WHICH WOULD NOT IMPACT INSTALLED PATIENT VALVE DISC. NO ABNORMALITIES WERE FOUND IN THE PRODUCTION RECORDS, THUS IT IS CONCLUDED THAT DURING MANUFACTURING, THE DISC WAS INSTALLED CORRECTLY. THE CUSTOMER REPORTED THAT THE PATIENT VALVE DISC WAS FOUND INSIDE THE BAG NEXT TO A MASK. POSSIBLE CAUSES FOR THIS ISSUE TO OCCUR ARE: DETACHMENT OF THE EXPIRATORY CONNECTOR AND PATIENT VALVE DISC DURING TRANSPORTATION; OR DISASSEMBLING OF THE EXPIRATORY CONNECTOR AND PATIENT VALVE DISC BY THE USER, WITH SUBSEQUENT WRONG ASSEMBLY. AS PER THE TRANSPORTATION TESTING, THE EXPIRATORY CONNECTOR AND PATIENT VALVE DISC WOULD NOT DETACH UNDER NORMAL TRANSPORTATION CONDITIONS. SEVERAL ATTEMPTS WERE MADE TO RETRIEVE MORE INFORMATION FROM THE CUSTOMER, HOWEVER THEY WERE UNSUCCESSFUL. DUE TO LIMITED INFORMATION, THE ROOT CAUSE REMAINS UNKNOWN. THE REPORTED FAILURE IS DETECTABLE DURING A PRESCRIBED PRE-CHECK. AS PER THE INSTRUCTIONS FOR USE: 'ALWAYS INSPECT THE RESUSCITATOR AND PERFORM FUNCTIONAL TEST AFTER UNPACKING, CLEANING, ASSEMBLY AND PRIOR TO USE.' THIS ISSUE IS CONSIDERED TO BE AN ISOLATED CASE, SINCE THIS WAS THE FIRST TIME THAT A DETACHMENT OF PATIENT VALVE DISC WAS REPORTED IN THE PAST YEAR. PRODUCTION HAS BEEN INFORMED ABOUT THE FAILURE AND WE WILL KEEP MONITORING THE MARKET FOR SIMILAR INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808341&pc=BZD | 2518422-2025-042872 | 2022-10-07 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAD PASSED AWAY. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAD PASSED AWAY. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. SECTION H6 WAS UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808365&pc=DSQ | 3007042319-2025-00743 | 2025-01-23 05:00 | Death | HEARTWARE, INC. | 2025-04-09 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Multiple Organ Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING PEDIATRIC HEART TRANSPLANTATION IN PATIENTS WITH A VENTRICULAR ASSIST DEVICE (VAD). THE AUTHORS DESCRIBED PATIENT DEATHS WHILE ON THE TRANSPLANT WAITING LIST. CAUSES OF DEATH INCLUDED CEREBRAL BLEEDING AND MULTIORGAN FAILURE. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS FEMALE/8 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE IMPACT OF PRE-TRANSPLANT VENTRICULAR ASSIST DEVICE SUPPORT IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE ON THE OUTCOMES OF HEART TRANSPLANTATION¿ ¿A SINGLE CENTER EXPERIENCE¿. FRONTIERS IN CARDIOVASCULAR MEDICINE. 2025. 12:1515218. DOI: 10.3389/FCVM.2025.1515218 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808618&pc=DSQ | 2916596-2025-02174 | 2025-03-14 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE. THE PATIENT HAD RIGHT VENTRICULAR FAILURE PRIOR TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT AND LEFT THE OPERATING ROOM WITH EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) RIGHT VENTRICULAR SUPPORT. THE PATIENT'S OUTCOME WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED UNTIL THE END AND DID NOT SHOW ANY MALFUNCTIONS INCLUDING ALARMS AND PARAMETERS.Event Description: ALTHOUGH THE LVAD AND RVAD ECMO WERE INITIATED SUCCESSFULLY THE PATIENT'S CIRCULATION DID NOT RECOVER AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PATIENT OUTCOME WAS NOT CONSIDERED DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE PUMP WILL NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808701&pc=OZD | 1220648-2025-27340 | 2025-03-24 04:00 | Death | ABIOMED, INC. | 2025-04-09 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: A LITERATURE REPORT TITLED "FIVE CASES OF PREEMPTIVE IMPELLA INSERTION BEFORE WEANING FROM CARDIOPULMONARY BYPASS IN HIGH-RISK PATIENTS WITH CARDIOGENIC SHOCK AFTER OPEN HEART SURGERY" WAS RECEIVED BY ABIOMED WHICH STATED THAT "A PATIENT IN JAPAN UNDERWENT TOTAL ARCH REPLACEMENT, CABG, AND INSERTION OF AN IMPELLA 5.5 (THROUGH THE ASCENDING AORTA) FOR DISTAL AORTIC ARCH ANEURYSM AND ASYMPTOMATIC MYOCARDIAL ISCHEMIA. AFTER SURGERY, THE IMPELLA FLOW RATE GRADUALLY DECREASED, AND RE-THORACOTOMY WAS PERFORMED FOR CARDIAC TAMPONADE THE DAY AFTER SURGERY. HOWEVER, THE HEMODYNAMICS REMAINED UNSTABLE, AND THE PATIENT EXPIRED ON THE THIRD DAY AFTER SURGERY." THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. A.1, A.4 AND A.5 ARE UNKNOWN. D.4 CATALOG NUMBER, SERIAL NUMBER, LOT NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN. D.6A AND D.6B IMPLANT AND EXPLANT DATES ARE UNKNOWN. E.1 NAME PREFIX/TITLE, FIRST NAME AND LAST NAME ARE UNKNOWN. H.4 DEVICE MANUFACTURE DATE IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808739&pc=OJX | 2649622-2025-09593 | 2023-04-15 04:00 | Death | MPRI | 2025-04-09 04:00 | OJX | ATTAIN PERFORMA MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED TWO DAYS POST PROCEDURE AND UPGRADE TO CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED: (B)(6) 2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808742&pc=LWS | 2649622-2025-09594 | 2023-04-15 04:00 | Death | MPRI | 2025-04-09 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED TWO DAYS POST PROCEDURE AND UPGRADE TO CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED: (B)(6) 2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808789&pc=MNS | 2518422-2025-103895 | 2025-04-07 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21808868&pc=DSQ | 2916596-2025-01884 | 2022-05-25 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fever; Pain; Dementia; Heart Failure/Congestive Heart Failure; Skin Infection; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2022 THROUGH (B)(6) 2022 FOR DRIVELINE INFECTION. THE PATIENT WAS ADMITTED AGAIN FOR DRIVELINE INFECTION ON (B)(6) 2024. THE PATIENT HAD FEVER AND PAIN AROUND THE DRIVELINE. CULTURE SHOWED METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA). THE PATIENT WAS TREATED WITH INTRAVENOUS VANCOMYCIN AND TRANSITIONED TO ORAL LONG-TERM MINOCYCLINE, DOXYCYCLINE, AND BACTRIM (ULTIMATELY DRUG RESISTANT). THE PATIENT WAS FOLLOWED BY THE INFECTIOUS DISEASE (ID) DEPARTMENT OVER THE ENTIRE COURSE. THE INFECTION NEVER RESOLVED. ADDITIONALLY, THE PATIENT WAS ON HOSPICE AND HAD AN EXPECTED DEATH ON (B)(6) 2025. THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE CAUSE OF DEATH WAS DUE TO HOSPICE. THE PUMP WOULD NOT BE EXPLANTED. THE REASON FOR HOSPICE WAS END STAGE HEART FAILURE, DRIVELINE INFECTION, AND DEMENTIA. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED DRIVELINE INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHERMORE, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809405&pc=LWS | 3004209178-2025-06180 | 2025-03-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-09 04:00 | LWS | EVERA MRI XT DR SURESCAN | Failure to Deliver Shock/Stimulation | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT¿S SPOUSE INDICATED THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DID NOT DELIVER A SHOCK AND QUESTIONED IF THE PATIENT¿S HEART WAS ¿JUST TOO FAR GONE¿. THE PATIENT¿S SPOUSE NOTED THAT EXTERNAL SHOCKS WERE PROVIDED; HOWEVER, ¿THAT DIDN¿T WORK EITHER¿. THE PATIENT¿S CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809422&pc=DSQ | 2916596-2025-01982 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-04-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU, REV. D, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809502&pc=NPT | 2025587-2025-02528 | 2024-10-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Unspecified Infection; Sepsis; Cardiac Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER VALVE, A LEFT AND RIGHT VENTRICULAR PERFORATION WAS OBSERVED. SUBSEQUENTLY, THE PROCEDURE WAS TRANSITIONED TO A THORACOTOMY TO ADDRESS THE PERFORATION. FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DEVELOPED A NON-CARDIOVASCULAR INFECTION AND SUBSEQUENTLY DIED DUE TO SEPSIS. PER THE PHYSICIAN, THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE LEFT VENTRICULAR PERFORATION WAS DUE TO DRAINAGE BEING INSERTED TO RESOLVE THE PERICARDIAL FLUID GENERATED BY THE RIGHT VENTRICULAR PERFORATION. THE CAUSE OF THE RIGHT VENTRICULAR PERFORATION WAS DUE TO PACING OF NON-MEDTRONIC PACING WIRE. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM (DCS) DID NOT CAUSE THE VENTRICULAR PERFORATION. IT WAS UNKNOWN WHEN THE INFECTION OR SEPSIS OCCURRED AND THE CAUSE OF THE SEPSIS WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME DELIV SYS D-EV, PRODUCT ID D-EVOLUTFX-2329, SERIAL/LOT UNKNOWN, USE BY DATE UNKNOWN, UDI UNKNOWN. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 D4 H4. CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: HEART VALVES; IMPLANT DATE; EXPLANT DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809539&pc=CAF | 2518422-2025-042952 | 2023-12-21 05:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | CAF | I-NEB CF UK UNIT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102454 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE CORRECT B5 SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS BEEN SCRAPPED; THEREFORE, IT CANNOT BE RETRIEVED FOR EVALUATION. IN ADDITION, APPROPRIATE CODE UPDATED/CORRECTED IN THIS FOLLOW-UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809702&pc=NIQ | 2124215-2025-20336 | 2025-03-17 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS PRE-DILATED WITH A 2.0 X15 MM BALLOON CATHETER FOLLOWED BY INTRAVASCULAR ULTRASOUND. A 2.50 X 38MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS DEPLOYED IN THE DISTAL LEFT MAIN TO LEFT ANTERIOR DESCENDING ARTERY, ANOTHER 2.50 X 38MM SYNERGY XD DES WAS ALSO ADVANCED AND DEPLOYED IN THE MIDDLE LEFT CIRCUMFLEX ARTERY WITH 85% STENOSIS, AND A 2.25 X 32MM SYNERGY XD DES WAS DEPLOYED IN THE DIFFUSE LESION OF THE RIGHT CORONARY ARTERY. POST-DEPLOYMENT, IT WAS OBSERVED THAT THE PATIENT DEVELOPED HYPOTENSION AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS IMMEDIATELY PERFORMED BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST AND AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809714&pc=NIQ | 2124215-2025-20342 | 2025-03-17 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS PRE-DILATED WITH A 2.0 X15 MM BALLOON CATHETER FOLLOWED BY INTRAVASCULAR ULTRASOUND. A 2.50 X 38MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS DEPLOYED IN THE DISTAL LEFT MAIN TO LEFT ANTERIOR DESCENDING ARTERY, ANOTHER 2.50 X 38MM SYNERGY XD DES WAS ALSO ADVANCED AND DEPLOYED IN THE MIDDLE LEFT CIRCUMFLEX ARTERY WITH 85% STENOSIS, AND A 2.25 X 32MM SYNERGY XD DES WAS DEPLOYED IN THE DIFFUSE LESION OF THE RIGHT CORONARY ARTERY. POST-DEPLOYMENT, IT WAS OBSERVED THAT THE PATIENT DEVELOPED HYPOTENSION AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS IMMEDIATELY PERFORMED BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST AND AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809715&pc=NIQ | 2124215-2025-20339 | 2025-03-17 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS PRE-DILATED WITH A 2.0 X15 MM BALLOON CATHETER FOLLOWED BY INTRAVASCULAR ULTRASOUND. A 2.50 X 38MM SYNERGY XD DRUG-ELUTING STENT (DES) WAS DEPLOYED IN THE DISTAL LEFT MAIN TO LEFT ANTERIOR DESCENDING ARTERY, ANOTHER 2.50 X 38MM SYNERGY XD DES WAS ALSO ADVANCED AND DEPLOYED IN THE MIDDLE LEFT CIRCUMFLEX ARTERY WITH 85% STENOSIS, AND A 2.25 X 32MM SYNERGY XD DES WAS DEPLOYED IN THE DIFFUSE LESION OF THE RIGHT CORONARY ARTERY. POST-DEPLOYMENT, IT WAS OBSERVED THAT THE PATIENT DEVELOPED HYPOTENSION AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS IMMEDIATELY PERFORMED BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST AND AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21809816&pc=OUT | 3008881809-2025-00197 | 2025-03-18 04:00 | Death | STRYKER NEUROVASCULAR CORK | 2025-04-09 04:00 | OUT | UNKNOWN_NEUROVASCULAR_PRODUCT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P170024 |  | 1 | Event Description: IT WAS REPORTED THAT 4-5 DAYS FOLLOWING SUBJECT FLOW DIVERTER IMPLANTATION PROCEDURE THE PATIENT HAD A FATALITY. IT WAS REPORTED THAT THE PATIENT DIED FROM A CARDIAC ARREST. NO FURTHER INFORMATION IS AVAILABLE. THE RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT EVENT IS UNKNOWN.Event Description: IT WAS REPORTED THAT 4-5 DAYS FOLLOWING SUBJECT FLOW DIVERTER IMPLANTATION PROCEDURE THE PATIENT HAD A FATALITY. IT WAS REPORTED THAT THE PATIENT DIED FROM A CARDIAC ARREST. NO FURTHER INFORMATION IS AVAILABLE. THE RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT EVENT IS UNKNOWN. Manufacturer Narrative: DUE TO THE AUTOMATED MES (MANUFACTURING EXECUTION SYSTEM) SYSTEM THERE ARE CONTROLS IN THE MANUFACTURING PROCESS TO ENSURE THE PRODUCT MET SPECIFICATIONS UPON RELEASE. THE SUBJECT DEVICE WAS NOT AVAILABLE; THEREFORE, A VISUAL INSPECTION AS WELL AS A FUNCTIONAL EVALUATION COULD NOT BE PERFORMED. THE REPORTED EVENT IS COVERED IN THE DEVICE DIRECTIONS FOR USE (DFU). AS WELL, THE RISK OF THE REPORTED EVENT IS DOCUMENTED IN THE RISK DOCUMENTATION AND THERE ARE CURRENT CONTROLS TO MITIGATE THE RISK OF THE AS REPORTED EVENT. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED, AND IT COULD NOT BE DEFINITIVELY DETERMINED IF THE DEVICE FAILED TO MEET SPECIFICATIONS BECAUSE THE PRODUCT WAS NOT RETURNED. ADDITIONAL INFORMATION STATED "AT A PATIENT HAD A FATALITY FOLLOWING A PROCEDURE WHERE A STRYKER DEVICE WAS USED." THE DEATH OCCURRED 4-5 DAYS POST INTERVENTION AT THE PATIENTS HOME. PATIENT DIED FROM A CARDIAC ARREST. ANTECEDENT OF MULTIPLE SCLEROSIS. NO ACUTE CONDITION REPORTED PRE OR DURING INTERVENTION. AWAITING CORONERS REPORT. IT WAS STATED THAT THE USER DID NOT OBSERVE ANY MALFUNCTION OF THE DEVICE DURING THE PROCEDURE. BASED UPON MEDICAL REVIEW, THE HARM OBSERVED IN THIS COMPLAINT IS ANTICIPATED IN NATURE AS PER THE DEVICE RISK ASSESSMENT. AFTER REVIEWING THE ADDITIONAL INFORMATION RECEIVED, THERE IS NO NEW INFORMATION THAT IMPACTS OR CHANGES THE CURRENT ASSESSMENT DECISION OF THIS COMPLAINT, THEREFORE, THE REPORTABILITY ASSESSMENT DECISION REMAINS UNCHANGED. IF ANY ADDITIONAL INFORMATION IS RECEIVED THAT COULD POTENTIALLY IMPACT THE CURRENT ASSESSMENT DECISION OF THIS RECORD AND/OR THE ROOT CAUSE OF THE INVESTIGATION, THE RECORD WILL BE REOPENED TO INCORPORATE AND ASSESS THE INFORMATION ACCORDINGLY. AN ASSIGNABLE CAUSE OF ANTICIPATED PROCEDURAL COMPLICATION WILL BE ASSIGNED TO THE REPORTED EVENT AS A PRODUCT RELATED ROOT CAUSE DOES NOT APPLY AND THE ISSUE IS DUE TO A KNOWN PHYSIOLOGICAL EFFECT OF THE PROCEDURE AND/OR PATIENT CONDITION NOTED WITH THE DIRECTIONS FOR USE, PRODUCT LABELING AND/OR RISK DOCUMENTATION FILES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810088&pc=NPT | 2015691-2025-02762 | 2025-03-21 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-09 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Gradient Increase; Central Regurgitation | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, A 20MM SAPIEN 3 ULTRA VALVE, IMPLANTED IN THE AORTIC POSITION, IS FAILING AFTER APPROXIMATELY 4 YEARS AND 3 MONTHS. FAILURE MODE IS MODERATE-TO-SEVERE CENTRAL AORTIC INSUFFICIENCY WITH AN ECCENTRIC JET. AORTIC VALVE MEAN GRADIENT MEASURED 23MMHG, PEAK GRADIENT MEASURED 48MMHG. PATIENT SYMPTOMATIC WITH SHORTNESS OF BREATH AND CHEST PAIN. DUE TO MULTIPLE COMORBIDITIES AND PERSISTENT DISABILITY, LIMITED TREATMENT OPTIONS WERE AVAILABLE. A DISCUSSION REGARDING A TAVR VALVE-IN-VALVE PROCEDURE WAS HELD, BUT THE PATIENT OPTED FOR DE-ESCALATION OF CARE AND CONTINUED ON COMFORT MEASURES, ANTICIPATING PROGRESSION TO NATURAL DEATH. THE PATIENT PASSED AWAY IN THE HOSPITAL. CAUSE OF DEATH WAS LISTED AS TAVR VALVE WITH MOD-SEVERE REGURGITATION AS WELL AS CONGESTIVE HEART FAILURE (CHF) EXACERBATION, ACUTE KIDNEY INJURY (AKI), AND NSTEMI. Manufacturer Narrative: INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION BASED ON THE ENGINEER EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: ADDITIONAL CODES ADDED TO H6 TYPE OF INVESTIGATION, CORRECTED H6 INVESTIGATION FINDINGS, AND CORRECTED H6 INVESTIGATION CONCLUSIONS. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODES IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), AND TRAINING MANUALS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE INCREASED GRADIENTS AND CENTRAL REGURGITATION RESULTING IN PATIENT DEATH WERE CONFIRMED BY THE PROVIDED MEDICAL RECORD. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 ULTRA VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. PER THE INSTRUCTIONS FOR USE (IFU), VALVE REGURGITATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE THV PROCEDURE. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO SEVERAL ISSUES INCLUDING PATIENT-RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. IN ADDITION, ELEVATED GRADIENTS MAY INDICATE OBSTRUCTED FLOW ACROSS THE VALVE. AN INCREASE IN GRADIENTS MAY RESULT FROM PATIENT FACTORS SUCH AS HYPERTROPHIC CARDIOMYOPATHY (HCM) OR SUB-VALVULAR STENOSIS. ADDITIONALLY, AN INCREASE IN GRADIENTS CAN INDICATE THAT A LEAFLET IS NOT FUNCTIONING OPTIMALLY DUE TO THROMBOSIS OR EARLY VALVE DEGENERATION. AS NOTED, THE PATIENT HAD VAST COMORBIDITIES (E.G. HYPERLIPIDEMIA (HLD), DIABETES (DM II) CKD AND ETC.), WHICH ARE WELL-KNOWN FACTORS THAT MAY INCREASE THE RISK OF EARLY VALVE DEGENERATION, LEADING TO CENTRAL REGURGITATION AND INCREASED GRADIENT. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (PRE-EXISTING COMORBIDITIES, STENOSIS) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. HOWEVER, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810178&pc=DRB | 3007284006-2025-00067 | 2025-03-27 04:00 | Death | SPECTRANETICS CORPORATION | 2025-04-09 04:00 | DRB | LLD EZ LEAD LOCKING DEVICE | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | K142116 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD, RIGHT ATRIAL (RA) LEAD, AND LEFT VENTRICULAR (LV) LEAD DUE TO BACTEREMIA. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH ON EACH LEAD, ATTEMPTS WERE MADE TO GET THROUGH A BINDING SITE IN THE INNOMINATE. NEXT, SWITCHING TO A SPECTRANETICS 11F TIGHTRAIL SUB-C ROTATING DILATOR SHEATH AND SPECTRANETICS 11F TIGHTRAIL ROTATING DILATOR SHEATH (LONG), ADVANCEMENT WAS MADE INTO THE SUPERIOR VENA CAVA (SVC), ENSURING THE LEADS WERE FREE AND NOT BOUND TO ONE ANOTHER. WITH THE 11F TIGHTRAIL (LONG), THE RV LEAD WAS REMOVED, FOLLOWED BY THE RA LEAD. THEN, WITH THE 11F TIGHTRAIL (LONG) ON THE LV LEAD, PROGRESS WAS MADE DOWN TO THE RA. WHILE APPLYING TRACTION, THE TIGHTRAIL DEVICE WAS ACTUATED, BUT THE LV LEAD WOULD NOT RELEASE, AND THE PATIENT¿S BLOOD PRESSURE BEGAN TO GRADUALLY DROP. A PERFORATION OF THE CORONARY SINUS (CS) WAS CONFIRMED VIA ANGIOGRAM, AND RESCUE EFFORTS BEGAN, INCLUDING CPR, BYPASS, AND STERNOTOMY. THE CS PERFORATION WAS REPAIRED; HOWEVER, THE PATIENT CONTINUED TO HAVE SIGNIFICANT BLOOD LOSS, AND THE LOCATION OF THE BLEEDING WAS NOT IDENTIFIED. FURTHER INTERVENTION WAS CEASED, AND THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE LLD EZ PROVIDING TRACTION WITHIN THE LV LEAD WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A6) PATIENT RACE - NOT AVAILABLE FROM FACILITY. H3) THE LLD EZ WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE LLD DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810183&pc=CBK | 2518422-2025-103905 | 2025-03-12 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | CBK | TRILOGY DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY AND THE DEATH IS NOT BELIEVED TO BE DUE TO A PRODUCT ISSUE. THE CAUSE OF DEATH IS BELIEVED TO BE ASPHYXIATION, AND THE AUTOPSY RESULTS ARE PENDING. THERE IS A CRIMINAL INVESTIGATION AGAINST THE CAREGIVER DUE TO THE PATIENT'S DEATH AND THE DEVICE DATA HAS BEEN REQUESTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY AND THE DEATH IS NOT BELIEVED TO BE DUE TO A PRODUCT ISSUE. THE CAUSE OF DEATH IS BELIEVED TO BE ASPHYXIATION, AND THE AUTOPSY RESULTS ARE PENDING. THERE IS A CRIMINAL INVESTIGATION AGAINST THE CAREGIVER DUE TO THE PATIENT'S DEATH AND THE DEVICE DATA HAS BEEN REQUESTED. ADDITIONAL INFORMATION RECEIVED REPORTS THAT THE CAUSE OF DEATH IS STILL UNKNOWN AND NOT AVAILABLE TO THE HOME SERVICE TO PROVIDE. THE PATIENT HAD RESPIRATOR DEPENDENCE, UNSPECIFIED LEVEL OF CERVICAL SPINAL CORD INJURY, TRACHEOSTOMY, CHRONIC RESPIRATORY FAILURE, HYPOXIA OR HYPERCAPNIA AND GASTROSTOMY. THE PATIENT HAD A HISTORY OF CHRONIC RESPIRATORY FAILURE SECONDARY TO BEING IN AN AUTO PEDESTRIAN MOTOR VEHICLE ACCIDENT AND SUSTAINED A C 1-2 SPINAL CORD INJURY. THE DEVICE IS CURRENTLY UNAVAILABLE FOR RETURN AS IT IS CURRENTLY IN POSSESSION OF LAW ENFORCEMENT UNTIL INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810204&pc=NJK | 3007797756-2025-00013 | 2025-03-06 05:00 | Death | PULMONX CORPORATION | 2025-04-09 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pneumothorax |  |  | 1 | Event Description: ON (B)(6) 2025, A PROCEDURE WAS PERFORMED IN WHICH TWO VALVES WERE REMOVED AND REPLACED. ON (B)(6) 2025, THE PATIENT DEVELOPED A PNEUMOTHORAX AND REQUIRED CHEST TUBE PLACEMENT. DURING THE INTERVENTION, THE PATIENT EXPERIENCED CARDIAC ARREST, RECEIVED CHEST COMPRESSIONS, AND WAS INTUBATED. THE PATIENT PASSED AWAY LATER THAT DAY. FOLLOW-UP INFORMATION HAS BEEN REQUESTED BUT WAS NOT AVAILABLE AT THE TIME OF THIS REPORT. A FOLLOW-UP MDR WILL BE SUBMITTED IF ADDITIONAL DETAILS ARE RECEIVED. Manufacturer Narrative: PNEUMOTHORAX IS THE MOST COMMON SIDE EFFECT ASSOCIATED WITH THE ZEPHYR VALVE TREATMENT (CRINER ET AL. A MULTICENTER RANDOMIZED CONTROLLED TRIAL OF ZEPHYR ENDOBRONCHIAL VALVE TREATMENT IN HETEROGENEOUS EMPHYSEMA (LIBERATE). AM J RESPIR CRIT CARE MED. 2018; 198 (9): 1151-1164). TARGETED LOBAR DEFLATION LIKELY CAUSES INFLATION OF THE IPSILATERAL LOBE, WHICH CAN RESULT IN A TEAR OF THE ALREADY COMPROMISED PARENCHYMAL TISSUE OF THE EMPHYSEMATOUS IPSILATERAL LOBE, RESULTING IN A PNEUMOTHORAX (CRINER ET AL. A MULTICENTER RANDOMIZED CONTROLLED TRIAL OF ZEPHYR ENDOBRONCHIAL VALVE TREATMENT IN HETEROGENEOUS EMPHYSEMA (LIBERATE). AM J RESPIR CRIT CARE MED. 2018; 198 (9): 1151-1164). IN THE LIBERATE STUDY (IDE CLINICAL STUDY USED TO SUPPORT PMA P180002'S APPROVAL), 26.6% OF THE ZEPHYR VALVE SUBJECTS EXPERIENCED A PNEUMOTHORAX IN THE TREATMENT PERIOD ([LESS THAN OR EQUAL TO 45 DAYS). THESE WERE MANAGED USING STANDARD OF CARE PROCEDURES AS PER PREVIOUSLY PUBLISHED GUIDELINES (VALIPOUR, ARSCHANG, ET AL. RESPIRATION 87.6 (2014): 513-521). IN 17.4% OF THE EVENTS, THE PNEUMOTHORAX RESOLVED WITHOUT ANY ADDITIONAL INTERVENTION WITH SUBJECTS UNDER CAREFUL OBSERVATION. IN OVER HALF THE EVENTS (56.5%), THE PNEUMOTHORAX AS MANAGED WITH A CHEST-TUBE ONLY. AN ADDITIONAL 13% OF THE EVENTS WERE MANAGED WITH A CHEST-TUBE AND THE TEMPORARY REMOVAL OF ONE OR MORE VALVES, WHILE ANOTHER 13% OF THE EVENTS WERE MANAGED WITH A CHEST-TUBE AND REMOVAL OF ALL THE IMPLANTED VALVES. UPON SUCCESSFUL RESOLUTION OF THE PNEUMOTHORAX, REMOVED VALVES CAN BE REPLACED. PATIENTS THAT EXPERIENCED A PNEUMOTHORAX EXPERIENCED CLINICAL BENEFITS OF THE ZEPHYR VALVE TREATMENT THAT WERE SIMILAR TO THE BENEFITS EXPERIENCED BY PATIENTS WHO DID NOT HAVE A PNEUMOTHORAX. THE ZEPHYR EBV SYSTEM IFU AND PULMONX TRAINING PROGRAM BOTH SPECIFICALLY REFERENCE PNEUMOTHORAX AS A KNOWN SIDE EFFECT OF THIS PROCEDURE AND THE PUBLISHED GUIDELINES (VALIPOUR, ARSCHANG, ET AL. "EXPERT STATEMENT: PNEUMOTHORAX ASSOCIATED WITH ENDOSCOPIC VALVE THERAPY FOR EMPHYSEMA-POTENTIAL MECHANISMS, TREATMENT ALGORITHM, AND CASE EXAMPLES." RESPIRATION 87.6 (2014): 513-521). THE REPORTED EVENT ALIGNS WITH THE EXPERIENCE OBSERVED IN THE LIBERATE CLINICAL STUDY AND IS AN EXPECTED SIDE EFFECT TO THE ZEPHYR VALVE TREATMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810322&pc=MNS | 2518422-2025-103908 | 2023-05-30 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | MNS | DREAMSTATION BIPAP AUTOSV | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K090539 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION BIPAP SV HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810366&pc=BZD | 2518422-2025-103872 | 2025-03-25 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION 2 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: UDI- NO SN NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810491&pc=NVN | 2017865-2025-39792 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-09 04:00 | NVN | ULTIPACE¿ LEAD, 58CM MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810527&pc=QEW | 3020347218-2025-00026 | 2024-01-01 05:00 | Death | INARI MEDICAL, INC. | 2025-04-09 04:00 | QEW | INARI MEDICAL THROMBECTOMY CATHETERS | Use of Device Problem; Insufficient Information | Great Vessel Perforation |  |  | 1 | Event Description: ON MARCH 12, 2025 INARI RECEIVED A BRAND EQUITY SURVEY RESPONSE WHERE AN ANONYMOUS INDIVIDUAL REPORTED THEY OBSERVED TWO AUTOPSIES WHERE THE PATIENT HAD A CATHETER PERFORATION IN THEIR PULMONARY ARTERY. Manufacturer Narrative: THERE WERE NO SPECIFIC CLINICAL CASES, EVENT DATES, HOSPITAL, OR PHYSICIAN NAMES PROVIDED IN THIS SURVEY RESPONSE. FOLLOW UP WITH THE SURVEY RESPONDENT WAS ATTEMPTED, HOWEVER NO ADDITIONAL INFORMATION WAS PROVIDED. DUE TO LACK OF INFORMATION INARI IS UNABLE TO FURTHER INVESTIGATE THIS SPECIFIC CLAIM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810564&pc=BZD | 2518422-2025-103911 | 2021-09-06 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE HAS PASSED DUE TO AN UNKNOWN HEART CONDITION. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810697&pc=NPT | 2025587-2025-02531 | 2024-08-19 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-09 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Perivalvular Leak; Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING OUTCOMES OF CARDIAC TAMPONADE POST-TRANSCATHETER AORTIC VALVE REPLACEMENT. THE STUDY POPULATION INCLUDED 2,050 PATIENTS WHO WERE PREDOMINANTLY MALE WITH A MEAN AGE OF 81.7 YEARS OLD. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION; THREE PATIENTS WERE IMPLANTED WITH A MEDTRONIC COREVALVE, EVOLUT R, OR EVOLUT PRO BIOPROSTHETIC VALVE AND IN ONE PATIENT A CONFIDA GUIDEWIRE WAS UTILIZED. FOUR DEATHS OCCURRED WITHIN THE STUDY POPULATION: ONE IN-HOSPITAL, ONE WITHIN 30 DAYS, AND TWO WITHIN ONE YEAR OF IMPLANT. FOR ALL DEATHS BUT ONE, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. FOR THE ONE IN-HOSPITAL DEATH: THE PATIENT WHO UNDERWENT IMPLANT OF A MEDTRONIC EVOLUT VALVE HAD SEVERE PARAVALVULAR REGURGITATION WHICH REQUIRED IMPLANT OF A SECOND EVOLUT VALVE. SUBSEQUENTLY THE PATIENT BECAME PROFOUNDLY HYPOTENSIVE WITH EVIDENCE OF CARDIAC TAMPONADE. DESPITE EMERGENT PERICARDIOCENTESIS THE PATIENT REMAINED HEMODYNAMICALLY UNSTABLE. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION AND UNFORTUNATELY PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED ON THIS DEATH. AMONG ALL PATIENTS, CLINICAL OBSERVATIONS INCLUDED: CARDIAC TAMPONADE IN SOME CASES REQUIRING PERICARDIOCENTESIS OR OPEN SURGERY, SHOCK, STROKE, ARRHYTHMIA REQUIRING PERMANENT PACEMAKER IMPLANT, AND ACUTE KIDNEY INJURY. ADDITIONALLY, THE PREFORMED MEDTRONIC CONFIDA GUIDEWIRE WAS USED IN ONE PATIENT WHO DEVELOPED TAMPONADE DUE TO LEFT VENTRICLE LACERATION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: SIMRAT KAUR ET AL. OUTCOMES OF CARDIAC TAMPONADE POST-TRANSCATHETER AORTIC VALVE REPLACEMENT: RESULTS FROM A TERTIARY CARDIAC CENTER. STRUCTURAL HEART VOL 9, 100356 2025. 10.1016/J.SHJ.2024.100356. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810842&pc=BZD | 2518422-2025-103902 | 2025-03-28 04:00 | Death | RESPIRONICS, INC. | 2025-04-09 04:00 | BZD | DREAMSTATION AUTO CPAP, JP | Adverse Event Without Identified Device or Use Problem | Aspiration Pneumonitis | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS NOTIFIED ABOUT A PATIENT WHO DIED. THE CLAIM STATES THAT THE PATIENT CHOKED ON VOMIT WHILE USING AN AUTO CPAP A-FLEX DEVICE. IT IS ALLEGED THAT ASPIRATION PNEUMONIA CAUSED THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810906&pc=LJT | 3006260740-2025-02221 | 2025-04-02 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-04-09 04:00 | LJT | POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K060812 |  | 1 | Event Description: ON (B)(6) 2016, PATIENT UNDERWENT BARD TI POWERPORT IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DIAGNOSED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: ON (B)(6) 2016, PATIENT UNDERWENT BARD TI POWERPORT IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DIAGNOSED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS, VIDEOS, OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INFECTION AND PATIENT DEATH AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED INFECTION AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE A LABELING REVIEW IS NOT REQUIRED. D4 (UNIQUE IDENTIFIER (UDI) #), G3, H4, H6 (COMPONENT, METHOD). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810907&pc=LJT | 3006260740-2025-02220 | 2025-04-02 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-04-09 04:00 | LJT | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K063377 |  | 1 | Event Description: ON (B)(6) 2019, PATIENT UNDERWENT BARD POWERPORT CLEARVUE ISP IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DIAGNOSED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. THE CERTIFICATE OF STERILITY FOR THIS LOT CERTIFIES THAT THE MATERIALS LISTED HAVE MET ALL REQUIRED ACCEPTANCE CRITERIA AT THE TIME OF LOT RELEASE INCLUDING STERILIZATION AND ETHYLENE OXIDE RESIDUALS AND HAS BEEN ACCEPTED BY BARD QUALITY ASSURANCE FOR DISTRIBUTION. IN ADDITION, THE LOT PASSED ALL AND THE APPLICABLE BIOBURDEN AND LAL TESTING RESULTS ARE ACCEPTABLE. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO MEDICAL RECORDS WERE PROVIDED. THE REPORTED INFECTION AND DEATH THEREFORE CANNOT BE CONFIRMED. NO CAUSAL RELATIONSHIP BETWEEN THE PORT AND THE INFECTION OR THE PATIENT DEATH HAS BEEN ESTABLISHED. BASED ON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE OF THE INFECTION AND DEATH IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3, H6 (METHOD). H11: B5, D4 (UNIQUE IDENTIFIER (UDI)#, H6 (COMPONENT). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810926&pc=LJT | 3006260740-2025-02222 | 2025-04-02 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-04-09 04:00 | LJT | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Septic Shock; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K122899 |  | 1 | Event Description: ON (B)(6) 2018, PATIENT UNDERWENT BARD POWERPORT CLEARVUE ISP IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, PATIENT EXPIRED DUE TO SEPTIC SHOCK.Event Description: ON (B)(6) 2018, PATIENT UNDERWENT BARD POWERPORT CLEARVUE ISP IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, PATIENT EXPIRED DUE TO SEPTIC SHOCK. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER IS UNKNOWN. INVESTIGATION SUMMARY: THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS/VIDEOS OR MEDICAL RECORDS WERE PROVIDED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INFECTION, SEPTIC SHOCK, AND DEATH AS NO OBJECTIVE EVIDENCE HAS BEEN PROVIDED TO CONFIRM OR CHARACTERIZE THESE EVENTS. BASED ON THE COMPLAINT INVESTIGATION AND EVALUATION OF THE REPORTED EVENT, A CAUSAL RELATIONSHIP BETWEEN THE DEVICE, INFECTION AND DEATH COULD NOT BE ESTABLISHED. NO DEVICE FAILURES OR ROOT CAUSES COULD BE IDENTIFIED BASED ON THE AVAILABLE INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. G3, H6 (COMPONENT). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810952&pc=FDS | 9610877-2025-00033 | 2025-04-02 04:00 | Death | HOYA CORPORATION PENTAX TOKYO OFFICE | 2025-04-09 04:00 | FDS | PENTAX MEDICAL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Respiratory Arrest | K180292 |  | 1 | Event Description: ON 02-APR-2025, PENTAX MEDICAL BECAME AWARE OF A PATIENT DEATH INVOLVING A PENTAX MEDICAL VIDEO GASTROSCOPE, MODEL EG34-I10, SERIAL NUMBER (B)(6). ACCORDING TO THE INITIAL REPORT, THE PHYSICIAN PERFORMED A THERAPEUTIC ENDOSCOPY TO IDENTIFY A GASTROINTESTINAL (GI) BLEEDING SOURCE. THE BLEEDING SOURCE WAS FOUND, HEMOSTASIS WAS ACHIEVED USING ELECTROCOAGULATION, AND CLIPS WERE DEPLOYED TO PREVENT FURTHER BLEEDING. DURING THE PROCEDURE, THE PATIENT'S PHYSICAL CONDITION DETERIORATED. THE ENDOSCOPE WAS REMOVED, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND A CODE BLUE TEAM WAS ACTIVATED. THE PATIENT WAS PRONOUNCED DEAD ON THE SAME DAY, (B)(6) 2025. THE PENTAX MEDICAL VIDEO GASTROSCOPE REPORTEDLY FUNCTIONED NORMALLY DURING THE PROCEDURE. THERE WAS NO INDICATION OF DEVICE MALFUNCTION, AND THE DEATH WAS NOT ATTRIBUTED TO THE ENDOSCOPE. THE PROCEDURE WAS SCHEDULED IN ADVANCE AND WAS NOT AN EMERGENCY CASE. THE PATIENT HAD A HISTORY OF HEART SURGERY; HOWEVER, DETAILED MEDICAL HISTORY WAS NOT AVAILABLE. THIS EVENT MEETS THE REQUIREMENTS FOR FDA REPORTABILITY; HOWEVER, SUBMISSION OF THIS REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, MANUFACTURER OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. Manufacturer Narrative: H6 CONTINUED: HEALTH EFFECT - CLINICAL CODE: 1762 CARDIAC ARREST, 4461 RESPIRATORY ARREST. HEALTH EFFECT - IMPACT CODE: 1802 DEATH, 4635 UNEXPECTED DETERIORATION. MEDICAL DEVICE PROBLEM CODE: 2993 ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM COMPONENT CODE: 4755. TYPE OF INVESTIGATION: 4110 HISTORICAL DATA ANALYSIS, 3331 ANALYSIS OF PRODUCTION RECORDS, 4111 COMMUNICATION/INTERVIEWS. INVESTIGATION FINDINGS: 213 NO DEVICE PROBLEM FOUND. INVESTIGATION CONCLUSIONS: 4310 CAUSES TRACED TO NON-DEVICE RELATED FACTORS. CORRECTION INFORMATION B4: DATE OF THIS REPORT G6: FOLLOW-UP #: 1 H2: IF FOLLOW-UP, WHAT TYPE? H3: DEVICE EVALUATED BY MANUFACTURE H6: ADVERSE EVENT PROBLEM ADDITIONAL INFORMATION D4: PRIMARY UNIQUE DEVICE IDENTIFIER (UDI) H4: DEVICE MANUFACTURE DATE H11: EVALUATION SUMMARY EVALUATION SUMMARY PENTAX MEDICAL AMERICA AND THE JAPANESE INVESTIGATION TEAM CONDUCTED A GOOD FAITH EFFORT (GFE) TO OBTAIN ADDITIONAL INFORMATION REGARDING THIS COMPLAINT. THE FOLLOWING TWO QUESTIONS WERE SUBMITTED TO THE SITE REPRESENTATIVE MULTIPLE TIMES BETWEEN 11-APR-2025 AND 15-MAY-2025: - HAS THE SUBJECT ENDOSCOPE (EG34-I10_(B)(6)) CONTINUED TO BE USED AT THE FACILITY AFTER THE COMPLAINT? - WHO DETERMINED THAT THERE WAS NOTHING WRONG WITH THE SCOPE (A PHYSICIAN OR A FACILITY STAFF MEMBER)? AFTER A DELAY, A RESPONSE WAS PROVIDED BY THE SITE REPRESENTATIVE ON 15-MAY-2025. THE RESPONSE INDICATED THAT THE ENDOSCOPE IS STILL IN USE AT THE FACILITY AND NO COMPLAINTS OR ISSUES HAVE BEEN REPORTED. IT WAS CONFIRMED BY A FACILITY STAFF MEMBER (ENDOSCOPY SUPERVISOR) THAT THE DEVICE FUNCTIONED NORMALLY DURING THE PROCEDURE. AFTER THE PROCEDURE WAS ABORTED DUE TO THE PATIENT'S CONDITION, THE ENDOSCOPE WAS LEAK TESTED AND REPROCESSED, AND NO ABNORMALITIES WERE FOUND. UPON RECEIVING THIS INFORMATION, THE JAPANESE INVESTIGATION REPRESENTATIVE CONFIRMED ON 16-MAY-2025 THAT THERE WERE NO PROBLEMS WITH THE ENDOSCOPE AND THAT IT COULD CONTINUE TO BE USED. SINCE THE DEVICE HAD BEEN FUNCTIONING NORMALLY AND HAD UNDERGONE STANDARD REPROCESSING WITHOUT ANY ISSUES, IT WAS DETERMINED THAT FURTHER INSPECTION OF THE DEVICE WAS UNNECESSARY. THE EVENT THAT OCCURRED WAS A DETERIORATION IN THE PATIENT'S CONDITION. BASED ON THE INVESTIGATION AND GFE, THE GASTROINTESTINAL BLEEDING HAD OCCURRED BEFORE THE ENDOSCOPE WAS INSERTED, AND IT WAS DETERMINED THAT IT WAS NOT DUE TO EQUIPMENT FAILURE. FURTHERMORE, BASED ON THE TREND ANALYSIS, THERE IS NO SIGNIFICANT INCREASE OR UPWARD TREND IN REPAIR COMPLAINTS RELATED TO THIS FAILURE MODE/HAZARD. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED BY THE MANUFACTURER. THE REVIEW BY DHR CONFIRMED THAT THE INSTRUMENT WAS MANUFACTURED UNDER NORMAL CONDITIONS ON 05-JUL-2024, PASSED ALL REQUIRED INSPECTIONS, AND WAS RELEASED ACCORDINGLY. ALSO, THERE WERE NO REWORKS OR CONCESSIONS, AND THE DATES OF APPROVAL FOR SHIPMENT AND ACTUAL DATE SHIPPED WERE CONFIRMED ON 16-JUL-2024. PENTAX MEDICAL HAS NOT RECEIVED ANY FURTHER INFORMATION FOR THIS EVENT AND THEREFORE, CONSIDERS THIS MEDWATCH REPORT CLOSED. ECTIONS, AND WAS RELEASED ACCORDINGLY. ALSO, THERE WERE NO REWORKS OR CONCESSIONS, AND THE DATES OF APPROVAL FOR SHIPMENT AND ACTUAL DATE SHIPPED WERE CONFIRMED ON 16-JUL-2024. PENTAX MEDICAL HAS NOT RECEIVED ANY FURTHER INFORMATION FOR THIS EVENT AND THEREFORE, CONSIDERS THIS MEDWATCH REPORT CLOSED. Manufacturer Narrative: H6 CONTINUED: HEALTH EFFECT - CLINICAL CODE : 1762 CARDIAC ARREST, 4461 RESPIRATORY ARREST HEALTH EFFECT - IMPACT CODE : 1802 DEATH, 4635 UNEXPECTED DETERIORATION MEDICAL DEVICE PROBLEM CODE : 2993 ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM COMPONENT CODE : 4776 INSUFFICIENT INFORMATION TYPE OF INVESTIGATION : 4118 TYPE OF INVESTIGATION NOT YET DETERMINED INVESTIGATION FINDINGS : 3233 RESULTS PENDING COMPLETION OF INVESTIGATION INVESTIGATION CONCLUSIONS : 11 CONCLUSION NOT YET AVAILABLE. ADDITIONAL EQUIPMENT USED: A PENTAX MEDICAL VIDEO PROCESSOR (MODEL EPK-I8020C, SERIAL NUMBER (B)(6), FIRMWARE VERSION 0107C-1.0055) WAS USED IN CONJUNCTION WITH THE VIDEO GASTROSCOPE. THE PROCESSOR FUNCTIONED NORMALLY AND NO MALFUNCTION WAS NOTED. INVESTIGATION IS IN-PROCESS. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE FILED WITH THE NEW INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810953&pc=ONU | 2124215-2025-21051 | 2024-07-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | ONU | RANGER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2024, THE SUBJECT UNDERWENT TREATMENT WITH THE RANGER DRUG COATED BALLOONS AS A PART OF THE CLINICAL TRIAL. THE TARGET LESION WAS IN RIGHT DISTAL SUPERFICIAL FEMORAL ARTERY WITH A LENGTH OF 120 MM AND 88% STENOSIS. PRIOR TO THE TARGET LESION TREATMENT WITH THE STUDY DEVICE, PRE-DILATION WAS PERFORMED BY USING 5 MM X 120 MM CHOCOLATE PTA BALLOON. TREATMENT OF TARGET LESION WAS PERFORMED BY DILATION USING TWO 5 MM X 100 MM RANGER DRUG COATED BALLOONS STUDY DEVICES. POST TREATMENT, DILATION WAS PERFORMED BY USING 5 MM X 200 MM NON-BOSTON SCIENTIFIC (BSC) PTA BALLOON FOLLOWED BY THE PLACEMENT OF 5.5 MM X 50 MM NON-BSC DRUG ELUTING STENT. FOLLOWING TREATMENT, THE FINAL RESIDUAL STENOSIS WAS NOTED TO BE 0%. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED ON CLOPIDOGREL. ON (B)(6) 2024, 33 DAYS POST INDEX PROCEDURE, THE SUBJECT EXPIRED DUE TO UNKNOWN CAUSE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. THE PRODUCT IS UNKNOWN AS THIS IS A CLINICAL COMPLAINT; WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21810959&pc=ONU | 2124215-2025-21054 | 2024-07-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-09 04:00 | ONU | RANGER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2024, THE SUBJECT UNDERWENT TREATMENT WITH THE RANGER DRUG COATED BALLOONS AS A PART OF THE CLINICAL TRIAL. THE TARGET LESION WAS IN RIGHT DISTAL SUPERFICIAL FEMORAL ARTERY WITH A LENGTH OF 120 MM AND 88% STENOSIS. PRIOR TO THE TARGET LESION TREATMENT WITH THE STUDY DEVICE, PRE-DILATION WAS PERFORMED BY USING 5 MM X 120 MM CHOCOLATE PTA BALLOON. TREATMENT OF TARGET LESION WAS PERFORMED BY DILATION USING TWO 5 MM X 100 MM RANGER DRUG COATED BALLOONS STUDY DEVICES. POST TREATMENT, DILATION WAS PERFORMED BY USING 5 MM X 200 MM NON-BOSTON SCIENTIFIC (BSC) PTA BALLOON FOLLOWED BY THE PLACEMENT OF 5.5 MM X 50 MM NON-BSC DRUG ELUTING STENT. FOLLOWING TREATMENT, THE FINAL RESIDUAL STENOSIS WAS NOTED TO BE 0%. ON JUNE 20,2024, THE SUBJECT WAS DISCHARGED ON CLOPIDOGREL. ON JULY 2024, 33 DAYS POST INDEX PROCEDURE, THE SUBJECT EXPIRED DUE TO UNKNOWN CAUSE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. THE PRODUCT IS UNKNOWN AS THIS IS A CLINICAL COMPLAINT, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21794440&pc=PNJ | 9612164-2025-01785 | 2025-03-19 04:00 | Death | MEDTRONIC IRELAND | 2025-04-08 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF) AND AN APPARENT RESPIRATORY ARREST AND DIED. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21794593&pc=OZP | 2032227-2025-150834 | 2025-03-03 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-08 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hypoglycemia | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS CARDIAC ARREST. CUSTOMER EXPERIENCED HYPOGLYCEMIA BEFORE DEATH WHICH WAS TREATED BY GLUCOSE TABLETS. THE BLOOD GLUCOSE VALUE OF 59 MG/DL. TROUBLESHOOTING WAS PERFORMED. CUSTOMER HAD OTHER HEALTH ISSUES LIKE GIST TUMORS AND DIABETES. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-397A, MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. PRODUCT RETURN REQUESTED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21794634&pc=PNJ | 9612164-2025-01786 | 2025-03-16 04:00 | Death | MEDTRONIC IRELAND | 2025-04-08 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED A CARDIAC PERFORATION AND IT WAS NOTED THERE WAS A SLIGHT PERICARDIAL EFFUSION NOTED ON THE ECHOCARDIOGRAM PRIOR TO LEAVING THE OPERATING ROOM. IT WAS ALSO NOTED THE PATIENT PASSED AWAY THREE DAYS POST IMPLANT DUE TO AN ACUTE EXACERBATION OF PRE-EXISTING HEART FAILURE AND MULTI-ORGAN FAILURE FOLLOWING THE PROCEDURE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6); D9: <(><<)>NO>. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795018&pc=OZP | 2032227-2025-150977 | 2019-09-22 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-08 04:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Break | Cardiac Arrest; Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC LEGAL FROM THE ATTORNEY OF THE FAMILY ON (B)(6) 2025. THE EVENT INVOLVED PRODUCT(S) MMT-1780KPK, UNOMEDICAL, MMT-332A. THE REPORTER ALLEGED THAT THE USE OF ONE OR MORE MEDTRONIC MINI MED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING CAUSED THE CLAIMANT TO SUFFER PERSONAL INJURIES AND WAS HOSPITALIZED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, WERE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). WHEN AND IF THE AFFECTED SERIAL NUMBER WAS IDENTIFIED, ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY. IT WAS STATED THAT THERE WAS 3-7 DAY OF HOSPITALIZATION - ICU STAY. THE CUSTOMER PASSED AWAY DUE TO HEART FAILURE ON (B)(6) 2019. THE INSULIN PUMP WAS NOT WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795170&pc=OZP | 2032227-2025-151029 | 2025-03-11 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-08 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOME ON (B)(6) 2025. THE CUSTOMER WAS NOT ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS ILLNESS FROM COMPLICATIONS OF DIABETES. CUSTOMER BLOOD GLUCOSE WAS UNKNOWN AT THE TIME OF INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-342, MMT-441AK, MMT-1884. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, CUSTOMER NOT WORN INSULIN PUMP AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS UTILIZED WITHIN 48 HOURS OF THE EVENT AND ALSO IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-342. NO PRODUCT RETURN IS REQUIRED FOR MMT-441AK. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795427&pc=MVK | 3008642652-2025-03588 | 2025-02-10 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-08 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 21:30:17 ON (B)(6) 2025. AT 11:12:55 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 11:13:28, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 11:14:00, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 11:14:27, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 11:15:00, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 11:15:26, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 11:37:53 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795445&pc=OZP | 2032227-2025-151112 | 2025-03-12 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-08 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Appropriate Term/Code Not Available | Insufficient Information | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON UNKNOWN DATE OF DEATH. THE PRIMARY CAUSE OF DEATH DUE TO MALFUNCTION OF THE PUMP. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-332A, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. PRODUCT RETURN REQUIRED FOR MMT-332A. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. PRODUCT RETURN REQUIRED FOR MMT-1884. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. PRODUCT RETURN REQUIRED FOR UNOMEDICAL. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795528&pc=OZD | 1220648-2025-27290 | 2025-03-07 05:00 | Death | ABIOMED, INC. | 2025-04-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS SUPPORTED BY ECPELLA FOR MECHANICAL CIRCULATORY SUPPORT. EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS REMOVED AND THE NEXT DAY, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION. ECMO WAS REINSERTED, BUT THE PATIENT COULD NOT BE SAVED AND SUBSEQUENTLY THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795712&pc=DSQ | 2916596-2025-02276 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT HOME. THE REASON WAS UNKNOWN. ADDITIONAL INFORMATION WAS UNAVAILABLE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795728&pc=NKB | 1526439-2025-00482 |  | Death | MEDOS INTERNATIONAL SÃ RL CH | 2025-04-08 04:00 | NKB | UNK - CONSTRUCTS: EXPEDIUM POSTERIOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A DATABASE RELATED RESEARCH ACTIVITY (DRRA) "DEPUY SPINAL IMPLANTS, SUBREPORT 2025 ¿ EXPEDIUM ANTERIOR" FROM SWEDISH SPINE REGISTRY (SWESPINE) REPORT. A TOTAL OF 35 PATIENTS WERE OPERATED WITH EXPEDIUM ANTERIOR INSTRUMENTATION BETWEEN JANUARY 02, 2006 AND DECEMBER 31, 2024. IN 4 CASES IT WAS COMBINED WITH EXPEDIUM POSTERIOR, IN 2 CASES WITH A NON-DEPUY DEVICE. ALL CASES ARE INCLUDED IN THE ANALYSIS. MEAN AGE OF THE PATIENTS WAS 20 YEARS. FEMALES WERE 60%. THE FOLLOWING COMPLICATIONS HAVE BEEN IDENTIFIED INCLUDE: REGISTERED INTRA-OPERATIVE (N=10): - (N=1) DEATH, - (N=3) DURAL INJURY, - (N=2) MEDULLARY INJURY, - (N=1) ROOT INJURY, - (N=3) OTHER. REGISTERED POST-OP COMPLICATIONS WITHIN 1 YEAR (N=3): - (N=2) SURGICAL SITE INFECTION, - (N=1) THROMBOSIS. REOPERATIONS (N=4): - (N=1) DRAIN HEMATOMA, - (N=1) REMOVAL IMPLANT, - (N=2) OTHER. THIS IS FOR THE UNKNOWN DEPUY EXPEDIUM ANTERIOR INSTRUMENTATION AND UNKNOWN DEPUY EXPEDIUM POSTERIOR INSTRUMENTATION. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. ADDITIONAL NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. INVESTIGATION SUMMARY: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21795924&pc=NIP | 2024168-2025-03588 | 2016-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-04-08 04:00 | NIP | SUPERA | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P120020 |  | 1 | Event Description: THE ARTICLE AIMED TO EVALUATE THE CLINICAL AND PROCEDURAL OUTCOMES OF BIOMIMETIC STENT IMPLANTATION IN PATIENTS WITH FEMOROPOPLITEAL ATHEROSCLEROTIC DISEASE, FOCUSING ON PATENCY RATES, PROCEDURAL SUCCESS, AND MAJOR ADVERSE LIMB EVENTS (MALE). 294 PATIENTS WITH FEMOROPOPLITEAL STENOSIS OR OCCLUSION TREATED WITH THE SUPERA STENT FROM JANUARY 2016 TO DECEMBER 2023. PATIENT EFFECTS DOCUMENTED INCLUDED: DEATH. THE ARTICLE CONCLUDED SUPERA STENTS PROVIDE HIGH PATENCY RATES AND FAVORABLE CLINICAL OUTCOMES IN COMPLEX FEMOROPOPLITEAL LESIONS. HOWEVER, LESION COMPLEXITY AND OCCLUSION LENGTH SIGNIFICANTLY IMPACT LONG-TERM SUCCESS. THE FINDINGS HIGHLIGHT THE IMPORTANCE OF CAREFUL PATIENT SELECTION AND LESION ASSESSMENT FOR OPTIMIZING ENDOVASCULAR TREATMENT STRATEGIES IN PAD MANAGEMENT. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED ¿ENDOVASCULAR TREATMENT OF FEMORO-POPLITEAL DISEASE WITH THE SUPERA STENT: A SINGLE CENTER EXPERIENCE.¿ NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE SUPERA PERIPHERAL STENT SYSTEMS INSTRUCTIONS FOR USE (IFU) AS A POTENTIAL ADVERSE EFFECT OF PERIPHERAL PERCUTANEOUS INTERVENTION. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED "ENDOVASCULAR TREATMENT OF FEMORO-POPLITEAL DISEASE WITH THE SUPERA STENT: A SINGLE CENTER EXPERIENCE" B2: DATE OF DEATH WILL BE ESTIMATED AS (B)(6) 2016 AS FOLLOW UP WAS PERFORMED AT 1 MONTH POST PROCEDURE. B3: THE DATE OF PROCEDURE WILL BE ESTIMATED AS (B)(6) 2016 AS THIS IS THE DATE WHEREIN PATIENTS CONSENTED TO STUDY PARTICIPATION BETWEEN JANUARY 2016 TO DECEMBER 2023. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21796030&pc=NKB | 1526439-2025-00486 |  | Death | MEDOS INTERNATIONAL SÃ RL CH | 2025-04-08 04:00 | NKB | UNK - CONSTRUCTS: EXPEDIUM: ANTERIOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THIS REPORT IS BEING FILED AFTER THE REVIEW OF A CLINICAL EVALUATION REPORT (CER) FROM A DATABASE RELATED RESEARCH ACTIVITY (DRRA) "DEPUY SPINAL IMPLANTS SUBREPORT 2025 ¿ EXPEDIUM ANTERIOR" FROM SWEDISH SPINE REGISTRY (SWESPINE) REPORT. A TOTAL OF 35 PATIENTS WERE OPERATED WITH EXPEDIUM ANTERIOR INSTRUMENTATION BETWEEN JANUARY 02, 2006 AND DECEMBER 31, 2024. IN 4 CASES IT WAS COMBINED WITH EXPEDIUM POSTERIOR, IN 2 CASES WITH A NON-DEPUY DEVICE. ALL CASES ARE INCLUDED IN THE ANALYSIS. MEAN AGE OF THE PATIENTS WAS 20 YEARS. FEMALES WERE 60%. THE FOLLOWING COMPLICATIONS HAVE BEEN IDENTIFIED INCLUDE: REGISTERED INTRA-OPERATIVE (N=10): - (N=1) DEATH, - (N=3) DURAL INJURY, - (N=2) MEDULLARY INJURY, - (N=1) ROOT INJURY, - (N=3) OTHER. REGISTERED POST-OP COMPLICATIONS WITHIN 1 YEAR (N=3): - (N=2) SURGICAL SITE INFECTION - (N=1) THROMBOSIS REOPERATIONS (N=4): - (N=1) DRAIN HEMATOMA, - (N=1) REMOVAL IMPLANT, - (N=2) OTHER. THIS IS FOR THE UNKNOWN DEPUY EXPEDIUM ANTERIOR INSTRUMENTATION AND UNKNOWN DEPUY EXPEDIUM POSTERIOR INSTRUMENTATION. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. INVESTIGATION SUMMARY: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21796365&pc=KRA | 9612164-2025-01798 |  | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-04-08 04:00 | KRA | PHENOM CATHETER | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ruptured Aneurysm; Thrombosis/Thrombus | K151638 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO TREATMENT, THE PATIENT EXPERIENCED A RUPTURED ANEURYSM AND HEMORRHAGIC STROKE. THE PHYSICIAN DECIDED TO TREAT THE WIDE-NECK ANEURYSM LOCATED FROM THE LEFT MIDDLE CEREBRAL ARTERY TO THE LEFT INTERNAL CAROTID ARTERY USING A PIPELINE VANTAGE WITH SHIELD TECHNOLOGY AND ADJUNCTIVE TARGET COILS. THE PROCEDURE WAS TECHNICALLY SUCCESSFUL WITH ANGIOGRAPHIC CONFIRMATION. HOWEVER, WITHIN A FEW WEEKS POST-PROCEDURE, THE ANEURYSM RE-RUPTURED, THE PIPELINE THROMBOSED, AND THE PATIENT SUFFERED ANOTHER HEMORRHAGIC STROKE, ULTIMATELY RESULTING IN DEATH. THE PIPELINE WAS USED OFF-LABEL WITH THE DISTAL LANDING ZONE IN THE MIDDLE CEREBRAL ARTERY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ANCILLARY DEVICES: SHEATH PENUMBRA BMX96, GUIDE CATHETER NAVIEN 6F .072ID 115 CM, MICROCATHETER MEDTRONIC PHENOM21 150 CM MICROCATHETER LOT #229655747. Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE PATIENT HAD BEEN HOSPITALIZED PRIOR TO THEIR PASSING AND THEY WERE ADMITTED WITH SPONTANEOUS INTRACEREBRAL HEMORRHAGE (SICH). THE INFORMATION IS CONFIRMED BY THE PHYSICIAN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21796411&pc=DQY | 9612164-2025-01799 |  | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-08 04:00 | DQY | REFLEX CATHETER | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ruptured Aneurysm; Thrombosis/Thrombus |  |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO TREATMENT, THE PATIENT EXPERIENCED A RUPTURED ANEURYSM AND HEMORRHAGIC STROKE. THE PHYSICIAN DECIDED TO TREAT THE WIDE-NECK ANEURYSM LOCATED FROM THE LEFT MIDDLE CEREBRAL ARTERY TO THE LEFT INTERNAL CAROTID ARTERY USING A PIPELINE VANTAGE WITH SHIELD TECHNOLOGY AND ADJUNCTIVE TARGET COILS. THE PROCEDURE WAS TECHNICALLY SUCCESSFUL WITH ANGIOGRAPHIC CONFIRMATION. HOWEVER, WITHIN A FEW WEEKS POST-PROCEDURE, THE ANEURYSM RE-RUPTURED, THE PIPELINE THROMBOSED, AND THE PATIENT SUFFERED ANOTHER HEMORRHAGIC STROKE, ULTIMATELY RESULTING IN DEATH. THE PIPELINE WAS USED OFF-LABEL WITH THE DISTAL LANDING ZONE IN THE MIDDLE CEREBRAL ARTERY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ANCILLARY DEVICES: SHEATH PENUMBRA BMX96, GUIDE CATHETER NAVIEN 6F .072ID 115CM, MICROCATHETER MEDTRONIC PHENOM21 150CM MICROCATHETER LOT #229655747. Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE PATIENT HAD BEEN HOSPITALIZED PRIOR TO THEIR PASSING AND THEY WERE ADMITTED WITH SPONTANEOUS INTRACEREBRAL HEMORRHAGE (SICH). THE INFORMATION IS CONFIRMED BY THE PHYSICIAN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21796543&pc=DSQ | 2916596-2025-02028 | 2025-02-23 05:00 | Death | THORATEC CORPORATION | 2025-04-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: LOG FILES WERE SENT FOR REVIEW. THE EVENT LOG FILE CAPTURED LOW POWER ADVISORY EVENTS ON (B)(6) 2025 BETWEEN 19:06:06 AND 20:49:55. THE EVENT LOG FILE CAPTURED LOW POWER HAZARD EVENTS ON (B)(6) 2025 BETWEEN 20:49:55 AND 20:57:58. THE EVENT LOG FILE CAPTURED NO EXTERNAL POWER EVENTS ON (B)(6) 2025 BETWEEN 20:57:59 AND 22:41:07 WITH THE EMERGENCY BACKUP BATTERY (EBB) OPERATING THE SYSTEM DURING THESE TIMES. THE PUMP STOPPED ON (B)(6) 2025 AT 22:41:07 DUE TO THE EBB BEING COMPLETELY DRAINED. THE PUMP AND PERIPHERAL EQUIPMENT OPERATED AS INTENDED WHEN THERE WAS ADEQUATE POWER TO THE SYSTEM. IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 WITH UNKNOWN CAUSE OF DEATH. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025. A COMPLETE SYSTEM SHUTDOWN DUE TO FULL BATTERY DEPLETION WAS CAPTURED (CAPTURED UNDER MFR 2916596-2025-02032). PRIOR TO THE PUMP STOP, THE DEVICE APPEARED TO BE OPERATING AS EXPECTED AT THE FIXED SPEED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21796594&pc=BZD | 2518422-2025-103756 | 2025-03-11 04:00 | Death | RESPIRONICS, INC. | 2025-04-08 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21796980&pc=FSA | 1419652-2025-00017 | 2023-04-30 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-04-08 04:00 | FSA | MAXI 500 | Detachment of Device or Device Component | Laceration(s) |  |  | 1 | Event Description: ARJO WAS NOTIFIED OF THE EVENT BY HEALTH CANADA ON (B)(6)2025. IT WAS REPORTED THAT DURING A PATIENT TRANSFER, THE NON ARJO SLING LOOP DETACHED FROM THE SPREADER BAR OF THE DEVICE. AS A RESULT OF THE EVENT, THE PATIENT FELL OUT OF THE DEVICE AND HIT THEIR HEAD ON THE LIFT. AS A RESULT OF THE INCIDENT, THE PATIENT WAS TAKEN TO HOSPITAL. THE PATIENT'S INJURY REQUIRED STITCHES FOR A LACERATION. THE PATIENT PASSED AWAY 4 DAYS AFTER THE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21797033&pc=MIH | 2247858-2025-00101 | 2023-08-07 04:00 | Death | BOLTON MEDICAL, INC. | 2025-04-08 04:00 | MIH | RELAY PRO THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Vascular Dissection | P200045 |  | 1 | Event Description: "SUBJECT PRESENTED PENETRATING AORTIC ULCER, WHERE 2 RELAYPRO BARE STENTS WERE PLACED. IT WAS A PLANNED PROXIMAL EXTENSION OF DEGENERATION OF CHRONIC TYPE B DISSECTION AFTER PREVIOUS REPAIR, PAU. INDEX PROCEDURE TOOK PLACE ON (B)(6) 2023 AND THE SUBJECT WAS DISCHARGED ON THE (B)(6) 2023. SUBJECT COMPLETED THEIR EARLY F/U VISIT ON (B)(6) 2023. HOWEVER, AS THE SITE WAS TRYING TO REACH THE SUBJECT FOR THEIR 1Y FOLLOW-UP VISIT, THEY WERE UNSUCCESSFUL, AND DID AN OBITUARY SEARCH AND FOUND HER OBITUARY. OBITUARY STATES THAT THE SUBJECT PASSED AWAY ON (B)(6) 2023; UNKNOWN CAUSE OF DEATH; UNDETERMINED IF RELATED TO DEVICE OR PROCEDURE." PATIENT OUTCOME: "DECEASED ON (B)(6) 2023." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR 2247858-2025-00100, AND DEVICE 2 IS BEING REPORTED UNDER THIS MDR. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21797055&pc=MIH | 2247858-2025-00100 | 2023-08-07 04:00 | Death | BOLTON MEDICAL, INC. | 2025-04-08 04:00 | MIH | RELAY PRO THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Vascular Dissection | P200045 |  | 1 | Event Description: "SUBJECT PRESENTED PENETRATING AORTIC ULCER, WHERE 2 RELAYPRO BARE STENTS WERE PLACED. IT WAS A PLANNED PROXIMAL EXTENSION OF DEGENERATION OF CHRONIC TYPE B DISSECTION AFTER PREVIOUS REPAIR, PAU. INDEX PROCEDURE TOOK PLACE ON (B)(6) 2023 AND THE SUBJECT WAS DISCHARGED ON THE (B)(6) 2023. SUBJECT COMPLETED THEIR EARLY F/U VISIT ON (B)(6) 2023. HOWEVER, AS THE SITE WAS TRYING TO REACH THE SUBJECT FOR THEIR 1Y FOLLOW-UP VISIT, THEY WERE UNSUCCESSFUL AND DID AN OBITUARY SEARCH AND FOUND HER OBITUARY. OBITUARY STATES THAT THE SUBJECT PASSED AWAY (B)(6) 2023; UNKNOWN CAUSE OF DEATH; UNDETERMINED IF RELATED TO DEVICE OR PROCEDURE." PATIENT OUTCOME: "DECEASED (B)(6) 2023." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER AND DEVICE 2 IS BEING REPORTED UNDERMDR 2247858-2025-00101. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21797169&pc=NIQ | MW5168702 | 2025-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-08 04:00 | NIQ | SYNERGY XD MONORAIL EVEROLIMUS-ELUTING PLATINUM CHROMIUM CORONARY STENT SYSTEM | Break; Physical Resistance/Sticking | Cardiac Arrest; Chest Pain; Discomfort; Respiratory Arrest |  |  | 1 | Event Description: PATIENT ARRIVED IN CATH LAB FOR A CARDIAC CATHETERIZATION. PROCEDURE BEGAN WITH ACCESS TO THE RIGHT GROIN AND DIAGNOSTIC IMAGING SHOWED A LESION IN THE PROXIMAL LAD. LESION WAS PREDILATED WITH A 2.0 X 20MM TAKERU BALLOON. A 2.75 X 24MM SYNERGY DES WAS THEN PASSED DOWN AND ACROSS THE LESION. STENT DEPLOYED TO 20 ATMS. UPON ATTEMPTED REMOVAL OF DELIVERY CATHETER, RESISTANCE WAS MET. CATHETER THEN RELEASED AND WAS REMOVED WHERE IT WAS NOTED THE CATHETER WAS BROKEN. ANGIOGRAPHY SHOWED THE BALLOON MARKERS WERE STILL IN THE VESSEL. PATIENT BEGAN TO COMPLAIN OF CHEST DISCOMFORT AND PATIENT WAS INTUBATED. VASCULAR SURGEON WAS CALLED TO ASSIST AND SUCCESSFULLY RETRIEVED BROKEN PIECE FROM BODY. PATIENT DISTRESS CONTINUED TO ESCALATE. CARDIAC AND RESPIRATORY ARREST SHORTLY FOLLOWED. CPR BEGAN AND ATTEMPTS TO PLACE AN IMPELLA DEVICE UNSUCCESSFUL. PATIENT EXPIRED IN CATH LAB.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21797525&pc=DSQ | 2916596-2025-02061 | 2025-03-17 04:00 | Death | THORATEC CORPORATION | 2025-04-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Pulmonary Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOPULMONARY FAILURE. THE PATIENT WAS HYPERCAPNIC WITH INCREASING PULMONARY VASCULAR CONGESTION. DIURETIC REQUIREMENTS INCREASED, AND THEY LATER BECAME HYPOTENSIVE. THEY WERE STARTED ON VASOPRESSORS. THE PATIENT'S SON REQUESTED NO FURTHER ESCALATION OF CARE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. NO DEVICES WOULD RETURN. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21797544&pc=OZD | 1220648-2025-27296 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED LIMB ISCHEMIA AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT AND DUE TO EXTENSIVE PAST CARDIAC HISTORY AND PATIENT FOUND WITH NEW LOW EJECTION FRACTION, THE PATIENT WAS PLACED ON AN IMPELLA CP DEVICE. THE IMPELLA CP WAS PRIMED AND BACKLOADED OVER 0.18 WIRE INTO THE LEFT VENTRICLE AND STARTED ON AUTO FLOWING 3.7 LITERS PER MINUTE ONCE THE WIRE WAS REMOVED. THE POSITION WAS OPTIMIZED UNDER FLUOROSCOPY AND TRANSESOPHAGEAL ECHOCARDIOGRAM. POST IMPLANT, THE IMPELLA CP DEVICE WAS REPOSITIONED ONCE THE PEEL AWAY SHEATH WAS REMOVED, AND THE REPO SHEATH WAS SUTURED INTO PLACE. DISTAL FLOW ANGIOGRAM WAS TAKEN WITH LIMITED FLOW PAST THE IMPELLA SHEATH, PUT PULSES WERE DOPPLERABLE. FEET WERE BLUE WHEN THE PATIENT ARRIVED IN THE ROOM. PREVIOUS BELOW BILATERAL LOWER EXTREMITIES ULTRASOUND WAS DONE IN THE EMERGENCY ROOM BEFORE GOING TO THE CATHETERIZATION LAB. HOWEVER, RESULTS HAD NOT RETURN. TOUHY BORST WAS LOCKED AND TIGHTENED AND ANGLE MATCHED DRESSING AT 90 CENTIMETERS. THE PATIENT WAS TAKEN TO THE UNIT FOR MANAGEMENT. THE FOLLOWING DAY THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 AND WOULD EXPIRE APPROXIMATELY 17 DAYS LATER. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED LIMB ISCHEMIA COMING IN IS UNCLEAR. THE IMPELLA DEVICE MAY HAVE EXACERBATED THE ISCHEMIA. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME (DEMISE). THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21797651&pc=OZD | 1220648-2025-27298 | 2025-03-12 04:00 | Death | ABIOMED, INC. | 2025-04-08 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Connection Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA SUPPORT AND DURING SUPPORT, THE AUTOMATED IMPELLA CONTROLLER (AIC) WAS NOT CHARGING WHEN PLUGGED INTO POWER. THE AIC PLUG WAS NOTED TO HAVE A VISIBLE FRACTURED WIRE. AIC TO AIC CHANGEOUT WAS PREFORMED WITHOUT ISSUE. THE PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY ARREST AND CARDIOPULMONARY RESUSCITATION WAS PREFORMED. HOWEVER, RETURN OF SPONTANEOUS CIRCULATION WAS UNABLE TO BE ACHIEVED. THE PATIENT EXPIRED AND THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: H.6 CODE 4647 WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27298. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21798874&pc=DQY | 2135147-2025-01846 | 2025-03-14 04:00 | Death | ABBOTT MEDICAL | 2025-04-08 04:00 | DQY | AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | K163000 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 20MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN FOR IMPLANT USING A 14F AMPLATZER TORQVUE 45X45 DELIVERY SHEATH. DURING PROCEDURE, THE ACTIVATED CLOTTING LEVELS WERE 283S AND THE VENOUS ACCESS WAS UNEVENTFUL, UPON GETTING TRANSEPTAL THE PHYSICIAN FLOSSED THE SEPTUM 7 TIMES WITH THE AMPLATZER DELIVERY SHEATH OVER A GUIDEWIRE. THE PHYSICIAN THEN BROUGHT TORQVUE SHEATH INTO THE RIGHT ATRIUM AND PASSED ICE CATHETER INTO THE LEFT ATRIUM, ONCE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WAS IN LEFT ATRIUM A LARGE LEFT SIDED CIRCUMFERENTIAL PERICARDIAL EFFUSION WAS NOTED ON ICE. A CARDIAC TAMPONADE WAS ALSO PRESENT. A PERICARDIOCENTESIS WAS PERFORMED TO STABILIZE PATIENT BUT, PATIENT LOST PULSES TWICE AND CARDIOPULMONARY RESUSCITATION (CPR) WAS DONE FOR APPROXIMATELY 6 MINUTES. THE TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) PROBE WAS PASSED INTO PATIENTS ESOPHAGUS AND EFFUSION WAS NOTED WITH LITTLE TO NO PULSE. THEY CONFIRMED EFFUSION WAS COMING FROM A HOLE PUNCTURED IN THE LEFT ATRIAL APPENDAGE, PHYSICIAN DECIDED TO CONTINUE WITH AMULET IMPLANT IN ORDER TO TRY TO SEAL OFF EFFUSION. THE 20MM AMULET DEVICE WAS ATTEMPTED BUT RESULTED IN LEAK AROUND LOBE AND CONTINUOUS FLOW OUT OF HOLE IN LEFT ATRIAL APPENDAGE HOLE. THERE WAS NO OTHER PRODUCT EXPERIENCE NOTED ON THE 20MM AMULET OTHER THAN MIS-SIZING AND THE AMULET WAS REMOVED FROM THE PATIENT PRIOR TO RELEASE FROM THE DELIVERY CABLE. THE 20MM AMULET WAS REMOVED AND A 22MM AMULET DEVICE WAS IMPLANTED WITH THE SAME DELIVERY SYSTEM, THAT SEEMED TO SEAL OFF APPENDAGE AND RESULTED IN NO LONGER WORSENING PERICARDIAL EFFUSION. A RECOMMENDATION WAS MADE TO CHANGE THE DELIVERY SYSTEM BUT DUE TO PATIENT'S STATUS AND IMMEDIATE NEED TO TRY TO SLOW BLEEDING SAME SHEATH WAS USED TO EXPEDITE TRYING TO STABILIZE PATIENT. A CONTINUOUS DRAIN OF PERICARDIAL EFFUSION WAS MAINTAINED AND PATIENT WAS ADMITTED TO INTENSIVE CARE UNIT. THE PHYSICIAN MENTIONED THE CAUSE OF EFFUSION WAS THE FELLOW WHILE FLOSSING THE SEPTUM WITH OUR SHEATH OVER A CIRCULAR GUIDEWIRE MAY HAVE PASSED THE WIRE AND SHEATH/DILATOR INTO THE LEFT ATRIAL APPENDAGE CAUSING A PERFORATION. AFTER THE PROCEDURE, THE PATIENT REPORTED STABLE. ON (B)(6) 2025, THE PATIENT WAS REPORTED PASSED AWAY. THERE WAS NO PERICARDIAL EFFUSION WAS NOTED ON DEATH OR POST LEAVING THE CARDIAC CATHETERIZATION (CATH) LAB BUT PATIENT PASSED AWAY FROM RIGHT VENTRICLE FAILURE (RV) FAILURE.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS ADMINISTERED 6000 UNITS OF HEPARIN FOR PROCEDURE. THE LARGEST DIMENSION OF THE PERICARDIAL EFFUSION WAS UNABLE TO BE OBTAINED DUE TO THE EMERGENCY OF THE SITUATION AND THE PATIENT CODING. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF PERICARDIAL EFFUSION, CARDIAC ARREST, CARDIAC TAMPONADE, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION CAN LEAD TO CARDIAC TAMPONADE, WHICH CAN CAUSE CARDIAC ARREST, APPEARS TO BE RELATED PROCEDURAL CONDITIONS. THE CAUSE OF THE REPORTED DEATH APPEARS TO BE RELATED TO RIGHT VENTRICLE FAILURE. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21798909&pc=OUT | 2029214-2025-00917 | 2025-03-11 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-08 04:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Stenosis; Hemorrhagic Stroke; Ischemia Stroke; Respiratory Arrest; Respiratory Insufficiency | P100018 |  | 1 | Event Description: IT WAS REPORTED THAT THE SUBJECT EXPERIENCED A PH1 HEMORRHAGIC TRANSFORMATION STROKE AFTER MECHANICAL THROMBECTOMY WITH A DECLINE IN NIHSS FROM 6 TO 15 FOLLOWING THE STUDY PROCEDURE INVOLVING THE PIPELINE FLEX SHIELD 3.25MM X 20MM. THE ANEURYSM WAS LOCATED IN THE LEFT HEMISPHERE, VERTEBRAL, AND WAS RUPTURED WITH A FUSIFORM MORPHOLOGY AND A MAXIMUM DIAMETER OF 3MM. THE LANDING ZONE ARTERY SIZE WAS 2.2MM DISTAL AND 3.2MM PROXIMAL. THE PATIENT HAD A MEDICAL HISTORY OF HYPERTENSION, GERD, RHEUMATOID ARTHRITIS, OSTEOARTHRITIS, HYPERLIPIDEMIA, RESPIRATORY DISTRESS, ALTERED MENTAL STATUS, EMESIS, HEADACHE, ABDOMINAL PAIN, NAUSEA, ANEMIA, PREDIABETES, HYDROCEPHALUS, AND ACUTE DEEP VEIN THROMBOSIS OF THE LEFT LOWER EXTREMITY. NO DUAL ANTIPLATELET TREATMENT WAS ADMINISTERED, AND THE ANGIOGRAPHIC RESULT POST-PROCEDURE WAS UNKNOWN. THE EVENT RESULTED IN THE PATIENT'S DEATH ON (B)(6) 2025.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT ON FOLLOW UP DSA IMAGING OBTAINED (B)(6) 2024, THE SUBJECT HAD 75-95% STENOSIS OF THE POSTERIOR INFERIOR CEREBELLAR ARTERY (PICA). THIS WAS A BRANCH ARTERY THAT WAS COVERED BY THE STUDY DEVICE DURING THE INDEX PROCEDURE AND THEREFORE RELATED TO THE STUDY DEVICE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED CORRECTION THAT THE PATIENT DIED ON (B)(6) 2024 DUE TO ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS). ADDITIONAL REPORTED EVENTS FOR THIS SUBJECT: PULMONARY EDEMA, BACTEREMIA, DELAYED CEREBRAL ISCHEMIA, ANEMIA REQUIRING TRANSFUSION, INCREASING RIGHT FRONTOPARIETAL SUBDURAL HEMORRHAGE, INTRAPARENCHYMAL HEMORRHAGE, PICA STENOSIS. NO ACTION OR INTERVENTION WAS TAKEN TO ADDRESS THE OBSERVED STENOSIS. Manufacturer Narrative: B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H6. IME CODING UPDATED BASED ON ADDITIONAL INFORMATION RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21798976&pc=DYB | 2017865-2025-39610 | 2025-03-19 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | DYB | AVEIR¿ INTRODUCER KIT ¿ 50 CM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K211989 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PRESENTED FOR IMPLANT. DURING SURGERY, THE INTRODUCER PERFORATED THROUGH THE HEART STRUCTURE. THE PATIENT'S BLOOD PRESSURE DROPPED AND TAMPONADE OCCURRED. THE PATIENT DECEASED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21799235&pc=DSQ | 2916596-2025-02153 | 2025-03-18 04:00 | Death | THORATEC CORPORATION | 2025-04-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Brain Injury; Syncope/Fainting; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT BEGAN SHOWING LOW FLOWS ON (B)(6) 2025; IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF LOW FLOW ALARMS. THE SUFFERED A STROKE ON (B)(6) 2025. A HEAD COMPUTED TOMOGRAPHY (CT) WAS PERFORMED AND THE STROKE WAS DIAGNOSED AS AN ISCHEMIC STROKE. NO CHANGES THAT MAY HAVE CONTRIBUTED TO THE PATIENT CONDITION OR ANTICOAGULATION STATUS WERE NOTED. THE PATIENT EXPERIENCED UNRESPONSIVENESS, HAD TO BE INTUBATED AND SUFFERED SEVERE BRAIN DAMAGE. THERE WAS NO NEUROLOGICAL INTERVENTION PERFORMED SINCE THE PATIENT WAS DO NOT RESUSCITATE (DNR) AND THERE WOULD BE NO ESCALATION OF CARE. LOG FILES WERE SENT FOR EVALUATION AND REVEALED LOW POWER HAZARDS, POWER CABLE DISCONNECTS AND NO EXTERNAL POWER ALARMS ON (B)(6) 2025 AT APPROXIMATELY 9:14AM WHILE TETHERED ON THE MOBILE POWER UNIT (MPU). THE ALARMS APPEARED TO HAVE BEEN CAUSED BY POWER TO THE MPU BEING BRIEFLY INTERRUPTED, BUT WHETHER THERE WERE ANY ENVIRONMENTAL CIRCUMSTANCES THAT AFFECTED AC POWER AT THE TIME OF THE EVENT WAS UNKNOWN. THERE DID NOT APPEAR TO HAVE BEEN ANY INTERRUPTION IN PUMP SUPPORT. ADDITIONALLY, THE LOG FILES DISCOVERED LOW POWER ADVISORY(S) ON (B)(6) 2025 AT APPROXIMATELY 8:46PM DUE TO NORMAL BATTERY DEPLETION. MULTIPLE STRINGS OF LOW FLOW FLAGS AND OTHER LOW FLOWS EVENTS THAT DID NOT PERSIST LONG ENOUGH TO ACTIVATE AN ALARM ON (B)(6) 2025 WERE CAPTURED, WHICH SEEMED TO HAVE BEEN PHYSIOLOGICAL IN NATURE WHERE THE LOW FLOW ESTIMATOR DROPPED BELOW 2.5 LITERS PER MINUTE. THE SITE DID NOT IDENTIFY A CAUSE FOR THE LOW FLOWS. THE PUMP APPEARED TO HAVE SEEN A REDUCTION IN BLOOD VOLUME. THE REPORTED EVENTS DID NOT APPEAR TO HAVE BEEN THE CAUSE OF ANY TYPE OF MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT ISSUES. THE LOW FLOW EVENTS SEEMED TO HAVE BEEN A PATIENT RELATED ISSUE AND NOT AN EQUIPMENT RELATED ISSUE. THERE WERE NO OTHER UNUSUAL EVENTS SEEN WITHIN THE LOG FILES. ON (B)(6) 2025, THE PATIENT UNDERWENT CARDIAC ARREST AND PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WOULD NOT BE PERFORMED. THE PUMP WAS NOT EXPLANTED. THE MCS EQUIPMENT SEEMED TO HAVE OPERATED AS EXPECTED, AND THE MPU WOULD NOT BE RETURNED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ANALYSIS OF THE LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE EVENTS COULD NOT BE DETERMINED THROUGH THIS EVALUATION. THE SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 17MAR2025 THROUGH 24MAR2025. LOW FLOW HAZARD ALARMS WERE CAPTURED ON (B)(6) 2025. NO OTHER NOTABLE EVENTS OR ALARMS ASSOCIATED WITH THE PUMP WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING STROKE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 6 OF THE IFU (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21799471&pc=NPT | 2025587-2025-02458 | 2025-03-27 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-08 04:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE ATTEMPTED IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO THE PATIENT HAVING A GRADIENT OF OVER 50 MILLIMETERS OF MERCURY (MMHG). DURING THE ATTEMPTED IMPLANT OF THE TRANSCATHETER VALVE, AS THE DELIVERY CATHETER SYSTEM (DCS) WAS ADVANCING, THERE WAS DIFFICULTY NAVIGATING THE AORTIC ARCH. SUBSEQUENTLY, AN AORTIC DISSECTION OCCURRED AND THE PATIENT DIED. IT WAS REPORTED THAT THE DCS CONTRIBUTED TO THE AORTIC DISSECTION, WHICH CAUSED THE PATIENT TO EXPIRE. IT WAS NOTED THAT A 1-HOUR PROCEDURAL DELAY OCCURRED AS A RESULT OF THIS EVENT. PER THE PHYSICIAN, THE PATIENT HAD TORTUOUS ANATOMY WHICH CONTRIBUTED TO THE EVENT. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {{EVFXPLUS-23}}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {{1095-HEART VALVES}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}} PRODUCT ID {{L-EVOLUTFX-2329}}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {{1095-HEART VALVES}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21799507&pc=DSQ | 3007042319-2025-00733 | 2025-04-01 04:00 | Death | HEARTWARE, INC. | 2025-04-08 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT PASSED AWAY. THE DETAILS REGARDING THE CLINICAL CONDITION OR CIRCUMSTANCES SURROUNDING THE DEATH HAVE NOT BEEN PROVIDED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS COULD NOT BE PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AND DETAILS SURROUNDING THE DEATH HAVE NOT BEEN PROVIDED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800506&pc=FKX | 0008030665-2025-00804 | 2025-03-17 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-08 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: A HOME THERAPY PROGRAM MANAGER AND PERITONEAL REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO A CARDIAC ARREST. IT WAS AFFIRMED THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF THE CARDIAC ARREST AND HER SUBSEQUENT DEATH. THE PATIENT¿S CARDIAC ARREST WAS ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOPULMONARY DISEASE THAT WAS UNRELATED TO CCPD THERAPY. IT WAS EXPLAINED THAT THE PATIENT WAS FOUND UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED TO THE HOSPITAL AND PRONOUNCED DECEASED IN HER HOME. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S).Event Description: A HOME THERAPY PROGRAM MANAGER AND PERITONEAL REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO A CARDIAC ARREST. IT WAS AFFIRMED THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF THE CARDIAC ARREST AND HER SUBSEQUENT DEATH. THE PATIENT¿S CARDIAC ARREST WAS ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOPULMONARY DISEASE THAT WAS UNRELATED TO CCPD THERAPY. IT WAS EXPLAINED THAT THE PATIENT WAS FOUND UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED TO THE HOSPITAL AND PRONOUNCED DECEASED IN HER HOME. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO HER DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CAN BE ATTRIBUTED TO A CARDIAC ARREST WITH NO INDICATION THESE EVENTS WERE DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. Manufacturer Narrative: CORRECTION PROVIDED IN G4. ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL NO INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800662&pc=NIK | 2017865-2025-39635 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS NOTED THAT THE PATIENT EXPERIENCED MULTIPLE CARDIAC ARRESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800709&pc=NVY | 2017865-2025-39634 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS NOTED THAT THE PATIENT EXPERIENCED MULTIPLE CARDIAC ARRESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800715&pc=NIK | 2017865-2025-39632 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS NOTED THAT THE PATIENT EXPERIENCED MULTIPLE CARDIAC ARRESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800716&pc=DTB | 2017865-2025-39633 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | DTB | BI,TINED,IS-1, J ,ENDO LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS NOTED THAT THE PATIENT EXPERIENCED MULTIPLE CARDIAC ARRESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800723&pc=NVY | 2649622-2025-09437 | 2025-03-05 05:00 | Death | MPRI | 2025-04-08 04:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Electrical /Electronic Property Problem; Over-Sensing; Under-Sensing; Low impedance; Decreased Sensitivity | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND DIMINISHED SENSING. THE PHYSICIAN NOTED THAT DUE TO THE UNDERSENSING, THE PATIENT MAY NOT HAVE RECEIVED NEEDED THERAPY FOR VENTRICULAR FIBRILLATION (VF). THE PATIENT IS DECEASED.Event Description: IT WAS FURTHER REPORTED THAT ON THE DAY OF THE PATIENT PASSING, THE RV LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR ELEVATED SENSING INTEGRITY COUNTER (SIC) AND HIGH RATE NON-SUSTAINED EPISODES. THE RV LEAD ALSO TRIGGERED AN ALERT FOR LOW PACE/SENSE IMPEDANCE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21800825&pc=ITI | 9615350-2025-00002 | 2025-02-05 05:00 | Death | MOTION CONCEPTS LP | 2025-04-08 04:00 | ITI | POWERED WHEELCHAIR SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Heart Failure/Congestive Heart Failure | K182995 |  | 1 | Event Description: ON MARCH 12, 2025, INVACARE NOTIFIED MOTION CONCEPTS LP OF A INCIDENT INVOLVING ONE OF OUR WHEELCHAIRS, IN WHICH THE END-USER TRAGICALLY PASSED AWAY FROM A HEART ATTACK, AFTER BECOMING PINNED UNDER THE WEIGHT OF THE CHAIR. THE PRODUCT HAD BEEN SOLD TO THE END-USER IN THE UNITED STATES THROUGH THE VAMC - SALISBURY, NC (659/121). AS REPORTED BY INVACARE, THE INCIDENT TOOK PLACE WHILE THE END-USER WAS ATTEMPTING TO DRIVE THE WHEELCHAIR UP A RAMP LEADING INTO HIS RESIDENCE. DURING THE PROCESS, THE WHEELCHAIR REPORTEDLY TIPPED OVER, PINNING THE USER UNDERNEATH. THE END-USER SUFFERED A HEART ATTACK AND SUBSEQUENTLY PASSED AWAY. THE END-USER'S NIECE, WHO WAS PRESENT AT THE TIME OF THE INCIDENT, CALLED EMERGENCY SERVICES. IT WAS FURTHER REPORTED THAT AN AUTOPSY WAS CONDUCTED. AT THE TIME OF THE EVENT, THE SYSTEM HAD BEEN IN USE FOR APPROXIMATELY FIVE MONTHS. Manufacturer Narrative: THE MENTIONED SYSTEM WAS SHIPPED FROM MOTION CONCEPTS TO (B)(4) ((B)(4) IMPORTER/DISTRIBUTOR), AS PER DEALER REQUIREMENTS, ON 18-OCT-2024. THE SYSTEM WAS SUBSEQUENTLY SOLD TO THE (B)(6) ON NOVEMBER 19, 2024. A REVIEW OF THE DEVICE HISTORY CONFIRMED THAT NO ISSUES OR CONCERNS HAD BEEN REPORTED BY EITHER THE DEALER OR THE END-USER, PRIOR TO THE REPORTED INCIDENT. MOTION CONCEPTS BECAME AWARE OF THE INCIDENT WHEN AN (B)(6) REPRESENTATIVE VISITED THE END-USER'S RESIDENCE, BY SCHEDULED APPOINTMENT, TO RETURN THE USER'S PREVIOUS WHEELCHAIR. DURING THE VISIT, THE END-USER'S NIECE INFORMED THE REPRESENTATIVE ABOUT THE INCIDENT AND SPECIFICALLY REQUESTED THAT THE SYSTEM INVOLVED IN THE INCIDENT NOT BE HANDLED. DESPITE MOTION CONCEPTS REQUESTING FOR FURTHER DETAILS, NO ADDITIONAL INFORMATION WAS PROVIDED. THIS LACK OF INFORMATION HAS LIMITED OUR ABILITY TO PERFORM AN EVALUATION OR DETERMINE WHETHER THE INCIDENT WAS RELATED TO A SYSTEM MALFUNCTION. THE (B)(6) REPRESENTATIVE CAPTURED SEVERAL PHOTOS OF THE WHEELCHAIR RAMP AT THE RESIDENCE, WHICH WERE SHARED WITH MOTION CONCEPTS WHEN (B)(6) REPORTED THE INCIDENT. UPON A REVIEW OF THE IMAGES, IT APPEARS THAT THE RAMP WAS CONSTRUCTED FROM ORIENTED STRAND BOARD AND WAS LACKING A CURB RAMPS OR CURB CUTS, WHICH IS AN ESSENTIAL DESIGN SAFEGUARD TO HELP TO PREVENT A WHEELCHAIR FROM ACCIDENTALLY TIPPING OVER OR DRIVING OFF THE EDGE. BASED ON THE INFORMATION MADE AVAILABLE TO MOTION CONCEPTS, THERE IS NO INDICATION THAT THERE WAS A SYSTEM MALFUNCTION. HOWEVER, SINCE THE INCIDENT INVOLVED ONE OF OUR WHEELCHAIRS, WE ARE CONSIDERING THIS AS A REPORTABLE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801037&pc=LWS | 3004209178-2025-06089 | 2024-12-26 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-08 04:00 | LWS | EVERA MRI XT DR SURESCAN | Device Sensing Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT RECEIVED A SINGLE SHOCK FROM THEIR DEVICE AT A TIME WHEN IT WAS SUSPECTED BY A REPRESENTATIVE OF THE PATIENT THAT THEY SHOULD HAVE RECEIVED ADDITIONAL THERAPY FROM THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT ID 5076-52; (SERIAL: (B)(6)); PRODUCT TYPE: 0270-LEAD-LV-PACE; IMPLANT DATE (B)(6) 2018. PRODUCT ID 6935M55; (SERIAL: (B)(6)); PRODUCT TYPE: 0264-LEAD-HV; IMPLANT DATE (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801377&pc=LYJ | 1644487-2025-10019 | 2025-03-08 05:00 | Death | LIVANOVA USA, INC. | 2025-04-08 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY AT HOME AFTER GOING TO BED AS USUAL THE NIGHT PRIOR. THE PATIENT WAS PROCLAIMED DEAD AT 7:03 AM WITH NO TRAUMA OR INJURY WHICH GAVE THE IMPRESSION THAT THE PATIENT PASSED ¿NATURALLY¿. THE MEDICAL EXAMINER PREDICTS SOMETHING MAY HAVE HAPPENED BETWEEN 5:30 AM TO 7:03 AM SINCE THE PATIENT HAD REPORTEDLY WOKEN UP AND GONE BACK TO SLEEP BETWEEN THOSE TIMES. THE MEDICAL EXAMINER FURTHER EXPLAINED THEY DID NOT OBSERVE THE SIGNS THEY USUALLY ASSOCIATE WITH SEIZURE DEATHS WHEN PERFORMING THE INVESTIGATION BUT CLARIFIED THAT THIS DOES NOT DETERMINE THE PATIENT DID NOT HAVE A SEIZURE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S GENERATOR AND LEAD WERE EXPLANTED AFTER DEATH. PER THE MEDICAL EXAMINER, THE BELIEVED CAUSE OF DEATH IS UNKNOWN, PENDING INVESTIGATION. THE SUSPECT DEVICE HAS NOT BEEN RECEIVED TO DATE.Event Description: THE SUSPECT DEVICE HAS BEEN RECEIVED AND IS UNDERGOING PRODUCT ANALYSIS.Event Description: AN INTERROGATION, A SYSTEM DIAGNOSTIC TEST, THE OUTPUT SIGNAL WAS MONITORED FOR MORE THAN 24-HRS, AND A COMPREHENSIVE AUTOMATED ELECTRICAL EVALUATION (SHOWS AN IFI=NO CONDITION) WERE PERFORMED. NO ANOMALIES WERE SEEN, AND THE DEVICE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. PA WORKSHEET WAS REVIEWED, AND NO ANOMALIES WERE SEEN. WANDCOMM DATA REVIEWED, AND NO ANOMALIES WERE SEEN. DATA DUMP WAS REVIEWED, AND NO ANOMALIES WERE SEEN. Manufacturer Narrative: H3 CODE 02 - ANALYSIS IS OF THE SUSPECT DEVICE IS UNDERWAY. H6, ADVERSE EVENT PROBLEM CODES, CORRECTED DATA: SUPPLEMENTAL REPORT #1 INADVERTENTLY SUBMITTED WITH THE INCORRECT HEALTH EFFECT IMPACT CODE. Manufacturer Narrative: H11, ADDITIONAL MANUFACTURER NARRATIVE, CORRECTED DATA: INITIAL REPORT INADVERTENTLY SUBMITTED WITHOUT ADDITIONAL MANUFACTURER NARRATIVE VERBIAGE. LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801381&pc=LWP | 2017865-2025-39397 | 2025-03-18 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | LWP | ENDURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN, ALTHOUGH THE PATIENT EXPERIENCED KIDNEY FAILURE AND NEEDED A VENTILATOR BEFORE PASSING AWAY. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801384&pc=NVY | 2017865-2025-39398 | 2025-03-18 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | NVY | ISOFLEX OPTIM LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960030 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN, ALTHOUGH THE PATIENT EXPERIENCED KIDNEY FAILURE AND NEEDED A VENTILATOR BEFORE PASSING AWAY. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801385&pc=NVN | 2017865-2025-39400 | 2025-03-18 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-08 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN, ALTHOUGH THE PATIENT EXPERIENCED KIDNEY FAILURE AND NEEDED A VENTILATOR BEFORE PASSING AWAY. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801705&pc=FKX | 0008030665-2025-00810 | 2025-03-22 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-08 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) NURSE REPORTED THAT A PD PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. REPORTEDLY, THE CYCLER WAS NOT ACTIVELY RUNNING. THE PATIENT¿S TREATMENT HAD BEEN COMPLETED. THE PATIENT WAS FOUND BY SPOUSE DECEASED IN THE BEDROOM AT APPROXIMATELY 7AM. IN A FOLLOW UP, ADDITIONAL DETAILS VERIFIED THAT THE PATIENT WAS FOUND DECEASED AT APPROXIMATELY 7AM ON (B)(6) 2025. THE PATIENT¿S SPOUSE FOUND PATIENT COLD AND PULSELESS. THE SPOUSE ADMINISTERED CARDIOPULMONARY RESUSCITATION (CPR) UNTIL THE RESCUE SQUAD ARRIVED. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME AND TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S TREATMENT WAS RECORDED AS COMPLETED AT 5:29AM, HOWEVER; THE PATIENT HAD NOT DISCONNECTED FROM THE LIBERTY SELECT CYCLER. THE EXACT TIME OF DEATH WAS NOT REPORTED. NO ISSUES WERE REPORTED WITH THE CYCLER OR THE PATIENT¿S TREATMENT. THERE WERE 0 DRAIN ALARMS AND 0 FILL ALARMS PER THE TREATMENT SHEET. THE MACHINE HAD A HISTORY OF 5 ALARMS (UNKNOWN TYPE OR DATE). THE PATIENT HAD REPORTED SLOW DRAINS PREVIOUSLY. THE CAUSE OF DEATH WAS DOCUMENTED AS CARDIAC ARREST PER THE DOCTOR. THE DEATH WAS NOT RELATED TO ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY.Event Description: A PERITONEAL DIALYSIS (PD) NURSE REPORTED THAT A PD PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. REPORTEDLY, THE CYCLER WAS NOT ACTIVELY RUNNING. THE PATIENT¿S TREATMENT HAD BEEN COMPLETED. THE PATIENT WAS FOUND BY SPOUSE DECEASED IN THE BEDROOM AT APPROXIMATELY 7AM. IN A FOLLOW UP, ADDITIONAL DETAILS VERIFIED THAT THE PATIENT WAS FOUND DECEASED AT APPROXIMATELY 7AM ON (B)(6) 2025. THE PATIENT¿S SPOUSE FOUND PATIENT COLD AND PULSELESS. THE SPOUSE ADMINISTERED CARDIOPULMONARY RESUSCITATION (CPR) UNTIL THE RESCUE SQUAD ARRIVED. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME AND TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S TREATMENT WAS RECORDED AS COMPLETED AT 5:29AM, HOWEVER; THE PATIENT HAD NOT DISCONNECTED FROM THE LIBERTY SELECT CYCLER. THE EXACT TIME OF DEATH WAS NOT REPORTED. NO ISSUES WERE REPORTED WITH THE CYCLER OR THE PATIENT¿S TREATMENT. THERE WERE 0 DRAIN ALARMS AND 0 FILL ALARMS PER THE TREATMENT SHEET. THE MACHINE HAD A HISTORY OF 5 ALARMS (UNKNOWN TYPE OR DATE). THE PATIENT HAD REPORTED SLOW DRAINS PREVIOUSLY. THE CAUSE OF DEATH WAS DOCUMENTED AS CARDIAC ARREST PER THE DOCTOR. THE DEATH WAS NOT RELATED TO ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. THE PATIENT WAS FOUND DECEASED WHILE STILL CONNECTED TO THE LIBERTY SELECT CYCLER, SO IT IS UNKNOWN AT WHAT TIME THE PATIENT EXPIRED. HOWEVER, THE PATIENT DID COMPLETE A FULL TREATMENT PRIOR TO THE DEATH. THERE WERE NO REPORTED ISSUES WITH THE TREATMENT AND NO DOCUMENTED FILL OR DRAIN ALARMS DURING THE TREATMENT. THE PATIENT¿S CAUSE OF DEATH WAS CARDIAC ARREST. DIALYSIS PATIENTS ARE AT EXTRAORDINARILY HIGH RISK FOR DEATH WITH CARDIAC DISEASE BEING THE MAJOR CAUSE OF DEATH AND THE SINGLE, LARGEST, SPECIFIC CAUSE OF DEATH BEING ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA). THE PATIENT HAD A HISTORY OF CARDIAC DISEASE INCLUDING ATRIAL FIBRILLATION, HYPERTENSION, CONGESTIVE HEART FAILURE AND CORONARY ARTERY DISEASE. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT ANY FAILURES OR PROBLEMS. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. SYSTEM AIR LEAK TEST PASSED. TEACH PUMP TEST PASSED. VALVE ACTUATION TEST PASSED. AN INTERNAL VISUAL OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801810&pc=NRY | 2029214-2025-00922 | 2024-01-12 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-08 04:00 | NRY | SOLITAIRE FR4 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Ischemia Stroke | K183022 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: "BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE." THE OBJECTIVE IS TO PRESENT A CASE OF MECHANICAL THROMBECTOMY AND SIMULTANEOUS DISTAL VERTEBRAL ARTERY EMBOLIZATION FOR BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 71 CATHETER, SOLITAIRE X, AND PHENOM 17 CATHETER. AMONG PATIENT ADVERSE EVENTS INCLUDED: THE STUDY DESCRIBES A CASE INVOLVING A PATIENT WHO EXPERIENCED A FALL AND SUBSEQUENT MEDICAL INTERVENTIONS. THE PATIENT DEVELOPED ALTERED CONSCIOUSNESS SIX HOURS AFTER THE CERVICAL SPINE TRAUMA AND HAD A RELATIVELY EARLY ONSET OF BASILAR ARTERY OCCLUSION CAUSED BY VERTEBRAL ARTERY INJURY. MECHANICAL THROMBECTOMY WAS PERFORMED USING THE REACT AND SOLITAIRE DEVICE, AND EMBOLIZATION OF THE INJURED VERTEBRAL ARTERY AT THE SOURCE OF THE EMBOLISM USING COILS DELIVERED THROUGH THE PHENOM CATHETER. COMPLETE RECANALIZATION WAS ACHIEVED DURING THE THROMBECTOMY. NO IMPROVEMENT IN NEUROLOGICAL SYMPTOMS WAS OBSERVED AFTER THE SURGERY. DIFFUSION-WEIGHTED IMAGING PERFORMED 13 H AFTER THE ONSET OF ALTERED CONSCIOUSNESS SHOWED INFARCTION IN THE MIDBRAIN, PONS, BOTH OCCIPITAL LOBES, BOTH CEREBELLAR HEMISPHERES, AND THE LEFT THALAMUS. PATENCY OF THE RIGHT VERTEBRAL AND BASILAR ARTERIES AND OCCLUSION OF THE LEFT VERTEBRAL ARTERY WERE CONFIRMED. GIVEN THE EXTENT OF THE BRAIN INFARCTION, A POOR NEUROLOGICAL AND LIFE PROGNOSIS WAS ANTICIPATED. THE PATIENT DIED TWO DAYS AFTER SURGERY DUE TO HYPOTENSION AND CARDIAC ARREST. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED NO MEDTRONIC DEVICES DIRECTLY RELATED TO THE ADVERSE EVENTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID REACT-71 (UNKNOWN); PRODUCT TYPE: PRODUCT ID (UNKNOWN); PRODUCT TYPE: ; G2: CITATION: AUTHORS: TANIMURA, M., IKEDA, H., FUJIWARA, T., UEZATO, M., OSUKI, T., KINOSADA, M., KUROSAKI, Y., CHIN, M. BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE. SURGICAL NEUROLOGY INTERNATIONAL 15:12 2024. DOI:10.25259/SNI_948_2023 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801853&pc=DQY | 2029214-2025-00923 | 2024-01-12 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-08 04:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Ischemia Stroke | K182097 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: "BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE." THE OBJECTIVE IS TO PRESENT A CASE OF MECHANICAL THROMBECTOMY AND SIMULTANEOUS DISTAL VERTEBRAL ARTERY EMBOLIZATION FOR BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 71 CATHETER, SOLITAIRE X, AND PHENOM 17 CATHETER. AMONG PATIENT ADVERSE EVENTS INCLUDED: THE STUDY DESCRIBES A CASE INVOLVING A PATIENT WHO EXPERIENCED A FALL AND SUBSEQUENT MEDICAL INTERVENTIONS. THE PATIENT DEVELOPED ALTERED CONSCIOUSNESS SIX HOURS AFTER THE CERVICAL SPINE TRAUMA AND HAD A RELATIVELY EARLY ONSET OF BASILAR ARTERY OCCLUSION CAUSED BY VERTEBRAL ARTERY INJURY. MECHANICAL THROMBECTOMY WAS PERFORMED USING THE REACT AND SOLITAIRE DEVICE, AND EMBOLIZATION OF THE INJURED VERTEBRAL ARTERY AT THE SOURCE OF THE EMBOLISM USING COILS DELIVERED THROUGH THE PHENOM CATHETER. COMPLETE RECANALIZATION WAS ACHIEVED DURING THE THROMBECTOMY. NO IMPROVEMENT IN NEUROLOGICAL SYMPTOMS WAS OBSERVED AFTER THE SURGERY. DIFFUSION-WEIGHTED IMAGING PERFORMED 13 H AFTER THE ONSET OF ALTERED CONSCIOUSNESS SHOWED INFARCTION IN THE MIDBRAIN, PONS, BOTH OCCIPITAL LOBES, BOTH CEREBELLAR HEMISPHERES, AND THE LEFT THALAMUS. PATENCY OF THE RIGHT VERTEBRAL AND BASILAR ARTERIES AND OCCLUSION OF THE LEFT VERTEBRAL ARTERY WERE CONFIRMED. GIVEN THE EXTENT OF THE BRAIN INFARCTION, A POOR NEUROLOGICAL AND LIFE PROGNOSIS WAS ANTICIPATED. THE PATIENT DIED TWO DAYS AFTER SURGERY DUE TO HYPOTENSION AND CARDIAC ARREST. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED NO MEDTRONIC DEVICES DIRECTLY RELATED TO THE ADVERSE EVENTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID SFR4-4-40-10 (UNKNOWN); IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID (UNKNOWN); G2: CITATION: AUTHORS: TANIMURA, M., IKEDA, H., FUJIWARA, T., UEZATO, M., OSUKI, T., KINOSADA, M., KUROSAKI, Y., & CHIN, M. BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE. SURGICAL NEUROLOGY INTERNATIONAL 15:12 2024. DOI:10.25259/SNI_948_2023 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21801906&pc=KRA | 2029214-2025-00924 | 2024-01-12 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-08 04:00 | KRA | CATHERA | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Ischemia Stroke |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: "BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE." THE OBJECTIVE IS TO PRESENT A CASE OF MECHANICAL THROMBECTOMY AND SIMULTANEOUS DISTAL VERTEBRAL ARTERY EMBOLIZATION FOR BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 71 CATHETER, SOLITAIRE X, AND PHENOM 17 CATHETER. AMONG PATIENT ADVERSE EVENTS INCLUDED: THE STUDY DESCRIBES A CASE INVOLVING A PATIENT WHO EXPERIENCED A FALL AND SUBSEQUENT MEDICAL INTERVENTIONS. THE PATIENT DEVELOPED ALTERED CONSCIOUSNESS SIX HOURS AFTER THE CERVICAL SPINE TRAUMA AND HAD A RELATIVELY EARLY ONSET OF BASILAR ARTERY OCCLUSION CAUSED BY VERTEBRAL ARTERY INJURY. MECHANICAL THROMBECTOMY WAS PERFORMED USING THE REACT AND SOLITAIRE DEVICE, AND EMBOLIZATION OF THE INJURED VERTEBRAL ARTERY AT THE SOURCE OF THE EMBOLISM USING COILS DELIVERED THROUGH THE PHENOM CATHETER. COMPLETE RECANALIZATION WAS ACHIEVED DURING THE THROMBECTOMY. NO IMPROVEMENT IN NEUROLOGICAL SYMPTOMS WAS OBSERVED AFTER THE SURGERY. DIFFUSION-WEIGHTED IMAGING PERFORMED 13 H AFTER THE ONSET OF ALTERED CONSCIOUSNESS SHOWED INFARCTION IN THE MIDBRAIN, PONS, BOTH OCCIPITAL LOBES, BOTH CEREBELLAR HEMISPHERES, AND THE LEFT THALAMUS. PATENCY OF THE RIGHT VERTEBRAL AND BASILAR ARTERIES AND OCCLUSION OF THE LEFT VERTEBRAL ARTERY WERE CONFIRMED. GIVEN THE EXTENT OF THE BRAIN INFARCTION, A POOR NEUROLOGICAL AND LIFE PROGNOSIS WAS ANTICIPATED. THE PATIENT DIED TWO DAYS AFTER SURGERY DUE TO HYPOTENSION AND CARDIAC ARREST. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED NO MEDTRONIC DEVICES DIRECTLY RELATED TO THE ADVERSE EVENTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: SFR4-4-40-10; PRODUCT ID: REACT-71; G2: CITATION: AUTHORS: TANIMURA, M., IKEDA, H., FUJIWARA, T., UEZATO, M., OSUKI, T., KINOSADA, M., KUROSAKI, Y., CHIN, M. BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE. SURGICAL NEUROLOGY INTERNATIONAL 15:12 2024. DOI:10.25259/SNI_948_2023 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21802304&pc=OAE | 2029046-2025-01077 | 2025-03-14 04:00 | Death | BIOSENSE WEBSTER INC | 2025-04-08 04:00 | OAE | QDOT MICRO¿ CATHETER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (PERSISTENT) ABLATION WITH A QDOT MICRO¿ CATHETER AND THE PATIENT DIED. AT THE END OF A PERSISTENT ATRIAL FIBRILLATION CASE, AS THE PHYSICIAN WAS PREPARING TO PULL THE CATHETERS OUT, THE PATIENT WAS CARDIOVERTED INTO SINUS RHYTHM AND THE BLOOD PRESSURE DROPPED. ULTRASOUND SHOWED NO EFFUSIONS OR PERFORATIONS. THE PATIENT EXPIRED. THE TEAM WAS TRYING TO FIND A POSSIBLE CAUSE. THE PHYSICIAN WAS LOOKING AT THE POSSIBILITY OF A STROKE BUT IT WAS NOT CONFIRMED. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21802465&pc=NPT | 2025587-2025-02479 | 2025-03-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-08 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem; Central Regurgitation | Low Blood Pressure/ Hypotension; Rupture; Pericardial Effusion; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH MODERATE/LOW CALCIFICATION ON THE NON-CORONARY CUSP (NCC), NO PRE-DILATION WAS PERFORMED. THERE WERE NO COMPLICATIONS DURING THE IMPLANT OF THE VALVE. AFTER THE VALVE IMPLANT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A BLOOD PRESSURE OF 50/30 MMHG. AN ECHOCARDIOGRAM SHOWED A PERICARDIAL EFFUSION AND A PERICARDIAL PUNCTURE WAS PERFORMED. A CONTRAST X-RAY REVEALED AN ANNULAR RUPTURE OF THE NCC. CARDIAC SURGERY WAS PROVIDED.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH MODERATE/LOW CALCIFICATION ON THE NON-CORONARY CUSP (NCC), NO PRE-DILATION WAS PERFORMED. THERE WERE NO COMPLICATIONS DURING THE IMPLANT OF THE VALVE. AFTER THE VALVE IMPLANT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A BLOOD PRESSURE OF 50/30 MMHG. AN ECHOCARDIOGRAM SHOWED A PERICARDIAL EFFUSION AND A PERICARDIAL PUNCTURE WAS PERFORMED. A CONTRAST X-RAY REVEALED AN ANNULAR RUPTURE OF THE NCC. CARDIAC SURGERY WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED TO CORRECT THE PREVIOUSLY REPORTED INFORMATION: AN ANNULAR RUPTURE DID NOT OCCUR. A THORACOTOMY WAS PERFORMED TO ADDRESS THE PERICARDIAL EFFUSION AND A VENTRICAR RUPTURE WAS IDENTIFIED AT THE APEX OF THE HEART. IT WAS NOTED A NON-MEDTRONIC (SYMEDRIX INNOWI) GUIDEWIRE WAS USED FOR THE PROCEDURE.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH MODERATE/LOW CALCIFICATION ON THE NON-CORONARY CUSP (NCC), NO PRE-DILATION WAS PERFORMED. THERE WERE NO COMPLICATIONS DURING THE IMPLANT OF THE VALVE. AFTER THE VALVE IMPLANT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A BLOOD PRESSURE OF 50/30 MMHG. AN ECHOCARDIOGRAM SHOWED A PERICARDIAL EFFUSION AND A PERICARDIAL PUNCTURE WAS PERFORMED. A CONTRAST X-RAY REVEALED AN ANNULAR RUPTURE OF THE NCC. CARDIAC SURGERY WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED TO CORRECT THE PREVIOUSLY REPORTED INFORMATION: AN ANNULAR RUPTURE DID NOT OCCUR. A THORACOTOMY WAS PERFORMED TO ADDRESS THE PERICARDIAL EFFUSION AND A VENTRICAR RUPTURE WAS IDENTIFIED AT THE APEX OF THE HEART. IT WAS NOTED A NON-MEDTRONIC (SYMEDRIX INNOWI) GUIDEWIRE WAS USED FOR THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN; THE VENTRICULAR RUPTURE WAS CAUSED BY THE NON-MEDTRONIC GUIDEWIRE AND NOT THE DELIVERY CATHETER SYSTEM (DCS). THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH MODERATE/LOW CALCIFICATION ON THE NON-CORONARY CUSP (NCC), NO PRE-DILATION WAS PERFORMED. THERE WERE NO COMPLICATIONS DURING THE IMPLANT OF THE VALVE. AFTER THE VALVE IMPLANT, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A BLOOD PRESSURE OF 50/30 MMHG. AN ECHOCARDIOGRAM SHOWED A PERICARDIAL EFFUSION AND A PERICARDIAL PUNCTURE WAS PERFORMED. A CONTRAST X-RAY REVEALED AN ANNULAR RUPTURE OF THE NCC. CARDIAC SURGERY WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED TO CORRECT THE PREVIOUSLY REPORTED INFORMATION: AN ANNULAR RUPTURE DID NOT OCCUR. A THORACOTOMY WAS PERFORMED TO ADDRESS THE PERICARDIAL EFFUSION AND A VENTRICAR RUPTURE WAS IDENTIFIED AT THE APEX OF THE HEART. IT WAS NOTED A NON-MEDTRONIC (SYMEDRIX INNOWI) GUIDEWIRE WAS USED FOR THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE VENTRICULAR RUPTURE WAS CAUSED BY THE NON-MEDTRONIC GUIDEWIRE AND NOT THE DELIVERY CATHETER SYSTEM (DCS). THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DEATH WAS DUE TO THE RUPTURE, AND NEITHER THE VALVE NOR THE DCS CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONALLY, THE HYPOTENSION AND HEMODYNAMIC INSTABILITY WERE RESULTS OF THE REPORTED RUPTURE. Manufacturer Narrative: UPDATED DATA: H6. CORRECTED DATA: A5A AND A5B; CORRECTED FROM BLANK TO NO INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. DATE OF DEATH UPDATED B5. FOURTH PARAGRAPH ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B1. ADV EVNT/PROD PROB B5. A SECOND PARAGRAPH WAS ADDED. H6. ADDITIONAL CODES. PRODUCT ANALYSIS: THE VALVE REMAINS IMPLANTED IN THE PATIENT, AND; THEREFORE, WAS NOT AVAILABLE FOR PRODUCT ANALYSIS. ADDITIONALLY, NO FLUOROSCOPIC IMAGES WERE SUBMITTED OF THE VALVE LOAD INSPECTION AND NO COMPUTED TOMOGRAPHY IMAGES WERE SUBMITTED TO MEDTRONIC FOR REVIEW. HOWEVER, THE EXECUTIVE SUMMARY WAS PROVIDED FOR ANATOMICAL CONSIDERATIONS AND ONE PROCEDURAL IMAGE, A CINE WAS SUBMITTED FOR REVIEW. REVIEW OF THE SUBMITTED IMAGE SHOWED EVIDENCE OF A POSSIBLE RUPTURE OR DISSECTION WITH WIDE OPEN AORTIC INSUFFICIENCY. CONCLUSION: THE INVESTIGATION REMAINS IN PROGRESS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B5, H1, H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21802710&pc=NPT | 2025587-2025-02482 | 2025-02-13 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-08 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 YEARS AND 3 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC AORTIC VALVE CARDIOPULMONARY ARREST OCCURRED SECONDARY TO CORONARY ARTERY DISEASE. SUBSEQUENTLY SUDDEN CARDIAC DEATH OCCURRED. THE PHYSICIAN INDICATED THE EVENT WAS ALSO POSSIBLY RELATED TO THE VALVE. AN AUTOPSY WAS NOT PERFORMED.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THAT A POST-SURGICAL WOUND INFECTION OF THE HIP HAD OCCURRED. THE PATIENT WAS IN REHABILITATION AND WAS RECEIVING INTRAVENOUS ANTIBIOTICS FOR A FEW WEEKS. PRIOR TO THE PATIENT¿S DEATH THE PATIENT WAS REPORTED TO HAVE BEEN FAILING TO THRIVE. AS REPORTED, THIS HAD CONTRIBUTED TO THE PATIENT¿S DEATH. THE PATIENT WAS MADE DNR (DO NOT RESUSCITATE) PRIOR TO THE DEATH, THEREFORE RESUSCITATION EFFORTS WERE NOT MADE. FURTHER DETAILS PERTAINING TO THE VALVE FUNCTION WERE NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21783072&pc=BZD | 2518422-2025-041499 | 2024-09-06 04:00 | Death | RESPIRONICS, INC. | 2025-04-07 04:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF A REP DREAMSTATION AUTO CPAP DEVICE ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21783355&pc=PNJ | 9612164-2025-01753 | 2025-02-04 05:00 | Death | MEDTRONIC IRELAND | 2025-04-07 04:00 | PNJ | MICRA¿ VR2 | Device Alarm System; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P150033 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT THE PATIENT RECEIVED A PACEMAKER SIX WEEKS AGO AND THE PATIENT PASSED AWAY FROM A HEART ATTACK AND THERE WAS NO ALARM FROM THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG).Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT THE PATIENT RECEIVED A PACEMAKER SIX WEEKS AGO AND THE PATIENT PASSED AWAY FROM A HEART ATTACK AND THERE WAS NO ALARM FROM THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). Manufacturer Narrative: PRODUCT EVENT SUMMARY: ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 20-MAY-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: CODING MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21783718&pc=NVZ | 3008973940-2025-02427 | 2024-10-21 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-07 04:00 | NVZ | ATTESTA S DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT TEN DAYS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST, RETROPERITONEAL HEMORRHAGE AND HYPOPLASTIC LEFT HEART SYNDROME. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21783722&pc=DTB | 2649622-2025-09191 | 2024-10-21 04:00 | Death | MPRI | 2025-04-07 04:00 | DTB | CAPSURE EPICARDIAL LEAD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | P950024 |  | 1 | Event Description: IT WAS REPORTED THAT TEN DAYS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST, RETROPERITONEAL HEMORRHAGE AND HYPOPLASTIC LEFT HEART SYNDROME. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21783730&pc=DTB | 2649622-2025-09194 | 2024-10-21 04:00 | Death | MPRI | 2025-04-07 04:00 | DTB | CAPSURE EPICARDIAL LEAD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | P950024 |  | 1 | Event Description: IT WAS REPORTED THAT TEN DAYS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST, RETROPERITONEAL HEMORRHAGE AND HYPOPLASTIC LEFT HEART SYNDROME. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21784038&pc=NVZ | 3008973940-2025-02431 | 2025-03-13 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-07 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED, AND THE PATIENT'S REPRESENTATIVE INDICATED THAT PATIENT DEATH WAS DUE TO THE PATIENT'S IMPLANTABLE PULSE GENERATOR (IPG). CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21784558&pc=MIH | 2017233-2025-06035 | 2025-03-11 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-07 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P040043 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE REPORTED THE FOLLOWING: ON (B)(6) 2025, THE PATIENT WAS IMPLANTED WITH GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM WITH A GORE® TRI-LOBE BALLOON CATHETER TO TREAT A THORACIC DISSECTION WITH AN INTRAMURAL HEMATOMA PRESENT. IT WAS REPORTED THAT AFTER DEPLOYING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313110/(B)(6), IMAGES REVEALED A PROXIMAL TYPE I ENDOLEAK. THE PHYSICIAN CHOSE TO USE A GORE® TRI-LOBE BALLOON CATHETER TO BALLOON THE ENDOLEAK. REPORTEDLY THE PHYSICIAN WAS ¿AGGRESSIVELY¿ BALLOONING THE PROXIMAL ENDOLEAK AND PERFORATED THE AORTA. WHILE RETRACTING THE BALLOON THE PATIENT CODED. THE PHYSICIAN IMPLANTED AN ADDITIONAL GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR313110/(B)(6) TO TREAT THE PERFORATED AORTA. A BLOOD TRANSFUSION WAS REQUIRED. THE PATIENT NEVER RECOVERED FROM THE CODE AND EXPIRED IN THE OPERATING ROOM. REPORTEDLY THE PATIENT HAD A FRIABLE AORTA WITH HISTORY OF CARDIAC ISSUES AND PERIPHERAL VASCULAR DISEASE. DURING THE CODE THE PATIENT DEVELOPED A PLEURAL EFFUSION FROM THE BLEEDING. Manufacturer Narrative: PATIENT MEDICATIONS: ASPIRIN, MOXIFLOXACIN, DOXYCYCLINE, CIPRO, VANCOMYCIN, PERCOCET AND PENICILLIN. THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU) STATES THAT ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO ENDOLEAK, VASCULAR TRAUMA, AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21784692&pc=OZD | 1220648-2025-27260 | 2025-03-13 04:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem; Positioning Problem | Perforation of Vessels | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT EXPERIENCED A LEFT VENTRICLE PERFORATION DURING IMPELLA 5.5 IMPLANT. THE PERFORATION WAS REPAIRED AND THE PUMP WAS PLACED FOR PATIENT SUPPORT. TEN DAYS INTO SUPPORT, THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH THE NOTED COMPLICATION. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ADDED FAILURE MODE (FM) POSITIONING ISSUES SINCE THE PUMP WAS REPORTED TO BE PLACED TOO DEEP. NO PRODUCT OR LOGS WERE RETURNED FOR EVALUATION. CLINICAL DETAILS NOTED THAT THE PHYSICIAN DID NOT USE THE INSERTION TEMPLATE PRIOR TO ATTACHING THE GRAFT DIRECT TO THE AORTA AND IMPELLA CAUSED A VENTRICLE PERFORATION. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS MOST LIKELY A USE RELATED ISSUE AS THE INSERTING PHYSICIAN DID NOT MEASURE AND PLACE THE GRAFT IN THE PROPER POSITION, RESULTING IN THE PUMP BEING PLACED TOO DEEP IN THE VENTRICLE UPON INSERTION. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21784872&pc=GAT | 9612501-2025-00966 | 2025-03-07 05:00 | Death | DAVIS & GECK CARIBE LTD | 2025-04-07 04:00 | GAT | TI-CRON | Expiration Date Error; Adverse Event Without Identified Device or Use Problem | Insufficient Information | K930591 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN OPEN EMERGENCY EXPLORATION OF MEDIASTINUM AND EVACUATION OF PERICARDIAL TAMPONADE, THE SUTURE WAS USED WHILE BEING EXPIRED. THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING AN OPEN EMERGENCY EXPLORATION OF MEDIASTINUM AND EVACUATION OF PERICARDIAL TAMPONADE, THE SUTURE WAS USED WHILE BEING EXPIRED. THE PATIENT DIED. THE PATIENT'S DEATH WAS NOT RELATED TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H6 ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21784888&pc=LWS | 2649622-2025-09221 | 2025-03-24 04:00 | Death | MPRI | 2025-04-07 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) DUE TO HIGH RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). ADDITIONALLY, THE LEAD DISPLAYED T-WAVE OVERSENSING (TWOS). THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT REPROGRAMMING WAS COMPLETED. NINE DAYS LATER IT WAS REPORTED THE PATIENT PASSED AWAY. IT WAS SUSPECTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS FUNCTIONING AS DESIGNED AND THAT TREATMENT OF THE PATIENT'S ARRHYTHMIAS OCCURRED APPROPRIATELY WITHIN THE DETECTIONS ZONES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785016&pc=MIH | 9612164-2025-01774 | 2025-04-01 04:00 | Death | MEDTRONIC IRELAND | 2025-04-07 04:00 | MIH | VALIANT CAPTIVIA - FF | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100040 |  | 1 | Event Description: VALIANT CAPTIVIA STENT GRAFTS WERE IMPLANTED DURING AN EMERGENCY ZONE 2 THORACIC ENDOVASCULAR AORTIC REPAIR OF A RUPTURED THORACIC AORTIC DISSECTION. THE PROXIMAL LANDING ZONE WAS 15 MM, BUT THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS, AND NO ENDOLEAKS WERE OBSERVED. THE PATIENT REMAINED STABLE FOLLOWING THE PROCEDURE. IT WAS REPORTED POSTOPERATIVELY DURING THE NIGHT ON (B)(6)2025, THE PATIENTS CONDITION SUDDENLY DETERIORATED AND THE PATIENT EXPIRED. PER THE PHYSICIAN THE CAUSE OF DEATH CANNOT BE DETERMINED DUE TO A CT SCAN NOT BEING PERFORMED. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: IT WAS SAID THE DEATH WAS PROBABLY RELATED TO THE RUPTURE BUT THE CAUSE IS COMPLETELY UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: V AMC2424C150TJ, SERIAL/LOT #: (B)(6), UBD: 13-JAN-2027, UDI#: (B)(4) MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785167&pc=NFK | MW5168644 | 2025-03-14 04:00 | Death | MEDTRONIC/MOZARC MEDICAL US LLC. | 2025-04-07 04:00 | NFK | CATHETER REPAIR KIT | Disconnection | Hemorrhage/Bleeding; Ventricular Fibrillation; Diminished Pulse Pressure; Asystole |  |  | 1 | Event Description: PER REPORT, VENOUS PORT DISCONNECTED DURING DIALYSIS WHICH LED TO SIGNIFICANT BLOOD LOSS. PATIENT BECAME PULSELESS AND CPR (CARDIOPULMONARY RESUSCITATION) STARTED. NOTED RHYTHM WAS VFIB AND THEN ASYSTOLE. PATIENT UNDERWENT MASSIVE BLOOD TRANSFUSION PROTOCOL. ROSC UNABLE TO BE ACHIEVED. PATIENT PASSED AWAY (B)(6) 2025 1641. DIAGNOSIS FOR USE: LONG TERM DIALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785201&pc=NIK | 9614453-2025-01353 | 2025-04-01 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-07 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Ventricular Fibrillation | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DEVICE SYSTEM UPGRADE PROCEDURE, THE PATIENT'S BLOOD PRESSURE BEGAN TO DROP. PROGRAMMING WAS DONE TO INCREASE THE PROGRAMMED LOWER RATE, AND EPINEPHRINE WAS DELIVERED AND COMPRESSIONS WERE STARTED. THE PATIENT'S PRESSURE RECOVERED TEMPORARILY, BUT DROPPED AGAIN. THE PATIENT STARTED GOING INTO VENTRICULAR FIBRILLATION (VF) AND AN EMERGENCY SHOCK WAS PROVIDED. MULTIPLE MANUAL EMERGENCY SHOCKS WERE PROVIDED THROUGH THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D), BUT THE PATIENT CONTINUED TO REENTER INTO VF AFTER PERIODS OF ASYSTOLE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS RECORDED AS CARDIAC ARREST AND VF. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD IMPLANTED: (B)(6) 2020; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785216&pc=LWS | 2649622-2025-09234 | 2025-04-01 04:00 | Death | MPRI | 2025-04-07 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DEVICE SYSTEM UPGRADE PROCEDURE, THE PATIENT'S BLOOD PRESSURE BEGAN TO DROP. PROGRAMMING WAS DONE TO INCREASE THE PROGRAMMED LOWER RATE, AND EPINEPHRINE WAS DELIVERED AND COMPRESSIONS WERE STARTED. THE PATIENT'S PRESSURE RECOVERED TEMPORARILY BUT DROPPED AGAIN. THE PATIENT STARTED GOING INTO VENTRICULAR FIBRILLATION (VF) AND AN EMERGENCY SHOCK WAS PROVIDED. MULTIPLE MANUAL EMERGENCY SHOCKS WERE PROVIDED THROUGH THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D), BUT THE PATIENT CONTINUED TO REENTER INTO VF AFTER PERIODS OF ASYSTOLE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS RECORDED AS CARDIAC ARREST AND VF. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD IMPLANTED: (B)(6) 2020; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785230&pc=NVN | 2649622-2025-09235 | 2025-04-01 04:00 | Death | MPRI | 2025-04-07 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Ventricular Fibrillation | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DEVICE SYSTEM UPGRADE PROCEDURE, THE PATIENT'S BLOOD PRESSURE BEGAN TO DROP. PROGRAMMING WAS DONE TO INCREASE THE PROGRAMMED LOWER RATE, AND EPINEPHRINE WAS DELIVERED AND COMPRESSIONS WERE STARTED. THE PATIENT'S PRESSURE RECOVERED TEMPORARILY, BUT DROPPED AGAIN. THE PATIENT STARTED GOING INTO VENTRICULAR FIBRILLATION (VF) AND AN EMERGENCY SHOCK WAS PROVIDED. MULTIPLE MANUAL EMERGENCY SHOCKS WERE PROVIDED THROUGH THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D), BUT THE PATIENT CONTINUED TO REENTER INTO VF AFTER PERIODS OF ASYSTOLE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS RECORDED AS CARDIAC ARREST AND VF. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD IMPLANTED: (B)(6) 2020; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785397&pc=BZD | 2518422-2025-041501 | 2024-11-12 05:00 | Death | RESPIRONICS, INC. | 2025-04-07 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785497&pc=PNJ | 9612164-2025-01780 | 2025-03-09 05:00 | Death | MEDTRONIC IRELAND | 2025-04-07 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Pneumonia; Hematuria | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED APPROXIMATELY TWO WEEKS POST LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT. DURING IMPLANT A GOOD THRESHOLD COULD NOT BE ACHIEVED, SO EXPANSION WAS PERFORMED FIVE TIMES. THE LEADLESS IPG WAS IMPLANTED. THE HEMATURIA DID NOT STOP AFTER THE LEADLESS IPG IMPLANTATION AND THEY DEVELOPED ASPIRATION PNEUMONIA AND PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT HAEMATURIA WAS NOT PRESENT PRIOR TO IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785516&pc=MLV | 2135147-2025-01800 | 2015-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-07 04:00 | MLV | AMPLATZER SEPTAL OCCLUDER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P000039 |  | 1 | Event Description: THE ARTICLE, "SURGICAL VERSUS PERCUTANEOUS CLOSURE OF POST-INFARCTION VENTRICULAR SEPTAL RUPTURE; REVIEW OF LITERATURE AND SINGLE-CENTER EXPERIENCE", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY EXPERIENCE IN THE PRESENTATION, MANAGEMENT, AND OUTCOMES OF ACUTE MYOCARDIAL INFARCTION (AMI) PATIENTS COMPLICATED BY POST-INFARCTION VENTRICULAR SEPTAL RUPTURE (PIVSR). DEVICES INCLUDED IN THE STUDY WERE AMPLATZER SEPTAL OCCLUDER, LIFETECH KONAR MULTIFUNCTIONAL OCCLUDER (MFO), AND AMPLATZER MUSCULAR VENTRICULAR SEPTAL DEFECT (VSD) OCCLUDER. THE ARTICLE CONCLUDED THAT PIVSR STILL CARRIES A GRAVE PROGNOSIS. EARLY SURGICAL OR PERCUTANEOUS INTERVENTION SEEMS ASSOCIATED WITH HIGHER MORTALITY, WHILE UPFRONT INSERTION OF IABP FOR A SAFE DEFERRAL OF REPAIR BEYOND THE ACUTE PHASE MAY LEAD TO BETTER OUTCOMES. LARGER RANDOMIZED STUDIES ARE REQUIRED TO DICTATE THE BEST MANAGEMENT. [THE PRIMARY AND CORRESPONDING AUTHOR WAS HOSSAMELDIN HUSSEIN, DEPARTMENT OF CARDIOLOGY, KASR AL-AINY FACULTY OF MEDICINE, CAIRO UNIVERSITY, KASR AL-AINY STREET, CAIRO, EGYPT, WITH CORRESPONDING EMAIL: HOSSAMELDIN.HUSSEIN.24@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM APRIL 2015 TO APRIL 2023. A TOTAL OF 32 PATIENTS WERE INCLUDED IN THE STUDY WHERE 9 UNDERWENT PERCUTANEOUS CLOSURE, OF WHICH 7 RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 65 YEARS AND THE GENDER RATIO WAS EVENLY DISTRIBUTED (16 OUT OF 32 WERE FEMALE). COMORBIDITIES INCLUDED DIABETES MELLITUS, SYSTEMIC HYPERTENSION, DYSLIPIDEMIA, SMOKING, FAMILY HISTORY OF CORONARY ARTERY DISEASE, MYOCARDIAL INFRACTION, MITRAL REGURGITATION, RIGHT VENTRICULAR DYSFUNCTION, VENTRICULAR SEPTAL RUPTURE, AND SINGLE/MULTI VESSEL DISEASE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER SEPTAL OCCLUDER WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES MELLITUS, SYSTEMIC HYPERTENSION, DYSLIPIDEMIA, SMOKING, FAMILY HISTORY OF CORONARY ARTERY DISEASE, MYOCARDIAL INFRACTION, MITRAL REGURGITATION, RIGHT VENTRICULAR DYSFUNCTION, VENTRICULAR SEPTAL RUPTURE, AND SINGLE/MULTI VESSEL DISEASE. COMPLICATIONS REPORTED INCLUDED PROCEDURAL MORTALITY, 30-DAY MORTALITY, IN-HOSPITAL MORTALITY, SURGICAL INTERVENTION (EXPLANT), UNEXPECTED MEDICAL INTERVENTION (DIALYSIS), CARDIOGENIC SHOCK, ATRIAL FIBRILLATION, VENTRICULAR ARRHYTHMIA, HEART BLOCK, RENAL FAILURE, BLEEDING, SEPSIS, VENTRICULAR FIBRILLATION/TACHYCARDIA, DEVICE MIGRATION, DEFORMATION, RESIDUAL SHUNT. THE REPORTED IFU DEVIATION WAS ASSOCIATED WITH IMPLANTING AMPLATZER SEPTAL OCCLUDER THAT DID NOT RECONFIGURE TO ITS ORIGINAL SHAPE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: SURGICAL VERSUS PERCUTANEOUS CLOSURE OF POST-INFARCTION VENTRICULAR SEPTAL RUPTURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SURGICAL VERSUS PERCUTANEOUS CLOSURE OF POST-INFARCTION VENTRICULAR SEPTAL RUPTURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785680&pc=BZD | 2518422-2025-041379 | 2024-06-11 04:00 | Death | RESPIRONICS, INC. | 2025-04-07 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN THIS REPORT IN BOX G: DATE RECEIVED BY MFG, IN BOX H: EVAL METHOD CODE, EVAL RESULT CODE, CONCLUSION CODE HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785745&pc=MLV | 2135147-2025-01803 | 2015-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-07 04:00 | MLV | AMPLATZER MUSCULAR VSD OCCLUDER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P040040 |  | 1 | Event Description: THE ARTICLE, "SURGICAL VERSUS PERCUTANEOUS CLOSURE OF POST-INFARCTION VENTRICULAR SEPTAL RUPTURE; REVIEW OF LITERATURE AND SINGLE-CENTER EXPERIENCE", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY EXPERIENCE IN THE PRESENTATION, MANAGEMENT, AND OUTCOMES OF ACUTE MYOCARDIAL INFARCTION (AMI) PATIENTS COMPLICATED BY POST-INFARCTION VENTRICULAR SEPTAL RUPTURE (PIVSR). DEVICES INCLUDED IN THE STUDY WERE AMPLATZER SEPTAL OCCLUDER, LIFETECH KONAR MULTIFUNCTIONAL OCCLUDER (MFO), AND AMPLATZER MUSCULAR VENTRICULAR SEPTAL DEFECT (VSD) OCCLUDER. THE ARTICLE CONCLUDED THAT PIVSR STILL CARRIES A GRAVE PROGNOSIS. EARLY SURGICAL OR PERCUTANEOUS INTERVENTION SEEMS ASSOCIATED WITH HIGHER MORTALITY, WHILE UPFRONT INSERTION OF IABP FOR A SAFE DEFERRAL OF REPAIR BEYOND THE ACUTE PHASE MAY LEAD TO BETTER OUTCOMES. LARGER RANDOMIZED STUDIES ARE REQUIRED TO DICTATE THE BEST MANAGEMENT. [THE PRIMARY AND CORRESPONDING AUTHOR WAS HOSSAMELDIN HUSSEIN, DEPARTMENT OF CARDIOLOGY, KASR AL-AINY FACULTY OF MEDICINE, CAIRO UNIVERSITY, KASR AL-AINY STREET, CAIRO, EGYPT, WITH CORRESPONDING EMAIL: HOSSAMELDIN.HUSSEIN.24@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM APRIL 2015 TO APRIL 2023. A TOTAL OF 32 PATIENTS WERE INCLUDED IN THE STUDY WHERE 9 UNDERWENT PERCUTANEOUS CLOSURE, OF WHICH 7 RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 65 YEARS AND THE GENDER RATIO WAS EVENLY DISTRIBUTED (16 OUT OF 32 WERE FEMALE). COMORBIDITIES INCLUDED DIABETES MELLITUS, SYSTEMIC HYPERTENSION, DYSLIPIDEMIA, SMOKING, FAMILY HISTORY OF CORONARY ARTERY DISEASE, MYOCARDIAL INFRACTION, MITRAL REGURGITATION, RIGHT VENTRICULAR DYSFUNCTION, VENTRICULAR SEPTAL RUPTURE, AND SINGLE/MULTI VESSEL DISEASE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER MUSCULAR VSD OCCLUDER WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES MELLITUS, SYSTEMIC HYPERTENSION, DYSLIPIDEMIA, SMOKING, FAMILY HISTORY OF CORONARY ARTERY DISEASE, MYOCARDIAL INFRACTION, MITRAL REGURGITATION, RIGHT VENTRICULAR DYSFUNCTION, VENTRICULAR SEPTAL RUPTURE, AND SINGLE/MULTI VESSEL DISEASE. COMPLICATIONS REPORTED INCLUDED PROCEDURAL MORTALITY, 30-DAY MORTALITY, IN-HOSPITAL MORTALITY, SURGICAL INTERVENTION (EXPLANT), UNEXPECTED MEDICAL INTERVENTION (DIALYSIS), CARDIOGENIC SHOCK, ATRIAL FIBRILLATION, VENTRICULAR ARRHYTHMIA, HEART BLOCK, RENAL FAILURE, BLEEDING, SEPSIS, VENTRICULAR FIBRILLATION/TACHYCARDIA, DEVICE MIGRATION, DEFORMATION, RESIDUAL SHUNT. THE REPORTED IFU DEVIATION WAS ASSOCIATED WITH IMPLANTING AMPLATZER MUSCULAR VSD OCCLUDER THAT DID NOT RECONFIGURE TO ITS ORIGINAL SHAPE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SURGICAL VERSUS PERCUTANEOUS CLOSURE OF POST-INFARCTION VENTRICULAR SEPTAL RUPTURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SURGICAL VERSUS PERCUTANEOUS CLOSURE OF POST-INFARCTION VENTRICULAR SEPTAL RUPTURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21785785&pc=NPT | 2025587-2025-02400 | 2025-02-19 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA; Encephalopathy; Intracranial Hemorrhage; Low Blood Pressure/ Hypotension; Hypoxia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, A CEREBRAL INFARCTION OCCURRED.Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, A CEREBRAL INFARCTION OCCURRED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT WAS TRANSFERRED TO A DIFFERENT HOSPITAL FOR REHABILITATION. AT THE HOSPITAL WHERE THE PATIENT WAS TRANSFERRED, THE PATIENT DEVELOPED A DECREASE IN CONSCIOUSNESS DUE TO ASPHYXIATION CAUSED BY CEREBRAL HEMORRHAGE AND HYPOXIC ENCEPHALOPATHY. THE PATIENT WAS TRANSFERRED TO A NEW HOSPITAL WHERE TRACHEAL INTUBATION WAS PERFORMED. NORADRENALINE WAS ADMINISTERED TO TREAT HYPOTENSION, BUT THE PATIENT'S BLOOD PRESSURE COULD NOT BE MAINTAINED. SUBSEQUENTLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS BLEEDING FROM THE CEREBRAL HEMORRHAGE. IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. ADDED SECOND PARAGRAPH. H1. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786050&pc=NPT | 2025587-2025-02404 | 2024-07-02 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Stroke/CVA; Non specific EKG/ECG Changes; Pneumonia; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, A CEREBRAL INFARCTION OCCURRED WHICH RESULTED IN THE PATIENT'S ACTIVITIES OF DAILY LIVING REQUIRING FULL ASSISTANCE. THE PATIENT WAS TRANSFERRED TO A LONG-TERM CARE HOSPITAL. SUBSEQUENTLY, AN UNSPECIFIED SEVERE CONDUCTION DISTURBANCE WAS OBSERVED WHICH REQUIRED THE IMPLANT OF A PERMANENT PACEMAKER. THE PATIENT'S CONDITION DETERIORATED, AND THE PATIENT SUBSEQUENTLY DIED DUE TO PNEUMONIA. PER THE PHYSICIAN, THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CEREBRAL INFARCTION WAS CONFIRMED VIA CONDITION CHECK, AND STROKE SYMPTOMS PRESENTED AFTER THE PATIENT RETURNED TO THEIR ROOM FOLLOWING THE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE. IT WAS FURTHER REPORTED THAT THE TYPE OF CONDUCTION DISTURBANCE OBSERVED WAS COMPLETE HEART BLOCK (CHB). ADDITIONALLY, THE CHB FIRST OCCURRED 3 DAYS FOLLOWING THE IMPLANT OF THE VALVE. Manufacturer Narrative: UPDATED DATA: D4. FULL UDI H11. UPDATED LOT, UBD, AND FULL UDI FOR SYSTEM REPORTED DCS. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329, PRODUCT ID: D-EVOLUTFX-2329, SERIAL/LOT: (B)(6), USE BY DATE: 2025-12-05, UDI: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. H6. H11. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329, PRODUCT ID: D-EVOLUTFX-2329, SERIAL/LOT: UNKNOWN, USE BY DATE: UNKNOWN, UDI: UNKNOWN. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786155&pc=NPT | 2025587-2025-02405 | 2024-10-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED FOR HEART FAILURE. A FTER DIURETIC MEDICATION ADJUSTMENT AND REHABILITATION, THE HEART FAILURE STABILIZED AND THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL. SUBSEQUENTLY, THE PATIENT DIED APPROXIMATELY FOUR MONTHS AFTER THE VALVE IMPLANT. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THERE WAS NO RELATION BETWEEN THE DEATH AND THE DEVICE NOR THE IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4., H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786331&pc=DRA | 2135147-2025-01809 | 2025-03-13 04:00 | Death | ABBOTT MEDICAL | 2025-04-07 04:00 | DRA | MITRACLIP | Difficult to Advance; Physical Resistance/Sticking | Hematoma; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Perforation of Vessels | K190167 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT MITRAL REGURGITATION (MR) WITH A GRADE OF UNKNOWN. IT WAS NOTED DURING THE INSERTION OF THE STEERABLE GUIDE CATHETER (SGC), THE DEVICE WAS GETTING HUNG UP ON THE PERCLOSES THAT WERE DEPLOYED PRIOR TO THE GUIDE INSERTION. THE SGC WAS ABLE TO BE ADVANCE INTO THE RIGHT ATRIUM, BLEEDING WAS OBSERVED AROUND THE SGC AND A HEMATOMA WAS FORMING AT THE SITE. IT WAS DECIDED TO IMPLANT ONE MITRACLIP. THE PATIENT'S BLOOD PRESSURE DECREASED ONCE THE CLIP WAS INTRODUCED. THE SGC WAS THEN REMOVED AND THE ACCESS SITE WAS STITCHED UP. PRESSURE WAS APPLIED. HOWEVER, BLEEDING CONTINUED AND FURTHER STITCHING WAS PERFORMED. THE FEMORAL ILIAC VEIN WAS THEN TIED OFF TO PREVENT BLEEDING. BLEEDING CONTINUED. THE PATIENT EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. THE INVESTIGATION DETERMINED THE REPORTED PHYSICAL RESISTANCE DURING ADVANCING AND DIFFICULT TO ADVANCE APPEAR TO BE DUE TO THE PROCEDURAL CONDITION DUE TO THE PERCLOSES. A CAUSE FOR THE REPORTED PATIENT EFFECTS OF PERFORATION OF VESSELS (RUPTURE) AND DEATH CANNOT BE DETERMINED. THE BLEEDING/HEMORRHAGE WAS DUE TO THE RUPTURE. THE HEMATOMA AND HYPOTENSION APPEAR TO BE CASCADING EFFECTS OF THE BLEEDING. PERFORATION OF VESSELS (RUPTURE), BLEEDING, HEMATOMA, DEATH, AND HYPOTENSION ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES AS PRESSURE WAS APPLIED AND THE VASCULAR OPENED THE PATIENTS GROIN AND TIED OFF THE FEMORAL ILIAC VEIN. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNED FOR EVALUATION, THE DEVICE WAS REPORTEDLY DISCARDED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS AN ESTIMATED DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786506&pc=NPT | 2025587-2025-02408 | 2024-07-22 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Anemia; Bradycardia; Ischemia; Pneumonia; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, ISCHEMIA WAS OBSERVED IN THE LOWER LIMB CAUSED BY PERCUTANEOUS CORONARY INTERVENTION (PCI). 12 DAYS FOLLOWING THE INITIAL IMPLANT PROCEDURE, THE PATIENT WAS HOSPITALIZED FOR CONGESTIVE HEART FAILURE (CHF). 19 DAYS FOLLOWING THE INITIAL IMPLANT PROCEDURE, A SEVERE CONDUCTION DISTURBANCE OCCURRED THAT RESULTED IN THE IMPLANT OF A PERMANENT PACEMAKER. SUBSEQUENTLY, THE PATIENT DIED DUE TO BRADYCARDIA, ANEMIA, COMPLICATION OF PNEUMONIA AND INABILITY TO EAT. IT WAS NOTED THAT DIALYSIS WAS ALSO DIFFICULT. PER THE PHYSICIAN, THE DEATH WAS NOT RELATED TO THE TAVR.Event Description: CORRECTED INFORMATION: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PCI WAS PERFORMED DUE TO ISCHEMIA IN THE LOWER LIMB FOLLOWING THE TAVR PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CONFIRMED THAT THE PCI WAS PERFORMED FOLLOWING THE TAVR PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUT FX DCS; PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012146911); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE. UPDATED DATA: H6. H11. LOT NUMBER ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: {D-EVOLUTFX-2329}; PRODUCT TYPE: {0195-HEART VALVES}. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. H6. H11. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [DELIV SYS D-EVOLUTFX-2329] PRODUCT ID [D-EVOLUTFX-2329] SERIAL/LOT [UNKNOWN] USE BY DATE [IUNKNOWN] UDI [IUNKNOWN]. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786566&pc=NPT | 2025587-2025-02410 | 2025-02-02 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY NINE MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THERE WAS NO RELATION BETWEEN THE VALVE NOR THE IMPLANT PROCEDURE AND THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786932&pc=PYX | 1220648-2025-27268 | 2025-03-08 05:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | PYX | IMPELLA RP FLEX | Pumping Stopped | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS PLACED ON BIPELLA SUPPORT FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT DEVELOPED OOZING AT THE RP INSERTION SITE. PRESSURE WAS HELD AND BLEEDING WAS RESOLVED. ADDITIONALLY, THE WHITE CONNECTOR CABLE HAD SEVERAL TWISTS WITH A SIGNIFICANT KINK BUT WAS CORRECTLY ATTACHED TO THE AUTOMATED IMPELLA CONTROLLER, BUT PLACEMENT SIGNAL NOT RELIABLE ALARM OCCURRED. FLOW CALCULATION WAS DISABLED AND THERE WAS ROTATION OF THE IMPELLAR ICON. IN ADDITION, THE MOTOR CURRENT HAD NOT CHANGED. ECHO PERFORMED LATER IN THE EVENING SHOWED THAT THE RP WAS WORKING WELL. THE PATIENT CONTINUED ON SUPPORT. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786941&pc=NGV | 2135147-2025-01815 | 2025-03-13 04:00 | Death | ABBOTT MEDICAL | 2025-04-07 04:00 | NGV | AMPLATZER AMULET | Patient-Device Incompatibility | Cardiac Arrest; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25 MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING UNKNOWN DELIVERY SYSTEM. SIZING WAS DETERMINED VIA CT AND WAS DETERMINED TO BE 22/25 MM. THE 25 MM AMULET WAS PLACED IN PROXIMAL POSITION. THE DEVICE WAS DEPLOYED BUT CANTED. PARTIAL RECAPTURE PERFORMED AND DEPLOYMENT ATTEMPTED AGAIN. AS THE DEVICE ADVANCED INTO LAA IN BALL CONFIGURATION THERE WAS A SUDDEN JUMP AT THE GROIN. THE DEVICE WAS PULLED BACK AND PATIENT EVALUATED - NO ISSUES OR EFFUSION WERE OBSERVED. IT WAS DECIDED TO THEN ATTEMPT IMPLANT MORE DISTAL INTO THE LAA USING A SMALL 18 MM AMPLATZER AMULET LAA OCCLUDER. PRIOR TO ENTERING THE PATIENT WITH THE 18 MM OCCLUDER, IT WAS THOUGHT A THROMBUS WAS FORMING IN LAA. UPON REVIEW WITH ECHOCARDIOGRAPHY THE THROMBUS SEEMED TO GO AWAY. THE 18 MM AMULET WAS THEN IMPLANTED WITH NO RECAPTURES. THE RESULT WAS EVALUATED UNDER ECHOCARDIOGRAPHY AND AN EFFUSION WAS PRESENT. PERICARDIOCENTESIS WAS PERFORMED AND 1500 CC OF FLUID WAS DRAINED. THE PATIENT WAS THEN SENT TO CARDIOVASCULAR SURGERY WHERE THE LAA LATERAL WALL WAS REPAIRED. A SMALL PUNCTURE SITE WAS PRESENT. THE AMULET WAS LEFT IN PLACE AND SURGERY CONCLUDED. THE PATIENT WAS DOING WELL INITIALLY, BUT ON THE AFTERNOON OF (B)(6) 2025 THE PATIENT DIED. CAUSE OF DEATH IS CURRENTLY UNKNOWN.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING A 14F AMPLATZER TORQVUE 45X45 DELIVERY SYSTEM. TRANSSEPTAL PUNCTURE WAS MID/MID. HEPARIN WAS ADMINISTERED AND ACTIVATED CLOTTING TIME WAS 325 SECONDS. LEFT ATRIAL PRESSURE WAS >12. A WIRE WAS NEVER PLACED IN THE LAA, ALL EXCHANGES PERFORMED IN LEFT SUPERIOR PULMONARY VEIN. SIZING WAS DETERMINED VIA CT AND WAS 22/25MM. THE 25MM AMULET WAS PLACED IN PROXIMAL POSITION. THE DEVICE WAS DEPLOYED BUT CANTED. PARTIAL RECAPTURE PERFORMED AND DEPLOYMENT ATTEMPTED AGAIN. AS THE DEVICE ADVANCED INTO LAA IN BALL CONFIGURATION TENSION BUILT UP ON THE DELIVERY SYSTEM NEAR THE GROIN AND THERE WAS A SUDDEN JUMP AS TENSION RELEASED, CAUSING THE OCCLUDER TO ENTER DEEPER INTO THE LAA THAN EXPECTED OR PAST SILLOUHETTE. THE DEVICE WAS PULLED BACK AND PATIENT EVALUATED - NO ISSUES OR EFFUSION WERE OBSERVED. IT WAS DECIDED TO THEN ATTEMPT IMPLANT MORE DISTAL INTO THE LAA USING A SMALL 18MM AMPLATZER AMULET LAA OCCLUDER UTILIZING THE SAME 14F AMPLATZER TORQVUE DELIVERY SYSTEM. PRIOR TO ENTERING THE PATIENT WITH THE 18MM OCCLUDER, IT WAS THOUGHT A THROMBUS WAS FORMING IN LAA UPON REVIEW WITH ECHOCARDIOGRAPHY IT WAS DETERMINED NOT BE A THROMBUS AND PROBABLY PECTINATE. THE 18MM AMULET WAS THEN IMPLANTED WITH NO RECAPTURES. AS THE RESULT WAS EVALUATED UNDER ECHOCARDIOGRAPHY AND AN EFFUSION WAS PRESENT. PERICARDIOCENTESIS WAS PERFORMED AND 1500 CC OF FLUID WAS DRAINED. THE PATIENT WAS THEN SENT TO CARDIOVASCULAR SURGERY WHERE THE LAA LATERAL WALL WAS REPAIRED. A SMALL PUNCTURE SITE WAS PRESENT. THE AMULET WAS LEFT IN PLACE AND SURGERY CONCLUDED. THE PATIENT WAS DOING WELL INITIALLY, BUT ON THE AFTERNOON OF (B)(6) 2025 24 HOURS AFTER THE PROCEDURE, THE PATIENT CODED AND DIED. CAUSE OF DEATH IS UNKNOWN. PATIENT WAS SAID TO BE IN GOOD CONDITION FOR AMULET PROCEDURE AND OPEN HEART SURGERY. Manufacturer Narrative: AN EVENT OF PATIENT-DEVICE INCOMPATIBILITY (IMPLANT-RELATED TISSUE DAMAGE), PERICARDIAL EFFUSION, CARDIAC ARREST AND PATIENT DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED PATIENT-DEVICE INCOMPATIBILITY COULD NOT BE DETERMINED. THE REPORTED PERICARDIAL EFFUSION AND DEATH COULD BE RELATED TO PROCEDURE. HOWEVER, THIS CANNOT BE CONFIRMED. THE REPORTED INTERVENTION WAS UNDER CASE-SPECIFIC CIRCUMSTANCES AS PERICARDIOCENTESIS WAS PERFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FURTHERMORE, THIS COMPLAINT WAS REVIEWED BY ABBOTT. THE REVIEWER STATED THAT, "THE EXACT CAUSE OF DEATH IS UNKNOWN BUT CONSIDERING THE PROXIMITY TO THE IMPLANT PROCEDURE AND THE NEED FOR OPEN SURGERY TO REPAIR THE LAA IT IS LIKELY THAT THE DEATH MAY BE PROCEDURE RELATED. IT IS NOT POSSIBLE TO DETERMINE THE EXACT CAUSE OF DEATH." CODES ADDED: H6 HEALTH EFFECT CLINICAL CODE: 1762 PERICARDIAL EFFUSION. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786942&pc=DQK | 3018490434-2025-00002 | 2025-02-22 05:00 | Death | BIOFOURMIS SINGAPORE LTE | 2025-04-07 04:00 | DQK | RHYTHM ANALYTICS - SOFTWARE (SAMD) | Incorrect, Inadequate or Imprecise Result or Readings | Tachycardia | K182344 |  | 1 | Event Description: COMPLAINT RECEIVED THAT PATIENT PASSED AWAY ON (B)(6) 2025 WHILE ENROLLED IN A HOSPITAL MONITORING PROGRAM ALLEGEDLY DUE TO MYOCARDIAL INFARCTION OR PULMONARY EMBOLISM. Manufacturer Narrative: PLEASE NOTE THAT THE INVESTIGATION IS STILL IN PROCESS AT THIS TIME. PLEASE NOTE THAT THE INITIAL ATTEMPT TO FILE THE MEDWATCH FORM 3500A FOR THIS EVENT OCCURED ON 27 MARCH, 2025 IN THE FORM OF AN EMAIL TO THE EMDR AT MDRPOLICY@FDA.HHS.GOV. THE HELP TEAM REDIRECTED THE EMAIL AND INFORMED BIOFOURMIS THAT THE REPORT REQUIRED SUBMISSION THROUGH THE ESG WITHIN THE WEBTRADER SYSTEM. THE BIOFOURMIS TEAM INITIATED THE PROCESS TO APPLY FOR AN ACCOUNT UNDER THE WEBTRADER / ESG SYSTEM BUT WAS UNABLE TO COMPLETE THE ACCOUNT REGISTRATION PROCESS WITHIN THE REQUIRED 30 DAYS FROM BECOMING AWARE OF THE EVENT. ON 1 APRIL, THE NEXT GEN ESG WAS MADE AVAILABLE AND THE LEGACY TEST DATABASE SYSTEM WAS NO LONGER ACTIVE REQUIRING ADDITIONAL STEPS TO COMPLETE THE REGISTRATION PROCESS. EMAILS OF THE EXCHANGE BETWEEN THE ESG AND ESGNG SUPPORT TEAM AND BIOFOURMIS ARE AVAILABLE FOR REVIEW. AT THE TIME OF THIS REPORT THAT WAS GENERATED USING THE E-SUBMITTER AND MOST CURRENT APPLICATIONS UTILIZING THE SAME INFORMATION ORIGINALLY SENT TO THE MDR TEAM ON 27 MAR 2025, BIOFOURMIS WAS STILL AWAITING ASSISTANCE FROM THE ESGNG SUPPORT TEAM TO REGISTER AN ACCOUNT WITHIN THE LEGACY ESG PRODUCTION SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21786972&pc=OZD | 1220648-2025-27270 | 2025-03-09 05:00 | Death | ABIOMED, INC | 2025-04-07 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Positioning Problem | Bradycardia; Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT). THE PATIENT WAS CONVERTED WITHOUT DEFIBRILLATION AND AN EPISODE OF BRADYCARDIA WARRANTING ADMINISTRATION OF ATROPINE. SUPPORT WAS CONTINUED. ADDITIONALLY, THE PATIENT HAD ACCESS SITE BLEEDING. PRESSURE WAS APPLIED AND HEPARIN WAS STOPPED TEMPORARILY UNTIL IN THERAPEUTIC RANGE. ANGLE-MATCH DRESSING WAS REINFORCED. SUPPORT WAS CONTINUED. FURTHERMORE, THE IMPELLA WAS SLIGHTLY REPOSITIONED UNDER ECHOCARDIOGRAM HELPING WITH SUCTION ALARMS, BUT THE PATIENT CONTINUOUSLY REQUIRED MORE SUPPORT. THE PATIENT EXPIRED ON SUPPORT. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO THE PATIENTS DEATH. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, LOW PUMP FLOW POSITIONING ISSUES, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DATA LOG REVIEWED: NO MOTOR CURRENT (MC) SPIKE OUTSIDE OF P-LEVEL CHANGES OBSERVED. MANY SUCTION ALARMS OCCURRED ON THE DAY LOW FLOW ISSUE REPORTED. IMPELLA POSITION IN AORTA ALARMS OCCURRED AROUND END OF THE IMPELLA SUPPORT. CLINICAL STATED THE IMPELLA WAS SLIGHTLY REPOSITIONED UNDER ECHOCARDIOGRAM ON (B)(6) HELPING WITH SUCTION ALARMS BUT PATIENT CONTINUOUSLY REQUIRED MORE SUPPORT. UNCLEAR EVIDENCE OBTAINED FROM THE ACCOUNT DATA WHY THE POSITION ISSUE HAPPENED AND WHAT OTHER PATIENT SUPPORT PERFORMED TO ATTEMPT IMPROVE LOW FLOW ISSUE. NO INDICATIONS FROM DATA LOGS OBTAINED. THE CAUSE OF THE ACCESS SITE BLEEDING MOST LIKELY WAS ANTICOAGULATION MANAGEMENT SINCE THE HEPARIN STOPPED TEMPORARILY UNTIL IN THERAPEUTIC RANGE TO SOLVE THE ISSUE. THE CAUSE OF THE LOW PUMP FLOW ISSUE CANNOT BE DETERMINED SINCE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. THE CAUSE OF THE POSITIONING ISSUES CANNOT BE DETERMINED SINCE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787049&pc=NRY | 2029214-2025-00900 | 2024-01-01 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-07 04:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion; Embolism/Embolus |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING 'THE FREQUENCY AND LOCATIONS OF SYSTEMIC EMBOLISM IN LARGE VESSEL OCCLUSION (LVO) STROKE PATIENTS RECEIVING REVASCULARIZATION THERAPY'. THE TIME FRAME OF THIS STUDY WAS 'MAY 2019 TO JULY 2021'. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: SOLITAIRE STENT RETRIEVER AND REACT CATHETER. IT WAS NOT SPECIFIED WHICH DEVICES WERE USED IN THE EVENTS BELOW. DEATHS OCCURRED IN THE STUDY POPULATION. THERE WERE 4 TOTAL CASES OF IN-HOSPITAL DEATH. THE 3-MONTH OUTCOME WAS NOTED TO BE SEVERE SEQUELAE OR DEATH FOR 14 PATIENTS. AMONG PATIENT ADVERSE EVENTS INCLUDED: FOUR PATIENTS WITH SYSTEMIC EMBOLISM EXHIBITED LARGE VESSEL OCCLUSION (LVO) RESULTING FROM EMBOLISM AS THE UNDERLYING MECHANISM. ALL FOUR PATIENTS UNDERWENT THROMBECTOMY, AND SUCCESSFUL RECANALIZATION (MTICI SCORE OF 2B, OR 3) WAS ACHIEVED IN 3 OF THEM. ONE OF THE FOUR PATIENTS DIED IN HOSPITAL, AND TWO EXPERIENCED SEVERE SEQUELAE (3-MONTH MRS SCORE OF 5). MULTIPLE ORGAN EMBOLISM WAS DETECTED IN ONLY 1 PATIENT, WHO HAD SPLEEN, RIGHT KIDNEY, LEFT UPPER LIMB (BRACHIAL ARTERY), AND RIGHT LOWER LIMB (EXTERNAL ILIAC ARTERY) EMBOLISMS. THE PATIENT WAS IN A COMA ON ARRIVAL AND IT WAS NOT CLEAR WHETHER THE EMBOLISM WAS SYMPTOMATIC. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: B3: DATE IS APPROXIMATE. YEAR IS CONFIRMED VALID. CONTINUATION OF D10: PRODUCT ID (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE ; EXPLANT DATE PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE ; EXPLANT DATE G2: CITATION: AUTHORS: SAKUTA, K., YAGUCHI, H., NAKADA, R., MIYAGAWA, S., HASEGAWA, I., OKUNO, K., TESHIGAWARA, A., FUGA, M., SHIMIZU, K., & IGUCHI, Y.. YIELD OF WHOLE BODY COMPUTED TOMOGRAPHY IN HYPER-ACUTE STROKE PATIENTS WITH LARGE VESSEL OCCLUSION. VASCULAR AND ENDOVASCULAR SURGERY 58(3):287-293 2024. DOI:10.1177/15385744231209877. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787063&pc=DQY | 2029214-2025-00901 | 2024-01-01 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-07 04:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion; Embolism/Embolus |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING 'THE FREQUENCY AND LOCATIONS OF SYSTEMIC EMBOLISM IN LARGE VESSEL OCCLUSION (LVO) STROKE PATIENTS RECEIVING REVASCULARIZATION THERAPY'. THE TIME FRAME OF THIS STUDY WAS 'MAY 2019 TO JULY 2021'. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: SOLITAIRE STENT RETRIEVER AND REACT CATHETER. IT WAS NOT SPECIFIED WHICH DEVICES WERE USED IN THE EVENTS BELOW. DEATHS OCCURRED IN THE STUDY POPULATION. THERE WERE 4 TOTAL CASES OF IN-HOSPITAL DEATH. THE 3-MONTH OUTCOME WAS NOTED TO BE SEVERE SEQUELAE OR DEATH FOR 14 PATIENTS. AMONG PATIENT ADVERSE EVENTS INCLUDED: FOUR PATIENTS WITH SYSTEMIC EMBOLISM EXHIBITED LARGE VESSEL OCCLUSION (LVO) RESULTING FROM EMBOLISM AS THE UNDERLYING MECHANISM. ALL FOUR PATIENTS UNDERWENT THROMBECTOMY, AND SUCCESSFUL RECANALIZATION (MTICI SCORE OF 2B, OR 3) WAS ACHIEVED IN 3 OF THEM. ONE OF THE FOUR PATIENTS DIED IN HOSPITAL, AND TWO EXPERIENCED SEVERE SEQUELAE (3-MONTH MRS SCORE OF 5). MULTIPLE ORGAN EMBOLISM WAS DETECTED IN ONLY 1 PATIENT, WHO HAD SPLEEN, RIGHT KIDNEY, LEFT UPPER LIMB (BRACHIAL ARTERY), AND RIGHT LOWER LIMB (EXTERNAL ILIAC ARTERY) EMBOLISMS. THE PATIENT WAS IN A COMA ON ARRIVAL AND IT WAS NOT CLEAR WHETHER THE EMBOLISM WAS SYMPTOMATIC. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: AUTHORS: SAKUTA, K., YAGUCHI, H., NAKADA, R., MIYAGAWA, S., HASEGAWA, I., OKUNO, K., TESHIGAWARA, A., FUGA, M., SHIMIZU, K., IGUCHI, Y. YIELD OF WHOLE BODY COMPUTED TOMOGRAPHY IN HYPER-ACUTE STROKE PATIENTS WITH LARGE VESSEL OCCLUSION. VASCULAR AND ENDOVASCULAR SURGERY 58(3):287-293 2024. DOI:10.1177/15385744231209877 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787092&pc=OZD | 1220648-2025-27276 | 2025-03-12 04:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ATTEMPTED TO BE IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS NOTED THAT IT WAS NOT POSSIBLE TO CROSS THE AORTIC VALVE DUE TO PATIENT'S SEVERE STENOSIS AND PRESENCE OF AN AORTIC VALVE REPLACEMENT SCAFFOLDING. THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA REQUIRING DEFIBRILLATION AND THEN SUBSEQUENT ROUNDS OF CPR AND CODE DRUGS. THE PATIENT REMAINED IN PULSELESS ELECTRICAL ACTIVITY WITH NO PULSE DESPITE 15 MINUTES OF CPR. THE PHYSICIAN WAS SUCCESSFUL AT GETTING GUIDEWIRE AND CATHETER INTO LEFT VENTRICLE, BUT THE IMPELLA CATHETER WOULD NEVER CROSS THE SEVERELY STENOTIC AORTIC VALVE AND THE SCAFFOLDING OF THE AVR HE HAD PREVIOUSLY. PT WAS PRONOUNCED EXPIRED ON THE CATH LAB TABLE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES, ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787093&pc=NIK | 9614453-2025-01366 | 2025-02-13 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-07 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Septic Shock; Heart Failure/Congestive Heart Failure | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM PATIENT EXPERIENCED SEVERE SEPTIC SHOCK AND DIED. IT WAS ALSO NOTED THAT THE PATIENT HAD ACUTE HEART FAILURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787102&pc=OJX | 2649622-2025-09284 | 2025-02-13 05:00 | Death | MPRI | 2025-04-07 04:00 | OJX | ATTAIN PERFORMA S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Septic Shock; Heart Failure/Congestive Heart Failure | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM PATIENT EXPERIENCED SEVERE SEPTIC SHOCK AND DIED. IT WAS ALSO NOTED THAT THE PATIENT HAD ACUTE HEART FAILURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787151&pc=LWS | 2649622-2025-09287 | 2025-02-13 05:00 | Death | MPRI | 2025-04-07 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Septic Shock; Heart Failure/Congestive Heart Failure | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM PATIENT EXPERIENCED SEVERE SEPTIC SHOCK AND DIED. IT WAS ALSO NOTED THAT THE PATIENT HAD ACUTE HEART FAILURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787166&pc=DTB | 2649622-2025-09288 | 2025-02-13 05:00 | Death | MPRI | 2025-04-07 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Septic Shock; Heart Failure/Congestive Heart Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM PATIENT EXPERIENCED SEVERE SEPTIC SHOCK AND DIED. IT WAS ALSO NOTED THAT THE PATIENT HAD ACUTE HEART FAILURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787187&pc=OZD | 1220648-2025-27282 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE 86 YEAR OLD FEMALE PATIENT ENCOUNTERED LIMB ISCHEMIA AND VENTRICULAR FIBRILLATION (VF). A FINAL GROIN SHOT WAS TAKEN TO CHECK FOR DISTAL LIMB FLOW AND FLOW APPEARED TO BE SLUGGISH. THE PATIENT WAS THEN TRANSFERRED TO THE UNIT. THE PATIENT'S DISTAL FLOW IN THE RIGHT LEG WAS CHECKED AND BEDSIDE TEAM WAS UNABLE TO FIND A PULSE SO AN ULTRASOUND OF THE LEG WAS ORDERED. AT THIS TIME, THE DOCTOR DID NOT FEEL A FEM FEM BY PASS WAS A GOOD OPTION AS THE PUMP WAS ONLY PLACED IN RELATION TO THE PATIENT CODING POST CATH. WHILE PREPPING THE PATIENT TO GO BACK TO THE UNIT, THE PATIENT BRADIED DOWN TO THE FOURTIES, ATROPINE WAS ADMINISTERED AND RHYTHM WAS CONVERTED. HOWEVER PATIENT BLOOD PRESSURE WAS 60S SYSTOLIC SO LEVOPHED WAS STARTED, DOBUTAMINE DOSING WAS INCREASED, AND DOPAMINE WAS STARTED IN THE CATH LAB. HOWEVER, THE PATIENT CONVERTED TO V-FIB AND CPR WAS INITIATED. THE PATIENT ULTIMATELY EXPIRED IN THE LAB. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27282 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 COMPONENT CODE 3038 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27282. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787212&pc=DSQ | 2916596-2025-02148 | 2025-02-26 05:00 | Death | THORATEC CORPORATION | 2025-04-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Renal Failure; Sepsis; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND MULTI ORGAN FAILURE. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PRE-IMPLANT THE PATIENT WAS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). AFTER IMPLANT THE PATIENT REQUIRED CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT), AND A RIGHT VENTRICULAR ASSIST DEVICE (RVAD). THE PATIENT DEVELOPED ACINOBACTER & KLEBSIELLA IN ENDOTRACHEAL TUBE. THE PATIENT WAS TREATED WITH ANTIBIOTICS BUT COULD NOT BE SAVED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR THE PUMP WAS EXPLANTED. THE DEVICE OPERATED AS EXPECTED. THE PUMP WOULD NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR IF THE DEVICE WAS EXPLANTED. NO PRODUCT WAS RETURNING FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, SEPSIS, AND INFECTION, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "ONGOING PATIENT ASSESSMENT AND CARE") ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21787827&pc=DSQ | 2916596-2025-01905 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-04-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿THE HEARTMATE 3 DEVICE IN SEVERELY OBESE PATIENTS¿ THAT THE HEARTMATE 3 (HM3) MAY BE RELATED TO DRIVELINE INFECTION, HOSPITAL READMISSIONS, AND DEATH. THIS WAS A RETROSPECTIVE COHORT STUDY CONDUCTED VIA REVIEW OF 193 PATIENT¿S CHARTS. THE PATIENTS RECEIVED A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) BETWEEN 01JUL2016 AND 01JUL2023. PATIENTS WITH A BODY MASS INDEX (BMI) OF LESS THAN 40 WERE CLASSIFIED AS THE CONTROL GROUP (N =166) AND HAD AN AVERAGE BMI OF 27.4. THOSE WITH A BMI OF 40 OR GREATER WERE CLASSIFIED AS SEVERELY OBESE (N=27) AND HAD AN AVERAGE BMI OF 45.0. THERE WAS NO SIGNIFICANT DIFFERENCE IN MORTALITY AT 30 DAYS (P=0.3), 1 YEAR (P=0.8), OR 2 YEARS (P=0.9). Manufacturer Narrative: SECTION B: DATE OF DEATH WAS ESTIMATED AS DATE OF EVENT. SECTION E1: CUSTOMER SITE WAS (B)(6) UNIVERSITY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, DOCUMENT ARE CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01JUL2023 SINCE THE PATIENTS WERE IMPLANTED BETWEEN JULY 1ST, 2016 AND JULY 1ST, 2023. BENJAMIN N. FRENCH, ET AL. JOURNAL OF CARDIAC FAILURE 31.1: 189. HTTP://DX.DOI.ORG/10.1016/J.CARDFAIL.2024.10.030. EMORY UNIVERSITY SCHOOL OF MEDICINE, NASHVILLE, TN, USA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21788299&pc=DSQ | 2916596-2025-01916 | 2025-03-06 05:00 | Death | THORATEC CORPORATION | 2025-04-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT HEART FAILURE (RHF) ON (B)(6) 2025. RHF DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. A DEVICE RELATED ISSUE WAS NOT CONSIDERED TO HAVE CAUSED OR CONTRIBUTED TO THE RHF AS IT WAS PATIENT RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PUMP WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS RIGHT HEART FAILURE AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAD. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21789219&pc=LMG | 2210968-2025-03662 | 2024-07-15 04:00 | Death | ETHICON INC. | 2025-04-07 04:00 | LMG | SURGICEL ABSORBABLE HEMOSTAT UNKNOWN | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | N12159 |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SHIFERAW MY, BALEH AS, GIZAW A, TEKLEMARIAM TL, AKLILU AT, AWEDEW AF, ANLEY DT, MEKURIA BH, YESUF EF, YIGZAW MA, MOLLA HT, ADISE MLDIE A, AWANO MM, TEYM A. PREDICTORS OF OPERATIVE ISCHEMIC CEREBROVASCULAR COMPLICATIONS IN SKULL BASE TUMOR RESECTIONS: EXPERIENCE IN LOW-RESOURCE SETTING. NEUROONCOL PRACT. 2024 JUL 15;11(6):790-802. DOI: 10.1093/NOP/NPAE063. PMID: 39554795; PMCID: PMC11567745. THE AIM OF THIS STUDY IS TO IMPROVE RISK STRATIFICATION, SURGICAL DECISION-MAKING, AND POSTOPERATIVE CARE PROTOCOLS. BETWEEN 2018 TO 2023, A MULTI-CENTRIC RETROSPECTIVE COHORT STUDY WAS CONDUCTED ON THE OCCURRENCE AND PREDICTORS OF ISCHEMIC CEREBROVASCULAR COMPLICATIONS AMONG PATIENTS WHO UNDERWENT SKULL BASE TUMOR RESECTION. IN A RETROSPECTIVE COHORT OF SURGICALLY TREATED 266 PATIENTS WITH SKULL BASE TUMORS, THE MEDIAN AGE OF THE PATIENTS WAS 37 (± IQR = 17) YEARS, INDICATING THAT THE AVERAGE AGE OF THE STUDY POPULATION WAS RELATIVELY YOUNG. ADDITIONALLY, THE MAJORITY OF PATIENTS WERE FEMALE, ACCOUNTING FOR 65% OF THE COHORT. THIS GENDER DISTRIBUTION SUGGESTS A POTENTIAL GENDER-BASED DIFFERENCE IN THE INCIDENCE OR PRESENTATION OF SKULL BASE TUMORS. AROUND 15.8% OF PATIENTS HAD COMORBIDITY. THE COMORBIDITIES ASSESSED INCLUDE HYPERTENSION, DIABETES MELLITUS, HIV, BRONCHIAL ASTHMA, CHRONIC KIDNEY, ISCHEMIC CHRONIC DISEASE, CONGESTIVE HEART DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, AND PRIOR ISCHEMIC STROKE. THE SKULL BASE RECONSTRUCTION COMMONLY UTILIZES MULTI-LAYERED TISSUE RECONSTRUCTION, INCLUDING FAT, NASO-SEPTAL FLAP, OSTEO-CARTILAGE TISSUE, AND HEMOSTATIC AGENTS LIKE GEL FOAM &/SURGICEL. REPORTED COMPLICATIONS ARE THE FOLLOWING: BRAIN CONTUSION (N=?) TREATMENT: NOT REPORTED BRAIN SWELLING (N=?) TREATMENT: NOT REPORTED UNKNOWN EVENT (N=?) TREATMENT: NOT SPECIFIED BUT THERE WERE PATIENTS HAD HEMOSTATIC AGENT USE WITH UNKNOWN CAUSE. UNKNOWN EVENT (N=?) TREATMENT: NOT SPECIFIED BUT THERE WERE PATIENT HAD DIFFICULT HEMOSTASIS WITH UNKNOWN CAUSE. SUBARACHNOID HEMORRHAGE (SAH) (N=?) TREATMENT: NOT REPORTED. ACUTE SUBDURAL HEMATOMA (ASDH) (N=?) TREATMENT: NOT REPORTED. INTRACEREBRAL HEMORRHAGE (ICH) (N=?) TREATMENT: NOT REPORTED. INTRAVENTRICULAR HEMORRHAGE (IVH) (N=?) TREATMENT: NOT REPORTED. EPIDURAL HEMATOMA (EDH) (N=?) TREATMENT: NOT REPORTED. PNEUMOCEPHALUS (N=?) TREATMENT: NOT REPORTED. MASS EFFECT ON CONTROL CT (N=?) TREATMENT: NOT REPORTED. IATROGENIC INJURY (N=?) TREATMENT: NOT REPORTED. VASOSPASM (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR ANTERIOR CEREBRAL ARTERY (ACA) (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR BASILAR ARTERY (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR MAJOR CORTICAL ARTERY (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR ARTERY OF HEUBNER (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR INTERNAL CAROTID ARTERY (ICA) (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR MIDDLE CEREBRAL ARTERY (MCA) (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR MULTIPLE ARTERY TERRITORY (N=?) TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR PERFORATORS (N=?) TREATMENT: NOT REPORTED. CSF LEAK DURING TRANS-SPHENOIDAL SURGERY (N=?) TREATMENT: NOT REPORTED. DEVELOPED MENINGITIS (N=?) TREATMENT: ANTIBIOTICS AND IMPROVED. SUFFERED FROM VASCULAR INSULTS RESULTING TO IN HOSPITAL MORTALITY/DEATH (N=?) TREATMENT: NONE. DIED OF SURGERY RELATED (IATROGENIC) ISCHEMIC STROKE ON THE 4TH POSTOPERATIVE DAY (N=1; 48-YEAR-OLD, FEMALE PATIENT) TREATMENT: NONE. IN CONCLUSION, SURGERY-RELATED ISCHEMIC CEREBROVASCULAR COMPLICATIONS ARE COMMON. THOROUGH PERIOPERATIVE RISK STRATIFICATION AND PROACTIVE TREATMENT PLANNING ARE CRUCIAL TO MITIGATE VASCULAR INSULTS ASSOCIATED WITH IT. IN LOW-RESOURCE SETTINGS, NEUROSURGICAL SERVICES ARE PROVIDED WITHOUT ADVANCED INSTRUMENTS, LEADING TO MORE COMPLICATIONS. THEREFORE, IT IS IMPORTANT TO FOCUS ON IMPROVING NEUROSURGICAL SETUP TO ENHANCE PATIENT OUTCOMES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). H6. COMPONENT CODE: G07002 - DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE MANUFACTURING RECORD EVALUATION CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. 1. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER(S). 2. DOES THE SURGEON BELIEVE THAT ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE PATIENT CONSEQUENCES DESCRIBED IN THE ARTICLE? 3. DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH THE ETHICON PRODUCTS USED IN THIS PROCEDURE? 4. WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED (PRODUCT REMOVED; RE-OPERATION; RE-CLOSURE; PRESCRIPTION STEROIDS; ANTIBIOTICS PRESCRIBED)? IF SO, PLEASE CLARIFY. 5. CAN SPECIFIC PATIENT DEMOGRAPHICS: INITIALS; AGE OR DATE OF BIRTH; BMI; GENDER; PATIENT PRE-EXISTING MEDICAL CONDITIONS (I.E. ALLERGIES, HISTORY OF REACTIONS), ALL CONCOMITANT MEDICATIONS, PAST MEDICAL HISTORY, ANY TREATMENT REQUIRED FOR EVENTS, DOSE, FREQUENCY, AND THERAPY DATES OF STUDY DRUGS BE PROVIDED? CITATION: NEUROSURGICAL REVIEW (2024) 47:885 HTTPS://DOI.ORG/10.1007/S10143-024-03132-2. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). CORRECTED INFORMATION: B1, B2, H1 - THIS STUDY WAS DESIGNED TO LOOK INTO THE CAUSES OF ISCHEMIC CEREBROVASCULAR COMPLICATIONS AFTER SKULL BASE TUMOR RESECTIONS. MULTIPLE FACTORS WERE CONSIDERED AS POSSIBLE CONTRIBUTING FACTORS TO THOSE COMPLICATIONS. ALTHOUGH THE USE OF HEMOSTATIC AGENTS WAS MENTIONED IN THE ARTICLE AS ONE STEP FREQUENTLY USED THE STEPS IN CLOSURE OF THE BASE OF THE SKULL, THE USE OF SUCH AGENTS WAS NOT EVALUATED AS ONE OF THE DEPENDENT OR INDEPENDENT VARIABLES IN THE ULTIMATE DEVELOPMENT OF SUCH POSTOPERATIVE COMPLICATIONS. AS SUCH NO RELATIONSHIP BETWEEN THE USE OF SUCH HEMOSTATIC AGENTS AND THE REPORTED COMPLICATIONS CAN BE DRAWN FROM THIS PUBLICATION. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21789524&pc=MOZ | 9612392-2025-00002 | 2025-03-01 05:00 | Death | SEIKAGAKU CORPORATION | 2025-04-07 04:00 | MOZ | GEL-ONE | Insufficient Information | Insufficient Information | P080020 |  | 1 | Event Description: ON (B)(6) 2025, AN 87-YEAR-OLD FEMALE PATIENT, WHO HAD BEEN PRESCRIBED GEL-ONE, PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE REPORTER. Manufacturer Narrative: THIS IS A DEFINITIVE REPORT. NO DETAILED INFORMATION IS AVAILABLE AS NO CONTACT INFORMATION IS PROVIDED. THIS CASE WAS RECEIVED FROM THE FDA AS THE MESSAGE TILED "MDR REPORT: MW5167799" ADDRESSED TO OUR US AGENT DATED ON MARCH 25, 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21789551&pc=NPT | 2025587-2025-02424 | 2024-11-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, THE PATIENT WAS NOT DISCHARGED FROM THE HOSPITAL. SUBSEQUENTLY, THE PATIENT DEVELOPED GRAM-POSITIVE COCCI SEPSIS, WHICH WAS BELIEVED TO HAVE BEEN CAUSED BY BACTERIAL TRANSLOCATION DUE TO CONSTIPATION. THERE WERE NO FINDINGS SUSPICIOUS OF PROSTHETIC VALVE INFECTIOUS ENDOCARDITIS. ANTIBIOTICS WERE ADMINISTERED, BUT THE PATIENT SUBSEQUENTLY DIED 2 MONTHS FOLLOWING THE TAVR PROCEDURE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS SEPSIS. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} . MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21790107&pc=LMG | 2210968-2025-03665 | 2024-07-15 04:00 | Death | ETHICON INC. | 2025-04-07 04:00 | LMG | SURGICEL ABSORBABLE HEMOSTAT UNKNOWN | Adverse Event Without Identified Device or Use Problem | Cardiovascular Insufficiency | N12159 |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: SHIFERAW MY, BALEH AS, GIZAW A, TEKLEMARIAM TL, AKLILU AT, AWEDEW AF, ANLEY DT, MEKURIA BH, YESUF EF, YIGZAW MA, MOLLA HT, ADISE MLDIE A, AWANO MM, TEYM A. PREDICTORS OF OPERATIVE ISCHEMIC CEREBROVASCULAR COMPLICATIONS IN SKULL BASE TUMOR RESECTIONS: EXPERIENCE IN LOW-RESOURCE SETTING. NEUROONCOL PRACT. 2024 JUL 15;11(6):790-802. DOI: 10.1093/NOP/NPAE063. PMID: 39554795; PMCID: PMC11567745. THE AIM OF THIS STUDY IS TO IMPROVE RISK STRATIFICATION, SURGICAL DECISION-MAKING, AND POSTOPERATIVE CARE PROTOCOLS. BETWEEN 2018 TO 2023, A MULTI-CENTRIC RETROSPECTIVE COHORT STUDY WAS CONDUCTED ON THE OCCURRENCE AND PREDICTORS OF ISCHEMIC CEREBROVASCULAR COMPLICATIONS AMONG PATIENTS WHO UNDERWENT SKULL BASE TUMOR RESECTION. IN A RETROSPECTIVE COHORT OF SURGICALLY TREATED 266 PATIENTS WITH SKULL BASE TUMORS, THE MEDIAN AGE OF THE PATIENTS WAS 37 (± IQR = 17) YEARS, INDICATING THAT THE AVERAGE AGE OF THE STUDY POPULATION WAS RELATIVELY YOUNG. ADDITIONALLY, THE MAJORITY OF PATIENTS WERE FEMALE, ACCOUNTING FOR 65% OF THE COHORT. THIS GENDER DISTRIBUTION SUGGESTS A POTENTIAL GENDER-BASED DIFFERENCE IN THE INCIDENCE OR PRESENTATION OF SKULL BASE TUMORS. AROUND 15.8% OF PATIENTS HAD COMORBIDITY. THE COMORBIDITIES ASSESSED INCLUDE HYPERTENSION, DIABETES MELLITUS, HIV, BRONCHIAL ASTHMA, CHRONIC KIDNEY, ISCHEMIC CHRONIC DISEASE, CONGESTIVE HEART DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, AND PRIOR ISCHEMIC STROKE. THE SKULL BASE RECONSTRUCTION COMMONLY UTILIZES MULTI-LAYERED TISSUE RECONSTRUCTION, INCLUDING FAT, NASO-SEPTAL FLAP, OSTEO-CARTILAGE TISSUE, AND HEMOSTATIC AGENTS LIKE GEL FOAM &/SURGICEL. REPORTED COMPLICATIONS ARE THE FOLLOWING: BRAIN CONTUSION (N=?). TREATMENT: NOT REPORTED. BRAIN SWELLING (N=?). TREATMENT: NOT REPORTED. UNKNOWN EVENT (N=?). TREATMENT: NOT SPECIFIED BUT THERE WERE PATIENTS HAD HEMOSTATIC AGENT USE WITH UNKNOWN CAUSE. UNKNOWN EVENT (N=?) TREATMENT: NOT SPECIFIED BUT THERE WERE PATIENT HAD DIFFICULT HEMOSTASIS WITH UNKNOWN CAUSE. SUBARACHNOID HEMORRHAGE (SAH) (N=?). TREATMENT: NOT REPORTED. ACUTE SUBDURAL HEMATOMA (ASDH) (N=?). TREATMENT: NOT REPORTED. INTRACEREBRAL HEMORRHAGE (ICH) (N=?). TREATMENT: NOT REPORTED. INTRAVENTRICULAR HEMORRHAGE (IVH) (N=?). TREATMENT: NOT REPORTED. EPIDURAL HEMATOMA (EDH) (N=?). TREATMENT: NOT REPORTED. PNEUMOCEPHALUS (N=?). TREATMENT: NOT REPORTED. MASS EFFECT ON CONTROL CT (N=?). TREATMENT: NOT REPORTED. IATROGENIC INJURY (N=?). TREATMENT: NOT REPORTED. VASOSPASM (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR ANTERIOR CEREBRAL ARTERY (ACA) (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR BASILAR ARTERY (N=?). TREATMENT: NOT REPORTED ISCHEMIC VASCULAR MAJOR CORTICAL ARTERY (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR ARTERY OF HEUBNER (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR INTERNAL CAROTID ARTERY (ICA) (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR MIDDLE CEREBRAL ARTERY (MCA) (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR MULTIPLE ARTERY TERRITORY (N=?). TREATMENT: NOT REPORTED. ISCHEMIC VASCULAR PERFORATORS (N=?). TREATMENT: NOT REPORTED. CSF LEAK DURING TRANS-SPHENOIDAL SURGERY (N=?). TREATMENT: NOT REPORTED. DEVELOPED MENINGITIS (N=?). TREATMENT: ANTIBIOTICS AND IMPROVED. SUFFERED FROM VASCULAR INSULTS RESULTING TO IN HOSPITAL MORTALITY/DEATH (N=?). TREATMENT: NONE. DIED OF SURGERY RELATED (IATROGENIC) ISCHEMIC STROKE ON THE 4TH POSTOPERATIVE DAY (N=1; 48 YEAR OLD, FEMALE PATIENT). TREATMENT: NONE. IN CONCLUSION, SURGERY-RELATED ISCHEMIC CEREBROVASCULAR COMPLICATIONS ARE COMMON. THOROUGH PERIOPERATIVE RISK STRATIFICATION AND PROACTIVE TREATMENT PLANNING ARE CRUCIAL TO MITIGATE VASCULAR INSULTS ASSOCIATED WITH IT. IN LOW-RESOURCE SETTINGS, NEUROSURGICAL SERVICES ARE PROVIDED WITHOUT ADVANCED INSTRUMENTS, LEADING TO MORE COMPLICATIONS. THEREFORE, IT IS IMPORTANT TO FOCUS ON IMPROVING NEUROSURGICAL SETUP TO ENHANCE PATIENT OUTCOMES. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). CORRECTED INFORMATION: B1, B2, H1 - THIS STUDY WAS DESIGNED TO LOOK INTO THE CAUSES OF ISCHEMIC CEREBROVASCULAR COMPLICATIONS AFTER SKULL BASE TUMOR RESECTIONS. MULTIPLE FACTORS WERE CONSIDERED AS POSSIBLE CONTRIBUTING FACTORS TO THOSE COMPLICATIONS. ALTHOUGH THE USE OF HEMOSTATIC AGENTS WAS MENTIONED IN THE ARTICLE AS ONE STEP FREQUENTLY USED THE STEPS IN CLOSURE OF THE BASE OF THE SKULL, THE USE OF SUCH AGENTS WAS NOT EVALUATED AS ONE OF THE DEPENDENT OR INDEPENDENT VARIABLES IN THE ULTIMATE DEVELOPMENT OF SUCH POSTOPERATIVE COMPLICATIONS. AS SUCH NO RELATIONSHIP BETWEEN THE USE OF SUCH HEMOSTATIC AGENTS AND THE REPORTED COMPLICATIONS CAN BE DRAWN FROM THIS PUBLICATION. THEREFORE, THIS MEDWATCH REPORT IS NOT REPORTABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT#: (B)(4). H6. COMPONENT CODE: G07002 - DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE MANUFACTURING RECORD EVALUATION CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. 1. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER(S). 2. DOES THE SURGEON BELIEVE THAT ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE PATIENT CONSEQUENCES DESCRIBED IN THE ARTICLE? 3. DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH THE ETHICON PRODUCTS USED IN THIS PROCEDURE? 4. WAS THERE ANY MEDICAL OR SURGICAL INTERVENTION PERFORMED (PRODUCT REMOVED; RE-OPERATION; RE-CLOSURE; PRESCRIPTION STEROIDS; ANTIBIOTICS PRESCRIBED)? IF SO, PLEASE CLARIFY. 5. CAN SPECIFIC PATIENT DEMOGRAPHICS: INITIALS; AGE OR DATE OF BIRTH; BMI; GENDER; PATIENT PRE-EXISTING MEDICAL CONDITIONS (I.E. ALLERGIES, HISTORY OF REACTIONS), ALL CONCOMITANT MEDICATIONS, PAST MEDICAL HISTORY, ANY TREATMENT REQUIRED FOR EVENTS, DOSE, FREQUENCY, AND THERAPY DATES OF STUDY DRUGS BE PROVIDED? CITATION: NEUROSURGICAL REVIEW (2024) 47:885 HTTPS://DOI.ORG/10.1007/S10143-024-03132-2. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21790279&pc=GDW | 2955842-2025-13704 | 2025-03-10 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-07 04:00 | GDW | SUREFORM | Insufficient Information | Hemorrhage/Bleeding; Perforation of Vessels | K183224 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY FOLLOWING NEOADJUVANT IMMUNOTHERAPY, SIGNIFICANT BLEEDING FROM AN UNSPECIFIED VESSEL OCCURRED RESULTING IN CONVERSION TO AN OPEN THORACOTOMY APPROACH. THE PATIENT EXPIRED DURING THE PROCEDURE. THE SURGEON REPORTED THAT THE BLEEDING BEGAN WHILE STAPLING THE APICAL BRANCH OF THE LEFT UPPER LOBE WITH A SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT WITH A WHITE RELOAD ACCESSORY. INITIAL HEMOSTASIS WAS ACHIEVED THROUGH COMPRESSION DURING A CONTROLLED UNDOCKING OF THE DA VINCI SYSTEM. ONCE OPEN, THE SURGEON IDENTIFIED A LARGE HOLE IN A VESSEL. THE PATIENT LOST OUTPUT FROM THE HEART AND MANUAL HEART MASSAGE WAS PERFORMED, BUT THEY WERE UNABLE TO RESTART THE HEART. THE SURGEON REPORTED THAT THE PATIENT MAY HAVE EXPERIENCED A MYOCARDIAL INFARCTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: UPDATED SECTIONS: D9, H2, H3, H6, H11. DEVICE EVALUATION: THE WHITE SUREFORM 45 RELOAD WAS RETURNED FOR FAILURE ANALYSIS. UPON VISUAL INSPECTION ALL STAPLES WERE FIRED CORRECTLY, AND BOTH THE SHUTTLE AND THE SHUTTLE KNIFE WERE IN PROPER CONDITION. NO PRODUCT ISSUE WAS IDENTIFIED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT WAS RECEIVED FOR FAILURE ANALYSIS. THE WHITE RELOAD WAS NOT RETURNED FOR INVESTIGATION WITH THE STAPLER. VISUAL INSPECTION DID NOT FIND EVIDENCE OF ANY PROBLEMS. THE INSTRUMENT WAS TESTED ON AN IN-HOUSE SYSTEM, PASSED ALL TESTING AND WAS FULLY FUNCTIONAL. THE FIRING TEST WAS COMPLETED TWO TIMES IN DIFFERENT POSITIONS AND PASSED EACH TIME. NO PRODUCT ISSUE WAS IDENTIFIED. Manufacturer Narrative: A REVIEW OF THE DA VINCI SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND NO RELATED SYSTEM ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE ONE SUBSEQUENT PROCEDURE PERFORMED ON THE SYSTEM AT THE TIME OF REVIEW FOUND NO SYSTEM ERRORS OCCURRED AND NO COMPLAINTS WERE REPORTED. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, PERMANENT CAUTERY SPATULA, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, SUREFORM 45 CURVED TIP STAPLER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. REVIEW OF THE SUREFORM 45 CURVED-TIP STAPLER LOGS FOUND THAT THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 4 TIMES AND FIRED 4 WHITE RELOADS. THERE WERE NO INCOMPLETE CLAMPS OR UNCLAMPS. ON EACH INSTALL, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE EACH COMPLETED WITH NO PAUSES FOR COMPRESSION. THERE WERE NO STAPLER RELATED ERRORS IN THE MSC LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF BLEEDING FROM AN INJURY TO AN UNSPECIFIED BLOOD VESSEL WHILE ATTEMPTING A PULMONARY LOBECTOMY. THE DETAILS ON HOW THE INJURY MAY HAVE OCCURRED IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21790494&pc=DSQ | 3007042319-2025-00729 | 2025-03-29 04:00 | Death | HEARTWARE, INC. | 2025-04-07 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem; Pumping Problem | Dyspnea; Fatigue; Fever; Intracranial Hemorrhage; Low Blood Pressure/ Hypotension; Pneumonia; Renal Failure; Irregular Pulse; Respiratory Failure; Sleep Dysfunction; Weight Changes; Heart Failure/Congestive Heart Failure; Cough; Decreased Appetite | P100047 |  | 1 | Event Description: DECREASE IN CIRCADIAN RHYTHM AND AN INCREASE IN PULSATILE TROUGH WERE OBSERVED STARTING THE NIGHT OF (B)(6). WHEN INQUIRING ABOUT THE PATIENT'S CONDITION, IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT HAD BEEN HOSPITALIZED FOR TREATMENT DUE TO ILLNESS/POOR PHYSICAL CONDITION. COMPARED TO HOME CARE, THE BLOOD PRESSURE HAS BEEN TRENDING LOWER, BUT THE BLOOD DATA REMAINS LARGELY WITHIN NORMAL RANGES. THE VAD REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT VISITED THE HOSPITAL WITH INITIAL SYMPTOMS OF FEVER, FATIGUE, COUGH, LOSS OF APPETITE, SHORTNESS OF BREATH AND WEIGHT GAIN. THE PATIENT WAS DIAGNOSED WITH WORSENING HEART FAILURE, AND RIGHT SIDED HEART FAILURE ASSOCIATED WITH PNEUMONIA. THE PATIENT WAS TREATED WITH ANTIBIOTICS, DIURETICS, AND INTRAVENOUS INOTROPES. ON APRIL 3, RENAL FAILURE PROGRESSED, AND BLOOD PRESSURE DROPPED, SO THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT. THE PATIENT UNDERWENT CONTINUOUS RENAL REPLACEMENT THERAPY CRRT), BUT SINCE BEEN REMOVED. THE PATIENT'S RESPIRATORY CONDITION WORSENED, SO CONTINUOUS HEMODIAFILTRATION (CHDF) AND VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (V-V ECMO) WERE INTRODUCED. ON APRIL 7TH, THERE WAS A PULMONARY HEMORRHAGE, AND A BRAIN/HEAD COMPUTERIZED TOMOGRAPHY (CT) SCAN REVEALED A SUBARACHNOID HEMORRHAGE. ANOTHER CT SCAN ON APRIL 10TH SHOWED THE EXPANSION OF THE SUBARACHNOID HEMORRHAGE, AND ON APRIL 13TH, PUPIL DILATION WAS CONFIRMED. APPROXIMATELY SIX DAYS LATER THE VAD EXHIBITED MULTIPLE LOW FLOW ALARMS. THE PATIENT SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED INTERMITTENT DECREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS WITHIN THE ANALYZED PERIOD, FOLLOWED BY A SUDDEN DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS ON (B)(6) 2025, LEADING TO PARAMETERS BELOW NORMAL OPERATING RANGE, AND TWO (2) LOW FLOW ALARMS LOGGED ON (B)(6) 2025 AT 09:02:26 AND AT 09:07:51. AS A RESULT, THE REPORTED LOW FLOW EVENT WAS CONFIRMED. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING WORSENING HEART FAILURE, RIGHT VENTRICULAR FAILURE, PNEUMONIA, PULMONARY HEMORRHAGE, RENAL FAILURE AND SUBARACHNOID HEMORRHAGE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, INAPPROPRIATE PUMP ROTATIONAL SPEED, AND/OR POOR VAD FILLING. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION, RENAL FAILURE, BLEEDING, WORSENING HEART FAILURE, RIGHT VENTRICULAR FAILURE, NEUROLOGICAL DYSFUNCTION, INFECTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF BLEEDING AND INFECTION EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDING THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21790569&pc=BZD | 2518422-2025-103739 | 2025-03-07 05:00 | Death | RESPIRONICS, INC. | 2025-04-07 04:00 | BZD | REMSTAR AUTO A-FLEX W/HUMID, SD CARD | Degraded | No Clinical Signs, Symptoms or Conditions | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21791038&pc=LYJ | 1644487-2025-00374 | 2024-12-09 05:00 | Death | LIVANOVA USA, INC. | 2025-04-07 04:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO CONGESTIVE HEART FAILURE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: ADDITIONAL INFORMATION RECEIVED NOTING THAT THE DEATH IS NOT RELATED TO THE VNS DEVICE. THE CAUSE OF DEATH WAS NOTED TO BE RESPIRATORY FAILURE, CONGESTIVE HEART FAILURE, AND THE VNS WAS NOT EXPLANTED AFTER DEATH. IT WAS ALSO NOTED THAT THE VNS PROVIDED SEIZURE REDUCTION FOR THE PATIENT. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21791924&pc=OZD | 1220648-2025-27284 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Display or Visual Feedback Problem; Unexpected Shutdown | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR APPROXIMATELY EIGHT DAYS FOR MECHANICAL CIRCULATORY SUPPORT UNTIL ESCALATED TO AN IMPELLA 5.5 WITH EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) PLACEMENT FOR CONTINUED THERAPY. APPROXIMATELY 41 DAYS AFTER THE IMPELLA 5.5 MCS STARTED, THE AUTOMATED IMPELLA CONTROLLER SUDDENLY WENT BLACK WITHOUT ANY WARNING OR ALARM AND RESTARTED ON ITS OWN. PER THE REPORT, IT TOOK APPROXIMATELY 30 SECONDS FOR THE AIC TO RESTART. AFTER REBOOTING THE AIC WORKED WITHOUT ANY PROBLEMS. HOWEVER, AS A PRECAUTION, THE AIC WAS SUBSEQUENTLY REPLACED. THE PATIENT CONTINUED IMPELLA AND ECMO THERAPIES FOR AN ADDITIONAL 10 DAYS BEFORE EXPIRING NOTED TO BE DUE TO MULTI-ORGAN FAILURE. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE WAS INTERRUPTION WITH THE AUTOMATED IMPELLA CONTROLLER (AIC), AND THE AIC WAS EXCHANGED AS PRECAUTION. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE AIC AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THERE WAS INTERRUPTION (POSSIBLY BRIEF); IMPACT TO THE EVENT OUTCOME IS UNKNOWN. THIS IS ONE OF TWO DEVICES RELATED TO THE PATIENT'S IMPLANT EXPERIENCE. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-27272 FOR THE IMPELLA 5.5 REPORT. THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - DISPLAY ISSUE HAS BEEN COMPLETED. THE CONSOLE LOGS FROM THE REPORTED DAY OF EVENT ARE CONSISTENT WITH COMPLAINT AND SHOW THAT AROUND 13:03 ON (B)(6) 2025 CONSOLE PERFORMED FULL RESET (REBOOT), PRECEDED BY UNRESPONSIVE SOFTWARE PROCESS (CALCULATOR). PRIOR TO REBOOT, THE LOGS INDICATE SYSTEM¿S INABILITY TO PERFORM SOFT-RESET (TO RESTART UNRESPONSIVE PROCESS), AS WELL AS MALFUNCTION OF OTHER SOFTWARE ELEMENTS DEPENDENT ON SYSTEM RESOURCES. ALTHOUGH THE LOG IS INCOMPLETE AND DOES NOT INCLUDE ENTRIES BEFORE 20:32 ON (B)(6) 2025, IT CLEARLY SHOWS INTERMITTENT, AND LATER PERMANENT, LOSS OF DATA STREAMING (USB DATA TRANSMISSION FROM AIC TO RLM) AS EVIDENCED BY MULTIPLE FAILED HANDSHAKE ATTEMPTS OF THE ICB. AFTER THE REBOOT, PUMP (B)(6) RESUMED OPERATION. ANALYSIS OF RLM JOURNAL COULD NOT BE PERFORMED AS THE UNIT WAS UNABLE TO FULLY BOOT-UP AND KEPT REBOOTING. THIS CORRELATES WITH AIC¿S INABILITY TO COMMUNICATE WITH RLM. THERE WAS NO ISSUE FOUND DURING THE CASE. THE DISPLAY FUNCTIONED/OPERATED TO SPECIFICATION. CONSOLE¿S BEHAVIOR IN THE FIELD IS CONSISTENT WITH UNAVAILABILITY OF AIC SYSTEM RESOURCES CAUSED BY RLM¿S PROLONGED ERRATIC BEHAVIOR. THE AIC SOFTWARE BEHAVIOR WAS INVESTIGATED SEPARATELY. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27284 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 COMPONENT CODE 765 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27284.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21792542&pc=NPT | 2025587-2025-02440 | 2025-04-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT PRO PLUS VALVE | Difficult to Insert; Material Invagination; Patient Device Interaction Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE WAS LOADED AND CHECKED UNDER FLUOROSCOPY CONFIRMING A GOOD LOAD. A PRE-IMPLANT BALLOON DILATATION WAS PERFORMED WITH A 26 MILLIMETER (MM) BALLOON. DIFFICULTY INTRODUCING THE DELIVERY CATHETER SYSTEM (DCS) INTO THE PATIENT'S FEMORAL ARTERY. SEVERAL ATTEMPTS AND SKIN OPENING WERE NECESSARY TO MAKE PROGRESS. THERE WAS RESISTANCE NOTED WHEN MOVING THE NOSE CONE INTO THE VENTRICULAR OUTFLOW TRACT AND A LOT OF THE GUIDEWIRE COULD BE SEEN INSIDE OF THE RIGHT VENTRICLE THAT THE IMPLANTING TEAM COULD NOT PULL BACK. THE IMPLANTING TEAM POSITIONED THE VALVE AT THE HEIGHT OF THE AORTIC ANNULUS TO BEGIN OPENING THE DCS. WHEN THE VALVE WAS AT 80% DEPLOYED, THE VALVE WOULD NOT OPEN AND APPEARED TO BE INFOLDED CAUSING INSTABILITY IN THE PATIENT. IT WAS DECIDED TO RETRIEVE THE VALVE TO VISUALLY INSPECT AND AT THAT MOMENT, A PERFORATION TO THE RIGHT VENTRICLE WAS OBSERVED WITH THE GUIDEWIRE USED. INTERVENTION BY THE CARDIAC SURGERY TEAM WAS PERFORMED WITH DRAINAGE OF BLOOD AND SUBSEQUENT OPENING OF CHEST. OF NOTE, THE PATIENT HAD A CALCIFIED ANNULUS AND FUSION OF THE RIGHT AND LEFT L EAFLETS (BICUSPID). SUBSEQUENTLY THE PATIENT DIED AND THE CAUSE OF DEATH WAS NOT REPORTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THE VALVE WAS RECAPTURED TWO TIMES AS THE VALVE DID NOT EXPAND ON THE INFLOW. A PROCEDURAL DELAY RESULTED FROM THE VALVE INFOLD. PER THE PHYSICIAN, THE CAUSE OF THE RIGHT VENTRICLE PERFORATION WAS THE NON-MEDTRONIC GUIDEWIRE (PTC-LUNDERQUIST ¿ SP MEDICAL) AS IT WAS BENT AT THE TRANSITION POINT BETWEEN IT'S SOFT TO HARD TRANSITION. THE IMPLANTING TEAM HAD DIFFICULTY IN ASCENDING THE DELIVERY CATHETER SYSTEM (DCS) DUE TO THE INGUINAL LIGAMENT BUT IT DID NOT PRESENT TORTUOSITY OR SIGNIFICANT CALCIFICATION. PER THE PHYSICIAN, THE DCS DID NOT CAUSE OR CONTRIBUTE TO THE DEATH BUT THE VALVE COULD NOT BE EXCLUDED FROM CONTRIBUTING TO THE DEATH.Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE WAS LOADED AND CHECKED UNDER FLUOROSCOPY CONFIRMING A GOOD LOAD. A PRE IMPLANT BALLOON DILATATION WAS PERFORMED WITH A 26 MILLIMETER (MM) BALLOON. DIFFICULTY INTRODUCING THE DELIVERY CATHETER SYSTEM (DCS) INTO THE PATIENT'S FEMORAL ARTERY. SEVERAL ATTEMPTS AND SKIN OPENING WERE NECESSARY TO MAKE PROGRESS. THERE WAS RESISTANCE NOTED WHEN MOVING THE NOSE CONE INTO THE VENTRICULAR OUTFLOW TRACT AND A LOT OF THE GUIDEWIRE COULD BE SEEN INSIDE OF THE RIGHT VENTRICLE THAT THE IMPLANTING TEAM COULD NOT PULL BACK. THE IMPLANTING TEAM POSITIONED THE VALVE AT THE HEIGHT OF THE AORTIC ANNULUS TO BEGIN OPENING THE DCS. WHEN THE VALVE WAS AT 80% DEPLOYED, THE VALVE WOULD NOT OPEN AND APPEARED TO BE INFOLDED CAUSING INSTABILITY IN THE PATIENT. IT WAS DECIDED TO RETRIEVE THE VALVE TO VISUALLY INSPECT AND AT THAT MOMENT, A PERFORATION TO THE RIGHT VENTRICLE WAS OBSERVED WITH THE GUIDEWIRE USED. INTERVENTION BY THE CARDIAC SURGERY TEAM WAS PERFORMED WITH DRAINAGE OF BLOOD AND SUBSEQUENT OPENING OF CHEST. OF NOTE, THE PATIENT HAD A CALCIFIED ANNULUS AND FUSION OF THE RIGHT AND LEFT L EAFLETS (BICUSPID). SUBSEQUENTLY THE PATIENT DIED AND THE CAUSE OF DEATH WAS NOT REPORTED. ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THE VALVE WAS RECAPTURED TWO TIMES AS THE VALVE DID NOT EXPAND ON THE INFLOW. A PROCEDURAL DELAY RESULTED FROM THE VALVE INFOLD. PER THE PHYSICIAN, THE CAUSE OF THE RIGHT VENTRICLE PERFORATION WAS THE NON-MEDTRONIC GUIDEWIRE (PTC LUNDERQUIST ¿ SP MEDICAL) AS IT WAS BENT AT THE TRANSITION POINT BETWEEN IT'S SOFT TO HARD TRANSITION. THE IMPLANTING TEAM HAD DIFFICULTY IN ASCENDING THE DELIVERY CATHETER SYSTEM (DCS) DUE TO THE INGUINAL LIGAMENT BUT IT DID NOT PRESENT TORTUOSITY OR SIGNIFICANT CALCIFICATION. PER THE PHYSICIAN, THE DCS DID NOT CAUSE OR CONTRIBUTE TO THE DEATH BUT THE VALVE COULD NOT BE EXCLUDED FROM CONTRIBUTING TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PERFORATION OCCURRED IN THE LEFT VENTRICLE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT EXPERIENCED HYPOTENSION AND PERICARDIAL EFFUSION AS A RESULT OF THE PERFORATION. CARDIAC ARREST SUBSEQUENTLY OCCURRED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. THE VALVE WAS WITHDRAWN FROM THE PATIENT, THE SAME VALVE WAS URGENTLY RE-LOADED ONTO THE SAME DELIVERY CATHETER SYSTEM (DCS), AND THE SYSTEM WAS RE-INSERTED INTO THE PATIENT. THE DCS WAS ADVANCED TO THE PATIENT'S ANNULUS AND DEPLOYMENT WAS ATTEMPTED. HOWEVER, THE VALVE WAS NEVER FULLY DEPLOYED. Manufacturer Narrative: UPDATED: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H2 H3 H11 IMAGE REVIEW: ONLY FIVE INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW WHICH APPEARS TO BE INCOMPLETE. HOWEVER, ADDITIONAL IMAGES AND ADDITIONAL DOCUMENTATION WERE PROVIDED TO SUPPORT THE REVIEW OF THIS EVENT VIA A SEPARATE METHOD. AS REPORTED BY THE FILED, FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. FURTHERMORE, MEDTRONIC RECOMMENDS MAINTAINING CONTROL OF GUIDEWIRE THROUGHOUT PROCEDURE TO ENSURE STABLE DEPLOYMENT AND PREVENT INJURY TO THE VENTRICLE WALL: FOR ALL WIRES WITH A TRANSITION POINT, ENSURE TRANSITION POINT IS HELD ABOVE THE APEX AND POINTING AWAY FROM THE VENTRICLE WALL; MAINTAIN STRICT FLUOROSCOPIC SURVEILLANCE OF THE GUIDEWIRE IN THE LEFT VENTRICLE. THERE IS EVIDENCE OR ANOTHER FLUOROSCOPIC VALVE LOAD INSPECTION AND A MISLOAD APPEARED TO BE PRESENT BY OVERLAP BEYOND NODE 4. IT WAS REPORTED THAT THE VALVE WAS NOT IMPLANTED; A PERFORATION OF THE LEFT VENTRICLE WAS REPORTED AND THE PATIENT SUBSEQUENTLY DIED. AS STATED IN THE IFU, POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT PRO+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: DEATH, CARDIAC TAMPONADE, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: PRE-IMPLANT COMPUTED TOMOGRAPHY (CT) IMAGING PROVIDED FROM (B)(6) 2025. AORTIC VALVE APPEARS TO BE A SIEVERS 1 BICUSPID WITH FUSION OF THE RCC AND LCC. AORTIC VALVE LEAFLETS AND RAPHE APPEAR MODERATELY CALCIFIED. LARGE ANNULUS MEASURED WITH A PERIMETER IS 99.2 MM WITH A PERIMETER DERIVED DIAMETER OF 31.6 MM SUGGESTING A 34 MM EVOLUT. EFFECTIVE ORIFICE AREA AT 5 MM MEASURED A PERIMETER OF 98.6 MM AND PERIMETER DERIVED DIAMETER OF 31.4 CONSISTENT WITH LARGE ANNULUS, NOT WITHIN LIMITS OF THE EVOLUT SIZING MATRIX. SINUS OF VALSALVA (SOV) DIAMETERS AND SINUS HEIGHTS ARE WITHIN THE RECOMMENDED RANGES. CALCIUM SEEN IN AO ANNULUS BELOW NON-CORONARY CUSP (NCC) AND LEFT CORONARY CUSP (LCC), EXTENDING INTO LEFT VENTRICULAR OUTFLOW TRACT (LVOT). ASCENDING AORTA IS DILATED MEASURING AN AVERAGE OF 43.6 MM AT 50 MM DISTANCE FROM THE BASAL PLANE. AORTIC ROOT ANGULATION IS 57°. CALCIFICATION SEEN IN RIGHT COMMON ILIAC ARTERY AND LEFT COMMON ILIAC ARTERY. FEMORAL DIAMETERS ARE WITHIN THE RECOMMENDED RANGES. REPORT REVIEW: FIELD TEAM, PRE-IMPLANT CASE REPORT PROVIDED FROM 3/27/2025. REPORT CONFIRMS LARGE ANNULUS PERIMETER (98.0 MM WITH A PERIMETER DERIVED DIAMETER OF 31.2 MM). POSSIBLE BICUSPID WITH MODERATE AV LEAFLET CALCIFICATION NOTED. AORTIC ROOT ANGULATION IS 52°. NO SUPRA-ANNULAR MEASUREMENT PROVIDED. IMAGE OF DILATED ASCENDING AORTA PROVIDED. INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW. FLUOROSCOPY VALVE LOAD INSPECTION WAS PERFORMED AND A MISLOAD APPEARED TO BE PRESENT BY OVERLAP BEYOND NODE 4. PER MEDTRONIC BEST PRACTICES, OVERLAP PRIOR TO NODE 4 IS CONSIDERED A GOOD LOAD; OVERLAP AT NODE 4 AND BEYOND IS CONSIDERED A MISLOAD. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), IF A MISLOAD IS DETECTED, DO NOT ATTEMPT TO RELOAD THE BIOPROSTHESIS. DISCARD THE ENTIRE SYSTEM. THE VALVE, CATHETER, LOADING SYSTEM, LOADING TRAY, AND SALINE MUST ALL BE REPLACED WITH NEW STERILE COMPONENTS. HOWEVER, THE MISLOAD WAS NOT IDENTIFIED AND THE VALVE WAS USED FOR THE PROCEDURE. THE IMAGES SHOW THAT A CURVED GUIDEWIRE WAS USED FOR THE PROCEDURE AND APPEARED TO BE WELL POSITIONED IN THE LEFT VENTRICLE. A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANT ATTEMPT. THE IMAGES SUPPORT THAT LEFT TRANSFEMORAL ACCESS WAS USED FOR THE PROCEDURE. IT WAS REPORTED THAT RESISTANCE WAS ENCOUNTERED DURING ADVANCEMENT OF THE DELIVERY CATHETER SYSTEM. AS STATED IN THE IFU: CAUTION: THERE WILL BE SOME RESISTANCE WHEN THE CATHETER IS ADVANCED THROUGH THE VASCULATURE. IF THERE IS A SIGNIFICANT INCREASE IN RESISTANCE, STOP ADVANCEMENT AND INVESTIGATE THE CAUSE OF THE RESISTANCE (FOR EXAMPLE, MAGNIFY THE AREA OF RESISTANCE) BEFORE PROCEEDING. DO NOT FORCE PASSAGE. FORCING PASSAGE COULD INCREASE THE RISK OF VASCULAR COMPLICATIONS (FOR EXAMPLE, VESSEL DISSECTION OR RUPTU RE). CAUTION: PERSISTENT FORCE ON THE CATHETER CAN CAUSE THE CATHETER TO KINK, WHICH COULD INCREASE THE RISK OF VASCULAR COMPLICATIONS (FOR EXAMPLE, VESSEL DISSECTION OR RUPTURE). NOTE: WHEN CROSSING THE AORTIC ARCH, IT IS CRITICAL THAT THE GUIDEWIRE IS CONTROLLED TO PREVENT IT FROM MOVING FORWARD. WITHOUT PROPER MANAGEMENT OF THE DISTAL TIP OF THE GUIDEWIRE, THE GUIDEWIRE COULD MOVE FORWARD AND CAUSE TRAUMA TO THE LV. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND AN INFOLD WAS EVIDENT. THE INFOLD IS CHARACTERIZED BY AN INWARD FOLD OR CREASE IN THE VALVE, EXTENDING FROM THE INFLOW. INFOLDING COULD OCCUR DUE A MISLOADED VALVE, ANATOMICAL CHARACTERISTICS SUCH AS CALCIUM, AND RECAPTURES. IF AN INFOLD IS IDENTIFIED, THE VALVE SHOULD BE RECAPTURED, REMOVED FROM THE BODY, AND DISCARDED. NEW STERILE COMPONENTS MUST BE USED. IT WAS REPORTED THAT THE VALVE WAS NOT IMPLANTED, AND A PERFORATION OF THE RIGHT VENTRICLE WAS NOTED DUE TO THE GUIDEWIRE; HOWEVER, THE GUIDEWIRE WAS IN THE LEFT VENTRICLE AND NOT THE RIGHT VENTRICLE. NO FURTHER IMAGES WERE INCLUDED FOR REVIEW THEREFORE IT IS NOT POSSIBLE TO VALIDATE THE VENTRICULAR PERFORATION NOR THE CAUSE OF DEATH. AS STATED IN THE IFU, POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT PRO+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: DEATH, CARDIAC TAMPONADE, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION). UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT PRO PLUS DCS; PRODUCT ID D-EVPROP34; (LOT: 0012516782); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 CONTINUATION OF D10: PRODUCT ID PTC-J3-C-F7-035-260; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: GUIDEWIRE; CORRECTED DATA: H6 - DEVICE CODE CORRECTED TO INCLUDE A150206 FOR THE CONCOMITANT DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21792612&pc=OZD | 1220648-2025-27286 | 2025-03-09 05:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Asystole | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT EXPERIENCED LIMB ISCHEMIA AS THE PATIENT ONLY HAD A RIGHT POSTERIOR TIB PULSE IN THE IMPELLA LEG. THE LEGS WERE COOL AND THE RIGHT FOOT STARTED TO BECOME DUSKY. THE LIMB ISCHEMIA IMPROVED WITH MANIPULATION OF THE ANGLE OF INSERTION BUT IT WAS STILL PRESENT. THE PATIENT WAS NOTED TO HAVE WORSENING MULTIPLE ORGAN DYSFUNCTION SYNDROME. MAXIMAL IMPELLA SUPPORT EVENTUALLY PROGRESSED AND COULD NOT INCREASE BEYOND P5. THE PATIENT WAS DEEMED NOT TO BE A CANDIDATE FOR ADDITIONAL MECHANICAL CIRCULATORY SUPPORT IN PATIENT DISCUSSIONS AND CONTINUED SUPPORT WAS NO LONGER WANTED BY THE PATIENT. THE PATIENT BECAME NON-PULSATILE AND PROGRESSED TO ASYSTOLE AND EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿ Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27286 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 COMPONENT CODE 3038 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27286.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21792665&pc=DSQ | 2916596-2025-02017 | 2025-03-16 04:00 | Death | THORATEC CORPORATION | 2025-04-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Embolism/Embolus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO STROKE ON (B)(6) 2025. THERE WAS NO ALARM. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM}. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21792736&pc=DSQ | 2916596-2025-02043 | 2025-03-15 04:00 | Death | THORATEC CORPORATION | 2025-04-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CEREBRAL BLEEDING. THE PATIENT HAD AN INFECTION AT THE SAME TIME, AND IT WAS REPORTED THAT THEY DID NOT KNOW WHERE THE INFECTION CAME FROM. THE PATIENTS INTERNATIONAL NORMALIZED RATIO (INR) WAS AT 10 WITHOUT POSSIBILITY TO DECREASE THE INR VALUE. THE CAUSE OF THE CEREBRAL BLEEDING WAS UNKNOWN BUT PROBABLY DUE TO HIGH LEVEL OF INR. THERE WAS NO FALL OR ANY CHANGE TO PATIENT CONDITION OR ANTICOAGULATION STATUS. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED INFECTION COULD NOT BE DETERMINED. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED CEREBRAL BLEEDING AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, AND PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), STROKE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, UNDER ¿ANTICOAGULATION¿, OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE HEARTMATE 3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION. THE IFU ALSO PROVIDES SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21792885&pc=NPW | 2015691-2025-02672 | 2025-03-13 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-07 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device | Cardiac Perforation | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE PROCEDURE IN TRICUSPID POSITION WHERE IMMEDIATELY UPON VALVE DEPLOYMENT, THE PATIENT'S PRESSURE DROPPED, THEY WERE PUT ON ECMO, AND TAKEN TO SURGERY. PRIOR TO FINAL RELEASE OF THE VALVE, IT WAS NOTED THAT THE MILD PERICARDIAL EFFUSION THE PATIENT HAD WHEN THE CASE STARTED WAS NOW A SEVERE PERICARDIAL EFFUSION WITH THE START OF TAMPONADE. THE VALVE WAS DEPLOYED, PATIENT HAD A PERICARDIAL TAP, THEIR PRESSURE STARTED TO RAPIDLY DECREASE AND THE GLOBAL FUNCTION OF THE HEART STOPPED. THE PATIENT HAD TO HAVE CONTINUOUS DRAINAGE OF THE EFFUSION WHILE RECEIVING EPINEPHRINE TO KEEP THE BLOOD PRESSURE UP. CT SURGERY WAS CALLED, PATIENT WAS PUT ON ECMO TO BE MADE STABLE FOR SURGERY. PHYSICIAN INFORMED US THAT IT WAS ABOUT A 1 CM LACERATION IN THE ANTERO-SUPERIOR LEAFLET (A/L) OF THE RIGHT VENTRICLE (RV) THAT WAS BEING SUTURED AND THAT THE PERFORATION WAS CAUSED BY THE WIRE. ON POST-OPERATIVE DAY (POD) 4, THE PATIENT WAS REPORTED TO BE STABLE AND OFF ECMO.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT NEVER LEFT THE HOSPITAL POST PROCEDURE; BACK ON ECMO/UNSTABLE RESULTING IN DEATH. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B2, B4, B5, G3, G6, H1, H2, H6 AND H11. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. THE COMPLAINT FOR DELIVERY SYSTEM AND/OR GUIDEWIRE PUSHED INTO VENTRICULAR WALL WAS CONFIRMED WITH OBJECTIVE EVIDENCE. NO MANUFACTURING NON-CONFORMITIES WERE IDENTIFIED FROM THE IMAGING EVALUATION. BASED ON THE DHR REVIEW, THERE WERE NO NON-CONFORMANCES RELATED TO THE ISSUE OBSERVED AND THE DEVICE PASSED ALL MANUFACTURING SPECIFICATIONS. SINCE NO EDWARDS DEFECT WAS IDENTIFIED, CORRECTIVE OR PREVENTATIVE ACTIONS ARE NOT REQUIRED. AVAILABLE INFORMATION SUGGESTS THAT PRE-EXISTING MILD EFFUSION AND THE INTERACTION WITH THE GUIDEWIRE ON THE HEAVILY TRABECULATED RIGHT VENTRICLE MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21793032&pc=DSQ | 2916596-2025-02005 | 2024-08-01 04:00 | Death | THORATEC CORPORATION | 2025-04-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Dysphagia/ Odynophagia; Intracranial Hemorrhage; Hemorrhagic Stroke; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO PALLIATION. THEY WERE ON POST STROKE PALLIATIVE CARE AFTER MINIMAL NEUROLOGY RECOVERY. THE PATIENT PRESENTED (B)(6) 2025 FOR IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) SHOCKS AND CONSTITUTIONAL SYMPTOMS DUE TO CHRONIC DRIVELINE INFECTION, WITH THE ONSET DATE OF AUGUST 2024. ON (B)(6) 2025, THEY RECEIVED STERNAL DEBRIDEMENT FOR A DRIVELINE INFECTION COMPLICATED BY A LEFT MIDDLE CEREBRAL ARTERY (MCA) STROKE WITH LEFT SECOND SEGMENT (M2) CUT-OFF, MODERATE ISCHEMIC BURDEN. ON (B)(6) 2025, A COMPUTED TOMOGRAPHY (CT) HEAD WAS REPEATED FOR WORSENING NEUROLOGICAL SYMPTOMS WHICH REVEALED HEMORRHAGIC TRANSFORMATION OF HIS KNOWN INFARCT. FURTHER, THEY WERE COMPLICATED BY INFLUENZA A, LEFT PSOAS RETROPERITONEAL BLEED, AND HEMORRHAGIC TRANSFORMATION OF THE STROKE. THE PATIENT REMAINED APHASIC, DYSPHAGIC, AND HIGHLY DEPENDENT FOR SELF-CARE. WITH ACTIVE INFECTION, PATIENT'S ANTICOAGULATION BRIDGE WAS CHANGED FROM WARFARIN TO INTRAVENOUS (IV) HEPARIN AND THEN HELD FOR DEBRIDEMENT. THE PATIENT HAD A HISTORY OF ATRIAL FIBRILLATION/FLUTTER PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD). AN ICD WAS IMPLANTED PRIOR TO VAD. THE ARRHYTHMIA WAS CONSIDERED CONSTITUTIONAL SYMPTOM DUE TO THE INFECTION. AFTER MULTIPLE DISCUSSIONS WITH THE PATIENT, FAMILY, AND PALLIATIVE CARE TEAM, AND IN LINE WITH THE PATIENT'S PREVIOUS WISHES, THE DECISION WAS MADE TO PROCEED WITH COMFORT MEASURES, INCLUDING DEVICE DEACTIVATION, IN THE PRESENCE OF THE FAMILY. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT HAD NON-DEVICE RELATED COMPLICATIONS; THEY DEVELOPED INTRA CRANIAL BLEEDING AND DEVELOPED RESIDUAL LESIONS WHICH WAS DECIDED BY THE PATIENT AND THEIR FAMILY TO STOP THE SUPPORT FOR PALLIATION (B)(6) 2025. Manufacturer Narrative: SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. C ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), CARDIAC ARRHYTHMIA, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AND INFECTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. SECTION 6 OF THE IFU, (UNDER ¿ANTICOAGULATION¿), ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21793212&pc=NPT | 2025587-2025-02445 | 2025-03-19 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-07 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Fluid/Blood Leak; Insufficient Information; Patient Device Interaction Problem; Central Regurgitation | Cardiogenic Shock; Heart Failure/Congestive Heart Failure; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 7 YEARS AND 10 MONTHS FOLLOWING THE EVOLUTR-34-US TRANSCATHETER AORTIC VALVE IMPLANT, REGURGITATION OF AN UNKNOWN TYPE WAS REPORTED. 7 YEARS AND 11 MONTHS FOLLOWING THE VALVE IMPLANT THE PATIENT WAS HOSPITALIZED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CLARIFIED PREVIOUSLY REPORTED INFORMATION INDICATING THAT THE PATIENT WAS HOSPITALIZED 6 YEARS AND 10 MONTHS FOLLOWING THE IMPLANT OF THE VALVE DUE TO AN UNKNOWN REASON. UNSPECIFIED TREATMENT WAS PERFORMED, AND THE PATIENT WAS DISCHARGED. SUBSEQUENTLY, 1 MONTH LATER, THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT IN CARDIOGENIC SHOCK, AND THE PATIENT WAS RE-HOSPITALIZED. THE PATIENT SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONALLY, THE TYPE OF AORTIC REGURGITATION THAT WAS INITIALLY REPORTED WAS MODERATE CENTRAL AORTIC REGURGITATION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS INITIALLY HOSPITALIZED DUE TO CARDIOGENIC SHOCK CAUSED BY HEART FAILURE. THE TREATMENT THAT WAS PROVIDED WAS EMERGENT LIFE SAVING TREATMENT. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO HEART FAILURE. Manufacturer Narrative: UPDATED DATA: B2. B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. THE DATE OF DEATH IS APPROXIMATE. B5. H1. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21793379&pc=NPT | 2015691-2025-02678 | 2024-09-07 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-04-07 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Device Stenosis | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, APPROXIMATELY FOUR YEARS POST TRANSFEMORAL TAVR PROCEDURE WITH A 20 MM SAPIEN 3 VALVE, THE PATIENT DIED DUE TO END-STAGE HEART FAILURE AND THE ASSOCIATED DECLINE IN ACTIVITIES OF DAILY LIVING (ADL), COGNITIVE FUNCTION, AND SWALLOWING FUNCTION. IT WAS REPORTED THAT THE PATIENT HAD HEART FAILURE AND AORTIC VALVE STENOSIS PRIOR TO THEIR DEATH. THE CAUSE OF DEATH WAS REPORTED AS CARDIAC DEATH. ACCORDING TO MEDICAL OPINION, THE EVENT WAS DETERMINED TO BE DEVICE-RELATED, NOT PROCEDURE-RELATED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR THE COMPLETED ENGINEERING EVALUATION. SECTIONS B.4, G.3, G.6, AND H.2 HAVE BEEN UPDATED. CORRECTIONS HAVE BEEN MADE TO H.6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSION. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, STENOSIS/INCREASED GRADIENT EVENTS FOR THE SAPIEN 3, SAPIEN 3 ULTRA, SAPIEN 3 ULTRA RESILIA VALVES POST-IMPLANTATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAS HISTORICALLY BEEN DUE TO PATIENT FACTORS (ATHEROSCLEROSIS, THROMBOSIS, ENDOCARDITIS, RHEUMATIC FEVER, PANNUS FORMATION) AND/OR PROCEDURAL FACTORS (UNDER-DEPLOYED VALVE, VALVE SIZE SELECTION, PATIENT-PROSTHESIS MISMATCH). THE COMPLAINT FOR STENOSIS WAS UNABLE TO BE CONFIRMED DUE TO THE UNAVAILABILITY OF THE RETURNED DEVICE, RELEVANT IMAGERY, OR RELEVANT MEDICAL RECORDS. IT IS POSSIBLE THAT PATIENT OR PROCEDURAL FACTORS IDENTIFIED IN THE TECHNICAL SUMMARY CONTRIBUTED TO THE COMPLAINT EVENT. HOWEVER, DUE TO INSUFFICIENT INFORMATION, A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21793580&pc=OZD | 1220648-2025-27287 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-07 04:00 | OZD | IMPELLA CP | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE COMPLAINANT REPORTED THAT THEM PUMP PRESENTED HIGH PURGE PRESSURE. THE HIGH PURGE PRESSURE ALARM WAS DISCUSSED WITH THE NURSE AND A KINK WAS FOUND IN LINE NEAR THE PURGE CASSETTE DOOR. HOWEVER, A FEW DAYS LATER THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN." Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG SHOWS A QUICK RISE IN PURGE PRESSURE WITH A PURGE SYSTEM BLOCKED ALARM, WHICH RESOLVED IN 10 MINUTES. ACCORDING TO THE CLINICAL DETAILS, A PURGE LINE KINK WAS NOTED NEAR THE PURGE CASSETTE DOOR. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY KINKED PURGE LINE NEAR PURGE CASSETTE DOOR, BASED ON GIVEN CLINICAL. B2 DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27287. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27287 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21782674&pc=QZI | 2124215-2025-21059 | 2025-03-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-06 04:00 | QZI | FARAWAVE | Off-Label Use | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT DIED. AN UNSPECIFIED EMPLOYEE OF AN UNKNOWN NON-BOSTON SCIENTIFIC MEDICAL DEVICE COMPANY TOLD A PHYSICIAN ABOUT A 'PATIENT DEATH DUE TO PFA ON THE CTI LINE'. NO FURTHER INFORMATION WAS PROVIDED AND NO DEVICE RETURN IS EXPECTED. Manufacturer Narrative: FURTHER INFORMATION REGARDING THE EVENT OR SPECIFIC DEVICE INFORMATION COULD NOT BE OBTAINED AS THE EVENTS WERE NOT REPORTED TO BOSTON SCIENTIFIC BY ANY PERSON THAT WAS PRESENT OR HAD SPECIFIC KNOWLEDGE ABOUT THE ALLEGED EVENT. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21781945&pc=DQY | 2029214-2025-00890 | 2023-10-01 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-05 04:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Paresis; Respiratory Failure; Cancer; Movement Disorder; Respiratory Insufficiency | K182097 |  | 1 | Event Description: MATSUMOTO, S., SUZUKI, T., EMOTO, S., NAKASE, H., MATSUMURA, Y.; INTERNAL MEDICINE; 2024; 63: 3087-3091; A CEREBRAL EMBOLISM CAUSED BY A MALIGNANT PERIPHERAL NERVE SHEATH TUMOR IN A PATIENT WITH NEUROFIBROMATOSIS TYPE 1; DOI: 10.2169/INTERNALMEDICINE.2996-23 ARTICLE OBJECTIVE: THE DOCUMENT REPORTS A CASE OF A CEREBRAL EMBOLISM CAUSED BY A MALIGNANT PERIPHERAL NERVE SHEATH TUMOR (MPNST) IN A PATIENT WITH NEUROFIBROMATOSIS TYPE 1 (NF-1). THE PATIENT WAS A 31-YEAR-OLD MAN WHO UNDERWENT RESECTION OF THE MPNST ON THE BUTTOCK 3 MONTHS PRIOR TO PRESENTING WITH A HYPERACUTE LEFT MIDDLE CEREBRAL ARTERY EMBOLISM. THE CASE REPORT DETAILS EVENTS THAT OCCURRED OVER A HOSPITALIZATION PERIOD OF 15 DAYS. THE DOCUMENT DOES NOT EXPLICITLY STATE WHETHER MULTIPLE MANUFACTURERS' DEVICES WERE USED. IT PRIMARILY MENTIONS THE USE OF A SPECIFIC MEDTRONIC PRODUCT FOR MECHANICAL THROMBECTOMY. THE DOCUMENT MENTIONS THE USE OF THE ASPIRATION CATHETER REACT-71 FOR MECHANICAL THROMBECTOMY. THE PATIENT DIED ON THE 15TH DAY OF HOSPITALIZATION DUE TO RESPIRATORY FAILURE CAUSED BY LUNG METASTASIS OF THE MPNST. ADVERSE EVENTS/COMPLICATIONS: ON THE 14TH DAY OF ADMISSION, HIS CONSCIOUSNESS LEVEL DECREASED TO GCS E3V2M5, AND NEUROLOGICAL EXAMINATION DEMONSTRATED CONJUGATE DEVIATION OF THE LEFT EYE WITH RIGHT UPPER AND LOWER LIMB PARALYSIS. IN ADDITION, HIS RESPIRATORY CONDITION WORSENED TO SPO2 94% (O2 1 L/MIN). BRAIN COMPUTED TOMOGRAPHY (CT) REVEALED AN INTRATUMORAL HEMORRHAGE IN THE LEFT FRONTAL LOBE NEOPLASM. IN ADDITION, BODY CONTRAST CT REVEALED MULTIPLE TUMORAL LESIONS IN THE LEFT LUNG AND MASSIVE PLEURAL EFFUSION WITH ATELECTASIS IN THE RIGHT LUNG. CT DID NOT REVEAL ARTERIOSCLEROTIC LESIONS IN THE AORTA. ANOTHER TUMOR LESION WAS OBSERVED IN THE RIGHT KIDNEY. AS IT WAS CONSIDERED THAT THE PATIENT WOULD NOT BE ABLE TO TOLERATE SURGERY DUE TO SEVERE ATELECTASIS, CONSERVATIVE TREATMENT FOR THE INTRATUMORAL HEMORRHAGE AND DRAINAGE OF THE PLEURAL FLUID WAS CONDUCTED. ON THE 15TH DAY AFTER ADMISSION, HIS RESPIRATORY STATUS ABRUPTLY WORSENED AND HE DIED OF RESPIRATORY FAILURE. ON AUTOPSY, MULTIPLE NODULES OF THE MPNST OCCUPIED THE RIGHT LUNG. MOREOVER, A LUNG HISTOLOGICAL EXAMINATION OF THE AUTOPSY CONFIRMED TUMOR INVASION OF THE PULMONARY VEINS. THE RESULTS OF ELASTICA VAN GIESON STAINING CONFIRMED THAT THE LESION WAS VENOUS INVASION. MOST TUMOR CELLS HAD SPINDLE-SHAPED AND HYPERCHROMATIC NUCLEI THAT WERE IDENTIFIED AS MPNST. MULTIPLE VASCULAR INVASIONS OF THE MPNST, SUCH AS IN THE KIDNEYS AND RIGHT THYROID GLAND, WERE ALSO OBSERVED. VERRUCAE, CALCIFICATION OF THE VALVES, AND PATENT FORAMEN OVALE WERE NOT OBSERVED IN THE HEART. FURTHERMORE, THERE WAS NO EVIDENCE OF CEREBROVASCULAR ATHEROSCLEROSIS. IN ADDITION, THE HEMORRHAGE IN THE LEFT FRONTAL LOBE WAS HISTOPATHOLOGICALLY IDENTIFIED TO BE THE RESULT OF MPNST. THEREFORE, RESPIRATORY FAILURE DUE TO METASTATIC MPNST IN THE LUNGS AND CEREBRAL HEMORRHAGE CAUSED BY INTRACRANIAL MPNST WERE REGARDED AS THE CAUSE OF DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21781951&pc=OZP | 2032227-2025-149554 | 2025-03-06 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-05 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Neuropathy; Insufficient Information | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER EXPERIENCED HYPOGLYCEMIA WITH A BLOOD GLUCOSE VALUE OF 54 MG/DL CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS NOT SURE. THE PUMP WAS WORN AT THE TIME OF PASSING. THE REASON FOR ADMISSION TO THE EMERGENCY ROOM VISIT AND HAD NO MORE PULSE. THE HEALTH ISSUES THAT LED UP TO PASSING WERE CHRONIC KIDNEY DISEASE, DIABETES, AND NEUROPATHY. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-397, MMT-1884, MMT-332A, MMT-7040A. NO PRODUCT RETURN IS REQUIRED FOR MMT-397. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: HEALTH EFFECT-CLINICAL CODE-2562, HEALTH EFFECT-IMPACT CODE-4614. THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21770665&pc=DTB | 2649622-2025-09073 | 2025-02-25 05:00 | Death | MPRI | 2025-04-04 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Syncope/Fainting; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT WAS AT A STORE AND EXPERIENCED A SYNCOPE EPISODE AND WAS HAVING TROUBLE BREATHING. THE PATIENT PASSED AWAY ONE WEEK LATER.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21770667&pc=DTB | 2649622-2025-09074 | 2025-02-25 05:00 | Death | MPRI | 2025-04-04 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Syncope/Fainting; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT WAS AT A STORE AND EXPERIENCED A SYNCOPE EPISODE AND WAS HAVING TROUBLE BREATHING. THE PATIENT PASSED AWAY ONE WEEK LATER.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21770671&pc=NVZ | 3008973940-2025-02388 | 2025-02-25 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-04 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Syncope/Fainting; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT WAS AT A STORE AND EXPERIENCED A SYNCOPE EPISODE AND WAS HAVING TROUBLE BREATHING. THE PATIENT PASSED AWAY ONE WEEK LATER.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21770781&pc=OZP | 2032227-2025-148272 | 2025-02-17 05:00 | Death | MEDTRONIC MINIMED | 2025-04-04 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Loss of consciousness |  |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CUSTOMER BLOOD GLUCOSE WAS PLUMMETED AND THEN BECAME UNCONSCIOUS AND THEN DIED AT HOME. THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNOMEDICAL, UNKNOWN NGP, MMT-368A(INF SET). NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR UNKNOWN NGP. NO PRODUCT RETURN IS REQUIRED FOR MMT-368A(INF SET). Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21771427&pc=BZD | 2518422-2025-041239 | 2024-06-25 04:00 | Death | RESPIRONICS, INC. | 2025-04-04 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21771559&pc=NVZ | 3008973940-2025-02401 | 2025-03-17 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-04 04:00 | NVZ | AZURE XT SR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Bacterial Infection; Erythema; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY THREE WEEKS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM THAT THE SYSTEM WAS REMOVED FOR AN UNKNOWN REASON AND REPLACED WITH A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PRESENTED TO THE CLINIC WITH A REDDENED AREA AT DISTAL END IPG POCKET. A POTENTIAL INFECTION WAS SUSPECTED. POST EXPLANT OF THE IPG SYSTEM THE POCKET HAD NO OBSERVABLE DRAINAGE OR SIGNS OF INFECTION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DECEASED. IT WAS FURTHER NOTED THAT BLOOD CULTURES WERE NEGATIVE, HOWEVER POCKET CULTURE SHOWED KLEBSIELLA PNEUMONIAE. IT WAS MENTIONED THAT THERE WAS NOTHING IN PROGRESS NOTES RELATING TO THE IPG SYSTEM AS BEING CAUSATIVE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21771565&pc=DTB | 2649622-2025-09093 | 2025-03-17 04:00 | Death | MPRI | 2025-04-04 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Bacterial Infection; Erythema; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY THREE WEEKS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM THAT THE SYSTEM WAS REMOVED FOR AN UNKNOWN REASON AND REPLACED WITH A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PRESENTED TO THE CLINIC WITH A REDDENED AREA AT DISTAL END IPG POCKET. A POTENTIAL INFECTION WAS SUSPECTED. POST EXPLANT OF THE IPG SYSTEM THE POCKET HAD NO OBSERVABLE DRAINAGE OR SIGNS OF INFECTION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DECEASED. IT WAS FURTHER NOTED THAT BLOOD CULTURES WERE NEGATIVE, HOWEVER POCKET CULTURE SHOWED KLEBSIELLA PNEUMONIAE. IT WAS MENTIONED THAT THERE WAS NOTHING IN PROGRESS NOTES RELATING TO THE IPG SYSTEM AS BEING CAUSATIVE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21771830&pc=NVN | 2649622-2025-09106 |  | Death | MPRI | 2025-04-04 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT TWO DAYS POST IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21771944&pc=BZD | 2518422-2025-041254 | 2023-06-08 04:00 | Death | RESPIRONICS, INC. | 2025-04-04 04:00 | BZD | DREAMSTATION 2 ADVANCE AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE THAT A USER OF THE DREAMSTATION 2 ADVANCE AUTO CPAP HAD PASSED AWAY BEFORE RECEIVING THE REPLACEMENT MACHINE. IT WAS NOT CONFIRMED WHETHER THE DEVICE HAD CAUSED THE DEATH OR NOT. THE USER'S DAUGHTER WANT TO USE THE MACHINE UNDER HER NAME. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21772625&pc=NVN | 2649622-2025-09124 | 2025-02-07 05:00 | Death | MPRI | 2025-04-04 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Arrhythmia | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED AND DID HAVE SOME ARRYTHMIAS PRIOR TO DEATH. NO FURTHER INFORMATION.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN BUT WAS NOT BELIEVED TO BE RELATED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED B2 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21772644&pc=DTB | 2649622-2025-09126 | 2025-02-07 05:00 | Death | MPRI | 2025-04-04 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Arrhythmia | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED AND DID HAVE SOME ARRYTHMIAS PRIOR TO DEATH. NO FURTHER INFORMATION.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN BUT WAS NOT BELIEVED TO BE RELATED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: CORRECTED B2 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21772653&pc=NVZ | 3008973940-2025-02410 | 2025-02-07 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-04 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Arrhythmia | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED AND DID HAVE SOME ARRYTHMIAS PRIOR TO DEATH. NO FURTHER INFORMATION.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN BUT WAS NOT BELIEVED TO BE RELATED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED B2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21772680&pc=OZD | 1220648-2025-27232 | 2025-03-11 04:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem; Material Split, Cut or Torn | Vascular Dissection | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT UNDERGOING HIGH RISK PERCUTANEOUS CORONARY INTERVENTION (HRPCI) WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE TEAM MADE THE DECISION TO ALSO PLACE A 7 FRENCH COMPANION SHEATH FOR THE HRPCI. THE TEAM OBSERVED A VASCULAR DAMAGE/PERFORATION OF THE ILIAC ARTERY FROM THE WIRELESS INSERTION OF THE 7 FRENCH SHEATH. THE VESSEL DAMAGE PROMPTED THE PUMP TO BE EXPLANTED AND COVERED WITH TWO STENTS TO REPAIR THE ARTERY. THE PATIENT LATER EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. VASCULAR DAMAGE - PERIPHERAL VESSEL: CLINICAL STATES THAT, AT THE TIME OF IMPLANT, THE COMPANION SHEATH WAS INSERTED WITHOUT A GUIDEWIRE RESULTING IN ARTERY DISSECTION. COMPANION SHEATH, 14F INTRODUCER, AND PUMP WERE RETURNED. THE COMPANION SHEATH HAD A DAMAGED TIP AND THE PARTIAL TIP MISSING. THESE SUPPORTS THE CLINICAL THAT THE COMPANION SHEATH WAS USED WITHOUT GUIDEWIRE. THEREFORE, THE ROOT CAUSE OF THE VASCULAR DAMAGE ISSUE WAS USE RELATED TO IMPROPER TECHNIQUE AS EVIDENCED IN THE PRODUCT INVESTIGATION. Manufacturer Narrative: ADDITION FINDINGS DURING THE INVESTIGATION THAT THERE WAS INTRODUCER DAMAGE ISSUE. THERE WAS A SHEATH TIP SPLIT WHICH HAPPENED DUE TO THE IMPROPER IMPLANT TECHNIQUE. H10 ADDITIONAL INVESTIGATION FINDINGS WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27232-1. H6 MEDICAL DEVICE PROBLEM CODE 4008 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27232-1. Manufacturer Narrative: THE PUMP AND INTRODUCER WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21772946&pc=NPT | 2025587-2025-02354 | 2025-04-02 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-04 04:00 | NPT | EVOLUT FX PLUS VALVE | Material Invagination; Device Dislodged or Dislocated; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Loss of consciousness | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, UPON THE INITIAL DEPLOYMENT ATTEMPT, THE VALVE WAS TOO HIGH AND WAS RECAPTURED. THE SECOND DEPLOYMENT WAS ADEQUATE AND THE VALVE WAS RELEASED. AN INFOLD AND UNDER-EXPANSION WERE OBSERVED AND A POST-IMPLANT BALLOON VALVULOPLASTY WAS PLANNED; HOWEVER, WITHIN ONE MINUTE OF THE VALVE RELEASE, THE VALVE DISLODGED OUT OF PLACE. THE PATIENT LOST CONSCIOUSNESS AND BLOOD PRESSURE WAS LOW. SUBSEQUENTLY, RESUSCITATION MEASURES WERE PERFORMED WHILE INTRODUCING A SNARE TO CAPTURE THE VALVE. THE VALVE WAS LIFTED VIA THE SNARE THAT CAUGHT ONE OF THE PADDLES. THE PLAN WAS TO IMPLANT A SECOND VALVE BUT THE PATIENT WAS TOO UNSTABLE. THE INITIAL VALVE WAS UNINTENTIONALLY PULLED THROUGHOUT THE AORTA AND ALL THE WAY DOWN THE DESCENDING AORTA UNTIL IT REACHED THE SHEATH. IT WAS DECIDED THAT THE PATIENT SHOULD BE TRANSFERRED TO SURGERY, BUT THE PATIENT PASSED AWAY PRIOR TO GETTING TO THE OPERATING ROOM. PER THE PHYSICIAN, THE CAUSE OF THE EVENT WAS NOT DETERMINED.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, UPON THE INITIAL DEPLOYMENT ATTEMPT, THE VALVE WAS TOO HIGH AND WAS RECAPTURED. THE SECOND DEPLOYMENT WAS ADEQUATE AND THE VALVE WAS RELEASED. AN INFOLD AND UNDER-EXPANSION WERE OBSERVED AND A POST-IMPLANT BALLOON VALVULOPLASTY WAS PLANNED; HOWEVER, WITHIN ONE MINUTE OF THE VALVE RELEASE, THE VALVE DISLODGED OUT OF PLACE. THE PATIENT LOST CONSCIOUSNESS AND BLOOD PRESSURE WAS LOW. SUBSEQUENTLY, RESUSCITATION MEASURES WERE PERFORMED WHILE INTRODUCING A SNARE TO CAPTURE THE VALVE. THE VALVE WAS LIFTED VIA THE SNARE THAT CAUGHT ONE OF THE PADDLES. THE PLAN WAS TO IMPLANT A SECOND VALVE BUT THE PATIENT WAS TOO UNSTABLE. THE INITIAL VALVE WAS UNINTENTIONALLY PULLED THROUGHOUT THE AORTA AND ALL THE WAY DOWN THE DESCENDING AORTA UNTIL IT REACHED THE SHEATH. IT WAS DECIDED THAT THE PATIENT SHOULD BE TRANSFERRED TO SURGERY, BUT THE PATIENT PASSED AWAY PRIOR TO GETTING TO THE OPERATING ROOM. PER THE PHYSICIAN, THE CAUSE OF THE EVENT WAS NOT DETERMINED. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH SINCE IT WAS PULLED THROUGH THE AORTA, BUT THE DELIVERY CATHETER SYSTEM (DCS) WAS NOT A FACTOR. NO PRE-IMPLANT BALLOON DILATION WAS PERFORMED AND A NON-MEDTRONIC (SAFARI) GUIDEWIRE WAS USED FOR THE PROCEDURE. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER AND PRIOR TO DISLODGEMENT, THE IMPLANT DEPTH WAS APPROXIMATELY 3 MM ON THE NON-CORONARY CUSP (NCC) AND LEFT CORONARY CUSP (LCC). THE DIRECTION OF DISLODGEMENT WAS AORTIC AND ABOVE THE ANNULUS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, NO MISLOAD WAS IDENTIFIED DURING LOADING. Manufacturer Narrative: IMAGE REVIEW: EIGHTEEN CINE IMAGES OF THE EVENT WERE PROVIDED FOR REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PROVIDED AND SHOWS A MISLOAD APPEARED TO BE PRESENT BY OVERLAP BEYOND NODE 4. AN INFOLD AND UNDER-EXPANSION WERE OBSERVED ON FULL RELEASE. EVIDENCE SHOWS THE VALVE WAS LIFTED VIA THE SNARE THAT CAUGHT ONE OF THE PADDLES. THE INITIAL VALVE WAS UNINTENTIONALLY PULLED THROUGHOUT THE AORTA AND ALL THE WAY DOWN THE DESCENDING AORTA UNTIL IT REACHED THE SHEATH. UPDATED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012253666); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012275361); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21772962&pc=MNS | 2518422-2025-041259 | 2021-10-14 04:00 | Death | RESPIRONICS, INC. | 2025-04-04 04:00 | MNS | RENTAL, DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION, THE DEVICE WAS VISUALLY INSPECTED INTERNALLY AND EXTERNALLY, AND NO FAULTS WERE FOUND. IN ADDITION, THE DEVICE PASSED ALL TESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773078&pc=FBQ | 9611102-2025-00018 | 2025-03-18 04:00 | Death | RICHARD WOLF GMBH | 2025-04-04 04:00 | FBQ | TROCAR SLEEVE STRAI. CAP 5.8MM WL 62MM | Material Separation | Pulmonary Edema | EXEMPT |  | 1 | Event Description: ACCORDING TO THE INFORMATION RECEIVED, DURING THE THORACOSCOPY PROCEDURE, SURGEON INITIALLY INSERTED THE TROCAR THROUGH WHICH THE SCOPE WAS PASSED TO VISUALIZE THE PLEURAL SPACE. UNFORTUNATELY, THE VISUALIZATION WAS POOR, AND THE USER DECIDED TO REMOVE THE SCOPE FOR FURTHER ASSESSMENT. HOWEVER, ONCE THE SCOPE WAS REMOVED, THE SURGEON DISCOVERED THAT ONLY THE METALLIC HEAD OF THE TROCAR REMAINED IN THE CHEST, WITH THE PLASTIC CANNULA COMPLETELY MISSING. A SUBSEQUENT CT SCAN OF THE CHEST REVEALED THAT THE MISSING PART OF THE TROCAR WAS LODGED IN THE POSTERIOR PLEURAL CAVITY. AS A RESULT, THE PATIENT WAS TAKEN FOR THORACOTOMY TO REMOVE THE BROKEN PIECE FROM THE POSTERIOR PLEURAL CAVITY. UNFORTUNATELY, FOLLOWING THE SURGERY, THE PATIENT DEVELOPED PULMONARY EDEMA, WAS INTUBATED, AND TRAGICALLY PASSED AWAY ONE DAY AFTER THE OPERATION. Manufacturer Narrative: THE DEVICE, TROCAR SLEEVE STRAI. CAP 5.8 MM WL 62 MM, TYPE 8905331, BATCH# 1539651 WAS MANUFACTURED ON FEBRUARY 20, 2023, AND DELIVERED TO THE CUSTOMER IN JULY 2023. SINCE DELIVERY, THE CUSTOMER HAS NOT REPORTED ANY FURTHER OR OTHER COMPLAINTS REGARDING THIS DEVICE. MICROSCOPIC EXAMINATIONS OF THE PRODUCT SHOW SEVERAL MECHANICAL DAMAGES TO THE TROCAR SLEEVE: PLASTIC DEFORMATIONS CAN BE SEEN IN SEVERAL PLACES AT THE DISTAL END. IN THE AREA OF THE SHAFT, THERE ARE SEVERAL PLACES WHERE THE PLASTIC HAS BEEN SCRAPED OFF OR CUT INTO. IN TWO PLACES IN THE AREA OF THE THREAD, THE SHAFT HAS EVEN BEEN COMPLETELY CUT OPEN. SUCH DAMAGE IS ONLY CONCEIVABLE AS A RESULT OF A MASSIVE FORCE APPLIED WITH A SHARP, PROBABLY METALLIC OBJECT. THE SHAFT OF THE TROCAR SLEEVE HAS BEEN TORN OUT OF THE HOUSING. THE TORN-OFF PARTS OF THE PLASTIC SLEEVE ARE STILL IN THE HOUSING. HERE TOO, A MASSIVE FORCE CAN BE ASSUMED. COMPARATIVE TENSILE TESTS REVEALED A FORCE OF 161 N, WHICH IS REQUIRED TO PULL THE SLEEVE OUT OF THE HOUSING. IN SUMMARY, THE TROCAR SLEEVE IN QUESTION SHOWS SEVERAL SIGNS OF DAMAGE THAT CANNOT OCCUR WHEN USED FOR ITS INTENDED PURPOSE. IT CAN THEREFORE BE ASSUMED THAT THE PRODUCT WAS USED CONTRARY TO ITS INTENDED PURPOSE. AS THE PRODUCT ALSO EXHIBITS SEVERAL SERIOUS DAMAGES, WHICH COULD NOT ALL HAVE BEEN CAUSED BY THE SAME ACTION, IT IS ALMOST CERTAIN THAT A PRE-DAMAGED PRODUCT WAS USED IN THE INCIDENT. FOR THIS PURPOSE, THE IFU GA-B199 / USA / 2012-11 V3.0 INSTRUCTIONS FOR USE CONTAIN THE FOLLOWING WARNINGS AND SAFETY INSTRUCTIONS: 1 INTENDED USE: TROCAR SLEEVES AND TROCARS ARE USED FOR CREATING AND SECURING ENDOSCOPIC ACCESS TO BODY CAVITIES. 4 COMBINATIONS: THE TROCAR SLEEVES ARE USED IN CONJUNCTION WITH TROCARS, INSUFFLATORS, DILATION SLEEVES AND GUIDE RODS, ENDOSCOPES AS WELL AS ENDOSCOPIC ACCESSORIES (E.G. FORCEPS, ELECTRODES, ETC.). CAUTION! DO NOT COMBINE PRODUCTS INCORRECTLY! INJURIES OF THE PATIENT, USER OR THIRD PARTIES AS WELL AS DAMAGE TO THE PRODUCT ARE POSSIBLE. COMBINE THE DIFFERENT PRODUCTS ONLY IF THEIR INTENDED USES AND RELEVANT TECHNICAL DATA (WORKING LENGTH, DIAMETER, ETC.) ARE THE SAME. FOLLOW THE INSTRUCTION MANUALS OF THE PRODUCTS USED IN COMBINATION WITH THIS PRODUCT. 6 USE: CAUTION! THE PRODUCTS HAVE ONLY LIMITED STRENGTH! EXCESSIVE FORCE WILL CAUSE DAMAGE, IMPAIR THE FUNCTION AND THEREFORE ENDANGER THE PATIENT. IMMEDIATELY BEFORE AND AFTER EACH USE, CHECK THE PRODUCTS FOR DAMAGE, LOOSE PARTS AND COMPLETENESS. MAKE SURE THAT NO MISSING PARTS REMAIN IN THE PATIENT. DO NOT USE THE PRODUCTS IF THEY ARE DAMAGED OR INCOMPLETE OR HAVE LOOSE PARTS. 6.2 ADDITIONAL NOTES AND INSTRUCTIONS FOR USE: WARNING! ADHESIONS, HIDDEN VESSELS AND ORGANS AT THE PUNCTURE SITE! DANGER OF PERFORATING ORGANS OR VESSELS. DANGER OF SEEPING HEMORRHAGES. MAKE SURE THE PUNCTURE SITE IS FREE OF VESSELS AND ADHESIONS! INSERT THE TROCAR SLEEVE WITH THE TROCAR IN THE BODY UNTIL THE TROCAR TIP IS POSITIONED FREELY IN THE BODY CAVITY. ENDOSCOPIC INTRAABDOMINAL CHECK OF THE PUNCTURE SIGHT! CHECK PNEUMOPERITONEUM FOR INTRAABDOMINAL, INTRAPELVIC PROCEDURES! 7 CHECKS: CAUTION! BE CAREFUL IF PRODUCTS ARE DAMAGED OR INCOMPLETE! INJURIES OF THE PATIENT, USER OR OTHERS ARE POSSIBLE. PERFORM THE CHECKS BEFORE AND AFTER EACH USE. DO NOT USE THE PRODUCTS IF THEY ARE DAMAGED, INCOMPLETE OR HAVE LOOSE PARTS. RETURN DAMAGED PRODUCTS TOGETHER WITH ANY LOOSE PARTS FOR REPAIR. DO NOT ATTEMPT TO DO ANY REPAIRS YOURSELF. CHECK INSTRUMENTS FOR DAMAGE, SHARP EDGES AND ROUGH SURFACES. CHECK MEMBRANE VALVE (2) AND SHEATH (1) FOR DAMAGE AND DEFORMATION. REPLACE DAMAGED MEMBRANE VALVE (E.G. TORN ENDS AT THE SLOTS) (2) AND SHEATH (1) (FOR SPARE PARTS SEE SECTION 10.0). 8 REPROCESSING AND MAINTENANCE. 8.1 DISASSEMBLY BEFORE CLEANING. 8.5 ASSEMBLE BEFORE STERILIZATION. IN THE RISK ASSESSMENT D2 "REUSABLE TROCARS, TROCAR SLEEVES AND NEEDLES" REV.02, THE DAMAGE "DEATH" CAUSED BY MECHANICAL ENERGY, WHICH CAN OCCUR DUE TO HANDLING, E.G. PRODUCT DAMAGE, WAS EVALUATED UNDER THE CORRESPONDING PROBABILITY OF OCCURRENCE AND ASSESSED AS ACCEPTABLE WITH THE CORRESPONDING RISK REDUCTION MEASURES "OBSERVANCE OF THE INSTRUCTIONS IN THE INSTRUCTIONS FOR USE".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773338&pc=OZD | 1220648-2025-27238 | 2025-03-12 04:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT HAD AN IMPELLA CP PUMP PLACED FOR SUPPORT OF ACUTE MYOCARDIAL INFARCTION. THE PUMP WAS PLACED TO ALLOW FOR PERCUTANEOUS CORONARY INTERVENTION. THE PATIENT LOST PULSE BILATERALLY BUT, THE PUMP REMAINED ON FOR SUPPORT UNTIL THE FAMILY MADE THE DECISION TO WITHDRAW CARE AFTER 27 HOURS. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. G4 PMA/510(K) NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27238. H6 COMPONENT CODE 3038 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27238. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773378&pc=OZD | 1220648-2025-27239 | 2025-03-13 04:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Pericardial Effusion | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A CP TO SUPPORT A 84 YEAR OLD MALE PATIENT UNDERGOING CARDIAC PROCEDURE FOR CORONARY ARTERY DISEASE. THE TEAM HAD A GROIN SITE BLEED THAT PROMPTED TREATMENT. THE SITE WAS RE-SUTURED DUE TO HEAVY BLEEDING AND ANGLE MATCHING DONE. THE PUMP REMAINS ON FOR SUPPORT AND THE PATIENT WAS TRANSFERRED TO THE TERTIARY CENTER. ADDITIONALLY, PATIENT WAS FOUND TO HAVE A LARGE PERICARDIAL EFFUSION AND WAS OPENED EMERGENTLY. AFTER MULTIPLE ATTEMPTS TO REPAIR LV, FAMILY WAS CONTACTED AND PATIENT WAS SENT TO UNIT IN PREPARATION FOR PASSING. IMPELLA HAD BEEN PLACED IN THE AORTA IN SURGICAL MODE BEFORE CASE WAS STARTED. SURGEON UNCERTAIN OF WHAT CAUSED PERFORATION, AS THE IMPELLA WAS SHALLOWED OUT TO AORTIC VALVE BEFORE PERFORATION. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773509&pc=MAE | 2135147-2025-01725 | 2020-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-04-04 04:00 | MAE | AMPLATZER DUCT OCCLUDER | Off-Label Use | Insufficient Information | P020024 |  | 1 | Event Description: THE ARTICLE, "PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, SINGLE CENTER STUDY TO DETERMINE THE SAFETY AND EFFICACY OF PERCUTANEOUS DEVICE CLOSURE FOR SIGNIFICANT PARAVALVULAR LEAKS (PVLS). DEVICES INCLUDED IN THE STUDY WERE AMPLATZER VASCULAR PLUG II, AMPLATZER VASCULAR PLUG III, AND AMPLATZER DUCT OCCLUDER. THE ARTICLE CONCLUDED THAT TRANSCATHETER PVL CLOSURE IS A SAFE AND EFFECTIVE ALTERNATIVE TO SURGICAL CORRECTION AND SHOULD BE USED IN SELECTED PATIENTS FOR SIGNIFICANT AMELIORATION OF HEART FAILURE SYMPTOMS AS WELL AS HEMOLYTIC ANEMIA WITH CLINICAL AND HEMODYNAMIC IMPROVEMENT. [THE PRIMARY AUTHOR WAS UDAY B KHANOLKAR, NARAYANA HRUDYALAYA, BENGALURU, INDIA. THE CORRESPONDING AUTHOR WAS MEEMANSA KASHYAP BUCH, NARAYANA HRUDYALAYA, BENGALURU, INDIA.] THE TIME FRAME OF THE STUDY WAS FROM 2020 TO 2023. A TOTAL OF 42 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL PATIENTS RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 50.11 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HEMOLYTIC ANEMIA AND CONGESTIVE CARDIAC FAILURE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER DUCT OCCLUDER WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEMOLYTIC ANEMIA AND CONGESTIVE CARDIAC FAILURE. COMPLICATIONS REPORTED INCLUDED DEATH, CARDIAC ARREST, PSEUDOANEURYSM, CARDIAC TAMPONADE, ENDOCARDITIS. THE REPORTED IFU DEVIATION/OFF-LABEL USE WAS ASSOCIATED WITH IMPLANTING AMPLATZER DUCT OCCLUDER FOR PARAVALVULAR LEAK CLOSURE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. IT SHOULD BE NOTED THAT THE AMPLATZER DUCT OCCLUDER, INSTRUCTION FOR USE (IFU), ARTMT600034251, REV B, STATES: THE AMPLATZER¿ DUCT OCCLUDER IS A PERCUTANEOUS, TRANSCATHETER OCCLUSION DEVICE INTENDED FOR THE NONSURGICAL CLOSURE OF A PATENT DUCTUS ARTERIOSUS (PDA). B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE: ARTICLE TITLE: "PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE." Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. ARTICLE TITLE: "PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773639&pc=KRD | 2135147-2025-01728 | 2020-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-04-04 04:00 | KRD | AMPLATZER VASCULAR PLUG II | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K071699 |  | 1 | Event Description: THE ARTICLE, "PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, SINGLE CENTER STUDY TO DETERMINE THE SAFETY AND EFFICACY OF PERCUTANEOUS DEVICE CLOSURE FOR SIGNIFICANT PARAVALVULAR LEAKS (PVLS). DEVICES INCLUDED IN THE STUDY WERE AMPLATZER VASCULAR PLUG II, AMPLATZER VASCULAR PLUG III, AND AMPLATZER DUCT OCCLUDER. THE ARTICLE CONCLUDED THAT TRANSCATHETER PVL CLOSURE IS A SAFE AND EFFECTIVE ALTERNATIVE TO SURGICAL CORRECTION AND SHOULD BE USED IN SELECTED PATIENTS FOR SIGNIFICANT AMELIORATION OF HEART FAILURE SYMPTOMS AS WELL AS HEMOLYTIC ANEMIA WITH CLINICAL AND HEMODYNAMIC IMPROVEMENT. [THE PRIMARY AUTHOR WAS UDAY B KHANOLKAR, NARAYANA HRUDYALAYA, BENGALURU, INDIA. THE CORRESPONDING AUTHOR WAS MEEMANSA KASHYAP BUCH, NARAYANA HRUDYALAYA, BENGALURU, INDIA.] THE TIME FRAME OF THE STUDY WAS FROM 2020 TO 2023. A TOTAL OF 42 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL PATIENTS RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 50.11 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HEMOLYTIC ANEMIA AND CONGESTIVE CARDIAC FAILURE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER VASCULAR PLUG II WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEMOLYTIC ANEMIA AND CONGESTIVE CARDIAC FAILURE. COMPLICATIONS REPORTED INCLUDED DEATH, CARDIAC ARREST, PSEUDOANEURYSM, CARDIAC TAMPONADE, ENDOCARDITIS. THE REPORTED IFU DEVIATION/OFF-LABEL USE WAS ASSOCIATED WITH IMPLANTING AMPLATZER VASCULAR PLUG II FOR PARAVALVULAR LEAK CLOSURE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. IT SHOULD BE NOTED THAT THE AMPLATZER VASCULAR PLUG II, INSTRUCTION FOR USE (IFU), ARTMT600034447, REV B, STATES: THE AMPLATZER¿ VASCULAR PLUG II IS INDICATED FOR ARTERIAL AND VENOUS EMBOLIZATIONS IN THE PERIPHERAL VASCULATURE. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: ARTICLE TITLED " PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE." Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ARTICLE TITLED " PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773770&pc=NGV | 2124215-2025-20695 | 2025-03-26 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-04 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Pericardial Effusion | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PERICARDIAL EFFUSION. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT AND THEN INSERTED THE WDS. THE CLOSURE DEVICE WAS DEPLOYED AND RECAPTURED THREE TIMES BEFORE ALL RELEASE CRITERIA WAS MET. THE CLOSURE DEVICE WAS RELEASED FROM THE CORE WIRE OF THE WDS AND SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. THE PATIENT WAS CHECKED FOR AN EFFUSION POST RELEASE AND A PERICARDIAL EFFUSION WAS DISCOVERED. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS WITH 800ML OF BRIGHT RED BLOOD REMOVED FROM THE PATIENT. THE PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION FOR A BRIEF AMOUNT OF TIME WHILE THEY WERE RECEIVING TREATMENT. THE EFFUSION RESOLVED AND THE PATIENT WAS STABILIZED. THE PATIENT WAS ADMITTED TO THE HOSPITAL TO RECOVER FROM THIS EVENT. NO OTHER COMPLICATIONS WERE REPORTED TO HAVE OCCURRED, AND THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PERICARDIAL EFFUSION. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT AND THEN INSERTED THE WDS. THE CLOSURE DEVICE WAS DEPLOYED AND RECAPTURED THREE TIMES BEFORE ALL RELEASE CRITERIA WAS MET. THE CLOSURE DEVICE WAS RELEASED FROM THE CORE WIRE OF THE WDS AND SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. THE PATIENT WAS CHECKED FOR AN EFFUSION POST RELEASE AND A PERICARDIAL EFFUSION WAS DISCOVERED. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS WITH 800ML OF BRIGHT RED BLOOD REMOVED FROM THE PATIENT. THE PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION FOR A BRIEF AMOUNT OF TIME WHILE THEY WERE RECEIVING TREATMENT. THE EFFUSION RESOLVED AND THE PATIENT WAS STABILIZED. THE PATIENT WAS ADMITTED TO THE HOSPITAL TO RECOVER FROM THIS EVENT. NO OTHER COMPLICATIONS WERE REPORTED TO HAVE OCCURRED, AND THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY. IT WAS FURTHER REPORTED THAT THE PATIENT WAS GIVEN A BLOOD TRANSFUSION DURING TREATMENT OF THE EFFUSION. TWO (2) DAYS POST PROCEDURE THE PATIENT DIED AS A RESULT OF THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773818&pc=OZD | 1220648-2025-27243 | 2025-03-13 04:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA 5.5 PUMP PLACED FOR SUPPORT, ALONG WITH EXTRACORPOREAL MEMBRANE OXYGENATION, FOR CARDIOMYOPATHY. IT WAS REPORTED THAT DURING SUPPORT, THERE WAS A CLOT ON THE PUMP OUTLET AND IN THE RIGHT LEG AND MEDIASTINUM THAT PROMPTED A SURGICAL WASHOUT. ADDITIONALLY, THE PATIENT BEGAIN HAVING HIGH PURGE PRESSURES AND BLOCKED ALARMS. THE PURGE WAS CHANGED FROM BICARBONATE TO TISSUE PLASMINOGEN ACTIVATOR AND THE PURGE PRESSURES IMPROVED. THE PATIENT REMAINED ON SUPPORT IN CRITICAL CARE IN A TENUOUS STATE. AFTER SOME TIME, THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH.". IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21773873&pc=DSQ | 2916596-2025-02201 | 2025-03-16 04:00 | Death | THORATEC CORPORATION | 2025-04-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE DEVICE WILL NOT BE EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21774124&pc=NPT | 2015691-2025-02593 | 2025-03-13 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-04 04:00 | NPT | SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS FRANCE AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE WITH A 26MM SAPIEN 3 ULTRA VALVE, THE PACING WAS PERFECTLY STABLE, AND THE VALVE WAS DEPLOYED IN A GOOD POSITION (70/30) WITHOUT PARAVALVULAR LEAKAGE. HOWEVER, 5 MINUTES LATER, THE PATIENT'S PRESSURE DROPPED, AND THE VALVE WAS OBSERVED TO BE IN THE LEFT VENTRICULAR OUTFLOW TRACT. THE VALVE HAD TURNED AROUND AND WAS UPSIDE DOWN. SURGERY WAS ATTEMPTED TO EXPLANT THE VALVE, BUT PATIENT PASSED AWAY. PER REPORT, THE PATIENT'S NATIVE ANNULUS HAD AORTIC INSUFFICIENCY WITHOUT CALCIFICATION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MIGRATION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER VALVE REPLACEMENT (TVR). ACCORDING TO THE LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE VENTRICLE. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT AND PROCEDURAL FACTORS (LACK OF CALCIFICATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN A PATIENT WITH AORTIC INSUFFICIENCY WITHOUT CALCIFICATION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO AORTIC INSUFFICIENCY PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21774337&pc=OZD | 1220648-2025-27246 | 2025-03-09 05:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED A GASTROINTESTINAL BLEED DURING IMPELLA 5.5 SUPPORT. HEPARIN WAS BEING ADMINISTERED THROUGH THE PURGE AT THE TIME OF THE BLEED. ONE UNIT OF PACKED RED BLOOD CELLS WAS GIVEN TO COUNTERACT THE BLEED. THE BLEED PERSISTED FOR ROUGHLY TWO WEEKS FOLLOWING THE INITIAL INTERVENTION. CARE WAS EVENTUALLY WITHDRAWN AFTER THIRTY DAYS OF SUPPORT AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH THE NOTED CONDITION DURING SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27246.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21774440&pc=NKM | 2135147-2025-01746 | 2013-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-04-04 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER EDGE-TO-EDGE REPAIR OF ATRIAL SECONDARY MITRAL REGURGITATION POSITIVELY INFLUENCES ATRIAL REMODELLING" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE ALL CONSECUTIVE PATIENTS TREATED WITH TEER FOR MITRAL REGURGITATION (MR) IN OUR CENTRE. DEVICES MENTIONED INCLUDE MITRACLIP AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT TRANSCATHETER MITRAL VALVE REPAIR BY EDGE-TO-EDGE THERAPY REPRESENTS A SAFE AND EFFECTIVE THERAPEUTIC OPTION IN SYMPTOMATIC PATIENTS WITH ATRIAL SECONDARY MITRAL REGURGITATION AND MIGHT HAVE THE POTENTIAL TO INDUCE LEFT ATRIAL REVERSE REMODELLING. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS ANIELA PETRESCU AT UNIVERSITY MEDICAL CENTER MAINZ OF THE JOHANNES GUTENBERG INSTITUTION WITH CORRESPONDING EMAIL ANIELA.M.POPESCU@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS JANUARY 2013 TO OCTOBER 2023. A TOTAL OF 103 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 88 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, PREVIOUS MYOCARDIAL INFARCTION, PREVIOUS STROKE. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, ARRHYTHMIA, RECURRENT MR, ADDITIONAL MITRACLIP PROCEDURE, HEART FAILURE, SHOCK. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, PREVIOUS MYOCARDIAL INFARCTION, PREVIOUS STROKE. COMPLICATIONS REPORTED INCLUDED DEATH, ARRHYTHMIA, RECURRENT MR, ADDITIONAL MITRACLIP PROCEDURE, HEART FAILURE, SHOCK; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21774855&pc=OZD | 1220648-2025-27248 | 2025-03-07 05:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA CP | Pumping Stopped | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT UNDERGOING A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT, AND EXPERIENCED A PUMP STOP UNEXPECTEDLY DURING THE SUPPORT. IT WAS NOTED THAT THE AUTOMATED IMPELLA CONTROLLER SCREEN RETURNED TO THE HOME SCREEN AND SHOWED A PROMPT STATING "PRESS OK TO RESUME DISPLAYED PERFORMANCE LEVEL" AT THE TIME OF THE FAILURE. THE PUMP RESTARTED ON ITS OWN. HOWEVER, IT WAS EXPLANTED FOLLOWING PCI COMPLETION. THE IMPELLA CP WAS UNEVENTFULLY REMOVED. THE PATIENT NOTED TO BE HEMODYNAMICALLY STABLE, AND HEMOSTASIS WAS ACHIEVED WITH TWO PERCLOSE. HOWEVER, FOLLOWING THE EXPLANT, THE PATIENT EXPIRED. NO DETAILS SURROUNDING THE DEMISE OUTCOME WAS PROVIDED. THE IMPACT OF THE TEMPORARY PUMP STOP IS UNCLEAR. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-27153 FOR THE AUTOMATED IMPELLA CONTROLLER DEVICE REPORT. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: ENTERING CARDIAC OUTPUT: ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: TROUBLESHOOTING THE PURGE SYSTEM: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED: ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21774886&pc=NPT | 2015691-2025-02601 | 2025-03-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-04 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM BRAZIL, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. THE PATIENT PRESENTED WITH SEVERE AORTIC CALCIFICATION WITH A CORONARY HEIGHT OF 8.3MM. THE VALVE WAS IMPLANTED AND THE PROCEDURE WAS SUCCESSFULLY COMPLETED. APPROXIMATELY 5 HOURS AFTER THE END OF THE PROCEDURE, THE PATIENT HAD A RIGHT CORONARY OCCLUSION REQUIRING A PCI; HOWEVER, THE PATIENT PASSED AWAY. AS PER MEDICAL OPINION, THE ROOT CAUSE OF THE CORONARY OCCLUSION WAS THE LOW CORONARY HEIGHT. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (SEVERE AORTIC CALCIFICATION & LOW CORONARY HEIGHT) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21774930&pc=LMH | 3005113652-2025-00314 |  | Death | PRINGY | 2025-04-04 04:00 | LMH | JUVEDERM (VOLUME/CONCENTRATION/ADDITIVE UNKNOWN) | Adverse Event Without Identified Device or Use Problem | Anaphylactic Shock; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P050047 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL (HCP) REPORTED A PATIENT DIED OF ANAPHYLAXIS FROM BOTOX® AND UNSPECIFIED JUVÉDERM PRODUCT. THE PATIENT RECEIVED BOTOX® AND UNSPECIFIED JUVÉDERM PRODUCT AT THEIR HOME AND LATER WENT INTO THE HOSPITAL WITH ANAPHYLAXIS AND THEY THOUGHT THE PATIENT HAD A HEART ATTACK BUT THEN SAID IT WAS AN ALLERGIC REACTION. Manufacturer Narrative: CLARIFICATION TO H6: TYPE OF INVESTIGATION CODE: THE FILLER WAS INJECTED INTO THE PATIENT AND IS NOT ACCESSIBLE FOR RETURN. THE SYRINGE WAS NOT RETURNED FOR EVALUATION. THE EVENT IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ABBVIE IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT. FURTHER INFORMATION REGARDING EVENT, PRODUCT, OR PATIENT DETAILS HAS BEEN REQUESTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. ANAPHYLAXIS A KNOWN POTENTIAL ADVERSE EVENT ADDRESSED IN THE PRODUCT LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21775108&pc=DSQ | 2916596-2025-02218 | 2025-03-25 04:00 | Death | THORATEC CORPORATION | 2025-04-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS NOT REPORTED WHETHER THIS WAS DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE CAUSE OF DEATH WAS NOT PROVIDED DUE TO HOSPITAL PRIVACY REASONS. THIS WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21775827&pc=NPT | 2025587-2025-02368 | 2024-06-25 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-04 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Inadequacy of Device Shape and/or Size | Insufficient Information | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) IN PATIENTS WITH SMALL AORTIC ANNULI (SAA). THE STUDY POPULATION CONSISTED OF 258 PATIENTS WITH SAA WHO UNDERWENT TAVR WITH EITHER A NON-MEDTRONIC BALLOON-EXPANDABLE VALVE (BEV = 90) OR A MEDTRONIC SELF-EXPANDING VALVE (SEV = 168). IN THE SEV GROUP, THE FOLLOWING ADVERSE OUTCOMES OCCURRED: AORTIC REGURGITATION (MODERATE TO SEVERE), PATIENT-PROSTHESIS MISMATCH (PPM), STROKE/DISABLING STROKE, MYOCARDIAL INFARCTION, REHOSPITALIZATION FOR HEART FAILURE, AND REINTERVENTION REQUIRED (SURGICAL EXPLANT OR VALVE-IN-VALVE REPLACEMENT) DUE TO VALVE MALPOSITION OR STRUCTURAL VALVE DETERIORATION. ADDITIONALLY, AMONG SEV PATIENTS, THE AUTHORS OBSERVED 4 IN-HOSPITAL DEATHS AND AN ADDITIONAL 38 DEATHS WITHIN THREE YEARS OF TAVR. WHILE THE AUTHORS FOUND THAT THE OCCURRENCE OF PPM HAD ¿NO DISCERNIBLE EFFECT¿ ON THE OUTCOMES OF DEATH, STROKE, AND MYOCARDIAL INFARCTION, A POTENTIAL DEVICE RELATIONSHIP TO THESE OUTCOMES CANNOT BE EXCLUDED. Manufacturer Narrative: CITATION: BASHIR H, SCHMIDT CW, ANSAH K, ET AL. TRANSCATHETER AORTIC VALVE REPLACEMENT IN PATIENTS WITH SMALL AORTIC ANNULUS: AN OBSERVATIONAL STUDY. STRUCT HEART. 2024;9(2):100338. PUBLISHED 2024 JUN 25. DOI:10.1016/J.SHJ.2024.100338 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. MEDTRONIC VALVE BRANDS USED IN THE STUDY: EVOLUT R (PRODUCT CODE NPT, PMA# P130021) AND EVOLUT PRO (PRODUCT CODE NPT, PMA# P130021). EARLIEST APPROVED PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21776353&pc=OZD | 1220648-2025-27250 | 2025-03-11 04:00 | Death | ABIOMED, INC. | 2025-04-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Renal Failure; Ventricular Fibrillation; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT WAS SUPPORTED FOR 92 HOURS BEFORE THE PATIENT WAS EXPLANTED AND SUPPORTED BY EXTRACORPOREAL MEMBRANE OXYGENATION. IN 2025, A PUBLISHED ARTICLE WAS IDENTIFIED. THE AUTHORS NOTED THE FOLLOWING: "WE PERFORMED THE BENTALL PROCEDURE, MITRAL VALVE REPAIR, TRICUSPID ANNULOPLASTY, AND CORONARY ARTERY BYPASS GRAFT ON A 73-YEAR-OLD MAN ADMITTED WITH CONGESTIVE HEART FAILURE CAUSED BY ANNULOAORTIC ECTASIA, ALONG WITH SEVERE AORTIC AND MITRAL REGURGITATION. VA-ECMO AND IMPELLA WERE REQUIRED, SINCE THE CARDIOPULMONARY BYPASS WEANING WAS DIFFICULT. IMPELLA WAS REMOVED ON POSTOPERATIVE DAY 4. ON POSTOPERATIVE DAYS 5 AND 6, LABORATORY DATA SHOWED WORSENING RENAL DYSFUNCTION, LACTATE LEVELS, AND ACIDOSIS. CONTRAST-ENHANCED COMPUTED TOMOGRAPHY SHOWED THROMBOSIS IN THE CELIAC AND SUPERIOR MESENTERIC ARTERIES. AORTIC THROMBECTOMY WAS PERFORMED. HYPERKALEMIA, CAUSED BY A REPERFUSION INJURY, RESULTED IN VENTRICULAR FIBRILLATION. CONTINUOUS HEMODIAFILTRATION IMPROVED THE HYPERKALEMIA. HOWEVER, IRREVERSIBLE ACIDOSIS PROGRESSED, AND THE VA-ECMO FLOW RATE COULD NOT BE SUSTAINED. ON POSTOPERATIVE DAY 7, THE PATIENT DIED." NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. LITERATURE CITATION: YAMANA, F., DOMAE, K., KAWASUMI, R., SAKAMOTO, T., HATA, M., SHIRAKAWA, Y., MASAI, T., & SAWA, Y. (2023). AORTIC THROMBOSIS WITH VISCERAL MALPERFUSION DURING CIRCULATORY SUPPORT WITH A COMBINATION OF IMPELLA AND EXTRACORPOREAL MEMBRANE OXYGENATION FOR POSTCARDIOTOMY CARDIOGENIC SHOCK. JOURNAL OF ARTIFICIAL ORGANS, 26(4), 330¿334. HTTPS://DOI.ORG/10.1007/S10047-023-01382-3. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21776967&pc=OAE | 2029046-2025-01038 |  | Death | BIOSENSE WEBSTER INC | 2025-04-04 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pleural Effusion; Ventricular Fibrillation; Cardiac Tamponade | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CARDIAC ABLATION WITH A QDOT MICRO AND THE PATIENT EXPERIENCED CARDIAC TAMPONADE, VENTRICULAR FIBRILLATION, CARDIAC ARREST THAT REQUIRED CPR AND ECMO; HOWEVER, IT RESULTED IN DEATH. IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC TAMPONADE, VENTRICULAR FIBRILLATION, AND CARDIAC ARREST. THE PATIENT EXPIRED AFTER THE PROCEDURE. A LEFT ACCESSORY PATHWAY MAP WAS PERFORMED FOLLOWED BY ABLATION OF THE LEFT LATERAL LEFT ATRIUM (LA) WITH A NON-BOSTON SCIENTIFIC RF CATHETER. AFTER PERFORMING THIS ABLATION, A PLEURAL EFFUSION WAS NOTICED THAT BECAME VERY LARGE VERY FAST. A "PERICARDIAL PERFUSION" WAS PERFORMED TO CLEAR THE EFFUSION, BUT THEN THE PATIENT WENT INTO CARDIAC ARREST WITH VENTRICULAR FIBRILLATION (VF). CARDIOPULMONARY RESUSCITATION WAS PERFORMED ALONG WITH MULTIPLE SHOCKS FROM CARDIOVERSION. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), THEN TRANSFERRED TO ANOTHER HOSPITAL AND ADMITTED. THE PATIENT WAS NOT CONSIDERED STABLE AT THE TIME OF ADMISSION AND REMAINED ON LIFE SUPPORT. THE PATIENT LATER EXPIRED. Manufacturer Narrative: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21777075&pc=NKM | 2135147-2025-01771 | 2025-03-13 04:00 | Death | ABBOTT MEDICAL | 2025-04-04 04:00 | NKM | MITRACLIP | Positioning Failure; Difficult to Remove | Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4 AND AN ENLARGED ATRIUM. ONE CLIP WAS SUCCESSFULLY IMPLANTED. TO FURTHER REDUCE MR, AN ADDITIONAL CLIP WAS INSERTED, BUT AFTER SEVERAL GRASPING ATTEMPTS, TISSUE DAMAGE WAS OBSERVED ON ONE OF THE LEAFLET. THEREFORE, THE PHYSICIAN DECIDED TO REMOVE THE CLIP. IT WAS NOTED THAT WHILE REMOVING THE CLIP, RESISTANCE WAS OBSERVED AND THE CLIP WAS UNABLE TO BE REMOVED. TROUBLESHOOTING WAS PERFORMED AND THE CLIP WAS ABLE TO BE REMOVED. ONCE REMOVED, TISSUE WAS OBSERVED ON THE GRIPPER OF THE CLIP. THE PROCEDURE WAS DISCONTINUED. MR WAS REDUCED TO A GRADE OF 3. THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. ON (B)(6) 2025, THE PATIENT EXPERIENCED RECURRENT MR, HEART FAILURE. THE PATIENT EXPIRED. Manufacturer Narrative: ALL AVAILABLE INFORMATION WAS INVESTIGATED, AND THE REPORTED DIFFICULT REMOVING CDS FROM SGC WAS NOT CONFIRMED VIA DEVICE ANALYSIS. THE REPORTED UNABLE TO GRASP LEAFLETS COULD NOT BE REPLICATED IN A TESTING ENVIRONMENT. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON AVAILABLE INFORMATION AND THE RETURNED DEVICE ANALYSIS, THE CAUSE OF THE REPORTED DIFFICULT REMOVING CDS FROM SGC AND UNABLE TO GRASP LEAFLETS WERE UNABLE TO BE DETERMINED. THE CAUSE OF THE REPORTED TISSUE INJURY, RECURRENT MR, HEART FAILURE, AND DEATH WERE UNABLE TO BE DETERMINED. THE REPORTED PATIENT EFFECTS OF TISSUE INJURY, MITRAL REGURGITATION, HEART FAILURE, AND DEATH, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED SERIOUS INJURY/ ILLNESS/ IMPAIRMENT WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE IS EXPECTED TO BE RETURNED FOR INVESTIGATION. IT HAS NOT YET BEEN RECEIVED. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21777902&pc=JXG | 3013886523-2025-00069 |  | Death | INTEGRA LIFESCIENCES MANSFIELD | 2025-04-04 04:00 | JXG | HAKIM PROGRAMMABLE VALVE | Infusion or Flow Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K221840 |  | 1 | Event Description: THIS IS 1 OF 2 REPORTS LINKED TO MFG REPORT NUMBERS: 3013886523-2025-00068. A PHYSICIAN REPORTED: ON AN UNSPECIFIED DATE, A CODMAN HAKIM PROGRAMMABLE VALVE (CHPV) WAS IMPLANTED FOR AN UNREPORTED INDICATION. ON AN UNKNOWN DATE, THE PATIENT HAD INEFFICIENCY MANAGING THE CSF (CEREBROSPINAL FLUID) FLOW. THE PATIENT TRAGICALLY PASSED AWAY ON AN UNKNOWN DATE LAST YEAR DUE TO VENTRICULAR COLLAPSE, WHICH WAS ATTRIBUTED TO THE VALVE'S PERFORMANCE BY THE PHYSICIAN. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. NO PATIENT INFORMATION PROVIDED INCLUDING: AGE, GENDER, CONCURRENT CONDITION, FAMILY HISTORY PAST MEDICATION, CONCOMITANT MEDICATION AND CO-SUSPECT MEDICATIONS, DRUG ALLERGIES. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED AFTER SEVERAL ATTEMPTS.Event Description: N/A Manufacturer Narrative: UPDATED FIELDS: D9, G3, G6, H2, H3, H6, H11. THE HAKIM VALVE (UNKNOWN ID) WAS NOT RETURNED FOR EVALUATION AFTER THREE GOOD FAITH ATTEMPTS (GFES) WERE MADE. LOT NUMBER INFORMATION HAS NOT BEEN PROVIDED; THEREFORE, AN EVALUATION OF THE DEVICE COULD NOT BE PERFORMED, AND DEVICE HISTORY RECORD (DHR) COULD NOT BE REVIEWED. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED. A MEDICAL ASSESSMENT WAS REQUESTED TO INTEGRA'S MEDICAL SAFETY TEAM: "THERE IS NO INDICATION THAT THE DEVICES MALFUNCTIONED BUT DO SUGGEST THAT PROCEDURAL DEVIATIONS FROM RECOMMENDED INSTRUCTIONS MAY HAVE CONTRIBUTED TO THE EVENTS, AS DEVICES WERE NOT USED AND MONITORED IN ACCORDANCE WITH ALL INSTRUCTIONS INCLUDED IN THE INSTRUCTIONS FOR USE (IFU). THE INFORMATION GATHERED DOES SUGGEST USE ERROR ONLY. IT APPEARS PROGRAMMING WAS PERFORMED BY UNTRAINED RESIDENTS WHO MAY NOT HAVE POSITIONED THE TRANSMITTER PROPERLY ON THE VALVE NOR RESTRICTED MOVEMENT OF THE COIL DURING PROGRAMING, OR OBSERVED THE PROPER OPERATING TEMPERATURE OF THE PROGRAMMER. ADDITIONALLY, THEY DID NOT TAKE CONFORMATIONAL RADIOGRAPHS POST-PROGRAMMING NOR PRE-PROGRAMMING CONFIRM THE SETTING AND IDENTIFY IF THERE HAD BEEN ANY CHANGE OF VALVE SETTING DUE TO STRONG MAGNETIC FORCE. THE HOSPITAL DOES NOT USE EVDÍS ROUTINELY DURING STABILIZATION OF THEIR TRAUMA PATIENTS BUT INSTEAD USE SHUNTS WHERE FIBRIN CLOTS MAY FORM AND DISRUPT VALVE PROGRAMMING. AS PER A CONFERENCE CALL WITH DR. (B)(6) MEDICAL ADVISOR DR. (B)(6). ON (B)(6) 2025, DR. (B)(6) GROUP WILL BE MEETING WITH DR.(B)(6). TO DISCUSS TRAINING, AND CONFIRM ANY ADDITIONAL COMPLAINTS THAT MAY HAVE OCCURRED AND SHOULD BE REPORTED TO INTEGRA." Manufacturer Narrative: AN INVESTIGATION HAS BEEN INITIATED BASED ON THE REPORTED INFORMATION. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW-UP REPORT WILL BE SUBMITTED. Manufacturer Narrative: UPDATED FIELDS: D9, G3, G6, H2, H3, H11. A MEDICAL ASSESSMENT WAS PROVIDED BY INTEGRA'S MEDICAL AFFAIRS DIRECTOR AS A RESULT OF A CONVERSATION WITH THE FACILITY PHYSICIAN: "THE EVENTS, AS DESCRIBED BY DR. (B)(6), DO NOT IDENTIFY NOR DO THEY INDICATE A CAUSAL EFFECT FROM THE CODMAN HAKIM PROGRAMMABLE VALVES. THERE IS NO INDICATION THAT THE DEVICES MALFUNCTIONED BUT DO SUGGEST THAT PROCEDURAL DEVIATIONS FROM RECOMMENDED INSTRUCTIONS MAY HAVE CONTRIBUTED TO THE EVENTS, AS DEVICES WERE NOT USED AND MONITORED IN ACCORDANCE WITH ALL INSTRUCTIONS INCLUDED IN THE INSTRUCTIONS FOR USE (IFU). THE INFORMATION GATHERED DOES SUGGEST USE ERROR ONLY. IT APPEARS PROGRAMMING WAS PERFORMED BY UNTRAINED RESIDENTS WHO MAY NOT HAVE POSITIONED THE TRANSMITTER PROPERLY ON THE VALVE NOR RESTRICTED MOVEMENT OF THE COIL DURING PROGRAMING, OR OBSERVED THE PROPER OPERATING TEMPERATURE OF THE PROGRAMMER. ADDITIONALLY, THEY DID NOT TAKE CONFORMATIONAL RADIOGRAPHS POST-PROGRAMMING NOR PRE-PROGRAMMING CONFIRM THE SETTING AND IDENTIFY IF THERE HAD BEEN ANY CHANGE OF VALVE SETTING DUE TO STRONG MAGNETIC FORCE. THE HOSPITAL DOES NOT USE EVD S ROUTINELY DURING STABILIZATION OF THEIR TRAUMA PATIENTS BUT INSTEAD USE SHUNTS WHERE FIBRIN CLOTS MAY FORM AND DISRUPT VALVE PROGRAMMING. AS PER A CONFERENCE CALL WITH DR. (B)(6) MEDICAL ADVISOR, DR. (B)(6) GROUP WILL BE MEETING WITH THE PHYSICIAN TO DISCUSS TRAINING, AND CONFIRM ANY ADDITIONAL COMPLAINTS THAT MAY HAVE OCCURRED AND SHOULD BE REPORTED TO INTEGRA."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778018&pc=LWS | 9614453-2025-01347 | 2025-03-07 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-04 04:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) THE PATIENT WAS GOING IN AND OUT VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION. THE PATIENT WAS WEARING A LIFEVEST UPON ARRIVAL FOR THE PROCEDURE AND IT WAS KNOWN THEY HAD PRIOR HEART RELATED COMPLICATIONS. THE DEVICE IMPLANT WAS SUCCESSFUL AND TESTING NUMBERS WERE NOTED TO BE ADEQUATE. HOWEVER, THE PATIENT REQUIRED THREE EXTERNAL SHOCKS AND CARDIOPULMONARY RESUSCITATION. A CODE WAS CALLED. UNFORTUNATELY, THE PATIENT WAS PRONOUNCED DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778027&pc=DXO | 3008061490-2025-00001 | 2025-03-07 05:00 | Death | OPSENS | 2025-04-04 04:00 | DXO | SAVVYWIRE | Capturing Problem; Insufficient Information | Cardiac Arrest | K213854 |  | 1 | Event Description: EVENT AS DESCRIBED BY OPSENS SALES REPRESENTATIVE: PROCEDURE TYPE: TAVR. VALVE: EDWARDS SAPIEN 26MM. LOSS OF CAPTURE DURING DEPLOYMENT OF THE VALVE. THE EDWARDS VALVE EMBOLIZED. INTERVENTION WAS PERFORMED TO MOVE THE VALVE TO AN APPROPRIATE AREA IN THE AORTA. THE PATIENT DID NOT SURVIVE DESPITE THEIR EFFORTS. PACING RATE 180, OUTPUT OF 20MA. VOO WAS CONFIRMED. BEFORE THE DEPLOYMENT OF THE VALVE TEST PACING WAS CONDUCTED. PACING RATE OF 120, OUTPUT OF 20MA WALKED DOWN UNTIL LOSS OF CAPTURE AT 10MA. Manufacturer Narrative: THE SAVVYWIRE WAS DISCARDED AND NOT RETURNED TO THE MANUFACTURER. NO INSPECTION WAS POSSIBLE. NO IMAGING AND NO ADDITIONAL DATA WERE PROVIDED BY THE ACCOUNT AS OF APRIL 4TH, 2025. CHECK OF THE DEVICE HISTORY RECORDS (DHR) OF GUIDEWIRE LOT NUMBER OSW-0323 SHOWED THAT IT WAS RELEASED PER SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION, PACING TEST WAS PERFORMED BEFORE VALVE DEPLOYMENT, AS PRESCRIBED, CONFIRMING VENTRICULAR CAPTURE AND THRESHOLDS. WITHOUT INSPECTING THE WIRE, OPSENS IS NOT ABLE TO FULLY INVESTIGATE THEREFORE NO DEFINITIVE ROOT CAUSE CAN BE CONCLUDED IN THIS REPORT. IF THERE IS ANY FURTHER RELEVANT INFORMATION, A FOLLOW UP REPORT WILL BE SUBMITTED TO THE FDA. THE ADVERSE EVENT RESULT OF THIS INCIDENT IS WELL MENTIONED IN THE SAVVYWIRE INSTRUCTIONS OF USE (IFU): ADVERSE EVENTS THAT MAY RESULT FROM THE USE OF THIS DEVICE INCLUDE, BUT ARE NOT LIMITED TO: ACCESS SITE OR VESSELS COMPLICATIONS, ADDITIONAL SURGICAL PROCEDURE, ALLERGIC REACTIONS, AMPUTATION, ANEURYSM, ANGINA, ARRHYTHMIA, BLEEDING, CARDIAC OR VESSEL PERFORATION/DISSECTION, CORONARY OBSTRUCTION, DEATH, EMBOLISM, FIBRILLATION, FOREIGN BODY/WIRE FRACTURE, HEART BLOCK, HEMATOMA, HYPOTENSION/HYPERTENSION, INFECTION, KIDNEY INJURY/FAILURE, MYOCARDIAL INFARCTION, NEED FOR PERMANENT PACEMAKER, PERICARDIAL EFFUSION, PNEUMOTHORAX, STROKE OR OTHER NEUROLOGIC EVENT, SPASM, TAMPONADE, THROMBUS, VALVE DYSFUNCTION OR COMPLICATIONS, VALVE MALPOSITIONING OR EMBOLIZATION, VASOSPASM, VESSEL OCCLUSION, WIRE ENTRAPMENT/ENTANGLEMENT, X-RAY RADIATION EXPOSURE COMPLICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778566&pc=FKX | 0008030665-2025-00781 | 2025-03-31 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-04 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED TO THE CLINIC. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, TREATMENT RECORD, HOSPITAL RECORDS, DEATH CERTIFICATE) WERE UNSUCCESSFUL.Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED TO THE CLINIC. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, TREATMENT RECORD, HOSPITAL RECORDS, DEATH CERTIFICATE) WERE UNSUCCESSFUL. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE CAUSE OF DEATH IS CURRENTLY UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THE LACK OF CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE INVESTIGATION. HOWEVER, THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. WHILE THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. A LACK OF CLINICAL FOLLOW-UP INFORMATION (E.G., PATIENT DEMOGRAPHICS, TREATMENT RECORD, HOSPITAL RECORDS, DEATH CERTIFICATE), AND NO EVALUATION OF THE LIBERTY SELECT CYCLER IN QUESTION, RENDERS THIS CLINICAL INVESTIGATION UNABLE TO EXCLUDE THE LIBERTY SELECT CYCLER FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER AND SHOWED BEZEL DAMAGE. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. AN AS-RECEIVED WITH REDUCED DWELL SIMULATED TREATMENT WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST AND VALVE ACTUATION TEST. AN INVESTIGATION OF THE CYCLER MUSHROOM HEADS VERIFIED THAT THE SURFACE CONDITIONS AND ALIGNMENTS WERE WITHIN SPECIFICATION. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778630&pc=MVK | 3008642652-2025-03553 | 2025-03-03 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-04 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise; Over-Sensing; Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHT APPROPRIATE TREATMENTS, FOUR INAPPROPRIATE TREATMENTS, AND FIVE NON-LIFEVEST DEFIBRILLATIONS. THE DEVICE WAS STARTED UP AT 11:23:27 ON (B)(6) 2025. AT 08:08:13 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 210 BPM WITH CPR/MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VT. CPR/MOTION ARTIFACT DELAYED THE LIFEVEST FROM DELIVERING A TREATMENT. AT 08:09:07, THE PATIENT RECEIVED THE FIRST NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE FIRST NON-LIFEVEST DEFIBRILLATIONS WAS VT @ 250 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 60 BPM. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 54 SECONDS INTO THE DETECTION SEQUENCE. AT 08:09:45, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS AF @ 110 BPM WITH PVCS AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 110 BPM WITH PVCS/NSVT AND CPR/MOTION ARTIFACT. AT 08:10:23, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS AF @ 90 BPM WITH PVCS/NSVT AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 110 BPM WITH PVCS/NSVT AND MOTION ARTIFACT. AT 08:11:48, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 220 BPM WITH MOTION ARTIFACT. AT 08:13:15, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 08:13:49, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 100 BPM WITH PVCS/NSVT AND CPR/MOTION ARTIFACT. AT 08:17:35, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 08:19:30, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 08:20:04, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 240 BPM. AT 08:22:00, THE PATIENT RECEIVED THE SECOND NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE SECOND NON-LIFEVEST DEFIBRILLATIONS WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 116 SECONDS AFTER THE PREVIOUS LIFEVEST SHOCK. AT 08:22:36, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 250 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 08:23:12 THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 08:23:44, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 60 BPM WITH PVCS/NSVT AND CPR/MOTION ARTIFACT. AT 08:26:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 08:26:42, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 40 BPM. AT 08:27:17, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 45 BPM WITH CPR/MOTION ARTIFACT. AT 08:31:18, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 08:32:12, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 130 BPM. AT 08:32:56, THE PATIENT RECEIVED THE THIRD NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE THIRD NON-LIFEVEST DEFIBRILLATIONS WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 50 BPM WITH CPR/MOTION ARTIFACT. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 44 SECONDS AFTER THE PREVIOUS LIFEVEST SHOCK. AT 08:37:55, THE PATIENT RECEIVED THE FOURTH NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE FOURTH NON-LIFEVEST DEFIBRILLATIONS WAS NSVT WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PVCS/NSVT AND CPR/MOTION ARTIFACT. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 31 SECONDS INTO THE DETECTION SEQUENCE. AT 08:46:45, THE PATIENT RECEIVED THE FIFTH NON-LIFEVEST DEFIBRILLATION. THE RHYTHM AT THE TIME OF THE FIFTH NON-LIFEVEST DEFIBRILLATIONS WAS VT @ 160 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF/SINUS RHYTHM FROM 5-60 BPM WITH PVCS AND BBB. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION 16 SECONDS INTO THE DETECTION SEQUENCE. THE ELECTRODE BELT WAS DISCONNECTED AT 08:47:31 ON 3/3/2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778636&pc=MVK | 3008642652-2025-03554 | 2025-03-02 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-04 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 04:18:53 ON (B)(6) 2025. AT 04:21:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 13:30:02 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778683&pc=NKM | 2135147-2025-01787 | 2010-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-04 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "RIGHT HEART FAILURE AND MORTALITY IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE INTERVENTIONS" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN RIGHT HEART FAILURE (RHF) AND MORTALITY IN PATIENTS WITH SEVERE TRICUSPID REGURGITATION (TR) UNDERGOING TRANSCATHETER TRICUSPID VALVE INTERVENTION (TTVI) AND TO DETERMINE WHETHER CLINICAL RHF STATUS REDUCES THE SURVIVAL BENEFIT OF SUCCESSFUL VERSUS FAILED TTVI. DEVICES MENTIONED INCLUDE (MITRACLIP, TRIALIGN,TRICINCH, CAVI, FORMA, CARDIOBAND, PASCAL. THE ARTICLE CONCLUDED THAT THE PRESENCE OF AT LEAST 2 RHF CLINICAL CRITERIA WAS INDEPENDENTLY ASSOCIATED WITH AN INCREASED RISK OF 1 -YEAR MORTALITY. PROCEDURAL SUCCESS WAS ASSOCIATED WITH A LOWER RISK OF MORTALITY REGARDLESS OF RHF STATUS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DR. MARIANNA ADAMO AT UNIVERSITY OF BRESCIA, BRESCIA, ITALY WITH CORRESPONDING EMAIL MARIANNAADAMO@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS AUGUST 2010 AND JANUARY 2023. A TOTAL OF 498 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 76. THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDE PRIOR CARDIA SURGERY, PRIOR MYOCARDIAL INFARCTION, DIABETES, ATRIAL FIBRILLATION, CARDIAC IMPLANTABLE ELECTRONIC DEVICE, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, COPD, HEART FAILURE, TRICUSPID REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, STROKE, INFECTION, ATRIAL FIBRILLATION, HEART FAILURE, PERICARDIAL EFFUSION, SURGICAL INTERVENTION ACUTE KIDNEY INJURY, BLOOD TRANSFUSION, BLEEDING.Event Description: THE ARTICLE "RIGHT HEART FAILURE AND MORTALITY IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE INTERVENTIONS" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN RIGHT HEART FAILURE (RHF) AND MORTALITY IN PATIENTS WITH SEVERE TRICUSPID REGURGITATION (TR) UNDERGOING TRANSCATHETER TRICUSPID VALVE INTERVENTION (TTVI) AND TO DETERMINE WHETHER CLINICAL RHF STATUS REDUCES THE SURVIVAL BENEFIT OF SUCCESSFUL VERSUS FAILED TTVI. DEVICES MENTIONED INCLUDE (MITRACLIP, TRIALIGN,TRICINCH, CAVI, FORMA, CARDIOBAND, PASCAL. THE ARTICLE CONCLUDED THAT THE PRESENCE OF AT LEAST 2 RHF CLINICAL CRITERIA WAS INDEPENDENTLY ASSOCIATED WITH AN INCREASED RISK OF 1 -YEAR MORTALITY. PROCEDURAL SUCCESS WAS ASSOCIATED WITH A LOWER RISK OF MORTALITY REGARDLESS OF RHF STATUS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DR. MARIANNA ADAMO AT UNIVERSITY OF BRESCIA, BRESCIA, ITALY WITH CORRESPONDING EMAIL MARIANNAADAMO@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS AUGUST 2010 AND JANUARY 2023. A TOTAL OF 498 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 76. THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDE PRIOR CARDIA SURGERY, PRIOR MYOCARDIAL INFARCTION, DIABETES, ATRIAL FIBRILLATION, CARDIAC IMPLANTABLE ELECTRONIC DEVICE, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, COPD, HEART FAILURE, TRICUSPID REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, STROKE, INFECTION, ATRIAL FIBRILLATION, HEART FAILURE, PERICARDIAL EFFUSION, SURGICAL INTERVENTION ACUTE KIDNEY INJURY, BLOOD TRANSFUSION, BLEEDING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE: ARTICLE TITLE: RIGHT HEART FAILURE AND MORTALITY IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE INTERVENTIONS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR CARDIA SURGERY, PRIOR MYOCARDIAL INFARCTION, DIABETES, ATRIAL FIBRILLATION, CARDIAC IMPLANTABLE ELECTRONIC DEVICE, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, COPD, HEART FAILURE, TRICUSPID REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, STROKE, INFECTION, ATRIAL FIBRILLATION, HEART FAILURE, PERICARDIAL EFFUSION, SURGICAL INTERVENTION ACUTE KIDNEY INJURY, BLOOD TRANSFUSION, BLEEDING; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED IFU DEVIATION/OFF-LABEL USE WAS ASSOCIATED WITH IMPLANTATION OF A MITRACLIP IN A TRICUSPID VALVE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE ADDITIONAL SERIOUS INJURIES REFERENCED IN B5 WILL BE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. B2: THE DEATH DATE IS ESTIMATED. B3: THE EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: RIGHT HEART FAILURE AND MORTALITY IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE INTERVENTIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21778711&pc=LPB | 2029046-2025-01042 | 2025-03-10 04:00 | Death | BIOSENSE WEBSTER INC | 2025-04-04 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A RIGHT SIDED ATRIAL FLUTTER ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS. THE PATIENT ALSO EXPERIENCED CARDIAC ARREST THAT REQUIRED CPR. THE PATIENT DIED. DURING A RIGHT-SIDED ATRIAL FLUTTER CASE AND AN AFIB ABLATION WITH A FLUTTER LINE, A PERICARDIAL EFFUSION WAS NOTICED IN THE PATIENT, BEFORE THEY PASSED AWAY. THEY HAD FINISHED ABLATING WITH THE THERMOCOOL SMARTTOUCH SF FOR THE FLUTTER AND THEY WENT TRANSSEPTAL, WHEN THE OCTARAY¿ CATHETER WAS ADVANCED INTO THE LEFT ATRIUM, THE PATIENT'S OXYGEN (O2) READINGS WERE LOW, AND THE PATIENT'S BLOOD PRESSURE DROPPED. THEY CONFIRMED ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WITH THE SOUNDSTAR® CATHETER, THAT THERE WAS A "LARGE" PERICARDIAL EFFUSION PRESENT. THE MEDICAL INTERVENTION PROVIDED WAS A PERICARDIOCENTESIS. AS THEY WERE TAPPING THE PATIENT, THEY CONFIRMED THE PERICARDIAL EFFUSION VIA ECHOCARDIOGRAM (ECHO). THEY THEN STARTED TO ADMINISTER CPR OR CHEST COMPRESSIONS TO THE PATIENT AS THEY WERE DRAINING FLUID. CATH AND OPERATING ROOM (OR) SUPPORT WERE THEN CALLED INTO THE ROOM, AND THEY BEGAN TO REVERSE HEPARIN TO THE PATIENT. EVENTUALLY THE PATIENT WAS SHOWING PULSELESS ELECTRICAL ACTIVITY(PEA). THEY WERE ABLE TO GET THE PATIENT'S PULSE BACK FOR ABOUT 3 MINUTES, BUT THE FLUID WAS STILL BUILDING UP IN THE PERICARDIAL SPACE. THEY ADMINISTERED FURTHER CHEST COMPRESSIONS, BUT THERE WAS NO PULSE IN THE PATIENT. THEY CALLED CODE FOR THE PATIENT, AND THE PATIENT THEN PASSED AWAY ON (B)(6) 2025. THEY DID NOT PERFORM OPEN-HEART SURGERY OR OPEN THE PATIENT'S CHEST, AND THEY PERFORMED CHEST COMPRESSIONS FOR ABOUT 30-40 MINUTES IN TOTAL. THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS THE EFFUSION, POSSIBLY WHEN GOING TRANSSEPTAL TO DO THE PULMONARY VEIN ISOLATION (PVI), HE THINKS HE COULD HAVE CAUSED IT ADJUSTING THE SHEATH. THE VERSACROSS RF WIRE AND STEERABLE SHEATH WERE USED FOR TRANSSEPTAL PUNCTURE AND THE FARADRIVE SHEATH WAS EXCHANGED AS WELL. IT WAS REPORTED THAT THE EVENT OCCURRED DURING THE TRANSSEPTAL PHASE; HOWEVER, ABLATION WAS PERFORMED PRIOR TO NOTING THE PERICARDIAL EFFUSION CAVOTRICUSPID ISTHMUS(CTI) WAS COMPLETED. THERE WAS NO EVIDENCE OF STEAM POP. NO ERROR MESSAGES OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THE PUMP PERFORMED AS EXPECTED. Manufacturer Narrative: THE BWI PRODUCT ANALYSIS LAB RECEIVED THE DEVICE FOR EVALUATION ON 25-APR-2025. THE ANALYSIS HAS BEGUN BUT IS NOT COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT A PATIENT UNDERWENT A RIGHT SIDED ATRIAL FLUTTER ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS. THE PATIENT ALSO EXPERIENCED CARDIAC ARREST THAT REQUIRED CPR. THE PATIENT DIED. THE DEVICE EVALUATION WAS COMPLETED ON 30-APR-2025. THE PRODUCT WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING JOHNSON & JOHNSON MEDTECH PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE DEVICE FEATURES WERE REVIEWED. NO MAGNETIC, FORCE, IRRIGATION, DEFLECTION, TEMPERATURE OR ELECTRICAL ISSUES WERE FOUND DURING THE PRODUCT INVESTIGATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO INTERNAL ACTIONS WERE IDENTIFIED. THE ROOT CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS THE PROCEDURE. IN ADDITION, NO DEVICE MALFUNCTION WAS REPORTED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE ANALYSIS. THE INSTRUCTION FOR USE (IFU) STATES THAT CAREFUL CATHETER MANIPULATION MUST BE PERFORMED TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. AS PART OF THE JOHNSON & JOHNSON MEDTECH QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. DURING THE DEVICE EVALUATION, RETRIEVED THE LOT NUMBER OF 31499754L. THEREFORE, PROCESSED D4. LOT, D4. EXPIRATION DATE, D4. PRIMARY UDI NUMBER AND H4. DEVICE MANUFACTURE DATE. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21780146&pc=NPT | 2025587-2025-02378 | 2024-05-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-04 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO SUDDEN CARDIAC ARREST AFTER BEING TRANSFERRED TO THE HOSPITAL.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, D4, H4, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21780660&pc=QLG | 2954323-2025-12136 | 2025-02-28 05:00 | Death | ABBOTT DIABETES CARE INC | 2025-04-04 04:00 | QLG | FREESTYLE LIBRE 2 | Insufficient Information | Insufficient Information | K193371 |  | 1 | Event Description: ABBOTT DIABETES CARE RECEIVED A MEDWATCH REPORT #MW5167182 WHICH REPORTED THE FOLLOWING INFORMATION PERTAINING TO AN ABBOTT DIABETES CARE (ADC) DEVICE: A HEALTHCARE PROFESSIONAL REPORTED ON (B)(6) 2025 THAT A CUSTOMER PASSED AWAY WHILE WEARING AN ADC DEVICE; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. NO CONTACT INFORMATION WAS PROVIDED TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. THE HCP MADE NO ALLEGATION OF DEFICIENCY WITH THE ADC DEVICE; THEREFORE, BASED ON THE LIMITED INFORMATION AVAILABLE, THERE IS NO EVIDENCE TO INDICATE THAT THE ABBOTT DIABETES CARE (ADC) DEVICE CONTRIBUTED TO THE REPORTED INCIDENT. Manufacturer Narrative: THIS COMPLAINT WAS RECEIVED VIA USER REPORT # MW5167182. ABBOTT DIABETES CARE (ADC) IS UNABLE TO FURTHER INVESTIGATE THE REPORTED COMPLAINT DUE TO THE ABSENCE OF AN ALLEGED PRODUCT ISSUE. ADDITIONALLY, NO CONTACT INFORMATION WAS PROVIDED IN THE USER REPORT, RENDERING ADC UNABLE TO CONDUCT ANY FOLLOW-UP ACTIVITIES. THE DEVICE MANUFACTURING DATE IS UNKNOWN, AND THE DATE ENTERED IN SECTION H4 REFLECTS THE DATE ADC BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ADC HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21780850&pc=MNS | 2518422-2025-041391 | 2021-09-24 04:00 | Death | RESPIRONICS, INC. | 2025-04-04 04:00 | MNS | DREAMSTATION ST30 | Degraded; Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: "THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972,Z-1973, AND Z-1974." Manufacturer Narrative: AS PER ADDITIONAL INFORMATION RECEIVED ON 01/12/2022: THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND NO FAULT FOUND. ADDITIONALLY, ALL TESTS PASSED. IN THIS REPORT, BOX H HAS BEEN UPDATED/CORRECTED. Manufacturer Narrative: CORRECTION TO THE PREVIOUSLY REPORTED INFORMATION: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND NO FAULT FOUND. ADDITIONALLY, ALL TESTS PASSED. THE "PROBLEM CODE GRID" HAS BEEN CORRECTED IN THIS REPORT. IN THIS REPORT, BOX H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21759344&pc=OZP | 2032227-2025-147379 | 2025-02-22 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-03 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Insufficient Flow or Under Infusion; Adverse Event Without Identified Device or Use Problem | Pneumonia; Coma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN BUT THE CUSTOMER EXPERIENCED HYPERGLYCEMIA WITH THE BLOOD GLUCOSE VALUES OF 500 MG/DL AND DIABETIC COMA AT THE TIME OF EMERGENCY ROOM VISIT. THE ILLNESS THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS ONLY PNEUMONIA. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-243A, MMT-332A, MMT-1780KL. TROUBLESHOOTING WAS PERFORMED FOR THE REPORTED EVENT WITH THE CUSTOMER'S NEXT OF KIN AND IT WAS FOUND THAT THE PUMP WAS NOT WORN BY THE CUSTOMER AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-243A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. MMT-1780KL WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21759540&pc=GDW | 2955842-2025-13048 | 2025-02-14 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-03 04:00 | GDW | SUREFORM | Failure to Fire | Sepsis; Brain Injury; Respiratory Failure | K173721 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE PATIENT WAS RE-HOSPITALIZED ON POST-OPERATIVE DAY (POD) 11 AND ULTIMATELY EXPIRED ON POD 18. THE PATIENT WAS READMITTED ON POD 11 WITH HEMATEMESIS AND SEPSIS WITHOUT SHOCK. THE PATIENT HAD A BMI OF 82 AND WAS INTUBATED FOR AIRWAY PROTECTION. A CT SCAN SHOWED A TINY AMOUNT OF CONTRAST LEAKAGE, BUT IT COULD NOT BE DETERMINED IF THE STAPLE LINE WAS OPEN OR NOT. SOMETIME AFTER THE CT SCAN, THE PATIENT CODED FROM RESPIRATORY FAILURE AND WAS RESUSCITATED. EXPLORATORY SURGERY FOUND A 7CM OPENING AT THE STAPLE LINES. THE SURGEON THOUGHT THAT THE STAPLE LINE WAS NOT OPEN BEFORE THE PATIENT CODED; IT WAS THOUGHT TO BE CAUSED BY, OR BECAME WIDER, DURING THE RESUSCITATIVE EFFORTS. THE SURGEON STATED THAT IF THE PATIENT HAD NOT CODED DUE TO RESPIRATORY FAILURE, THE FINDINGS WOULD HAVE BEEN TREATABLE. THE PATIENT SUFFERED RESPIRATORY FAILURE, CARDIOGENIC SHOCK, RENAL FAILURE, AND AN ANOXIC BRAIN INJURY. THE ULTIMATE CAUSE OF DEATH WAS REPORTED TO BE THE ANOXIC BRAIN INJURY THAT WAS IRREVERSIBLE. DURING THE PROCEDURE, A SUREFORM 60 STAPLER FIRED FOUR BLACK RELOADS WITH BUTTRESS MATERIAL WITHOUT INCIDENT. DURING THE FIFTH BLACK RELOAD FIRE A ¿TISSUE TOO THICK TO CONTINUE¿ MESSAGE APPEARED. THE SURGEON DID NOT THINK THE TISSUE WAS TOO THICK AS THE OTHER FIRINGS WERE UNEVENTFUL. A SIXTH SUREFORM 60 BLACK RELOAD WITHOUT BUTTRESS MATERIAL FIRED SUCCESSFULLY TO FINISH THE STAPLE LINE. AT THE END OF THE PROCEDURE, THE SURGEON INSPECTED THE STOMACH STAPLE LINES VIA ENDOSCOPY AND NO ISSUES WERE FOUND. THE SURGEON STATED THAT THE LEAK WAS LIKELY FROM THE STAPLE LINE THAT HAD THE MISFIRE, BUT THE STAPLE LINES HAD APPEARED IN GOOD CONDITION WHEN INSPECTED DURING THE PROCEDURE; THE LEAK AND THE DEATH ¿ARE TWO SEPARATE ISSUES¿. THE SURGEON ALSO STATED THAT ¿THE STAPLER MISFIRE AND LEAK WERE THE ONLY OBVIOUS ABNORMALITIES AND IT WAS LIKELY RELATED SOMEHOW TO THE PATIENT OUTCOME.¿ Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE 30 DEGREE PLUS, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE ADVANCED STAPLER LOGS SHOWED THAT THE SUREFORM 60 STAPLER WAS USED TO FIRE SIX BLACK RELOADS WITH NO INCOMPLETE CLAMPS OR UNCLAMPS. THE FIFTH BLACK RELOAD FIRE STOPPED AT ABOUT 82%. THE SAME STAPLER INSTRUMENT WAS INSTALLED TO FIRE A SIXTH BLACK RELOAD SUCCESSFULLY. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BARIATRIC SLEEVE GASTRECTOMY PROCEDURE FOR EXTREME MORBID OBESITY AND BMI OF 82. THE INTRAOPERATIVE FINDINGS AT THE TIME OF THE INITIAL PROCEDURE SHOWED NO EVIDENCE OF LEAK. THE PATIENT RETURNED 11 DAYS AFTER SURGERY WITH A SMALL LEAK NOTED ON CT. THE CAUSE AND EXACT LOCATION OF THIS LEAK WAS NOT PROVIDED. THE PATIENT BEGAN HAVING ADDITIONAL COMPLICATIONS INCLUDING RESPIRATORY FAILURE, RENAL FAILURE AND CARDIOGENIC SHOCK. THE PATIENT EVENTUALLY CODED. ON RE-EXPLORATION, THE STAPLE LINE HAD NOW DEHISCED WHICH WAS PRESENT ON THE CONTRAST CT SCAN PREVIOUSLY DONE. THE CAUSE OF THIS DISRUPTION IS UNKNOWN HOWEVER THE SURGEON SPECULATED THIS DISRUPTION MAY HAVE OCCURRED AS A RESULT OF COMPRESSIONS WHEN THE PATIENT CODED. THE PATIENT EVENTUALLY DEVELOPED AN ANOXIC BRAIN INJURY AND DIED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION A4 PATIENT WEIGHT: THE PATIENT BMI OF 82 KG/M2 WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21759834&pc=DSP | 2249723-2025-0001575 | 2025-03-23 04:00 | Death | DATASCOPE CORP. | 2025-04-03 04:00 | DSP | CARDIOSAVE HYBRID W/ E/F PLUG | Electrical Power Problem; Power Problem | Heart Failure/Congestive Heart Failure; Insufficient Information | K112372 |  | 1 | Event Description: ON (B)(6) 2025, IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) DID NOT SWITCH TO MAINS OPERATION ONCE THE POWER CONNECTION WAS PROPERLY ESTABLISHED. THE DEVICE REMAINED ON BATTERY POWER AND THE CONTROL BLOCK WAS CORRECTLY SECURED IN THE TROLLEY. ON (B)(6) 02025 ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT LATER DIED DUE TO POOR GENERAL CONDITION, NOT BECAUSE THE THERAPY WAS NOT STARTED. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS - B2, B4, B7, G1, G3, G6, H2, H6(HEALTH EFFECT ¿ CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSIONS), H11. A GETINGE FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO INVESTIGATE THE ISSUE. THE FSE REPLACED THE DEFECTIVE POWER CORD (D012-00-1688-02). THE UNIT PASSED ALL SAFETY AND FUNCTIONAL TESTS AND WAS RETURNED TO THE CUSTOMER FOR CLINICAL USE. Manufacturer Narrative: CORRECTED FIELDS - D10, H6(MEDICAL DEVICE ¿ PROBLEM CODE).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760172&pc=LWS | 2649622-2025-08969 | 2025-03-12 04:00 | Death | MPRI | 2025-04-03 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PHYSICIAN COMMENTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED SHORT V-V INTERVALS. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND NOT RECEIVED AT THIS TIME.Event Description: IT WAS FURTHER REPORTED THE PATIENT CAUSE OF DEATH WAS END STAGE RENAL FAILURE AND HYPOTENSION. Manufacturer Narrative: CONTINUATION OF D10: 407652 LEAD, IMPLANTED: (B)(6) 2012. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760273&pc=OZP | 2032227-2025-147685 | 2025-02-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-03 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Memory Loss/Impairment | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNOMEDICAL, MMT-332A, MMT-1884. THE CUSTOMER HAD NOT WORN THE PUMP AT THE TIME OF PASSING AND IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. THE CUSTOMER'S SPOUSE REPORTED THAT THE CUSTOMER EXPERIENCED HIGH BLOOD GLUCOSE OF OVER 700 MG/DL. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE CAUSE OF PASSING WAS ALZHEIMER'S DECEASED AND DIABETES. THE INSULIN PUMP WILL NOT BE RETURNED FOR ANALYSIS. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760352&pc=MIH | 2017233-2025-06027 | 2025-03-15 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-03 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Paralysis | P040043 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE ON (B)(6) 2025 FROM THE IMEDIDATA STUDY DATABASE: A PATIENT WITH TYPE B (DISTAL TO LEFT SUBCLAVIAN ARTERY) THORACIC / THORACOABDOMINAL DISSECTION UNDERWENT EMERGENCY TREATMENT ON (B)(6), 2025 USING GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE) AND GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG). THE AORTIC COMPONENT OF THE TBE DEVICE (TAC084515E) WAS IMPLANTED IN ZONE 1, WITH THE SIDE BRANCH (TSB081006) IN THE LEFT COMMON CAROTID ARTERY (LCCA), AND A SECOND SIDE BRANCH (TSB121706) IN THE LEFT SUBCLAVIAN ARTERY (LSA) USING SNORKELING TECHNIQUE WITH SELF EXPANDABLE STENT. THE 2 CTAG DEVICES ARE IMPLANTED IN THE PROXIMAL DESCENDING THORACIC AORTA (TGMR373710), AND THE DISTAL DESCENDING THORACIC AORTA (TGMR312610). THERE WERE NO TYPE I OR TYPE III ENDOLEAK OBSERVED AT THE CONCLUSION OF THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THAT A 24 F GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS USED TO INTRODUCE THE TBE DEVICE DURING THE INITIAL PROCEDURE ((B)(6) 2025). DUE TO THE LARGE SHEATH SIZE AND THE SMALL ACCESS POINTS OF THE PATIENT, THE PERCUTANEOUS ACCESS AT THE RIGHT FEMORAL SITE FAILED AND LEAD TO A DISSECTION OF THE RIGHT FEMORAL AND EXTERNAL ILIAC ARTERY. A RIGHT ILIAC ARTERY STENTING AND A SURGICAL CUT DOWN WAS PERFORMED. ON THE SAME DAY, THE PATIENT SUFFERED A MAJOR HEMISPHERIC STROKE WITH RIGHT-SIDE HEMIPLEGIA, SYSTEMIC THROMBOLYSIS WAS BEGUN. WHILE ON SYSTEMIC THROMBOLYSIS, FEMORAL AND RETROPERITONEAL BLEEDING OCCURRED, LEADING TO ANOTHER SURGICAL CUTDOWN ON THE RIGHT FEMORAL ARTERY. IT IS SUSPECTED THAT THE THROMBOLYSIS TRIGGERED A RETROPERITONEAL HEMATOMA (LIKELY DUE TO THE PREVIOUS VASCULAR ACCESS INJURY). THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PRIMARY RELATIONSHIP IS REPORTEDLY RELATED TO TREATMENT/PROCEDURE. Manufacturer Narrative: THE AVAILABLE INFORMATION REPORTED IN THE COMPLAINT DOES NOT REASONABLY SUGGEST A POTENTIAL MALFUNCTION OR PRODUCT PACKAGING AND/OR LABELING ISSUE HAS OCCURRED. BASED ON THE INCIDENT DESCRIPTION AND THE SUBSEQUENT INVESTIGATION, NO FURTHER INFORMATION WAS PROVIDED TO GORE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THIS INCIDENT AND ASSIGN A ROOT CAUSE. PER THE CONFORMABLE GORE® TAG® THORACIC ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE® TAG® THORACIC ENDOPROSTHESIS MAY INCLUDE BUT ARE NOT LIMITED TO: NEUROLOGIC DAMAGE, LOCAL OR SYSTEMIC (E.G., STROKE, PARAPLEGIA, PARAPARESIS). UPDATED H6: ADDED HEALTH EFFECT - IMPACT CODE F1901. UPDATED COMPONENT CODE TO G04122, CODE G07001 WAS INADVERTENTLY ENTERED AND SHOULD BE REMOVED. UPDATED INVESTIGATION FINDINGS CODE TO C19, CODE C21 IS NO LONGER APPLICABLE. UPDATED INVESTIGATION CONCLUSIONS CODE TO D15 AND D12, CODE D16 IS NO LONGER APPLICABLE. Manufacturer Narrative: D10: CONCOMITANT MEDICAL PRODUCTS: (RELEVANT ASSOCIATED DEVICES USED WITH THE DEVICE BEING REPORTED ON): GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR312610), GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TAC084515E, TSB081006, TSB121706E). GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF2433). AS PART OF OUR ROUTINE QUALITY PROCEDURES, EACH BATCH OF DEVICES UNDERGOES COMPREHENSIVE QUALITY CONTROL TESTING AND INSPECTIONS PRIOR TO RELEASE FOR DISTRIBUTION. GORE REVIEWED THE MANUFACTURING RECORDS ASSOCIATED WITH THE REPORTED LOT NUMBER AND VERIFIED THAT ALL RELEASE CRITERIA HAD BEEN MET. ENGINEERING EVALUATION COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED. THE PATIENT REFERENCED IN THIS EVENT FILE IS ENROLLED IN THE TGR23-02TB TOGETHER AORTIC REGISTRY AND INFORMATION REGARDING THIS EVENT WAS GATHERED FROM THE IMEDIDATA RAVE CLINICAL STUDY DATABASE (CSD). W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760355&pc=CBK | 2518422-2025-041095 | 2021-12-16 05:00 | Death | RESPIRONICS, INC. | 2025-04-03 04:00 | CBK | TRILOGY 100 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT WHO HAD BEEN USING A TRILOGY 100 VENTILATOR PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION AGAINST THE DEVICE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE EVALUATION CONFIRMED THAT THE DEVICE OPERATED AND ALARMED AS DESIGNED. THE DEVICE WAS CLEANED AND PASSED ALL FINAL TESTING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760363&pc=MIH | 2017233-2025-06028 | 2025-03-15 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-03 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Paralysis | P210032 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE ON (B)(6) 2025 FROM THE IMEDIDATA STUDY DATABASE: A PATIENT WITH TYPE B (DISTAL TO LEFT SUBCLAVIAN ARTERY) THORACIC / THORACOABDOMINAL DISSECTION UNDERWENT EMERGENCY TREATMENT ON (B)(6) 2025 USING GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE) AND GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG). THE AORTIC COMPONENT OF THE TBE DEVICE (TAC084515E) WAS IMPLANTED IN ZONE 1, WITH THE SIDE BRANCH (TSB081006) IN THE LEFT COMMON CAROTID ARTERY (LCCA), AND A SECOND SIDE BRANCH (TSB121706) IN THE LEFT SUBCLAVIAN ARTERY (LSA) USING SNORKELING TECHNIQUE WITH SELF EXPANDABLE STENT. THE 2 CTAG DEVICES ARE IMPLANTED IN THE PROXIMAL DESCENDING THORACIC AORTA (TGMR373710), AND THE DISTAL DESCENDING THORACIC AORTA (TGMR312610). THERE WERE NO TYPE I OR TYPE III ENDOLEAK OBSERVED AT THE CONCLUSION OF THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THAT A 24 F GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS USED TO INTRODUCE THE TBE DEVICE DURING THE INITIAL PROCEDURE ((B)(6)2025). DUE TO THE LARGE SHEATH SIZE AND THE SMALL ACCESS POINTS OF THE PATIENT, THE PERCUTANEOUS ACCESS AT THE RIGHT FEMORAL SITE FAILED AND LEAD TO A DISSECTION OF THE RIGHT FEMORAL AND EXTERNAL ILIAC ARTERY. A RIGHT ILIAC ARTERY STENTING AND A SURGICAL CUT DOWN WAS PERFORMED. ON THE SAME DAY, THE PATIENT SUFFERED A MAJOR HEMISPHERIC STROKE WITH RIGHT-SIDE HEMIPLEGIA, SYSTEMIC THROMBOLYSIS WAS BEGUN. WHILE ON SYSTEMIC THROMBOLYSIS, FEMORAL AND RETROPERITONEAL BLEEDING OCCURRED, LEADING TO ANOTHER SURGICAL CUTDOWN ON THE RIGHT FEMORAL ARTERY. IT IS SUSPECTED THAT THE THROMBOLYSIS TRIGGERED A RETROPERITONEAL HEMATOMA (LIKELY DUE TO THE PREVIOUS VASCULAR ACCESS INJURY). THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PRIMARY RELATIONSHIP IS REPORTEDLY RELATED TO TREATMENT/PROCEDURE. Manufacturer Narrative: THE AVAILABLE INFORMATION REPORTED IN THE COMPLAINT DOES NOT REASONABLY SUGGEST A POTENTIAL MALFUNCTION OR PRODUCT PACKAGING AND/OR LABELING ISSUE HAS OCCURRED. BASED ON THE INCIDENT DESCRIPTION AND THE SUBSEQUENT INVESTIGATION, NO FURTHER INFORMATION WAS PROVIDED TO GORE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THIS INCIDENT AND ASSIGN A ROOT CAUSE. PER THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INCLUDE, BUT ARE NOT LIMITED TO: NEUROLOGIC DAMAGE, LOCAL OR SYSTEMIC (E.G., STROKE, PARAPLEGIA, PARAPARESIS). UPDATED H6: ADDED HEALTH EFFECT - IMPACT CODE F1901. UPDATED COMPONENT CODE TO G04122, CODE G07001 WAS INADVERTENTLY ENTERED AND SHOULD BE REMOVED. UPDATED INVESTIGATION FINDINGS CODE TO C19, CODE C21 IS NO LONGER APPLICABLE. UPDATED INVESTIGATION CONCLUSIONS CODE TO D15 AND D12, CODE D16 IS NO LONGER APPLICABLE. Manufacturer Narrative: C1: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. Manufacturer Narrative: D10: CONCOMITANT MEDICAL PRODUCTS: (RELEVANT ASSOCIATED DEVICES USED WITH THE DEVICE BEING REPORTED ON): GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (TGMR373710E, TGMR312610), GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TAC084515E, TSB081006). GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF2433). AS PART OF OUR ROUTINE QUALITY PROCEDURES, EACH BATCH OF DEVICES UNDERGOES COMPREHENSIVE QUALITY CONTROL TESTING AND INSPECTIONS PRIOR TO RELEASE FOR DISTRIBUTION. GORE REVIEWED THE MANUFACTURING RECORDS ASSOCIATED WITH THE REPORTED LOT NUMBER AND VERIFIED THAT ALL RELEASE CRITERIA HAD BEEN MET. ENGINEERING EVALUATION COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED. THE PATIENT REFERENCED IN THIS EVENT FILE IS ENROLLED IN THE TGR23-02TB TOGETHER AORTIC REGISTRY AND INFORMATION REGARDING THIS EVENT WAS GATHERED FROM THE IMEDIDATA RAVE CLINICAL STUDY DATABASE (CSD). W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760702&pc=MIH | 9612164-2025-01701 | 2025-03-19 04:00 | Death | MEDTRONIC IRELAND | 2025-04-03 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P100021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN INDEX PROCEDURE INVOLVING THE HELI-FX ENDOSYSTEM AND AN ENDURANTIIS, A RUPTURE WAS SUSPECTED WHEN THE NON-MEDTRONIC WIRE APPEARED TO DEFLECT OUTSIDE THE AORTIC LUMEN UPON REMOVAL OF THE HELI-FX GUIDE. PRE-OPERATIVELY THE DIAMETER OF THE AORTA AT THE RENAL ARTERY MEASURED 26MM WHILE THE LENGTH OF THE PROXIMAL AORTIC NECK WAS 9MM. MODERATE VESSEL TORTUOSITY AND SEVERE VESSEL CALCIFICATION WERE REPORTED. THE PROCEDURE INITIALLY INVOLVED THE SUCCESSFUL PLACEMENT OF A 32MM BIFURCATED DEVICE AND 16X16MM DIAMETER LIMBS, ALONG WITH SIX ENDOANCHORS DUE TO THE PATIENT'S SHORT AND CONICAL NECK ANATOMY. A SELECTIVE ANGIOGRAM WAS PERFORMED, AND NO RUPTURE LEAKS WERE IDENTIFIED AT THAT TIME. HOWEVER, APPROXIMATELY ONE HOUR LATER, THE PATIENT WAS REPORTED TO BE UNSTABLE, PROMPTING ANOTHER ANGIOGRAM WHICH REVEALED A TEAR ABOVE THE LEFT RENAL ARTERY. THE LEFT RENAL ARTERY WAS EMBOLIZED VIA BRACHIAL APPROACH AND COVERED WITH A 32X32X49MM AORTIC CUFF, WHILE THE RIGHT RENAL ARTERY WAS STENTED AND "SNORKELED" BEHIND A 36X36X49MM AORTIC CUFF, MAINTAINING BLOOD FLOW THROUGH THE STENT. THE PATIENT RECEIVED THREE UNITS OF BLOOD AND WAS TRANSFERRED TO CRITICAL CARE. MULTIPLE STROKES WERE CONFIRMED AND THE PATIENT WAS NOT MOVING THE RIGHT SIDE OF THEIR BODY FOR SEVERAL DAYS WHILE INTUBATED. THE PHYSICIAN CONFIRMED THAT THE RUPTURE AND STROKE WERE NOT ATTRIBUTED TO THE HELI-FX GUIDE. IT IS BELIEVED THAT IT WAS THE NON-MEDTRONIC WIRE THAT CAUSED THE RUPTURE AND IT WAS STATED THAT ¿AS THE WIRE WAS IN FRONT OF THE HELI-FX GUIDE AND THE GUIDE WOULD ONLY FOLLOW THE WIRE". THE STROKES WERE BELIEVED TO BE CAUSED FROM THE BRACHIAL ACCESS AND MULTIPLE EXCHANGES OVER THE ARCH DURING THE SECONDARY PROCEDURE TO FIX THE RUPTURE. SEVERAL EMBOLI WERE REPORTED TO AFFECT THE BRAIN STEM. THE PATIENT¿S FAMILY DECIDED TO COMPASSIONATELY REMOVE INTUBATION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: ENDURANT II ILIAC STENT GRAFT; PRODUCT ID: ETLW1616C124E, SERIAL#: (B)(6) BRAND NAME: ENDURANT II ILIAC STENT GRAFT; PRODUCT ID: ETLW1616C93E, SERIAL#: (B)(6) B2: THE EXACT DATE OF DEATH IS UNKNOWN. MONTH AND YEAR ARE VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760739&pc=MIH | 9612164-2025-01702 | 2025-03-19 04:00 | Death | MEDTRONIC IRELAND | 2025-04-03 04:00 | MIH | ENDURANT II EXTENSION CUFF | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P100021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN INDEX PROCEDURE INVOLVING THE HELI-FX ENDOSYSTEM AND A N ENDURANT IIS, A RUPTURE WAS SUSPECTED WHEN THE NON-MEDTRONIC WIRE APPEARED TO DEFLECT OUTSIDE THE AORTIC LUMEN UPON REMOVAL OF THE HELI-FX GUIDE. PRE-OPERATIVELY THE DIAMETER OF THE AORTA AT THE RENAL ARTERY MEASURED 26MM WHILE THE LENGTH OF THE PROXIMAL AORTIC NECK WAS 9MM. MODERATE VESSEL TORTUOSITY AND SEVERE VESSEL CALCIFICATION WERE REPORTED. THE PROCEDURE INITIALLY INVOLVED THE SUCCESSFUL PLACEMENT OF A 32MM BIFURCATED DEVICE AND 16X16MM DIAMETER LIMBS, ALONG WITH SIX ENDOANCHORS DUE TO THE PATIENT'S SHORT AND CONICAL NECK ANATOMY. A SELECTIVE ANGIOGRAM WAS PERFORMED, AND NO RUPTURE LEAKS WERE IDENTIFIED AT THAT TIME. HOWEVER, APPROXIMATELY ONE HOUR LATER, THE PATIENT WAS REPORTED TO BE UNSTABLE, PROMPTING ANOTHER ANGIOGRAM WHICH REVEALED A TEAR ABOVE THE LEFT RENAL ARTERY. THE LEFT RENAL ARTERY WAS EMBOLIZED VIA BRACHIAL APPROACH AND COVERED WITH A 32X32X49MM AORTIC CUFF, WHILE THE RIGHT RENAL ARTERY WAS STENTED AND "SNORKELED" BEHIND A 36X36X49MM AORTIC CUFF, MAINTAINING BLOOD FLOW THROUGH THE STENT. THE PATIENT RECEIVED THREE UNITS OF BLOOD AND WAS TRANSFERRED TO CRITICAL CARE. MULTIPLE STROKES WERE CONFIRMED AND THE PATIENT WAS NOT MOVING THE RIGHT SIDE OF THEIR BODY FOR SEVERAL DAYS WHILE INTUBATED. THE PHYSICIAN CONFIRMED THAT THE RUPTURE AND STROKE WERE NOT ATTRIBUTED TO THE HELI-FX GUIDE. IT IS BELIEVED THAT IT WAS THE NON-MEDTRONIC WIRE THAT CAUSED THE RUPTURE AND IT WAS STATED THAT ¿AS THE WIRE WAS IN FRONT OF THE HELI-FX GUIDE AND THE GUIDE WOULD ONLY FOLLOW THE WIRE". THE STROKES WERE BELIEVED TO BE CAUSED FROM THE BRACHIAL ACCESS AND MULTIPLE EXCHANGES OVER THE ARCH DURING THE SECONDARY PROCEDURE TO FIX THE RUPTURE. SEVERAL EMBOLI WERE REPORTED TO AFFECT THE BRAIN STEM. THE PATIENT¿S FAMILY DECIDED TO COMPASSIONATELY REMOVE INTUBATION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: ENDURANT II EXTENSION CUFF; PRODUCT ID: ETCF3636C49E, SERIAL#: (B)(6) B2; THE EXACT DATE OF DEATH IS UNKNOWN. MONTH AND YEAR ARE VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21760743&pc=OTD | 9612164-2025-01703 | 2025-03-19 04:00 | Death | MEDTRONIC IRELAND | 2025-04-03 04:00 | OTD | ENDO ANCHOR SYSTEM - HELI-FX AAA | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | K140036 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN INDEX PROCEDURE INVOLVING THE HELI-FX ENDOSYSTEM AND A N ENDURANT IIS, A RUPTURE WAS SUSPECTED WHEN THE NON-MEDTRONIC WIRE APPEARED TO DEFLECT OUTSIDE THE AORTIC LUMEN UPON REMOVAL OF THE HELI-FX GUIDE. PRE-OPERATIVELY THE DIAMETER OF THE AORTA AT THE RENAL ARTERY MEASURED 26MM WHILE THE LENGTH OF THE PROXIMAL AORTIC NECK WAS 9MM. MODERATE VESSEL TORTUOSITY AND SEVERE VESSEL CALCIFICATION WERE REPORTED. THE PROCEDURE INITIALLY INVOLVED THE SUCCESSFUL PLACEMENT OF A 32MM BIFURCATED DEVICE AND 16X16MM DIAMETER LIMBS, ALONG WITH SIX ENDOANCHORS DUE TO THE PATIENT'S SHORT AND CONICAL NECK ANATOMY. A SELECTIVE ANGIOGRAM WAS PERFORMED, AND NO RUPTURE LEAKS WERE IDENTIFIED AT THAT TIME. HOWEVER, APPROXIMATELY ONE HOUR LATER, THE PATIENT WAS REPORTED TO BE UNSTABLE, PROMPTING ANOTHER ANGIOGRAM WHICH REVEALED A TEAR ABOVE THE LEFT RENAL ARTERY. THE LEFT RENAL ARTERY WAS EMBOLIZED VIA BRACHIAL APPROACH AND COVERED WITH A 32X32X49MM AORTIC CUFF, WHILE THE RIGHT RENAL ARTERY WAS STENTED AND "SNORKELED" BEHIND A 36X36X49MM AORTIC CUFF, MAINTAINING BLOOD FLOW THROUGH THE STENT. THE PATIENT RECEIVED THREE UNITS OF BLOOD AND WAS TRANSFERRED TO CRITICAL CARE. MULTIPLE STROKES WERE CONFIRMED AND THE PATIENT WAS NOT MOVING THE RIGHT SIDE OF THEIR BODY FOR SEVERAL DAYS WHILE INTUBATED. THE PHYSICIAN CONFIRMED THAT THE RUPTURE AND STROKE WERE NOT ATTRIBUTED TO THE HELI-FX GUIDE. IT IS BELIEVED THAT IT WAS THE NON-MEDTRONIC WIRE THAT CAUSED THE RUPTURE AND IT WAS STATED THAT ¿AS THE WIRE WAS IN FRONT OF THE HELI-FX GUIDE AND THE GUIDE WOULD ONLY FOLLOW THE WIRE". THE STROKES WERE BELIEVED TO BE CAUSED FROM THE BRACHIAL ACCESS AND MULTIPLE EXCHANGES OVER THE ARCH DURING THE SECONDARY PROCEDURE TO FIX THE RUPTURE. SEVERAL EMBOLI WERE REPORTED TO AFFECT THE BRAIN STEM. THE PATIENT¿S FAMILY DECIDED TO COMPASSIONATELY REMOVE INTUBATION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: ENDO ANCHOR SYSTEM - HELI-FX AAA; PRODUCT ID: SG-64, SERIAL#: (B)(6). B2: THE EXACT DATE OF DEATH IS UNKNOWN. MONTH AND YEAR ARE VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761312&pc=PIJ | 1213643-2025-00295 | 2022-01-01 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-04-03 04:00 | PIJ | XENMATRIX AB | Patient-Device Incompatibility | Abdominal Pain; Adhesion(s); Hernia; Obstruction/Occlusion; Fluid Discharge | K133223 |  | 1 | Event Description: PER MEDWATCH REPORT (MW5167617), "ALLEGED EVENT: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH 1ST STRATTICE ON (B)(6) 2016 AND 2ND STRATTICE ON (B)(6) 2017. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN REMOVED OR REVISED. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT HAS DIED ON (B)(6) 2022 AND THERE IS NO FURTHER DETAILS ON DEATH CAUSALITY. THIS RECORD CAPTURED 1ST STRATTICE. ON (B)(6) 2025, (B)(6) RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHO UNDERWENT A VENTRAL HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT SP100381-101 ON (B)(6) 2016. THE PATIENT UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT SP100537-078 ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT SP100595." THE PATIENT THEN UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH ON (B)(6) 2018 AND IMPLANTED WITH A NONABBVIE DEVICE, XENMATRIX LOT HUCR0748. THE MESH WAS REVISED DUE TO RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, BARD MESH LOT #HUCQ0919. THE MESH WAS REVISED ON (B)(6) 2022 DUE TO OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH ON OR ABOUT (B)(6) 2016. THE PATIENT ALSO UNDERWENT A OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) ON OR ABOUT(B)(6) 2017. THE PATIENT WAS TREATED FOR INCISIONAL/ VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) BETWEEN (B)(6) 2018. THE PATIENT UNDERWENT OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN BETWEEN(B)(6) 2018 AND (B)(6) 2018. THE PATIENT WAS TREATED FOR RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON OR ABOUT (B)(6) 2018. THE PATIENT WAS SEEN FOR ABDOMINAL PAIN BETWEEN 2018 AND (B)(6) 2022. THE PATIENT WAS TREATED FOR BOWEL OBSTRUCTION BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ABDOMINAL PAIN/ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR HERNIA SYMPTOMS FROM APPROXIMATELY 2018 TO (B)(6) 2022. THE PATIENT WAS TREATED OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF REHERNIATION BETWEEN (B)(6) 2022 AND (B)(6) 2022. THE PATIENT RECEIVED POST-OP REHAB BETWEEN 2018-2019. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)". REFERENCE REPORT: MW5167618. LOT #: HUCR0748". Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. AS REPORTED, PATIENT HAD A COMPLEX MEDICAL HISTORY WITH MULTIPLE IMPLANTS AND SURGICAL PROCEDURES. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. REVIEW OF MANUFACTURING RECORDS FOR THE XENMATRIX AB IMPLANTS USED TO TREAT THE PATIENT CONFIRMED THE PRODUCTS WERE MANUFACTURED TO SPECIFICATION. TO DATE, THIS IS THE ONLY REPORTED COMPLAINT FOR THIS MANUFACTURING LOT OF 38 UNITS. NOTE, THE ACTUAL DATE OF DEATH AND THE DATE OF EVENT WERE NOT REPORTED IN THE REPORT SENT TO BD. WE HAVE DOCUMENTED THESE REQUIRED DATE FIELDS WITH A BEST ESTIMATE. THIS MDR REPRESENTS THE XENMATRIX AB (DEVICE #2). AN ADDITIONAL MDR WAS SUBMITTED TO REPRESENT THE XENMATRIX AB (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761314&pc=PIJ | 1213643-2025-00294 | 2022-01-01 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-04-03 04:00 | PIJ | XENMATRIX AB | Patient-Device Incompatibility | Abdominal Pain; Adhesion(s); Hernia; Obstruction/Occlusion; Fluid Discharge | K133223 |  | 1 | Event Description: PER MEDWATCH REPORT (MW5167617), "ALLEGED EVENT: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH 1ST STRATTICE ON (B)(6) 2016 AND 2ND STRATTICE ON (B)(6) 2017. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN REMOVED OR REVISED. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT HAS DIED ON (B)(6) 2022 AND THERAILS OE IS NO FURTHER DETN DEATH CAUSALITY. THIS RECORD CAPTURED 1ST STRATTICE. ON (B)(6) 2025, (B)(6) RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHO UNDERWENT A VENTRAL HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT SP100381-101 ON (B)(6) 2016. THE PATIENT UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT SP100537-078 ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT SP100595." THE PATIENT THEN UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH ON (B)(6) 2018 AND IMPLANTED WITH A NONABBVIE DEVICE, XENMATRIX LOT HUCR0748. THE MESH WAS REVISED DUE TO RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, BARD MESH LOT #HUCQ0919. THE MESH WAS REVISED ON (B)(6) 2022 DUE TO OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH ON OR ABOUT (B)(6) 2016. THE PATIENT ALSO UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) ON OR ABOUT(B)(6) 2017. THE PATIENT WAS TREATED FOR INCISIONAL/ VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) BETWEEN (B)(6) 2018. THE PATIENT UNDERWENT OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN BETWEEN (B)(6) 2018 AND (B)(6) 2018. THE PATIENT WAS TREATED FOR RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON OR ABOUT (B)(6) 2018. THE PATIENT WAS SEEN FOR ABDOMINAL PAIN BETWEEN 2018 AND (B)(6) 2022. THE PATIENT WAS TREATED FOR BOWEL OBSTRUCTION BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ABDOMINAL PAIN/ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR HERNIA SYMPTOMS FROM APPROXIMATELY 2018 TO (B)(6) 2022. THE PATIENT WAS TREATED OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF REHERNIATION BETWEEN (B)(6) 2022 AND (B)(6) 2022. THE PATIENT RECEIVED POST-OP REHAB BETWEEN 2018-2019. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)". REFERENCE REPORT: MW5167618. LOT #: HUCR0748". Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. AS REPORTED, PATIENT HAD A COMPLEX MEDICAL HISTORY WITH MULTIPLE IMPLANTS AND SURGICAL PROCEDURES. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. REVIEW OF MANUFACTURING RECORDS FOR THE XENMATRIX AB IMPLANTS USED TO TREAT THE PATIENT CONFIRMED THE PRODUCTS WERE MANUFACTURED TO SPECIFICATION. TO DATE, THIS IS THE ONLY REPORTED COMPLAINT FOR THIS MANUFACTURING LOT OF 31 UNITS. NOTE, THE ACTUAL DATE OF DEATH AND THE DATE OF EVENT WERE NOT REPORTED IN THE REPORT SENT TO BD. WE HAVE DOCUMENTED THESE REQUIRED DATE FIELDS WITH A BEST ESTIMATE. THIS MDR REPRESENTS THE XENMATRIX AB (DEVICE #1). AN ADDITIONAL MDR WAS SUBMITTED TO REPRESENT THE XENMATRIX AB (DEVICE #2). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761431&pc=OZP | 2032227-2025-148009 | 2024-12-19 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-03 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Infusion or Flow Problem | Insufficient Information | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS PASSED AWAY ON (B)(6) 2024 AND HAS REPORTED FOR PUMP DELIVERY VOLUME ACCURACY DURING CHANGES IN AIR PRESSURE. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-397A, MMT-332A, MMT-1884. PRODUCT RETURN REQUIRED FOR MMT-397A. TROUBLESHOOTING WAS NOT PERFORMED. TROUBLESHOOTING WAS NOT PERFORMED AND IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP WITHIN 48 HOURS OF THE REPORTED EVENT AND UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. PRODUCT RETURN REQUIRED FOR MMT-332A. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. PRODUCT RETURN REQUIRED FOR MMT-1884. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED.Event Description: SUMMARY UPDATE: THE CASE WAS REPORTED IN ERROR. SAFETY NOTIFICATION WAS SENT OUT TO THE CUSTOMER BUT THERE WAS NO ALLEGATION THAT THEY EXPERIENCED PUMP DELIVERY VOLUME ACCURACY DURING CHANGES IN AIR PRESSURE. IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS PASSED AWAY ON (B)(6) 2024 AND HAS REPORTED FOR PUMP DELIVERY VOLUME ACCURACY DURING CHANGES IN AIR PRESSURE. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-397A, MMT-332A, MMT-1884. PRODUCT RETURN REQUIRED FOR MMT-397A. TROUBLESHOOTING WAS NOT PERFORMED. TROUBLESHOOTING WAS NOT PERFORMED AND IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP WITHIN 48 HOURS OF THE REPORTED EVENT AND UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. PRODUCT RETURN REQUIRED FOR MMT-332A. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. PRODUCT RETURN REQUIRED FOR MMT-1884. IT IS UNKNOWN WHETHER PRODUCT WILL BE RETURNED. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT AND/OR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761436&pc=OZD | 1220648-2025-27204 | 2025-03-08 05:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | IMPELLA 5.5 | Failure to Pump | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING A DIFFICULT IMPELLA 5.5 DEVICE IMPLANT, COMPLICATIONS WERE ENCOUNTERED WHEN ATTEMPTING TO START THE PUMP. THE STAFF NOTED THAT A SPIKE IN MOTOR CURRENT OCCURRED DURING EACH ATTEMPT FOLLOWED BY AN IMMEDIATE STOP. THE STAFF WAS ADVISED TO EXCHANGE THE PUMP; HOWEVER, THIS OPTION WAS NOT AVAILABLE. THE IMPELLA 5.5 DEVICE IMPLANT ULTIMATELY FAILED. THE PATIENT'S OUTCOME WAS SUBSEQUENTLY UPDATED TO EXPIRED. NO FURTHER DETAILS SURROUNDING THE DEMISE WAS PROVIDED. THE IMPACT OF THE PUMP'S FAILURE TO START TO THE EVENT IS UNCLEAR. Manufacturer Narrative: THE INVESTIGATION OF PUMP DID NOT START HAS BEEN COMPLETED. PUMP DID NOT START: UPON VISUAL INSPECTION, A SEVERE KINKING OF THE CATHETER PRESENT JUST BELOW MOTOR HOUSING. ELECTRICAL TESTING WAS PERFORMED AND OPEN LINES WERE PRESENT IN ALL THREE MOTOR CABLE LINES. BROKEN MOTOR CABLE LINES WERE PRESENT AT KINKING IN CATHETER BELOW MOTOR HOUSING. TWO MOTOR CABLE LINES WERE FOUND TO BE COMPLETELY BROKEN, AND THIRD MOTOR CABLE LINE SHOWED EXCESSIVE NECKING. DATA LOGS WERE REVIEWED AND REAL TIME (RT) LOG AT BEGINNING OF CASE SHOWS PUMP WAS ATTEMPTED TO BE STARTED MULTIPLE TIMES WITH IMMEDIATE MOTOR CURRENT SPIKES TO 30000 WITH "IMPELLA STOPPED (RPM DEVIATIONS)" ALARMS BEING TRIGGERED. CLINICAL DETAILS NOTED THAT THE CLINICAL TEAM HAD DIFFICULTY INSERTING THE PUMP DUE TO PATIENTS ANATOMY AND ATTEMPTED INSERTION MULTIPLE TIMES BEFORE INSERTION WAS SUCCESSFUL. THE ROOT CAUSE OF THE PUMP NOT STARTING WAS DUE TO AN ELECTRICAL OPEN OF THE MOTOR CABLE LINES AS A RESULT OF EXCESSIVE FORCE WHEN INSERTING THE PUMP BASED ON REVIEW OF RETURNED PRODUCT AND DATA LOG ANALYSIS. B2 LIFE-THREATENING WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27204. G1 MANUFACTURING SITE NAME, ADDRESS, COUNTRY AND FAX NUMBER WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27204. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF THE ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761468&pc=NPT | 2015691-2025-02541 | 2025-03-07 05:00 | Death | EDWARDS LIFESCIENCES | 2025-04-03 04:00 | NPT | SAPIEN 3 ULTRA RESILIA VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM THE UNITED KINGDOM, IT WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. THE PATIENT PRESENTED WITH A HEAVY CALCIFIED AORTIC VALVE. THE VALVE WAS DEPLOYED SUCCESSFULLY WITH NOMINAL VOLUME WITH ONLY SLIGHTLY HEMODYNAMIC INSTABILITY. POST IMPLANTATION, ARTERIAL PRESSURE DROPPED AND ECHO SHOWED PERICARDIAL EFFUSION. A PERICARDIOCENTESIS WAS PERFORMED DRAINING APPROXIMATELY 300ML OF BLOOD BUT NO TRANSFUSION WAS REQUIRED. THE SUSPECTED CAUSE OF THE EFFUSION WAS AN ANNULAR RUPTURE. AS PER MEDICAL OPINION, THE LEAK WAS SETTLED AND NO FURTHER ACTIONS WERE REQUIRED SO THE PATIENT WAS TRANSFERRED IN STABLE CONDITION TO THE CCU. HOWEVER, THE PATIENT PASSED AWAY LATER THAT DAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF THE REPORTED EVENT WAS UNABLE TO BE DETERMINED BUT MAY HAVE BEEN DUE TO PROCEDURAL AND OR PATIENT FACTORS OR THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761653&pc=DSQ | 2916596-2025-01960 | 2024-10-21 04:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Pneumonia; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE PATIENT DIED DUE TO ONGOING COMPLICATIONS SINCE THE IMPLANT DATE. THE PATIENT HAD BEEN ADMITTED IN THE HOSPITAL SINCE THE IMPLANTATION. THEY HAD PNEUMONIA, WAS ON A VENTILATOR, NO RECOVERY, AND WEAKNESS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. ADDITIONALLY, A SPECIFIC CAUSE FOR THE PATIENT'S PNEUMONIA COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AND LOCALIZED INFECTION AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761759&pc=OZD | 1220648-2025-27206 | 2025-03-10 04:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE 85-YEAR-OLD MALE PATIENT WAS ON IMPELLA CP SUPPORT FOR SUPPORT DURING A CARDIAC PROCEDURE. IT WAS NOTED THAT AFTER THE IMPLANT, THE PATIENT HAD A LARGE HEMATOMA AT THE ACCESS SITE. AT THE SAME TIME THE DEVICE WAS UNABLE TO BE REPOSITIONED OUT OF THE MITRAL APPARATUS AND WAS IMPEDING FLOWS WITH THE CALCIFICATION ON BOTH VALVES. THE PUMP WAS EXPLANTED. IT WAS LATER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21761906&pc=DSQ | 2916596-2025-02050 | 2025-03-01 05:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fever; Renal Failure; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE CLINICIAN NOTED THAT THE PATIENT PASSED AWAY WHILE STAYING AT THE NURSING HOSPITAL WITH DIALYSIS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED UNTIL THE END AND NO DEVICE SPECIFIC ISSUES WERE NOTIFIED. THE DEVICE WAS NOT EXPLANTED.Event Description: POSSIBLE PATIENT SYMPTOMS WERE FEVER AND BIOMARKER. ANTIBIOTICS AND NECESSARY DRUGS MIGHT HAVE APPLIED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, SEPSIS, AND RENAL DYSFUNCTION. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, SEVERAL SECTIONS OF THE HEARTMATE 3 LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762078&pc=NPT | 2025587-2025-02316 | 2024-04-21 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 MONTHS AND 11 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPSIS. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762117&pc=DSQ | 2916596-2025-02031 | 2024-12-12 05:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUFFERED FROM PROGRESSIVE MULTI-ORGAN FAILURE (MOF), AND PASSED AWAY DUE TO MOF. THE MOF WAS NOT CONSIDERED TO BE RELATED TO ANY DEVICE RELATED ISSUE. THE DEVICE WAS RUNNING AS EXPECTED. NO DEVICE ISSUES WERE MENTIONED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762272&pc=NPT | 2025587-2025-02317 | 2024-07-13 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AFTER HAVING BEEN TRANSFERRED TO ANOTHER HOSPITAL FOR REHABILITATION, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: CORRECTED B.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762431&pc=NPT | 2025587-2025-02318 | 2025-03-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A "FRAIL PATIENT" WITH A P RE-EXISTING NON-MEDTRONIC SURGICAL AORTIC VALVE REPLACEMENT, THE VALVE WAS LOADED PROPERLY AND A FRAME OVERLAP TO THE 2ND NODE WAS DETECTED. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED AND DEPLOYMENT WAS ATTEMPTED; HOWEVER, HYPOTENSION WAS NOTED AS THE VALVE WAS OPENED AT APPROXIMATELY 90%. THE VALVE WAS POSITIONED TOO LOW AND A RECAPTURE WAS PERFORMED. THE GUIDEWIRE WAS EXCHANGED TO A STIFFER NON-MEDTRONIC GUIDEWIRE. A SECOND DEPLOYMENT WAS ATTEMPTED AND THE VALVE WAS IN A GOOD POSITION, BUT THE BLOOD PRESSURE DROPPED AGAIN. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PHYSICIAN DECIDED TO PERFORM A RECAPTURE AND WITHDRAW THE VALVE AND DCS FROM THE PATIENT AS VALVE FUNCTION WAS UNCERTAIN. A NEW VALVE WAS PROPERLY LOADED WHILE CPR WAS ONGOING. THE NEW VALVE WAS IMPLANTED IN GOOD POSITION, AND THE BLOOD PRESSURE RECOVERED TEMPORARILY. AFTER SEVERAL MINUTES, THE BLOOD PRESSURE DROPPED AGAIN WITH AND CPR RESUMED. SEVERAL DEFIBRILLATIONS WERE PERFORMED, BUT WERE UNSUCCESSFUL. AFTER APPROXIMATELY 20 MINUTES OF CPR, THE TEAM DECIDED TO STOP ALL ACTIONS AND THE PATIENT SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS NOT REPORTED, BUT IT WAS NOTED THAT A PERICARDIAL BLEED WAS NOT DETECTED AND THE CORONARY ARTERIES WERE PATENT.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A "FRAIL PATIENT" WITH A PRE-EXISTING NON-MEDTRONIC SURGICAL AORTIC VALVE REPLACEMENT, THE VALVE WAS LOADED PROPERLY AND A FRAME OVERLAP TO THE 2ND NODE WAS DETECTED. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED AND DEPLOYMENT WAS ATTEMPTED; HOWEVER, HYPOTENSION WAS NOTED AS THE VALVE WAS OPENED AT APPROXIMATELY 90%. THE VALVE WAS POSITIONED TOO LOW AND A RECAPTURE WAS PERFORMED. THE GUIDEWIRE WAS EXCHANGED TO A STIFFER NON-MEDTRONIC GUIDEWIRE (SAFARI). A SECOND DEPLOYMENT WAS ATTEMPTED AND THE VALVE WAS IN A GOOD POSITION, BUT THE BLOOD PRESSURE DROPPED AGAIN. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PHYSICIAN DECIDED TO PERFORM A RECAPTURE AND WITHDRAW THE VALVE AND DCS FROM THE PATIENT AS VALVE FUNCTION WAS UNCERTAIN. A NEW VALVE WAS PROPERLY LOADED WHILE CPR WAS ONGOING. THE NEW VALVE WAS IMPLANTED IN GOOD POSITION, AND THE BLOOD PRESSURE RECOVERED TEMPORARILY. AFTER SEVERAL MINUTES, THE BLOOD PRESSURE DROPPED AGAIN WITH AND CPR RESUMED. SEVERAL DEFIBRILLATIONS WERE PERFORMED, BUT WERE UNSUCCESSFUL. AFTER APPROXIMATELY 20 MINUTES OF CPR, THE TEAM DECIDED TO STOP ALL ACTIONS AND THE PATIENT SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS NOT REPORTED, BUT IT WAS NOTED THAT A PERICARDIAL BLEED WAS NOT DETECTED AND THE CORONARY ARTERIES WERE PATENT. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN THAT THE CAUSE OF DEATH WAS DUE TO WEAK LEFT VENTRICLE, RIGHT VENTRICLE AND PULMONARY HYPERTENSION. ACCORDING TO THE PHYSICIAN, THE ATTEMPTED FIRST VALVE, THE IMPLANTED SECOND VALVE, AND THE DCS CAN ALL BE EXCLUDED AS CONTRIBUTING CAUSES TO THE DEATH; HOWEVER, THE PHYSICIAN ALSO NOTED THAT THE PATIENT "DIDN¿T TOLERATE THE DROP IN BLOOD PRESSURE DURING RECAPTURE AND DIDN¿T RECOVER." FURTHER, THE PHYSICIAN INDICATED THAT THE PATIENT'S UNDERLYING MEDICAL HISTORY (POOR LEFT VENTRICLE AND RIGHT VENTRICLE FUNCTION, PULMONARY HYPERTENSION, HIGH RISK STATUS) CONTRIBUTED TO THE DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (SECOND PARAGRAPH), D4, AND H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H2 H3 H6 IMAGE REVIEW: STATIC IMAGES AND MEDIA FILES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. IMAGES CONFIRMED THAT THE PATIENT HAD A PREVIOUSLY IMPLANTED SURGICAL AORTIC VALVE OF UNKNOWN TYPE AND SIZE. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. THERE IS EVIDENCE THAT AT SOME POINT DURING VALVE DEPLOYMENT ATTEMPT CARDIOPULMONARY RESUSCITATION WAS BEING PERFORMED. IT WAS REPORTED THAT THE FIRST VALVE WAS RECAPTURED, WITHDRAWN FROM THE PATIENT AND REPLACED WITH A NEW SYSTEM. IMAGES OF THE RECAPTURES AND FLUOROSCOPIC VALVE LOAD INSPECTION OF THE NEW VALVE WERE NOT PROVIDED FOR REVIEW. IMAGES OF A RELEASED VALVE REVEAL A DEEP IMPLANTATION, APPROXIMATELY 8 MILLIMETER (MM). IT WAS REPORTED THAT CARDIOPULMONARY RESUSCITATION EFFORTS WERE RESUMED AFTER VALVE IMPLANTATION DUE TO HYPOTENSION. A POST IMPLANT ANGIOGRAM WAS PERFORMED WHICH SHOWED EVIDENCE OF CORONARY PERFUSION TO THE LEFT CORONARY ARTERY; HOWEVER, IT IS DIFFICULT TO ASSESS RIGHT CORONARY PERFUSION DUE TO THE CHEST COMPRESSION SYSTEM THAT APPEARS TO BEING USED AT THE TIME OF THE ANGIOGRAM. IT WAS REPORTED THAT THE PATIENT HAD POOR VENTRICULAR FUNCTION THOUGH THE EJECTION FRACTION IS NOT KNOWN. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), SEVERE VENTRICULAR DYSFUNCTION WITH LEFT VENTRICULAR EJECTION FRACTION (LVEF) <(><<)>20% IS LISTED AS A PRECAUTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: EVOLUTFX-26, SERIAL #: (B)(6), UBD: 22-FEB-2026, UDI#: (B)(4). OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: L-EVOLUTFX-2329; PRODUCT LOT NUMBER: UNKNOWN; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE, PRODUCT ID: EVOLUTFX-26; SERIAL #: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762438&pc=NPT | 2025587-2025-02320 | 2025-03-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A "FRAIL PATIENT" WITH A P RE-EXISTING NON-MEDTRONIC SURGICAL AORTIC VALVE REPLACEMENT, THE VALVE WAS LOADED PROPERLY AND A FRAME OVERLAP TO THE 2ND NODE WAS DETECTED. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED AND DEPLOYMENT WAS ATTEMPTED; HOWEVER, HYPOTENSION WAS NOTED AS THE VALVE WAS OPENED AT APPROXIMATELY 90%. THE VALVE WAS POSITIONED TOO LOW AND A RECAPTURE WAS PERFORMED. THE GUIDEWIRE WAS EXCHANGED TO A STIFFER NON-MEDTRONIC GUIDEWIRE. A SECOND DEPLOYMENT WAS ATTEMPTED AND THE VALVE WAS IN A GOOD POSITION, BUT THE BLOOD PRESSURE DROPPED AGAIN. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PHYSICIAN DECIDED TO PERFORM A RECAPTURE AND WITHDRAW THE VALVE AND DCS FROM THE PATIENT AS VALVE FUNCTION WAS UNCERTAIN. A NEW VALVE WAS PROPERLY LOADED WHILE CPR WAS ONGOING. THE NEW VALVE WAS IMPLANTED IN GOOD POSITION, AND THE BLOOD PRESSURE RECOVERED TEMPORARILY. AFTER SEVERAL MINUTES, THE BLOOD PRESSURE DROPPED AGAIN WITH AND CPR RESUMED. SEVERAL DEFIBRILLATIONS WERE PERFORMED, BUT WERE UNSUCCESSFUL. AFTER APPROXIMATELY 20 MINUTES OF CPR, THE TEAM DECIDED TO STOP ALL ACTIONS AND THE PATIENT SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS NOT REPORTED, BUT IT WAS NOTED THAT A PERICARDIAL BLEED WAS NOT DETECTED AND THE CORONARY ARTERIES WERE PATENT.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A "FRAIL PATIENT" WITH A P RE-EXISTING NON-MEDTRONIC SURGICAL AORTIC VALVE REPLACEMENT, THE VALVE WAS LOADED PROPERLY AND A FRAME OVERLAP TO THE 2ND NODE WAS DETECTED. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED AND DEPLOYMENT WAS ATTEMPTED; HOWEVER, HYPOTENSION WAS NOTED AS THE VALVE WAS OPENED AT APPROXIMATELY 90%. THE VALVE WAS POSITIONED TOO LOW AND A RECAPTURE WAS PERFORMED. THE GUIDEWIRE WAS EXCHANGED TO A STIFFER NON-MEDTRONIC GUIDEWIRE (SAFARI). A SECOND DEPLOYMENT WAS ATTEMPTED AND THE VALVE WAS IN A GOOD POSITION, BUT THE BLOOD PRESSURE DROPPED AGAIN. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PHYSICIAN DECIDED TO PERFORM A RECAPTURE AND WITHDRAW THE VALVE AND DCS FROM THE PATIENT AS VALVE FUNCTION WAS UNCERTAIN. A NEW VALVE WAS PROPERLY LOADED WHILE CPR WAS ONGOING. THE NEW VALVE WAS IMPLANTED IN GOOD POSITION, AND THE BLOOD PRESSURE RECOVERED TEMPORARILY. AFTER SEVERAL MINUTES, THE BLOOD PRESSURE DROPPED AGAIN WITH AND CPR RESUMED. SEVERAL DEFIBRILLATIONS WERE PERFORMED, BUT WERE UNSUCCESSFUL. AFTER APPROXIMATELY 20 MINUTES OF CPR, THE TEAM DECIDED TO STOP ALL ACTIONS AND THE PATIENT SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS NOT REPORTED, BUT IT WAS NOTED THAT A PERICARDIAL BLEED WAS NOT DETECTED AND THE CORONARY ARTERIES WERE PATENT. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN THAT THE CAUSE OF DEATH WAS DUE TO WEAK LEFT VENTRICLE, RIGHT VENTRICLE AND PULMONARY HYPERTENSION. ACCORDING TO THE PHYSICIAN, THE ATTEMPTED FIRST VALVE, THE IMPLANTED SECOND VALVE, AND THE DCS CAN ALL BE EXCLUDED AS CONTRIBUTING CAUSES TO THE DEATH; HOWEVER, THE PHYSICIAN ALSO NOTED THAT THE PATIENT "DIDN¿T TOLERATE THE DROP IN BLOOD PRESSURE DURING RECAPTURE AND DIDN¿T RECOVER." FURTHER, THE PHYSICIAN INDICATED THAT THE PATIENT'S UNDERLYING MEDICAL HISTORY (POOR LEFT VENTRICLE AND RIGHT VENTRICLE FUNCTION, PULMONARY HYPERTENSION, HIGH RISK STATUS) CONTRIBUTED TO THE DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (SECOND PARAGRAPH) AND H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H2 H3 H6 IMAGE REVIEW: STATIC IMAGES AND MEDIA FILES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. IMAGES CONFIRMED THAT THE PATIENT HAD A PREVIOUSLY IMPLANTED SURGICAL AORTIC VALVE OF UNKNOWN TYPE AND SIZE. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. THERE IS EVIDENCE THAT AT SOME POINT DURING VALVE DEPLOYMENT ATTEMPT CARDIOPULMONARY RESUSCITATION WAS BEING PERFORMED. IT WAS REPORTED THAT THE FIRST VALVE WAS RECAPTURED, WITHDRAWN FROM THE PATIENT AND REPLACED WITH A NEW SYSTEM. IMAGES OF THE RECAPTURES AND FLUOROSCOPIC VALVE LOAD INSPECTION OF THE NEW VALVE WERE NOT PROVIDED FOR REVIEW. IMAGES OF A RELEASED VALVE REVEAL A DEEP IMPLANTATION, APPROXIMATELY 8 MILLIMETER (MM). IT WAS REPORTED THAT CARDIOPULMONARY RESUSCITATION EFFORTS WERE RESUMED AFTER VALVE IMPLANTATION DUE TO HYPOTENSION. A POST IMPLANT ANGIOGRAM WAS PERFORMED WHICH SHOWED EVIDENCE OF CORONARY PERFUSION TO THE LEFT CORONARY ARTERY; HOWEVER, IT IS DIFFICULT TO ASSESS RIGHT CORONARY PERFUSION DUE TO THE CHEST COMPRESSION SYSTEM THAT APPEARS TO BEING USED AT THE TIME OF THE ANGIOGRAM. IT WAS REPORTED THAT THE PATIENT HAD POOR VENTRICULAR FUNCTION THOUGH THE EJECTION FRACTION IS NOT KNOWN. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), SEVERE VENTRICULAR DYSFUNCTION WITH LEFT VENTRICULAR EJECTION FRACTION (LVEF) <(><<)>20% IS LISTED AS A PRECAUTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: D-EVOLUTFX-2329; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID: L-EVOLUTFX-2329; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID: EVOLUTFX-26, SERIAL #: (B)(6), UBD: 22-FEB-2026, UDI#: (B)(4). SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762458&pc=OZD | 1220648-2025-27209 | 2025-03-11 04:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Pumping Stopped; Application Program Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT HAD BEEN ESCALATED FROM AN IMPELLA CP TO AN IMPELLA 5.5 PUMP. THE PUMP WAS SUCCESSFULLY PLACED AFTER SOME ISSUES IN GAINING THE LEFT VENTRICLE ACCESS. DURING SUPPORT, THERE WERE MULTIPLE CONTROLLER ALARMS WHICH MADE THE PUMP STOP. THE AUTOMATED IMPELLA CONTROLLER WAS REPLACED AND THERE WERE NO ISSUES. SUBSEQUENTLY, THE PATIENT EXPIRED. Manufacturer Narrative: THE CONSOLE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - CONTROLLER FAILURE (RED ALARM) AND AIC - PUMP STOP HAS BEEN COMPLETED. AIC - CONTROLLER ERROR (YELLOW ALARM) WAS ADDED AS A FAILURE MODE. CONSOLE LOGS FROM THE REPORTED DAY OF THE EVENT ARE CONSISTENT WITH THE COMPLAINT. THE CASE STARTED ON (B)(6) 2025 AND ON (B)(6) 2025 MULTIPLE EPM RESET MESSAGES WAS SEEN AFTER WHICH MULTIPLE "CONTROLLER ERROR (POWER BATTERY MANAGER (PBM) VOLTAGE OUT-OF-SPEC) WAS TRIGGERED ALONG WITH THEM MULTIPLE CONTROLLER ERROR (LOSS OF (EPM). SUBSEQUENTLY, IMPELLA STOPPED (MECHANICAL/ELECTRICAL ISSUE) ALARM WAS TRIGGERED ON WHICH THE PURGE CASSETTE WAS REMOVED AS WELL AS SHUTDOWN WAS REQUESTED, HOWEVER "SHUTDOWN REQUESTED BY USER BUT PUMP IS CONNECTED" ERROR WAS DISPLAYED. ADDITIONALLY, FURTHER ERRORS WERE LOGGED, INCLUDING "PLACEMENT SIGNAL NOT RELIABLE. PUMP MOTOR DRIVER (PMD) HAS DETECTED A PUMP SPEED DEVIATION - ALARM ISSUED AND "CONTROLLER ERROR (OPM COMM CRC INVALID) [OPTICAL PUMP CONNECTED]." AFTER MULTIPLE REQUESTS FOR SHUTDOWN, A HARD SHUTDOWN WAS ULTIMATELY INITIATED AT 09:38:22 ON 11TH MARCH 2025. DEVICE ANALYSIS: THE CONSOLE WAS RETURNED TO FIELD SERVICE FOR FURTHER INVESTIGATION. A THOROUGH VISUAL INSPECTION OF THE INTERNAL COMPONENTS WAS CONDUCTED, BUT NO PHYSICAL PROBLEMS WERE FOUND. THE POWER SUPPLY FROM THE PBM TO THE IMPELLATRONIC PCA AND THE 4-IN-1 UNIT WAS STABLE. WHILE TEST THE AIC THE GREY PUSH BUTTON STOPPED WORKING, AND WHILE INVESTIGATING FOUND THAT THE 4IN1 WAS MESSING UP THE COMMUNICATION WITH ALL SUBSYSTEMS WITHIN THE AIC, FURTHER LOOKING INTO THE REAL TIME LOGS, WAS ABLE TO SEE THE SOME OF THE SIMILAR ERROR MESSAGES AS DISPLAYED DURING THE CASE. REPLACING THE 4IN1, AIC WORKED AS REQUIRED AND NO ISSUES WERE OBSERVED. ROOT CAUSE: THE ROOT CAUSE OF THE CONTROLLER ERRORS "LOSS OF COMM (EPM) AND PBM VOLTAGE OUT-OF-SPEC" WAS MOST LIKELY DUE TO THE DEFECTIVE 4IN1 BOARD. THE ROOT CAUSE OF THE IMPELLA STOPPED "MECHANICAL/ELECTRICAL ISSUE AND PMD HAS DETECTED A PUMP SPEED DEVIATION" WAS ALSO MOST LIKELY DUE TO THE DEFECTIVE 4IN1 BOARD. THERE WERE NO CONTROLLER FAILURE ALARM WAS DISCOVERED DURING THE CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762841&pc=OZD | 1220648-2025-27211 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND LATER THIS DAY, VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION WAS PLACED (VA ECMO). ADDITIONALLY, A HEMATOMA WAS NOTICED AROUND IMPELLA SITE. MANUAL PRESSURE WAS HELD FOR ONE HOUR AND HEMATOMA WAS RESOLVED. THE PATIENT REQUIRED TWO UNITS OF BLOOD. POSITIVE PEDAL PULSES REMAINED. VA ECMO AND IMPELLA WERE CONSISTENTLY HAVING SUCTION EVENTS. IMPELLA UNABLE TO RUN ABOVE P-1 PERFORMANCE LEVEL. A GASTROINTESTINAL BLEED WAS NOTED AND THE PATIENT RECEIVED EIGHT MORE UNITS OF BLOOD. TRANSFER TO A LARGER FACILITY WAS RECOMMENDED AND COMPLETED THE NEXT DAY. UPON ARRIVAL, THE PATIENT WAS STILL CRITICALLY ILL. FLOWS ON VA ECMO 3.5, HIGHEST OF 5.0. IMPELLA P-1 TO P-2, FLOWS 0.7 HIGHEST OF 2.4 WITH ECMO CIRCUIT CHANGE. THERE WAS A CONCERN ABOUT RETROPERITONEAL BLEED. CONTINUOUS RENAL REPLACEMENT THERAPY WAS STARTED. THE PATIENT RECEIVED MULTIPLE PUSHES OF EPINEPHRINE EN ROUTE FOR BRADYCARDIA. IMPELLA INSERTION SITE WAS DRY WITH KNOWN BRUISING, SOFT TO PALPATION. DISTAL PERFUSION WAS POOR WITH PATIENT¿S CURRENT STATE WITH NO PLANS TO ESCALATE TO IMPELLA 5.5 AS LEFT VENTRICLE WAS WELL VENTED WITH THE IMPELLA CP. THE OVERALL PLAN WAS TO GET THE PATIENT HEMODYNAMICALLY STABLE OVERNIGHT, EVALUATE THE BLEEDING. THE FOLLOWING DAY, IT WAS NOTED THAT THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR WITH THE EVENT OUTCOME. HOWEVER, TO NOTE, THE NURSE NOTED THE PATIENT CAME IN ASYSTOLE, CHEMICALLY CODED FOR FOUR HOURS AND MASSIVE TRANSFUSION PROTOCOL FOLLOWED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BLEEDING REPORTED AROUND THE ACCESS SITE, BLOOD GIVEN AND MANUAL PRESSURE HELD. THE ROOT CAUSE OF ACCESS SITE BLEEDING CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. LOW PUMP FLOW: SHORTLY AFTER ARRIVAL TO THE UNIT PATIENT HAD SUCTION ALARMS. THEY HAD TO BE PUT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THEY ALSO HAD ACCESS SITE BLEEDING REQUIRING 2 UNITS OF BLOOD. THE NEXT DAY BOTH CP AND ECMO WITH REPORTS OF SUCTION. DATA LOGS CONFIRM SUCTION ALARMS AND SHOW THE PLACEMENT SIGNAL (PS) TRENDING DOWN. NO MOTOR CURRENT (MC) SPIKES OUTSIDE OF P-LEVEL CHANGES. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED TO THE PATIENT'S DETERIORATING CONDITION, BLEEDING ISSUES REQUIRING BLOOD TO BE GIVEN, AND ECMO ALSO HAVING SUCTION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762843&pc=NPT | 2015691-2025-02553 | 2025-03-12 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-03 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN BRAZIL, DURING THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA VALVE, POST VALVE IMPLANTATION, THE PATIENT WAS NOTED TO BE PROGRESSING TO SHOCK AND CARDIAC TAMPONADE DUE TO SIGNIFICANT PERICARDIAL EFFUSION. THE SCINTIGRAPHY REVEALED AN AORTIC ANNULUS RUPTURE. A PUNCTURE WITH PERICARDIAL DRAINAGE WAS PERFORMED. ALTHOUGH THERE WAS CONTINUOUS HIGH-VOLUME DRAINAGE, BLEEDING PERSISTED WHICH INDICATED AN UNCONTROLLED LESION. SUBSEQUENTLY, THE PATIENT WENT INTO CARDIOPULMONARY ARREST WHICH PROMPTED AN EMERGENCY INTERVENTION. A PERICARDIAL WINDOW PROCEDURE WAS PERFORMED BUT WAS UNSUCCESSFUL. AFTER APPROXIMATELY 35 MINUTES OF CARDIOPULMONARY RESUSCITATION (CPR), THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (SUCH AS CALCIFICATION) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21762911&pc=KNT | 9611594-2025-00074 |  | Death | AVANOS MEDICAL INC. | 2025-04-03 04:00 | KNT | CORTRAK 2 NG/NI FEED, ELECTROMAGNETIC STYLET, ENFIT, CONNECTOR | Break; Malposition of Device | Foreign Body In Patient | K220588 |  | 1 | Event Description: IT WAS REPORTED THE NASO-JEJUNAL (NJ) TUBE WAS PLACED WITH THE NEW MAGNETIC DEVICE. THE CAREGIVERS WERE "CONCERNED AS THE TUBE BECAME STIFF VERY QUICKLY." THE PATIENT ENDED UP HAVING A REPLACEMENT ON (B)(6) 2025 WHICH WAS "ONLY AROUND 3-MONTHS LATER." MOREOVER, WHEN INTERVENTIONAL RADIOLOGIST (IR) WENT TO REPLACE IT, THE TUBE HAD RUPTURED, AND A SEGMENT HAD BROKEN OFF AND WAS LODGED IN THE DUODENUM MAKING IT DIFFICULT FOR A FURTHER NJ TO BE SITED BY IR. THE PATIENT WAS "NOT FIT ENOUGH FOR [A] OGD [OESOPHAGO-GASTRO-DUODENOSCOPY] TO REMOVE IT." THE PATIENT "DIED SHORTLY AFTER THE PROCEDURE FROM ASPIRATION PNEUMONIA. IT WAS NOTED THE PATIENT WAS "ADMITTED WITH AN UGI [UPPER GASTROINTESTINAL] BLEED."Event Description: ADDITIONAL INFORMATION RECEIVED 22-APRIL-2025 STATING THE CAUSE OF DEATH WAS RESPIRATORY INFECTION AND HEART FAILURE. THE PATIENT ALSO HAD A GASTRO-INTESTINAL BLEEDING. THE END SECTION OF THE TUBE HAD FRACTURED OFF WHEN CLINCIANS REMOVED IT AND COULD BE SEEN TO BE RETAINED IN THE DUODENUM. Manufacturer Narrative: THE DEVICE HISTORY RECORD FOR THE REPORTED LOT NUMBER, 30277235, IN THIS COMPLAINT WAS REVIEWED AND THE MATERIAL WAS PRODUCED ACCORDING TO THE MANUFACTURING PROCEDURES AND MET THE QUALITY REQUIREMENTS. ALL INFORMATION REASONABLY KNOWN AS OF 22-AUG-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY. ATTACHMENTS & OUTPUTS Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD AND UDI NUMBER ARE IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 03-APR-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5. ALL INFORMATION REASONABLY KNOWN AS OF 23-JUL-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763036&pc=NKE | 2182208-2025-01675 | 2024-08-17 04:00 | Death | MEDTRONIC, INC. | 2025-04-03 04:00 | NKE | MEDTRONIC CRT-P | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP) COMPARED TO BIVENTRICULAR PACING (BIVP) IN PATIENTS WITH MID-RANGE EJECTION FRACTION. THE AUTHORS DESCRIBED TWO PATIENT DEATHS IN THE BIVP GROUP AND ONE PATIENT DEATH IN THE CSP GROUP; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS IN BOTH GROUPS WHO WERE DIAGNOSED WITH A SUSPECTED INFECTION AND ALL OF THEM UNDERWENT AN EXTRACTION AND REIMPLANTATION. ONE PATIENT IN THE BIVP GROUP UNDERWENT REIMPLANTATION DUE TO TWIDDLER¿S SYNDROME. ONE PATIENT WHO UNDERWENT HIS BUNDLE PACING (HBP) IN THE CSP GROUP WAS FOUND TO HAVE ASCITES AND PLEURAL EFFUSION CAUSED BY TRICUSPID REGURGITATION SUSPECTED TO BE RELATED TO LEAD INTERACTION. THERE WERE HEART FAILURE, ALL-CAUSE, AND CARDIOVASCULAR RELATED HOSPITALIZATIONS. IN BOTH GROUPS, THERE WERE DEVICES EXPLANTED AND REPLACED DUE TO EARLY ELECTIVE REPLACEMENT INDICATOR (ERI) AND LEADS WITH RISING THRESHOLDS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: POTENTIAL MODEL NUMBERS INCLUDE: 8042, W1TR03, ADDRL1, CRT01, C4TR01, W4TR01, W1SR01, W1DR01, A2DR01, C2TR01, DTMA1QQ, AND W1TR01. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/73 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL OUTCOMES OF CONDUCTION SYSTEM PACING COMPARED TO BIVENTRICULAR PACING IN PATIENTS WITH MID-RANGE EJECTION FRACTION. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY. 2025. 68:111¿116. DOI: 10.1007/S10840-024-01882-Z MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763039&pc=NVN | 2182208-2025-01676 | 2024-08-17 04:00 | Death | MEDTRONIC, INC. | 2025-04-03 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP) COMPARED TO BIVENTRICULAR PACING (BIVP) IN PATIENTS WITH MID-RANGE EJECTION FRACTION. THE AUTHORS DESCRIBED TWO PATIENT DEATHS IN THE BIVP GROUP AND ONE PATIENT DEATH IN THE CSP GROUP; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS IN BOTH GROUPS WHO WERE DIAGNOSED WITH A SUSPECTED INFECTION AND ALL OF THEM UNDERWENT AN EXTRACTION AND REIMPLANTATION. ONE PATIENT IN THE BIVP GROUP UNDERWENT REIMPLANTATION DUE TO TWIDDLER¿S SYNDROME. ONE PATIENT WHO UNDERWENT HIS BUNDLE PACING (HBP) IN THE CSP GROUP WAS FOUND TO HAVE ASCITES AND PLEURAL EFFUSION CAUSED BY TRICUSPID REGURGITATION SUSPECTED TO BE RELATED TO LEAD INTERACTION. THERE WERE HEART FAILURE, ALL-CAUSE, AND CARDIOVASCULAR RELATED HOSPITALIZATIONS. IN BOTH GROUPS, THERE WERE DEVICES EXPLANTED AND REPLACED DUE TO EARLY ELECTIVE REPLACEMENT INDICATOR (ERI) AND LEADS WITH RISING THRESHOLDS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/73 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL OUTCOMES OF CONDUCTION SYSTEM PACING COMPARED TO BIVENTRICULAR PACING IN PATIENTS WITH MID-RANGE EJECTION FRACTION. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY. 2025. 68:111¿116. DOI: 10.1007/S10840-024-01882-Z. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763058&pc=DXY | 2182208-2025-01677 | 2024-08-17 04:00 | Death | MEDTRONIC, INC. | 2025-04-03 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP) COMPARED TO BIVENTRICULAR PACING (BIVP) IN PATIENTS WITH MID-RANGE EJECTION FRACTION. THE AUTHORS DESCRIBED TWO PATIENT DEATHS IN THE BIVP GROUP AND ONE PATIENT DEATH IN THE CSP GROUP; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS IN BOTH GROUPS WHO WERE DIAGNOSED WITH A SUSPECTED INFECTION AND ALL OF THEM UNDERWENT AN EXTRACTION AND REIMPLANTATION. ONE PATIENT IN THE BIVP GROUP UNDERWENT REIMPLANTATION DUE TO TWIDDLER¿S SYNDROME. ONE PATIENT WHO UNDERWENT HIS BUNDLE PACING (HBP) IN THE CSP GROUP WAS FOUND TO HAVE ASCITES AND PLEURAL EFFUSION CAUSED BY TRICUSPID REGURGITATION SUSPECTED TO BE RELATED TO LEAD INTERACTION. THERE WERE HEART FAILURE, ALL-CAUSE, AND CARDIOVASCULAR RELATED HOSPITALIZATIONS. IN BOTH GROUPS, THERE WERE DEVICES EXPLANTED AND REPLACED DUE TO EARLY ELECTIVE REPLACEMENT INDICATOR (ERI) AND LEADS WITH RISING THRESHOLDS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: POTENTIAL MODEL NUMBERS INCLUDE: 5075, 4024, 5024M, 6935, 6935M, 4968, 4193, 4196, 4298, 4798, 4598. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/73 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL OUTCOMES OF CONDUCTION SYSTEM PACING COMPARED TO BIVENTRICULAR PACING IN PATIENTS WITH MID-RANGE EJECTION FRACTION. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY. 2025. 68:111¿116. DOI: 10.1007/S10840-024-01882-Z MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763096&pc=BZD | 2518422-2025-103663 | 2022-07-15 04:00 | Death | RESPIRONICS, INC. | 2025-04-03 04:00 | BZD | DREAMSTATION AUTO | Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THROAT CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT WAS ADMITTED TO HOSPITAL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763107&pc=NPT | 2025587-2025-02327 | 2024-04-08 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Unspecified Infection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 MONTHS AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS A SHUNT INFECTION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763306&pc=CBK | 2518422-2025-103669 | 2025-03-08 05:00 | Death | RESPIRONICS, INC. | 2025-04-03 04:00 | CBK | TRILOGY | Insufficient Information | Insufficient Information | K093905 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ON (B)(6) 2025 REGARDING A PATIENT DEATH THAT OCCURRED ON (B)(6) 2025 WHILE THE PATIENT WAS ON VENTILATION THERAPY USING A TRILOGY 202 VENTILATOR. THE REPORTER STATED THE CUSTOMER WOULD LIKE TO KNOW IF THE TRILOGY 202 VENTILATOR ALARMED FOR APNEA DURING USE SINCE THE PATIENT DIED WHILE RECEIVING VENTILATION THERAPY. THE REPORTER STATED THEY ARE NOT SURE IF THE VENTILATOR ALARMED OR NOT. IT WAS REPORTED THE PATIENT WAS ON THE VENTILATOR AND HAD BEEN FOR A NUMBER OF DAYS WITH NO ISSUES. THE REPORTER INDICATED THE ERROR LOGS HAD BEEN REVIEWED BUT PROVIDED NO ADDITIONAL INFORMATION REGARDING THE FINDINGS OF THAT REVIEW. THE REPORTER INDICATED THEY WERE UNAWARE OF ANY OTHER DEVICES USED AT THE TIME OF THE EVENT OTHER THAN THE TRILOGY 202 VENTILATOR. MEDICAL ATTENTION WAS REPORTEDLY SOUGHT AS A DIRECT RESULT OF THE ISSUE, BUT THOSE DETAILS HAVE NOT BEEN PROVIDED. Manufacturer Narrative: MDR REPORT 2518422-2025-103669 IS A DUPLICATE OF MDR REPORT 2518422-2025-103169. MDR REPORT 2518422-2025-103669 WAS FILED FOR THE SAME ISSUE REPORTED ON MDR REPORT 2518422-2025-103169. 2518422-2025-103669 B5: THE MANUFACTURER RECEIVED INFORMATION ON 03/14/2025 REGARDING A PATIENT DEATH THAT OCCURRED ON 03/08/2025 WHILE THE PATIENT WAS ON VENTILATION THERAPY USING A TRILOGY 202 VENTILATOR. THE REPORTER STATED THE CUSTOMER WOULD LIKE TO KNOW IF THE TRILOGY 202 VENTILATOR ALARMED FOR APNEA DURING USE SINCE THE PATIENT DIED WHILE RECEIVING VENTILATION THERAPY. THE REPORTER STATED THEY ARE NOT SURE IF THE VENTILATOR ALARMED OR NOT. IT WAS REPORTED THE PATIENT WAS ON THE VENTILATOR AND HAD BEEN FOR A NUMBER OF DAYS WITH NO ISSUES. THE REPORTER INDICATED THE ERROR LOGS HAD BEEN REVIEWED BUT PROVIDED NO ADDITIONAL INFORMATION REGARDING THE FINDINGS OF THAT REVIEW. THE REPORTER INDICATED THEY WERE UNAWARE OF ANY OTHER DEVICES USED AT THE TIME OF THE EVENT OTHER THAN THE TRILOGY 202 VENTILATOR. MEDICAL ATTENTION WAS REPORTEDLY SOUGHT AS A DIRECT RESULT OF THE ISSUE, BUT THOSE DETAILS HAVE NOT BEEN PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763359&pc=OZD | 1220648-2025-27214 | 2025-02-13 05:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | IMPELLA 5.5 | Material Split, Cut or Torn | Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED AN IMPELLA 5.5 TO SUPPORT A PATIENT ADMITTED IN WITH CARDIOMYOPATHY. THE PUMP WAS PLACED AND HEART TRANSPLANT DECISIONS WERE BEING MADE. THE PUMP SUPPORTED AND ON DAY TWO, THERE WAS AN ACCESS SITE/MUSCLE BLEED THAT PROMPTED SURGICAL INTERVENTION AND RED BLOOD CELLS WERE GIVEN. THE PUMP ADDITIONALLY WAS DAMAGED AT THE STERILE SLEEVE AND WAS REMEDIED BY USE OF TEGADERM. THE PUMP REMAINED ON FOR OVER 42 DAYS. HOWEVER, THE PATIENT WAS PUT ON COMFORT CARE AND EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. THE ROOT CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL INFORMATION AND SINCE THE PRODUCT WAS NOT RETURNED. THE ROOT CAUSE OF THE PRODUCT DAMAGE (STERILE SLEEVE) ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND DUE TO INSUFFICIENT CLINICAL INFORMATION. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763611&pc=LWS | 2182208-2025-01682 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-03 04:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P980016 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ANTITACHYCARDIA PACING (ATP) AND COMPARING INTRINSIC ANTITACHYCARDIA PACING (IATP) TO CONVENTIONAL ATP (CATP). THE AUTHORS DESCRIBED PATIENT DEATHS IN THE IATP AND CATP GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN BUT DESCRIBED AS ALL-CAUSE, CARDIAC, OR ARRHYTHMIA RELATED. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED INAPPROPRIATE ATP; THE CAUSES OF THE INAPPROPRIATE THERAPY WERE UNKNOWN, AND REPROGRAMMING WAS PERFORMED. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE EFFICACY AND SAFETY OF INTRINSIC ANTITACHYCARDIA PACING. JOURNAL OF ARRHYTHMIA. 2025; 41:E13221. DOI: 10.1002/JOA3.13221 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763627&pc=NIK | 2182208-2025-01683 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-03 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P010031 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ANTITACHYCARDIA PACING (ATP) AND COMPARING INTRINSIC ANTITACHYCARDIA PACING (IATP) TO CONVENTIONAL ATP (CATP). THE AUTHORS DESCRIBED PATIENT DEATHS IN THE IATP AND CATP GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN BUT DESCRIBED AS ALL-CAUSE, CARDIAC, OR ARRHYTHMIA RELATED. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED INAPPROPRIATE ATP; THE CAUSES OF THE INAPPROPRIATE THERAPY WERE UNKNOWN, AND REPROGRAMMING WAS PERFORMED. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE EFFICACY AND SAFETY OF INTRINSIC ANTITACHYCARDIA PACING. JOURNAL OF ARRHYTHMIA. 2025; 41:E13221. DOI: 10.1002/JOA3.13221 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763653&pc=DXY | 2182208-2025-01684 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-03 04:00 | DXY | MEDTRONIC ICD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ANTITACHYCARDIA PACING (ATP) AND COMPARING INTRINSIC ANTITACHYCARDIA PACING (IATP) TO CONVENTIONAL ATP (CATP). THE AUTHORS DESCRIBED PATIENT DEATHS IN THE IATP AND CATP GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN BUT DESCRIBED AS ALL-CAUSE, CARDIAC, OR ARRHYTHMIA RELATED. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED INAPPROPRIATE ATP; THE CAUSES OF THE INAPPROPRIATE THERAPY WERE UNKNOWN, AND REPROGRAMMING WAS PERFORMED. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE EFFICACY AND SAFETY OF INTRINSIC ANTITACHYCARDIA PACING. JOURNAL OF ARRHYTHMIA. 2025; 41:E13221. DOI: 10.1002/JOA3.13221. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763701&pc=CBK | 2518422-2025-103659 | 2025-04-01 04:00 | Death | RESPIRONICS, INC. | 2025-04-03 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT ON (B)(6) 2025 AT 08:31, THE PATIENT'S MOTHER DETACHED THE CATHETER MOUNT TO PERFORM SPUTUM SUCTION. AT THAT TIME, A "CIRCUIT DISCONNECTED" ALARM SOUNDED. THE MOTHER WENT OUT AFTER THAT. AT 08:36 THE SAME MORNING, AN ALARM OCCURRED ON THE MONITOR (THE CATHETER MOUNT MAY HAVE COME OFF DURING THAT TIME). IT WAS REPORTED THAT THE PATIENT'S FATHER DID NOT NOTICE THIS. IT WAS REPORTED THERE WAS NO ALARM LOGGED ON THE TRILOGY EVO VENTILATOR. THE SAME MORNING, AT 09:30, THE PATIENT'S FATHER NOTICED THE CATHETER MOUNT WAS DETACHED BECAUSE A "LOW MINUTE" VENTILATION ALARM 1.8L/MIN. HAD OCCURRED. NO MEDICAL INTERVENTION WAS INITIATED. IT WAS CONFIRMED THE PATIENT HAD CARDIOPULMONARY ARREST. THE PATIENT'S CAUSE OF DEATH HAS BEEN REPORTED AS CARDIAC ARREST. CAUSAL RELATIONSHIP BETWEEN THE MEDICAL DEVICE AND THE PATIENT HARM IN THE OPINION OF THE DOCTOR WAS REPORTED AS: OCCURRENCE RECORD OF AN ALARM (OR THE CIRCUIT DISCONNECT ALARM) WHICH OCCURRED TUESDAY, (B)(6) 2025, IS PRESENT ON THE MONITOR. THE FATHER DID NOT NOTICE IT AT THE TIME. THERE IS A POSSIBILITY THAT THE CATHETER MOUNT CAME OFF DURING THE TIME INDICATED, BUT THERE IS NO ALARM LOGGED ON THE TRILOGY EVO VENTILATOR. THE DOCTOR REPORTEDLY CHECKED THE ALARM LOG ON SITE AND PROVIDED THEIR OPINION AS STATED HERE. ADDITIONAL INFORMATION PROVIDED: ON 01 APRIL 2025, THE PHILIPS SALES REPRESENTATIVE CHECKED THE OPERATING STATUS OF THE TRILOGY EVO VENTILATOR AT THE PATIENT'S HOME TOGETHER WITH THE POLICE. THEY CONFIRMED THAT THE DEVICE OPERATED NORMALLY AND ANNUNCIATED ALARMS. IT WAS REPORTED IT MAY TAKE TIME FOR THE DEVICE TO BE RETRIEVED FOR MANUFACTURER'S INVESTIGATION SINCE THE POLICE MAY COLLECT IT. THE DEVICE WAS NOT AVAILABLE AT THE TIME OF REPORTING. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER AND EVALUATED ON 24 APRIL 2025 WITH THE FOLLOWING FINDINGS: WHEN THE CATHETER MOUNT WAS REMOVED USING THE SAME CIRCUIT WITH THE SETTING OF THE ISSUE OCCURRING, IT WAS CONFIRMED THAT "LOW MINUTE VENTILATION" AND "CIRCUIT DISCONNECTED" ALARM OCCURRED. PERFORMED AN OPERATIONAL CHECK AS USUAL, AND THERE WERE NO ABNORMALITIES FOUND. UPON CHECKING THE ERROR LOG, IT WAS CONFIRMED THAT "LOW MINUTE VENTILATION" ALARM OCCURRED FOR 3482 SECONDS AND "CIRCUIT DISCONNECTED" ALARM OCCURRED FOR 3475 SECONDS. THERE WERE NO ERRORS INDICATING ABNORMALITIES OF THE DEVICE. WHEN TESTS FOR THE FLOW OR PRESSURE WERE PERFORMED AS A VERIFICATION, NO ABNORMALITIES WERE OBSERVED. THERE WERE NO ABNORMALITIES FOUND DURING AN INTERNAL INSPECTION WITH DISASSEMBLING THE DEVICE. BASED ON THE ABOVE, IT IS BELIEVED THAT THERE WERE NO ABNORMALITIES IN THE DEVICE, AND THE CATHETER MOUNT WAS REMOVED, THEN "LOW MINUTE VENTILATION" AND "CIRCUIT DISCONNECTED" ALARM KEPT SOUNDING. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763723&pc=DSQ | 2916596-2025-02167 | 2025-03-04 05:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF THE DEATH WAS UNKNOWN AND WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763906&pc=DSQ | 2916596-2025-01915 | 2025-03-14 04:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THERE NO WAS ADDITIONAL INFORMATION. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21763907&pc=DSQ | 2916596-2025-02168 | 2025-02-03 05:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF THE DEATH WAS UNKNOWN AND WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21764009&pc=QAN | 2183870-2025-00161 | 2025-01-27 05:00 | Death | COVIDIEN | 2025-04-03 04:00 | QAN | ABRE VENOUS SELF-EXPANDING STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200026 |  | 1 | Event Description: LITERATURE EVENT WAS RECEIVED: THIS STUDY AIMED TO EVALUATE THE CLINICAL OUTCOMES, TECHNICAL SUCCESS, AND COMPLICATIONS ASSOCIATED WITH THE PLACEMENT OF ABRE¿ VENOUS STENTS FOR CENTRAL VENOUS STENOSIS/OCCLUSION (CVS/CVO). MATERIALS AND METHODS A RETROSPECTIVE REVIEW WAS CONDUCTED, IDENTIFYING 21 PATIENTS WHO UNDERWENT ABRE¿ VENOUS STENT PLACEMENT FOR SVC SYNDROME AT A SINGLE INSTITUTION. DEMOGRAPHIC AND CLINICAL DATA WERE COLLECTED, INCLUDING STENT CONFIGURATIONS, PROCEDURAL DETAILS, AND FOLLOW-UP OUTCOMES. PRIMARY STENT PATENCY, SYMPTOM RESOLUTION, AND SURVIVAL RATES WERE ANALYZED. SURVIVAL CURVES WERE GENERATED USING KAPLAN-MEIER ANALYSIS, AND COMPLICATIONS WERE RECORDED. RESULTS FACIAL AND/OR UPPER EXTREMITY EDEMA WAS PRESENT IN ALL PATIENTS, AND MALIGNANCY ACCOUNTED FOR SVC SYNDROME IN 71.4% OF CASES, WITH LUNG CANCER AS THE PREDOMINANT ETIOLOGY (66.6%). THE TECHNICAL SUCCESS RATE WAS 100%. CLINICAL SYMPTOM RESOLUTION WAS ACHIEVED IN ALL PATIENTS. THE 30-DAY MORTALITY RATE WAS 23.8%, WITH ALL CASES INVOLVING ONCOLOGY PATIENTS. AT SIX MONTHS, CROSS-SECTIONAL IMAGING SHOWED A PRIMARY STENT PATENCY RATE OF 93%, AND THIS REMAINED STABLE THROUGH 12 MONTHS. THE MEAN SURVIVAL TIME FOR THE CANCER SUBGROUP WAS 337.2 ± 343 DAYS, WHILE THE OVERALL COHORT MEAN WAS 885.8 ± 453.7 DAYS. CONCLUSION THE ABRE¿ VENOUS STENT DEMONSTRATES POTENTIAL AS AN EFFECTIVE STENT FOR SVC SYNDROME, ACHIEVING HIGH SYMPTOM RELIEF AND PATENCY RATES. HOWEVER, FURTHER LONG-TERM STUDIES AND RANDOMIZED CONTROLLED TRIALS ARE NECESSARY TO VALIDATE THESE FINDINGS. A TOTAL OF 21 PATIENTS WHO UNDERWENT ABRE¿ VENOUS STENT PLACEMENT FOR SUPERIOR VENA CAVA (SVC) SYNDROME WERE IDENTIFIED THROUGH A RETROSPECTIVE INSTITUTIONAL DATABASE REVIEW. SYMPTOM RESOLUTION WAS ACHIEVED IN ALL PATIENTS (100%). THE 30-DAY POST-PROCEDURAL MORTALITY RATE WAS 23.8% (5/21), ALL OF WHOM WERE ONCOLOGY PATIENTS PRESENTING WITH ACUTE SVC SYNDROME AND ASSOCIATED AIRWAY COMPROMISE. AT SIX MONTHS, CROSS-SECTIONAL CT IMAGING WAS AVAILABLE. AT SIX MONTHS, CROSS-SECTIONAL CT IMAGING WAS AVAILABLE FOR 15 PATIENTS (71.4%). SIX PATIENTS WERE EXCLUDED FROM PRIMARY STENT PATENCY ANALYSIS DUE TO DEATH. PRIMARY STENT PATENCY AT SIX MONTHS WAS 93.3% (14/15). ONE PATIENT EXPERIENCED STENT OCCLUSION 59 DAYS POST-DEPLOYMENT DUE TO TUMOR INFILTRATION. THIS PATIENT UNDERWENT ASPIRATION THROMBECTOMY, BALLOON MACERATION, AND PLACEMENT OF A NEW OVERLAPPING STENT. THE RETRIEVED SPECIMEN CONFIRMED MALIGNANT INFILTRATION. THE STENT REMAINED PATENT FOR THE REMAINDER OF THE STUDY PERIOD, SPANNING AN ADDITIONAL 84 DAYS. AT 12 MONTHS, STENT PATENCY REMAINED UNCHANGED FOR THE COHORT. THREE ADDITIONAL PATIENTS WERE EXCLUDED FROM PATENCY ANALYSIS DUE TO DEATH. A PATIENT WITH END-STAGE RENAL DISEASE (ESRD) PRESENTED WITH A CRUSHED SVC STENT AND MULTIFOCAL AREAS OF SEVERE STENOSIS AFTER 4 YEARS. THIS WAS SUCCESSFULLY TREATED WITH ANGIOPLASTY AND THE PLACEMENT OF A NEW STENT WITHIN THE EXISTING SVC STENT. Manufacturer Narrative: JOURNAL TITLE: SAFETY AND EFFICACY OF THE ABRE¿ VENOUS STENT IN TREATING THORACIC VENOUS STENOSIS/OCCLUSION: A SINGLE-CENTER EXPERIENCE EMERGENCY RADIOLOGY HTTPS://DOI.ORG/10.1007/S10140-025-02316-7 A2 AVERAGE AGE A3 MAJORITY GENDER B3 DATE OF PUBLICATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21764020&pc=NPT | 2025587-2025-02333 | 2024-01-05 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Aspiration/Inhalation; Stroke/CVA; Paralysis; Pneumonia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DEVELOPED A CEREBRAL INFRACTION. SYMPTOMS IMPROVED; HOWEVER, THE LEFT SIDE WAS LEFT WITH INCOMPLETE PARALYSIS. THE PATIENT WAS TRANSFERRED TO A DIFFERENT HOSPITAL FOR REHABILITATION. APPROXIMATELY 5 MONTHS AND 1 WEEK LATER, THE PATIENT DEVELOPED ASPIRATION PNEUMONIA AND PASSED AWAY. THE CAUSE OF DEATH WAS ASPIRATION PNEUMONIA.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PARALYSIS OF THE LOWER LIMB WAS OBSERVED 4 DAYS FOLLOWING IMPLANT OF THE VALVE, AND THE STROKE WAS CONFIRMED ON MAGNETIC RESONANCE IMAGING (MRI). THE PATIENT WAS DISCHARGED FROM THE HOSPITAL FOLLOWING REHABILITATION. IT WAS FURTHER REPORTED THAT THE STROKE DID NOT CONTRIBUTE TO THE ASPIRATION PNEUMONIA. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21764064&pc=LWS | 2017865-2025-39163 | 2025-03-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-03 04:00 | LWS | DURATA STS OPTIM ACTIVE FIXATION | High impedance | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH VOLTAGE (HV) IMPEDANCE ANOMALY. ON (B)(6)2025, THE PATIENT PRESENTED TO THE HOSPITAL TO EXTRACT THE RIGHT ATRIAL (RA) AND RV LEADS. DURING THE PROCEDURE, THE SUPERIOR VENA CAVA (SVC) WAS OCCLUDED AND THE PHYSICIAN WAS ONLY ABLE TO EXTRACT THAT RA LEAD AND A PORTION OF THE RV LEAD. THE PHYSICIAN DECIDED TO IMPLANT AN ATRIAL LEADLESS PACEMAKER (LP). THE LP IMPLANT WAS SUCCESSFUL, HOWEVER, AFTER AN HOUR THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT 'CODED.' THE PATIENT WAS TREATED WITH ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND LATER PUT ON LIFE SUPPORT. IT WAS NOTED THAT THERE WAS AN EFFUSION AT THE RIGHT ATRIUM AND INFERIOR VENA CAVA (IVC) JUNCTION. THE PHYSICIAN SUSPECTED THAT THIS WAS CAUSED BY LEAD EXTRACTION FROM THE GROIN. THE PATIENT WAS REMOVED FROM LIFE SUPPORT A FEW DAYS AFTER. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21765712&pc=DSQ | 2916596-2025-01902 | 2024-12-14 05:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Chest Pain; Purulent Discharge; Fever; Cough; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FROM (B)(6) 2024 FOR A RECURRENT DRIVELINE INFECTION (ONSET REPORTED UNDER MFR 2916596-2024-07317). THE PATIENT WAS DETERMINED NOT TO BE A SURGICAL CANDIDATE AND DECIDED TO BE PLACED INTO HOSPICE CARE AND CONTINUE LEFT VENTRICULAR ASSIST DEVICE (LVAD) SUPPORT. THE PATIENT PASSED AWAY ON (B)(6) 2025Event Description: THE PATIENT'S DRIVELINE INFECTION HAD CULTURES OF PROTEUS MIRABILIS AND ENTEROCOCCUS FAECALIS, AND PATIENT SYMPTOMS OF FEVER, COUGH, LEFT CHEST PAIN, AND PURULENT DRAINAGE FROM THE DRIVELINE. THE INFECTION WAS TREATED WITH INTRAVENOUS (IV) ANTIBIOTICS, ALTHOUGH THE INFECTION DID NOT RESOLVE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: EVENT DATE CORRECTED BACK. SECTION E1; REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21766051&pc=DSQ | 2916596-2025-01959 | 2025-03-09 05:00 | Death | THORATEC CORPORATION | 2025-04-03 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Renal Failure; Sepsis; Septic Shock; Cardiogenic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT REMOVED THEIR OWN RIGHT VENTRICULAR ASSIST DEVICE (RVAD) CANNULA ON (B)(6) 2025, AND SHORTLY AFTER REQUIRED REINTUBATION RELATED TO RESPIRATORY DISTRESS AND DEVELOPED WORSENING CARDIOGENIC SHOCK AND AN ACUTE KIDNEY INJURY (AKI) THAT REQUIRED CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) STARTING ON (B)(6) 2025. THE PATIENT LATER DEVELOPED SEPTIC SHOCK RELATED TO PNEUMONIA AND CHOLECYSTITIS. A PERCUTANEOUS DRAIN WAS PLACED ON (B)(6) 2025, WITH CONCERN NOTED FOR PERITONITIS AND ISCHEMIC GUT FROM A BILE LEAK. INOTROPES WERE INITIATED. THE PATIENT PROGRESSED TO MULTISYSTEM ORGAN FAILURE AND THE FAMILY DECIDED TO PURSUE COMFORT CARE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE DEVICE WAS NOTED TO HAVE FUNCTIONED AS INTENDED.Event Description: IT WAS REPORTED THAT PATIENT HAD DISTRESS TOLERANCE DUE TO TOBACCO AND THAT THEIR LEFT VENTRICLE APEX WAS CALCIFIED. A RIGHT VENTRICULAR ASSIST DEVICE WAS PLACED POST OPERATION AND WAS SET TO 3 LPM. IT WAS SAID THAT THE PUMP RAN WELL THROUGHOUT AND THAT THE PATIENTS POST OPERATIONAL VITALS WERE AT A HEART RATE OF 80BPM, BLOOD PRESSURE OF 108/75 MMHG, A PULMONARY ARTERIAL PRESSURE OF 47/36, AND A CENTRAL VENOUS PRESSURE OF 20 MMHG. Manufacturer Narrative: ADDITIONAL CODES HEALTH EFFECT - CLINICAL CODES: 1758 - CALCIUM DEPOSITS/CALCIFICATION, 1696 - RESPIRATORY DISTRESS, 2011 - PNEUMONIA, 2252 - PERITONITIS, 4820 - BILIARY LEAK, 1932 - INFLAMMATION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WOULD NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH, SEPSIS, HYPERTENSION, AND LOCALIZED INFECTION AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ ALSO INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. SECTION 6 ALSO STATES THAT POST-IMPLANTATION HYPERTENSION MAY BE TREATED AT THE DISCRETION OF THE ATTENDING PHYSICIAN, AND ANY THERAPY THAT CONSISTENTLY MAINTAINS MEAN ARTERIAL BLOOD PRESSURE LESS THAN 90 MM HG (MILLIMETERS OF MERCURY) SHOULD BE CONSIDERED ADEQUATE. THE HEARTMATE 3 LVAS PATIENT HANDBOOK CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6 - ADDITIONAL HEALTH EFFECT CLINICAL CODES: 1696 - RESPIRATORY DISTRESS 2011 - PNEUMONIA 2252 - PERITONITIS 4820 - BILIARY LEAK. 1932 - INFLAMMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL CODES: HEALTH EFFECT - CLINICAL CODE: 1758 - CALCIUM DEPOSITS/CALCIFICATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21766263&pc=CBK | 2518422-2025-041168 | 2022-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-04-03 04:00 | CBK | TRILOGY 100 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY WHILE USING A VENTILATOR. THE INFORMATION INDICATES THAT DUE TO A MISTAKE BY A DOCTOR IN THE TREATMENT FOR REPLACING THE TRACHEOSTOMY STOMA CANNULA EVERY TWO WEEKS. THE UNIT HAS BEEN SEIZED BY THE POLICE. THERE IS NO ALLEGATION THE VENTILATOR MALFUNCTIONED OR FAILED TO DELIVER THERAPY AS DESIGNED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21766459&pc=BZD | 2518422-2025-103667 | 2025-03-28 04:00 | Death | RESPIRONICS, INC. | 2025-04-03 04:00 | BZD | DREAMSTATION AUTO CPAP, JP | Patient Device Interaction Problem | Aspiration Pneumonitis | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE USER DEVELOPED ASPIRATION PNEUMONIA FROM DEVICE DUE TO VOMIT AND LATER PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21766510&pc=OZD | 1220648-2025-27219 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis; Sepsis | P140003 |  | 1 | Event Description: INFORMATION RECEIVED VIA LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. HE WAS ADMITTED WHILE AWAITING ADVANCED THERAPY DECISIONS, SUCH AS LEFT VENTRICULAR ASSIST DEVICE (LVAD) OR NOT. IT WAS REPORTED THAT THE PATIENT ENDURED HEMOLYSIS WITH "POSSIBLY RELATED" TO THE USE OF THE PUMP. PER LOQI: HEMOGLOBIN DROPPED ON (B)(6) 2025 FROM 7.6 FROM 8.6. PLASMA HGB CHANGED. PATIENT WAS DIAGNOSED WITH HEMOLYTIC ANEMIA ON (B)(6) TO IMPELLA. 2 UNITS OF RBCS TOTAL WERE GIVEN DURING ADMISSION. PATIENT CONTINUED TO SHOW HEMOLYSIS UNTIL DEATH PER NOTES AND PI. PLAMA FREE HGB NADIR WAS 7.5 (ON (B)(6) 2025) PEAK 67.1 (ON (B)(6) 25); BASELINE UNKNOWN FOR LDH AND PLASMA HGB. PEAK LDH 759 (ON (B)(6) 2025) TOTAL BIL PEAK 5.4 (ON (B)(6) 2025). ADDITIONALLY, THE PATIENT ENDURED SEPSIS WITH "POSSIBLY RELATED" TO THE PUMP USE. LOQI STATED: PATIENT HAD B/C ON (B)(6) 2025 +MRSE. BC ON (B)(6) 2025 GPC CLUSTERS. PATIENT'S CAUSE OF INFECTION, PER ID, LIKELY RELATED TO IMPELLA. PATIENT TREATED WITH ANTIBIOTICS FOR BACTEREMIA. Manufacturer Narrative: THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. B2 DATE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21766629&pc=OZD | 1220648-2025-27222 | 2025-03-11 04:00 | Death | ABIOMED, INC | 2025-04-03 04:00 | OZD | IMPELLA CP | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT UNDERGOING AORTIC VALVULOPLASTY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT HIGH PURGE PRESSURE/LOW PURGE FLOWS WERE ENCOUNTERED DURING IMPELLA CP SUPPORT. THE PURGE CASSETTE WAS REPLACED DURING TROUBLESHOOTING, BUT THE ISSUE PERSISTED. IT WAS NOTED THAT THE PURGE WAS RUNNING WITHOUT SODIUM BICARBONATE AT THE TIME OF THE NOTED ISSUE. A TISSUE PLASMINOGEN ACTIVATOR WAS ORDERED TO HELP MANAGE THE CONDITION. SUPPORT WITH THE IMPLANTED PUMP CONTINUED UNINTERRUPTED. TWO DAYS LATER, CARE WAS WITHDRAWN AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21766786&pc=OZD | 1220648-2025-27225 | 2025-03-11 04:00 | Death | ABIOMED, INC. | 2025-04-03 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IMPELLA INSERTED VIA LEFT FEMORAL ARTERY PRIOR TO PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH FLOWS ON AUTO OF 3.8L THROUGHOUT CASE. PATIENT WAS GIVEN BOLUS TWICE AT BEGINNING OF CASE DUE TO SMALL SIZE OF LEFT VENTRICLE AND ONLY REQUIRED SMALL DOSE OF NEO DURING PROLONGED HYPOTENSION. THE PATIENT TOLERATED IMPELLA WEAN AT THE CONCLUSION OF PCI AND REMOVAL OF THE SHEATH WAS SLOW AND WITH CAUTION DUE TO PATIENT'S INABILITY TO RECEIVE BLOOD PRODUCTS (JEHOVAH WITNESS). HEMOSTASIS WAS ACHIEVED WITH PERCLOSE X2 AND LIGHT MANUAL PRESSURE FOR 10 MINUTES. THE PATIENT WAS FEELING WELL POST PCI. ON 11-MAR-2025, A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY INDICATING THAT THE PATIENT EXPERIENCED VENTRICULAR ARRHYTHMIA WITH A "REMOTE (UNLIKELY)" RELATIONSHIP TO THE IMPELLA DEVICE USED. THE PATIENT EXPIRED ON (B)(6) 2024. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767018&pc=NPT | 2025587-2025-02345 | 2024-01-25 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Unspecified Infection; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH SEVERE AORTIC STENOSIS AND DE CREASED CARDIAC FUNCTION, THE PATIENT'S CONDITION GRADUALLY WORSENED DUE TO POOR FOOD INTAKE AND CONCURRENT INFECTION. THIS LED TO AN EXACERBATION OF HEART FAILURE. THE PATIENT PASSED AWAY OF INFECTION APPROXIMATELY 2 WEEKS AFTER THE PROCEDURE. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE INFECTION OR DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767035&pc=MIH | 3007284313-2025-03878 | 2025-02-28 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-03 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Migration or Expulsion of Device; Malposition of Device; Migration | Hemorrhage/Bleeding | P020004 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT A PATIENT DEATH OCCURRED AROUND BEGINNING OF (B)(6) 2025 (EXACT DATE UNKNOWN) WITH A 72-YEAR-OLD MALE PATIENT. IMPLANTATION WAS WITH THE GORE® EXCLUDER®AAA ENDOPROSTHESIS (RLT261412) ON (B)(6) 2025. THE IMPLANT PROCEDURE WAS REPORTEDLY NOT SUCCESSFUL. IT WAS REPORTED THAT AN EXPLANT WAS NEEDED AT THE SAME DAY. DURING THE IMPLANT PROCEDURE, AFTER THE FIRST DEPLOYMENT OF THE TRUNK-IPSILATERAL-LEG (RLT261412) AND WITH THE REMOVED RECONSTRAINING LOOP (REPOSITIONING), AN UNINTENTIONAL MIGRATION HAS OCCURRED. UNFORTUNATELY, THE DEVICE RLT261412 COULD NOT BE SNARED AND PHYSICIAN DECIDED TO DO A CONVERSION. REPORTEDLY, THE MIGRATION IN CRANIAL DIRECTION / "JUMP" OF THE WHOLE TRUNK OCCURRED AFTER THE REPOSITIONING WAS REMOVED. THE DEVICE MIGRATED UPWARDS IN TOWARDS THE THORACIC-ABDOMINAL AORTA, >20MM FROM THE INTENDED POSITION AT THE INFRARENAL POSITION. SO, NO DETAILED EXPLANATION HAS BEEN PROVIDED, A "JUMP" SEEMED TO HAVE HAPPENED WHICH REQUIRED EMERGENT REINTERVENTION. THE PATIENT WAS IN GENERAL ANESTHETIC DURING THE IMPLANT PROCEDURE. REPORTEDLY, THE PATIENT DIED IN ICU (INTENSIVE CARE UNIT), AFTER A FEW DAYS AT AN UNKNOWN DATE, WITH A DEATH CAUSED BY AN UNINTENTIONAL MASSIVE-BLOOD FLOW DURING EXPLANT PROCEDURE, FOLLOWED BY A HYPOXIC BRAIN DAMAGE. REPORTEDLY, THE FAMILY WANTED TO STOP THE LIFE-SUSTAINING DEVICES IN ICU AS THIS WAS THE PATIENT'S LAST WILL. THE EXPLANT WAS DISCARDED AT THE IMPLANT FACILITY. ANGIOGRAM DATA IS BEING FINALIZED AT THE HEALTHCARE PROFESSIONAL (HCP) AND REQUESTED FOR EVALUATION AT GORE, THEN EVALUATE THE POSITIONING AND SEE THE CTA AFTER REPOSITIONING WAS REMOVED.Event Description: IT WAS REPORTED TO GORE THAT ON (B)(6) 2025, THIS PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE AND WAS IMPLANTED WITH GORE® EXCLUDER® AAA ENDOPROSTHESIS DEVICES TO TREAT AN ABDOMINAL AORTIC ANEURYSM. ENSUING INITIAL DEPLOYMENT OF THE GORE® EXCLUDER® TRUNK-IPSILATERAL LEG COMPONENT RLT261412, THE PHYSICIAN REPOSITIONED THE DEVICE, DISENGANGED THE CONSTRAINING MECHANISM AND REMOVED THE TRANSPARENT KNOB FROM THE CATHETER HANDLE, WHICH REMOVED THE CONSTRAINING LOOP FROM THE DEVICE. IMMEDIATELY THEREAFTER, THE DEVICE COMPONENT WAS REPORTED TO HAVE JUMPED AND MIGRATED PROXIMALLY FOR A LENGTH OF MORE THAN 20 MM. AS IT WAS NOT POSSIBLE TO CORRECT THE POSITION OF THE DEVICE COMPONENT BY ENDOVASCULAR MEANS, THE PHYSICIAN DECIDED TO PERFORM OPEN SURGERY TO REMOVE THE DEVICE. HOWEVER, THE PATIENT EXPERIENCED SUBSTANTIAL BLOOD LOSS DURING THE SURGERY, RESULTING IN HYPOXIC BRAIN DAMAGE. FOLLOWING THE PATIENT¿S WILL, HIS FAMILY AS A HEALTH CARE PROXY DECIDED TO NO LONGER EXTEND HIS LIFE. THE PATIENT EXPIRED A FEW DAYS POST PROCEDURE ON AN UNKNOWN DATE IN THE INTENSIVE CARE UNIT. Manufacturer Narrative: H3: NO DEVICE IS AVAILABLE FOR THE EVALUATION, IT WAS DISCARDED AT IMPLANT FACILITY. H6: B14: THE INVESTIGATION IS ONGOING. PRELIMINARY RESULTS ARE NOT YET AVAILABLE. B15: IMAGING SERIES IS REQUESTED, WAITING FOR THE HEALTHCARE PROFESSIONAL TO SEND TO GORE. B13: INTERVIEWS WERE DONE WITH COMPANY REPRESENTATIVE AND THE HEALTHCARE PROFESSIONAL / IMPLANT PHYSICIAN. COMMUNICATION LOOPS WERE CLOSED TO GATHER SUFFICIENT INFORMATION ABOUT THIS EVENT. DATE OF DEATH REMAINS UNKNOWN. THE OPERATOR HAD DISCOVERED A BRAIN DAMAGE AS A RESULT OF HEMORRHAGE DURING THE EXPLANT PROCEDURE. E0506: THIS WAS REPORTED BY PHYSICIAN, AND IT MIGHT IS THE CAUSE OF THE ADVERSE OUTCOME. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: UPDATED G3. B5: DESCRIBE EVENT OR PROBLEM, UPDATED THE EVENT DESCRIPTION. H6: CODE C19: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE WAS NOT RETURNED TO GORE AS IT WAS DISCARDED AT THE HCP SITE. H6: CODING: MEDICAL DEVICE PROBLEM (ANNEX A), CODE A010402 ADDED. CODE A0104 IS NO LONGER APPLICABLE. TYPE OF INVESTIGATION, CODE B11. THE HISTORICAL DATA ANALYSIS WAS COMPLETED. COMPONENT CODE, G04122 ADDED. INVESTIGATION FINDINGS CODE (ANNEX C), CODE C19 ADDED. C21 IS NO LONGER APPLICABLE. INVESTIGATION CONCLUSIONS (ANNEX D), CODE D15 ADDED. D16 IS NO LONGER APPLICABLE. INVESTIGATION SUMMARY: THE REPORTED PROXIMAL MIGRATION OF THE DEVICE WAS CONFIRMED DURING EVALUATION OF THE RETURNED CLINICAL IMAGES. THE REPORTED INFORMATION INDICATES THE ENDOPROSTHESIS ¿JUMPED¿ PROXIMALLY DURING REMOVAL OF THE CONSTRAINING MECHANISM. THE PROVIDED CLINICAL IMAGING CONFIRMS THE DEVICE, WITH A PARTIALLY CONSTRAINED PROXIMAL END, WAS LOCATED DISTAL TO THE RENAL ARTERIES BEFORE SUBSEQUENT IMAGING SHOWS THE DEVICE APPROXIMATELY 10CM ABOVE THE RENAL ARTERIES. THE CAUSE FOR THE PROXIMAL DEVICE MIGRATION COULD NOT BE ESTABLISHED WITH THE AVAILABLE INFORMATION. THE RISK MANAGEMENT FILE FOR THE GORE® EXCLUDER® AAA ENDOPROSTHESIS DEVICE WAS REVIEWED AND DETERMINED TO BE ACCURATE AND COMPLETE. NO UPDATE IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767329&pc=MVK | 3008642652-2025-03501 | 2025-02-16 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-03 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 09:20:50 ON (B)(6) 2025. AT 12:48:48 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SEVERE BRADYCARDIA @ 10 BPM WITH MOTION ARTIFACT. AT 12:49:39, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SEVERE BRADYCARDIA @ 10 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS SEVERE BRADYCARDIA @ 10 BPM WITH MOTION ARTIFACT. AT 12:51:58, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM. AT 12:52:29, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM WITH CPR/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 12:53:01 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767334&pc=MVK | 3008642652-2025-03502 | 2025-01-31 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-03 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON AN UNKNOWN DATE. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS ON THE LAST WEAR DATE. THE DEVICE WAS STARTED UP AT 00:08:24 ON (B)(6) 2025. AT 16:11:35, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 100 BPM WITH TALL T WAVES. AT 16:12:59, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. DOUBLE COUNTING CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 100 BPM WITH TALL T WAVES. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM WITH TALL T WAVES. AT 16:14:34, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 90 BPM WITH TALL T WAVES. AT 16:14:34, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. DOUBLE COUNTING CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS RHYTHM @ 90 BPM WITH TALL T WAVES AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 90 BPM WITH TALL T WAVES. AT 16:28:39, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM WITH TALL T WAVES. AT 16:14:34, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. DOUBLE COUNTING CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS RHYTHM @ 60 BPM WITH TALL T WAVES. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH TALL T WAVES. AT 16:30:44, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 100 BPM WITH TALL T WAVES. AT 16:31:14, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. DOUBLE COUNTING CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 100 BPM WITH TALL T WAVES. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 90 BPM WITH TALL T WAVES. THE ELECTRODE BELT WAS DISCONNECTED AT 17:02:57 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767341&pc=MVK | 3008642652-2025-03504 | 2025-03-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-03 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 14:13:37 ON (B)(6) 2025. AT 03:38:54 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 03:40:08, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIA SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 04:26:20 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767344&pc=MVK | 3008642652-2025-03505 | 2024-12-27 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-03 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Structural Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 15:35:10 ON (B)(6) 2024. AT 11:37:26 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF @ 60 BPM DEGRADING TO VT/VF WITH VARYING AMPLITUDES. AT 11:38:17, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH VARYING AMPLITUDES AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 11:40:35 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF ELECTRODE BELT HAS BEEN COMPLETED. THE REPORTED PROBLEM (WIRES EXPOSED) HAS BEEN CONFIRMED. UPON INVESTIGATION THE CABLE CONNECTING THE DISTRIBUTION NODE (DN) TO THE REAR THERAPY ELECTRODES WAS PULLED FROM THE STRAIN RELIEF. THE ROOT CAUSE FOR THE STRAINED CABLE COULD NOT BE POSITIVELY IDENTIFIED. THERE IS NO INDICATION THAT THE STRAINED CABLE CAUSED OR CONTRIBUTED TO AN ADVERSE EVENT AS THE DEVICE WAS ABLE TO DETECT AND TREAT VF RHYTHM ON THE DATE OF THE EVENT. DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767528&pc=NPT | 2025587-2025-02352 | 2024-03-21 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-03 04:00 | NPT | EVOLUT FX VALVE | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Aspiration/Inhalation; Bacterial Infection; Diarrhea; Inflammation; Pneumonia; Vomiting; Fungal Infection; Low Oxygen Saturation; Aspiration Pneumonitis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS THE IMPLANT OF A TRANSCATHETER VALVE, A STANFORD TYPE A AORTIC DISSECTION WAS OBSERVED, AND THE PATIENTS BLOOD PRESSURE WAS CONTROLLED WITHOUT SURGERY. FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DEVELOPED ACUTE CHOLECYSTITIS AND WAS TREATED WITH ANTIBIOTICS, FASTING, AND FLUID REPLACEMENT; HOWEVER, DIARRHEA PERSISTED. SUBSEQUENTLY, THE PATIENT TESTED POSITIVE FOR CLOSTRIDIUM DIFFICILE (C. DIFF), AND WAS TREATED WITH ORAL VANCOMYCIN. IT WAS REPORTED THAT YEAST-LIKE FUNGI WAS DETECTED ON THE TIP OF THE PATIENT'S CATHETER, AND THE PATIENT WAS TREATED WITH MICAFUNGIN. ADDITIONALLY, THE PATIENT PRESENTED WITH VOMITING THAT LED TO AN ABRUPT REDUCTION IN OXYGENATION, WHICH IMPROVED WITH SUCTIONING; HOWEVER, THE PATIENT DIED DUE TO ASPIRATION PNEUMONIA. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND VALVE OR IMPLANT PROCEDURE.Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS THE IMPLANT OF A TRANSCATHETER VALVE, A STANFORD TYPE A AORTIC DISSECTION WAS OBSERVED, AND THE PATIENTS BLOOD PRESSURE WAS CONTROLLED WITHOUT SURGERY. FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DEVELOPED ACUTE CHOLECYSTITIS AND WAS TREATED WITH ANTIBIOTICS, FASTING, AND FLUID REPLACEMENT; HOWEVER, DIARRHEA PERSISTED. SUBSEQUENTLY, THE PATIENT TESTED POSITIVE FOR CLOSTRIDIUM DIFFICILE (C. DIFF), AND WAS TREATED WITH ORAL VANCOMYCIN. IT WAS REPORTED THAT YEAST-LIKE FUNGI WAS DETECTED ON THE TIP OF THE PATIENT'S CATHETER, AND THE PATIENT WAS TREATED WITH MICAFUNGIN. ADDITIONALLY, THE PATIENT PRESENTED WITH VOMITING THAT LED TO AN ABRUPT REDUCTION IN OXYGENATION, WHICH IMPROVED WITH SUCTIONING; HOWEVER, THE PATIENT DIED DUE TO ASPIRATION PNEUMONIA. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND VALVE OR IMPLANT PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PATIENT ANATOMY INCLUDED TWO TORTUOUS SECTIONS ALONG THE ADVANCEMENT PATH FOR THE DELIVERY CATHETER SYSTEM (DCS): 1) FROM THE FEMORAL ARTERY TO THE ABDOMINAL AORTA (LEGS TO THE ABDOMEN), WHERE THE AORTA WAS IN THE SHAPE OF A QUESTION MARK, AND 2) FROM THE DESCENDING AORTA TO THE ASCENDING AORTA. PER THE PHYSICIAN, THE PATIENT'S ANATOMY WAS A CONTRIBUTING FACTOR, BUT THE DCS CAUSED THE AORTIC DISSECTION. INFORMATION WAS INITIALLY REPORTED, BUT NOT CAPTURED IN THE ABOVE EVENT NARRATIVE: THE PATIENT DIED APPROXIMATELY TWO AND A HALF MONTHS FOLLOWING THE VALVE IMPLANT PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY THE CAUSE OF THE AORTIC DISSECTION WAS THE DEVICE PREVIOUSLY REPORTED AS A SYSTEM REPORTED DEVICE. IT IS UNCLEAR WHETHER THE VALVE CONTRIBUTED TO THE ADVERSE EVENT; HOWEVER, IT WILL REMAIN REPORTED AS A POTENTIAL CONTRIBUTING FACTOR. UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE, SERIAL #, AND FULL UDI # D. THE SYSTEM REPORTED DEVICE: LOT #: 0011957865; MANUFACTURING DATE: 2023-09-18; USE BEFORE DATE: 2025-09-17; FULL UDI #:(B)(4), H4. DEVICE MFG DATE H6. THE VALVE DEVICE CODE WAS UPDATED BASED ON THE ADDITIONAL INFORMATION. THE RETURN STATUS OF THE DCS WAS CONFIRMED AS DISCARDED. CORRECTED DATA; THIS INFORMATION WAS AVAILABLE AT THE TIME OF INITIAL MEDWATCH SUBMISSION AND ERRONEOUSLY OMITTED. B5. INFORMATION WAS ADDED TO THE SECOND PARAGRAPH E. FACILITY STREET D10. H6. PATIENT CODES ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329 PRODUCT ID: D-EVOLUTFX-2329 SERIAL/LOT: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767557&pc=NGV | 2124215-2025-20717 | 2025-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-03 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT EMBOLIZATION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. THE LAA HAD A BROCCOLI-LIKE SHAPE WITH PROXIMAL PECTINATE AND A SHARP BEND IN THE ANTERIOR/SUPERIOR DIRECTION. A WATCHMAN DOUBLE CURVE ACCESS SYSTEM (WAS) AND 31MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED FOR USE. ACCESS WAS SUCCESSFULLY OBTAINED VIA THE LEFT FEMORAL VEIN (LFV), AND THE SEPTUM WAS CROSSED WITHOUT COMPLICATIONS AT AN INFERIOR/MID LOCATION. A CONTRAST INJECTION OF THE LAA WAS PERFORMED USING A 6F PIGTAIL CATHETER. THE PHYSICIAN PROPERLY FLUSHED AND PREPARED THE 31MM WDS BEFORE DEPLOYMENT. AFTER PERFORMING THE TUG TEST, THE PHYSICIAN BEGAN RELEASING THE CLOSURE DEVICE PRIOR TO FINAL POSITION, ANCHOR, SIZE AND SEAL (PASS) CRITERIA ASSESSMENT. HOWEVER, A WIRE BIAS WAS DETECTED, AND THE CLOSURE DEVICE STARTED TO BECOME DISLODGED FROM THE APPENDAGE PRIOR TO BEING FULLY RELEASED, EMBOLIZING INTO THE LEFT VENTRICLE OUTFLOW TRACT (LVOT). ATTEMPTS TO SNARE THE DEVICE WERE UNSUCCESSFUL. THE PATIENT WAS THEN TRANSFERRED TO THE CARDIOVASCULAR (CV) OPERATING ROOM (OR) FOR SURGICAL RETRIEVAL OF THE DEVICE. FOURTEEN (14) HOURS AFTER THE SURGERY, THE PATIENT COULD NOT BE WEANED OFF THE PUMP AND SUBSEQUENTLY PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767626&pc=NIK | 2124215-2025-20758 | 2025-03-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-03 04:00 | NIK | MOMENTUM X4 CRT-D | High impedance; Incorrect, Inadequate or Imprecise Result or Readings; Failure to Convert Rhythm; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P010012/S436 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) RECEIVED MULTIPLE INEFFECTIVE SHOCKS FOR AN ARRHYTHMIC STORM REQUIRING EXTERNAL DEFIBRILLATION. THE PATIENT WAS SUBSEQUENTLY HOSPITALIZED AND INTUBATED IN THE INTENSIVE CARE UNIT. UPON INTERROGATION, IT WAS NOTED THAT A CODE 1005, INDICATIVE OF AN OPEN CIRCUIT CONDITION. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) WAS CONTACTED AND CONFIRMED THAT THE CODE 1005 WAS DECLARED WHEN A DELIVERED SHOCK PRODUCED AN ELEVATED, OUT-OF-RANGE IMPEDANCE OF GREATER THAN 145 OHMS. TS STRONGLY RECOMMENDED REPLACING THE RIGHT VENTRICULAR (RV) LEAD TO ENSURE ADEQUATE THERAPY DELIVERY AND CONVERSION. HOWEVER, A FEW DAYS LATER, THE PATIENT UNFORTUNATELY PASSED AWAY DUE TO CEREBRAL INJURY. AT THIS TIME, THIS CRT-D REMAINS IMPLANTED. THE RV LEAD IS NOT A BOSTON SCIENTIFIC PRODUCT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21767816&pc=MKJ | 3030677-2025-001085 | 2025-04-02 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-04-03 04:00 | MKJ | HS1 (W B) DEFIB SIMP. CHINESE EXCHG | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P160029 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT DURING A PATIENT USE CASE, THE DEVICE FAILED TO DELIVERY THERAPY. THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: SERIAL NUMBER UPDATED. Manufacturer Narrative: UPDATED CONCLUSION CODE GRID.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21749314&pc=NAY | 2955842-2025-12753 | 2025-01-11 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-02 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem |  | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED PARTIAL NEPHRECTOMY PROCEDURE, THE PATIENT EXPIRED. THERE WAS A 1.3 HOUR DELAY TO THE PROCEDURE START DUE TO DIFFICULTIES PLACING A URETHRAL CATHETER DUE TO UNFORESEEN ANATOMICAL DEVIATIONS FROM PREVIOUS PROCEDURES. INTRAOPERATIVELY, FURTHER COMPLICATIONS WERE ENCOUNTERED WITH OPTIMAL DOCKING DUE TO THE PATIENT'S HEIGHT AND MAINTAINING THE PATIENT¿S OXYGEN SATURATION; THE PROCEDURE WAS SUBSEQUENTLY COMPLETED WITH NO ADDITIONAL ISSUES. APPROXIMATELY 1.5 HOURS POST-OPERATIVELY, THE PATIENT EXHIBITED SYMPTOMS OF TACHYCARDIA AND ABDOMINAL RIGIDITY, INDICATING A POTENTIAL HEMORRHAGE. THE PATIENT WAS RETURNED TO THE OPERATING ROOM, AND AN EXPLORATORY LAPAROTOMY WAS PERFORMED, REVEALING NO HEMORRHAGE. IT WAS DETERMINED BY THE PHYSICIANS THAT THE SYMPTOMS WERE REACTIONS TO THE ANESTHETIC DRUGS. THE PATIENT EXPIRED LATER THE SAME DAY AS THE PROCEDURE.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS. ALL MULTI-USE INSTRUMENTS USED IN THE CASE HAVE BEEN USED IN SUBSEQUENT PROCEDURES. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED SHORTLY FOLLOWING A ROBOTIC PROCEDURE THAT WAS UNEVENTFUL INTRAOPERATIVELY. THE PATIENT WAS EMERGENTLY RE-EXPLORED AND NO SURGICAL OR OTHER INTRAABDOMINAL ISSUES WERE NOTED. ALTHOUGH OTHER POSSIBILITIES EXIST, AN ANESTHETIC COMPLICATION WAS SUSPECTED AS THE CLINICAL SCENARIO IS CONSISTENT WITH MALIGNANT HYPERTHERMIA. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. . Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21750143&pc=DTB | 2649622-2025-08853 | 2023-10-21 04:00 | Death | MPRI | 2025-04-02 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED, AND THE PATIENT'S REPRESENTATIVE REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM NEVER WORKED. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE..MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21750145&pc=NVZ | 3008973940-2025-02334 | 2023-10-21 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-02 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED, AND THE PATIENT'S REPRESENTATIVE REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM NEVER WORKED. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21750147&pc=DTB | 2649622-2025-08855 | 2023-10-21 04:00 | Death | MPRI | 2025-04-02 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED, AND THE PATIENT'S REPRESENTATIVE REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM NEVER WORKED. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21750159&pc=DTQ | 8010762-2025-0000151 | 2024-12-02 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-02 04:00 | DTQ | CARDIOHELP | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding | K133598 |  | 1 | Event Description: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM (B)(6) ON 2025-03-31: THE VIGILANCE AREA OF MEDICAL DEVICES OF (B)(6) HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON (B)(6) 2024. THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" AND THE CONSEQUENCE FOR THE PATIENT WAS "HEMORRHAGE AFTER REMOVING THE DEVICE AT THE INSERTION SITE. THIS RESULTED IN THE PATIENT'S DEATH. COMPLAINT ID#: (B)(4).Event Description: THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE. COMPLAINT# (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE SPAIN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER: 701048012. Manufacturer Narrative: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-03-31: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2024-12-02 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" AND THE CONSEQUENCE FOR THE PATIENT WAS "HEMORRHAGE AFTER REMOVING THE DEVICE AT THE INSERTION SITE. THIS RESULTED IN THE PATIENT'S DEATH. RENACER REGISTRY PATIENT (B)(6).¿ NEW INFORMATION WAS RECEIVED ON 2025-04-02 THAT NO MALFUNCTION ON THE CARDIOHELP DEVICE OCCURRED. THE EVENT WAS NOT RELATED TO THE CARDIOHELP DEVICE. THE TREATMENT WITH THE CARDIOHELP MACHINE WAS FINALIZED ON 2024-11-18. ON (B)(6) 2024, AND THE PATIENT SUFFERED A TEAR OF THE SUPERFICIAL FEMORAL ARTERY WITH MASSIVE HEMORRHAGE AND WAS REVIEWED BY THE SURGICAL TEAM. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" COULD NOT BE LINKED TO A DEVICE DEFECT/FAILURE. THEREFORE THE COMPLAINT COULD NOT BE CONFIRMED. THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21750369&pc=FMI | 9616791-2025-00001 | 2025-03-18 04:00 | Death | WAISMED LTD | 2025-04-02 04:00 | FMI | NIO ADULT | Firing Problem | Insufficient Information | K142086 |  | 1 | Event Description: COMPLAINT DESCRIPTION FROM CUSTOMER ON (B)(6) 2025: "PATIENT WAS IN CARDIAC ARREST WITH CPR BEING PERFORMED. CREW ARRIVED, TOOK OVER CPR. IV WAS ATTEMPTED AND FAILED. NIO WAS BROUGHT OUT, SITE LOCATED AND WHEN WE SPUN THE TOP ON THE NIO TO UNLOCK IT, IT WAS TIGHT TO TURN. PRESSURE PLACED WITH DOMINANT HAND, OTHER HAND STABILIZING NIO. NIO DID NOT DEPLOY, SECOND CREWMEMBER ATTEMPTED TO DEPLOY IT AND IT DID NOT WORK. SECOND NIO WAS BROUGHT OUT AND THE SAME THING HAPPENED. TOP WAS TIGHT TO SPIN AND IT DID NOT DEPLOY FOR EITHER THE RN OR EMT. MANUAL IO WAS PLACED IN THIS PATIENT FIRST ATTEMPT. CPR CONTINUED. PATIENT DIED. ONE OF THE NIO THAT DID NOT DEPLOY WAS BROUGHT BACK TO THE STATION AND IS CURRENTLY BEING SENT BACK WITH THE PROVIDED SHIPPING LABEL.Event Description: ADDITIONAL INFORMATION FROM CUSTOMER RECEIVED ON MARCH 30, 2025: "PATIENT WAS IN CARDIAC ARREST WITH CPR BEING PERFORMED. CREW ARRIVED, TOOK OVER CPR. IV WAS ATTEMPTED AND FAILED. NIO WAS BROUGHT OUT, SITE LOCATED AND WHEN WE SPUN THE TOP ON THE NIO TO UNLOCK IT, IT WAS TIGHT TO TURN. PRESSURE PLACED WITH DOMINANT HAND, OTHER HAND STABILIZING NIO. NIO DID NOT DEPLOY, SECOND CREWMEMBER ATTEMPTED TO DEPLOY IT AND IT DID NOT WORK. SECOND NIO WAS BROUGHT OUT AND THE SAME THING HAPPENED. TOP WAS TIGHT TO SPIN AND IT DID NOT DEPLOY FOR EITHER THE RN OR AEMT. MANUAL IO WAS PLACED IN THIS PATIENT FIRST ATTEMPT. CPR CONTINUED. PATIENT DIED. ONE OF THE NIO THAT DID NOT DEPLOY WAS BROUGHT BACK TO THE STATION AND IS CURRENTLY BEING SENT BACK WITH THE PROVIDED SHIPPING LABEL." Manufacturer Narrative: PLEASE NOTE THAT IN THE INITIAL REPORT, THE BATCH NUMBER PROVIDED BY THE CUSTOMER WAS INCORRECT (2440043). DURING THE INVESTIGATION OF THE DEVICE, IT WAS CLARIFIED THAT THE CORRECT BATCH NUMBER IS: 2440040. AS PART OF THE INVESTIGATION PROCESS, THE FOLLOWING EXAMINATIONS WERE PERFORMED: VISUAL INSPECTION: THE DEVICE WAS VISUALLY EXAMINED, AND IT APPEARED THAT THE TRIGGER WAS UNCLIPPED. X-RAY EXAMINATION ((B)(6) 2025): THE DEVICE WAS X-RAYED, AND ALL METAL COMPONENTS WERE FOUND TO BE IN THEIR EXPECTED POSITIONS. THERE WAS NOTHING IN THE X-RAY THAT WOULD CLEARLY INDICATE A CAUSE FOR THE DEVICE'S FAILURE TO ACTIVATE. POST-X-RAY FUNCTIONAL TEST (04/11/2025): THE DEVICE WAS TESTED TO DETERMINE IF IT COULD BE ACTIVATED. IT WAS NOTED THAT THE SAFETY CAP WAS SIGNIFICANTLY MORE DIFFICULT TO ROTATE THAN EXPECTED. THE DEVICE WAS PLACED ON A TORQUE TABLE TO MEASURE THE FORCE REQUIRED TO ROTATE THE SAFETY CAP FROM THE "LOCKED" TO "UNLOCKED" POSITION. THE TORQUE MEASUREMENTS RANGED FROM 500 N·MM TO OVER 900 N·MM, MORE THAN TWICE THE VALUES MEASURED ON A DEVICE WHERE THE SAFETY CAP ROTATED AS INTENDED. HOWEVER, THE INCREASED TORQUE REQUIRED TO TURN THE SAFETY CAP WOULD NOT IMPACT THE USER'S ABILITY TO FIRE THE DEVICE PROVIDED THEY COULD MOVE IT FROM THE LOCKED TO UNLOCKED POSITION. POST-ACTIVATION INSPECTION: AFTER ACTIVATING THE DEVICE, THE SAFETY CAP WAS REMOVED AND BOTH THE CAP AND BODY WERE EXAMINED. ADDITIONAL SCRATCHES WERE OBSERVED ON THE BODY OF THE DEVICE. THE MOST LIKELY CAUSE WAS DETERMINED TO BE EXCESS FLASH (MATERIAL LEFT ON THE COMPONENT AFTER MANUFACTURING) NEAR THE PRONG AREA OF THE SAFETY CAP. ROOT CAUSE ANALYSIS: A FISHBONE DIAGRAM WAS COMPLETED TO ANALYZE THE FAILURE MODE. THE MOST LIKELY ROOT CAUSE OF THE INCREASED TORQUE REQUIRED TO TURN THE SAFETY CAP IS FLASH FROM THE MANUFACTURING PROCESS. HOWEVER, THIS IS NOT LIKELY TO HAVE LED TO THE DEVICE NOT FIRING SINCE THE SAFETY CAP COULD BE ROTATED TO THE UNLOCKED POSITION. THE MOST LIKELY ROOT CAUSE FOR THE DEVICE NOT FIRING IS EITHER DUE TO THE USER NOT FULLY PRESSING DOWN ON THE DEVICE OR NOT PRESSING STRAIGHT DOWN ON THE DEVICE. IF THE USER PRESSED DOWN AT AN ANGLE AND THEN PULLED THE TRIGGER BACK WITH EXCESSIVE FORCE WITHOUT FIRST UNLOCKING THE TRIGGER, IT MAY HAVE CAUSED DEFORMATION. CONCLUSIONS: THE CUSTOMER COMPLAINT REGARDING THE DIFFICULTY IN ROTATING THE TOP OF THE DEVICE WAS REPLICATED. THE ISSUE WAS MOST LIKELY CAUSED BY FLASH (MATERIAL LEFT ON THE COMPONENT AFTER MANUFACTURING) IN THE SAFETY CAP, OR DEBRIS (BROKEN FLASH) LOCATED BETWEEN THE SAFETY CAP AND THE DEVICE BODY. THE COMPLAINT REGARDING DIFFICULTY IN FIRING THE DEVICE WAS ALSO REPLICATED. THIS WAS MOST LIKELY CAUSED BY NOT FOLLOWING THE INSTRUCTIONS AS OUTLINED IN THE IFU, AND BY SQUEEZING THE TRIGGER WITH EXCESSIVE FORCE EITHER PRIOR TO FULLY UNLOCKING THE TRIGGER, OR BY ONLY PARTIALLY UNLOCKING THE TRIGGER BY PRESSING DOWN AT AN ANGLE. *NOTE- UPDATED B2 TO INCLUDE "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS". UPDATED B7 TO INCLUDE RELEVANT INFORMATION. Manufacturer Narrative: THERE WERE TWO MEDICICAL DEVICES INVOLVED IN THIS REPORT. ONE MEDICAL DEVICE ONLY REMAIN IN USER HANDS, WHICH HAS BEEN REQUESTED TO BE RETURNED TO MANUFACTURER SITE TO PERFORM AN ENGINERING INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21750941&pc=OZP | 2032227-2025-147048 | 2025-03-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-02 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON AN UNKNOWN DATE OF DEATH. THE CAUSE OF DEATH WAS UNKNOWN. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1780KL, UNOMEDICAL, MMT-332A. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751102&pc=OZD | 1220648-2025-27175 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-04-02 04:00 | OZD | IMPELLA CP | Malposition of Device; Deformation Due to Compressive Stress |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING REPOSITIONING OF THE IMPELLA CP DEVICE, THE CANNULA KINKED, AND THE PUMP COULD NO LONGER BE ADJUSTED. THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB WHERE IT WAS DISCOVERED THAT THE INLET WAS STUCK IN THE MITRAL VALVE. ATTEMPTS TO STRAIGHTEN THE IMPELLA PUMP WITH A LASOO CATHETER WERE UNSUCCESSFUL AND SUBSEQUENTLY PULLING ON THE PUMP RESULTED IN A PAPILLARY MUSCLE INJURY. FOLLOWING A CARDIAC COMPUTED TOMOGRAPHY, THE PATIENT WAS DEEMED TO BE BRAIN DEAD, EXPIRED ON SUPPORT. THE IMPELLA CP DEVICE REMAINED IN THE AORTA AND WAS REMOVED POST-MORTEM. Manufacturer Narrative: PATIENT-SPECIFIC INFORMATION IS UNAVAILABLE AT THE TIME OF THIS MDR WRITING; THEREFORE, RESPECTIVE FIELDS REFLECT "UNKNOWN" OR LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF PRODUCT DAMAGE (CANNULA KINK), POSITIONING ISSUES, AND VENTRICLE PERFORATION HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INVESTIGATION. THE RETURNED PUMP HAD SIGNIFICANT CANNULA KINKS AND A PIGTAIL KINK. NO OTHER PHYSICAL DAMAGE WAS NOTED ON THE RETURNED PRODUCT. THE PUMP WAS ABLE TO RUN IN A BUCKET OF WATER. A DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. ACCORDING TO THE GIVEN CLINICAL DETAILS, THE IMPELLA WAS REPOSITIONED FOR UNKNOWN REASON, AND IT WAS PUSHED IN DEEP IN LEFT VENTRICLE (LV), IT CAUSED THE CANNULA TO KINK. THE PUMP'S INLET VALVE WAS STUCK AT THE MITRAL VALVE, AND PULLING THE PUMP CAUSED AN INJURY TO THE PAPILLARY MUSCLE. THE CAUSE OF THE CANNULA KINK ISSUE WAS USE-RELATED, SINCE THE CANNULA WAS PUSHED DURING REPOSITIONING. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THE CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE, BASED ON GIVEN CLINICAL INFORMATION. D9 DATE DEVICE RETURNED TO MANUFACTURER ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751148&pc=MIH | 3007284313-2025-03874 | 2025-03-07 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-02 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Material Deformation | Ischemia; Obstruction/Occlusion | P020004 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT A GORE® EXCLUDER® AAA ENDOPROSTHESIS AND A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS WAS IMPLANTED ON (B)(6) 2025 DURING AN ENDOVASCULAR PROCEDURE ON A PATIENT WITH AN ABDOMINAL AORTIC ANEURYSM ASSOCIATED WITH A RIGHT COMMON ILIAC ANEURYSM. DURING THE PROCEDURE THE FINAL ANGIOGRAPHY SHOWED A PERFECT RESULT WITH NO DELAY OF THE CONTRAST PRODUCT AND GOOD ARTERIAL PULSES ON BOTH AXES. 9 DAYS POST-IMPLANT, ON (B)(6) 2025, THE PATIENT WAS ADMITTED IN EMERGENCY FOR ACUTE ISCHEMIA. THE CT SCAN SHOWS THAT THE GORE® EXCLUDER® AAA ENDOPROSTHESIS HAS COMPLETELY COLLAPSED AND THE CONTRAST PASSING OUTSIDE THE ENDOPROSTHESIS. THE PATIENT UNDERWENT A SURGERY FOR AN AXILLARY BIFEMORAL BYPASS. THE PATIENT PAST AWAY 2 DAYS LATER FROM ISCHEMIA SYNDROMES. Manufacturer Narrative: CORRECTED H6: ADDED TYPE OF INVESTIGATION CODE B20. CODE B13 AND B17 WAS INADVERTENTLY ENTERED AND SHOULD BE REMOVED. UPDATED INVESTIGATION FINDINGS CODE TO C24, CODE C21 IS NO LONGER APPLICABLE. UPDATED INVESTIGATION CONCLUSIONS CODE TO D15, CODE D16 IS NO LONGER APPLICABLE. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. AS PART OF OUR ROUTINE QUALITY PROCEDURES, EACH BATCH OF DEVICES UNDERGOES COMPREHENSIVE QUALITY CONTROL TESTING AND INSPECTIONS PRIOR TO RELEASE FOR DISTRIBUTION. GORE REVIEWED THE MANUFACTURING RECORDS ASSOCIATED WITH THE REPORTED LOT NUMBER AND VERIFIED THAT ALL RELEASE CRITERIA HAD BEEN MET. THE PRIMARY REPORTED DEVICE FAILURE MODE, COLLAPSE OR COMPRESSION OF THE EXCLUDER C3 DEVICE IPSILATERAL LEG WITHIN NINE DAYS OF INITIAL IMPLANT, IS CONSISTENT WITH EVALUATION OF THE PROVIDED CLINICAL IMAGES. THE IMAGING EVALUATION SHOWS COMPRESSION AND OCCLUSION OF THE ENDOPROSTHESIS, BUT THE IMAGING IS UNABLE TO CONFIRM THE REPORT OF CONTRAST OUTSIDE THE ENDOPROSTHESIS. THE IMAGING EVALUATION INDICATES THE PATIENT¿S AORTA IS TORTUOUS NEAR THE RENAL ARTERIES, BUT THE SPECIFIC CAUSE FOR THE DEVICE COLLAPSE COULD NOT BE ESTABLISHED WITH THE AVAILABLE INFORMATION. BASED ON THE INCIDENT DESCRIPTION AND THE SUBSEQUENT INVESTIGATION, NO FURTHER INFORMATION WAS PROVIDED TO GORE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THIS INCIDENT AND ASSIGN A ROOT CAUSE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751291&pc=OZD | 1220648-2025-27178 | 2025-03-07 05:00 | Death | ABIOMED, INC. | 2025-04-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Bradycardia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT UNDERGOING HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE SUPPORT WAS SUCCESSFUL FOR UNDER THREE HOURS, AND THEN THE IMPELLA CP DEVICE WAS EXPLANTED. AT THE TIME OF EXPLANT THERE WAS AN OOZING AT THE ACCESS SITE, WHICH WAS TREATED BY CONTRALATERAL LEG BALLOONING AND EVENTUAL VASCULAR SURGERY. THE SITE ACHIEVED HEMOSTASIS, BUT IN LATER HOURS THE PATIENT EXPERIENCE BRADYCARDIA AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751376&pc=OZD | 1220648-2025-27179 | 2025-03-09 05:00 | Death | ABIOMED, INC. | 2025-04-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS NOTED THAT PATIENT WAS ALREADY IN RESUSCITATION BEFORE IMPELLA IMPLANTATION. AFTER SUCCESSFULLY IMPLANTATION AND PERFORMED PERCUTANEOUS CORONARY INTERVENTION AND UPON TRANSFER OF THE PATIENT TO THE INTENSIVE CARE UNIT, THE PATIENT HAD TO BE RESUSCITATED AGAIN FOR APPROXIMATELY ONE HOUR. DUE TO A HEMATOMA IN THE GROIN/PUNCTURE SITE, THE PUMP WAS REMOVED. IT WAS NOTED THAT THE PATIENT WAS ON A HIGH DOSE OF HEPARIN AND HAD MULTIPLE ORGAN FAILURE AND THIS WAS THE SUSPECTED CAUSE OF THE HEMATOMA. SURVIVAL OUTCOME AT EXPLANT NOTED THAT THE PATIENT EXPIRED A FEW HOURS AFTER EXPLANT AND CAUSE OF DEATH WAS SUSPECTED AS MULTIPLE ORGAN FAILURE. HOWEVER, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE.. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751564&pc=DTZ | 8010762-2025-0000154 | 2025-03-21 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-02 04:00 | DTZ | TUBING SET | Fluid/Blood Leak; Gas/Air Leak | Cardiac Arrest | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN FRANCE. IT WAS REPORTED THAT THE CIRCUIT WAS PRIMED WITH WATER IN ACCORDANCE TO STANDARD PROCEDURES TO PREPARE FOR ECMO SUPPORT. THE FLOW RATE WAS OBTAINED WITHOUT DIFFICULTY. A FEW MINUTES LATER THERE WAS A LOSS OF FLOW, A MACHINE ALARM AND VISIBLE BUBBLES WITHIN THE OXYGENATOR. THE SYSTEM COULD NOT BE DE-AIRED. THE CIRCUIT WAS CLAMPED AND THE CIRCUIT REPLACED. THE PATIENT WAS ADMITTED FOR EMERGENCY CARDIAC ARREST. THE CARE WAS DELAYED FOR THIS PATIENT. THE ECMO WAS INITIATED AT 8 PM AND THE PATIENT PASSED AWAY AT 3 AM. COMPLAINT ID# (B)(4). Event Description: COMPLAINT ID: (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE FRANCE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET" WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: THE EVENT OCCURRED IN FRANCE DURING TREATMENT. IT WAS REPORTED THAT THE CIRCUIT WAS PRIMED WITH WATER IN ACCORDANCE WITH STANDARD PROCEDURES TO PREPARE FOR ECMO SUPPORT. THE FLOW RATE WAS OBTAINED WITHOUT DIFFICULTY. A FEW MINUTES LATER THERE WAS A LOSS OF FLOW, A MACHINE ALARM AND VISIBLE BUBBLES WITHIN THE OXYGENATOR. THE SYSTEM COULD NOT BE DE-AIRED. THE CIRCUIT WAS CLAMPED AND THE CIRCUIT REPLACED. THE PATIENT WAS ADMITTED FOR EMERGENCY CARDIAC ARREST. THE CARE WAS DELAYED FOR THIS PATIENT. THE ECMO WAS INITIATED AT 8 PM ON (B)(6) 2025 AND THE PATIENT PASSED AWAY AT 3 AM ON (B)(6) 2025. NEW INFORMATION WAS RECEIVED ON 2025-04-03 THAT THE PATIENT WAS CONNECTED TO THE FIRST HLS SET WHICH CONTAINED AIR IN THE SYSTEM. LATER, IT WAS CONFIRMED THAT THE REPLACEMENT SET HAD NO FAULT AND WAS DISCARDED BY THE CUSTOMER. THE FIRST SET WAS AVAILABLE FOR INVESTIGATION, WHICH ACCORDING TO THE CUSTOMER HAD NO VISIBLE DEFECT NOR ANY LEAKS DURING/AFTER PRIMING. FURTHER INFORMATION WAS RECEIVED THAT THERE WAS NO AIR ENTRY AFTER THE REPLACEMENT OF THE AFFECTED SET. THE USED CANNULAS WERE NOT REPLACED AND WERE THEREFORE WORKING AS EXPECTED. NO INFORMATION WAS AVAILABLE IF THE USED CANNULAS WERE GETINGE CANNULAS. THE CANNULAS WERE DISCARDED BY THE CUSTOMER. AS PATIENT DEATH WAS REPORTED, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS INVESTIGATED BY THE GETINGE LABORATORY ON 2025-06-27 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. THERE WERE NO VISIBLE DAMAGES ON THE PRODUCT. THE PRODUCT COULD BE PRIMED WITHOUT ANY PROBLEMS AND THERE WAS NO AIR ENTRY. THE PRODUCT FUNCTIONALITY WAS AS EXPECTED. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-07-04 WITH FOLLOWING CONCLUSION: "ON (B)(6) 2025, AN INCIDENT OCCURRED INVOLVING AN HLS SET ADVANCED 7.0 USED WITH THE CARDIOHELP-I SYSTEM DURING VA-ECLS. THE PATIENT, A 45-YEAR-OLD FEMALE ADMITTED AFTER CARDIAC ARREST, RECEIVED ECLS SUPPORT BEGINNING AT APPROXIMATELY 20:15. WHILE INITIAL STARTUP WAS UNEVENTFUL AND TARGET FLOW WAS ACHIEVED, A SUDDEN LOSS OF FLOW OCCURRED SHORTLY THEREAFTER, ACCOMPANIED BY BUBBLE DETECTION ALARMS. AIR BUBBLES WERE OBSERVED IN THE OXYGENATOR AND COULD NOT BE REMOVED. NO AIR ENTRY SOURCE WAS IDENTIFIED AT THE TIME. THE CLINICAL TEAM REPLACED THE CIRCUIT, WHICH FUNCTIONED NORMALLY. THE PATIENT EXPIRED AT AROUND 03:00; NO DIRECT CAUSAL LINK TO THE INCIDENT WAS ESTABLISHED, THOUGH TREATMENT WAS DELAYED. THE SUBSEQUENT INVESTIGATION FOUND NO PRODUCT-RELATED DEFECTS. A MINOR LEAK AT THE VENOUS SAMPLING PORT WAS ATTRIBUTED TO DRIED BLOOD RESIDUE AND WAS NO LONGER PRESENT AFTER CLEANING. THE TUBING, INCLUDING THE CRITICAL VENOUS LINE, SHOWED NO DAMAGE OR KINKS. FUNCTIONAL TESTING, INCLUDING A 60-MINUTE LEAK TEST, CONFIRMED SYSTEM INTEGRITY. A PRESSURE SENSOR MALFUNCTION WAS NOTED BUT DETERMINED TO BE RELATED TO TRANSPORT AND NOT CONNECTED TO THE COMPLAINT. IMPORTANTLY, THE VENOUS CANNULA USED DURING ECMO INITIATION - A CRITICAL COMPONENT IN ASSESSING THE ROOT CAUSE OF AIR INGRESS - WAS NOT RETURNED FOR FURTHER EVALUATION. THIS OMISSION LIMITS THE ABILITY TO FULLY RULE OUT A CONNECTION-RELATED ISSUE OR A DEFECT AT THE INSERTION SITE. LITERATURE SUPPORTS THE RISK OF AIR INGRESS VIA THE VENOUS LINE OR CANNULA DURING ECMO, PARTICULARLY IN THE PRESENCE OF LOOSE CONNECTIONS, NEGATIVE PRESSURE, OR INSUFFICIENT PRIMING. IN ADDITION, AIR INGRESS THROUGH A CENTRAL VENOUS CATHETER, IF PRESENT, COULD BE A CONTRIBUTING FACTOR, PARTICULARLY IF CONNECTIONS WERE NOT FULLY SECURED. THE ASSUMPTION OF VENOUS AIR ENTRAINMENT APPEARS TO CONFORM TO THE EVENT PATTERN AS DESCRIBED: LOSS OF FLOW (DUE TO AIR INVASION OF THE CENTRIFUGAL PUMP), AIR IN THE OXYGENATOR, THEN BUBBLE ALARM. IT IS PRESUMED THAT THE REPORTED BUBBLE ALARM WAS ARTERIAL. NO MENTION OF VENOUS BUBBLE SENSOR DEPLOYMENT, HOWEVER, WAS CITED IN THE COMPLAINT NARRATIVE. IN CONCLUSION AND GIVEN THE ABSENCE OF THE VENOUS CANNULA FOR INVESTIGATION AND THE LIMITED PROCEDURAL DETAILS AVAILABLE, NO DEFINITIVE ROOT CAUSE OF THE AIR INGRESS CAN BE CONFIRMED. HOWEVER, POSSIBLE CAUSES INCLUDE A LEAK AT THE VENOUS CANNULA SITE, AIR ENTRY VIA A CENTRAL VENOUS CATHETER, OR A TRANSIENT LEAK AT THE VENOUS SAMPLING PORT, OBSERVED DURING INITIAL FLUSHING AND ATTRIBUTED TO DRIED BLOOD RESIDUES, THOUGH THIS WAS NOT REPRODUCIBLE UNDER CONTROLLED TESTING CONDITIONS." ACCORDING TO THE INSTRUCTION OF USE OF THE HLS SET IT IS MENTIONED THAT INCORRECT POSITIONING OF THE CARDIOHELP CAN CAUSE AIR PERMEATION ON THE BLOOD SIDE OF THE HLS MODULE ADVANCED. THE CARDIOHELP MUST BE POSITIONED AT A LEVER LOWER THAT THE PATIENT AND SECURED CLOSE TO THE PATIENT. THE FLOW/BUBBLE SENSOR MUST ALWAYS BE AFFIXED ON THE ARTERIAL SIDE OF THE SET IF YOU ARE USING THE CARDIOHELP-I. WITH AN ACTIVATED INTERVENTION, THE DETECTION OF BUBBLES BY THE FLOW/BUBBLE SENSOR TRIGGERS A PUMP STOP. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-07-03. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "AIR IN SYSTEM" COULD NOT BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE HLS SET. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751595&pc=FKX | 0008030665-2025-00741 | 2025-03-07 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-02 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Malaise | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). IT WAS CONFIRMED THE PATIENT WAS IN ACTIVE TREATMENT UTILIZING THE LIBERTY SELECT CYCLER AT THE TIME OF THE EVENT. IT IS UNKNOWN WHAT PHASE AND CYCLE THE PATIENT WAS IN. THE PATIENT CALLED THEIR SPOUSE INTO THE BEDROOM AND STATED THEY WERE NOT FEELING WELL. THE PATIENT WAS SITTING ON THE EDGE OF THE BED AND BECAME QUIET. THE PATIENT THEN BECAME UNRESPONSIVE. 911 WAS CALLED. THE 911 OPERATOR INSTRUCTED THE SPOUSE TO PLACE THE PATIENT ON THE FLOOR. DESPITE THE PATIENT BEING A DO NOT RESUSCITATE (DNR) ORDER, THE SPOUSE BEGAN CARDIOPULMONARY RESUSCITATION (CPR). EMERGENCY MEDICAL SERVICES (EMS) ARRIVED AND CONTINUED TO PERFORM RESUSCITATIVE EFFORTS. PER THE PATIENT¿S SPOUSE, THEY ADMINISTERED CPR AND MEDICATIONS (UNSPECIFIED). RESUSCITATION EFFORTS WERE UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED AT HOME. NO AUTOPSY WAS COMPLETED. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S CAUSE OF DEATH IS REPORTEDLY CARDIAC ARRHYTHMIA. THE PATIENT DID HAVE A HISTORY OF UNSPECIFIED CARDIAC ARRHYTHMIA. THERE WAS NO REPORTED ISSUE WITH THE PATIENT¿S TREATMENT, OR THE LIBERTY SELECT CYCLER IN RELATION TO THE EVENT. THE LAST REPORTED ISSUES WERE DRAIN COMPLICATIONS WHICH WERE SUBSEQUENTLY FOUND TO BE CAUSED BY SEVERE CONSTIPATION AS A RESULT OF THE PATIENT BEING BEDRIDDEN. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE REPORTED CAUSE OF DEATH WAS CARDIAC ARRHYTHMIA. CARDIAC DISEASE IS THE MAJOR CAUSE OF DEATH, ACCOUNTING FOR APPROXIMATELY 37 PERCENT OF ALL-CAUSE MORTALITY IN PATIENTS RECEIVING EITHER HEMODIALYSIS OR PERITONEAL DIALYSIS WITH THE LARGEST SPECIFIC CAUSE ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA). THE PATIENT REPORTEDLY HAD A HISTORY OF UNSPECIFIED CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751658&pc=PNJ | 9612164-2025-01677 | 2025-03-13 04:00 | Death | MEDTRONIC IRELAND | 2025-04-02 04:00 | PNJ | MICRA¿ AV2 | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Asystole | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE LEADLESS IPG EXHIBITED NON CAPTURE. DURING LEADLESS IPG PLACEMENT AND REPOSITIONING EXTERNAL PACING WAS REQUIRED. THE PATIENT RHYTHM RESUMED IN PULSELESS ELECTRICAL ACTIVITY. THE PATIENT EXPERIENCED ASYSTOLE AND CARDIAC ARREST. THE PATIENT REMAINED UNCONSCIOUS. A CODE WAS CALLED AND THE PATIENT NEVER REVIVED. IT WAS BELIEVED HIS HEART MUSCLE WAS NOT VIABLE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: [MICRA], PRODUCT ID: [MC2AVR1-DELSYS], (SERIAL: (B)(6), D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21751939&pc=OZD | 1220648-2025-27187 | 2025-03-07 05:00 | Death | ABIOMED, INC. | 2025-04-02 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Cardiac Arrest; Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED AN ACCESS SITE HEMATOMA, POSITIONING ISSUE AND LOW FLOWS REQUIRING INTERVENTION. HOWEVER, THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT UNDERWENT A CATHETERIZATION WHICH SHOWED 100% OCCLUDED LEFT ANTERIOR DESCENDING ARTERY, 95% OCCLUDED CIRCUMFLEX AND RIGHT CORONARY ARTERIES WITH AN EJECTION FRACTION OF 35%. DURING THE PERCUTANEOUS CORONARY INTERVENTION, BALLOONING THE RIGHT CORONARY ARTERY, THE PATIENT HAD A SEIZURE, AND THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE. APPROXIMATELY FIVE MINUTES AFTER THE IMPELLA CP DEVICE IMPLANT, A "LARGE" HEMATOMA DEVELOPED AT THE ACCESS SITE. MANUAL PRESSURE WAS APPLIED; HOWEVER, THE PATIENT'S BLOOD PRESSURE BEGAN TO RAPIDLY DROP. LOW PUMP FLOWS WERE THEN ENCOUNTERED, AND REPOSITIONING PROVED TO BE DIFFICULT AS THE LEFT VENTRICLE REPEATEDLY PUSHED THE DEVICE OUT OF THE HEART. CARDIOPULMONARY RESUSCITATION WAS SUBSEQUENTLY INITIATED; HOWEVER, THE FLOWS KEPT DROPPING AND A CODE WAS CALLED. ULTIMATELY, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS: "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY" (INCLUDING VENTRICULAR PERFORATION). Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND LOW PUMP FLOW, HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE POSITIONING ISSUE MOST LIKELY WAS PATIENT CONDITION RELATED SINCE PATIENT HAD LEFT VENTRICLE (LV) ANEURYSM AND THE LEFT VENTRICLE REPEATEDLY PUSHED THE DEVICE OUT OF THE HEART. THE CAUSE OF THE LOW PUMP FLOW ISSUE CANNOT BE DETERMINED SINCE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27187 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 CODE 1884 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27187.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752166&pc=NPW | 2015691-2025-02506 | 2025-03-05 05:00 | Death | EDWARDS LIFESCIENCES | 2025-04-02 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device | Cardiac Perforation | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE DELIVERY SYSTEM (DS) UTILIZED IN TRICUSPID POSITION WHERE, DURING THE PROCEDURE, THE WIRE WAS PUSHED OUT OF DS WHICH CAUSED AN EFFUSION AND ULTIMATELY THE PATIENT EXPIRED. DELIVERY SYSTEM INSERTED THROUGH RIGHT JUGULAR VEIN AND NOTED SOME LEAD INTERACTION IN ATRIUM AND RIGHT VENTRICLE (RV) LOOPED AROUND DS IN VENTRICLE. DS MANIPULATION WAS NOT LIMITED BY LEAD INTERACTION, BUT ALL AGREED IT WOULD BE MORE OPTIMAL, FOR ANCHOR REVEAL AND LEAFLET CAPTURE, TO ATTEMPT TO HAVE LEAD DRAPE ON OTHER SIDE OF DS. WHILE BACKING DS OUT AND CLOCKING/COUNTER-CLOCKING, THE WIRE WAS PULLED BACK WITH MINIMAL CURL EXPOSURE. THE TEAM WAS ADVISED TO NOT PULL WIRE BACK ANY FURTHER. ANOTHER MANEUVER WAS MADE AND WIRE WAS RETRACTED INTO SYSTEM BY SECONDARY OPERATOR. IMMEDIATELY AGAIN, TEAM WAS ADVISED NOT TO PUSH WIRE BACK OUT OF THE DELIVERY SYSTEM. UNFORTUNATELY, WIRE WAS PUSHED/PULLED OUT OF DELIVERY SYSTEM ANOTHER TWO TIMES. BLOOD PRESSURE BEGAN TO DROP. RAPID EFFUSION AND TAMPONADE NOTED ON ECHO WITH FIRST MEASUREMENT OF 1.5 CM THEN IMMEDIATELY (20 SECONDS) 4.0 CM. ECHO TEAM WAS INSTRUCTED TO LOOK FOR WIRE/DEVICE IN PERICARDIAL SPACE, BOTH WERE CONFIRMED. COMPRESSIONS STARTED AND PROCEDURE FOR PERICARDIOCENTESIS UNDERWAY. SURGICAL TEAM PREPPED THE PATIENT FOR OPEN CONVERSION. INSERTION OF CANNULAS INTO SUPERIOR VENA CAVA (SVC) RUPTURED A PORTION OF THE VESSEL. SURGICAL TEAM WAS ABLE TO REPAIR RV. MORE TIME WAS SPENT TO REPAIR SVC, AND WHEN TAKING PATIENT OFF PUMP, THE RV CONTINUED FILLING, BUT NOT EMPTYING UNTIL THE RV RUPTURED AND WAS CONSIDERED INOPERABLE. PATIENT EXPIRED. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. ADDITIONAL H6 TYPE OF INVESTIGATION CODE: 'ANALYSIS OF IMAGES' AND 'LABELLING REVIEW.' THE COMPLAINT FOR DELIVERY SYSTEM AND/OR GUIDEWIRE PUSHED INTO VENTRICULAR WALL WAS CONFIRMED WITH OBJECTIVE EVIDENCE. THE IMAGING EVALUATION CONFIRMED THAT THE SAFARI GUIDEWIRE PERFORATED THE RIGHT VENTRICLE AND WAS OBSERVED IN THE PERICARDIAL SPACE. THE GUIDEWIRE WAS FULLY RETRACTED INTO THE DELIVERY SYSTEM AND THEN PUSHED FULLY OUT OF THE DELIVERY SYSTEM, WHICH IS IMPROPER GUIDEWIRE TECHNIQUE. NO MANUFACTURING NON-CONFORMITIES WERE IDENTIFIED FROM THE IMAGING EVALUATION. THE DEVICE HISTORY RECORD REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING AND STERILIZATION INSPECTIONS. NO NONCONFORMANCES RELATED TO THE COMPLAINT EVENT WERE IDENTIFIED. SINCE NO EDWARDS DEFECT WAS IDENTIFIED, CORRECTIVE OR PREVENTATIVE ACTIONS ARE NOT REQUIRED. PROCEDURE MANUAL AND IFU PROVIDE INSTRUCTION AND BEST PRACTICES TO AVOID EXCESSIVE MOVEMENT AND WIRE PRESSURE TO PREVENT THE OCCURRENCE OF ANY DAMAGE TO THE CARDIAC VASCULATURE WITH THE GUIDEWIRE. WHILE THESE DOCUMENTS ARE WRITTEN WITH INTENT FOR ON-LABEL FEMORAL ACCESS USE, THE GUIDELINES ARE STILL APPLICABLE. FURTHERMORE, ALL PHYSICIANS PRESENT AT THIS CASE ARE TRAINED TO THE LATEST REVISION OF THE IFU AND PROCEDURE MANUAL REFERENCED IN THIS COMPLAINT AND THEREFORE TRAINED TO BEST PRACTICES FOR GUIDEWIRE MANAGEMENT. THE EDWARDS PERSONNEL SUPPORTING THE CASE WERE PROVIDING INSTRUCTIONS TO THE PHYSICIAN REGARDING GUIDEWIRE MANAGEMENT, AND THE PHYSICIAN DID NOT FOLLOW THESE INSTRUCTIONS. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752224&pc=NIP | 2124215-2025-17332 | 2023-07-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-02 04:00 | NIP | INNOVA | Fracture | Cardiac Arrest; Hemorrhage/Bleeding; Perforation of Vessels; Pseudoaneurysm; Ruptured Aneurysm |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2023, THE PATIENT UNDERWENT TREATMENT OF A MODERATELY CALCIFIED EXTERNAL ILIAC ARTERY. A GUIDEWIRE WAS POSITIONED AND AN 8X40X75 INNOVA SELF-EXPANDING STENT WAS IMPLANTED. THE STENT WAS POST DILATED VIA BALLOON ANGIOPLASTY. THE PATIENT WAS SENT TO A RECOVERY ROOM. WHILE IN RECOVERY, THE PATIENT EXPERIENCED A DECREASE IN BLOOD PRESSURE AND RETURNED TO THE OPERATING ROOM. THE INNOVA STENT WAS FOUND TO BE FRACTURED, WHICH CAUSED AN ANEURYSM THAT HAD RUPTURED. AN ADDITIONAL STENT WAS PLACED, AND ANGIOPLASTY WAS PERFORMED FOR 5 MINUTES TO TREAT THE ANEURYSM. THE PATIENT WAS SENT TO THE RECOVERY ROOM AND LATER MOVED TO THE INTENSIVE CARE UNIT. ON (B)(6) 2023, THE PATIENT HAD A CARDIAC ARREST AND DIED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2023, THE PATIENT UNDERWENT TREATMENT OF A MODERATELY CALCIFIED EXTERNAL ILIAC ARTERY. A GUIDEWIRE WAS POSITIONED AND AN 8X40X75 INNOVA SELF-EXPANDING STENT WAS IMPLANTED. THE STENT WAS POST DILATED VIA BALLOON ANGIOPLASTY. THE PATIENT WAS SENT TO A RECOVERY ROOM. WHILE IN RECOVERY, THE PATIENT EXPERIENCED A DECREASE IN BLOOD PRESSURE AND RETURNED TO THE OPERATING ROOM. THE INNOVA STENT WAS FOUND TO BE FRACTURED, WHICH CAUSED AN ANEURYSM THAT HAD RUPTURED. AN ADDITIONAL STENT WAS PLACED, AND ANGIOPLASTY WAS PERFORMED FOR 5 MINUTES TO TREAT THE ANEURYSM. THE PATIENT WAS SENT TO THE RECOVERY ROOM AND LATER MOVED TO THE INTENSIVE CARE UNIT. ON (B)(6) 2023, THE PATIENT HAD A CARDIAC ARREST AND DIED. IT WAS FURTHER REPORTED THE PATIENT UNDERWENT A LEFT EXTERNAL ENDARTERECTOMY AND RIGHT EXTERNAL ILIAC ARTERY STENTING. THE RIGHT EXTERNAL ILIAC ARTERY STENTING WAS COMPLICATED BY STENT FRACTURE. THERE WAS A SMALL PSEUDOANEURYSM THAT WAS REPAIRED INTRA OPERATIVELY. THE PATIENT BLED FROM THE SITE IN THE POST ANESTHETIC CARE UNIT AND WAS IMMEDIATELY BROUGHT BACK TO THE OPERATING ROOM FOR PLACEMENT OF A COVERED STENT. THE PATIENT WAS THEN ADMITTED TO THE INTENSIVE CARE UNIT (ICU), WHERE THEY EXPERIENCED PROFOUND HEMORRHAGIC SHOCK FOLLOWED BY DISTRIBUTIVE SHOCK. THE PATIENT DEVELOPED MULTI-ORGAN FAILURE AND PERSISTENT ELEVATED LACTATE DESPITE AGGRESSIVE BLOOD PRODUCT RESUSCITATION AND SUPPORTIVE CARE, RESULTING IN CARDIAC ARREST AND DEATH. Manufacturer Narrative: CORRECTED: D4: LOT NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI) #. H4: DEVICE MANUFACTURE DATE. UPDATED: B5: DESCRIBE EVENT OR PROBLEM. H6: PATIENT CODES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752274&pc=NPW | 2015691-2025-02508 | 2025-02-02 05:00 | Death | EDWARDS LIFESCIENCES | 2025-04-02 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Patient Device Interaction Problem | Heart Block | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE VALVE IN TRICUSPID POSITION WHERE THE PROCEDURE WAS COMPLETED. ON POST-OPERATIVE DAY (POD) 10, PATIENT EXPERIENCED COMPLETE HEART BLOCK AND EXPIRED. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. ADDITIONAL CODE FOR TYPE OF INVESTIGATION: LABELLING REVIEW. THE COMPLAINT FOR VALVE INTERACTS WITH CONDUCTION SYSTEM WAS CONFIRMED WITH OTHER EMPIRICAL EVIDENCE VIA INFORMATION PROVIDED BY EDWARDS CLINICAL SPECIALIST. THE DEVICE HISTORY RECORD REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING AND STERILIZATION INSPECTIONS. NO NON-CONFORMANCES RELATED TO THE COMPLAINT EVENT WERE IDENTIFIED. SINCE NO EDWARDS DEFECT WAS IDENTIFIED, CORRECTIVE OR PREVENTATIVE ACTIONS ARE NOT REQUIRED. INFORMATION PROVIDED PER THE INSTRUCTIONS FOR USE (IFU), CONDUCTION SYSTEM DISTURBANCE AND/OR HEART BLOCK WHICH MAY REQUIRE TREATMENT (INCLUDING PERMANENT PACEMAKER) ARE POTENTIAL ADVERSE EVENTS. HEART BLOCK AND OTHER CONDUCTION DISTURBANCES ARE COMMON IN PATIENTS WITH UNDERLYING CARDIOVASCULAR DISEASE AND CAN BE EXACERBATED OR AGGRAVATED WITH STANDARD INTRA-OPERATIVE MEDICATIONS OR ANESTHESIA. PATIENT WAS REPORTED TO HAVE EXPERIENCED COMPLETE HEART BLOCK ON POD 10 AND PASSED AWAY. NO INFORMATION PROVIDED ON PATIENT STATUS IMMEDIATELY POST PROCEDURE. MECHANISMS FOR CONDUCTION ABNORMALITIES COULD BE RELATED TO CORONARY OBSTRUCTION DUE TO AIR EMBOLISM, CORONARY SPASM AND/OR LOCAL EDEMA AFFECTING THE AV NODE. CONDUCTION DISTURBANCES HAVE ALSO BEEN DOCUMENTED DURING TRANSCATHETER CARDIAC PROCEDURES AS A RESULT OF DIRECT INTERACTION WITH CARDIAC ANATOMY, ESPECIALLY IN PATIENTS WITH UNDERLYING CONDUCTION ABNORMALITIES. OTHER POTENTIAL CAUSES INCLUDE TRAUMA BY A GUIDEWIRE OR DELIVERY SYSTEM. ADDITIONALLY, CARDIAC CONDUCTION SYSTEM COMPLICATIONS ARE KNOWN RISKS WITH TRICUSPID VALVE INTERVENTIONS DUE TO THE PROXIMITY OF THE ATRIOVENTRICULAR NODE (AV NODE) TO THE SEPTAL LEAFLET OF THE TRICUSPID VALVE. THESE CONDUCTION DISTURBANCES CAN LEAD TO POST-OPERATIVE HEART BLOCK REQUIRING PACEMAKER IMPLANTATION - IN EXTREME CONDITIONS DEATH. THIS EVENT MAY HAVE POTENTIALLY BEEN INFLUENCED BY PATIENT HEALTH FACTORS; HOWEVER, A DEFINITE ROOT CAUSE IS UNABLE TO BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752278&pc=DTZ | 3013876692-2025-00025 | 2025-03-21 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-02 04:00 | DTZ | TUBING SET | Fluid/Blood Leak; Gas/Air Leak | Cardiac Arrest | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN FRANCE. IT WAS REPORTED THAT THE CIRCUIT WAS PRIMED WITH WATER IN ACCORDANCE WITH STANDARD PROCEDURES TO PREPARE FOR ECMO SUPPORT. THE FLOW RATE WAS OBTAINED WITHOUT DIFFICULTY. A FEW MINUTES LATER THERE WAS A LOSS OF FLOW, A MACHINE ALARM AND VISIBLE BUBBLES WITHIN THE OXYGENATOR. THE SYSTEM COULD NOT BE DE-AIRED. THE CIRCUIT WAS CLAMPED AND THE CIRCUIT REPLACED. THE PATIENT WAS ADMITTED FOR EMERGENCY CARDIAC ARREST. THE CARE WAS DELAYED FOR THIS PATIENT. THE ECMO WAS INITIATED AT 8 PM AND THE PATIENT PASSED AWAY AT 3 AM. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE FRANCE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: THE EVENT OCCURRED IN FRANCE DURING TREATMENT. IT WAS REPORTED THAT THE CIRCUIT WAS PRIMED WITH WATER IN ACCORDANCE TO STANDARD PROCEDURES TO PREPARE FOR ECMO SUPPORT. THE FLOW RATE WAS OBTAINED WITHOUT DIFFICULTY. A FEW MINUTES LATER THERE WAS A LOSS OF FLOW, A MACHINE ALARM AND VISIBLE BUBBLES WITHIN THE OXYGENATOR. THE SYSTEM COULD NOT BE DE-AIRED. THE CIRCUIT WAS CLAMPED AND THE CIRCUIT REPLACED. THE PATIENT WAS ADMITTED FOR EMERGENCY CARDIAC ARREST. THE CARE WAS DELAYED FOR THIS PATIENT. THE ECMO WAS INITIATED AT 8 PM ON (B)(6) 2025 AND THE PATIENT PASSED AWAY AT 3 AM ON (B)(6) 2025. NEW INFORMATION WAS RECEIVED ON 2025-04-03 THAT THE PATIENT WAS CONNECTED TO THE FIRST HLS SET WHICH CONTAINED AIR IN THE SYSTEM. LATER, IT WAS CONFIRMED THAT THE REPLACEMENT SET HAD NO FAULT AND WAS DISCARDED BY THE CUSTOMER. THE FIRST SET WAS AVAILABLE FOR INVESTIGATION, WHICH ACCORDING TO THE CUSTOMER HAD NO VISIBLE DEFECT NOR ANY LEAKS DURING/AFTER PRIMING. FURTHER INFORMATION WAS RECEIVED THAT THERE WAS NO AIR ENTRY AFTER THE REPLACEMENT OF THE AFFECTED SET. THE USED CANNULAS WERE NOT REPLACED AND WERE THEREFORE WORKING AS EXPECTED. NO INFORMATION WAS AVAILABLE IF THE USED CANNULAS WERE GETINGE CANNULAS. THE CANNULAS WERE DISCARDED BY THE CUSTOMER. AS PATIENT DEATH WAS REPORTED, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS INVESTIGATED BY THE GETINGE LABORATORY ON 2025-06-27 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. THERE WERE NO VISIBLE DAMAGES ON THE PRODUCT. THE PRODUCT COULD BE PRIMED WITHOUT ANY PROBLEMS AND THERE WAS NO AIR ENTRY. THE PRODUCT FUNCTIONALITY WAS AS EXPECTED. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-07-04 WITH FOLLOWING CONCLUSION: "ON (B)(6) 2025, AN INCIDENT OCCURRED INVOLVING AN HLS SET ADVANCED 7.0 USED WITH THE CARDIOHELP-I SYSTEM DURING VA-ECLS. THE PATIENT, A 45-YEAR-OLD FEMALE ADMITTED AFTER CARDIAC ARREST, RECEIVED ECLS SUPPORT BEGINNING AT APPROXIMATELY 20:15. WHILE INITIAL STARTUP WAS UNEVENTFUL AND TARGET FLOW WAS ACHIEVED, A SUDDEN LOSS OF FLOW OCCURRED SHORTLY THEREAFTER, ACCOMPANIED BY BUBBLE DETECTION ALARMS. AIR BUBBLES WERE OBSERVED IN THE OXYGENATOR AND COULD NOT BE REMOVED. NO AIR ENTRY SOURCE WAS IDENTIFIED AT THE TIME. THE CLINICAL TEAM REPLACED THE CIRCUIT, WHICH FUNCTIONED NORMALLY. THE PATIENT EXPIRED AT AROUND 03:00; NO DIRECT CAUSAL LINK TO THE INCIDENT WAS ESTABLISHED, THOUGH TREATMENT WAS DELAYED. THE SUBSEQUENT INVESTIGATION FOUND NO PRODUCT-RELATED DEFECTS. A MINOR LEAK AT THE VENOUS SAMPLING PORT WAS ATTRIBUTED TO DRIED BLOOD RESIDUE AND WAS NO LONGER PRESENT AFTER CLEANING. THE TUBING, INCLUDING THE CRITICAL VENOUS LINE, SHOWED NO DAMAGE OR KINKS. FUNCTIONAL TESTING, INCLUDING A 60-MINUTE LEAK TEST, CONFIRMED SYSTEM INTEGRITY. A PRESSURE SENSOR MALFUNCTION WAS NOTED BUT DETERMINED TO BE RELATED TO TRANSPORT AND NOT CONNECTED TO THE COMPLAINT. IMPORTANTLY, THE VENOUS CANNULA USED DURING ECMO INITIATION - A CRITICAL COMPONENT IN ASSESSING THE ROOT CAUSE OF AIR INGRESS - WAS NOT RETURNED FOR FURTHER EVALUATION. THIS OMISSION LIMITS THE ABILITY TO FULLY RULE OUT A CONNECTION-RELATED ISSUE OR A DEFECT AT THE INSERTION SITE. LITERATURE SUPPORTS THE RISK OF AIR INGRESS VIA THE VENOUS LINE OR CANNULA DURING ECMO, PARTICULARLY IN THE PRESENCE OF LOOSE CONNECTIONS, NEGATIVE PRESSURE, OR INSUFFICIENT PRIMING. IN ADDITION, AIR INGRESS THROUGH A CENTRAL VENOUS CATHETER, IF PRESENT, COULD BE A CONTRIBUTING FACTOR, PARTICULARLY IF CONNECTIONS WERE NOT FULLY SECURED. THE ASSUMPTION OF VENOUS AIR ENTRAINMENT APPEARS TO CONFORM TO THE EVENT PATTERN AS DESCRIBED: LOSS OF FLOW (DUE TO AIR INVASION OF THE CENTRIFUGAL PUMP), AIR IN THE OXYGENATOR, THEN BUBBLE ALARM. IT IS PRESUMED THAT THE REPORTED BUBBLE ALARM WAS ARTERIAL. NO MENTION OF VENOUS BUBBLE SENSOR DEPLOYMENT, HOWEVER, WAS CITED IN THE COMPLAINT NARRATIVE. IN CONCLUSION AND GIVEN THE ABSENCE OF THE VENOUS CANNULA FOR INVESTIGATION AND THE LIMITED PROCEDURAL DETAILS AVAILABLE, NO DEFINITIVE ROOT CAUSE OF THE AIR INGRESS CAN BE CONFIRMED. HOWEVER, POSSIBLE CAUSES INCLUDE A LEAK AT THE VENOUS CANNULA SITE, AIR ENTRY VIA A CENTRAL VENOUS CATHETER, OR A TRANSIENT LEAK AT THE VENOUS SAMPLING PORT, OBSERVED DURING INITIAL FLUSHING AND ATTRIBUTED TO DRIED BLOOD RESIDUES, THOUGH THIS WAS NOT REPRODUCIBLE UNDER CONTROLLED TESTING CONDITIONS." ACCORDING TO THE INSTRUCTION OF USE OF THE HLS SET IT IS MENTIONED THAT INCORRECT POSITIONING OF THE CARDIOHELP CAN CAUSE AIR PERMEATION ON THE BLOOD SIDE OF THE HLS MODULE ADVANCED. THE CARDIOHELP MUST BE POSITIONED AT A LEVER LOWER THAT THE PATIENT AND SECURED CLOSE TO THE PATIENT. THE FLOW/BUBBLE SENSOR MUST ALWAYS BE AFFIXED ON THE ARTERIAL SIDE OF THE SET IF YOU ARE USING THE CARDIOHELP-I. WITH AN ACTIVATED INTERVENTION, THE DETECTION OF BUBBLES BY THE FLOW/BUBBLE SENSOR TRIGGERS A PUMP STOP. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-07-03. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "AIR IN SYSTEM" COULD NOT BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE HLS SET. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752360&pc=DSQ | 2916596-2025-01869 | 2025-03-01 05:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE SYSTEM ORGAN FAILURE (MSOF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS FURTHER COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752380&pc=DSQ | 2916596-2025-02163 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE.Event Description: THE PATIENTS MSOF WAS NOT CONSIDERED TO BE A PROGRESSION OF THEIR UNDERLYING CARDIAC DISEASE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: A4: PATIENT WEIGHT NOT PROVIDED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752410&pc=LWS | 2649622-2025-08915 | 2025-03-10 04:00 | Death | MPRI | 2025-04-02 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Appropriate Term/Code Not Available | Bacterial Infection; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED FOR AN UNSPECIFIED REASON. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE DEVICE SYSTEM WAS EXPLANTED DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA INFECTION. THE PATIENT PASSED AWAY APPROXIMATELY SIXTEEN DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752429&pc=DTB | 2649622-2025-08916 | 2025-03-10 04:00 | Death | MPRI | 2025-04-02 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Bacterial Infection; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED FOR AN UNSPECIFIED REASON. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE DEVICE SYSTEM WAS EXPLANTED DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA INFECTION. THE PATIENT PASSED AWAY APPROXIMATELY SIXTEEN DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752436&pc=DSQ | 3007042319-2025-00716 | 2025-03-21 04:00 | Death | HEARTWARE, INC. | 2025-04-02 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY EXPIRED DUE TO RIGHT VENTRICULAR FAILURE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IN USE. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS RIGHT VENTRICULAR FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RIGHT VENTRICULAR FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF RIGHT HEART FAILURE. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752444&pc=OJX | 2649622-2025-08919 | 2025-03-10 04:00 | Death | MPRI | 2025-04-02 04:00 | OJX | ATTAIN STABILITY QUAD MRI SURESCAN | Appropriate Term/Code Not Available | Bacterial Infection; No Clinical Signs, Symptoms or Conditions | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED FOR AN UNSPECIFIED REASON. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE DEVICE SYSTEM WAS EXPLANTED DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA INFECTION. THE PATIENT PASSED AWAY APPROXIMATELY SIXTEEN DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752457&pc=NIK | 9614453-2025-01301 | 2025-03-10 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-02 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Bacterial Infection; No Clinical Signs, Symptoms or Conditions | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED FOR AN UNSPECIFIED REASON. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE DEVICE SYSTEM WAS EXPLANTED DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA INFECTION. THE PATIENT PASSED AWAY APPROXIMATELY SIXTEEN DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752662&pc=DTQ | 3013876692-2025-00026 | 2024-12-02 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-02 04:00 | DTQ | CARDIOHELP | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding | K133598 |  | 1 | Event Description: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON (B)(6) 2025: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2024-12-02. THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" AND THE CONSEQUENCE FOR THE PATIENT WAS "HEMORRHAGE AFTER REMOVING THE DEVICE AT THE INSERTION SITE. THIS RESULTED IN THE PATIENT'S DEATH. COMPLAINT ID# (B)(4).Event Description: THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE. COMPLAINT# (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE SPAIN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012. Manufacturer Narrative: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-03-31: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2024-12-02 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" AND THE CONSEQUENCE FOR THE PATIENT WAS "HEMORRHAGE AFTER REMOVING THE DEVICE AT THE INSERTION SITE. THIS RESULTED IN THE PATIENT'S DEATH. RENACER REGISTRY PATIENT (B)(6).¿ NEW INFORMATION WAS RECEIVED ON 2025-04-02 THAT NO MALFUNCTION ON THE CARDIOHELP DEVICE OCCURRED. THE EVENT WAS NOT RELATED TO THE CARDIOHELP DEVICE. THE TREATMENT WITH THE CARDIOHELP MACHINE WAS FINALIZED ON 2024-11-18. ON (B)(6) 2024, AND THE PATIENT SUFFERED A TEAR OF THE SUPERFICIAL FEMORAL ARTERY WITH MASSIVE HEMORRHAGE AND WAS REVIEWED BY THE SURGICAL TEAM. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" COULD NOT BE LINKED TO A DEVICE DEFECT/FAILURE. THEREFORE, THE COMPLAINT COULD NOT BE CONFIRMED. THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752699&pc=MWS | 2184009-2025-00472 | 2024-11-13 05:00 | Death | PERFUSION SYSTEMS | 2025-04-02 04:00 | MWS | OCTOPUS EVOLUTION TISSUE STABILIZER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | N/A |  | 1 | Event Description: A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALITY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 G4: DEVICE IS CLASS 1 EXEMPT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752711&pc=DSQ | 2916596-2025-02029 | 2025-03-05 05:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Tachycardia; Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO VENTRICULAR TACHYCARDIA, TRANSIENT RECURRENT VENTRICULAR FIBRILLATION (TRVF), AND RENAL FAILURE. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRYTHMIA, RENAL FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT", LIST ARRYTHMIA AS A POTENTIAL LATE IMPLANT COMPLICATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752735&pc=DSQ | 2916596-2025-01989 | 2025-03-17 04:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752737&pc=DSQ | 2916596-2025-02192 | 2025-02-24 05:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE PATIENT'S SON CALLED TO INFORM THE TEAM THAT THE PATIENT WAS FOUND DOWN IN THEIR HOME ON THE EVENING OF (B)(6) 2025 AND HAD PASSED AWAY. IT APPEARED THEY MIGHT HAVE FALLEN BUT THE SON WAS UNAWARE OF THE FULL DETAILS AS THE PATIENT WAS HOME ALONE. IT WAS UNKNOWN OF THE OUTCOME WAS DEVICE OR THERAPY RELATED, OR IF IT OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WOULD NOT RETURN FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21752962&pc=DSQ | 2916596-2025-02228 | 2006-01-01 05:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT, US, EP | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED PREVIOUSLY UNDER MFR # 2916596-2025-00764 THAT WERE CONCERNS FOR NEEDING A TOUGHER DRIVELINE. IT WAS NOTED THAT THERE WERE DEATHS IN THE HEARTMATE II TRIAL FROM CAR DOOR TRANSECTIONS. THE HEARTMATE II TRIAL WAS A BRIDGE TO TRANSPLANT (BTT) TRIAL. THE PATIENT'S HEARTWARE WAS "WEAR" TESTED BY SLAMMING DOORS DUE TO THIS EVENT. THEY WERE NOTED TO BE MUCH TOUGHER COMPARED TO HEARTMATE DRIVELINES. HEARTMATE 3 WAS NOTED TO BE BETTER THAN HEARTMATE II WHICH WERE DELICATE AND FLIMSY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE II DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE (HM) II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C AND THE HEARTMATE II PATIENT HANDBOOK REV. C ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. THE PATIENT HANDBOOK INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE PATIENT THINKS THAT, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. SECTIONS 6 AND 8 OF THE HMII IFU, AS WELL AS SECTIONS 4 AND 6 OF THE HMII PATIENT HANDBOOK, PROVIDE INFORMATION REGARDING HOW TO CARE FOR THE DRIVELINE. THE PATIENT HANDBOOK ALSO INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED). NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION B3: EVENT DATE IS UNKNOWN AND HAS BEEN ESTIMATED. SECTIONS A AND D: PATIENT AND DEVICE INFORMATION IS UNKNOWN. DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753096&pc=IZQ | 1518293-2025-00008 |  | Death | LIEBEL- FLARSHEIM | 2025-04-02 04:00 | IZQ | INJECTOR OPVSU NPS | Use of Device Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K063503 |  | 1 | Event Description: THIS INCIDENT WAS REPORTED BY A FACILITY IN (B)(6) ON (B)(6) 2025. "(B)(6) 2025, GUERBET WAS CONTACTED (B)(6) ASKING FOR DOCUMENTATION OF TRAINING FOR THE STAFF ON (B)(6) 2021. SHE REPORTED A FATAL AIR BOLUS USING THE OPTIVANTAGE, SHE DOES NOT BLAME THE INJECTOR BUT CONSIDERS THIS TECHNOLOGY ERROR WHO HAS BEEN USING THE INJECTOR SINCE INSTALL IN 2021. THEY USE GUERBET SELF FILL SYRINGES.". Manufacturer Narrative: OVERALL INVESTIGATION SUMMARY: GUERBET WAS CONTACTED BY A CUSTOMER REPORTING A FATAL AIR BOLUS OCCURRED USING OPTIVANTAGE INJECTOR 849002 SERIAL NUMBER (B)(6). THE CUSTOMER DID NOT BLAME THE INJECTOR AND CONSIDERS THIS A TECHNOLOGIST ERROR. ADDITIONAL DETAILS WERE REQUESTED FROM THE CUSTOMER BUT HAVE NOT BEEN RECEIVED AT THE TIME OF COMPLAINT INVESTIGATION CLOSURE. THIS COMPLAINT WILL BE REOPENED WHEN THE ADDITIONAL DETAILS ARE RECEIVED. THERE WAS NO REPORT OF MALFUNCTION OR REQUEST FOR EVALUATION OF THE INJECTOR; THEREFORE, THE INVESTIGATION REVIEWED HISTORICAL RECORDS AND DESIGN RECORDS. OPTIVANTAGE WAS NOT DESIGNED AND IS NOT CAPABLE OF DETECTING OR PREVENTING AIR INJECTIONS; HOWEVER, AS LISTED BELOW, THE IFU INCLUDES INSTRUCTIONS FOR THE PROPER PURGING OF AIR FROM THE SYSTEM PRIOR TO INJECTION, AS WELL AS PRECAUTIONS AND STEPS TO PREVENT AIR INJECTIONS. HISTORICAL COMPLAINTS FOR OPTIVANTAGE AIR INJECTIONS WERE REVIEWED AND ALL INSTANCES WERE DETERMINED TO BE USER-RELATED, INDEPENDENT OF THE INJECTOR. THE RISK ASSESSMENT AND CONTROL DOCUMENTATION FOR THE OPTIVANTAGE WAS REVIEWED AND CONFIRMED THAT THE HAZARDOUS SITUATION OF AIR INJECTIONS LEADING TO AIR EMBOLISMS HAVE BEEN MITIGATED WITHIN THE MANUFACTURER'S CAPACITY. THE MITIGATIONS IN PLACE INCLUDE, BUT ARE NOT LIMITED TO WARNING STATEMENTS IN THE INSTRUCTIONS FOR USE (IFU), SYRINGE ADAPTERS PROVIDING A CLEAR VIEW OF THE SYRINGE, PROMPTS APPEARING ON SCREEN TO ENSURE THE USER FOLLOWS A PROPER FILL/PURGE SEQUENCE ON THE SYRINGES, A "CHECK FOR AIR..." MESSAGE ON THE POWERHEAD AND CONSOLE DISPLAYS, AND THE IFU CONTAINING INSTRUCTIONS ON HOW TO CLEAR AIR BUBBLES FROM THE CONSUMABLES. A REVIEW OF GUERBET'S COMPLAINT TRACKING SYSTEM RELATED TO THIS UNIT WAS CONDUCTED BUT FOUND NO RELATED INCIDENTS. GUERBET TECHNICAL SERVICE TEAM HAS NEVER SERVICED THIS INJECTOR OUTSIDE OF THE INITIAL INSTALLATION AND APPLICATIONS TRAINING. THE HEALTH HAZARD EVALUATION (HHE) (B4) WAS OPENED IN RELATION TO THIS COMPLAINT. THE HHE CONCLUSION INDICATES NO ACTIONS ARE WARRANTED AT THIS TIME. IMPACT ASSESSMENT SUMMARY: THE AIR BOLUS WAS FATAL AND THE PATIENT PASSED. IMDRF CODES = B13, B11; C19, C23; D11. ROOT/PROBABLE CAUSE CODE: PERSONNEL - PERFORMANCE - FAILED TO FOLLOW PROCEDURE. ROOT/ PROBABLE CAUSE SUMMARY: AIR INJECTIONS ARE GENERALLY USER-RELATED AS THE OPTIVANTAGE INJECTOR REQUIRES THE USERS TO MANUALLY PURGE THE AIR FROM THE SYSTEM AND CONFIRM THAT THEY HAVE CHECKED FOR AIR. BASED ON THIS AND THE INVESTIGATION, THERE IS NO NEED FOR A CAPA AT THIS TIME. COMPLAINTS ARE TRENDED AND REPORTED DURING MANAGEMENT AND METRIC REVIEWS TO CONSIDER CORRECTIVE ACTION. QUALITY WILL CONTINUE TO MONITOR AND TREND FOR SIMILAR ISSUES. DISPOSITION SUMMARY: THE UNIT REMAINED IN USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753278&pc=DWF | 2184009-2025-00474 | 2024-11-13 05:00 | Death | PERFUSION SYSTEMS | 2025-04-02 04:00 | DWF | CLEARVIEW ARTERIOTOMY SHUNT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K120612 |  | 1 | Event Description: A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALLY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753337&pc=NRY | 2029214-2025-00846 | 2024-03-04 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-02 04:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: MECHANICAL THROMBECTOMY IN STROKE¿RETROSPECTIVE COMPARISON OF METHODS: ASPIRATION VS. STENT RETRI EVERS VS. COMBINED METHOD¿IS ASPIRATION THE BEST STARTING POINT? THE OBJECTIVE IS TO COMPARE THE TECHNICAL AND CLINICAL EFFICACY OF THREE MAIN MECHANICAL THROMBECTOMY METHODS IN PATIENTS WITH LARGE VESSEL OCCLUSION. THE TIME FRAME OF THIS STUDY WAS: FROM JUNE 2019 TO MAY 2021. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION: MEDTRONIC, MICROVENTION, PENUMBRA, AND OTHERS WERE USED. THE FOLLOWING MEDTRONIC DEVICES WERE USED: STENT RETRIEVERS: SOLITAIRE AND CATCH. ASPIRATION CATHETERS: REACT. DEATHS OCCURRED IN THE STUDY POPULATION: THE DOCUMENT MENTIONS IN-HOSPITAL MORTALITY RATES AS FOLLOWS: 12 (28%) FOR ASPIRATION ONLY (AO), 34 (27%) FOR STENT RETRIEVER ONLY (SO), 34 (32%) FOR COMBINED METHOD (CM), WITH NO SIGNIFICANT DIFFERENCES AMONG THE GROUPS. THE CAUSES OF DEATH WERE: THE DOCUMENT DOES NOT SPECIFY INDIVIDUAL CAUSES OF DEATH. AMONG PATIENT ADVERSE EVENTS INCLUDED: SYMPTOMATIC INTRACRANIAL HEMORRHAGES (SICH) AND INTRACRANIAL HEMORRHAGES (ICH). RATES OF SICH: 2 (5%) FOR AO, 11 (9%) FOR SO, AND 9 (8%) FOR CM, WITH NO SIGNIFICANT DIFFERENCES AMONG THE GROUPS. CLINICAL OUTCOMES OF INTEREST INCLUDED THE 90-DAY MODIFIED RANKIN SCALE (MRS) SCORES, SHOWING NO SIGNIFICANT DIFFERENCES AMONG THE GROUPS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: UNK-NV-SOLITAIRE; DATE G2: CITATION: AUTHORS: MEDER, G., ZUCHOWSKI, P., SKURA, W., PLESZKA, P., DURA, M., RAJEWSKI, P., NOWACZEWSKA, M., MEDER, M., ALEXANDRE, M., PEDICELLI, A. MECHANICAL THROMBECTOMY IN STROKE¿RETROSPECTIVE COMPARISON OF METHODS: ASPIRATION VS. STENT RETRIEVERS VS. COMBINED METHOD¿IS ASPIRATION THE BEST STARTING POINT? J. CLIN. MED. 13, 1477 2024. DOI: 10.3390/JCM13051477 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753361&pc=DQY | 2029214-2025-00848 | 2024-03-04 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-02 04:00 | DQY | CATHETER, PERCUTANEOUS | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: MECHANICAL THROMBECTOMY IN STROKE¿RETROSPECTIVE COMPARISON OF METHODS: ASPIRATION VS. STENT RETRI EVERS VS. COMBINED METHOD¿IS ASPIRATION THE BEST STARTING POINT? THE OBJECTIVE IS TO COMPARE THE TECHNICAL AND CLINICAL EFFICACY OF THREE MAIN MECHANICAL THROMBECTOMY METHODS IN PATIENTS WITH LARGE VESSEL OCCLUSION. THE TIME FRAME OF THIS STUDY WAS: FROM JUNE 2019 TO MAY 2021. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION: MEDTRONIC, MICROVENTION, PENUMBRA, AND OTHERS WERE USED. THE FOLLOWING MEDTRONIC DEVICES WERE USED: STENT RETRIEVERS: SOLITAIRE AND CATCH. ASPIRATION CATHETERS: REACT. DEATHS OCCURRED IN THE STUDY POPULATION: THE DOCUMENT MENTIONS IN-HOSPITAL MORTALITY RATES AS FOLLOWS: 12 (28%) FOR ASPIRATION ONLY (AO), 34 (27%) FOR STENT RETRIEVER ONLY (SO), 34 (32%) FOR COMBINED METHOD (CM), WITH NO SIGNIFICANT DIFFERENCES AMONG THE GROUPS. THE CAUSES OF DEATH WERE: THE DOCUMENT DOES NOT SPECIFY INDIVIDUAL CAUSES OF DEATH. AMONG PATIENT ADVERSE EVENTS INCLUDED: SYMPTOMATIC INTRACRANIAL HEMORRHAGES (SICH) AND INTRACRANIAL HEMORRHAGES (ICH). RATES OF SICH: 2 (5%) FOR AO, 11 (9%) FOR SO, AND 9 (8%) FOR CM, WITH NO SIGNIFICANT DIFFERENCES AMONG THE GROUPS. CLINICAL OUTCOMES OF INTEREST INCLUDED THE 90-DAY MODIFIED RANKIN SCALE (MRS) SCORES, SHOWING NO SIGNIFICANT DIFFERENCES AMONG THE GROUPS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: UNK-NV-SOLITAIRE; PRODUCT ID: UNK-NV-SOLITAIRE; DATE G2: CITATION: AUTHORS: MEDER, G., ZUCHOWSKI, P., SKURA, W., PLESZKA, P., DURA, M., RAJEWSKI, P., NOWACZEWSKA, M., MEDER, M., ALEXANDRE, M., PEDICELLI, A. MECHANICAL THROMBECTOMY IN STROKERETROSPECTIVE COMPARISON OF METHODS: ASPIRATION VS. STENT RETRIEVERS VS. COMBINED METHOD¿IS ASPIRATION THE BEST STARTING POINT? J. CLIN. MED. 13, 1477 2024. DOI: 10.3390/JCM13051477. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753467&pc=FQH | 2184009-2025-00476 | 2024-11-13 05:00 | Death | PERFUSION SYSTEMS | 2025-04-02 04:00 | FQH | CLEARVIEW BLOWER/ MISTER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K973485 |  | 1 | Event Description: A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALLY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753538&pc=CAC | 2184009-2025-00478 | 2024-11-13 05:00 | Death | PERFUSION SYSTEMS | 2025-04-02 04:00 | CAC | DISTAL PERFUSION KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K994171 |  | 1 | Event Description: A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALLY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753561&pc=LWS | 2124215-2025-20365 | 2025-03-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-02 04:00 | LWS | EMBLEM S-ICD | Failure to Convert Rhythm; Low impedance | Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE WAS AN ATTEMPTED IMPLANT. DURING THE IMPLANT PROCEDURE, THE ELECTRODE WAS CONNECTED TO THE DEVICE, AND TESTING INDICATED SYSTEM IMPEDANCE WAS 70 OHMS. A VENTRICULAR FIBRILLATION (VF) WAS TRIGGERED, PROVIDING A 65J SHOCK DELIVERY. THE DEVICE DETECTED THE VF AND WAS NOT ABLE TO TERMINATE RHYTHM. EXTERNAL DEFIBRILLATION WAS NEEDED TO RETURN TO NORMAL SINUS RHYTHM, WITH A SHOCK IMPEDANCE OF LESS THAN 30 OHMS. THE SICD WAS EXPLANTED PRIOR TO WOUND CLOSURE AND THE ELECTRODE WAS EXPLANTED. THE POCKET WAS CLOSED, AND NO DEVICE IMPLANTED. AT THIS TIME THE PATIENT REMAINS IN THE HOSPITAL FOR MONITORING. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE WAS AN ATTEMPTED IMPLANT. DURING THE IMPLANT PROCEDURE, THE ELECTRODE WAS CONNECTED TO THE DEVICE, AND TESTING INDICATED SYSTEM IMPEDANCE WAS 70 OHMS. A VENTRICULAR FIBRILLATION (VF) WAS TRIGGERED, PROVIDING A 65J SHOCK DELIVERY. THE DEVICE DETECTED THE VF AND WAS NOT ABLE TO TERMINATE RHYTHM. EXTERNAL DEFIBRILLATION WAS NEEDED TO RETURN TO NORMAL SINUS RHYTHM, WITH A SHOCK IMPEDANCE OF LESS THAN 30 OHMS. THE SICD WAS EXPLANTED PRIOR TO WOUND CLOSURE AND THE ELECTRODE WAS EXPLANTED. THE POCKET WAS CLOSED, AND NO DEVICE IMPLANTED. UNFORTUNATELY, THE PATIENT PASSED AWAY SEVERAL DAYS AFTER THIS PROCEDURE. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS AND IS PENDING THE INVESTIGATION CONCLUSION. ONCE THE ANALYSIS IS COMPLETED, THIS RECORD WILL BE UPDATED WITH RESULTS. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS B2 (OUTCOMES ATTRIB TO ADV EVENT), B5 (EVENT DESCRIPTION), D6B (EXPLANT DATE), H1 (TYPE OF REPORTABLE EVENT), AND H6 (IMPACT CODES).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753645&pc=BZD | 2518422-2025-103625 | 2025-02-04 05:00 | Death | RESPIRONICS, INC. | 2025-04-02 04:00 | BZD | BIPAP SYSTEM ONE 60 SERIES | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION DUE TO A USER OF BIPAP AUTO BIFLEX. DAUGHTER CALLED IN STATING PATIENT PASSED AWAY AND DOES NOT NEED THE REPLACEMENT DEVICE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE HAS BEEN RETURNED FROM THE THIRD-PARTY SERVICE CENTER. EVALUATION SHOWS THE UNIT IS FUNCTIONAL. THERE ARE NO VISIBLE FOAM PARTICLES AND HAS CONTAMINATION FINDING WHICH IS DUST. THE DEVICE WAS ROUTED TO SCRAP. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753830&pc=DSQ | 2916596-2025-01947 | 2025-03-13 04:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD MULTIPLE SYSTEM ORGAN FAILURE, WITH A DECISION TO MOVE TO COMFORT MEASURES AND TURN OFF THE VENTRICULAR ASSIST DEVICE (VAD). THE PATIENT THEN PASSED AWAY DUE TO CARDIAC ARREST. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED AND AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21753992&pc=NPT | 2025587-2025-02291 | 2023-09-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-02 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO WORSENING HEART FAILURE. AFTER THE HOSPITALIZATION, A DECREASED ACTIVITIES OF DAILY LIVING SCORE WAS NOTED AND THE PATIENT WAS TRANSFERRED TO REHABILITATION. APPROXIMATELY 2 MONTHS AND 2 WEEKS LATER, THE PATIENT PASSED AWAY IN ANOTHER HOSPITAL. THE CAUSE OF DEATH IS UNKNOWN. THE RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE OR THE PROCEDURE IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21754152&pc=DSQ | 2916596-2025-01840 | 2025-03-08 05:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS BROUGHT INTO THE EMERGENCY DEPARTMENT ON (B)(6) 2025 FOR A MASSIVE LEFT SUBDERMAL HEMATOMA WITH A 2+CM MIDLINE SHIFT WITH PONTOMEDULLARY HEMORRHAGE. ADDITIONALLY, THE CENTER NOTICED NO EXTERNAL POWER ALARMS AND DRIVELINE DISCONNECTS ON (B)(6) 2025 FROM INTERROGATION. LOG FILES WERE SUBMITTED FOR REVIEW AND REVEALED THE DRIVELINE DISCONNECT ALARMS AND PUMP OFF EVENTS APPEARED TO HAVE BEGUN BETWEEN 10:00 AM AND 11:00 AM ON (B)(6) 2025. THE DRIVELINE WAS RECONNECTED AT 01:05 PM ON (B)(6) 2025. THE LOG FILES DID NOT REVEAL WHAT LED TO THE DRIVELINE BEING DISCONNECTED. THE LOG FILE ALSO CAPTURED MULTIPLE LOW POWER AND NO EXTERNAL POWER ALARMS BETWEEN (B)(6) 2025. THE LOW POWER ALARMS WERE OCCURRING ON BATTERY POWER. THE NO EXTERNAL POWER ALARMS WERE OCCURRING WHILE ON MOBILE POWER UNIT (MPU) POWER. THE LOG FILE APPEARS TO SHOW A LOSS OF POWER TO THE MPU. THE PUMP CONTINUED TO OPERATE AT THE SET SPEED AND WAS SUPPORTED BY THE SYSTEM CONTROLLER'S EMERGENCY BACKUP BATTERY UNTIL EXTERNAL POWER WAS RESTORED. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PATIENT PASSED AWAY FROM STROKE ON 11MAR2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE PUMP STOP THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 08MAR2025 THROUGH 11MAR2025. THE FILE CAPTURED THE DRIVELINE PREVIOUSLY NOT CONNECTED TO THE SYSTEM CONTROLLER AND WAS LATER CONNECTED ON 08MAY2025. FOLLOWING THE DRIVELINE BEING CONNECTED, THE PUMP RAMPED BACK UP TO THE FIXED SPEED WITHOUT ISSUE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED WHILE THE DRIVELINE WAS CONNECTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. SECTION 2 OF THE IFU AND SECTION 2 OF THE PATIENT HANDBOOK INSTRUCT THE USER TO CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK ALSO EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. ADDITIONALLY, SECTION 2 OF THE IFU, UNDER "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON ALL SYSTEM CONTROLLER ALARMS, INCLUDING DRIVELINE DISCONNECT ALARMS, AS WELL AS THE PROPER ACTIONS ASSOCIATED WITH THEM. THESE DOCUMENTS ALSO WARN OF EVENTS THAT MAY CAUSE THE PUMP TO STOP AND HOW TO PREVENT PUMP STOPS FROM OCCURRING. ADDITIONALLY, THE ¿HANDLING EMERGENCIES¿ SECTION OF THE PATIENT HANDBOOK LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE. FURTHERMORE, THE PATIENT HANDBOOK CAUTIONS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755159&pc=DSQ | 2916596-2025-02023 | 2025-03-13 04:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS WORKING AS EXPECTED FOR THE WHOLE TIME OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED MULTI ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755170&pc=NPT | 2025587-2025-02300 | 2025-02-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-02 04:00 | NPT | EVOLUT FX PLUS VALVE | Activation Failure; Patient Device Interaction Problem | Intracranial Hemorrhage; Tachycardia; Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS EVFXPLUS-23 DEVICE IN A PATIENT WITH A PREVIOUSLY IMPLANTED SURGICAL AORTIC VALVE, THE RIGHT CORONARY ARTERY WAS PRE-WIRED DUE TO RISK OF OCCLUSION. SUBSEQUENTLY, THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT AND THE VALVE WAS FULLY DEPLOYED. AFTER DEPLOYMENT, UNDER-EXPANSION OF THE VALVE FRAME WAS OBSERVED. CONTRAST IMAGING REVEALED A DISSECTION OF THE RIGHT CORONARY ARTERY. A STENT WAS PLACED AT THE RIGHT CORONARY ARTERY, AND A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 24 MM NON-MEDTRONIC BALLOON. AN INTERACTION BETWEEN THE BALLOON AND THE WORKING GUIDEWIRE WAS NOTED. PER REQUEST OF THE IMPLANTING PHYSICIAN, A SECOND VALVE WAS LOADED INTO A NEW DCS. THE IMPLANTED VALVE WAS "PULLED" FREE USING THE ENTANGLED BALLOON AND GUIDEWIRE. VENTRICULAR TACHYCARDIA WAS NOTED, AND DEFIBRILLATION WAS PERFORMED. AFTER DEFIBRILLATION, THE IMPLANTING PHYSICIAN PULLED THE EXPANDED VALVE DOWN TO THE FEMORAL BIFURCATION AND ATTEMPTED TO PULL VALVE THROUGH A 14 FR CONNECTING TUBE USING VASCULAR GRABBER DEVICE. THIS ATTEMPT WAS UNSUCCESSFUL; HOWEVER, THE SECOND VALVE WAS SUCCESSFULLY DELIVERED THROUGH SECONDARY LEFT FEMORAL ACCESS SITE. IT WAS NOTED THAT THE INITIAL VALVE WAS OBSTRUCTING FEMORAL ARTERY BLOOD FLOW. THE DECISION WAS MADE TO SNARE THE OBSTRUCTIVE VALVE AND DEPLOY AN ABDOMINAL GRAFT TO RESTORE BLOOD FLOW. THE PATIENT REQUIRED AN ARTIFICI AL HEART PUMP INTRA-PROCEDURALLY. DESPITE THESE INTERVENTIONS, THE PATIENT REMAINED INTUBATED AND NEVER REGAINED CONSCIOUSNESS, THE CAUSE OF DEATH WAS BELIEVED TO BE A STROKE OR BRAIN BLEED, AS INDICATED BY NEUROLOGICAL ASSESSMENT. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755189&pc=NPT | 2025587-2025-02301 | 2025-02-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-02 04:00 | NPT | EVOLUT FX DCS | Activation Failure; Patient Device Interaction Problem | Intracranial Hemorrhage; Tachycardia; Obstruction/Occlusion; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS EVFXPLUS-23 DEVICE IN A PATIENT WITH A PREVIOUSLY IMPLANTED SURGICAL AORTIC VALVE, THE RIGHT CORONARY ARTERY WAS PRE-WIRED DUE TO RISK OF OCCLUSION. SUBSEQUENTLY, THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT AND THE VALVE WAS FULLY DEPLOYED. AFTER DEPLOYMENT, UNDER-EXPANSION OF THE VALVE FRAME WAS OBSERVED. CONTRAST IMAGING REVEALED A DISSECTION OF THE RIGHT CORONARY ARTERY. A STENT WAS PLACED AT THE RIGHT CORONARY ARTERY, AND A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 24 MM NON-MEDTRONIC BALLOON. AN INTERACTION BETWEEN THE BALLOON AND THE WORKING GUIDEWIRE WAS NOTED. PER REQUEST OF THE IMPLANTING PHYSICIAN, A SECOND VALVE WAS LOADED INTO A NEW DCS. THE IMPLANTED VALVE WAS "PULLED" FREE USING THE ENTANGLED BALLOON AND GUIDEWIRE. VENTRICULAR TACHYCARDIA WAS NOTED, AND DEFIBRILLATION WAS PERFORMED. AFTER DEFIBRILLATION, THE IMPLANTING PHYSICIAN PULLED THE EXPANDED VALVE DOWN TO THE FEMORAL BIFURCATION AND ATTEMPTED TO PULL VALVE THROUGH A 14 FR CONNECTING TUBE USING VASCULAR GRABBER DEVICE. THIS ATTEMPT WAS UNSUCCESSFUL; HOWEVER, THE SECOND VALVE WAS SUCCESSFULLY DELIVERED THROUGH SECONDARY LEFT FEMORAL ACCESS SITE. IT WAS NOTED THAT THE INITIAL VALVE WAS OBSTRUCTING FEMORAL ARTERY BLOOD FLOW. THE DECISION WAS MADE TO SNARE THE OBSTRUCTIVE VALVE AND DEPLOY AN ABDOMINAL GRAFT TO RESTORE BLOOD FLOW. THE PATIENT REQUIRED AN ARTIFICI AL HEART PUMP INTRA-PROCEDURALLY. DESPITE THESE INTERVENTIONS, THE PATIENT REMAINED INTUBATED AND NEVER REGAINED CONSCIOUSNESS, THE CAUSE OF DEATH WAS BELIEVED TO BE A STROKE OR BRAIN BLEED, AS INDICATED BY NEUROLOGICAL ASSESSMENT. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VFXPLUS-23, SERIAL/LOT #: (B)(6), UBD: 25-JUL-2026, UDI #: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755609&pc=NPT | 2025587-2025-02303 | 2023-12-12 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-02 04:00 | NPT | EVOLUT FX VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Chronic Obstructive Pulmonary Disease (COPD); No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS AN ACUTE CHRONIC OBSTRUCTIVE PULMONARY DISEASE EXACERBATION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE PATIENT HAD CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AT BASELINE, WHICH CAUSED/CONTRIBUTED TO THE DEATH. ADDITIONALLY, THE DEATH WAS NOT RELATED TO THE VALVE OR VALVE IMPLANT PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755755&pc=NVN | 2017865-2025-39169 | 2025-03-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-02 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH VOLTAGE (HV) IMPEDANCE ANOMALY. ON (B)(6) 2025, THE PATIENT PRESENTED TO THE HOSPITAL TO EXTRACT THE RIGHT ATRIAL (RA) AND RV LEADS. DURING THE PROCEDURE, THE SUPERIOR VENA CAVA (SVC) WAS OCCLUDED AND THE PHYSICIAN WAS ONLY ABLE TO EXTRACT THAT RA LEAD AND A PORTION OF THE RV LEAD. THE PHYSICIAN DECIDED TO IMPLANT AN ATRIAL LEADLESS PACEMAKER (LP). THE LP IMPLANT WAS SUCCESSFUL, HOWEVER, AFTER AN HOUR THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT 'CODED.' THE PATIENT WAS TREATED WITH ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND LATER PUT ON LIFE SUPPORT. IT WAS NOTED THAT THERE WAS AN EFFUSION AT THE RIGHT ATRIUM AND INFERIOR VENA CAVA (IVC) JUNCTION. THE PHYSICIAN SUSPECTED THAT THIS WAS CAUSED BY LEAD EXTRACTION FROM THE GROIN. THE PATIENT WAS REMOVED FROM LIFE SUPPORT A FEW DAYS AFTER. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755816&pc=PNJ | 2017865-2025-39164 | 2025-03-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-02 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH VOLTAGE (HV) IMPEDANCE ANOMALY. ON (B)(6) 2025, THE PATIENT PRESENTED TO THE HOSPITAL TO EXTRACT THE RIGHT ATRIAL (RA) AND RV LEADS. DURING THE PROCEDURE, THE SUPERIOR VENA CAVA (SVC) WAS OCCLUDED AND THE PHYSICIAN WAS ONLY ABLE TO EXTRACT THAT RA LEAD AND A PORTION OF THE RV LEAD. THE PHYSICIAN DECIDED TO IMPLANT AN ATRIAL LEADLESS PACEMAKER (LP). THE LP IMPLANT WAS SUCCESSFUL, HOWEVER, AFTER AN HOUR THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT 'CODED.' THE PATIENT WAS TREATED WITH ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND LATER PUT ON LIFE SUPPORT. IT WAS NOTED THAT THERE WAS AN EFFUSION AT THE RIGHT ATRIUM AND INFERIOR VENA CAVA (IVC) JUNCTION. THE PHYSICIAN SUSPECTED THAT THIS WAS CAUSED BY LEAD EXTRACTION FROM THE GROIN. THE PATIENT WAS REMOVED FROM LIFE SUPPORT A FEW DAYS AFTER. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755853&pc=NVZ | 2017865-2025-39167 | 2025-03-12 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-02 04:00 | NVZ | ELLIPSE DR ICD | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion; No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH VOLTAGE (HV) IMPEDANCE ANOMALY. ON (B)(6) 2025, THE PATIENT PRESENTED TO THE HOSPITAL TO EXTRACT THE RIGHT ATRIAL (RA) AND RV LEADS. DURING THE PROCEDURE, THE SUPERIOR VENA CAVA (SVC) WAS OCCLUDED AND THE PHYSICIAN WAS ONLY ABLE TO EXTRACT THAT RA LEAD AND A PORTION OF THE RV LEAD. THE PHYSICIAN DECIDED TO IMPLANT AN ATRIAL LEADLESS PACEMAKER (LP). THE LP IMPLANT WAS SUCCESSFUL, HOWEVER, AFTER AN HOUR THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT 'CODED.' THE PATIENT WAS TREATED WITH ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND LATER PUT ON LIFE SUPPORT. IT WAS NOTED THAT THERE WAS AN EFFUSION AT THE RIGHT ATRIUM AND INFERIOR VENA CAVA (IVC) JUNCTION. THE PHYSICIAN SUSPECTED THAT THIS WAS CAUSED BY LEAD EXTRACTION FROM THE GROIN. THE PATIENT WAS REMOVED FROM LIFE SUPPORT A FEW DAYS AFTER. Manufacturer Narrative: THIS REPORT IS TO RETRACT REPORT SUBMITTED ON 02 APR 2025. THIS IS NOT A REPORTABLE EVENT. ALL PREVIOUSLY SUBMITTED INFORMATION SHOULD BE CORRECTED TO NOT APPLICABLE AND NULL FIELDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755872&pc=DSQ | 2916596-2025-01920 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD HEART FAILURE SYMPTOMS FOR A MONTH. THE PATIENT LACTATE DEHYDROGENASE (LDH) LEVEL HAD GRADUALLY INCREASED FROM FEBRUARY. THE PREVIOUS DAY, THE PATIENTS LDH WAS ABOUT 1000. THE FACILITY HAD A SUSPICION OF PUMP THROMBOSIS. LOG FILES WERE SENT FOR REVIEW TO RULE OUT THE POSSIBILITY OF PUMP THROMBOSIS AND CHECK FOR ROTOR DISPLACEMENT AND NOISE. THE LOG FILE CAPTURED THE POWER AND FLOW SLIGHTLY TRENDING UPWARD. NO LOW FLOW EVENTS WERE SEEN IN THE EVENT LOG FILE. OTHERWISE, THE EVENT LOG FILE CAPTURED ROUTINE EVENTS AND THERE WERE NO NOTABLE ALARM CONDITIONS. ALTHOUGH THE RECENT LOG FILE ANALYSIS SHOWED NO PROBLEMS WITH THE PUMP, AN ULTRASOUND SCAN REVEALED WHAT APPEARED TO BE A BLOOD CLOT. ON (B)(6) 2025 THE PATIENT UNDERWENT A PUMP REPLACEMENT DUE TO SUSPECTED THROMBUS. BLOOD CLOT-LIKE MATERIAL WAS FOUND TO BE ADHERING TO BOTH THE INFLOW AND OUTFLOW OF BLOOD. ON (B)(6) 2025 WHILE IN THE OPERATING ROOM, THE PATIENT DEVELOPED A CEREBRAL HEMORRHAGE AND PASSED AWAY, FOLLOWING THE PUMP EXCHANGE.Event Description: THE CAUSAL RELATIONSHIP FOR THE RENAL DYSFUNCTION THE DEVICE WAS CLEARLY RELATED AS THE RENAL FAILURE WAS DUE TO INABILITY TO MAINTAIN FLOW CAUSED BY VEGETATION IN THE OUTFLOW GRAFT. THE PATIENT'S MAXIMUM CREATININE LEVEL WAS 3.00 MG/DL. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE MODULAR CABLE WAS RETURNED DISCONNECTED FROM THE PUMP CABLE. THE SEALED OUTFLOW GRAFT WAS SECURED TO THE PUMP COVER OUTLET PORT. THE OUTFLOW GRAFT BEND RELIEF WAS ATTACHED AT THE GRAFT HARDWARE. THE APICAL CUFF WAS NOT RETURNED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES DID NOT REVEAL ANY ADHERED DEPOSITIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE RETRIEVED LVAD EVENT LOG FILE CONTAINED RELEVANT DATA FROM 10:13:55 ON (B)(6) 2025 THROUGH 01:51:27 ON (B)(6) 2025. MEAN FLOW RANGED FROM 3.1-3.6 LPM FROM THE ONSET OF THE DATA THROUGH 14:17:58 ON (B)(6) 2025 AND WAS THEN CAPTURED AT 1.6-1.8 LPM BETWEEN 14:24:22 AND 21:08:38 ON (B)(6) 2025. FINALLY, MEAN FLOW WAS CAPTURED AT 0 LPM AT 01:51:27 ON (B)(6) 2025. THE SET SPEED WAS OBSERVED TO HAVE BEEN DECREASED FROM 5000 RPM TO 3900 RPM. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. INCIDENTAL FINDINGS: THE RETRIEVED LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT LOG FILE CAPTURED EVENTS INDICATIVE OF A LOSS OF POWER TO THE PUMP THAT WOULD HAVE CAUSED THE PUMP TO STOP. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND THE HEARTMATE 3 PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, BLEEDING, AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON SYSTEM CONTROLLER ALARMS, AS WELL AS THE PROPER ACTIONS ASSOCIATED WITH THEM. THESE DOCUMENTS ALSO WARN OF EVENTS THAT MAY CAUSE THE PUMP TO STOP AND HOW TO PREVENT PUMP STOPS FROM OCCURRING. THE ¿HANDLING EMERGENCIES¿ SECTION OF THE PATIENT HANDBOOK LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE. FURTHERMORE, THE PATIENT HANDBOOK CAUTIONS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT UNDERWENT A PUMP EXCHANGE FROM (B)(6) ON (B)(6) 2025 DUE TO SUSPECTED THROMBUS AND INFECTION. ON (B)(6) 2025 WHILE IN THE OPERATING ROOM THE PATIENT DEVELOPED A CEREBRAL HEMORRHAGE AND PASSED AWAY, FOLLOWING THE PUMP EXCHANGE. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE MODULAR CABLE WAS RETURNED DISCONNECTED FROM THE PUMP CABLE. THE SEALED OUTFLOW GRAFT WAS SECURED TO THE PUMP COVER OUTLET PORT. THE OUTFLOW GRAFT BEND RELIEF WAS ATTACHED AT THE GRAFT HARDWARE. THE APICAL CUFF WAS NOT RETURNED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES DID NOT REVEAL ANY ADHERED DEPOSITIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE RETRIEVED LVAD EVENT LOG FILE CONTAINED RELEVANT DATA FROM 10:13:55 ON (B)(6) 2025 THROUGH 01:51:27 ON (B)(6) 2025. MEAN FLOW RANGED FROM 3.1-3.6 LPM FROM THE ONSET OF THE DATA THROUGH 14:17:58 ON (B)(6) 2025 AND WAS THEN CAPTURED AT 1.6-1.8 LPM BETWEEN 14:24:22 AND 21:08:38 ON (B)(6) 2025. FINALLY, MEAN FLOW WAS CAPTURED AT 0 LPM AT 01:51:27 ON (B)(6) 2025. THE SET SPEED WAS OBSERVED TO HAVE BEEN DECREASED FROM 5000 RPM TO 3900 RPM. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. INCIDENTAL FINDINGS: THE RETRIEVED LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT LOG FILE CAPTURED EVENTS INDICATIVE OF A LOSS OF POWER TO THE PUMP THAT WOULD HAVE CAUSED THE PUMP TO STOP. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, BLEEDING, AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON SYSTEM CONTROLLER ALARMS, AS WELL AS THE PROPER ACTIONS ASSOCIATED WITH THEM. THESE DOCUMENTS ALSO WARN OF EVENTS THAT MAY CAUSE THE PUMP TO STOP AND HOW TO PREVENT PUMP STOPS FROM OCCURRING. THE ¿HANDLING EMERGENCIES¿ SECTION OF THE PATIENT HANDBOOK LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE. FURTHERMORE, THE PATIENT HANDBOOK CAUTIONS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21755873&pc=DSQ | 2916596-2025-02002 | 2025-03-19 04:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS COSMETIC DAMAGE TO THE DRIVELINE. LOG FILES WERE SENT FOR REVIEW, AND THERE WERE NO UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY. THE FILE CAPTURED SEVERAL PULSATILITY INDEX (PI) EVENTS ON 17MAR2025. THE EQUIPMENT WAS WORKING AS INTENDED. THE DRIVELINE PHASE DATA GRAPH WAS NORMAL. THIS INDICATED THAT THE WIRE'S CONDUCTIVE MATERIAL WAS IN VERY GOOD CONDITION. THE DAMAGE TO THE DRIVELINE OUTER JACKET WAS COSMETIC ONLY. THERE WERE NO ALARMS ASSOCIATED WITH THE DAMAGE, AND NO IMAGES WERE AVAILABLE. IT WAS FURTHER REPORTED THAT THE COSMETIC DRIVELINE REPAIR WAS NOT PERFORMED BECAUSE THE PATIENT HAD A BRAIN BLEED ON (B)(6) 2025. THE SOURCE OF THE BLEEDING WAS A HEMORRHAGIC STROKE. THE CAUSE OF DEATH WAS THE HEMORRHAGIC STROKE, WHICH WAS NOT SURVIVABLE, AND HEARTMATE 2 SUPPORT WAS WITHDRAWN TO ALLOW FOR PALLIATIVE PASSING. THE PATIENT WAS PLACED IN COMFORT CARE AND WITHDRAWN FROM PUMP SUPPORT. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORTED SUPERFICIAL DRIVELINE DAMAGE WAS CONFIRMED THROUGH REVIEW OF THE SUBMITTED IMAGES; HOWEVER, A SPECIFIC CAUSE FOR THE DAMAGE COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED IMAGES SHOWED A TEAR IN THE OUTER JACKET. ADDITIONALLY, THERE WAS AN AREA ON THE DRIVELINE WHERE A PORTION OF THE OUTER JACKET WAS MISSING. THE BIONATE LAYER WAS VISIBLE IN SOME OF THESE AREAS AND APPEARED DISCOLORED, BUT OTHERWISE INTACT. ADDITIONALLY, A MAJORITY OF THE OUTER JACKET WAS COVERED IN TAPE. THE SUBMITTED LOG FILE CONTAINED EVENTS FROM (B)(6) 2025. NO NOTABLE ALARMS OR EVENTS, OR SIGNIFICANT FLUCTUATIONS IN PUMP PARAMETERS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AND OPERATED AT OR ABOVE THE LOW SPEED LIMIT FOR THE DURATION OF THE FILE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) AND DRIVELINE, SERIAL NUMBER (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (HMII LVAS) INSTRUCTIONS FOR USE (IFU), REV. A, AND HMII LVAS PATIENT HANDBOOK, REV. C, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS STROKE, BLEEDING, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, UNDER ¿ANTICOAGULATION¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTIONS 6 AND 8 OF THE IFU, AS WELL AS SECTIONS 4 AND 6 OF THE HMII PATIENT HANDBOOK, PROVIDE INFORMATION REGARDING HOW TO CARE FOR THE DRIVELINE. THESE DOCUMENTS ADDRESS DAMAGE DUE TO WEAR AND FATIGUE OF THE DRIVELINE AND OUTLINE INDICATIONS OF DRIVELINE DAMAGE, AS WELL AS THE HOW TO RESPOND TO SUCH EVENTS. THESE DOCUMENTS INSTRUCT THE USER TO CHECK THEIR DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS. THE PATIENT HANDBOOK ALSO INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED). THE PATIENT HANDBOOK ALSO CAUTIONS THE USER NOT TO TWIST, KINK, OR SHARPLY BEND THE DRIVELINE, WHICH MAY CAUSE DAMAGE TO THE WIRES INSIDE, EVEN IF EXTERNAL DAMAGE IS NOT VISIBLE. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINE ALL SYSTEM CONTROLLER ALARMS AND PROVIDES INFORMATION REGARDING HOW TO RESPOND TO AND TROUBLESHOOT THE ALARMS. THE PATIENT HANDBOOK ALSO CONTAINS A SECTION ON HANDLING EMERGENCIES AND FURTHER INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE PATIENT THINKS THAT, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. DURING REVIEW OF THE SUBMITTED IMAGES OF THE DRIVELINE, IT WAS ALSO NOTED THAT THE DRIVELINE APPEARED KINKED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21756089&pc=KRA | 2029214-2025-00867 | 2024-06-06 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-02 04:00 | KRA | CRAGG-MCNAMARA | Adverse Event Without Identified Device or Use Problem | Ischemia | K940634 |  | 1 | Event Description: TORRES RUIZ, I., OOI, X. Y., HARRY, L., KOKSOY, C., PALLISTER, Z. S., GILANI, R., MILLS, J. L., BAILEY, C. J., <(>&<)> CHUNG, J. (2024). MULTILEVEL THROMBOTIC OR EMBOLIC BURDEN AND ITS ROLE IN SEX-RELATED OUTCOMES IN ACUTE LIMB ISCHEMIA. JOURNAL OF VASCULAR SURGERY, 80(6), 1796¿1803. HTTPS://DOI.ORG/10.1016/J.JVS.2024.06.007 MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A TWO-CENTER RETROSPECTIVE REVIEW OVER 9 YEARS OF 170 PATIENTS TREATED FOR ACUTE LIMB ISCHEMIA (ALI). THE PURPOSE OF THE STUDY WAS TO QUANTIFY THE EFFECT OF SEX UPON AMPUTATION-FREE SURVIVAL (AFS) AFTER A PERCUTANEOUS-FIRST APPROACH FOR ALI. IT WAS NOTED THAT IN SOME CDT CASES, CRAGG-MCNAMARA CATHETERS WERE USED FOR THROMBOLYTIC INFUSION. 92 PATIENTS RECEIVED CATHETER-DIRECTED THROMBOLYSIS (CDT) ALONE, AND 35 PATIENTS WERE TREATED WITH A COMBINATION OF CDT AND ASPIRATION THROMBECTOMY. FOR PATIENTS WHO EXPERIENCED COMPLICATIONS, IT WAS NOT SPECIFIED WHICH TYPE OF TREATMENT(S) THE PATIENT RECEIVED. THE ARTICLE REPORTS SOME COMPLICATIONS RELATED TO THE PROCEDURES BUT DOES NOT SPECIFICALLY MENTION DEVICE MALFUNCTIONS OR CATHETER-RELATED COMPLICATIONS. THE COMPLICATIONS OBSERVED INCLUDED: 13 BLEEDING EPISODES (8%) 4 CASES OF ATRIAL FIBRILLATION (2%) 3 RE-THROMBOSIS/CLOT EXTENSION EVENTS (1.7%) THE ARTICLE SPECIFIES THAT 11 PATIENTS REQUIRED A MAJOR AMPUTATION FOLLOWING TREATMENT. OF THESE, 2 WERE ABOVE-THE-KNEE AMPUTATIONS (AKAS), AND 9 WERE BELOW-THE-KNEE AMPUTATIONS (BKAS). ADDITIONALLY, THE STUDY REPORTED A 30-DAY MORTALITY RATE OF 8%, WHICH EQUATES TO 13 PATIENTS OUT OF THE 170 WHO PARTICIPATED IN THE STUDY. CAUSE OF DEATH WAS NOT SPECIFIED IN THE ARTICLE. THE ARTICLE NOTED THAT FEMALES WERE MORE LIKELY THAN MALES TO HAVE POOR OUTCOME OR OUTCOME OF MORTALITY. HOWEVER, THE FEMALES IN THE STUDY GROUP ALSO HAD A HIGHER AVERAGE AGE, HAD MORE LIKELY PRESENTED WITH ATRIAL FIBRILLATION, AND MORE OFTEN HAD MULTI-LEVEL OCCLUSIONS PRIOR TO TREATMENT. Manufacturer Narrative: A2. THE REPORTED PATIENT AGE IS THE MEDIAN AGE FROM ALL PATIENTS INCLUDED IN THE STUDY. A3A. THOUGH THE MAJORITY (51%) OF PATIENTS IN THE STUDY WERE MALE, IT WAS NOTED THAT FEMALES WERE MORE LIKELY THAN MALES TO HAVE POOR OUTCOME OR OUTCOME OF MORTALITY. THEREFORE, THE REPORTED SEX IS FEMALE. B4. REPORTED EVENT DATE IS DATE THE ARTICLE WAS ACCEPTED FOR PUBLICATION. E1. THE REPORTED FACILITY IS THE FACILITY ASSOCIATED WITH THE COMMUNICATING AUTHOR. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21756133&pc=QMG | 3015053838-2025-00055 | 2025-03-04 05:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-04-02 04:00 | QMG | SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Deflation Problem; Difficult to Remove; Detachment of Device or Device Component | Foreign Body In Patient | P200039/S008 |  | 1 | Event Description: A SHOCKWAVE C2+ INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A 2.5MM LESION IN THE ANTERIOR INTERVENTRICULAR ARTERY (IVA). 80 IVL PULSES WERE DELIVERED AT 4 ATM. AFTER 8 CYCLES, THE CATHETER BROKE, AND ITS BALLOON BECAME BLOCKED IN SITU, MAKING REMOVAL IMPOSSIBLE. THERE WAS NO EXCESSIVE FORCE USED, BUT A POTENTIAL BALLOON RUPTURE OR LOSS OF PRESSURE MAY HAVE OCCURRED, PREVENTING COMPLETE DEFLATION. THE BALLOON AND PART OF THE INTERNAL WIRES REMAINED IN THE PROXIMAL IVA. THE PATIENT UNDERWENT CARDIAC SURGERY FOR A BYPASS ON THE IVA, WITH THE CATHETER BALLOON LEFT IN PLACE. THE PATIENT IS STABLE IN THE CARDIAC SURGERY REANIMATION DEPARTMENT. THE PHYSICIAN SUGGESTED THAT THE EVENT MAY HAVE BEEN CAUSED BY BALLOON BREAKAGE AND SUBSEQUENT ANCHORING OF THE BALLOON TO THE CALCIUM IN SITU.Event Description: SHOCKWAVE MEDICAL RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. THE DATE OF DEATH IS UNKNOWN, AND NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: SHOCKWAVE MEDICAL RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. THE DATE OF DEATH IS UNKNOWN, AND NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. THE CAUSE OF THE DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE CAUSE OF THE DETACHMENT COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THERE WAS NO IVL DEVICE MALFUNCTION REPORTED. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21756150&pc=DSQ | 2916596-2025-01988 | 2025-03-13 04:00 | Death | THORATEC CORPORATION | 2025-04-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21756448&pc=FMI | 3008258694-2025-00001 | 2025-03-18 04:00 | Death | WAISMED LTD. | 2025-04-02 04:00 | FMI | NIO ADULT | Mechanical Problem | Insufficient Information |  |  | 1 | Event Description: PATIENT WAS IN CARDIAC ARREST WITH CPR BEING PERFORMED. CREW ARRIVED AND TOOK OVER CPR. IV WAS ATTEMPTED AND FAILED. NIO WAS BROUGHT OUT, SITE LOCATED AND WHEN WE SPUN THE TOP ON THE NIO TO UNLOCK IT, IT WAS TIGHT TO TURN. PRESSURE PLACED WITH DOMINANT HAND, OTHER HAND STABILIZING NIO. NIO DID NOT DEPLOY, SECOND CREWMEMBER ATTEMPTED TO DEPLOY IT AND IT DID NOT WORK. SECOND NIO WAS BROUGHT OUT AND THE SAME THING HAPPENED. TOP WAS TIGHT TO SPIN AND IT DID NOT DEPLOY FOR EITHER THE RN OR AEMT. MANUAL IO WAS PLACED IN THIS PATIENT FIRST ATTEMPT. CPR CONTINUED. PATIENT DIED. ONE OF THE NIO THAT DID NOT DEPLOY WAS BROUGHT BACK TO THE STATION AND IS CURRENTLY BEING SENT BACK WITH THE PROVIDED SHIPPING LABEL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21759011&pc=NPT | 2015691-2025-02536 | 2025-03-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-02 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Difficult to Open or Close; Central Regurgitation | Low Blood Pressure/ Hypotension | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), THE PATIENT UNDERWENT A TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN A PRE-EXISTING AORTIC 29MM NON-EDWARDS (EW) TRANSCATHETER HEART VALVE (THV) WITH STENOSIS AND PARAVALVULAR LEAK. THE PATIENT WAS DEEMED HIGH RISK FOR CORONARY OCCLUSION. PRIOR TO THE S3UR IMPLANT, A BASILICA PROCEDURE OF BOTH LEAFLETS WAS SUCCESSFULLY PERFORMED, AND WIRES AND BALLOONS WERE PLACED DOWN EACH CORONARY ARTERY FOR RESCUE PURPOSES. THE S3UR VALVE WAS IMPLANTED IN THE INTENDED POSITION. POST DEPLOYMENT, THE RIGHT CORONARY ARTERY (RCA) FLOW APPEARED COMPROMISED AND THE PATIENT'S SYSTOLIC BLOOD PRESSURE MEASURED LOW (<50MMHG). THUS, CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. ADDITIONALLY, THERE WAS CENTRAL AORTIC INSUFFICIENCY (AI) ''AS IT APPEARED THE SAPIEN LEAFLETS WERE NOT FUNCTIONING FULLY, POSSIBLY BECAUSE OF THE LOW BLOOD PRESSURE AND ALSO PERHAPS DUE TO THE OUTFLOW OF THE SAPIEN BEING CONSTRAINED RESULTING IN PINWHEELING/NON-FUNCTIONING LEAFLET(S)''. THE VALVE-IN-VALVE (S3UR IN NON-EW THV) WERE POST DILATED. CPR WAS PAUSED AND A SECOND 23MM S3UR VALVE WAS IMPLANTED. THERE WAS NO HEART ACTIVITY AT THIS POINT AND THE PATIENT WAS PRONOUNCED DECEASED.Event Description: ADDITIONAL CLARIFICATION PER FCS: THE REPORTED COMPROMISED FLOW AT THE RIGHT CORONARY ARTERY WAS CONFIRMED TO BE A CORONARY OCCLUSION. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADDITIONAL INFORMATION FROM THE CLINICAL SPECIALIST. SECTIONS B5, G3, G6, H2 AND H6: HEALTH EFFECT - CLINICAL CODE HAVE BEEN UPDATED. ADDITIONS TO SECTIONS B5 AND H6: HEALTH EFFECT - CLINICAL CODE. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6: COMPONENT CODE, HEALTH EFFECT - IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. ADDITIONS TO SECTIONS H6: COMPONENT CODE, HEALTH EFFECT - IMPACT CODE AND TYPE OF INVESTIGATION. CORRECTIONS TO SECTIONS H6: INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGING WAS PROVIDED AND REVEALED THE ANNULUS AREA WAS 368.1MM^2 (21.7MM), WHICH IS WITHIN THE RECOMMENDATION RANGE FOR A 23MM SAPIEN 3 ULTRA RESILIA VALVE. IN THIS CASE, THE COMPLAINTS OF LEAFLET MOTION RESTRICTED AND DEPLOYED VALVE EXHIBITS CENTRAL REGURGITATION WERE UNABLE TO BE CONFIRMED DUE TO THE UNAVAILABILITY OF RELEVANT IMAGERY AND/OR MEDICAL REPORT. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO A REPORT OF A RESTRICTED OR NON-FUNCTIONING LEAFLET, INCLUDING MALPOSITION OF THE VALVE, IMPINGEMENT OF A LEAFLET DUE TO THE CALCIFICATION, IMPROPER EXPANSION, OR POST-DILATION OF THE IMPLANTED VALVE. ALL THESE FACTORS HAVE THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE VALVE LEAFLETS, RESULTING IN CENTRAL REGURGITATION. FOR THIS EVENT, IT WAS REPORTED THAT THE VALVE LEAFLETS WERE PINWHEELING, INDICATING THAT THE VALVE WAS UNDER-EXPANDED. AN UNDER-EXPANDED VALVE CAN HAVE SUBOPTIMAL LEAFLET COAPTATION WITH RESTRICTED LEAFLET MOTION, LEADING TO CENTRAL REGURGITATION. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (UNDER-EXPANDED THV) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THE CORONARY OCCLUSION. INVESTIGATION RESULTS INDICATE THE PATIENT WAS AT HIGH RISK FOR CORONARY OCCLUSION (PATIENT FACTORS). IN ADDITION, PRIOR TO THE TAVR PROCEDURE, A BASILICA PROCEDURE WAS PERFORMED WITH WIRES AND BALLOONS PLACED DOWN EACH CORONARY ARTERY FOR RESCUE PURPOSES. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21736068&pc=OZO | 2032227-2025-145774 | 2025-03-01 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-01 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Heart Failure/Congestive Heart Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOSPITAL ON (B)(6) 2025. THE CAUSE OF DEATH WAS HEART ATTACK AND DIABETES. THE CUSTOMER'S BLOOD GLUCOSE WAS 300 MG/DL AT THE TIME OF THE INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-386A, MMT-332A, MMT-1715KL. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER WORN THE INSULIN PUMP AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND THE CUSTOMER WAS NOT USING SENSORS. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THE PUMP PASSED THE FUNCTIONAL TESTS, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST AND DAT WITHIN SPECIFICATION AT .08750 INCHES. SUCCESSFULLY DOWNLOADED HISTORY FILES AND TRACES USING THUS. SUCCESSFULLY UPLOADED PUMP TO CARELINK. ON THE EVENT DATE OF 01-MAR-2025, THERE IS NO UNEXPECTED ALARMS/SUSPENDS RECORDED IN THE FORMATTED HISTORY FILE. ON THE EVENT DATE OF 01-MAR-2025 OF THE DAILYTOTALCOLLECTIONSTARTTIME, THE DAILYTOTALOFBASALINSULINDELIVERED = 47.425 U AND DAILYTOTALOFBOLUSINSULINDELIVERED = 00.000 U WHICH IS EQUAL TO THE DAILYTOTALOFALLINSULINDELIVERED = 47.425 U. ALL BOLUS/BASAL WERE DELIVERED IN AUTO MODE/MANUAL MODE. IN FURTHER REVIEW OF THE FORMATTED HISTORY FILE 1 WEEK PRIOR TO THE EVENT DATE OF 01-MAR-2025, THESE PUMP ERROR(S)/ALARM(S) WERE NOTED: MULTIPLE NO DELIVERY ALARM/INSULIN FLOW BLOCKED ALARM WERE FOUND ON (B)(6) 2025 DURING BASAL DELIVERY. THE INSULIN FLOW BLOCKED ALARM FUNCTIONS PROPERLY DURING THE BASIC OCCLUSION TEST, OCCLUSION TEST AND FORCE SENSOR TEST. NO UNEXPECTED INSULIN FLOW BLOCKED ALARM/NO DELIVERY ALARM NOTED DURING TESTING. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE PCBA 1, PCBA 2, FORCE SENSOR, MOTOR AND VIBRATOR ASSEMBLY NOTED. FORCE SENSOR ZERO OFFSET WITHIN SPECIFICATION (23.29 MV). THE PUMP WAS RECEIVED WITHOUT A BATTERY. THE PUMP WAS RECEIVED WITH A BATTERY CAP. NO CRACKED/DAMAGED BATTERY CAP NOTED DURING VISUAL INSPECTION. THE TEST P-CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: A CRACKED BATTERY TUBE THREADS, A SCRATCHED CASE, A CRACKED CASE BEHIND THE PUMP NEAR THE BATTERY TUBE COMPARTMENT AND A BROKEN BELT CLIP RAILS. THE PUMP PASSED ALL THE REQUIRED TESTING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21737204&pc=NAY | 2955842-2025-12413 | 2025-02-13 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-01 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Myocardial Infarction; Great Vessel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY LOBECTOMY, A BRANCH OF THE PULMONARY ARTERY IN THE LEFT LUNG WAS INJURED RESULTING IN BLEEDING AND THE PATIENT EXPIRED. THE SURGEON STATED THAT THE PROCEDURE WAS GOING WELL UNTIL MANIPULATING THE LUNG AND NEARBY ANATOMY PREPARING THE ARTERY FOR STAPLING, WHEN TOO MUCH TENSION RESULTED IN A TEAR IN THE ARTERY. THE SURGEON CLAMPED THE BLEEDING VESSEL WITH TWO DA VINCI GRASPING INSTRUMENTS. THE SURGEON INSTRUCTED THE BEDSIDE ASSISTANT TO UNDOCK AND REMOVE THE TWO INSTRUMENTS THAT WERE NOT CLAMPING THE INJURED ARTERY WHILE HE WENT TO THE BEDSIDE TO INITIATE THE CONVERSION TO OPEN THORACOTOMY; HOWEVER, THE ASSISTANT UNDOCKED ALL 4 OF THE INSTRUMENTS RESULTING IN AN EMERGENT VS. CONTROLLED CONVERSION AND AN ADDITIONAL LOSS OF ONE LITER OF BLOOD. UPON CONVERSION, THE SURGEON WAS ABLE TO CONTROL THE BLEEDING. TWO UNITS OF BLOOD WERE ADMINISTERED AND THE LOBECTOMY PROCEDURE WAS RESUMED. ABOUT 45 MINUTES LATER, AFTER THE ANESTHESIOLOGIST ADMINISTERED IV COLLOID SOLUTION, THE PATIENT EXPERIENCED BRADYCARDIA FOLLOWED BY VENTRICULAR TACHYCARDIA. RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL, INCLUDING IV VASOPRESSORS, DIRECT PADDLE SHOCKS, AND DIRECT CARDIAC MASSAGE THAT RESULTED IN AN IRREPARABLE TEAR INJURY TO THE RIGHT VENTRICLE. THERE WAS NO ALLEGATION OF ANY ISSUES WITH THE DA VINCI SYSTEM OR INSTRUMENTS, AND IT WAS REPORTED THAT THE EVENT WAS NOT DUE TO THE DA VINCI PRODUCTS. THE SURGEON STATED THAT THE HOSPITAL WILL BE INSTITUTING QUARTERLY TRAINING OF THE EMERGENCY UNDOCKING PROCESS FOR ALL SURGEONS AND OR PERSONNEL. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, LONG BIPOLAR GRASPER X2. ALL MULTI-USE INSTRUMENTS USED IN THE CASE WERE USED IN SUBSEQUENT PROCEDURES. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE EVENT. REVIEW OF THE EVENT BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF A BLEEDING INCIDENT. HOW THIS TRACTION INJURY OCCURRED IS NOT SPECIFIED. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT. BY THE SURGEON¿S ACCOUNT, THERE WERE ADDITIONAL MISCOMMUNICATIONS BETWEEN THE SURGEON AND THE SURGEON¿S TEAM THAT MAY HAVE CONTRIBUTED TO THE OVERALL OUTCOME. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21737206&pc=NAY | 2955842-2025-12407 | 2025-02-27 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-01 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Great Vessel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED NEPHROURETERECTOMY PROCEDURE, THE AORTA WAS DAMAGED RESULTING IN BLEEDING; THE PATIENT SUBSEQUENTLY EXPIRED. THE SURGEON REPORTED THAT THE DA VINCI SYSTEM DID NOT CAUSE OR CONTRIBUTE TO THE EVENT. THE SURGEON WAS DISSECTING ALONG THE ARTERIA SACRALIS SUPERIOR WITHIN THE TUMOR MASS WITH A MARYLAND BIPOLAR FORCEPS, MONOPOLAR CURVED SCISSORS AND A VESSEL SEALER EXTEND WHEN THE ARTERY WAS DAMAGED. IT WAS THOUGHT THAT THE ARTERY WAS INADVERTENTLY TORN AS IT WAS COMPLETELY WITHIN THE TUMOR MASS. THE PROCEDURE WAS CONVERTED TO OPEN TO CONTROL THE BLEEDING. DESPITE A TOTAL AND FULLY COMPENSATED ESTIMATED BLOOD LOSS WAS 2700ML, THE PATIENT WENT INTO CARDIAC ARREST WHICH REMAINED REFRACTORY FOR 55 MINUTES OF RESUSCITATION EFFORTS. THE SURGEON STATED THAT THE EVENT WAS CAUSED BY THE PATIENT'S CRITICAL CONDITION DUE TO THE ANATOMICAL EXTENT OF THE TUMOR, WHICH HAD GROWN TEN TIMES IN SIZE SINCE THE LAST CT SCAN 4 WEEKS PRIOR. THE TUMOR WAS 8×7×6 CM, LOCATED PRIMARILY ON THE LEFT KIDNEY AND EXTENDED INTO THE RETROPERITONEUM, FIXATING THE KIDNEY, ADRENAL GLAND, AND MAJOR ARTERIES (AORTA, RENAL VEIN AND ARTERY). THE CAUSE OF DEATH WAS REPORTED AS HEMORRHAGIC SHOCK AND RAPID GROWTH, DEDIFFERENTIATED UROTHELIAL CARCINOMA. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MCS, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF BLEEDING WHILE TRYING TO REMOVE A LARGE RENAL TUMOR THAT EXTENDED INTO THE RETROPERITONEUM. THE SURGEON DESCRIBES THIS TUMOR AS RAPIDLY INCREASING WITH CHALLENGING ANATOMY INCLUDING A DIFFICULT VESSEL ENCOMPASSED BY THE MASS WHEN THEY GOT INTO BLEEDING TRYING TO FREE THE TUMOR FROM THIS VESSEL. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21737214&pc=NAY | 2955842-2025-12410 | 2022-06-24 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-01 04:00 | NAY | DAVINCI XI | Insufficient Information | Septic Shock; Multiple Organ Failure; Unspecified Tissue Injury | K131861 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT¿S FAMILY IN A NEWSPAPER ARTICLE THAT IN (B)(6) 2022, DURING A DA VINCI ASSISTED PROSTATECTOMY, A PART OF THE PATIENT'S INTESTINE WAS CUT BY MISTAKE AND WAS REPAIRED WITH SUTURE. THE PATIENT BEGAN TO FEEL ILL THE FOLLOWING EVENING. THE NEXT MORNING THE SURGEON EXAMINED THE PATIENT AND SUSPECTED SEPTIC SHOCK AND MULTIPLE ORGAN FAILURE. THE PATIENT UNDERWENT AN EMERGENCY REOPERATION AND SPENT MORE THAN ONE MONTH IN THE ICU POSTOPERATIVELY. APPROXIMATELY 2 MONTHS POST THE INITIAL PROSTATECTOMY PROCEDURE, THE PATIENT BECAME SEPTIC WITH MULTIPLE ORGAN FAILURE A SECOND TIME. THE FAMILY STATED THE PATIENT HAD TO UNDERGO 5 ADDITIONAL UNSPECIFIED SURGICAL PROCEDURES IN TOTAL. THE PATIENT WAS TRANSFERRED PER FAMILY REQUEST TO ANOTHER FACILITY AND WERE INFORMED THAT HE HAD TO UNDERGO A 6TH OPERATION BECAUSE ¿THE INTESTINAL AREA THAT CAUSED THE INITIAL SEPSIS HAD TO BE REPAIRED AGAIN.¿ AFTER THIS OPERATION, THEY WERE INFORMED THAT THE DAMAGE COULD NOT BE REPAIRED. THE PATIENT EXPIRED IN (B)(6) 2022, FIVE MONTHS AFTER THE INITIAL PROSTATECTOMY PROCEDURE. THE ARTICLE DID NOT PROVIDE ANY INFORMATION OR INDICATE ANY ISSUES WITH DA VINCI PRODUCTS. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THERE ARE NO DA VINCI INSTRUMENT LOGS AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED 5 MONTHS FOLLOWING A ROBOTIC PROCEDURE WITH A NUMBER OF COMPLICATIONS AND RE-OPERATIONS OCCURRING OVER THAT TIME PERIOD. A BOWEL INJURY WAS NOTED AT THE TIME OF THE INITIAL PROCEDURE. HOW THAT INJURY OCCURRED IS NOT PROVIDED. IT IS ALSO UNKNOWN HOW THE INITIAL INJURY MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENTUAL DEMISE OF THE PATIENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21737729&pc=OZD | 1220648-2025-27126 | 2025-03-05 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT UPON EVALUATION POST IMPELLA CP PLACEMENT, THERE WAS DISTAL LIMB ISCHEMIA TO THE LEFT ARM. THE DECISION WAS MADE TO PLACE A RADIAL IPSILATERAL REPERFUSION CIRCUIT WITH A SIX FRENCH SHEATH IN THE LEFT RADIAL. GOOD FLOW WAS OBTAINED TO THE ARM. ADDITIONALLY, THERE WAS BLEEDING FROM THE INSERTION SITE. GAUZE WAS PACKED AROUND THE SITE AND TWO UNITS OF PACKED RED BLOOD CELLS WERE ADMINISTERED. BLEEDING WAS RESOLVED. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDREW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: B.2 REQUIRED INTERVENTION RECEIVED PRIOR TO DEATH WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27126. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21737755&pc=NVZ | 3008973940-2025-02307 | 2024-03-22 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-01 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Premature Discharge of Battery | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT IS DECEASED. THE PATIENT'S REPRESENTATIVE NOTED THAT THE BATTERY OF THE IMPLANTABLE PULSE GENERATOR (IPG) DID NOT MEET EXPECTATIONS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738173&pc=OZD | 1220648-2025-27129 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE INSERTION OF AN IMPELLA 5.5 DEVICE IN A PATIENT WITH POSTCARDIOTOMY CARDIOGENIC SHOCK WAS UNSUCCESSFUL WITH ANASTOMOSIS BLEEDING. THE PATIENT WOULD ULTIMATELY EXPIRE BUT NOT BEFORE IMPLANTING AN IMPELLA CP DEVICE THAT EXPERIENCED LOW PUMP FLOW. THE PATIENT WAS ADMITTED FOR A PERICARDIECTOMY AND DURING THIS PROCEDURE, THE PATIENT ARRESTED. IT WAS SUSPECTED THAT THE LEFT ANTERIOR DESCENDING ARTERY HAD OCCLUDED AND THE PATIENT WAS TAKEN FOR AN EMERGENT CORONARY ARTERY BYPASS GRAFT SURGERY. HOWEVER, THE PATIENT WOULD NOT SEPARATE FROM THE BYPASS. THE DECISION WAS MADE TO IMPLANT AN IMPELLA 5.5 DEVICE. DURING THE IMPLANT, THE IMPELLA 5.5 MADE IT INTO THE GRAFT, PAST ANASTOMOSIS BUT WAS UNABLE TO MAKE THE TURN INTO THE HEART AFTER MULTIPLE ATTEMPTS. MULTIPLE BLOOD PRODUCTS WERE USED DUE TO AXILLARY ANASTOMOSIS BLEEDING A "GREAT DEAL." WHAT WAS INFUSED AND THE NUMBER OF UNITS IS UNKNOWN. THE PHYSICIAN FELT THERE WAS A STRICTURE PREVENTING THE IMPELLA 5.5 FROM ENTERING THE HEART. THE IMPELLA 5.5 IMPLANT WAS ABORTED, AND AN IMPELLA CP WAS CALLED FOR TO ATTEMPT TO HELP SEPARATE FROM THE PATIENT FROM THE BYPASS. THE IMPELLA CP DEVICE WAS IMPLANTED WITHOUT INCIDENT. HOWEVER, IT WAS NOTED THAT THERE WERE "GREAT FLOWS AT THE BEGINNING." THE BLOOD PRESSURE WOULD INCREASE AND DECREASE WITH BLOOD PRODUCTS BEING GIVEN. HOWEVER, DESPITE "AGGRESSIVE" TREATMENT, THE PATIENTS BLOOD PRESSURE FELL DESPITE VOLUME, AND THE PATIENT EXPIRED ON THE OPERATING ROOM TABLE. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THE IMPELLA CP TO HAVE "GREAT FLOWS AT BEGINNING¿ REASONABLY SUGGESTING OPTIMAL FLOWS WERE NOT SUSTAINED RESULTING IN AN OUTCOME OF SERIOUS INJURY. HOWEVER, THE IMPELLA 5.5 BLEEDING FROM THE ANASTOMOSIS REQUIRING MULTIPLE BLOOD PRODUCTS AS THE ONSET OF SERIOUS ADVERSE COMPLICATIONS AND LIKELY TO HAVE CONTRIBUTED TO THE PATIENT¿S DEMISE. THUS, CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THE IMPELLA CP TO HAVE "GREAT FLOWS AT BEGINNING¿ REASONABLY SUGGESTING OPTIMAL FLOWS WERE NOT SUSTAINED RESULTING IN AN OUTCOME OF SERIOUS INJURY. HOWEVER, THE IMPELLA 5.5 BLEEDING FROM THE ANASTOMOSIS REQUIRING MULTIPLE BLOOD PRODUCTS AS THE ONSET OF SERIOUS ADVERSE COMPLICATIONS AND LIKELY TO HAVE CONTRIBUTED TO THE PATIENT¿S DEMISE. THUS, CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MEDICAL DEVICE REPORT FOR THE IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 06-MAR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738217&pc=OZD | 1220648-2025-27130 | 2023-04-17 04:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED VENTRICULAR TACHYCARDIA WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: " ON (B)(6) 2023: VA ECMO DECANNULATION W L CFA AND CFV REPAIRS, PROTEK DUO W OXYGENATOR INSERTED C/B VTACH IN OR WHICH WAS TREATED W AMIO". THE PATIENT RECOVERED WITH SEQUELAE. IF WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON IMPELLA 5.5 SUPPORT. THE PROCEDURAL OUTCOME WAS THE PATIENT SURVIVED SURGERY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738246&pc=OZD | 1220648-2025-27131 | 2025-03-07 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANTS REPORTED A PATIENT CAME IN WITH CHEST PAIN AND COUGH AND WAS SENT TO THE MEDICAL INTENSIVE CARE UNIT FOR PNEUMONIA TREATMENT. THE PATIENT WAS HYPOTENSIVE AND HAD SHORTNESS OF BREATH AND SO THE PATIENT WAS INTUBATED AND TROPONIN WAS INCREASED. THE PATIENT WAS FOUND TO HAVE MICROVASCULAR DISEASE AND THREE STENTS TO THE RIGHT CORONARY ARTERY. THE PATIENT, DECOMPENSATED DURING PERCUTANEOUS CORONARY INTERVENTION (PCI). IT WAS REPORTED THAT DURING SUPPORT, THE AUTOMATED IMPELLA CONTROLLER (AIC) HAD A CONTROLLER FAILURE ERROR WHEN THE PUMP STOPPED. THE PUMP STOP ALARM HAPPENED SIMULTANEOUSLY AS THE CONTROLLER ERROR. THE PUMP WAS EXPLANTED AND THE AIC WAS REMOVED. OUTCOME AT THE END OF UNIT SUPPORT WAS NOTED THAT THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE CONSOLE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - CONTROLLER FAILURE (RED ALARM) HAS BEEN COMPLETED. THE ROOT CAUSE WAS UNABLE TO BE DETERMINE DUE TO THE INABILITY TO BE REPRODUCED IN THE DEVICE ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738689&pc=LJT | 3006260740-2025-01907 | 2025-03-25 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-04-01 04:00 | LJT | UNKNOWN POWER PORT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH HEMORRHAGE. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY DEVELOPED WITH HEMORRHAGE. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED HEMORRHAGE AND PATIENT DEATH AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED HEMORRHAGE AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. G3, H6 (COMPONENT). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738706&pc=DSQ | 2916596-2025-01798 | 2022-07-01 04:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED IN THE RESEARCH ARTICLE ¿NO SURVIVAL BENEFIT IN TRANSPLANTING STABLE LVAD PATIENTS (STATUS 4)¿ THAT HEARTMATE 3 (HM3) IMPLANTATION IS ASSOCIATED WITH DEATH AND UNKNOWN ADVERSE EVENTS REQUIRING TRANSPLANTATION. ADULT PATIENTS WHO WERE SUPPORTED BY DURABLE LEFT VENTRICULAR ASSIST DEVICES (LVAD) AT THE TIME OF LISTING AND AT THE TIME OF HEART TRANSPLANT (HT) WHO WERE REGISTERED IN THE UNITED NETWORK FOR ORGAN SHARING (UNOS) DATABASE BETWEEN 18OCT2018 AND 30JUN2022 WERE INCLUDED IN THE STUDY. PATIENTS LISTED FOR DUAL ORGANS OR RE-TRANSPLANTATION WERE EXCLUDED. ONE-YEAR SURVIVAL IN THE WAITLIST AND AFTER HT WAS CALCULATED AT EACH LISTING OR TRANSPLANT STATUS. THE NET BENEFIT WITHIN EACH STATUS WAS CALCULATED IN DAYS, BY SUBTRACTING THE SURVIVAL AFTER TRANSPLANTATION FROM THE SURVIVAL IN THE WAITLIST. A TOTAL OF 2117 HM3 PATIENTS WERE LISTED FOR HT DURING THE STUDY PERIOD. OF THESE, 899 WERE NOT TRANSPLANTED WITHIN 1 YEAR OF LISTING. THE SURVIVAL AT 1 YEAR OF THESE PATIENTS WAS 37.5%, 51.9%, 87.8%, AND 94.5% FOR STATUS 1, 2, 3, AND 4 RESPECTIVELY. DURING THE SAME STUDY PERIOD, A TOTAL OF 1440 HM3 PATIENTS RECEIVED HT. THE ONE-YEAR SURVIVAL AFTER HT WAS 94%, 90.1%, 89.5%, AND 88.9% FOR PATIENTS TRANSPLANTED AS STATUS 1, 2, 3, AND 4 RESPECTIVELY. PATIENTS TRANSPLANTED AS STATUS 1 AND 2 HAD A NET POSITIVE SURVIVAL BENEFIT AFTER TRANSPLANTATION OF 168.7 AND 135.3 DAYS RESPECTIVELY (P < 0.001). PATIENTS TRANSPLANTED AS STATUS 3 HAD A NEUTRAL BENEFIT FROM TRANSPLANTATION (P = 0.684), WHEREAS PATIENTS TRANSPLANTED AS STATUS 4 HAD A NET NEGATIVE SURVIVAL BENEFIT OF 15.9 DAYS (P < 0.01). IN TERMS OF SURVIVAL, TRANSPLANTING STABLE HM3 LVAD PATIENTS LISTED AS STATUS 4 DOESN'T OFFER AN ADVANTAGE TO PATIENTS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLE ¿NO SURVIVAL BENEFIT IN TRANSPLANTING STABLE LVAD PATIENTS (STATUS 4)¿ WAS RECEIVED. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JUL2022 AS PATIENTS WERE REGISTERED IN UNOS BETWEEN 18OCT2018 AND 30JUN2022. AUTHOR INFORMATION: LOYAGA-RENDON, R., LEE, S., MANANDHAR-SHRESTHA, N., ACHARYA, D., & LEACCHE, M. (2025). NO SURVIVAL BENEFIT IN TRANSPLANTING STABLE LVAD PATIENTS (STATUS 4). JOURNAL OF CARDIAC FAILURE, 31(1), 262¿263. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.10.209. COREWELL HEALTH WEST, GRAND RAPIDS, MI. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738738&pc=NGV | 2124215-2025-19775 | 2025-03-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-01 04:00 | NGV | WATCHMAN FLX PRO | Device Dislodged or Dislocated | Low Blood Pressure/ Hypotension | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION AND PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A 24 MM WATCHMAN FLX PRO LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. THE PATIENT WAS DISCHARGED THE FOLLOWING DAY AFTER TRANSTHORACIC ECHOCARDIOGRAM. THE PATIENT WAS ADMITTED TO THE HOSPITAL FIVE (5) WEEKS POST PROCEDURE DUE TO BEING IN ATRIAL FIBRILLATION AND SHOCK. THE PATIENT WAS CARDIOVERTED AND A DIAGNOSTIC COMPUTED TOMOGRAPHY (CT) WAS PERFORMED. THE CT IMAGING SHOWED THAT THE CLOSURE DEVICE HAD EMBOLIZED AND WAS POSITIONED INFRA-RENALLY AND WAS LIMITING BLOOD FLOW TO THE LOWER EXTREMITIES. THE PATIENT HAD EMERGENCY ABDOMINAL SURGERY. THE CLOSURE DEVICE WAS SUCCESSFULLY REMOVED DURING SURGERY AND THE PATIENT WAS BROUGHT TO THE INTENSIVE CARE UNIT TO RECOVER FROM THIS EVENT. THE PATIENT DID NOT MAKE A FULL RECOVERY, AND THE PATIENT DIED POST-SURGERY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738833&pc=DXY | 2182208-2025-01641 | 2025-02-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-01 04:00 | DXY | MEDTRONIC ICD | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Infection; Sepsis; Cardiogenic Shock; Ischemia Stroke; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING INAPPROPRIATE THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND A VENTRICULAR ASSIST DEVICE (VAD). THE AUTHORS DESCRIBED PATIENT DEATHS; CAUSES OF DEATH INCLUDED RIGHT VENTRICULAR (RV) FAILURE, CARDIOGENIC SHOCK, INFECTION/SEPSIS, INTRACRANIAL HEMORRHAGE, ISCHEMIC STROKE, AND OTHER UNKNOWN CAUSES. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO DETECTION OF SUPRA VENTRICULAR TACHYCARDIA (SVT), ATRIAL FIBRILLATION (AF), ATRIAL TACHYCARDIA (AT), AND SINUS TACHYCARDIA. OTHER CAUSES OF INAPPROPRIATE THERAPY INCLUDED DOUBLE COUNTING OF R WAVES, ELECTROMAGNETIC INTERFERENCE (EMI), LEAD FRACTURE, AND OVERSENSING OF UNCLEAR ETIOLOGY. PATIENTS WERE HOSPITALIZED, UNDERWENT VT ABLATION, OR HAD REPROGRAMMING PERFORMED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: APPROPRIATE AND INAPPROPRIATE ICD SHOCKS IN PATIENTS WITH LVADS: PREVALENCE, ASSOCIATED FACTORS, AND ETIOLOGIES. HEART RHYTHM. 2025; 22:394¿402. DOI: 10.1016/J.HRTHM.2024.07.099. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738836&pc=DSQ | 3007042319-2025-00709 | 2025-02-01 05:00 | Death | HEARTWARE, INC. | 2025-04-01 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Infection; Sepsis; Cardiogenic Shock; Ischemia Stroke; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING INAPPROPRIATE THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND A VENTRICULAR ASSIST DEVICE (VAD). THE AUTHORS DESCRIBED PATIENT DEATHS; CAUSES OF DEATH INCLUDED RIGHT VENTRICULAR (RV) FAILURE, ELECTRICAL STORM, CARDIOGENIC SHOCK, INFECTION/SEPSIS, INTRACRANIAL HEMORRHAGE, ISCHEMIC STROKE, AND OTHER UNKNOWN CAUSES. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY FROM THEIR ICD DUE TO SUPRA VENTRICULAR TACHYCARDIA (SVT), ATRIAL FIBRILLATION (AF), ATRIAL TACHYCARDIA (AT), AND SINUS TACHYCARDIA. PATIENTS WERE HOSPITALIZED, UNDERWENT VT ABLATION, OR HAD REPROGRAMMING OF THEIR ICD PERFORMED. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: APPROPRIATE AND INAPPROPRIATE ICD SHOCKS IN PATIENTS WITH LVADS: PREVALENCE, ASSOCIATED FACTORS, AND ETIOLOGIES. HEART RHYTHM. 2025; 22:394¿402. DOI: 10.1016/J.HRTHM.2024.07.099 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738872&pc=OCE | 3002836339-2025-00034 |  | Death | VERICEL | 2025-04-01 04:00 | OCE | EPICEL | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | 990200 |  | 1 | Event Description: ISCHEMIC BOWEL DISEASE [INTESTINAL ISCHAEMIA]. Manufacturer Narrative: CASE REFERENCE NUMBER: (B)(4) IS A SPONTANEOUS REPORT INITIALLY RECEIVED FROM A NON-HEALTHCARE PROFESSIONAL ON 03-MAR-2025, CONCERNING A 61-YEAR-OLD MALE PATIENT WHO EXPIRED AFTER GRAFTING WITH EPICEL. ON 12-FEB-2025, QC LOT RELEASE ASSAYS (GRAFT INSPECTION QC2-027, DUAL STAIN QC2-094, ENDOTOXIN QC2-010, STERILITY PRE-RELEASE QC2-005, STERILITY FINAL PRODUCT QC2-012) AND ENVIRONMENTAL RESULTS (MANUFACTURING: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT) AND GRADE B (VIABLE AIR PARTICULATES AND NONVIABLE AIR PARTICULATES) AND QC STERILITY: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT)) WERE PERFORMED. ALL RESULTS WERE REPORTED AS ''PASS''. QC LOT RELEASE ASSAY (IMPLEMENTED 29-SEP-2019) 3T3 RESIDUAL ASSAY Q2C-136 WAS REPORTED AS LESS THAN 0.1 - PASS. THE PATIENT'S MEDICAL HISTORY AND CONCOMITANT MEDICATION DETAILS WERE NOT REPORTED. ON (B)(6) 2025, THE PATIENT WAS CRITICALLY ILL UPON ADMISSION DUE TO SIGNIFICANT BURN INJURY OF 74% TOTAL BODY SURFACE AREA (TBSA) WITH INHALATION INJURY. DUE TO THE PATIENT'S INSTABILITY, THE INITIAL EPICEL GRAFTING WAS DELAYED. ON (B)(6) 2025, THE PATIENT RECEIVED EPICEL GRAFTS (CULTURED EPIDERMAL AUTOGRAFTS) (LOT NUMBER: EE03375-31, EXPIRATION DATE: UNKNOWN) FOR THIRD DEGREE BURN. AFTER THIS PROCEDURE, THE PATIENT UNDERWENT RE-BIOPSY DUE TO LACK OF CELL EXPANSION FOR A SECOND EPICEL CASE, THAT WAS SCHEDULED FOR ON (B)(6) 2025. ON AN UNKNOWN DATE, THE PATIENT EXPERIENCED ISCHEMIC BOWEL DISEASE (PT: INTESTINAL ISCHAEMIA). ON (B)(6) 2025, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE ISCHEMIC BOWEL DISEASE. IT WAS UNKNOWN IF THE AUTOPSY WAS PERFORMED. THE REPORTER ASSESSED THE EVENT OF ISCHEMIC BOWEL DISEASE AS NOT RELATED TO EPICEL. PER REPORTER, THE PATIENT'S DEATH WAS A RESULT OF CRITICAL ILLNESS, FROM THE INITIAL BURN INJURY. ADDITIONAL INFORMATION WAS RECEIVED ON 17-MAR-2025 AND 19-MAR-2025 AND PROCESSED TOGETHER WITH THE INITIAL. ADDITIONAL INFORMATION WAS RECEIVED ON 26-MAR-2025 AND PROCESSED TOGETHER WITH THE INITIAL. NO FURTHER INFORMATION WAS PROVIDED. COMPANY COMMENT: VERICEL ASSESSED THE EVENT OF INTESTINAL ISCHAEMIA AS SERIOUS (DUE TO FATAL OUTCOME), NOT RELATED AND UNEXPECTED FOR EPICEL. THIS CASE QUALIFIES AS A 15-DAY-REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21738965&pc=DXY | 2182208-2025-01642 | 2025-02-01 05:00 | Death | MEDTRONIC, INC. | 2025-04-01 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Infection; Sepsis; Cardiogenic Shock; Ischemia Stroke; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING INAPPROPRIATE THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND A VENTRICULAR ASSIST DEVICE (VAD). THE AUTHORS DESCRIBED PATIENT DEATHS; CAUSES OF DEATH INCLUDED RIGHT VENTRICULAR (RV) FAILURE, CARDIOGENIC SHOCK, INFECTION/SEPSIS, INTRACRANIAL HEMORRHAGE, ISCHEMIC STROKE, AND OTHER UNKNOWN CAUSES. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO DETECTION OF SUPRA VENTRICULAR TACHYCARDIA (SVT), ATRIAL FIBRILLATION (AF), ATRIAL TACHYCARDIA (AT), AND SINUS TACHYCARDIA. OTHER CAUSES OF INAPPROPRIATE THERAPY INCLUDED DOUBLE COUNTING OF R WAVES, ELECTROMAGNETIC INTERFERENCE (EMI), LEAD FRACTURE, AND OVERSENSING OF UNCLEAR ETIOLOGY. PATIENTS WERE HOSPITALIZED, UNDERWENT VT ABLATION, OR HAD REPROGRAMMING PERFORMED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: APPROPRIATE AND INAPPROPRIATE ICD SHOCKS IN PATIENTS WITH LVADS: PREVALENCE, ASSOCIATED FACTORS, AND ETIOLOGIES. HEART RHYTHM. 2025; 22:394¿402. DOI: 10.1016/J.HRTHM.2024.07.099 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739231&pc=DSQ | 2916596-2025-01789 | 2025-03-02 05:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT PRESENTED WITH PSEUDOMONAS BACTEREMIA AND ELEVATED INCREASED INTERNATIONAL NORMALIZED RATIO. THEN HAD LARGE BRAIN BLEED. FAMILY WITHDREW CARE. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. NO PRODUCT WOULD BE RETURNING. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION, BLEEDING, STROKE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 OF THE IFU (UNDER "ONGOING PATIENT ASSESSMENT AND CARE") ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739318&pc=QZI | 3002648230-2025-00259 | 2025-03-06 05:00 | Death | MEDTRONIC CRYOCATH LP | 2025-04-01 04:00 | QZI | PULSESELECT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P230017 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A CARDIAC ABLATION PROCEDURE USING THE PULSED FIELD ABLATION CATHETER, THE PATIENT DIED AT HOME DAYS AFTER THE PROCEDURE. THE CAUSE OF DEATH IS UNKNOWN. THE PATIENT FELT "UNWELL" DURING THE DAYS FOLLOWING THE PROCEDURE AND REFUSED TO GO TO THE EMERGENCY ROOM. THE CASE WAS COMPLETED WITH PULSED FIELD ABLATION. THE PATIENT WAS ADMITTED FOR OVERNIGHT OBSERVATION WITH IMPROVED POST-OPERATIVE VITALS, LEADING TO DISCHARGE THE FOLLOWING DAY. THE PATIENT LATER EXPIRED.Event Description: IT WAS LATER REPORTED THAT THE PATIENT HAD A "POOR" EJECTION FRACTION PRIOR TO THE PULSED FIELD ABLATION PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: NON-MEDTRONIC UNKNOWN, PRODUCT TYPE: MAPPING SYSTEM PRODUCT ID: NON-MEDTRONIC UNKNOWN, PRODUCT TYPE: MAPPING CATHETER PRODUCT ID: NON-MEDTRONIC UNKNOWN, PRODUCT TYPE: ULTRASOUND CATHETER MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: 10FCC13; PRODUCT TYPE: SHEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739322&pc=DRA | 9612164-2025-01644 | 2025-03-06 05:00 | Death | MEDTRONIC IRELAND | 2025-04-01 04:00 | DRA | FLEXCATH CONTOUR¿ STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K232321 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A CARDIAC ABLATION PROCEDURE USING THE PULSED FIELD ABLATION CATHETER, THE PATIENT DIED AT HOME DAYS AFTER THE PROCEDURE. THE CAUSE OF DEATH IS UNKNOWN. THE PATIENT FELT "UNWELL" DURING THE DAYS FOLLOWING THE PROCEDURE AND REFUSED TO GO TO THE EMERGENCY ROOM. THE CASE WAS COMPLETED WITH PULSED FIELD ABLATION. THE PATIENT WAS ADMITTED FOR OVERNIGHT OBSERVATION WITH IMPROVED POST-OPERATIVE VITALS, LEADING TO DISCHARGE THE FOLLOWING DAY. THE PATIENT LATER EXPIRED.Event Description: IT WAS LATER REPORTED THAT THE PATIENT HAD A "POOR" EJECTION FRACTION PRIOR TO THE PULSED FIELD ABLATION PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID PSCC100; PRODUCT TYPE: CATHETER; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: NON-MEDTRONIC UNKNOWN, PRODUCT TYPE: MAPPING SYSTEM PRODUCT ID: NON-MEDTRONIC UNKNOWN, PRODUCT TYPE: MAPPING CATHETER PRODUCT ID: NON-MEDTRONIC UNKNOWN, PRODUCT TYPE: ULTRASOUND CATHETER MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739446&pc=DSQ | 2916596-2025-01854 | 2025-02-17 05:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS OF DEATH NOT PROVIDED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE WAS NOT EXPLANTED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739454&pc=OZD | 1220648-2025-27138 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT HAD NO FLOW DOWN THE RIGHT LEG PAST THE SHEATH. PEEL AWAY WAS REMOVED AND REPO SHEATH POSITIONED. PENUMBRA THROMBECTOMY WAS PERFORMED AND FLOW RESTORED. THE PHYSICIAN ALSO PLACED A REPERCUSSION CATHETER AND TWO UNITS OF PACKED RED BLOOD CELLS WERE ADMINISTERED. THE "PATIENT'S" GENERAL OUTCOME WAS NOTED TO BE CRITICALLY ILL AND GRIM. THE PUMP SUPPORTED THE PATIENT FOR TWO MORE DAYS AND THEN CARE WAS WITHDRAWN AND PATIENT SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: IMPELLA CP WITH SMARTASSIST CATHETER INSERTION SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMARTASSIST CATHETER SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739481&pc=DSQ | 2916596-2025-02084 | 2024-11-30 05:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE HOSPITAL WOULD NOT COMMUNICATE ADDITIONAL INFORMATION DUE TO PATIENT PRIVACY LAWS. THERE WERE NO DEVICE ISSUES AT THE TIME OF DEATH ACCORDING TO PATIENT RECORDS. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739487&pc=LYJ | 1644487-2025-00359 | 2024-06-06 04:00 | Death | CYBERONICS - HOUSTON | 2025-04-01 04:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THE CAUSE OF DEATH AND VNS INVOLVEMENT IS UNKNOWN. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739526&pc=DSQ | 2916596-2025-02182 | 2025-03-21 04:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THIS WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739639&pc=DTE | 2182208-2025-01648 | 2022-09-10 04:00 | Death | MEDTRONIC, INC. | 2025-04-01 04:00 | DTE | MEDTRONIC EPG | Failure to Capture; No Device Output; Pacing Problem; Battery Problem; Output below Specifications | Bradycardia; Cardiac Arrest; Asystole |  |  | 1 | Event Description: IT WAS REPORTED THAT WHEN THE PATIENT WAS IN RECOVERY POST CORONARY ARTERY BYPASS GRAFT (CABG) SURGERY, THE PATIENT'S PACEMAKER EXHIBITED LOSS OF CAPTURE, BATTERY DEPLETION, PACING BELOW THE PROGRAMMED RATE, AND STOPPED PACING ENTIRELY LEADING TO CARDIAC ARREST, BRADYCARDIA, AND THE PATIENT CODING. IT WAS NOTED THAT THE PACEMAKER WAS PROGRAMMED VVI. THE PATIENT WAS INTUBATED AND RECEIVED MULTIPLE ROUNDS OF CARDIOPULMONARY RESUSCITATION AND MEDICATION WAS ADMINISTERED. PULSE RETURNED HOWEVER THE PATIENT'S CONDITION WORSENED OVER THE NEXT TWO DAYS. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED.Event Description: IT WAS REPORTED THAT WHEN THE PATIENT WAS IN RECOVERY POST CORONARY ARTERY BYPASS GRAFT (CABG) SURGERY, THE PATIENT'S EXTERNAL PULSE GENERATOR (EPG) EXHIBITED A LOSS OF CAPTURE, BATTERY DEPLETION, PACING BELOW THE PROGRAMMED RATE, AND STOPPED PACING ENTIRELY LEADING TO CARDIAC ARREST, BRADYCARDIA, AND THE PATIENT CODING. POST SURGERY THE PATIENT HAD BECOME FULLY PACER DEPENDENT AT THE BACKUP RATE OF 45 BEATS PER MINUTE(BPM). THE EPG WAS PROGRAMMED IN VENTRICULAR DEMAND MODE(VVI) WHEN THE PATIENT'S TELEMETRY ALARMED BRADYCARDIA WITH A HEART RATE OF 36 (BPM). IT WAS REPORTED THAT THE EPG BATTERY HAD DEPLETED. THE PATIENT WAS INTUBATED AND RECEIVED MULTIPLE ROUNDS OF CARDIOPULMONARY RESUSCITATION AND MEDICATION WAS ADMINISTERED. PULSE RETURNED HOWEVER THE PATIENT'S CONDITION WORSENED OVER THE NEXT TWO DAYS. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED TWO DAYS LATER. Manufacturer Narrative: CORRECTION B3, B5, D1, D2, H6 - ANNEX A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739696&pc=MIH | 2953161-2025-01235 | 2024-04-02 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-01 04:00 | MIH | GORE® EXCLUDER®AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Rupture | P020004 |  | 1 | Event Description: ON (B)(6) 2015, THIS PATIENT UNDERWENT TREATMENT OF ABDOMINAL AORTIC ANEURYSM RUPTURE WITH GORE® EXCLUDER® AAA ENDOPROSTHESES (RLT351414/11417444). ON AN UNKNOWN DATE, A PROXIMAL TYPE I ENDOLEAK WAS DISCOVERED. ON (B)(6) 2017, THE PATIENT UNDERWENT A PLANNED REINTERVENTION AND TWO GORE® EXCLUDER® AAA AORTIC EXTENDER COMPONENTS (PLA360400/13671846, PLA360400 / (B)(6) WERE IMPLANTED PROXIMALLY TO EXTEND COVERAGE. IT IS UNKNOWN IF THE ENDOLEAK WAS RESOLVED. THE PATIENT TOLERATED THE PROCEDURE AND WAS DISCHARGED ON (B)(6) 2017. ON (B)(6) 2019, FOLLOW-UP ULTRASOUND DETERMINED THE MAXIMUM DIAMETER OF THE ANEURYSM MEASURED 78.2MM. ON (B)(6) 2024, FOLLOW-UP COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) REPORTS A PROXIMAL TYPE I ENDOLEAK NOT RELATED TO THE DEVICE OR PROCEDURE. THE MAXIMUM DIAMETER OF THE ANEURYSM MEASURED 147.0MM, THE PATIENT WAS SCHEDULED FOR A PROCEDURE AT A FUTURE DATE HOWEVER IT WAS CANCELLED BY THE PATIENT WHO EXPRESSED NO ADDITIONAL PROCEDURES WERE WANTED. ON (B)(6) 2024, THE PATIENT PRESENTED TO EMERGENCY ROOM (ER) WITH A RUPTURED ANEURYSM AND THE FAMILY OPTED FOR COMFORT CARE ONLY. THE PATIENT EXPIRED IN THE ER. Manufacturer Narrative: THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU) STATES THAT ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO ENDOLEAK. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ADDITIONAL DEVICES IMPLANTED AND/OR RELATED TO THIS EVENT: PLA360400/(B)(6) UDI (B)(4). Manufacturer Narrative: H1: CORRECTED REPORTABILITY DETERMINATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739724&pc=DYB | MW5168448 | 2025-03-19 04:00 | Death | COOK MEDICAL LLC. | 2025-04-01 04:00 | DYB | FLEXOR CHECK-FLO INTRODUCER |  | Impaired Healing; Foreign Body In Patient; Arteriosclerosis/ Atherosclerosis |  |  | 1 | Event Description: THIS 88 Y/O FEMALE HAD PLANNED ANGIOGRAM WITH ATTEMPT FOR REVASCULARIZATION TO IMPROVE WOUND HEALING OF HER RIGHT FOOT; SHE HAS SIGNIFICANT ATHEROSCLEROTIC PERIPHERAL ARTERY DISEASE. DURING THE PROCEDURE, PROVIDER WAS PERFORMING A SHEATH EXCHANGE (5/90 COOK- SHEATH IS 90CM LONG) AND THE SHEATH AND WIRE SHEARED INSIDE THE VESSEL. THERE WAS RETAINED SHEATH/WIRE AND VASCULAR WAS CONTACTED FOR EMERGENT EVALUATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21739996&pc=NSA | 3004123209-2025-00119 | 2024-12-06 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-04-01 04:00 | NSA | PACKAGE 350P, 1X PAD-PAK-03, FRENCH, 350-STR-FR-10 | Battery Problem | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE ISSUED A LOW BATTERY PROMPT DURING A PATIENT INVOLVED EVENT. HEARTSINE'S INVESTIGATION FOUND THAT THE DEVICE POWERED ITSELF OFF SEVERAL TIMES DURING THE EVENT, LEADING TO A PATIENT DEATH. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: HEARTSINE'S INVESTIGATION FOUND THE DEVICE POWERED OFF DURING USE. THE CUSTOMER EVENT FORM SUPPLIED BY THE CUSTOMER STATED THAT THE FAULT LED TO A PATIENT DEATH. THE INVESTIGATION VERIFIED THE FAULT BUT COULD NOT DUPLICATE IT. THE INVESTIGATION WAS UNABLE TO DETERMINE THE ROOT CAUSE OF THE REPORTED FAULT AND CAN ONLY SUGGEST A POTENTIAL USE ERROR. THE DEVICE SHALL BE RETAINED BY HEARTSINE AND THE CUSTOMER PROVIDED WITH A REPLACEMENT DEVICE. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. THE CUSTOMER PROVIDED HEARTSINE WITH ALL AVAILABLE INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21740009&pc=OZD | 1220648-2025-27142 | 2024-11-27 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED VENTRICLE PERFORATION DURING IMPELLA CP SUPPORT. IT WAS NOTED THAT THE PATIENT'S COURSE WAS COMPLICATED BY PERICARDIAL EFFUSION AND LEFT VENTRICLE ANTERIOR WALL PERFORATION AS WELL AS RIGHT VENTRICLE PERFORATION. THE FIELD REPRESENTATIVE NOTED THAT THIS OCCURRED UPON ASPIRATION OF A PRIOR THROMBUS BY THE PHYSICIAN. THE PATIENT WAS TAKEN TO OPERATING ROOM EMERGENTLY FOR REPAIR WITH SUTURE AND PACKING AS WELL AS TRANSCATHETER AORTIC VALVE REPLACEMENT. BLOOD LOSS WAS SIGNIFICANT AND PATIENT RECEIVED SIX UNITS OF PACKED RED BLOOD CELLS, TWO UNITS OF PLACEMENTS, TWENTY UNITS OF PROTHROMBIN COMPLEX CONCENTRATE AND DESMOPRESSIN. THE PATIENT RECOVERED WITH SEQUELAE.Event Description: THE FAMILY DID MAKE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. AS PER THE CLINICAL DETAILS, THE LEFT VENTRICLE WAS PERFORATED DURING IMPELLA USE. THE CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. Manufacturer Narrative: B2 OUTCOME REVISED TO REFLECT DEATH AND DATE OF DEATH WHICH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27142. B5 REVISED TO REFLECT THE PATIENT OUTCOME DEATH AS IT WAS ERRONEOUSLY DOCUMENTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27142. H1 TYPE OF REPORTABLE EVENT WAS ERRONEOUSLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27142. H6 REVISED TO REFLECT DEATH WHICH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27142.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21740038&pc=LDD | 3003832357-2025-000245 | 2025-03-24 04:00 | Death | SCHILLER AG | 2025-04-01 04:00 | LDD | TEMPUS LS-MANUAL | Failure to Discharge; Loose or Intermittent Connection | Unspecified Heart Problem; No Clinical Signs, Symptoms or Conditions | K200849 |  | 1 | Event Description: IT WAS REPORTED TO PHILIPS THAT THE DEVICE DIDN'T SHOW RHYTHM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21740069&pc=NPT | 2025587-2025-02257 | 2024-03-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Patient Device Interaction Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Perforation; Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE DELIVERY OF THE FX DELIVERY CATHETER SYSTEM (DCS), A PERFORATION OF THE RIGHT COMMON ILIAC ARTERY OCCURRED. AN ANGIOGRAM CONFIRMED THE PERFORATION OF THE VESSEL. THE PATIENT EXPERIENCED A BLEEDING EVENT THAT COULD NOT BE CONTROLLED, RESULTING IN A SIGNIFICANT DROP IN BLOOD PRESSURE. THE PATIENT'S ANATOMY WAS DESCRIBED AS TORTUOUS, CALCIFIED AND ANEURYSMAL, WHICH MADE THE PASSAGE OF THE DCS DIFFICULT. DESPITE THE VASCULAR SURGEON'S ATTEMPTS TO CONTROL THE BLEEDING, THE PATIENT DID NOT SURVIVE. IT WAS NOTED THAT THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 9.9 MM. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM (DCS) CAUSED THE PERFORATION. PER THE PHYSICIAN, THE DCS CANNOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO A RUPTURE OF THE COMMON ILIAC ARTERY/AORTA WHILE ATTEMPTING TO ADVANCE THE DELIVERY CATHETER SYSTEM (DCS).Event Description: IT WAS REPORTED THAT DURING THE DELIVERY OF THE FX DELIVERY CATHETER SYSTEM (DCS), A PERFORATION OF THE RIGHT COMMON ILIAC ARTERY OCCURRED. AN ANGIOGRAM CONFIRMED THE PERFORATION OF THE VESSEL. THE PATIENT EXPERIENCED A BLEEDING EVENT THAT COULD NOT BE CONTROLLED, RESULTING IN A SIGNIFICANT DROP IN BLOOD PRESSURE. THE PATIENT'S ANATOMY WAS DESCRIBED AS TORTUOUS, CALCIFIED AND ANEURYSMAL, WHICH MADE THE PASSAGE OF THE DCS DIFFICULT. DESPITE THE VASCULAR SURGEON'S ATTEMPTS TO CONTROL THE BLEEDING, THE PATIENT DID NOT SURVIVE. IT WAS NOTED THAT THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 9.9 MM. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM (DCS) CAUSED THE PERFORATION. PER THE PHYSICIAN, THE DCS CANNOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO A RUPTURE OF THE COMMON ILIAC ARTERY/AORTA WHILE ATTEMPTING TO ADVANCE THE DELIVERY CATHETER SYSTEM (DCS). Manufacturer Narrative: UPDATED B.5, D.4, H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: MEDTRONIC RECEIVED VAGUE INFORMATION DETAILING AN ADVERSE EVENT WHILE USING A MEDTRONIC DEVICE. THIS EVENT WAS REPORTED WITH THE LIMITED INFORMATION AVAILABLE AT THAT TIME. ADDITIONAL INFORMATION WAS REQUESTED AND WHEN OBTAINED, IT WAS CONFIRMED THIS EVENT IS A DUPLICATE TO AN EVENT PREVIOUSLY REPORTED. LET THIS REGULATORY REPORT SUBMIT THESE EVENTS ARE ONE AND THE SAME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: {EVOLUTFX-34}; PRODUCT TYPE: {0195-HEART VALVES}; ELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21740245&pc=FKX | 0008030665-2025-00735 | 2025-03-24 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-01 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS PERITONEAL DIALYSIS (PD) PATIENT HAD EXPIRED. THE CLINIC NURSE CONFIRMED THE PATIENT EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH WAS UNKNOWN, BUT IT WAS POSSIBLY A SEIZURE. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT'S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS IN TREATMENT ON THE LIBERTY SELECT CYCLER ON (B)(6) 2025 AT APPROXIMATELY 11:45 PM. THE PATIENT'S SPOUSE NOTED THE PATIENT APPEARED TO DISPLAY SEIZURE-LIKE ACTIVITY AND THEN WAS UNRESPONSIVE. THE PATIENT HAD A DO NOT RESUSCITATE (DNR) ORDER. THE PATIENT WAS PRONOUNCED DECEASED AT HOME AND TRANSPORTED TO THE FUNERAL HOME. THE CAUSE OF DEATH WAS DOCUMENTED AS HEART FAILURE. THE PDRN STATED THERE WERE NO ISSUES WITH THE PATIENT'S TREATMENT OR THE LIBERTY SELECT CYCLER PRIOR TO THE DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE RELATED TO THE DEATH. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN THE PERITONEAL DIALYSIS (PD) UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A LIBERTY SELECT CYCLER DEFECT OR DEFICIENCY REPORTED FOR THIS EVENT. THIS PATIENT EXPERIENCED WHAT APPEARED TO BE SEIZURE-LIKE ACTIVITY PRIOR TO BECOMING UNRESPONSIVE. THE CAUSE OF DEATH WAS ATTRIBUTED TO HEART FAILURE (HF). CARDIOVASCULAR DISEASE, INCLUDING HF, ACCOUNTS FOR APPROXIMATELY 50% OF THE DEATHS IN PATIENTS UNDERGOING DIALYSIS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21740584&pc=DSQ | 2916596-2025-01860 | 2025-02-21 05:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21740841&pc=NVZ | 3008973940-2025-02323 | 2025-03-07 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-01 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECOMPENSATED TWO DAYS AND EXPERIENCED CARDIAC ARREST AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. EVENTUALLY, THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND DECOMPENSATED TWO DAYS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. EVENTUALLY, THE PATIENT PASSED AWAY. THE PHYSICIAN SPECIFIED THAT THE CAUSE OF DEATH WAS UNKNOWN BUT IT WAS DEEMED NOT RELATED TO IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741196&pc=DSQ | 2916596-2025-01827 | 2025-03-10 04:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Fever; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THERE WAS NO CLEAR CAUSE HOWEVER REPORTEDLY THE PATIENT EXPERIENCED FEVER AND HEMODYNAMIC INSTABILITY AND LOSS OF FLOW ON BOTH HEARTMATE 3 AND IMPELLA RIGHT VENTRICULAR SUPPORT. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE WAS OPERATING FINE UNTIL THE END WHEN WITHDRAWAL OF CARE. FLOWS WERE DROPPING IN SPITE OF AN ATTEMPT TO INCREASE SPEED. THE DEVICE WOULD NOT RETURN. LOG FILE REVIEW SHOWED THAT ON (B)(6) 2025 AT 12:28:22 THE FLOW WAS AT 2.4 LITER PER MINUTE (LPM), WITH LOW FLOW ALARMS. ON (B)(6) 2025 AT 12:46:10 THE FLOW WAS AT 2.5 LPM, WITH NO ALARMS. AT 20:55:50 THE FLOW WAS AT 3.8 LPM, THE PUMP WAS FUNCTIONING AS INTENDED. ON (B)(6) 2025 AT 8:33:36 THE FLOW DROPPED TO 2.4 LPM, WITH LOW FLOW ALARMS. THE PUMP WAS RUNNING AT SET SPEED. FROM THIS TIME FORWARD, THE FLOW CONTINUED ON A DOWNWARD TREND. ON (B)(6) 2025 AT 10:32:00 THE FLOW DROPPED TO .7 LPM WITH LOW FLOW ALARMS. ON (B)(6) 2025 AT 10:33:12, THE DRIVELINE WAS DISCONNECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: ANALYSIS OF THE SUBMITTED LOG FILE CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE EVENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 09MAR2025 THROUGH 19MAR2025. LOW FLOW HAZARD ALARMS WERE CAPTURED ON 09MAR2025 THROUGH 10MAR2025. ON 10MAR2025, A SUSTAINED LOW FLOW HAZARD WAS CAPTURED. APPROXIMATELY AN HOUR LATER, AN INTENTIONAL PUMP STOP WAS CAPTURED ON 10MAR2025 AND THEN THE DRIVELINE WAS THEN DISCONNECTED, WHICH IS CONSISTENT WITH THE REPORTED WITHDRAWAL OF CARE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LVAS IFU, REVISION D, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO ADDRESSES ALL PUMP PARAMETERS. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741234&pc=OZD | 1220648-2025-27153 | 2025-03-07 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Pumping Stopped | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. PUMP STOPPED UNEXPECTEDLY DURING SUPPORT. IT WAS NOTED THAT THE AIC SCREEN RETURNED TO THE HOME SCREEN AND SHOWED A PROMPT STATING "PRESS OK TO RESUME DISPLAYED PERFORMANCE LEVEL" AT THE TIME OF THE FAILURE. THE PUMP RESTARTED ON ITS OWN BUT WAS EXPLANTED FOLLOWING PCI COMPLETION. THERE WAS NO PATIENT HARM REPORTED.Event Description: AT THE TIME OF EXPLANT, THE PATIENT WAS DEEMED HEMODYNAMICALLY STABLE AND HEMOSTASIS WAS SUCCESSFULLY ACHIEVED. UNFORTUNATELY THE PATIENT EXPIRED IN THE PROCEDURE AREA. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - PUMP STOP HAS BEEN COMPLETED. DATA ANALYSIS: CP PUMP WAS CONNECTED TO CONSOLE AT 13:38 ON (B)(6) 2025. AT 15:04 ON (B)(6) THE IMCGUI PROCESS STOPPED TRIGGERING THE WATCHDOG AND A PARTIAL RESET. THE EPC REBOOTED AND THE PUMP RETURNED HEMODYNAMIC SUPPORT THROUGH THE CONCLUSION OF THE PROCEDURE. DEVICE ANALYSIS: THE SUSPECT CONSOLE WAS TESTED BUT THE FAILURE MODE DID NOT REPRODUCE. THE IMC GUI PROCESS COULD BE FORCED TO CRASH REPRODUCING THE SYMPTOMS IN THE FIELD BUT THE PROCESS DID NOT CRASH ON ITS OWN. ROOT CAUSE: THE CAUSE OF THE UNEXPECTED RESTART DURING SUPPORT WAS THE SOFTWARE IMCGUI PROCESS CRASHING. IT IS UNKNOWN WHAT TRIGGERED THE PROCESS TO CRASH AS THE FAILURE MODE WAS NOT REPRODUCED. B1 ADVERSE EVENT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. B2 THE SELECTION DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. B5 INFORMATION REGARDING PATIENT OUTCOME; DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H6 CODES 4582 AND 2199 WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. NEW CODES WERE ADDED TO HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741238&pc=BZD | 2518422-2025-037866 | 2024-02-13 05:00 | Death | RESPIRONICS, INC. | 2025-04-01 04:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741329&pc=OZD | 1220648-2025-27155 | 2025-02-28 05:00 | Death | ABIOMED, INC. | 2025-04-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED PULSELESS ELECTRICAL ACTIVITY (PEA) CARDIAC ARREST DURING IMPELLA CP SUPPORT. THE PATIENT REQUIRED AROUND 25 MINUTES OF CARDIOPULMONARY RESUSCITATION (CPR) WITH MULTIPLE DOSES OF EPINEPHERINE, 2 AMPS OF BICARB AND ALSO CALCIUM. THE PATIENT EVENTUALLY REGAINED PULSE. TWO ADDITIONAL UNITS OF BICARB WERE ADMINISTERED WITH SUBSTANTIAL IMPROVEMENT IN HEART RATE AND STABILIZED HEMODYNAMICS NOTED. IMPELLA WAS TURNED DOWN TO P2 DURING CPR AND BACK UP TO P8.¿ THE PATIENT RECOVERED WITH SEQUELAE. AS SUPPORT CONTINUED, PEA ARREST OCCURRED AGAIN. SIX ROUNDS ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) WERE PROVIDED WHILE THE IMPELLA WAS AT P2 AND CODE WAS CALLED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741379&pc=DSQ | 2916596-2025-01758 | 2022-07-23 04:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT'S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741398&pc=NPT | 2025587-2025-02264 | 2024-03-29 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS HEART FAILURE.Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT AWAY. THE CAUSE OF DEATH WAS HEART FAILURE. Manufacturer Narrative: UPDATED: D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741650&pc=DSQ | 2916596-2025-01879 | 2019-01-01 05:00 | Death | THORATEC CORPORATION | 2025-04-01 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿ASSOCIATION OF AMIODARONE WITH RVF POST LVAD¿ THAT THE HEARTMATE 3 IS ASSOCIATED WITH RIGHT VENTRICULAR FAILURE (RVF), END ORGAN DYSFUNCTION, PROLONGED HOSPITALIZATION, ATRIAL FIBRILLATION, VENTRICULAR TACHYCARDIA, SUPRAVENTRICULAR TACHYCARDIA, AND DEATH. AMIODARONE IS A FREQUENTLY USED MEDICATION IN THIS PATIENT POPULATION AND HAS SIGNIFICANT ADVERSE EFFECTS INCLUDING PNEUMONITIS, HEPATOTOXICITY AND EXACERBATION OF ARRHYTHMIAS. OUR AIM WAS TO INVESTIGATE WHETHER AMIODARONE IS ASSOCIATED WITH RVF POST LVAD. CONSECUTIVE PATIENTS IMPLANTED WITH HEARTMATE 3 (HM3) LVAD AT OUR INSTITUTION BETWEEN 01JAN2019 AND 31MAY2023 WERE ADJUDICATED FOR THE DEVELOPMENT OF EARLY RVF POST LEFT VENTRICULAR ASSIST DEVICE (LVAD). A TOTAL OF 177 PATIENTS UNDERWENT HM3 LVAD. THE MEDIAN AGE AT IMPLANT WAS 57 YEARS (IQR 20), 68.4% MALE. EIGHTY-SEVEN (49.2%) PATIENTS RECEIVED AMIODARONE WITHIN 6 MONTHS PRIOR TO HM3, MEDIAN TOTAL DOSE WAS 15.42GRAMS (IQR28.6), MEDIAN DURATION OF THERAPY 38 DAYS (IQR 113) AND 60 PATIENTS (69%) WERE ON AMIODARONE AT TIME OF IMPLANTATION. AMIODARONE USE WAS INDICATED FOR VENTRICULAR TACHYCARDIA IN 43 OF 87 (49%) PATIENTS, ATRIAL FIBRILLATION IN 36 (41%), AND SUPRAVENTRICULAR TACHYCARDIA IN 8 (9%). EARLY RVF OCCURRED IN 89 OF 177 PATIENTS (50.3%), WITH NO DIFFERENCE IN INCIDENCE BETWEEN PATIENTS EXPOSED TO AMIODARONE VS NO EXPOSURE (47.2% VS 52.8%, P=0.29). TOTAL LENGTH OF STAY (LOS) WAS SIGNIFICANTLY HIGHER IN THE AMIODARONE GROUP (31 VS 21.3 DAYS, P=0.02), ALTHOUGH THERE WAS NO DIFFERENCE IN POST LVAD LOS (20 VS 19.5 DAYS, P=0.955). THIRTY-FIVE PATIENTS DIED WITH LVAD DURING THE FOLLOW UP PERIOD, TIME TO DEATH WAS SHORTER AND TRENDED TOWARDS SIGNIFICANCE IN THE AMIODARONE GROUP (110 VS 503 DAYS, P=0.053). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. CORNELL UNIVERSITY, ITHACA, NY; MEDICAL UNIVERSITY OF SOUTH CAROLINA, CHARLESTON, SC. O¿CONNOR, S. K., HOUSTON, B., KILIC, A., & GRIFFIN, J. (2025). ASSOCIATION OF AMIODARONE WITH RVF POST LVAD. JOURNAL OF CARDIAC FAILURE, 31(1), 345. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.10.420 MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE SERIAL NUMBERS WERE NOT REPORTED AND WERE UNABLE TO BE DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), DOCUMENT 100169835, REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741699&pc=MHX | 1218950-2025-000201 | 2025-02-28 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-04-01 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Cardiac Arrest; Insufficient Information | K143057 |  | 1 | Event Description: IT WAS REPORTED THERE WAS NO RED ALARM GENERATED AT THE TIME OF SHIFT CHANGE AND THE PATIENT PASSED AWAY. AFTER SHIFT CHANGE, THE USER ENTERED THE PATIENT ROOM AND FOUND THE PATIENT DECEASED. A REVIEW OF THE WAVEFORMS AT THE CENTRAL STATION INDICATED THERE WAS NO DATA FROM 1800-2100 ASIDE FROM A FLAT TRACE. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: IT WAS REPORTED THERE WAS NO RED ALARM GENERATED AT THE TIME OF SHIFT CHANGE AND THE PATIENT PASSED AWAY. AFTER SHIFT CHANGE, THE USER ENTERED THE PATIENT ROOM AND FOUND THE PATIENT DECEASED. ADDITIONAL INFORMATION INDICATED THE CAUSE OF DEATH WAS CARDIAC ARREST AND THIS FRAGILE PATIENT WAS BEING MONITORED FOR ALTERATION IN EJECTION FRACTION. THE PATIENT WAS FOUND UNRESPONSIVE WITH DILATED PUPILS. THE STAFF EXPECTED A RED ASYSTOLE ALARM FOR THE PATIENT'S CONDITION; HOWEVER, THERE WERE INOP ALARMS FOR ECG LEADS OFF. A PHILIPS REMOTE SERVICE ENGINEER (RSE) REVIEWED THE LOGS PROVIDED AND DETERMINED THAT THE ALARMS WERE PRESENT AT THE CENTRAL UNIT AND NO PROBLEMS WERE OBSERVED AT THE PIC IX LEVEL. IT WAS NOTED THAT AT 7:12 P.M. THERE IS AN ECG CONTACT FAULT ALARM, AND THEN A CONTACT FAULT RA, WHICH WAS ACKNOWLEDGED AT THE CENTRAL STATION. AT THIS TIME THE ECG THEREFORE WAS NO LONGER BEING RECEIVED. THE BLUE TECHNICAL ALARM HAD NO REPETITION. THE LACK OF TREND CAN BE EXPLAINED BY THE CONTACT FAULT ALARMS. ALARMS ARE INDEED PRESENT AT THE CENTRAL UNIT, FOR EXAMPLE AT 7:13 P.M. AND 9:20 P.M., WE DO NOT OBSERVE ANY PROBLEM AT THE PICIX LEVEL. AT THE MONITOR LEVEL, NO DISCONNECTION IS OBSERVED, THE MONITOR HAD REMAINED ONLINE, IT DID NOT GO OFFLINE. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED THE INFORMATION PROVIDED AND CONFIRMED THE FINDINGS OF THE RSE. THE PSE DETERMINED THAT FROM THE CLINICAL AUDIT LOG, AT 19:13, THERE IS A CONTACT ISSUE WITH THE ECG CONNECTION WHICH WAS SILENCED AT 19:18. THERE WAS NO EVIDENCE TO SHOW THAT THE ECG CONNECTION WAS RESTORED. BASED ON THE AVAILABLE INFORMATION, THE PIC IX WOULD NOT BE RECEIVING ALARMS. THERE WAS NOTHING IN THE AUDIT LOGS PAST 21:31 TO SHOW IF ANYTHING WAS RESTORED. BASED ON THE AVAILABLE INFORMATION THE PIIC IX WAS FUNCTIONING AS SPECIFIED, BUT IT IS RECOMMENDED THAT THE CUSTOMER UPGRADE TO REVISION 4.3. BASED ON THE AVAILABLE INFORMATION THE PIIC IX WAS FUNCTIONING AS SPECIFIED, BUT IT IS RECOMMENDED THAT THE CUSTOMER UPGRADE TO REVISION 4.3. A REVIEW OF THE CLINICAL AUDIT LOGS REVEALED ALARMS WERE GENERATED AT THE CENTRAL STATION WITH NO DISCONNECTION ISSUES. THERE WAS NO FUNCTIONAL ISSUE NOTED. THE MONITOR LOGS REVEALED AN ECG CONTACT FAULT ALARM, WHICH WAS ACKNOWLEDGED AT THE CENTRAL STATION, AFTER WHICH TIME THE ECG WAS NO LONGER BEING RECEIVED AS THE BLUE INOP ALARMS FOR ECG DO NOT REPEAT. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. GOOD FAITH EFFORTS ARE BEING MADE TO OBTAIN THE SERIAL NUMBER AND UDI NUMBER. REPORTING INSTITUTION PHONE #: (B)(6). REPORTER PHONE #: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21741771&pc=NPT | 2025587-2025-02266 | 2025-03-05 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR, 20 MONTHS, AND 20 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21742512&pc=NPT | 2025587-2025-02269 | 2025-02-01 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT PRO PLUS VALVE | Calcified; Gradient Increase; Insufficient Information; Device Stenosis | Stenosis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS, 3 MONTHS, AND 20 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, VALVE STENOSIS AND VALVE CALCIFICATION WERE NOTED. THE MEAN GRADIENT WAS 41 MM HG. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS, 3 MONTHS, AND 20 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, VALVE STENOSIS AND VALVE CALCIFICATION WERE NOTED. THE MEAN GRADIENT WAS 41 MM HG. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE VALVE WAS IMPLANTED AT AN APPROXIMATE DEPTH OF 1MM WITHIN THE NATIVE AORTIC ANNULUS. IT WAS ALSO NOTED THROMBUS WAS POSSIBLE, BUT NOT CONFIRMED BY THE SITE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PRIOR TO THE IMPLANT OF THE VALVE, THE PATIENT HAD A GRADIENT OF 35 MILLIMETERS OF MERCURY (MM HG). IMMEDIATELY FOLLOWING THE IMPLANT OF THE VALVE, THE GRADIENT WAS 12 MM HG. APPROXIMATELY 1 YEAR FOLLOWING IMPLANT OF THE VALVE, THE GRADIENT WAS 11 MM HG. APPROXIMATELY 4 YEARS FOLLOWING THE IMPLANT OF THE VALVE, THE GRADIENT WAS 9 MM HG. 5 YEARS AND 2 MONTHS FOLLOWING IMPLANT OF THE VALVE, THE MEAN GRADIENT WAS 41 MM HG, AND THE PATIENT HAD CALCIFIED LEAFLETS AND SEVERE STENOSIS, BUT THERE WAS NO EVIDENCE OF LEAFLET THICKENING. THE PATIENT DID NOT RECEIVE A TRANSCATHETER VALVE OR SURGERY DUE TO CONCERNS OF CORONARY OBSTRUCTION. SUBSEQUENTLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: UPDATED DATA: B1. B2. B5. H1. H6. CORRECTED DATA: B3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. CORRECTED DATA: D. SUSPECT MEDICAL DEVICE FULL UDI# WAS ERRONEOUSLY OMITTED FROM THE INITIAL MEDWATCH REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21742520&pc=GEI | 3005075853-2025-02427 | 2025-03-13 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-01 04:00 | GEI | ENSEAL X1 CURVED 25CM SHAFT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K172580 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A SEGMENTECTOMIA A THIN VESSEL WAS SEALED WITH COVE AND AFTER 24-HOURS MASSIVE BLEEDING OCCURRED. WHEN OPENING THE PATIENT, IT WAS IDENTIFIED THAT THE SEALED VESSEL WAS OPEN. THE PATIENT DIED. Manufacturer Narrative: B)(4). DATE SENT: 4/1/2025. D4 BATCH #: UNKNOWN. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE, NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT IS THE SURGEON'S EXPERIENCE WITH DEVICE? WHAT WAS THE INDICATION OF THE ORIGINAL PROCEDURE? PLEASE PROVIDE A TIMELINE OF ALL OF THE EVENTS. WERE THERE ANY ISSUES WITH HEMOSTASIS INTRA-OP OF THE ORIGINAL PROCEDURE? DID THE PATIENT HAVE ANY PRIOR CHEMOTHERAPY OR RADIATION? WAS THE PATIENT ON AN ANTI COAGULATION REGIME? WHERE WAS THE LOCATION OF THE VESSEL? WAS THE VESSEL UNDER TENSION WHEN SEALED? WAS IT A THIN-WALLED VESSEL OR THIN WIDTH VESSEL? WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? WAS THE VESSEL SKELETONIZED OR AS PART OF A BUNDLE WHEN IT WAS SEALED? WAS THE COMPLETION TONE HEARD ON THE FIRING ON THE AFFECTED VESSEL? WERE THERE ANY GENERATOR ALERT SCREENS DURING THE ORIGINAL PROCEDURE? CAN THE GENERATOR LOG BE DOWNLOADED AND PROVIDED FOR ENGINEERING REVIEW? HOW WAS BLEEDING IDENTIFIED POST PROCEDURE? WHAT WAS DONE IN ATTEMPT TO CONTROL THE BLEEDING? WAS AN AUTOPSY PERFORMED? IF YES, CAN THE AUTOPSY BE PROVIDED TO ETHICON? WHAT WAS THE CAUSE OF DEATH? IS THE SURGEON INTERESTED IN HAVING A CALL WITH ETHICON MEDICAL AND ENGINEERING? AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 4/22/2025. ADDITIONAL INFORMATION WAS OBTAINED: ACCORDING TO THE SALES REP RESPONSIBLE FOR CREATING THIS COMPLAINT, (B)(4) REFERS TO THE SAME PATIENT AS (B)(4). THEREFORE, PRODUCT CODE NSLX125C BATCH C9EL72 MUST BE ADDED TO (B)(4) MUST BE VOIDED. UPON REVIEW OF THE INFORMATION PROVIDED, IT WAS CONCLUDED THAT THIS EVENT DOES NOT MEET THE FDA DEFINED CRITERIA FOR A REPORTABLE EVENT AND IS BEING CONSIDERED NOT REPORTABLE. SEE RELATED REPORT NUMBER: 3005075853-2025-02844.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21745000&pc=OAE | 3008452825-2025-00147 | 2025-03-26 04:00 | Death | ST. JUDE MEDICAL | 2025-04-01 04:00 | OAE | TACTICATH¿ CONTACT FORCE ABLATION CATHETER, SENSOR ENABLED¿ BID CURVE D-F | Communication or Transmission Problem | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P130026 |  | 1 | Event Description: DURING THE VENTRICULAR TACHYCARDIA PROCEDURE, IT WAS FOUND THAT THE CATHETER WAS NOT RECOGNIZED BY THE SYSTEM. THE CATHETER WAS REPLACED, AND THE PROCEDURE WAS COMPLETED WITH NO ADVERSE CONSEQUENCES TO THE PATIENT. IT WAS LATER REPORTED THAT FOLLOWING THIS INITIAL PROCEDURE ON "(B)(6) 2026," THE PATIENT RETURNED TO THE ICU BECAUSE HE WAS IN SERIOUS HEALTH CONDITION. ON (B)(6) 2025, THE PATIENT WAS NOT IMPROVING SO UNDERWENT ANOTHER PROCEDURE THAT MORNING AND RETURNING TO THE ICU TO RECOVER. LATER THAT SAME DAY, THE PATIENT WAS STILL NOT IMPROVING SO UNDERWENT ANOTHER PROCEDURE THAT AFTERNOON, BUT UNFORTUNATELY THERE WAS STILL NO IMPROVEMENT IN THE PATIENT'S CONDITION AND HE RETURNED TO THE ICU. THE PATIENT DIED ON (B)(6) 2025. THE PATIENT WAS ALREADY IN A SERIOUS HEALTH CONDITION AND DID NOT SHOW IMPROVEMENT FROM THE SURGICAL INTERVENTIONS.Event Description: DURING THE VENTRICULAR TACHYCARDIA PROCEDURE, IT WAS FOUND THAT THE CATHETER WAS NOT RECOGNIZED BY THE SYSTEM. THE CATHETER WAS REPLACED, AND THE PROCEDURE WAS COMPLETED WITH NO ADVERSE CONSEQUENCES TO THE PATIENT. IT WAS LATER REPORTED THAT FOLLOWING THIS INITIAL PROCEDURE ON (B)(6) 2025, THE PATIENT RETURNED TO THE ICU BECAUSE HE WAS IN SERIOUS HEALTH CONDITION. ON (B)(6) 2025, THE PATIENT WAS NOT IMPROVING SO UNDERWENT ANOTHER PROCEDURE THAT MORNING AND RETURNING TO THE ICU TO RECOVER. LATER THAT SAME DAY, THE PATIENT WAS STILL NOT IMPROVING SO UNDERWENT ANOTHER PROCEDURE THAT AFTERNOON, BUT UNFORTUNATELY THERE WAS STILL NO IMPROVEMENT IN THE PATIENT'S CONDITION AND HE RETURNED TO THE ICU. THE PATIENT DIED ON (B)(6) 2025. THE PATIENT WAS ALREADY IN A SERIOUS HEALTH CONDITION AND DID NOT SHOW IMPROVEMENT FROM THE SURGICAL INTERVENTIONS. THE CAUSE OF DEATH IS NOT RELATED TO ANY ABBOTT DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, D9, G3, H2, H3, H6. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21746082&pc=NGV | 2124215-2025-20062 | 2025-03-27 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-01 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Difficult to Open or Close; Device Dislodged or Dislocated | Dyspnea | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION AND EXPLANTATION OCCURRED. A PRE-PROCEDURE TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) RULED OUT THROMBI PRIOR TO THE PROCEDURE, YET MEASUREMENTS OF THE OSTIUM AND DEPTH OF THE LEFT ATRIAL APPENDAGE (LAA) WERE NOT TAKEN. A LAA CLOSURE PROCEDURE WAS BEING PERFORMED USING ONLY ANGIOGRAPHY FOR IMAGING. LOCAL SEDATION AT THE GROIN AREA WAS ADMINISTERED, WITH NO FURTHER SEDATION GIVEN. TRANSSEPTAL PUNCTURE WAS PERFORMED, AND THE LAA WAS VISUALIZED FOR THE FIRST TIME VIA ANGIOGRAPHY USING CONTRAST MEDIUM IN TWO PLANES. 7,000U OF HEPARIN WAS ADMINISTERED. A NON-BOSTON SCIENTIFIC (BSC) STIFF GUIDEWIRE WAS INSERTED AND PLACED IN THE PULMONARY VEIN. THE TRANSSEPTAL SHEATH WAS THEN REPLACED WITH A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). A PIGTAIL CATHETER WAS INSERTED AND NAVIGATED INTO THE LAA. THIS WAS FOLLOWED BY VISUALIZATION IN THE ANGIOGRAM USING CONTRAST MEDIUM IN TWO PLANES. A 31MM WATCHMAN FLX CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE ADVANCED, THE FLXBALL FORMED UNDER RETRACTION, AND DEPLOYED. THE WDS WAS THEN VISUALIZED IN TWO PLANES IN THE ANGIOGRAM USING CONTRAST MEDIUM. IT WAS NOTED THE ROW OF HOOKS ON THE WDS TO THE MITRAL VALVE WAS NOT ANCHORED AND NO COMPRESSION COULD BE SEEN ANGIOGRAPHICALLY. THE CLOSURE DEVICE WAS RELEASED FROM THE CORE WIRE OF THE WDS AND THE CLOSURE DEVICE IMMEDIATELY EMBOLIZED. THE PATIENT REMAINED STABLE. THE CLOSURE DEVICE SETTLED IN THE LEFT VENTRICLE. THE PATIENT WAS TRANSPORTED TO THE NEAREST HOSPITAL FOR SURGICAL INTERVENTION BUT DEVELOPED SHORTNESS OF BREATH AND REQUIRED INTUBATION. THE CLOSURE DEVICE WAS THEN NOTED TO BE STUCK IN THE MITRAL VALVE, SO THE PATIENT WAS STABILIZED AND FINALLY TAKEN TO THE CARDIAC SURGERY DEPARTMENT WHERE OPEN HEART SURGERY WAS PERFORMED, THE CLOSURE DEVICE EXPLANTED, AND THE LAA WAS SURGICALLY LIGATED. THE PATIENT THEN DIED FROM COMPLICATIONS POST OPEN HEART SURGERY. THE DATE OF DEATH OR OFFICIAL CAUSE OF DEATH IS UNKNOWN.Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION AND EXPLANTATION OCCURRED. A PRE-PROCEDURE TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) RULED OUT THROMBI PRIOR TO THE PROCEDURE, YET MEASUREMENTS OF THE OSTIUM AND DEPTH OF THE LEFT ATRIAL APPENDAGE (LAA) WERE NOT TAKEN. A LAA CLOSURE PROCEDURE WAS BEING PERFORMED USING ONLY ANGIOGRAPHY FOR IMAGING. LOCAL SEDATION AT THE GROIN AREA WAS ADMINISTERED, WITH NO FURTHER SEDATION GIVEN. TRANSSEPTAL PUNCTURE WAS PERFORMED, AND THE LAA WAS VISUALIZED FOR THE FIRST TIME VIA ANGIOGRAPHY USING CONTRAST MEDIUM IN TWO PLANES. 7,000U OF HEPARIN WAS ADMINISTERED. A NON-BOSTON SCIENTIFIC (BSC) STIFF GUIDEWIRE WAS INSERTED AND PLACED IN THE PULMONARY VEIN. THE TRANSSEPTAL SHEATH WAS THEN REPLACED WITH A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). A PIGTAIL CATHETER WAS INSERTED AND NAVIGATED INTO THE LAA. THIS WAS FOLLOWED BY VISUALIZATION IN THE ANGIOGRAM USING CONTRAST MEDIUM IN TWO PLANES. A 31MM WATCHMAN FLX CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE ADVANCED, THE FLXBALL FORMED UNDER RETRACTION, AND DEPLOYED. THE WDS WAS THEN VISUALIZED IN TWO PLANES IN THE ANGIOGRAM USING CONTRAST MEDIUM. IT WAS NOTED THE ROW OF HOOKS ON THE WDS TO THE MITRAL VALVE WAS NOT ANCHORED AND NO COMPRESSION COULD BE SEEN ANGIOGRAPHICALLY. THE CLOSURE DEVICE WAS RELEASED FROM THE CORE WIRE OF THE WDS AND THE CLOSURE DEVICE IMMEDIATELY EMBOLIZED. THE PATIENT REMAINED STABLE. THE CLOSURE DEVICE SETTLED IN THE LEFT VENTRICLE. THE PATIENT WAS TRANSPORTED TO THE NEAREST HOSPITAL FOR SURGICAL INTERVENTION BUT DEVELOPED SHORTNESS OF BREATH AND REQUIRED INTUBATION. THE CLOSURE DEVICE WAS THEN NOTED TO BE STUCK IN THE MITRAL VALVE, SO THE PATIENT WAS STABILIZED AND FINALLY TAKEN TO THE CARDIAC SURGERY DEPARTMENT WHERE OPEN HEART SURGERY WAS PERFORMED, THE CLOSURE DEVICE EXPLANTED, AND THE LAA WAS SURGICALLY LIGATED. THE PATIENT THEN DIED FROM COMPLICATIONS POST OPEN HEART SURGERY. THE DATE OF DEATH OR OFFICIAL CAUSE OF DEATH IS UNKNOWN. IT WAS FURTHER REPORTED THE LAA WAS A WINDSOCK SHAPE. THE LAA WAS ALSO NOT SEALED UPON DEPLOYMENT OF THE WDS AS ONE SIDE OF THE DEVICE WAS NOT ANCHORED IN THE OSTIUM CAUSING A LEAK, BUT LEAK WAS UNABLE TO BE MEASURED DUE TO THE ABSENCE OF TEE.Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION AND EXPLANTATION OCCURRED. A PRE-PROCEDURE TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) RULED OUT THROMBI PRIOR TO THE PROCEDURE, YET MEASUREMENTS OF THE OSTIUM AND DEPTH OF THE LEFT ATRIAL APPENDAGE (LAA) WERE NOT TAKEN. A LAA CLOSURE PROCEDURE WAS BEING PERFORMED USING ONLY ANGIOGRAPHY FOR IMAGING. LOCAL SEDATION AT THE GROIN AREA WAS ADMINISTERED, WITH NO FURTHER SEDATION GIVEN. TRANSSEPTAL PUNCTURE WAS PERFORMED, AND THE LAA WAS VISUALIZED FOR THE FIRST TIME VIA ANGIOGRAPHY USING CONTRAST MEDIUM IN TWO PLANES. 7,000U OF HEPARIN WAS ADMINISTERED. A NON-BOSTON SCIENTIFIC (BSC) STIFF GUIDEWIRE WAS INSERTED AND PLACED IN THE PULMONARY VEIN. THE TRANSSEPTAL SHEATH WAS THEN REPLACED WITH A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). A PIGTAIL CATHETER WAS INSERTED AND NAVIGATED INTO THE LAA. THIS WAS FOLLOWED BY VISUALIZATION IN THE ANGIOGRAM USING CONTRAST MEDIUM IN TWO PLANES. A 31MM WATCHMAN FLX CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE ADVANCED, THE FLXBALL FORMED UNDER RETRACTION, AND DEPLOYED. THE WDS WAS THEN VISUALIZED IN TWO PLANES IN THE ANGIOGRAM USING CONTRAST MEDIUM. IT WAS NOTED THE ROW OF HOOKS ON THE WDS TO THE MITRAL VALVE WAS NOT ANCHORED AND NO COMPRESSION COULD BE SEEN ANGIOGRAPHICALLY. THE CLOSURE DEVICE WAS RELEASED FROM THE CORE WIRE OF THE WDS AND THE CLOSURE DEVICE IMMEDIATELY EMBOLIZED. THE PATIENT REMAINED STABLE. THE CLOSURE DEVICE SETTLED IN THE LEFT VENTRICLE. THE PATIENT WAS TRANSPORTED TO THE NEAREST HOSPITAL FOR SURGICAL INTERVENTION BUT DEVELOPED SHORTNESS OF BREATH AND REQUIRED INTUBATION. THE CLOSURE DEVICE WAS THEN NOTED TO BE STUCK IN THE MITRAL VALVE, SO THE PATIENT WAS STABILIZED AND FINALLY TAKEN TO THE CARDIAC SURGERY DEPARTMENT WHERE OPEN HEART SURGERY WAS PERFORMED, THE CLOSURE DEVICE EXPLANTED, AND THE LAA WAS SURGICALLY LIGATED. THE PATIENT THEN DIED FROM COMPLICATIONS POST OPEN HEART SURGERY. THE DATE OF DEATH OR OFFICIAL CAUSE OF DEATH IS UNKNOWN. IT WAS FURTHER REPORTED THE LAA WAS A WINDSOCK SHAPE. THE LAA WAS ALSO NOT SEALED UPON DEPLOYMENT OF THE WDS AS ONE SIDE OF THE DEVICE WAS NOT ANCHORED IN THE OSTIUM CAUSING A LEAK, BUT LEAK WAS UNABLE TO BE MEASURED DUE TO THE ABSENCE OF TEE. Manufacturer Narrative: H6 EVALUATION METHOD CODE ADDED: ANALYSIS OF INFORMATION PROVIDED BY USER/THIRD PARTY H6 EVALUATION CONCLUSION CODE UPDATED FROM CMC-NO PROBLEM DETECTED TO FAILURE TO FOLLOW INSTRUCTIONS MEDICAL MEDIA WAS PROVIDED TO AID IN THE INVESTIGATION AND WAS REVIEWED BY MEMBERS OF THE BSC MEDICAL SAFETY TEAM. THE MEDIA REVIEW REPORT CONCLUDED: FLUOROSCOPY STILL IMAGES AND VIDEO CLIPS WERE SUBMITTED FOR REVIEW. THE MEDIA SHOWED CONTRAST INJECTION OF THE LAA IN SEVERAL PROJECTIONS AND MEASUREMENTS OF THE LAA (ABOUT 26-28MM, IF CORRECT WOULD YIELD LOW OUTSIDE OF RANGE COMPRESSION). A PIGTAIL CATHETER WAS ADVANCED INTO THE LAA. LATER THE WATCHMAN FLX CLOSURE DEVICE WAS DEPLOYED IN THE LAA. THE SHAPE OF THE CLOSURE DEVICE SHOWED NO COMPRESSION WITH A FLAT DISTAL FACE. THERE SEEMED TO BE A SHOULDER AND INSUFFICIENT ENGAGEMENT ON THE INFERIOR SIDE OF THE CLOSURE DEVICE. THE CLOSURE DEVICE EMBOLIZATION TO THE LEFT VENTRICLE WAS NOT SHOWN IN THE MEDIA PROVIDED. IN CONCLUSION, THIS WAS A FLUOROSCOPY ONLY PROCEDURE AND RELEASE CRITERIA WERE NOT MET. THE CLOSURE DEVICE SIZING AND POSITIONING WAS INADEQUATE AND CONTRIBUTED TO THE ACUTE DEVICE EMBOLIZATION. Manufacturer Narrative: B5: INFORMATION ADDED. H6 DEVICE CODE ADDED: DIFFICULT TO OPEN OR CLOSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21746656&pc=NPT | 2025587-2025-01814 | 2025-01-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT FX DCS | Fluid/Blood Leak; Failure to Advance; Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC HEART VALVE PROCEDURE, USING TRANSFEMORAL ACCESS, THE 29 MM VALVE WAS ATTEMPTED TO BE IMPLANTED BUT INFOLDED. THE VALVE WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. THE PATIENT BECAME HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY. A SECOND 29 MM VALVE WAS PREPPED BUT WOULD NOT CROSS THE PATIENT¿S VALVE DURING CARDIOPULMONARY RESUSCITATION (CPR). CPR CONTINUED FOR APPROXIMATELY 20 MINUTES. A 23 MM NON-MEDTRONIC (EDWARDS SAPIEN 3 ULTRA) TRANSCATHETER VALVE WAS PREPPED AND IMPLANTED WITHOUT ISSUE OR LEAKS. THE NON-MEDTRONIC VALVE IMPLANT WAS CONSIDERED A BAILOUT IMPLANT AS IT TOOK PLACE DURING AN ACTIVE CODE EVENT WITH ACTIVE RESUSCITATION EFFORTS FROM THE PREVIOUS FAILED IMPLANT ATTEMPTS. THE PATIENT DID NOT RECOVER AND EXPIRED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVOLUTFX-29 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0195-HEART VALVES; CONTINUATION OF H10: MW5165894, MW5165895 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21746883&pc=DTB | 2649622-2025-08838 | 2025-03-07 05:00 | Death | MPRI | 2025-04-01 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND DECOMPENSATED TWO DAYS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GE NERATOR (IPG) SYSTEM. EVENTUALLY, THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND DECOMPENSATED TWO DAYS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GE NERATOR (IPG) SYSTEM. EVENTUALLY, THE PATIENT PASSED AWAY. THE PHYSICIAN SPECIFIED THAT THE CAUSE OF DEATH WAS UNKNOWN BUT IT WAS DEEMED NOT RELATED TO IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21746944&pc=NVN | 2649622-2025-08839 | 2025-03-07 05:00 | Death | MPRI | 2025-04-01 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND DECOMPENSATED TWO DAYS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GE NERATOR (IPG) SYSTEM. EVENTUALLY, THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND DECOMPENSATED TWO DAYS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GE NERATOR (IPG) SYSTEM. EVENTUALLY, THE PATIENT PASSED AWAY. THE PHYSICIAN SPECIFIED THAT THE CAUSE OF DEATH WAS UNKNOWN BUT IT WAS DEEMED NOT RELATED TO IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748016&pc=BZD | 2518422-2025-103579 | 2025-01-30 05:00 | Death | RESPIRONICS, INC. | 2025-04-01 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Cancer | K113068 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED THAT A REMSTAR PRO C-FLEX+, SYS ONE 60 SRS, DOM DEVICE PATIENT HAS PASSED AWAY. IT WAS NOT SUGGESTED THAT THE DEVICE CAUSED THE DEATH. THE PATIENT'S SPOUSE CONTACTED THE MANUFACTURER REGARDING RETURNING THE DEVICE. THE PATIENT WAS IN HOSPICE FOR ONE MONTH BEFORE HE DIED. THE CAUSE OF DEATH WAS LISTED AS AN UNSPECIFIED CANCER. FURTHER INFORMATION REGARDING THE REPORTED DEATH IS NECESSARY. THE MANUFACTURER'S INVESTIGATION IS ONGOING, AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748242&pc=DXY | 2182208-2025-01653 | 2025-03-02 05:00 | Death | MEDTRONIC, INC. | 2025-04-01 04:00 | DXY | MEDTRONIC ICD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) HAD "STOPPED" WORKING NINE DAYS PRIOR TO THE PATIENT 'S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748593&pc=NVY | 2017865-2025-38986 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | NVY | ISOFLEX OPTIM LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748594&pc=DXY | 2017865-2025-38985 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | DXY | ENDURITY¿ CORE DR, MRI (US) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748595&pc=NVN | 2017865-2025-38987 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748599&pc=DYB | 2017865-2025-38988 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | DYB | PEEL-AWAY INTRODUCER 14 CM SHEATH INTRODUCER KIT, 7F | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K894431 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748600&pc=DYB | 2017865-2025-38989 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | DYB | PEEL-AWAY INTRODUCER 14 CM SHEATH INTRODUCER KIT, 7F | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K894431 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748602&pc=DSA | 2017865-2025-39062 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | DSA | DISPOSABLE CABLE EXTENSIONS | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748604&pc=DSA | 2017865-2025-39064 | 2025-03-11 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-01 04:00 | DSA | DISPOSABLE CABLE EXTENSIONS | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. THE DAY AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748827&pc=NPT | 2025587-2025-02275 |  | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Perivalvular Leak; Insufficient Information; Patient Device Interaction Problem | Hemorrhage/Bleeding; Myocardial Infarction; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TWO CASES WITH ACCELERATED STRUCTURAL VALVE DETERIORATION IN SYSTEMIC SCLEROSIS PATIENTS FOLLOWING TRANSCATHETER AORTIC VALVE REPLACEMENT. CASE 1: A 64-YEAR-OLD WOMAN PRESENTED WITH SYMPTOMS OF HEART FAILURE SECONDARY TO SEVERE STRUCTURAL VALVE DETERIORATION OF A MEDTRONIC 26-MM COREVALVE BIOPROSTHETIC VALVE, WHICH HAD BEEN IMPLANTED FOUR YEARS PREVIOUSLY DUE TO TREAT SEVERE SYMPTOMATIC AORTIC STENOSIS. TRANSTHORACIC ECHOCARDIOGRAPHY CONFIRMED PROSTHETIC VALVE STENOSIS WITH PEAK GRADIENTS OF 80 MMHG. MULTIDETECTOR COMPUTERIZED TOMOGRAPHY FOUND BASAL CALCIFICATION OF THE BIOPROSTHETIC AORTIC VALVE LEAFLETS WHICH IMPEDED THEIR MOVEMENT. SUBSEQUENTLY THE PATIENT UNDERWENT A VALVE-IN-VALVE PROCEDURE WITH IMPLANTATION OF A MEDTRONIC 26-MM EVOLUT R VALVE. POST-PROCEDURE TRANSTHORACIC ECHOCARDIOGRAPHY DEMONSTRATED OPTIMAL HEMODYNAMICS WITH MINIMAL PARAVALVULAR LEAK. DURING THE POSTOPERATIVE PERIOD, THE PATIENT RECEIVED BLOOD TRANSFUSIONS AND UNDERWENT A GASTROINTESTINAL ENDOSCOPY WHICH DID NOT IDENTIFY A SOURCE OF BLEEDING. THE PATIENT WAS DISCHARGED 5 DAYS POST-PROCEDURE. FIVE MONTHS LATER, THE PATIENT DEVELOPED ACUTE CORONARY SYNDROME AND SUBSEQUENTLY PASSED AWAY. CASE 2: A 72-YEAR-OLD WOMAN UNDERWENT TRANSCATHETER AORTIC VALVE REPLACEMENT DUE TO SEVERE SYMPTOMATIC AORTIC STENOSIS WITH IMPLANTATION OF A MEDTRONIC 26-MM EVOLUT PRO BIOPROSTHETIC VALVE. APPROXIMATELY FOUR YEARS LATER THE PATIENT PRESENTED WITH NEW-ONSET EXERTIONAL DYSPNEA DUE TO SEVERE STRUCTURAL VALVE DETERIORATION WITH PROSTHETIC VALVE STENOSIS. TRANSTHORACIC ECHOCARDIOGRAPHY MEASURED A PEAK GRADIENT OF 95 MMHG. MULTIDETECTOR COMPUTERIZED TOMOGRAPHY REVEALED A SEVERELY CALCIFIED BIOPROSTHETIC VALVE WITH REDUCED LEAFLET MOBILITY. SUBSEQUENTLY THE PATIENT UNDERWENT A VALVE-IN-VALVE PROCEDURE WITH IMPLANTATION OF A NON-MEDTRONIC 23-MM BIOPROSTHETIC VALVE. TRANSTHORACIC ECHOCARDIOGRAPHY INDICATED NORMAL VALVULAR FUNCTION WITH TRIVIAL PARAVALVULAR LEAK. THE PATIENT WAS DISCHARGED THREE DAYS LATER POST-PROCEDURE. AN ECHOCARDIOGRAPHIC EVALUATION AT SIX MONTHS POST-PROCEDURE CONFIRMED NORMAL HEMODYNAMICS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: AVIDAN ET AL. ACCELERATED STRUCTURAL VALVE DETERIORATION IN SYSTEMIC SCLEROSIS PATIENTS FOLLOWING TRANSCATHETER AORTIC VALVE REPLACEMENT: A CASE SERIES. EUR HEART J CASE REP. 2025 FEB 5;9(2):YTAF060. DOI: 10.1093/EHJCR/YTAF060. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748901&pc=DQX | 2025587-2025-02277 | 2025-03-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A SMALL, FRAIL LEFT VENTRICULAR CAVITY, THE DELIVERY CATHETER SYSTEM WAS ADVANCED OVER A MEDTRONIC GUIDEWIRE THROUGH THE AORTIC ARCH; HOWEVER, MOVEMENT OF THE GUIDEWIRE IN THE LEFT VENTRICLE WAS NOTED. SUBSEQUENTLY, SEVERE HYPOTENSION OCCURRED. AN ECHOCARDIOGRAM WAS PERFORMED AND REVEALED A PERICARDIAL EFFUSION. THE EFFUSION EVOLVED INTO CARDIAC TAMPONADE, CAUSED BY THE LEFT VENTRICULAR PERFORATION DUE TO THE GUIDEWIRE. AN EMERGENT PERICARDIOCENTESIS WAS PERFORMED AND APPROXIMATELY 1 LITER OF ARTERIAL BLOOD WAS DRAINED. THE PATIENT WAS STABILIZED AND TRANSFERRED TO THE INTENSIVE CARE UNIT; HOWEVER, BEGAN BLEEDING AGAIN, AND SUBSEQUENTLY DIED APPROXIMATELY ONE HOUR LATER.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A SMALL, FRAIL, LEFT VENTRICULAR CAVITY, THE DELIVERY CATHETER SYSTEM (DCS) WAS ADVANCED OVER A MEDTRONIC GUIDEWIRE THROUGH THE AORTIC ARCH; HOWEVER, MOVEMENT OF THE GUIDEWIRE IN THE LEFT VENTRICLE WAS NOTED. SUBSEQUENTLY, SEVERE HYPOTENSION OCCURRED. AN ECHOCARDIOGRAM WAS PERFORMED AND REVEALED A PERICARDIAL EFFUSION. THE EFFUSION EVOLVED INTO CARDIAC TAMPONADE, CAUSED BY THE LEFT VENTRICULAR PERFORATION DUE TO THE GUIDEWIRE. AN EMERGENT PERICARDIOCENTESIS WAS PERFORMED AND APPROXIMATELY 1 LITER OF ARTERIAL BLOOD WAS DRAINED. THE PATIENT WAS STABILIZED AND TRANSFERRED TO THE INTENSIVE CARE UNIT; HOWEVER, BEGAN BLEEDING AGAIN, AND SUBSEQUENTLY DIED APPROXIMATELY ONE HOUR LATER. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE GUIDEWIRE WAS NOT INTRODUCED INTO THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED THERE, BUT WAS PLACED IN THE APEX OF THE LEFT VENTRICLE USING A PIGTAIL CATHETER. THE GUIDEWIRE POSITION WAS MONITORED THROUGHOUT THE PROCEDURE, AND THE GUIDEWIRE WAS REPOSITIONED DURING THE PROCEDURE. IT WAS NOT TWISTED OR TORQUED, BUT THERE WAS DIFFICULTY IN POSITIONING THE WIRE CORRECTLY. THE POSITION OF THE GUIDEWIRE WAS VISUALLY MONITORED USING TWO PROJECTIONS TO ENSURE STABILITY, THE POSITION WAS NOT MANIPULATED WITHOUT FLUOROSCOPIC GUIDANCE, AND NO UNUSUAL RESISTANCE WAS FELT. DURING THE PROGRESSION OF THE DCS INTO THE AORTIC ARCH, THE WIRE WAS MOVED FORWARD TOWARDS THE APEX OF THE LEFT VENTRICLE. THE VALVE WAS SUCCESSFULLY IMPLANTED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS VENTRICULAR FAILURE AFTER PERICARDIAL EFFUSION, AND BOTH THE VALVE AND THE DCS CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED CORRECTED DATA: G3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVPROPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6)2025; PRODUCT ID D-EVPROP23-29 (LOT: 0012497854); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE PRODUCT ID L-EVPROP23-29 (LOT: 0012402712); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21748904&pc=NPT | 2025587-2025-02278 | 2025-03-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-01 04:00 | NPT | EVOLUT PRO PLUS DCS | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A SMALL, FRAIL, LEFT VENTRICULAR CAVITY, THE DELIVERY CATHETER SYSTEM WAS ADVANCED OVER A MEDTRONIC GUIDEWIRE THROUGH THE AORTIC ARCH; HOWEVER, MOVEMENT OF THE GUIDEWIRE IN THE LEFT VENTRICLE WAS NOTED. SUBSEQUENTLY, SEVERE HYPOTENSION OCCURRED. AN ECHOCARDIOGRAM WAS PERFORMED AND REVEALED A PERICARDIAL EFFUSION. THE EFFUSION EVOLVED INTO CARDIAC TAMPONADE, CAUSED BY THE LEFT VENTRICULAR PERFORATION DUE TO THE GUIDEWIRE. AN EMERGENT PERICARDIOCENTESIS WAS PERFORMED AND APPROXIMATELY 1 LITER OF ARTERIAL BLOOD WAS DRAINED. THE PATIENTWAS STABILIZED AND TRANSFERRED TO THE INTENSIVE CARE UNIT; HOWEVER, BEGAN BLEEDING AGAIN, AND SUBSEQUENTLY DIED APPROXIMATELY ONE HOUR LATER.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A SMALL, FRAIL, LEFT VENTRICULAR CAVITY, THE DELIVERY CATHETER SYSTEM (DCS) WAS ADVANCED OVER A MEDTRONIC GUIDEWIRE THROUGH THE AORTIC ARCH; HOWEVER, MOVEMENT OF THE GUIDEWIRE IN THE LEFT VENTRICLE WAS NOTED. SUBSEQUENTLY, SEVERE HYPOTENSION OCCURRED. AN ECHOCARDIOGRAM WAS PERFORMED AND REVEALED A PERICARDIAL EFFUSION. THE EFFUSION EVOLVED INTO CARDIAC TAMPONADE, CAUSED BY THE LEFT VENTRICULAR PERFORATION DUE TO THE GUIDEWIRE. AN EMERGENT PERICARDIOCENTESIS WAS PERFORMED AND APPROXIMATELY 1 LITER OF ARTERIAL BLOOD WAS DRAINED. THE PATIENT WAS STABILIZED AND TRANSFERRED TO THE INTENSIVE CARE UNIT; HOWEVER, BEGAN BLEEDING AGAIN, AND SUBSEQUENTLY DIED APPROXIMATELY ONE HOUR LATER. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE GUIDEWIRE WAS NOT INTRODUCED INTO THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED THERE BUT WAS PLACED IN THE APEX OF THE LEFT VENTRICLE USING A PIGTAIL CATHETER. THE GUIDEWIRE POSITION WAS MONITORED THROUGHOUT THE PROCEDURE, AND THE GUIDEWIRE WAS REPOSITIONED DURING THE PROCEDURE. IT WAS NOT TWISTED OR TORQUED, BUT THERE WAS DIFFICULTY IN POSITIONING THE WIRE CORRECTLY. THE POSITION OF THE GUIDEWIRE WAS VISUALLY MONITORED USING TWO PROJECTIONS TO ENSURE STABILITY, THE POSITION WAS NOT MANIPULATED WITHOUT FLUOROSCOPIC GUIDANCE, AND NO UNUSUAL RESISTANCE WAS FELT. DURING THE PROGRESSION OF THE DCS INTO THE AORTIC ARCH, THE WIRE WAS MOVED FORWARD TOWARDS THE APEX OF THE LEFT VENTRICLE. THE VALVE WAS SUCCESSFULLY IMPLANTED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS VENTRICULAR FAILURE AFTER PERICARDIAL EFFUSION, AND BOTH THE VALVE AND THE DCS CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED CORRECTED DATA: G3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. FLUOROSCOPY VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. THERE IS EVIDENCE OF A BALLOON ACROSS THE AORTIC VALVE FOR BALLOON AORTIC VALVULOPLASTY, BUT DILATATION WAS NOT CAPTURED ON FLUOROSCOPIC. DURING VALVE DEPLOYMENT, THE VALVE DEPTH WAS APPROXIMATELY 3MM ON THE NON-CORONARY CUSP IN THE CUSP OVERLAP VIEW AND GUIDEWIRE SEEMS TO HAVE PROLAPSED IN THE LEFT VENTRICLE. MEDTRONIC RECOMMENDS MAINTAINING CONTROL OF GUIDEWIRE THROUGHOUT PROCEDURE TO ENSURE STABLE DEPLOYMENT AND PREVENT INJURY TO THE VENTRICLE WALL: FOR ALL WIRES WITH A TRANSITION POINT, ENSURE TRANSITION POINT IS HELD ABOVE THE APEX AND POINTING AWAY FROM THE VENTRICLE WALL; MAINTAIN STRICT FLUOROSCOPIC SURVEILLANCE OF THE GUIDEWIRE IN THE LEFT VENTRICLE. THE VALVE WAS DEPLOYED, AND FINAL ANGIOGRAM REVEALS A WELL EXPANDED BIOPROSTHESIS IN A GOOD POSITION WITHOUT EVIDENCE OF AORTIC INSUFFICIENCY/PARAVALVULAR LEAK. IT WAS REPORTED THAT HYPOTENSION OCCURRED, AND AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION THAT WORSENED TO CARDIAC TAMPONADE AND SUBSEQUENT DEATH. IMAGES OF THE PERICARDIAL EFFUSION WERE NOT PROVIDED FOR REVIEW. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT PRO+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: DEATH, CARDIAC TAMPONADE, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION). UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVPROPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; PRODUCT ID GWBC30 (LOT: 92106996); NON-IMPLANTABLE DEVICE PRODUCT ID L-EVPROP23-29 (LOT: 0012402712); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21749265&pc=NAY | 2955842-2025-12757 | 2025-03-07 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-01 04:00 | NAY | DAVINCI XI | Insufficient Information | Bowel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED PARTIAL NEPHRECTOMY PROCEDURE, THE PATIENT WAS FOUND TO HAVE A SMALL BOWEL PERFORATION AND EXPIRED ON POST-OPERATIVE DAY (POD) 2. THE PARTIAL NEPHRECTOMY SURGERY WAS REPORTED TO BE DIFFICULT FOR UNSPECIFIED REASONS AND THE OPERATION TIME WAS LONG. THERE WAS INTRA-OPERATIVE BLEEDING WITH AN ESTIMATED BLOOD LOSS OF 1,600 ML; HOWEVER, THE PROCEDURE WAS COMPLETED WITHOUT OTHER ISSUES. NO DA VINCI PRODUCT ISSUE WAS NOTICED INTRA-OPERATIVELY. ON POD1, THE PATIENT COMPLAINED OF ABDOMINAL PAIN. EMERGENCY SURGERY FOUND A SMALL INTESTINE PERFORATION. THE SURGEON STATED IT WAS NOT KNOWN WHEN THE PERFORATION OCCURRED. THE PATIENT EXPIRED THE FOLLOWING DAY. THE SURGEON REPORTED THAT THE RELATIONSHIP BETWEEN THE EVENT AND THE DA VINCI PRODUCTS USED IS UNCLEAR. IT WAS FURTHER STATED BY THE SURGEON THAT THERE IS A POSSIBILITY THAT IT WAS PERFORATED WHEN THE FIRST PORT WAS INSERTED, OR THAT THE SMALL BOWEL WAS CAUGHT WHEN THE WOUND WAS CLOSED, AND ¿IT IS POSSIBLE THAT SOMETHING HAPPENED WITH DA VINCI DURING THE PROCEDURE¿. THE OFFICIAL CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 2 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE NEEDLE DRIVERS X2, PROGRASP FORCEPS, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BOWEL PERFORATION THAT WAS NOT DETECTED AT THE TIME OF THE ORIGINAL SURGERY. ALTHOUGH A TROCAR INJURY AT THE BEGINNING OF THE SURGERY IS SUSPECTED TO BE A POTENTIAL CAUSE OF THE INJURY, THE EXACT CAUSE OF THE INJURY AND HOW IT HAPPENED IS UNKNOWN. THE PATIENT EVENTUALLY DIED POSSIBLY SECONDARY TO THIS INJURY. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21889404&pc=OZP | 2032227-2025-159673 | 2019-03-25 04:00 | Death | MEDTRONIC MINIMED | 2025-04-23 04:00 | OZP | UNKNOWN NGP | Excess Flow or Over-Infusion; Adverse Event Without Identified Device or Use Problem | Hypoglycemia |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER'S FATHER PASSED AWAY IN 2019 DUE TO INSULIN PUMP OVERDOSED. PATIENT BLOOD GLUCOSE WAS UNKNOWN AT THE TIME OF INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNK_NGP, UNOMEDICAL, UNK_RESERVOIR. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, PATIENT WORN INSULIN PUMP AT THE TIME OF PASSING. PATIENT WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF REPORTED EVENT AND THERE WAS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. PRODUCT RETURN REQUESTED FOR UNK_NGP BUT CUSTOMER DECLINED TO RETURN, OR IS UNABLE TO RETURN. . NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE PROVIDED DUE TO REGIONAL PRIVACY REGULATIONS. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21889785&pc=OZD | 1220648-2025-27749 | 2025-04-03 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Bacterial Infection | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED AN IMPELLA 5.5 PUMP TO SUPPORT A PATIENT WHO NEEDED CARE ESCALATED FROM AN INTRA-AORTIC BALLOON PUMP AND TRANSFERRED TO A TERTIARY CENTER. THE PUMP WAS SUPPORTING WHEN ANTIBIOTICS WERE NEEDED DUE TO METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS BACTEREMIA GRAM POSITIVE LABS. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS: ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21889905&pc=GDW | 3005075853-2025-02884 | 2025-03-31 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-23 04:00 | GDW | CIRCULAR MECH STD 29MM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Burn(s); Cardiac Arrest | K983536 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SURGERY FOR DEEP INFILTRATIVE ENDOMETRIOSIS WITH DISCOID RESECTION OF THE COLON. DURING THE SURGERY, A HARMONIC AND A CIRCULAR STAPLER WERE USED. (B)(6), THE PATIENT WAS SCHEDULED FOR EMERGENCY SURGERY DUE TO AN UNDETECTED LEAK IN THE SIGMOID COLON. HE SUFFERED CARDIAC ARREST FOUR TIMES DURING THE PROCEDURE, BUT STILL MANAGED TO SURVIVE. HE DIED POSTOPERATIVELY. THE CAUSE OF THE LEAKAGE WAS ATTRIBUTED TO A THERMAL LESION IN THE SIGMOID COLON THAT OCCURRED DURING THE MANIPULATION AND RELEASE OF STRUCTURES DURING THE FIRST SURGERY. Manufacturer Narrative: (B)(4). DATE SENT: 4/23/2025. D4 BATCH #: UNKNOWN. INVESTIGATION SUMMARY: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: ON WHAT ANATOMICAL LOCATION WAS THE THERMAL INJURY IDENTIFIED POST-OPERATIVELY? WHAT WAS THE DISTANCE BETWEEN THE ANASTOMOTIC SITE AND THERMAL LESION? DOES THE SURGEON BELIEVE THERE WAS AN ALLEGED DEFICIENCY OF THE CDH29B THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE THERMAL INJURY, POST-OPERATIVE LEAK, AND/OR PATIENT¿S DEATH? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 6/20/2025. ADDITIONAL INFORMATION: ¿ THE SALES REP WAS PRESENT IN THE FIRST SURGERY THAT WAS PROGRAMED AND AS PART OF AN EVENT. THIS SURGERY WAS COMPLETED SUCCESSFULLY WITHOUT ANY ISSUES OR COMPLICATIONS. ¿ THE LEAK HAPPENED 2-3 DAYS AFTER THE SURGERY AND THE REP WAS NOT INFORMED, SHE KNEW ABOUT IT. SHE KNEW ABOUT IT BY CHANCE BECAUSE SHE SAW THE SCHEDULED SURGERY IN A DIFFERENT HOSPITAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21890002&pc=MIH | MW5169434 |  | Death | MEDTRONIC VASCULAR | 2025-04-23 04:00 | MIH | VALIANT CAPTIVIA | Device Dislodged or Dislocated; Migration | Cardiac Arrest; Chest Pain; Pain |  |  | 1 | Event Description: MY HUSBAND HAD THE SURGICAL PROCEDURE TO PLACE AN AORTIC STENT ON (B)(6) 2018, THAT NIGHT HE EXPERIENCED A SUDDEN PAIN IN HIS CHEST. THEN REPORTED ACID REFLUX ISSUES AND BLOODWORK BEGAN TO DECLINE. ON (B)(6) HE BEGAN DIALYSIS AFTER HIS KIDNEYS SHUT DOWN. 3:00 AM ON (B)(6), HE SUFFERED CARDIAC ARREST AND DIED AT 10:18PM ON (B)(6) BECAUSE THE STENT HAD BECOME DISLODGED AND TRAVELED TO HIS LOWER ABDOMEN. HE DIED SLOWLY, PAINFULLY AND IN A HOSPITAL. HE SPENT HIS LAST WEEK IN A HOSPITAL, WHEN HE HAD A WIFE AND PETS AND FAMILY AND FRIENDS THAT LOVED HIM DEARLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21890574&pc=DTZ | 2184009-2025-00570 | 2025-03-18 04:00 | Death | PERFUSION SYSTEMS | 2025-04-23 04:00 | DTZ | FUSION HOLLOW FIBER OXYGENATOR | Fluid/Blood Leak | Cardiac Arrest; Hemorrhage/Bleeding | K240666 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A CARDIOPULMONARY BYPASS PROCEDURE OF LESS THAN A 6 HOUR DURATION, USING THE FUSION OXYGENATOR, STRAW-COLORED FLUID AND FOAM WERE OBSERVED COMING FROM THE GAS EXHAUST OF THE OXYGENATOR. THE CUSTOMER DESCRIBED THIS AS PLASMA LEAKAGE. CIRCULATORY ARREST OCCURRED DURING THE PROCEDURE, AND THE OXYGENATOR WAS CHANGED OUT AND REPLACED WITH ANOTHER OXYGENATOR. THE REST OF THE ORIGINAL CUSTOM TUBING PACK WAS USED. THERE WAS MINIMAL BLOOD LOSS OF 250MLS, AND NO BLOOD TRANSFUSION WAS REQUIRED. NO DAMAGE WAS SEEN TO ANY COMPONENT OR PACKAGING. THE PROCEDURE WAS AN EMERGENCY DURING THE NIGHT. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CLINICIAN DESCRIBED THE LEAK AS A PLASMA LEAK. THE PROCEDURE COMMENCED ON (B)(6) 2025 AND CONTINUED INTO (B)(6) 2025. ALL ACTIVATED COAGULATION TIMES (ACT) WERE WITHIN PROTOCOL. FOR THE SECOND CARDIOPUMONARY BYPASS (CPB) RUN, ALL ACTS WERE GREATER THAN 1000 SECONDS. THE CIRCULATORY ARREST WAS NOT PLANNED. THE CLINICIAN COOLED THE PATIENT TO 18¿ IN A CONTROLLED TIMEFRAME TO USE CIRCULATORY ARREST TO REPLACE THE OXYGENATOR. IT WAS AN ELECTIVE AORTIC VALVE REPLACEMENT (AVR). THERE WERE NO KNOWN COAGULATION CONDITIONS. THE PRIMING INCLUDED PLASMALYTE, GELOFUSINE, MANNITOL AND 10 ,000IU HEPARIN. IT WAS AN INITIAL BYPASS RUN OF 253 MINUTES. BYPASS WAS TERMINATED AND THE PUMP WAS THEN CIRCULATING FOR 2 HOURS WHEN THE SURGEON DEEMED IT NECESSARY TO RECOMMENCE CPB. AT 30 MINUTES INTO THE SECOND CPB RUN, THE OXYGENATOR STARTED TO LEAK AND AN INCREASED SWEEP SPEED WAS REQUIRED. THE SURGEON STATED THAT THE OPERATION WAS LIKELY TO CONTINUE FOR SEVERAL HOURS, THEREFORE, THE DECISION WAS MADE TO COOL DOWN AND CHANGE THE OXYGENATOR OUT. THE TOTAL SECOND CPB WAS 294 MINUTES. IT TOOK 30 MINUTES TO COOL, CHANGE OUT THE OXYGENATOR AND CONTINUE BYPASS WITH A NEW OXYGENATOR. THE PATIENT PASSED AWAY IN THE OPERATING THEATRE FROM SURGICAL COMPLICATIONS. THE CLINICIAN BELIEVED THAT IT WAS HIGHLY UNLIKELY THAT CHANGING OUT THE OXYGENATOR AFFECTED THE OUTCOME, HOWEVER, THEY CAN NOT CONFIRM IT WITH 100% CERTAINTY. THE PRESSURES AND FLOWS WITHIN THE CPB CIRCUIT INCLUDING THE OXYGENATOR WERE ALL WITHIN NORMAL PARAMETERS. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CAUSE OF DEATH HAD NOT BEEN IDENTIFIED BUT IT WAS DESCRIBED AS COMPLICATIONS ARISING FROM SURGERY. IT WAS NOT THOUGHT TO BE DIRECTLY LINKED TO THE OXYGENATOR. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE FIRST PUMP RUN WAS A STANDARD CASE. THE SECOND PUMP RUN WAS REQUIRED DUE TO COMPLICATIONS WITH THE PATIENT REQUIRING A SECOND PROCEDURE. THE SECOND PUMP RUN WAS STARTED WITH THE ORIGINAL CIRCUIT AND OXYGENATOR. AFTER 30 MINUTES OF THE SECOND PUMP RUN, THE OXYGENATOR SHOWED SIGNS OF FAILING. THE PATIENT WAS COOLED AND THE OXYGENATOR WAS CHANGED. THE OPERATION WAS COMPLETED WITH THE SECOND OXYGENATOR WHICH HAD NO ISSUES. THEY WERE NOT ON PATIENT BYPASS FOR 120 MINUTES. THEY WERE ON STANDBY RECIRCULATING THROUGH THE RECIRCULATION LINE WITH DILUTE BLOOD/ HARTMANNS MIX. THE START OF SECOND BYPASS RUN WAS 30 MINUTES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21890802&pc=MIH | 9612164-2025-02026 | 2025-02-16 05:00 | Death | MEDTRONIC IRELAND | 2025-04-23 04:00 | MIH | ENDURANT STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿FROM PLANNING TO PRACTICE: IMPACT OF ACHIEVED PROXIMAL SEALING ZONE IN ENDOVASCULAR ANEURYSM REPAIR (EVAR)¿. THE TIME FRAME OF THIS STUDY WAS OVER A THIRTEEN YEAR PERIOD. ENDURANT STENT GRAFTS WERE IMPLANTED IN ALL PATIENTS (N= 275). PATIENT MORTALITY WAS REPORTED OF WHICH 5 DEATHS WERE DEEMED ANEURYSM RELATED. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿FROM PLANNING TO PRACTICE: IMPACT OF ACHIEVED PROXIMAL SEALING ZONE IN ENDOVASCULAR ANEURYSM REPAIR¿ ACCARINO, G.; SILVERIO, A.; BELLINO, M.; FURGIUELE, S.; FIMIANI, M.; SICA, M.; DE VUONO, F.; FORNINO, G.; TURCHINO, D.; ACCARINO, G.; ET AL JOURNAL OF CLINICAL MEDICINE. 2025, 14, 1309. HTTPS://DOI.ORG/10.3390/ JCM14041309. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21890816&pc=OZD | 1220648-2025-27760 | 2025-04-02 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED PRIOR TO THE IMPELLA PLACEMENT THE PATIENT FAILED TO COME OFF THE PUMP TWICE POST CORONARY ARTERY BYPASS GRAFT SURGERY AND THE PATIENT FAILED TO COME OFF THE PUMP TWICE TO THE INTRA-AORTIC BALLOON PUMP MECHANICAL CIRCULATORY SUPPORT (MCS). THE DECISION WAS MADE TO IMPLANT AN IMPELLA 5.5 DEVICE FOR MCS. SUBSEQUENTLY, A LEFT VENTRICLE PERFORATION WAS OBSERVED. EMERGENT CLOSURE OF PERFORATION WAS COMPLETED. THE IMPELLA 5.5 DEVICE WAS PULLED INTO THE AORTA AND WAS PLACED IN SURGICAL MODE FOR INTERVENTION WITH CROSS CLAMP AND ON CPB. THE PATIENT HAD FRIABLE VESSELS PER THE OPERATING ROOM STAFF. INTRA PROCEDURAL COMPLICATIONS WITH SEWING GRAFTS ON AND THOUGHT TO HAVE HAD SOME SORT OF CONNECTIVE TISSUE DISORDER. THE LEFT VENTRICLE PERFORATION WAS UNDETERMINED IF CAUSED BY IMPELLA OR OTHER PROCEDURAL COMPLICATIONS. THE PERFORATION WAS CLOSED AS NOTED BUT THE PATIENT EXPIRED AFTER APPROXIMATELY TWO HOURS OF SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891028&pc=DTZ | 8010762-2025-0000181 |  | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-23 04:00 | DTZ | HLS SET | Pressure Problem | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN THE US. IT WAS REPORTED THAT DURING PATIENT TREATMENT ISSUES WERE PRESENT WITH ERRONEOUS PRESSURE READINGS (ARTERIAL, INTERNAL AND DELTA) FROM AN HLS-SET. THE AFFECTED HLS SET WAS DISCARDED. ON 2025-04-16 THE INFORMATION WAS RECEIVED THAT THE FAMILY HAD DECIDED TO DISCONTINUE TREATMENT. HOWEVER, STATE NOW THERE IS NO INFORMATION ON DEATH OR SERIOUS INJURY. AS A PRESSURE READING ISSUE CAN LEAD TO A PUMP STOP, IF THE INTERVENTION IS SET BY THE USER, A REPORT IS REQUIRED. COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID # (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER. Manufacturer Narrative: NEW INFORMATION HAS BEEN RECEIVED ON 2025-05-14 BY THE SSU/ CUSTOMER, THAT THE INVOLVED PATIENT PASSED AWAY. THE PATIENT¿S FAMILY DECIDED TO DISCONTINUE THE TREATMENT, RESULTING IN THE PATIENT PASSING AWAY. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER. Manufacturer Narrative: THE EVENT OCCURRED IN THE US. IT WAS REPORTED THAT DURING PATIENT TREATMENT ISSUES WERE PRESENT WITH ERRONEOUS PRESSURE READINGS (ARTERIAL, INTERNAL AND DELTA) FROM AN HLS-SET. NEW INFORMATION HAS BEEN RECEIVED ON 2025-05-14 BY THE SSU/CUSTOMER, THAT THE INVOLVED PATIENT PASSED AWAY. THE PATIENT¿S FAMILY DECIDED TO DISCONTINUE THE TREATMENT, RESULTING IN THE PATIENT PASSING AWAY. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS NOT AVAILABLE FOR TECHNICAL INVESTIGATION AS IT WAS DISCARDED BY THE CUSTOMER. THUS, THE EXACT ROOT CAUSE REMAINS UNKNOWN. HOWEVER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-06-02 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES INACCURATE PRESSURE READINGS FROM A PRE-PRIMED HLS SET USED DURING ECMO THERAPY. THE CHIEF OF PERFUSION AT THE TREATING INSTITUTION, SUGGESTED DURING A CONFERENCE THAT FLUID MAY HAVE ENTERED THE WHITE CONNECTOR OF THE HLS SET PRIOR TO THERAPY, POSSIBLY CONTRIBUTING TO THE MALFUNCTION. THE TEAM REPORTEDLY ATTEMPTED TO CLEAN THE AREA BEFORE USE. ECMO THERAPY BEGAN ON (B)(6) 2025. NO ISSUES WERE NOTED DURING INITIAL PRIMING WITH PLASMALYTE, AND NO DEPOSITS OR CRYSTAL FORMATIONS WERE SEEN AT THE SENSOR PORT WHICH MAY BE IN CONFLICT WITH THE PREVIOUS SUGGESTION OF FLUID INTRUSION IN THE PORT THE ABSENCE OF VISIBLE DEPOSITS OR RESIDUE AT THE TIME OF SETUP MAY SUGGEST THAT INTRUSION MAY HAVE OCCURRED AFTER ECMO INITIATION AS INTRUSION WAS NOT IDENTIFIED DURING INITIAL DEVICE INSPECTION. WHILE THE CIRCUIT APPEARED CLEAN AT INITIATION, THE DISCOVERY OF BLOOD UPON DISASSEMBLY RAISES QUESTIONS ABOUT WHETHER CONTAMINATION MAY HAVE DEVELOPED SOMETIME DURING SUPPORT. THE REPORTED PRESSURE ISSUE WAS FIRST OBSERVED ON (B)(6), RESULTING IN THE REPLACEMENT OF THE SENSOR PROBE (CABLE) ON (B)(6). THE REPORTED BEHAVIOR SUGGESTS THE ROOT CAUSE MAY LIE IN EITHER THE CABLE ITSELF OR THE INTERNAL SENSOR PORT OF THE HLS SET. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER FOUND NEAR THE POST-OXYGENATOR MEMBRANE. LAB VALUES ON FEBRUARY 1 INDICATED STABLE PARAMETERS. THE HLS SET WAS NOT REPLACED, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES REPORTEDLY CORRELATED WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING BLOOD SAMPLING. NO REDUCTION IN BLOOD FLOW WAS OBSERVED. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER IDENTIFIED NEAR THE POST-OXYGENATOR MEMBRANE AT THE END OF THERAPY. LABORATORY VALUES OBTAINED ON FEBRUARY 1 SHOWED A HEMOGLOBIN LEVEL OF 8.6 G/DL, HEMATOCRIT OF 28%, D-DIMER CONCENTRATION OF 12.21 MICROGRAM/ML, PLATELET COUNT OF 126 × 1 BILLION/L, AND A PTT OF 26.7 SECONDS. WHILE THESE VALUES APPEARED RELATIVELY STABLE FOR AN ECMO PATIENT, THE MARKEDLY ELEVATED D-DIMER LEVEL MAY SUGGEST EARLY THROMBOTIC ACTIVITY WITHIN THE CIRCUIT, DESPITE THE ABSENCE OF VISIBLE CLOTS AT THAT TIME (1, 2). ACCORDING TO RAJSIC ET AL. (2022) (3), THE RECOMMENDED THERAPEUTIC RANGE FOR THE PREVENTION OF VENOUS THROMBOEMBOLISM IS 1.5 TO 2.5 TIMES THE PATIENT¿S BASELINE APTT (40¿80 SECONDS). A PTT OF 26.7 SECONDS MAY HAVE CONTRIBUTED TO THROMBUS FORMATION WITHIN THE ECMO CIRCUIT. THE HLS SET ITSELF WAS NOT REPLACED DURING THE ECMO RUN, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES WERE OBSERVED IN CORRELATION WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING PRE- AND POST-OXYGENATOR BLOOD SAMPLING. DESPITE THESE ANOMALIES, NO REDUCTION IN BLOOD FLOW WAS REPORTED, AND THE CIRCUIT CONTINUED TO PROVIDE EFFECTIVE SUPPORT. IN CONCLUSION, A PRODUCT MALFUNCTION INVOLVING THE INTERNAL PRESSURE SENSOR OF THE HLS SET WAS LIKELY, AS EVIDENCED BY PERSISTENT INACCURATE PRESSURE READINGS DESPITE CABLE REPLACEMENT. HOWEVER, THIS MALFUNCTION DID NOT RESULT IN A LOSS OF ECMO SUPPORT OR REDUCTION IN BLOOD FLOW. THAT SAID, REDUNDANT PRESSURE MONITORING DURING WAS, ASSUMEDLY, PRESENT VIA THE PART PRESSURE FROM THE HLS DISPOSABLE. NOTABLY, THE FIRST PRESSURE ANOMALY WAS REPORTED ON (B)(6)¿NINE DAYS AFTER ECMO INITIATION ON (B)(6)¿INDICATING THAT THE HLS SET INITIALLY PERFORMED AS EXPECTED. THE SUBSEQUENT DISCOVERY OF BLOOD RESIDUES ON THE DEVICE UPON DISASSEMBLY SUGGESTS THAT CONTAMINATION OF THE INTERNAL PRESSURE SENSOR PORT MAY HAVE OCCURRED WHILE THERAPY, POTENTIALLY CONTRIBUTING TO THE OBSERVED BEHAVIOR. THE PATIENT¿S CLINICAL COURSE ULTIMATELY ENDED WITH WITHDRAWAL OF CARE AT THE FAMILY¿S REQUEST. WHILE THROMBOTIC EVENTS MAY HAVE CONTRIBUTED TO OVERALL MORBIDITY, THE DECISION TO WITHDRAW CARE APPEARS TO BE INFLUENCED BY BROADER CLINICAL OR ETHICAL CONSIDERATIONS. BASED ON AVAILABLE INFORMATION, THE PATIENT¿S EXPIRATION WAS NOT INFLUENCED BY THE DEVICE MALFUNCTION. THE FOLLOWING IS A QUOTE FROM THE INITIATING COMPLAINT INDICATING APPARENT SYNCHRONY WITH THE ABOVE OPINION ON DEVICE INFLUENCE: "ECLS COORDINATOR AT (B)(6) HOSPITAL, MENTIONED THEY ARE HAVING ISSUES WITH ERRONEOUS PRESSURE READINGS FROM AN HLS ON A PATIENT. IT IS NOT NEGATIVELY IMPACTING THE PATIENT OR LEVEL OF CARE, NO PATIENT HARM." ACCORDING TO THE INSTRUCTION FOR USE (HLS SET, CHAPTER "PREPARATION AND INSTALLATION") THE PRESSURE SENSORS HAVE TO BE CHECKED BEFORE PRIMING. IN GENERAL, IT SHOULD BE NOTED THAT PRE-PRIMING CAN HAVE A NEGATIVE IMPACT ON THE SENSOR FUNCTION AND THE ASSOCIATED PRESSURE VALUES AND CAN LEAD TO THE OUTPUT OF UNEXPECTED / IMPLAUSIBLE PRESSURE VALUES SHORTLY BEFORE USE, AS INDICATED IN CHAPTER "PRIMING THE SYSTEM", WHERE IT SAYS "ONLY FILL THE SYSTEM SHORTLY BEFORE USE AND IN ACCORDANCE WITH THE PRIMING INSTRUCTIONS. CALIBRATION VALUES ARE LOST WHEN THE DRIVE UNIT IS SWITCHED OFF". FURTHER THE CARDIOHELP HAS A FLOW/BUBBLE SENSOR AND A VENOUS PROBE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN ADDITION, WITHIN THE CHAPTER ¿PRIMING THE SYSTEM¿ IT IS WRITTEN NOT TO PRE-PRIME THE CIRCUIT TOO EARLY. BASED ON THE RESULTS THE REPORTED FAILURE "PRESSURE ISSUES" COULD NOT BE CONFIRMED. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-04. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891250&pc=OZD | 1220648-2025-27762 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Restricted Flow rate; Increase in Pressure | Septic Shock | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THERE WERE SUCTION ALARMS AND THE PUMP WAS REPOSITIONED. SUPPORT CONTINUED. ADDITIONALLY, THE PATIENT HAD SEPTIC SHOCK WHILE ON SUPPORT. THE CAUSE AND TREATMENT OF THE SEPTIC SHOCK IS UNKNOWN. FURTHERMORE, HIGH PURGE PRESSURE WAS NOTED. TISSUE PLASMA ACTIVATOR WAS ADDED TO THE PURGE. THE PATIENT REMAINS ON SUPPORT.Event Description: ADDITIONAL INFORMATION RECEIVED THAT THE PATIENT EXPIRED WHILE ON SUPPORT. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE INFORMING THAT THE PATIENT EXPIRED WHILE ON SUPPORT. THE FOLLOWING FIELDS WERE UPDATED TO REFLECT THE NEW OUTCOME. B2 DEATH AND DATE OF DEATH. B5 ADDITIONAL INFORMATION ADDED. D6B EXPLANT DATE. H1 TYPE OF REPORT. H6 CODE 1802 ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891384&pc=OZD | 1220648-2025-27763 | 2025-04-01 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Material Twisted/Bent | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP PUMP WAS PLACED ALONG WITH AN EXTRACORPOREAL MEMBRANE OXYGENATION TO SUPPORT A PATIENT ADMITTED IN WITH CARDIAC HISTORY AND ST - ELEVATED MYOCARDIAL INFARCTION. THE CORONARIES WERE INTERVENED UPON AND THE PATIENT WAS SENT FOR COMPUTED TOMOGRAPHY DUE TO FREE FLUID IN THE ABDOMEN, WHICH PROMPTED MANY BLOOD PRODUCTS. DURING SUPPORT, THERE WAS SUCTION AND LIMB ISCHEMIA ISSUES. THE TEAM AND FAMILY MADE THE PATIENT DO NOT RESUSCITATE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND LIMB ISCHEMIA HAS BEEN COMPLETED. PRODUCT DAMAGE (CANNULA KINK) WAS ADDED SINCE A CANNULA KINK WAS FOUND ON THE RETURNED DEVICE. LOW FLOW - REPORTED EXTENDED CARDIOPULMONARY RESUSCITATION (CPR) AND HYPOPERFUSION REPORTED AS CONTRIBUTING FACTORS. THE PUMP WAS RETURNED AND A CANNULA KINK NEAR THE MOTOR HOUSING WAS OBSERVED. DATA LOGS SHOW PUMP FLUCTUATED BETWEEN P7 AND P2 ON AUTO FOR THE FIRST HOUR. THERE WERE SUCTION ALARMS FROM WHEN THE PUMP STARTED AND IMPELLA FLOW REDUCED ALARMS IN RESPONSE TO SUCTION. PUMP WAS THEN SET TO P-3, BRIEFLY TURNED UP TO P9 AND BACK DOWN BUT WITH FLOWS BELOW 0.5L/MIN. FLOWS THEN FLUCTUATE BETWEEN 0.3L/MIN AND 1L/MIN WITH MORE SUCTION ALARMS WHEN PUMP AT P-2 FOR THE NEXT 4 HOURS. AFTER TURNING THE P-LEVEL TO P-3, THEN BACK TO P-2, FLOWS STABILIZED AROUND 0.6L/MIN FOR 4 HOURS. PUMP WAS RAN AT P1 FOR THE LAST 21 HOURS WITH FLOWS BETWEEN 0.5L/MIN AND 0.7L/MIN. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY A KINKED CANNULA AS OBSERVED ON THE RETURNED PRODUCT DUE TO EXTENDED CPR REPORTED AS A CONTRIBUTING FACTOR. CANNULA KINK - REPORTED EXTENDED CPR AND HYPOPERFUSION REPORTED AS CONTRIBUTING FACTORS TO THE FAILURE MODES. IMPELLA CP WAS RETURNED AND A CANNULA KINK WAS FOUND. THE ROOT CAUSE OF THE PRODUCT DAMAGE WAS A KINKED CANNULA AS OBSERVED ON THE RETURNED PRODUCT DUE TO EXTENDED CPR REPORTED AS A CONTRIBUTING FACTOR. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891399&pc=KDJ | 1416980-2025-02433 | 2025-03-30 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-04-23 04:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS. THE CAUSE OF THE PERITONITIS WAS NOT REPORTED. IT WAS REPORTED THAT ON AN UNSPECIFIED DATE AFTER THE PERITONITIS DIAGNOSIS, THE PATIENT WAS HOSPITALIZED FOR AN UNRELATED EVENT. TREATMENT FOR THE PERITONITIS EVENT WAS NOT REPORTED. IT WAS REPORTED THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT IS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF THE PERITONITIS WAS RESOLVED PRIOR TO DEATH. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: B2, B3, B5, B6, D10, E1. B5: UPON FOLLOW UP IT WAS REPORTED THE CAUSE OF PERITONITIS WAS UNKNOWN. THE PERITONITIS WAS MANIFESTED BY CLOUDY FLUID. THE SAME DATE OF THE PERITONITIS DIAGNOSIS, THE PATIENT WAS TREATED WITH VANCOMYCIN INJECTION (2 GM, EVERY SEVEN DAYS, INTRAPERITONEAL, ONGOING) AND CEFTAZIDIME INJECTION (1 GM, AT BEDTIME INTRAPERITONEAL, ONGOING) FOR PERITONITIS. TWO DAYS AFTER THE PERITONITIS ONSET, THE PATIENT EXPERIENCED A BRAIN HEMORRHAGE. THE FOLLOWING DAY, THE PATIENT PASSED AWAY. THE CAUSE OF THE DEATH WAS REPORTED TO BE DUE TO A BRAIN HEMORRHAGE. AT THE TIME OF DEATH, THE PATIENT WAS RECOVERING FROM THE PERITONITIS EVENT AND ANTIBIOTIC AND PD THERAPY WERE ONGOING. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891526&pc=HIF | 3002808148-2025-06698 | 2025-03-22 04:00 | Death | SHIRAKAWA OLYMPUS CO., LTD. | 2025-04-23 04:00 | HIF | HIGH FLOW INSUFFLATION UNIT | Adverse Event Without Identified Device or Use Problem | Air Embolism; Cardiac Arrest; Low Blood Pressure/ Hypotension | K243527 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC SURGERY FOR DIVERTICULAR DISEASE, AFTER THE INTRODUCTION OF OPTIC TROCAR, INDUCTION OF PNEUMOPERITONEUM WITH CARBON DIOXIDE (CO2) AT 12 MMHG, PLACEMENT OF THREE ADDITIONAL TROCARS, EXPLORATION OF THE ABDOMINAL CAVITY, AND LYSIS OF ADHESIONS, THE PATIENT EXPERIENCED A DROP IN BLOOD PRESSURE AND CARDIAC ARREST, WHICH RESULTED IN DEATH. THE PATIENT'S FAMILY AGREED ONLY TO A POSTMORTEM COMPUTED TOMOGRAPHY (CT) SCAN, WHICH DOCUMENTED A PICTURE OF SYSTEMIC GAS EMBOLISM AND ABSENCE OF VASCULAR LESIONS. THE DEVICE WAS CHECKED BEFORE STARTING THE PROCEDURE WITHOUT DETECTING ANY ANOMALY. IT DID NOT GIVE ANY ERROR CODE UNTIL PATIENT¿S DEATH. NO OVERPRESSURE EVENT OCCURRED, AND NO DEVICE MALFUNCTION WAS REPORTED DURING THE PROCEDURE. IT WAS NOTED THAT CHECKS WERE PERFORMED ON THE DEVICES USED DURING THE PROCEDURE, AND THE RESULTS WERE IN NORMAL RANGE. THE CUSTOMER HYPOTHESIZED ROLE OF GRADIENT PARTIAL PRESSURE OF INSUFFLATED GAS. THE HIGH-FLOW INSUFFLATION UNIT AND THE NON-OLYMPUS ANESTHESIA VENTILATOR USED DURING THE PROCEDURE HAVE BEEN PLACED UNDER QUARANTINE. Manufacturer Narrative: D10: FRAME TROLLEY S/5 AISYS MONITOR RACK UPPER 3 DWR SN (B)(6) GE HEALTHCARE. ON MARCH 25, 2025, AN OLYMPUS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE USER FACILITY AND CHECKED THE DEVICE ON SITE IN THE PRESENCE OF THE HOSPITAL CLINICAL ENGINEER. NO ISSUE WAS FOUND. ON MARCH 27, 2025, AN OLYMPUS FSE WAS AT THE CUSTOMER'S SITE AGAIN FOR A DEEPER CHECK OF THE SUBJECT DEVICE WITH DEDICATED TOOLS. THE DEVICE PASSED ALL TESTS; NO FAILURES OR DEFECTS WERE FOUND. ALL CHECKS WERE PERFORMED IN THE PRESENCE OF THE HOSPITAL CLINICAL ENGINEER. OLYMPUS WAS NOT ABLE TO DETECT A FAILURE. THE DEVICE MEETS ITS SPECIFICATIONS. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891562&pc=DSQ | 3007042319-2025-00815 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-23 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Cardiac Arrest; Hemorrhage/Bleeding; Unspecified Infection; Respiratory Failure; Multiple Organ Failure; Unspecified Nervous System Problem; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING VENTRICULAR ASSIST DEVICES (VADS). THE ARTICLE DISCUSSED THE INTERAGENCY REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT (INTERMACS) REPORT, COVERING DATA ON IMPLANTATION AND FOLLOW-UP ON CLINICAL OUTCOMES IN THE RECENT (2017-2021) VERSUS PRIOR (2012-2016) ERAS. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE DUE TO BLEEDING, UNSPECIFIED CIRCULATORY SYSTEM ISSUES, HEART FAILURE, MAJOR INFECTIONS, MULTIPLE ORGAN FAILURE, NEUROLOGIC DYSFUNCTION, RESPIRATORY FAILURE, SUDDEN DEATH, WITHDRAWAL OF SUPPORT, UNSPECIFIED DEVICE MALFUNCTIONS AND OTHER UNSPECIFIED CAUSES. THERE WERE PATIENTS WHO EXPERIENCED ARTERIAL/VENOUS THROMBOEMBOLISMS, PUMP THROMBI, MAJOR DEVICE-RELATED/NON-DEVICE INFECTIONS, MAJOR BLEEDING, GASTROINTESTINAL BLEEDING, CARDIAC ARRHYTHMIAS, HEPATIC DYSFUNCTION, NEUROLOGIC DYSFUNCTION, STROKES, PERICARDIAL FLUID BUILDUP, RENAL DYSFUNCTION, RESPIRATORY FAILURE AND WOUND DEHISCENCE. PATIENTS REQUIRED RE-HOSPITALIZATIONS, SURGICAL INTERVENTIONS, AND PERICARDIAL FLUID DRAINAGE AS TREATMENT. THERE WERE ALSO UNKNOWN DEVICE MALFUNCTIONS. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: ### THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/57 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. REFERENCED ARTICLE: THE SOCIETY OF THORACIC SURGEONS INTERMACS 2022 ANNUAL REPORT: FOCUS ON THE 2018 HEART TRANSPLANT ALLOCATION SYSTEM. THE ANNALS OF THORACIC SURGERY. 2023; 115:311-28. DOI: 10.1016/J.ATHORACSUR.2022.11.023 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. ### MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891565&pc=DSQ | 2916596-2025-02302 | 2025-03-30 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WAS NOT EXPLANTED AND PRODUCT WILL NOT RETURN. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891570&pc=MCX | 2124215-2025-25744 | 2025-03-27 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-23 04:00 | MCX | ROTAPRO? | Entrapment of Device | Obstruction/Occlusion; Foreign Body In Patient | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH MYOCARDIAL INFARCTION (MI). THE 90% STENOSED TARGET LESION WAS LOCATED IN THE TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX ARTERY (LCX). A 1.50MM ROTAPRO WAS SELECTED FOR PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY (PTCA). DURING THE PROCEDURE, THE BURR BECAME STUCK IN A HIGHLY CALCIFIED LESION. THE PHYSICIAN ATTEMPTED TO REMOVE THE BURR FROM THE VESSEL BUT WAS UNSUCCESSFUL AND WAS NOT ABLE TO COMPLETE PROCEDURE. FLOW WITHIN THE VESSEL SLOWED AND EVENTUALLY STOPPED. THE PATIENT PASSED AWAY DUE TO THE VESSEL BLOCK.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH MYOCARDIAL INFARCTION (MI). THE 90% STENOSED TARGET LESION WAS LOCATED IN THE TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX ARTERY (LCX). A 1.50MM ROTAPRO WAS SELECTED FOR PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY (PTCA). DURING THE PROCEDURE, THE BURR BECAME STUCK IN A HIGHLY CALCIFIED LESION. THE PHYSICIAN ATTEMPTED TO REMOVE THE BURR FROM THE VESSEL BUT WAS UNSUCCESSFUL AND WAS NOT ABLE TO COMPLETE PROCEDURE. FLOW WITHIN THE VESSEL SLOWED AND EVENTUALLY STOPPED. THE PATIENT PASSED AWAY DUE TO THE VESSEL BLOCK. Manufacturer Narrative: H6 IMPACT CODE: CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891602&pc=DSQ | 2916596-2025-02326 | 2025-03-30 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER BEING TRANSITIONED TO COMFORT CARE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21891701&pc=NPT | 2025587-2025-02937 | 2023-10-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-23 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Unspecified Infection; Thromboembolism; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, THE PATIENT WAS RE-HOSPITALIZED FOR THE PURPOSE OF TREATING HEART FAILURE AFTER A FEMORAL NECK FRACTURE SURGERY. SUBSEQUENTLY, A PERIPHERAL VASCULAR THROMBOEMBOLISM OF THE LEFT FEMORAL ARTERY WAS OBSERVED, SO AN ENDOVASCULAR THROMBECTOMY WAS PERFORMED. TREATMENT WAS CONTINUED, HOWEVER, THE PATIENT SUFFOCATED. THE CONDITION WORSENED, AND THE PATIENT SUBSEQUENTLY DIED. PER THE PHYSICIAN, THE PERIPHERAL VASCULAR THROMBOEMBOLISM WAS RELATED TO THE TAVR PROCEDURE, BUT THE DEATH WAS REPORTED AS NOT RELATED TO THE VALVE OR PROCEDURE. THE CAUSE OF DEATH WAS REPORTED AS INFECTION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT 10 DAYS FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, CHOLECYSTITIS WAS OBSERVED. PER THE PHYSICIAN, THE CHOLECYSTITIS WAS NOT RELATED TO THE VALVE OR PROCEDURE. Manufacturer Narrative: CORRECTED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21892097&pc=DSQ | 3007042319-2025-00818 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-23 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CHILDREN WITH CONGENITAL HEART DISEASE (CHD) AND OUTCOMES AFTER LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. PATIENTS WERE STUDIED IN GROUPS OF UNIVENTRICULAR CHD, BIVENTRICULAR CHD, OR NON-CHD. CHILDREN WITH CHD HAD A SIGNIFICANTLY HIGHER RISK FOR POST¿VAD IMPLANTATION MORBIDITY AND MORTALITY THAN CHILDREN WITH NON-CHD. THE AUTHORS DESCRIBED PATIENT DEATHS IN ALL GROUPS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS IN ALL GROUPS WHO EXPERIENCED HEMORRHAGIC OR ISCHEMIC STROKES AND ISCHEMIC STROKES WITH HEMORRHAGIC CONVERSION, ACUTE ENCEPHALOPATHY, HYPOXIC ISCHEMIC ENCEPHALOPATHY, TRANSIENT ISCHEMIC ATTACK (TIA), MAJOR BLEEDING EVENTS, SEPSIS, MEDIASTINITIS, LOCALIZED NON-DEVICE INFECTIONS, SUSTAINED SUPRA VENTRICULAR TACHYCARDIA (SVT) WHICH REQUIRED DRUG TREATMENT OR CARDIOVERSION, SUSTAINED VENTRICULAR TACHYCARDIA (VT) WHICH REQUIRED DEFIBRILLATION OR CARDIOVERSION, RENAL DYSFUNCTION, HEMOLYSIS, HEPATIC DYSFUNCTION, UNKNOWN DEVICE MALFUNCTIONS, OR RESPIRATORY FAILURE. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/4 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: OUTCOMES FOR CHILDREN WITH CONGENITAL HEART DISEASE UNDERGOING VENTRICULAR ASSIST DEVICE IMPLANTATION AN ACTION REGISTRY ANALYSIS. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY. 2025; 85:804¿814. DOI: 10.1016/J.JACC.2024.10.083 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21892102&pc=NPT | 2025587-2025-02940 | 2024-03-09 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-23 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA; Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO MULTIPLE CEREBRAL INFARCTIONS AND SEPSIS. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21892268&pc=OZD | 1220648-2025-27766 | 2025-04-01 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT HAD AN ATTEMPTED IMPELLA CP IMPLANT DURING A HIGH RISK CARDIAC PROCEDURE WHEN THE PATIENT WENT INTO CARDIAC ARREST. THERE WERE MULTIPLE ROUNDS OF RESUSCITATION PERFORMED AND ACCESS WAS OBTAINED FOR IMPELLA CP INSERTION. DURING THE IMPLANT, THE STAFF WAS UNABLE TO ADVANCE PAST THE CALCIFIED ILIAC ARTERY. THE PATIENT REQUIRED CONTINUED CARDIOPULMONARY RESUSCITATION AND SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE PUMP WAS UNABLE TO ADVANCE PAST ILIAC ARTERY DUE TO CALCIFICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21892269&pc=BZD | 2518422-2025-104612 | 2025-04-09 04:00 | Death | RESPIRONICS, INC. | 2025-04-23 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21892533&pc=NPT | 2025587-2025-02945 | 2025-03-09 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-23 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN.Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE; SERIAL #; FULL UDI # H4. DEVICE MFG DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21892759&pc=NDN | 1526439-2025-00524 | 2025-01-01 05:00 | Death | DEPUY SPINE INC | 2025-04-23 04:00 | NDN | CONFIDENCE KIT, NO NEEDLES | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K060300 |  | 1 | Event Description: IN A REVISION SURGERY FOR SPINAL FIXATION DUE TO VERTEBRA T10 COLLAPSE AND A FRACTURE AT VERTEBRA L5 IN A PATIENT WITH SEVERE OSTEOPOROSIS, IT WAS DECIDED TO PERFORM FIXATION FROM VERTEBRA T8 DOWN TO THE PELVIS. AFTER INSERTING ALL 18 SCREWS, DECOMPRESSION WAS PERFORMED, AND THE SURGEON BEGAN INJECTING CEMENT INTO THE VERTEBRAE. 4 CEMENT KITS WERE INJECTED, DURING WHICH THE PATIENT EXPERIENCED, TWICE, A DROP IN THE PULSE AND BLOOD PRESSURE, WAS STABILIZED AND THE SURGEON DECIDED TO CONTINUE WITH THE SURGERY. THEREAFTER, AND PRIOR TO PROCEEDING TO INJECT AN ADDITIONAL CEMENT KIT, THE SURGEON NOTICED MASSIVE BLEEDING IN THE SURGICAL AREA THAT HAD BEEN COVERED WITH A STERILE DRAPE AND HAD NOT BEEN OBSERVED BEFORE. THE SURGEON TRIED TO CONTROL THE BLEEDING AND THE PATIENT'S BLOOD PRESSURE AND PULSE DROPPED SHARPLY. AN ADDITIONAL SENIOR SURGEON JOINED THE SURGERY AND RESUSCITATION EFFORTS BEGAN, BUT ULTIMATELY, THEY HAD TO DECLARE HER DEATH. BASED ON OUR INQUIRY, NO AUTOPSY WAS PERFORMED, AND THE CAUSE OF DEATH IS UNKNOWN.Event Description: ADDITIONAL INFORMATION RECEIVED: A. IT WAS MENTIONED THAT THIS WAS A REVISION SURGERY. WHAT WAS THE ORIGINAL SURGERY, AND WHEN WAS IT PERFORMED? WHAT WAS THE REASON FOR THE REVISION? REVISION SURGERY OF KYPHOPLASTY B. ARE ANY PREOP OR INTRAOPERATIVE IMAGES AVAILABLE? NO. C. ANY DEVICE MALFUNCTIONING WAS NOTICED? NO. D. WHAT IS THE SURGEON'S ASSESSMENT OF THE CAUSE OF DEATH? DAMAGE TO BLOOD VESSELS, LATER DETECTION OF BLEEDING. E. CAN THE AUTOPSY REPORT BE SHARED? IF NOT, WHAT¿S THE CONCLUSION OF THE REPORT? NO, THE FAMILY DECIDED NOT TO MAKE AUTOPSY.Event Description: ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: G1 (MANUFACTURING SITE NAME), H4. H3, H6: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. DEVICE HISTORY REVIEW (DHR): A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE. PRODUCT CODE: 283913000. LOT NUMBER: 424341. IT WAS ELECTRONICALLY REVIEWED AND NO NONCONFORMANCES/MANUFACTURING IRREGULARITIES WERE IDENTIFIED DURING THE MANUFACTURING PROCESS. THE PRODUCT WAS RELEASED ON: 13 FEB 2025. MANUFACTURING SITE: (B)(4). EXPIRY DATE: 31 JAN 2027. CEMENT NUMBER: PRODUCT CODE: 183901001/ BATCH NUMBER: 4597805. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: B5 AND D10. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. H11 ADDITIONAL NARRATIVE: ADDED: A2 (DOB, AGE), A3A (SEX), B5 CORRECTED: H6 (HEALTH EFFECT - IMPACT CODE) IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893008&pc=NPT | 2025587-2025-02950 | 2024-02-10 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-23 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO AN UNKNOWN REASON. Manufacturer Narrative: CORRECTED DATA: E1 - STREET ADDRESS CORRECTED FROM (B)(6) UPDATED DATA: D4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893105&pc=DSQ | 2916596-2025-02254 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Muscle Weakness; Numbness; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUFFERED A STROKE ON (B)(6) 2025, SHOWING WEAKNESS AND NUMBNESS ON THEIR LEFT SIDE, AND WAS ADMITTED TO THE EMERGENCY ROOM (ER). THE PATIENT'S COUMADIN (WARFARIN) WAS REVERSED, AND THE PATIENT WAS CONSERVATIVELY MANAGED AND MONITORED. A COMPUTED TOMOGRAPHY (CT) OF THE PATIENT'S HEAD WAS TAKEN, AND THEY WERE FOUND TO HAVE SUFFERED A SUBARACHNOID HEMORRHAGE (SAH), THE CAUSE OF WHICH WAS UNKNOWN. THE CUSTOMER REPORTED NO CHANGES TO THE PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED, AND THERE HAD BEEN NO RECENT CHANGES TO THE PATIENT'S PUMP PARAMETERS. THE PATIENT WAS TRANSFERRED TO INPATIENT HOSPICE. THE PATIENT ELECTED TO HAVE THEIR VENTRICULAR ASSIST DEVICE (VAD) TURNED OFF, AND THEY PASSED AWAY (B)(6) 2025. THE PATIENT WAS ALSO NOTED TO HAVE HAD EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO) (MFR 2916596-2024-52521). THE PUMP WAS FUNCTIONING AS EXPECTED AND WAS EXPLANTED AND RETURNED. Manufacturer Narrative: SECTION E3 CORRECTION. SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE UNEXPECTED MEDICAL INTERVENTION ADDED. SECTION H6 CORRECTION: MEDICAL DEVICE PROBLEM CODE 2993 - ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED HEMORRHAGIC STROKE AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, EVALUATION OF HEARTMATE 3 LVAS, SERIAL NUMBER: (B)(6), CONFIRMED EOGO. A SPECIFIC CAUSE FOR THIS FINDING COULD NOT BE CONCLUSIVELY DETERMINED. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE FULLY INTACT. THE MODULAR CABLE WAS ALSO RETURNED ATTACHED TO THE PUMP CABLE. APPROXIMATELY 4.0¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF PROPERLY ENGAGED TO THE GRAFT HARDWARE. TWO ADDITIONAL SEGMENTS OF THE OUTFLOW GRAFT WERE RETURNED DETACHED FROM THE PUMP, MEASURING 1.5¿ AND 5.5¿, RESPECTIVELY. THE APICAL CUFF WAS RETURNED WITH THE CUFF LOCK FULLY ENGAGED. EXAMINATION OF THE TEXTURED, BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE DURING SUPPORT. UPON REMOVAL OF THE OUTFLOW GRAFT BEND RELIEF, A SMOOTH, BIOLOGICAL ACCUMULATION WAS OBSERVED ALONG THE PROXIMAL END OF THE 4.0¿ OUTFLOW GRAFT SEGMENT UNDERNEATH THE BEND RELIEF. THE TISSUE WAS DISPLACED DURING DISASSEMBLY, REVEALING THAT IT HAD BEEN WELL-FORMED INTO THE GEOMETRY OF A LENGTHWISE CREASE ALONG THE PROXIMAL END OF THE GRAFT. THE REMAINDER OF THE OUTFLOW GRAFT WAS UNREMARKABLE. THE TEXTURED SURFACE OF THE GRAFT ATTACHMENT REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS. THE O-RING IN THE GRAFT ATTACHMENT WAS PRESENT AND SEATED PROPERLY. THE LUMEN OF THE SEALED OUTFLOW GRAFT SEGMENTS REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS. THESE FINDINGS APPEAR CONSISTENT WITH EOGO DURING SUPPORT. A SPECIFIC CAUSE FOR THE OBSERVED DEFORMATION IN THE OUTFLOW GRAFT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. THIS SECTION UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH. SECTION 5 OF THE IFU ALSO CAUTIONS THE USER THAT SHARP BENDS, TWISTS, OR KINKS IN THE DRIVELINE MAY MAKE IT MORE SUSCEPTIBLE TO WEAR AND FATIGUE OVER TIME. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT", CAUTIONS THE USER TO AVOID PULLING ON OR MOVING THE DRIVELINE. THIS SECTION FURTHER CAUTIONS: "DO NOT TWIST, KINK, OR SHARPLY BEND THE DRIVELINE, SYSTEM CONTROLLER POWER CABLES, THE POWER MODULE PATIENT CABLE, OR THE MOBILE POWER UNIT PATIENT CABLE, WHICH MAY CAUSE DAMAGE TO THE WIRES INSIDE, EVEN IF EXTERNAL DAMAGE IS NOT VISIBLE. DAMAGE TO THE DRIVELINE OR CABLES COULD CAUSE THE LEFT VENTRICULAR ASSIST DEVICE TO STOP. IF THE DRIVELINE OR CABLES BECOME TWISTED, KINKED, OR BENT, CAREFULLY UNRAVEL AND STRAIGHTEN." THIS SECTION (UNDER "CARING FOR THE DRIVELINE") INSTRUCTS THE USER TO CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS. FURTHERMORE, SECTION 6 OF THE IFU (UNDER "ANTICOAGULATION") OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING" PROVIDES ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE IN A SUB-SECTION ENTITLED "WHAT NOT TO DO: DRIVELINE AND CABLES". FINALLY, SECTION 8 OF THE IFU, "EQUIPMENT STORAGE AND CARE" (UNDER "CLEANING THE DRIVELINE") EXPLAINS THAT AS NEEDED, THE USER CAN CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, WARM WATER AND MILD DISH SOAP SHOULD BE USED. THE HEARTMATE 3 LVAS PATIENT HANDBOOK IS ALSO CURRENTLY AVAILABLE. SECTION 4 OF THE PATIENT HANDBOOK, "LIVING WITH THE HEARTMATE III" (UNDER "CARING FOR THE DRIVELINE") CONTAINS INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE. THIS SECTION INSTRUCTS THE USER: "CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE (CUTS, HOLES, TEARS). CALL YOUR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED)." AND "KEEP YOUR DRIVELINE CLEAN. WIPE OFF ANY DIRT OR GRIME. IF THE DRIVELINE GETS DIRTY, USE A TOWEL WITH MILD DISH SOAP AND WARM WATER TO GENTLY CLEAN IT. NEVER SUBMERGE THE DRIVELINE OR OTHER SYSTEM COMPONENTS IN WATER OR LIQUID." IN ADDITION, SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" CONTAINS A SUB-SECTION ENTITLED "WHAT NOT TO DO: DRIVELINE AND CABLES", WHICH EXPLICITLY CAUTIONS THE USER NOT TO TWIST, KINK, OR SHARPLY BEND THE DRIVELINE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893169&pc=NDN | 1526439-2025-00528 | 2025-01-01 05:00 | Death | DEPUY SPINE INC | 2025-04-23 04:00 | NDN | CONFIDENCE KIT, NO NEEDLES | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K060300 |  | 1 | Event Description: IN A REVISION SURGERY FOR SPINAL FIXATION DUE TO VERTEBRA T10 COLLAPSE AND A FRACTURE AT VERTEBRA L5 IN A PATIENT WITH SEVERE OSTEOPOROSIS, IT WAS DECIDED TO PERFORM FIXATION FROM VERTEBRA T8 DOWN TO THE PELVIS. AFTER INSERTING ALL 18 SCREWS, DECOMPRESSION WAS PERFORMED, AND THE SURGEON BEGAN INJECTING CEMENT INTO THE VERTEBRAE. 4 CEMENT KITS WERE INJECTED, DURING WHICH THE PATIENT EXPERIENCED, TWICE, A DROP IN THE PULSE AND BLOOD PRESSURE, WAS STABILIZED AND THE SURGEON DECIDED TO CONTINUE WITH THE SURGERY. THEREAFTER, AND PRIOR TO PROCEEDING TO INJECT AN ADDITIONAL CEMENT KIT, THE SURGEON NOTICED MASSIVE BLEEDING IN THE SURGICAL AREA THAT HAD BEEN COVERED WITH A STERILE DRAPE AND HAD NOT BEEN OBSERVED BEFORE. THE SURGEON TRIED TO CONTROL THE BLEEDING AND THE PATIENT'S BLOOD PRESSURE AND PULSE DROPPED SHARPLY. AN ADDITIONAL SENIOR SURGEON JOINED THE SURGERY AND RESUSCITATION EFFORTS BEGAN, BUT ULTIMATELY, THEY HAD TO DECLARE HER DEATH.Event Description: ADDITIONAL INFORMATION RECEIVED: A. IT WAS MENTIONED THAT THIS WAS A REVISION SURGERY. WHAT WAS THE ORIGINAL SURGERY, AND WHEN WAS IT PERFORMED? WHAT WAS THE REASON FOR THE REVISION? REVISION SURGERY OF KYPHOPLASTY B. ARE ANY PREOP OR INTRAOPERATIVE IMAGES AVAILABLE? NO. C. ANY DEVICE MALFUNCTIONING WAS NOTICED? NO. D. WHAT IS THE SURGEON'S ASSESSMENT OF THE CAUSE OF DEATH? DAMAGE TO BLOOD VESSELS, LATER DETECTION OF BLEEDING. E. CAN THE AUTOPSY REPORT BE SHARED? IF NOT, WHAT¿S THE CONCLUSION OF THE REPORT? NO, THE FAMILY DECIDED NOT TO MAKE AUTOPSY.Event Description: ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H3, H6: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE DEVICE HISTORY REVIEW (DHR): THIS COMBINATION OF PART#283913000 AND LOT NUMBER 421395 DOESN'T EXIST. THE CORRECT PART NR FOR THIS LOT SHOULD BE 283910000. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. H11 ADDITIONAL NARRATIVE: ADDED: A2 (DOB, AGE), A3A (SEX), B5 CORRECTED: H6 (HEALTH EFFECT - IMPACT CODE) IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: B5 AND D10. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893179&pc=DSQ | 2916596-2025-02478 | 2025-03-10 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Liver Damage/Dysfunction; Renal Failure; Septic Shock; Respiratory Failure; Thrombocytopenia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD SEVERE VASOPLEGIA SECONDARY TO SUSPECTED COMBINED CARDIOGENIC-SEPTIC SHOCK WITH RIGHT VENTRICULAR FAILURE, AND COMBINED HYPOXIC-HYPERCARBIC RESPIRATORY FAILURE SECONDARY TO ACUTE RESPIRATORY DISTRESS SYNDROME STATUS POST VENO-ARTERIAL (VA) EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CANNULATION ON (B)(6) 2025 AND DECANNULATION (B)(6) 2025. THE PATIENT HAD ACUTE RENAL FAILURE (ARF) REQUIRING CONTINUOUS RENAL REPLACEMENT (CRRT), ACUTE LIVER INJURY, WHICH WAS RESOLVED, ENCEPHALOPATHY WHICH WAS RESOLVED, POST-OPERATIVE ILEUS, WHICH WAS RESOLVED, HYPERVOLEMIA AND HEPARIN INDUCED THROMBOCYTOPENIA (HIT) (PF4+ BUT SRA NEGATIVE). THE PATIENT PASSED AWAY ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, INFECTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893247&pc=NDN | 1526439-2025-00529 | 2025-01-01 05:00 | Death | DEPUY SPINE INC | 2025-04-23 04:00 | NDN | CONFIDENCE KIT, NO NEEDLES | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K060300 |  | 1 | Event Description: ADDITIONAL INFORMATION RECEIVED: A. IT WAS MENTIONED THAT THIS WAS A REVISION SURGERY. WHAT WAS THE ORIGINAL SURGERY, AND WHEN WAS IT PERFORMED? WHAT WAS THE REASON FOR THE REVISION? REVISION SURGERY OF KYPHOPLASTY B. ARE ANY PREOP OR INTRAOPERATIVE IMAGES AVAILABLE? NO. C. ANY DEVICE MALFUNCTIONING WAS NOTICED? NO. D. WHAT IS THE SURGEON'S ASSESSMENT OF THE CAUSE OF DEATH? DAMAGE TO BLOOD VESSELS, LATER DETECTION OF BLEEDING. E. CAN THE AUTOPSY REPORT BE SHARED? IF NOT, WHAT¿S THE CONCLUSION OF THE REPORT? NO, THE FAMILY DECIDED NOT TO MAKE AUTOPSY.Event Description: ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS.Event Description: IN A REVISION SURGERY FOR SPINAL FIXATION DUE TO VERTEBRA T10 COLLAPSE AND A FRACTURE AT VERTEBRA L5 IN A PATIENT WITH SEVERE OSTEOPOROSIS, IT WAS DECIDED TO PERFORM FIXATION FROM VERTEBRA T8 DOWN TO THE PELVIS. AFTER INSERTING ALL 18 SCREWS, DECOMPRESSION WAS PERFORMED, AND THE SURGEON BEGAN INJECTING CEMENT INTO THE VERTEBRAE. 4 CEMENT KITS WERE INJECTED, DURING WHICH THE PATIENT EXPERIENCED, TWICE, A DROP IN THE PULSE AND BLOOD PRESSURE, WAS STABILIZED AND THE SURGEON DECIDED TO CONTINUE WITH THE SURGERY. THEREAFTER, AND PRIOR TO PROCEEDING TO INJECT AN ADDITIONAL CEMENT KIT, THE SURGEON NOTICED MASSIVE BLEEDING IN THE SURGICAL AREA THAT HAD BEEN COVERED WITH A STERILE DRAPE AND HAD NOT BEEN OBSERVED BEFORE. THE SURGEON TRIED TO CONTROL THE BLEEDING AND THE PATIENT'S BLOOD PRESSURE AND PULSE DROPPED SHARPLY. AN ADDITIONAL SENIOR SURGEON JOINED THE SURGERY AND RESUSCITATION EFFORTS BEGAN, BUT ULTIMATELY, THEY HAD TO DECLARE HER DEATH. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: B5 AND D10. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: H6 COMPONENT CODES: MOST RELEVANT COMPONENT CODE IS G07002 (APPROPRIATE TERM/CODE NOT AVAILABLE) TO CAPTURE NO FINDINGS AVAILABLE DUE TO NO PRODUCT RETURNED. H3, H6: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. DEVICE HISTORY LOT: A DHR EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE. PRODUCT CODE: 283913000, LOT NUMBER: 395676, NO NON-CONFORMANCE'S / MANUFACTURING IRREGULARITIES RELATED TO THE MALFUNCTION WERE IDENTIFIED. 2 NRS LINKED TO THIS LOT WERE FOUND: JBL-NR-0027666 IS RELATED TO 3 OUT OF 10 TESTS WERE POSITIVE ON DOSE AUDIT TESTS FOR FIRST QUARTER OF 2024 ON CONFIDENCE KITS. CONSIDERING THE CEMENT IS ALREADY STERILIZED BEFORE BEING SENT TO JABIL, THIS ISSUE IS OUT OF SCOPE. THE NR DEMONSTRATED THAT THE REST OF THE CEMENT KIT IS COMPLIANT WITH STERILIZATION REQUIREMENTS. JBL-NR-0027635 IS RELATED TO DIMENSIONAL FEATURE NON-CONFORM. THIS ISSUE IS IMPACTING THE TRANSFER ADAPTOR, WHICH IS NOT RELATED TO THE COMPLAINT DESCRIPTION. CONSIDERING THE CORRESPONDING DRAWING HAS BEEN UPDATED TO REFLECT THE MANUFACTURED PRODUCT, PARTS IMPACTED BY THIS ISSUE ARE CONSIDERED AS CONFORM SINCE THE INITIAL VALIDATION. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893365&pc=NDN | 1526439-2025-00530 | 2025-01-01 05:00 | Death | DEPUY SPINE INC | 2025-04-23 04:00 | NDN | CONFIDENCE KIT, NO NEEDLES | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K060300 |  | 1 | Event Description: IN A REVISION SURGERY FOR SPINAL FIXATION DUE TO VERTEBRA T10 COLLAPSE AND A FRACTURE AT VERTEBRA L5 IN A PATIENT WITH SEVERE OSTEOPOROSIS, IT WAS DECIDED TO PERFORM FIXATION FROM VERTEBRA T8 DOWN TO THE PELVIS. AFTER INSERTING ALL 18 SCREWS, DECOMPRESSION WAS PERFORMED, AND THE SURGEON BEGAN INJECTING CEMENT INTO THE VERTEBRAE. 4 CEMENT KITS WERE INJECTED, DURING WHICH THE PATIENT EXPERIENCED, TWICE, A DROP IN THE PULSE AND BLOOD PRESSURE, WAS STABILIZED AND THE SURGEON DECIDED TO CONTINUE WITH THE SURGERY. THEREAFTER, AND PRIOR TO PROCEEDING TO INJECT AN ADDITIONAL CEMENT KIT, THE SURGEON NOTICED MASSIVE BLEEDING IN THE SURGICAL AREA THAT HAD BEEN COVERED WITH A STERILE DRAPE AND HAD NOT BEEN OBSERVED BEFORE. THE SURGEON TRIED TO CONTROL THE BLEEDING AND THE PATIENT'S BLOOD PRESSURE AND PULSE DROPPED SHARPLY. AN ADDITIONAL SENIOR SURGEON JOINED THE SURGERY AND RESUSCITATION EFFORTS BEGAN, BUT ULTIMATELY, THEY HAD TO DECLARE HER DEATH.Event Description: ADDITIONAL INFORMATION RECEIVED: A. IT WAS MENTIONED THAT THIS WAS A REVISION SURGERY. WHAT WAS THE ORIGINAL SURGERY, AND WHEN WAS IT PERFORMED? WHAT WAS THE REASON FOR THE REVISION? REVISION SURGERY OF KYPHOPLASTY B. ARE ANY PREOP OR INTRAOPERATIVE IMAGES AVAILABLE? NO. C. ANY DEVICE MALFUNCTIONING WAS NOTICED? NO. D. WHAT IS THE SURGEON'S ASSESSMENT OF THE CAUSE OF DEATH? DAMAGE TO BLOOD VESSELS, LATER DETECTION OF BLEEDING. E. CAN THE AUTOPSY REPORT BE SHARED? IF NOT, WHAT¿S THE CONCLUSION OF THE REPORT? NO, THE FAMILY DECIDED NOT TO MAKE AUTOPSY.Event Description: ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: D4 (EXPIRY DATE), G1 (MANUFACTURING SITE NAME), H4 CORRECTED: D4 (PRIMARY UDI NUMBER). H3, H6: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. DEVICE HISTORY REVIEW (DHR): A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 283913000421387 LOT NUMBER: IT WAS ELECTRONICALLY REVIEWED AND NO NONCONFORMANCES/MANUFACTURING IRREGULARITIES WERE IDENTIFIED DURING THE MANUFACTURING PROCESS. THE PRODUCT WAS RELEASED ON: 23 JAN 2025 MANUFACTURING SITE: JABIL LE LOCLE EXPIRY DATE: 31 DEC 2026 CEMENT: PART # 183901001/ LOT # 4572175. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: B5 AND D10 IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. H11 ADDITIONAL NARRATIVE: ADDED: A2 (DOB, AGE), A3A (SEX), B5 CORRECTED: H6 (HEALTH EFFECT - IMPACT CODE) IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893419&pc=MKJ | 1220908-2025-01390 |  | Death | ZOLL MEDICAL CORPORATION | 2025-04-23 04:00 | MKJ | AED PLUS | Self-Activation or Keying | Insufficient Information | P160015 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE AN 82-YEAR-OLD FEMALE PATIENT, THE DEVICE SELF DISCHARGED AFTER PROMPTING SHOCK ADVISED. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE ZOLL AED PLUS IS A SEMI-AUTOMATIC DEFIBRILLATOR, MEANING THE USER MUST PRESS THE SHOCK BUTTON TO DELIVER TREATMENT. DURING THE REPORTED EVENT, THE DEVICE FUNCTIONED AS DESIGNED: IT DETECTED A SHOCKABLE RHYTHM, ALERTED THE USER, AND DELIVERED A 120-JOULE SHOCK AFTER THE USER PRESSED THE BUTTON. THE LOG SHOWED NORMAL OPERATION, NO ERRORS, AND A SUCCESSFUL SELF-TEST BEFORE THE EVENT. POST-INCIDENT TESTING CONFIRMED THE DEVICE, ELECTRODES, AND BATTERIES WERE FULLY FUNCTIONAL. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893487&pc=DSQ | 2916596-2025-02381 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem | Aortic Valve Stenosis; Hemorrhage/Bleeding; Shock | P160054 |  | 1 | Event Description: ON (B)(6) 2025 THE PATIENT UNDERWENT SURGICAL REVISION DUE TO WORSENING CLINICAL SITUATION. UNFORTUNATELY DURING INITIAL PHASE OF OPENING AND VISUALIZATION OF THE HEARTMATE 3 PUMP AND OUTFLOW GRAFT SUDDEN BLEEDING WAS NOTICED FROM THE AREA OF THE OUTFLOW GRAFT WITHOUT POSSIBILITY TO VISUALIZE SOURCE OF BLEEDING AND TERMINATE IT. THE DECISION WAS MADE TO START CARDIOPULMONARY BYPASS (CPB) FROM FEMORAL APPROACH. SINCE THERE WERE PREVIOUS MULTIPLE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO)'S USING FEMORAL ACCESS THE TEAM WAS UNABLE DUE TO SEVERAL COMPLICATIONS TO ESTABLISH CPB SUPPORT. THE PATIENT PASSED AWAY UNDER PICTURE OF HEMORRHAGIC SHOCK. AN AUTOPSY WAS SCHEDULED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PREVIOUS MULTIPLE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMOS) WERE RELATED TO PRE-LEFT VENTRICULAR ASSIST DEVICE THERAPY.Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED FOR WORSENING CLINICAL CONDITION. A FULL EVALUATION WAS DONE AND A COMPUTED TOMOGRAPHY (CT) SCAN SHOWED SUSPECTED EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO) LIMITATION AT THE OUTFLOW BEND RELIEF PART AND STENOSIS AT THE AORTIC STRUCTURE. THE PLAN WAS TO LIST THE PATIENT FOR HEART TRANSPLANT. IN THE CASE OF DETERIORATION SURGICAL REVISION WOULD BE CONSIDERED. AN ECHOCARDIOGRAM (ECHO) SHOWED LIMITED POSSIBILITY TO UNLOAD THE LEFT VENTRICLE AND REGULAR OPENING AORTIC VALVE. THE BLOOD TEST ON HEMOLYSIS WAS NEGATIVE. THERE WERE NO LOW FLOWS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORT OF BLEEDING AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED COMPUTED TOMOGRAPHY (CT) IMAGES CONFIRMED THE REPORTED EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO); HOWEVER, A SPECIFIC CAUSE FOR THE REPORTED EVENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED CT IMAGES CONFIRMED COMPRESSION OF THE OUTFLOW GRAFT BENEATH THE BEND RELIEF. THESE FINDINGS ARE CONSISTENT WITH EOGO. REVIEW OF THE SUBMITTED LOG FILES CAPTURED THE SYSTEM OPERATING AS INTENDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS BLEEDING AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP", INSTRUCTS THE USER TO VERIFY THAT THE OUTFLOW GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5 OF THE IFU, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893687&pc=NIE | 3006260740-2025-02531 | 2025-03-12 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-04-23 04:00 | NIE | POWERTRIALYSIS | Device Dislodged or Dislocated | Hemorrhage/Bleeding | K133456 |  | 1 | Event Description: IT WAS REPORTED, PATIENT REMOVED TRIALYSIS CATHETER THAT WAS PLACED IN FEMORAL VEIN AND BLED OUT. SUTUTRES WERE IN TACT ALONG WITH THE SUTURE WINGS BUT TRIALYSIS HAD BEEN PULLED THROUGH THE LOOP. PATIENT BLED OUT AND EXPIRED. NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF DEATH WAS NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE REFLECTED IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE EVENT. EACH EVENT REPORTED TO BD IS EVALUATED AND INVESTIGATED IN ACCORDANCE WITH OUR COMPLAINT INVESTIGATION PROCEDURES. THE INVESTIGATION PROCESS INCLUDES, BUT IS NOT LIMITED TO, EVALUATION OF THE EVENT DETAILS PROVIDED BY THE COMPLAINT FACILITY, A REVIEW OF COMPLAINT HISTORY, MANUFACTURING RECORDS AND RISK DOCUMENT WHERE APPLICABLE, AND AN EVALUATION OF THE SUBJECT DEVICE WHEN AVAILABLE TO IDENTIFY POTENTIAL CONTRIBUTING FACTORS. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DETACHED SUTURE WINGS IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF A 20 CM POWERTRIALYSIS CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED THE SUTURE WINGS OF THE DEVICE WERE DISPLACED AND POSITIONED ABOUT 3 CM DISTAL TO THEIR ORIGINAL LOCATION ON THE MOLDED JOINT. NO OBVIOUS DAMAGE COULD BE SEEN ON THE SUTURE WINGS OR MOLDED JOINT IN THE RETURNED PHOTOGRAPH; HOWEVER, THE RESOLUTION OF THE PHOTOGRAPH MADE IT DIFFICULT TO DISCERN SUCH DETAILS. USE RESIDUE WAS OBSERVED ON THE CATHETER IN THE RETURNED PHOTOGRAPH. THE DISTAL END OF THE CATHETER APPEARED TO BE MISSING; HOWEVER, NO DETAILS OF THE POSSIBLE BREAK IN THE CATHETER COULD BE DISCERNED FROM THE RETURNED PHOTOGRAPH. WHILE THE SUTURE WINGS OF THE CATHETER SHOWN IN THE RETURNED PHOTOGRAPH WERE OBSERVED TO BE DISPLACED, IT WAS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE OF THIS DISPLACEMENT FROM THE RETURNED PHOTOGRAPH. THE APPROPRIATE RISK DOCUMENTATION WAS REVIEWED, AND RISK MANAGEMENT TOOLS ARE IN PLACE TO ENSURE THAT POTENTIAL RISKS ASSOCIATED WITH DEVICE USE ARE IDENTIFIED AND REDUCED TO THE EXTENT POSSIBLE. NO BATCH INFORMATION WAS RECEIVED FROM THE CUSTOMER AND, CONSEQUENTLY, IT WAS NOT POSSIBLE TO PERFORM A HISTORICAL REVIEW OF SIMILAR COMPLAINTS NOR WAS IT POSSIBLE TO PERFORM A REVIEW OF MANUFACTURING RECORDS AND PROCESSES FOR THE BATCH INVOLVED. DETAILS REGARDING CATHETER PLACEMENT, INDWELL TIME, OR MAINTENANCE ARE UNKNOWN, AND NO FURTHER INFORMATION REGARDING PATIENT HISTORY OR CONDITION HAS BEEN COMMUNICATED TO BD. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT, AND THIS EVALUATION WILL BE UPDATED UPON RECEIPT OF A PHYSICAL SAMPLE FOR EVALUATION AND/OR ANY ADDITIONAL INFORMATION PROVIDED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893860&pc=OZD | 1220648-2025-27775 | 2025-03-19 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5 WAS INPLANTED IN A PATIENT ADMITTED WITH CARDIAC HISTORY. THE PATIENT WAS BEING SUPPORTED ON TWO PUMPS FOR OVER 59 DAYS. ON DAY 26 OF THE BEING ON THE SECOND PUMP, THERE WERE CONCERNS NOTED FOR VENTRICULAR TACHYCARDIA (VT) DUE TO POOR POSITION WITHIN THE LEFT VENTRICLE. THE VT TRIGGERED THE AUTOMATIC IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR TO SHOCK THE PATIENT. THE PUMP POSITION WAS ASSESSED BY COMPUTED TOMOGRAPHY IMAGING. HOWEVER, DESPITE ALL THE MEASURES, THE CARE TEAM DETERMINED THAT THERE WERE NO OPTIONS FOR THE PATIENT AND THE PATIENT WAS TRANSITIONED TO COMFORT CARE AND SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21893908&pc=DSQ | 2916596-2025-02338 | 2025-01-25 05:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR (RV) FAILURE LEADING TO MULTI-SYSTEM ORGAN FAILURE. THE PATIENT WAS SUBTHERAPEUTIC FOR OVER 6 MONTHS WHICH LED TO CLOT IN RV AND CAUSED RV FAILURE. THE DEVICE FUNCTIONED AS EXPECTED WITH NO ISSUES. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING THROMBOEMBOLISM, MULTI ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894105&pc=DSQ | 2916596-2025-02537 | 2025-02-25 05:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. IT WAS REPORTED THAT NO, THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION.Event Description: THE MULTISYSTEM ORGAN FAILURE (MSOF) WAS CONSIDERED A NATURAL PROGRESSION OF THE PATIENT'S DISEASE FOR WHICH THEY WERE RECEIVING LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894108&pc=MHX | 9710602-2025-00690 | 2025-04-02 04:00 | Death | GE MEDICAL SYSTEMS (CHINA) CO., LTD (WUXI) | 2025-04-23 04:00 | MHX | B125 | Defective Alarm | Insufficient Information | K171580 |  | 1 | Event Description: WHILE BEING USED TO MONITOR A PATIENT DURING TRANSPORT, THE BATTERY FAILED AND THE MONITOR SHUT DOWN. IT WAS ALLEGED THAT THE MONITOR DID NOT ALARM FOR LOW BATTERY. THE PATIENT EXPIRED. Manufacturer Narrative: LEGAL MANUFACTURER: HCS TROUT - 3114 N GRANDVIEW BLVD. USA WAUKESHA, WI 53188. A1-6: NOT PROVIDED BY CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894125&pc=DSQ | 2916596-2025-02473 | 2025-01-09 05:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO ISCHEMIC STROKE AND GASTROINTESTINAL BLEEDING. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER WAS NOT AUTHORIZED TO PROVIDE ANY FURTHER INFORMATION ON THE PATIENT; HOWEVER, IT WAS COMMUNICATED THAT THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS BLEEDING, STROKE, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894514&pc=DRB | 3007284006-2025-00081 | 2025-03-25 04:00 | Death | SPECTRANETICS CORPORATION | 2025-04-23 04:00 | DRB | LLD EZ LEAD LOCKING DEVICE | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Perforation | K142116 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO BACTEREMIA. A SPECTRANETICS LLD EZ LEAD LOCKING DEVICE (LLD EZ) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS GLIDELIGHT LASER SHEATH, THE LEAD WAS REMOVED; HOWEVER, THE PATIENT'S BLOOD PRESSURE DECLINED. RESCUE EFFORTS BEGAN, INCLUDING STERNOTOMY, AND A PERFORATION AT THE SUPERIOR VENA CAVA (SVC)/RA JUNCTION WAS DISCOVERED AND REPAIRED. THE PATIENT WAS TRANSFERRED TO THE ICU, BUT UNFORTUNATELY PASSED AWAY LATER THAT DAY. THIS REPORT CAPTURES THE LLD EZ PROVIDING TRACTION WITHIN THE RV LEAD WHEN THE SVC/ RA JUNCTION PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTED IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. B6) RELEVANT TEST/LABORATORY DATA - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED; THUS, THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE DEVICE WAS DISCARDED; THUS, NO DEVICE EVALUATION COULD BE COMPLETED. H6) GREAT VESSEL PERFORATION, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE LLD. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894525&pc=OZD | 1220648-2025-27777 | 2025-04-02 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE AS AN ELECTIVE PLACEMENT FOR MECHANICAL CIRCULATORY SUPPORT AND THE FOLLOWING DAY THE PATIENT EXPERIENCED A STROKE AND EXPIRED. THE PATIENT PRESENTED WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE PATIENT HAD MULTI-VESSEL DISEASE AND WAS PLANNED FOR CORONARY ARTERY BYPASS GRAFTING WITH INTRAOPERATIVE PLACEMENT OF THE IMPELLA 5.5 DEVICE. THE PATIENT HAD DIFFUSED POST-OP BLEEDING. HEPARIN WAS WITHHELD DUE TO THE BLEEDING WITH PLANS FOR WASHOUT AND EVALUATION OF HEPARIN DRIP THE NEXT MORNING. THE IMPELLA 5.5 DEVICE WAS SUCCESSFULLY IMPLANTED FOR MECHANICAL CIRCULATORY SUPPORT. TRANSESOPHAGEAL ECHOCARDIOGRAM SHOWED THE INLET 5CM TO AORTIC VALVE IN THE OPERATING ROOM. THE FOLLOWING DAY, THE PATIENT'S PUPILS WERE BLOWN. THE PATIENT WAS TAKEN FOR A COMPUTED TOMOGRAPHY SCAN, WHICH SHOWED AN ISCHEMIC STROKE WITH MULTIPLE EMBOLI. THE PATIENT SUBSEQUENTLY EXPIRED THIS SAME DAY. FURTHER INFORMATION NOTED THE IMPELLA WAS PRIMED FOR 20 MINUTES BEFORE INSERTION, THE GRAFT AND SHEATH WERE DE-AIRED TWICE BEFORE INSERTION, AND PRIOR TO IMPELLA PLACEMENT, THE PATIENT HAD AIR BUBBLES IN THE LEFT VENTRICLE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS, ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894698&pc=NIK | 2017865-2025-50825 | 2025-04-08 04:00 | Death | ABBOTT | 2025-04-23 04:00 | NIK | GALLANT DR | Over-Sensing; Pacing Problem; Under-Sensing | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. UPON INVESTIGATION, EPISODES OF OVERSENSING, AUTOMATIC MODE SWITCH, AND UNDERSENSING WERE OBSERVED ON THE DEVICE. THE PHYSICIAN DOES NOT ALLEGE THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. FURTHERMORE, TECHNICAL SUPPORT WAS CONTACTED AND DOES NOT ALLEGE A MALFUNCTION ON THE DEVICE. THE DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894700&pc=LWS | 9614453-2025-01603 | 2024-10-23 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-23 04:00 | LWS | EVERA MRI XT DR SURESCAN | Insufficient Information | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED. THE PATIENT'S SPOUSE INQUIRED AS TO WHY THE DEVICE DID NOT "GO OFF". THE DISPOSITION OF THE DEVICE REMAINS UNKNOWN. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894725&pc=DSQ | 2916596-2025-02414 | 2024-04-28 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21894796&pc=DSQ | 2916596-2025-02484 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO UNKNOWN CAUSE. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21895111&pc=NPW | 2015691-2025-03183 | 2025-02-26 05:00 | Death | EDWARDS LIFESCIENCES | 2025-04-23 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Patient Device Interaction Problem | Heart Block | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE IN TRICUSPID POSITION WHERE ON POST-OPERATIVE DAY 6, AFTER THE EVOQUE PROCEDURE, THE PATIENT DIED. THIS PATIENT WAS DISCHARGED HOME AFTER RECEIVING THE EVOQUE VALVE. THE PATIENT WENT TO THE EMERGENCY ROOM AT AN OUTSIDE SMALLER HOSPITAL IN COMPLETE HEART BLOCK (AV BLOCK), UNABLE TO TRANSCUTANEOUSLY PACE RESULTING IN DEATH. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. ADDITIONAL H6 TYPE OF INVESTIGATION CODE: LABELLING REVIEW. THE COMPLAINT FOR VALVE INTERACTS WITH CONDUCTION SYSTEM WAS CONFIRMED WITH OTHER EMPIRICAL EVIDENCE VIA INFORMATION REPORTED BY EDWARDS CLINICAL SPECIALIST. SIX DAYS AFTER THE VALVE WAS IMPLANTED AND AFTER THE PATIENT EXPERIENCED A COMPLETE ATRIO-VENTRICULAR BLOCK (AVB), THE PATIENT DIED; DUE TO ER UNABLE TO TRANSCUTANEOUSLY IMPLANT A PPM. NO MANUFACTURING NON-CONFORMITIES WERE FOUND DURING DHR REVIEW. SINCE NO EDWARDS DEFECT WAS IDENTIFIED, CORRECTIVE OR PREVENTATIVE ACTIONS ARE NOT REQUIRED. A DEFINITE ROOT CAUSE IS UNABLE TO BE DETERMINED. PER THE INSTRUCTIONS FOR USE (IFU), CONDUCTION SYSTEM DISTURBANCE AND/OR HEART BLOCK WHICH MAY REQUIRE TREATMENT (INCLUDING PERMANENT PACEMAKER) ARE POTENTIAL ADVERSE EVENTS. HEART BLOCK AND OTHER CONDUCTION DISTURBANCES ARE COMMON IN PATIENTS WITH UNDERLYING CARDIOVASCULAR DISEASE AND CAN BE EXACERBATED OR AGGRAVATED WITH STANDARD INTRA-OPERATIVE MEDICATIONS OR ANESTHESIA. SUCH EVENTS ARE RELATED TO THE PATIENT'S UNDERLYING CONDITION. OTHER MECHANISMS FOR CONDUCTION ABNORMALITIES COULD BE RELATED TO CORONARY OBSTRUCTION DUE TO AIR EMBOLISM, CORONARY SPASM AND/OR LOCAL EDEMA AFFECTING THE AV NODE. CONDUCTION DISTURBANCES HAVE ALSO BEEN DOCUMENTED DURING TRANSCATHETER CARDIAC PROCEDURES AS A RESULT OF DIRECT INTERACTION WITH CARDIAC ANATOMY, ESPECIALLY IN PATIENTS WITH UNDERLYING CONDUCTION ABNORMALITIES. OTHER POTENTIAL CAUSES INCLUDE TRAUMA BY A GUIDEWIRE OR DELIVERY SYSTEM. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21895341&pc=NVZ | 2182208-2025-02016 | 2025-03-14 04:00 | Death | MEDTRONIC, INC. | 2025-04-23 04:00 | NVZ | MYCARELINK PATIENT MONITOR | Communication or Transmission Problem | Ventricular Fibrillation | P890003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A LINQ DEVICE EXPERIENCED A TERMINAL EPISODE ON (B)(6) 2025, RESULTING IN DEATH. THE PATIENT PREVIOUSLY HAD A REPLACEMENT DEVICE WHEN THE PREVIOUS DEVICE REACHED THE RECOMMENDED REPLACEMENT TIME (RRT), AND IT APPEARED THAT WHEN THE DEVICE WAS REPLACED, THE PATIENT NEVER PERFORMED AN INITIAL MANUAL TRANSMISSION WITH THE REMOTE MONITOR, WHICH WAS REQUIRED TO PAIR WITH THE NEW DEVICE, SO NO TRANSMISSIONS WERE COMING THROUGH. THE ONLY TRANSMISSIONS ON THE NETWORK WERE FROM THE LINQ MOBILE MANAGER AND CARELINK EXPRESS. THE PATIENT RECORDED A SYMPTOM ON (B)(6) 2025, BUT IT WAS NOT SEEN UNTIL THE LINQ DEVICE WAS INTERROGATED BY CARELINK EXPRESS POST-MORTEM. THERE WERE NO ACTIONS TAKEN ON THE MONITOR. IT WAS NOTED THAT THE MONITOR HAD BEEN PLUGGED IN, SUCCESSFULLY UPDATED THE FIRMWARE, AND REGULARLY CHECKED IN, INDICATING THAT THE MONITOR WAS WORKING AS EXPECTED. THE REMOTE MONITOR WAS DISCARDED.Event Description: IT WAS REPORTED THAT A PATIENT WITH A LINQ DEVICE EXPERIENCED A TERMINAL EPISODE ON (B)(6) 2025, RESULTING IN DEATH. THE PATIENT'S PREVIOUS DEVICE REACHED THE RECOMMENDED REPLACEMENT TIME (RRT), AND IT APPEARED THAT WHEN THE DEVICE WAS REPLACED, THE PATIENT NEVER PERFORMED AN INITIAL MANUAL TRANSMISSION WITH THE REMOTE MONITOR, WHICH WAS REQUIRED TO PAIR WITH THE NEW DEVICE, SO NO TRANSMISSIONS WERE COMING THROUGH. THE ONLY TRANSMISSIONS ON THE NETWORK WERE FROM THE LINQ MOBILE MANAGER AND CARELINK EXPRESS. THE PATIENT RECORDED A SYMPTOM ON (B)(6) 2025, BUT IT WAS NOT SEEN UNTIL THE LINQ DEVICE WAS INTERROGATED BY CARELINK EXPRESS POST-MORTEM. THERE WERE NO ACTIONS TAKEN ON THE MONITOR. IT WAS NOTED THAT THE MONITOR HAD BEEN PLUGGED IN, SUCCESSFULLY UPDATED THE FIRMWARE, AND REGULARLY CHECKED IN, INDICATING THAT THE MONITOR WAS WORKING AS EXPECTED. THE REMOTE MONITOR WAS DISCARDED. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 24-APR-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B5 AND CORRECTED THE DATE THE PATIENT RECORDED A SYMPTOM FROM (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21895356&pc=NVY | 2649622-2025-11006 | 2025-03-25 04:00 | Death | MPRI | 2025-04-23 04:00 | NVY | EPSILA EV¿ MRI SURESCAN¿ | Appropriate Term/Code Not Available | Pain; Discomfort; Numbness; Paresthesia; No Clinical Signs, Symptoms or Conditions | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT AROUND FIVE WEEKS POST IMPLANT, THE EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (EV-ICD) SYSTEM WAS EXPLANTED FOR UNKNOWN REASONS. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT AT THE EXPLANT PROCEDURE THE PATIENT NOTED THAT SINCE EXPLANT THEY EXPERIENCED DIFFICULTIES WITH NUMBING, TINGLING, EXTREME PAIN, AND DISCOMFORT FROM HIS ABDOMINAL AND LEFT THORACIC INCISION SITES. THEY STATED THAT LIFTING THEIR ARMS CAUSED SHOOTING, NUMBING, AND TINGLING DISCOMFORT WHICH INHIBITED THEIR ACTIVITIES OF DAILY LIVING. THE EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (EV-ICD) SYSTEM WAS REPLACED WITH A TRANSVENOUS SYSTEM. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT AT THE EXPLANT PROCEDURE THE PATIENT NOTED THAT SINCE IMPLANT THEY EXPERIENCED DIFFICULTIES WITH NUMBING, TINGLING, EXTREME PAIN, AND DISCOMFORT FROM HIS ABDOMINAL AND LEFT THORACIC INCISION SITES. THEY STATED THAT LIFTING THEIR ARMS CAUSED SHOOTING, NUMBING, AND TINGLING DISCOMFORT WHICH INHIBITED THEIR ACTIVITIES OF DAILY LIVING. THE EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (EV-ICD) SYSTEM WAS REPLACED WITH A TRANSVENOUS SYSTEM. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: CORRECTION B5: "SINCE IMPLANT" MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID DVEA3E4 ((B)(6)); PRODUCT TYPE: 0235-ICD; IMPLANT DATE (B)(6) 2024; EXPLANT DATE (B)(6) 2025, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21895733&pc=NPW | 2015691-2025-03184 |  | Death | EDWARDS LIFESCIENCES | 2025-04-23 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device; Migration | Cardiogenic Shock | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE VALVE IN TRICUSPID POSITION WHERE, IMMEDIATELY UPON DEPLOYMENT, IMPLANT WAS NOTED TO BE HIGH ON THE LATERAL SIDE OF ANNULUS. MODERATE PVL (PARAVALVULAR LEAK) WAS NOTED IN THE LATERAL PORTION. OVERALL, THE VALVE POSITION WAS STABLE. ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SPECIALIST STATED THAT THE PATIENT NEVER LEFT THE HOSPITAL AND HAD A CARDIOGENIC SHOCK. THE PATIENT THEN PASSED AWAY. THE CARDIOGENIC SHOCK WAS ALLEGED TO BE RELATED TO THE EVOQUE VALVE.Event Description: AFTER REVIEWING THE POST-OPERATIVE DAY 2 TTE (DATED (B)(6) 2025), THERE IS CONFIRMATION THAT THE EVOQUE VALVE EMBOLIZED INTO THE RIGHT ATRIUM. THE EVOQUE VALVE IS POSITIONED MORE THAN 10MM FROM THE TV ANNULUS ON THE POSTERIOR, ANTERIOR, AND LATERAL SIDES. THE VALVE IS UNSTABLE, EXHIBITING A ROCKING MOTION INTO THE RIGHT ATRIUM. THERE IS TRACE TRANSVALVULAR TR AND MODERATE TO SEVERE LATERAL PVL. Manufacturer Narrative: ADDITIONAL H6 TYPE OF INVESTIGATION CODES: ANALYSIS OF IMAGES AND LABELLING REVIEW. THE COMPLAINT FOR MALPOSITION - VALVE EMBOLIZES INTO ATRIUM WAS CONFIRMED WITH OBJECTIVE EVIDENCE VIA THE IMAGING EVALUATION FROM THE POD 2 TTE. PER THE EVENT DESCRIPTION, THE PATIENT WAS REPORTED TO BE STABLE FOLLOWING INITIAL EVOQUE DEPLOYMENT, WITH A MODERATE PVL. HOWEVER, REVIEW OF POD2 IMAGING REVEALED THE EVOQUE VALVE POSITIONED LOW ON POSTERIOR, ANTERIOR AND LATERAL SIDES, CAUSING MOST ANCHORS TO LOSE CONTACT WITH THE ANNULUS. THE VALVE WAS OBSERVED TO HAVE SIGNIFICANT ROCKING MOTION INTO THE RIGHT ATRIUM. PATIENT WAS REPORTED TO HAVE PASSED DUE TO CARDIOGENIC SHOCK FROM THE VALVE EMBOLIZATION APPROXIMATELY 1 MONTH FOLLOWING THE CASE. MALPOSITION EVENTS SUCH AS EMBOLIZATION MAY BE CAUSED BY A VARIETY OF FACTORS SUCH AS DEVICE TO NATIVE ANATOMY SIZING (PERIMETER DERIVED DIAMETER), USE FACTORS AND PROCEDURAL FACTORS. PATIENT WAS NOTED TO HAVE A MINOR WEIGHT REDUCTION (9KG) FOLLOWING INITIAL TEE EVALUATION, CHANGES IN WEIGHT CAN POTENTIALLY IMPACT SIZING. IN CONJECTURE, THE PATIENT WAS REPORTED TO HAVE SIGNIFICANT TETHERING ON ALL LEAFLETS. THESE FACTORS MAY HAVE CONTRIBUTED TO THE EMBOLIZATION OF THE EVOQUE VALVE; HOWEVER, A ROOT CAUSE WAS UNABLE TO BE DETERMINED SINCE PROCEDURAL IMAGING WAS NOT PROVIDED FOR REVIEW TO CONFIRM REASON FOR EMBOLIZATION. THE DEVICE HISTORY RECORD REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING AND STERILIZATION INSPECTIONS. NO NONCONFORMANCES RELATED TO THE COMPLAINT EVENT WERE IDENTIFIED. SINCE NO EDWARDS DEFECT WAS IDENTIFIED, CORRECTIVE OR PREVENTATIVE ACTIONS ARE NOT REQUIRED. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, B5, G3, G6, H2, H6 AND H11. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. B2: OTHER SERIOUS- CARDIOGENIC SHOCK
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21895782&pc=LPB | 2029046-2025-01300 | 2025-03-26 04:00 | Death | BIOSENSE WEBSTER INC | 2025-04-23 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Cardiac Arrest; Cardiac Tamponade | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PULMONARY VEIN ISOLATION AND THE PATIENT EXPERIENCED PERICARDIAL EFFUSION THAT REQUIRED PERICARDIOCENTESIS. THE PATIENT ALSO SUFFERED CARDIAC ARREST AND REQUIRED RESUSCITATION. ULTIMATELY, THE PATIENT EXPIRED. IT WAS REPORTED THAT THERE WAS SUCCESSFUL COMPLETION OF PULMONARY VEIN ISOLATION FOLLOWED BY SUBSTRATE MODIFICATION OF THE ANTERIOR WALL OF THE LEFT ATRIUM. FOLLOWING SPONTANEOUS INDUCTION OF ATRIAL TACHYCARDIA, A MAP OF THE RIGHT ATRIUM WAS CREATED, DIAGNOSING ATYPICAL FLUTTER DEPENDENT ON THE TRICUSPID VALVE. DURING THE CREATION OF A CAVOTRICUSPID ISTHMUS (CTI) ABLATION LINE, THE STUDY WAS CLOSED SYSTEMICALLY DUE TO AN ERROR MESSAGE AND COULD LATER BE REOPENED WITHOUT DATA LOSS. SUBSEQUENTLY, AN AUDIBLE STEAM POP WAS NOTED BY THE ATTENDING PHYSICIAN. AFTER A RAPID DROP IN BLOOD PRESSURE, A PERICARDIAL EFFUSION WAS CONFIRMED VIA TRANSTHORACIC ECHOCARDIOGRAPHY (TTE), WHICH WAS SUCCESSFULLY DRAINED. DUE TO LOW BLOOD PRESSURE, THE PATIENT SUBSEQUENTLY REQUIRED RESUSCITATION, AND THE EMERGENCY TEAM WAS CALLED. TO COMPENSATE FOR BLOOD LOSS, THE ASPIRATED BLOOD WAS REINFUSED, AND BLOOD TRANSFUSIONS WERE ADMINISTERED. THE BLEEDING WAS CONTROLLED, AND THE BLOOD GAS ANALYSIS (BGA) OF THE PATIENT SHOWED NO SIGNIFICANT ABNORMALITIES GIVEN THE CIRCUMSTANCES. AFTER A PERIOD OF HEMODYNAMIC STABILITY, THE PATIENT AGAIN REQUIRED RESUSCITATION. A STABLE CIRCULATION COULD NOT BE RESTORED. AFTER SEVERAL HOURS, RESUSCITATION EFFORTS FOR THE PATIENT WERE DISCONTINUED, AND DEATH WAS DECLARED. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS PATIENT CONDITION. PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS GREATER THAN 300 SECONDS. Manufacturer Narrative: THE PRODUCT WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING JOHNSON & JOHNSON MEDTECH PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE DEVICE FEATURES WERE REVIEWED. NO MAGNETIC, FORCE, IRRIGATION, DEFLECTION, TEMPERATURE OR ELECTRICAL ISSUES WERE FOUND DURING THE PRODUCT INVESTIGATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE 31528814L NUMBER, AND NO INTERNAL ACTIONS RELATED TO THE REPORTED COMPLAINT CONDITION WERE IDENTIFIED. THE ROOT CAUSE OF THE ADVERSE EVENT AND STEAM POP REMAINS UNKNOWN. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS PATIENT CONDITION. IN ADDITION, NO DEVICE MALFUNCTION WAS REPORTED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE ANALYSIS. THE INSTRUCTION FOR USE (IFU) STATES THAT CAREFUL CATHETER MANIPULATION MUST BE PERFORMED TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. AS PART OF THE JOHNSON & JOHNSON MEDTECH QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21896037&pc=OZD | 1220648-2025-27788 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-04-23 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED DURING USE WITH A 56 YEAR OLD MALE PATIENT, THE AO POSITION WAVEFORM ON THE CONSOLE (AIC) WAS OVERSHOT WHEN THE P VALUE WAS INCREASED TO P8 OR HIGHER. THE AIC WAS REPLACED AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT HARM REPORTED.Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND POST IMPLANT, SAME DAY, IT WAS NOTED WHEN THE P VALUE WAS INCREASED TO PERFORMANCE LEVEL P-8 OR HIGHER, THE AORTIC POSITION WAVEFORM OVERSHOT. CONSEQUENTLY, THE AUTOMATED IMPELLA CONTROLLER WAS EXCHANGED, AND THE ISSUE RESOLVED. HOWEVER, FOLLOWING, APPROXIMATELY SIX DAYS LATER THE PATIENT EXPIRED NOTED FROM CIRCULATORY FAILURE. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT WAS COMPLETED AND NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE AUTOMATED IMPELLA CONTROLLER (AIC) AS A CONTRIBUTING FACTOR IN THE OUTCOME. OPTICAL SIGNAL ISSUE REQUIRED EXCHANGE IN AIC DEVICE, AND THERE WAS A BRIEF PERIOD OF INTERRUPTED SUPPORT TO THE PATIENT. UPDATES WERE MADE TO RESPECTIVE SECTIONS: B1, B2, H1, AND H6 CODING (HEALTH/PRODUCT PROBLEM) ACCORDINGLY GIVEN THE MEDICAL SAFETY REVIEW OUTCOME. ADDITIONALLY, UPDATES WERE MADE TO SECTION D9 AND H6 CODING (INVESTIGATION) ACCORDINGLY, AS THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF THE EVALUATION/ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION B5 WAS UPDATED TO INCLUDE COMPLETE DETAILS OF THE REPORTED EVENT. Manufacturer Narrative: THE INVESTIGATION OF OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. BASED ON THE DEVICE ANALYSIS, DATA ANALYSIS, AND THE CLINICAL DETAILS; THE ROOT CAUSE FOR THE PLACEMENT SIGNAL NOT RELIABLE ALARM (OPTICAL SIGNAL ISSUE) WAS DUE TO THE DEFECTIVE OPTICAL BENCH WITH LOW SIGNAL NOT RELIABLE (SNR). B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27788. D4 LOT NUMBER WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27788. D6A SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27788. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27788 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H5 WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27788. H8 WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27788. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21896052&pc=DSQ | 2916596-2025-02488 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-04-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED.Event Description: THE PATIENT HAD MULTIPLE MYELOMA WHICH WAS TREATED WITH IMMUNOSUPPRESSANTS WHICH MADE THE PATIENT MORE SUSCEPTIBLE TO INFECTIONS. THE ORIGIN OF THE SEPSIS WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21878931&pc=OZD | 1220648-2025-27718 | 2025-03-28 04:00 | Death | ABIOMED, INC. | 2025-04-22 04:00 | OZD | IMPELLA CP | Malposition of Device | Ischemia; Obstruction/Occlusion; Implant Pain | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND THE 14FR. SHEATH WAS 100% OCCLUDING THE ARTERY. THE PATIENT WAS SCREAMING IN PAIN FROM THE RIGHT LOWER EXTREMITY. THE PATIENT WAS GIVEN PAIN MEDICATION AND EVENTUALLY INTUBATED AS DURING THE PROCEDURE THE PATIENT WAS THRASHING AND ALMOST FELL OFF THE TABLE. THE WIRE POSITION WAS LOST; HOWEVER, AFTER INTUBATION, THE PHYSICIAN WAS ABLE TO COMPLETE THE REVASCULARIZATION. POST STENTING PEEL-AWAY SHEATH WAS REMOVED. RIGHT LOWER EXTREMITY ILIAC ANGIOGRAM WAS PERFORMED WITH REPOSITIONING SHEATH SIDE ACCESS. THE FLOW WAS RESTORED TO EXTREMITY. FOLLOWING, IN THE MORNING OVERNIGHT, THE PATIENT'S URINE BLOOD WAS TINGED AND THE ECHOCARDIOGRAM SHOWED IMPELLA PLACEMENT DEEP. THE PHYSICIANS DECIDED TO WAIT UNTIL LATER IN THE MORNING TO ADDRESS THE POSITIONING AND POSSIBLY REPOSITION. PULSATILITY WAS RETURNING TO IMPELLA. ALSO, NO PULSES COULD BE FOUND DISTALLY; THEREFORE, THE PLAN WAS NOTED TO HAVE A VASCULAR ULTRASOUND LATER IN THE MORNING. THAT AFTERNOON IT WAS NOTED THE PATIENT WENT ASYSTOLE FOUR TIMES AND CHEMICAL CODE WAS GIVEN TO ACHIEVE ELECTRICAL ACTIVITY. PATIENT WAS HAVING SIGNS OF MULTI-SYSTEM ORGAN FAILURE. THE PATIENT'S FEET REMAIN COLD BILATERALLY. THE PATIENT WOULD ULTIMATELY EXPIRE THE FOLLOWING DAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21878934&pc=OZD | 1220648-2025-27719 | 2025-03-29 04:00 | Death | ABIOMED, INC. | 2025-04-22 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Shock | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE VIA THE RIGHT AXILLARY ARTERY FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT WAS ALSO CANNULATED FOR EXTRACORPOREAL MEMBRANE OXYGENATION. THE FOLLOWING DAY IT WAS NOTED THE PATIENT WAS BLEEDING FROM CUTDOWN SITE AND A DEVELOPED A "MASSIVE" HEMATOMA FROM THE UPPER CHEST TO SHOULDER AND NECK. THREE UNITS OF BLOOD PRODUCTS WERE ADMINISTERED. OF NOTE, THE PATIENT WAS ALSO BLEEDING FROM GASTROINTESTINAL WITH OVER 300CC SUCTIONED FROM THE NASOGASTRIC TUBE. CONSULTS REQUESTED FROM CARDIOTHORACIC SURGERY AND GASTROINTESTINAL PHYSICIAN. IT WAS NOTED THE TEAM BELIEVED THE PATIENT WAS ENTERING HEMORRHAGIC SHOCK. THE PATIENT WAS NOTED NOT A CANDIDATE FOR ADVANCE OPTIONS. FOLLOWING THE EVENT THE PATIENT WAS NOTED TO BE IN CRITICAL CONDITION. WITHDRAW OF CARE DISCUSSION WAS PLANNED IN THE EVENT THE PATIENT DID NOT WAKE UP. PROGNOSIS REMAINED VERY POOR, NEUROLOGICAL STATUS WAS UNKNOWN, AND THE PATIENT WOULD EXPIRE APPROXIMATELY NINE DAYS LATER. CARE WAS WITHDRAWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, AS THERE WAS BLEEDING FROM THE ACCESS SITE AND PATIENT ALSO GASTROINTESTINAL BLEEDING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21879065&pc=MNS | 2518422-2025-104535 | 2025-04-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21879087&pc=MNS | 2518422-2025-104531 | 2025-04-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE TOTAL BLOWER HOURS WERE 1452.3 WITH NO ERRORS OR REPORTED FAULTS. THE BLOWER BOX ASSEMBLY WAS CONTAMINATED AND REPLACED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED AS PREVENTIVE MAINTENANCE. IN ADDITION, THE SD CARD WAS FITTED. THE DEVICE WAS CLEANED AND TESTED. THE DEVICE PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21879105&pc=BZD | 2518422-2025-104538 | 2025-03-04 05:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | BZD | DREAMSTATION CPAP PRO | Degraded; Adverse Event Without Identified Device or Use Problem | Hypoxia; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS REPORTED AS BEING RETURNED TO THE DME TWO YEARS AGO, HOWEVER THE DEVICE WAS NEVER RECEIVED BY THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS REPORTED AS BEING RETURNED TO THE DME TWO YEARS AGO, HOWEVER THE DEVICE WAS NEVER RECEIVED BY THE MANUFACTURER. THE MANUFACTURER RECEIVED ADDITIONAL INFORMATION FROM THE PATIENT'S DAUGHTER REGARDING THE EL SETTLEMENT PAYMENT IN REFERENCE TO THE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. SHE EXPRESSED, "MY DAD SHOULD NOT HAVE PASSED IF HE RECEIVED HIS NEW MACHINE ON TIME. WE TRIED FOR MONTHS." THE DAUGHTER SHARED THAT HER FATHER HAD BEEN USING THE CPAP DEVICE FOR APPROXIMATELY 4 OR 5 YEARS BEFORE HIS DEATH. SHE EXPLAINED THAT HER FATHER REGULARLY CLEANED HIS DEVICE ACCORDING TO THE INSTRUCTIONS, USING THE RECOMMENDED CLEANING SOLUTIONS AND WATER. WHILE SHE WAS UNSURE IF HE USED A SOCLEAN OR A SIMILAR DEVICE, SHE NOTED THAT HE ALWAYS KEPT UP WITH THE LATEST EQUIPMENT, SO IT IS POSSIBLE HE DID. THE DAUGHTER FURTHER MENTIONED THAT HER FATHER NEVER APPEARED WELL-RESTED, OFTEN REQUIRING OXYGEN AT 3 OR 4 LITERS AND USED THE CPAP MACHINE AT NIGHT. SHE RECALLED THAT HE HAD USED THE CPAP THE NIGHT BEFORE HIS PASSING BUT WAS STRUGGLING TO GET IT TO WORK PROPERLY. UPON WAKING, HIS OXYGEN LEVELS WERE IN THE 70S, AND HE HAD EXPERIENCED DIFFICULTIES WITH THE CPAP DEVICE OVER THE PAST FEW NIGHTS. SHE STATED THAT HER FATHER ALWAYS SEEMED TIRED AND GROGGY, AND ALTHOUGH HE INTENDED TO CONTACT PHILIPS ABOUT HIS DEVICE, SHE BELIEVES HE NEVER MANAGED TO DO SO BEFORE HE FELL ILL AND PASSED AWAY. THE DAUGHTER DESCRIBED A PARTICULARLY DIFFICULT NIGHT BEFORE HIS PASSING, WHEN HER FATHER HAD TROUBLE USING THE MACHINE, COULD NOT SLEEP WITH IT, AND FOUND IT MALFUNCTIONING. SHE ALSO REPORTED THAT HE RELIED ON OXYGEN, THE CPAP MACHINE, AND A NEBULIZER. THE DAUGHTER STATED HER FATHER PASSED AWAY FROM COPD AND "ALL THAT STUFF" RESTRICTIVE AIRWAYS, CHRONIC COPD, THE WHOLE REASON HE WAS ON THE NEBULIZER AND OXYGEN. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX A AGE AND SEX IN THIS REPORT. THE MANUFACTURER HAS UPDATED BOX B DATE OF DEATH AND OTHER RELEVANT HISTORY IN THIS REPORT. THE MANUFACTURER HAS UPDATED PRODUCT PROBLEM OR ADVERSE EVENT TO BOTH. THE PREVIOUSLY SUBMITTED REPORT WAS FILED AS AN ADVERSE EVENT ONLY. THE MANUFACTURER HAS UPDATED BOX E INITIAL REPORTER IN THIS REPORT. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21879107&pc=MNS | 2518422-2025-104533 | 2025-04-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21880205&pc=NVZ | 3008973940-2025-02814 | 2025-03-11 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-22 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT DIED UNEXPECTEDLY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM.Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT DIED UNEXPECTEDLY DUE TO SUDDEN CARDIAC ARREST. IT WAS LATER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: CORRECTED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: ETLW1613C93E STENT GRAFT IMPLANTED ON (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21880234&pc=NVN | 2649622-2025-10745 | 2025-03-11 04:00 | Death | MPRI | 2025-04-22 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT DIED UNEXPECTEDLY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM.Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT DIED UNEXPECTEDLY DUE TO SUDDEN CARDIAC ARREST. IT WAS LATER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: ETLW1613C93E STENT GRAFT IMPLANTED ON (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21880269&pc=DTB | 2649622-2025-10748 | 2025-03-11 04:00 | Death | MPRI | 2025-04-22 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT DIED UNEXPECTEDLY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM.Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT DIED UNEXPECTEDLY DUE TO SUDDEN CARDIAC ARREST. IT WAS LATER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: CORRECTED: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: ETLW1613C93E STENT GRAFT IMPLANTED ON (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21881315&pc=NPT | 2025587-2025-02886 | 2023-12-27 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Respiratory Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS RESPIRATORY FAILURE. CONTRIBUTING FACTORS TO THE RESPIRATORY FAILURE WERE UNSPECIFIED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21881376&pc=QNR | 3003306248-2025-00138 | 2025-02-04 05:00 | Death | THORATEC SWITZERLAND GMBH | 2025-04-22 04:00 | QNR | CENTRIMAG BLOOD PUMP | Gas/Air Leak; Adverse Event Without Identified Device or Use Problem | Air Embolism | K020271 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO). ON (B)(6) 2025 AT APPROXIMATELY 23:32, IT WAS IDENTIFIED THAT THE VENOUS CANNULA CRACKED AND DISCONNECTED FROM CIRCUIT. A CIRCUIT EXCHANGE WAS PERFORMED IN ATTEMPT TO STOP EXSANGUINOUS DRAINAGE. DURING TIME OF CIRCUIT, THE PATIENT CODED AND WAS PLACED BACK ON VV ECMO BUT BUBBLES WERE IDENTIFIED IN THE CIRCUIT WITH INABILITY TO GET OUT WITHOUT CAUSING ADDITIONAL HARM TO PATIENT. THE PATIENT PASSED AWAY AT 00:05 ON (B)(6) 2025 WHEN THE CENTRIMAG BLOOD PUMP WAS "BUMPED BY STAFF" WHICH CAUSED AIR TO GET IN THE CIRCUIT. THE PATIENT WAS NON-PULSATILE ALREADY AND THE AIR CAUSED THE PATIENT TO PASS AWAY. PER THE SITE, GIVEN THE CIRCUMSTANCES, IT WAS BELIEVED THAT THE HARDWARE WAS NOT AT FAULT AND WOULD NOT BE RETURNED. THE STAFF IN THE ROOM "KNOCKED INTO THE PUMP" ALLEGEDLY CAUSING A CRACK WHICH REQUIRED SWITCHING THE CIRCUIT DUE TO THE PATIENT EXSANGUINATING. A NEW PRE-PRIMED CIRCUIT WAS INITIATED, AND AIR WAS IDENTIFIED POST INITIATION. THE PATIENT DECEASED BECAUSE THE AIR WAS UNABLE TO BE SAFELY REMOVED. THE STAFF IN THE ROOM "KNOCKED INTO THE PUMP" ALLEGEDLY CAUSING A CRACK WHICH REQUIRED SWITCHING THE CIRCUIT DUE TO THE PATIENT EXSANGUINATING. A NEW PRE-PRIMED CIRCUIT WAS INITIATED, AND AIR WAS IDENTIFIED POST INITIATION. THE PATIENT DECEASED BECAUSE THE AIR WAS UNABLE TO BE SAFELY REMOVED. THE PUMP WAS A PRE-PRIMED CIRCUIT THAT WAS READY FOR EMERGENCIES PRIOR TO THE EVENT. PER TWO SEPARATE PHONE CALLS, AIR WAS IDENTIFIED IN THE CIRCUIT BEFORE INITIATING SUPPORT AND LIKELY ENTERED THE SYSTEM DURING PRIMING. AFTER FOLLOW-UP THE ECMO COORDINATOR COMMUNICATED "THE CHANGE OUT WAS ATTEMPTED BUT WAS UNSUCCESSFUL BECAUSE THE PATIENT [HAD PASSED AWAY]¿.Event Description: THE ECMO SPECIALIST AND REGISTERED NURSE (RN) WERE BATHING THE PATIENT, AFTER THE BATH, THE RN WALKED PAST THE MACHINE AND BUMPED IT WITH THEIR HIP. THE MACHINE THEN BEGAN ALARMING LOW FLOW AND THE ECMO SPECIALIST RESPONDED AND NOTICED EXSANGUINATION COMING FROM THE TUBING. REGARDING THE SECOND CIRCUIT THAT COULD NOT BE PLACED ON THE PATIENT, IT WAS UNABLE TO BE IDENTIFIED HOW THERE WAS AIR IN IT. IT BEGAN ALARMING WHEN THERE WERE BUBBLES AFTER PERFORMING THE CIRCUIT EXCHANGE Manufacturer Narrative: D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORT OF AIR IN THE CIRCUIT COULD NOT BE CONCLUSIVELY DETERMINED, AND A DIRECT CORRELATION BETWEEN THE CENTRIMAG BLOOD PUMP AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PATIENT REQUIRED AN EMERGENT CIRCUIT EXCHANGE TO A PRE-PRIMED BACKUP CIRCUIT WITH VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO). DURING THE EXCHANGE, THE PATIENT CODED. AFTER INITIATING SUPPORT ON THE BACKUP CIRCUIT, AIR WAS REPORTEDLY IDENTIFIED IN THE CIRCUIT THAT WAS UNABLE TO BE REMOVED WITHOUT CAUSING ADDITIONAL HARM TO THE PATIENT, AND THE PATIENT EXPIRED. IT WAS REPORTED THAT THE AIR HAD LIKELY ENTERED THE SYSTEM DURING PRIMING. IT WAS REPORTED THAT THE CENTRIMAG BLOOD PUMP WOULD NOT BE RETURNED FOR EVALUATION. THE CENTRIMAG BLOOD PUMP INSTRUCTIONS FOR USE (IFU) LISTS DEATH AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE CENTRIMAG CIRCULATORY SUPPORT SYSTEM UNDER ¿ADVERSE EVENTS¿. THIS IFU ALSO PROVIDES THE FOLLOWING WARNINGS AND CAUTIONS: IFU WARNING #2: ENSURE THAT THE PUMP AND CIRCUIT HAVE BEEN DEBUBBLED AND PRIMED PROPERLY PRIOR TO USE TO MINIMIZE THE RISK OF AIR ENTRY TO THE PATIENT. IFU WARNING #3: MASSIVE AIR ENTRY INTO THE PUMP WILL CAUSE THE PUMP TO DEPRIME AND BLOOD FLOW TO STOP. CLAMP THE OUTLET TUBING, STOP THE PUMP, AND REMOVE AIR PRIOR TO RESUMING CIRCULATION. IFU WARNING #10: FREQUENT PATIENT AND DEVICE MONITORING IS RECOMMENDED. IFU WARNING #14: DO NOT OPERATE THE PUMP WITH ITS INLET TUBING CLAMPED AS A NEGATIVE PRESSURE WILL BE GENERATED IN THE PUMP AND AIR BUBBLES MAY BE FORMED IN THE PRIMING FLUID OR BLOOD. IFU WARNING #16: AS WITH ALL CONTINUOUS FLOW PUMPS, OPERATING AT TOO HIGH A SPEED CAN RESULT IN NEGATIVE PRESSURE AT THE INLET WHICH CAN LEAD TO COLLAPSE OF THE VENTRICLE OR BLOOD VESSELS, INLET CANNULA OBSTRUCTION, INSPIRATION OF AIR, OUTGASSING, CAVITATION AND INCREASED RISK OF EMBOLISM. ALWAYS OPERATE THE SYSTEM AT THE LOWEST SPEED CONSISTENT WITH THE VOLUME OF BLOOD AVAILABLE TO BE PUMPED AND CLINICALLY ACCEPTABLE CIRCULATORY SUPPORT. IFU CAUTION #2: THIS DEVICE SHOULD ONLY BE USED BY PERSONS THOROUGHLY TRAINED IN EXTRACORPOREAL CIRCULATION PROCEDURES. IFU CAUTION #15: ALWAYS HAVE A BACKUP CENTRIMAG PUMP, CONSOLE, MOTOR, AND ACCESSORIES AVAILABLE FOR USE. THE SECTION TITLED ¿PUMP SETUP AND OPERATION¿ PROVIDES INSTRUCTIONS FOR INSERTING THE CANNULA AND DE-AIRING THE SYSTEM. THIS SECTION WARNS ¿DO NOT OVER TIGHTEN SUTURES WHEN SECURING CANNULAE TO TISSUES AND VESSELS. OVER TIGHTENED SUTURES MAY RESULT IN OBSTRUCTION AND INTERRUPTION OF BLOOD FLOW THROUGH THE CANNULAE. SUTURING USED TO SECURE CANNULA MUST BE MADE WITH SUFFICIENT TENSION TO HOLD THE CANNULA IN PLACE OVER THE FULL RANGE OF PATIENT ACTIVITY. FAILURE TO EFFECTIVELY SECURE CANNULAE IN PLACE POSES RISK OF DECANNULATION, BLEEDING, OR AIR EMBOLUS.¿ THIS SECTION ALSO WARNS ¿IF LEAKS OR OTHER ANOMALIES ARE FOUND IN THE CIRCUIT, REMOVE THE PUMP AND REPLACE WITH A NEW, STERILE PUMP, REPEATING THE ABOVE STEPS TO PRIME.¿ NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D2B: CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: CORRECTED DATA: SECTION H10: MEDWATCH NUMBER ADDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21881440&pc=NPT | 2025587-2025-02891 | 2024-11-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Renal Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 8 MONTHS AND 2 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. PER THE PATIENT'S FAMILY, THE CAUSE OF DEATH WAS RENAL FAILURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS NOT RELATED TO THE PATIENT'S UNDERLYING MEDICAL HISTORY. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21881652&pc=DQY | 2124215-2025-25174 | 2025-04-17 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-22 04:00 | DQY | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Bradycardia; Stroke/CVA; Loss of consciousness | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT AIR EMBOLISM AND CEREBRAL VASCULAR ACCIDENT (CVA) OCCURRED. THE PROCEDURE WAS CANCELLED. THE PATIENT DIED. A LEFT ATRIAL APPENDAGE CLOSURE (LAA) PROCEDURE WAS BEING PERFORMED USING INTRACARDIAC ECHOCARDIOGRAM (ICE) UNDER LOCAL AND CONSCIOUS SEDATION. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS INSERTED AND ADVANCED INTO THE LEFT ATRIUM (LA) ALONG WITH THE ICE CATHETER. AIR WAS THEN VISUALIZED IN THE LA ON ICE, BUT NOT FLUOROSCOPY AS IT WAS NOT IN USE DURING TIME OF WAS INSERTION. THE PATIENT BECAME BRADYCARDIC. HIGH FLOW OXYGEN WAS APPLIED, AND THE HEAD OF THE PATIENT'S BED RAISED. THE PHYSICIAN CONTINUED THE PROCEDURE. A 24MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY (WDS) SYSTEM WAS INSERTED AND ADVANCED TO THE LAA. THE PHYSICIAN ENCOUNTERED CHALLENGES WITH PATIENT ANATOMY AND THE LAA AND DECIDED IT WAS BEST TO ABORT THE PROCEDURE. AFTER THE CONCLUSION OF THE PROCEDURE, THE PATIENT WAS NOT RESPONDING TO NEUROLOGICAL STATUS CHECKS. A RAPID RESPONSE WAS CALLED, AND THE PATIENT HAD A COMPUTED TOMOGRAPHY (CT) SCAN PERFORMED WHICH FOUND A CVA. THE PATIENT WAS INTUBATED AND TAKEN TO THE INTENSIVE CARE UNIT (ICU). THE PLAN WAS TO TRANSFER THE PATIENT TO ANOTHER FACILITY TO TRY HYPERBARIC CHAMBER INTERVENTIONS, BUT THE PATIENT DIED ONE DAY POST INDEX PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21881899&pc=NGV | 2124215-2025-25112 | 2025-04-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-22 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 35MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED WITH NO RECAPTURES. IT WAS NOTED THERE WAS 1/3 SHOULDER IN THE 135-DEGREE VIEW ON IMAGING, BUT THE CLOSURE DEVICE HAD MET RELEASE CRITERIA. THE PATIENT WAS DISCHARGED AND DOING WELL. THE FOLLOWING DAY POST INDEX PROCEDURE, THE PATIENT STARTED TO FEEL SICK. TWO DAYS POST INDEX PROCEDURE, THE PATIENT DIED IN THEIR SLEEP. PER THE PHYSICIAN, THE CAUSE OF DEATH REMAINS UNKNOWN. NO FURTHER DETAILS WERE MADE AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21881970&pc=OZD | 1220648-2025-27727 | 2025-04-02 04:00 | Death | ABIOMED, INC. | 2025-04-22 04:00 | OZD | IMPELLA CP | Positioning Failure; Malposition of Device | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE IMPELLA CP WAS NOT ABLE TO GO ABOVE P3 WITHOUT SUCTION ALARMS. THE IMPELLA CP WAS REPOSITIONED THE NEXT DAY AS MULTIPLE POSITIONING ALARMS OCCURRED OVERNIGHT. ON DAY FOUR OF IMPELLA SUPPORT, THE IMPELLA CP WAS REPOSITIONED AGAIN. ON DAY SEVEN OF IMPELLA CP SUPPORT, THE PATIENT EXPERIENCED TORSADE REQUIRING CARDIOPULMONARY RESUSCITATION WITH RETURN OF SPONTANOUS CIRCULATION ACHIEVED. THE IMPELLA APPEARED MALPOSITION IN THE CHEST X-RAY AND WAS REPOSITIONED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ON IMPELLA SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE CAUSE OF THE POSITIONING ISSUES MOST LIKELY WAS USE RELATED DUE TO CARDIOPULMONARY RESUSCITATION (CPR). AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE TORSAD ARRHYTHMIA WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882008&pc=NVZ | 3004209178-2025-06937 |  | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-22 04:00 | NVZ | ADAPTA | Appropriate Term/Code Not Available | Sepsis; Septic Shock; Respiratory Failure | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. FOLLOW-UP WAS ATTEMPTED TO OBTAIN THE RELEVANT INFORMATION; WITH THE CAUSE OF DEATH COULD NOTED TO BE SEPSIS WITH ACUTE RESPIRATORY FAILURE. THIS EVENT IS BEING CONSERVATIVELY REPORTED IN AN ABUNDANCE OF CAUTION IN RESPONSE TO THE ADVISORY DESCRIBING THE POTENTIAL FOR DEVICE CIRCUIT ERROR TO OCCUR WHILE THE DEVICE IS PROCESSING AN ATRIAL-SENSED EPISODE. THIS PATIENT WAS CONFIRMED TO BE PROGRAMMED TO A PACING MODE SUSCEPTIBLE TO THE DESCRIBED ADVISORY ISSUE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED SEPTIC SHOCK DUE TO AN UNKNOWN ORGANISM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. ANALYSIS OF THE DEVICE REVEALED NORMAL BATTERY DEPLETION. THE BATTERY INDICATOR SIGNIFYING THAT IT IS TIME FOR DEVICE REPLACEMENT OCCURRED AFTER EXPLANT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882055&pc=LJT | 1317056-2025-00101 | 2023-03-31 04:00 | Death | ANGIODYNAMICS, INC | 2025-04-22 04:00 | LJT | SMART PORT CT | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Sepsis; Septic Shock | K062414 |  | 1 | Event Description: ON OR ABOUT (B)(6) 2023, (B)(6). UNDERWENT PLACEMENT OF AN ANGIODYNAMICS SMART PORT CT PRODUCT, MODEL NO. H787CT80STPDVI1; LOG NO. 4799362. THE DEVICE WAS IMPLANTED BY (B)(6), MD., AT (B)(6) HOSPITAL, TO BE USED FOR CHEMOTHERAPY. ON OR ABOUT (B)(6) 2023, (B)(6) PRESENTED HERSELF TO (B)(6) HOSPITAL WITH A CHIEF COMPLAINT OF PORT INFECTION, WHERE SHE WAS DIAGNOSED WITH SEPTIC SHOCK SECONDARY BLOODSTREAM INFECTION PORT-A-CATH WITH PULMONARY SEPTIC EMBOLI AND EXPIRED 24 HOURS AFTER ADMISSION INTO HOSPITAL. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT SEPSIS INFECTION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). NO PORT DEVICE WAS RETURNED FOR EVALUATION SINCE THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION, DAMAGE OR PERFORMANCE ISSUE DURING USE. A DEVICE HISTORY RECORD (DHR) REVIEW COULD BE PERFORMED SINCE THERE WAS NO LEGITIMATE REPORTED LOT NUMBER AND REPORTED PACKAGED GOOD ITEM NUMBER WAS NOT CONFIRMED TO BE SHIPPED TO THE REPORTED END USER FACILITY. THE PORT WAS IMPLANTED INTO THE PATIENT 45 DAYS PRIOR TO THE PATIENT'S INFECTION. THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING USE. IF THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE IN QUESTION WAS TO CONTRIBUTE TO AN INFECTION IN THE PATIENT IT WOULD HAVE LIKELY OCCURRED SHORTLY AFTER THE IMPLANT DATE AND NOT 45 DAYS LATER. DEVICE DIRECTIONS FOR USE CAUTIONS THAT EACH ACCESS OF THE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE, FOLLOWING THE INSTITUTIONS UNIVERSAL PRECAUTIONS. THERE IS NO INDICATION FROM THE REPORTED COMPLAINT THAT THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE COULD HAVE CONTRIBUTED TO THE PATIENT'S INFECTION. INFECTION IS CAUTIONED IN THE DEVICE DFU AS A POTENTIAL COMPLICATION OF PORT SYSTEM USE. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED IN THE SMARTPORT DEVICE CONTAINS THE FOLLOWING DIRECTIONS AND PRECAUTIONS: CONTRAINDICATIONS CATHETER INSERTION IN THE SUBCLAVIAN VEIN MEDIAL TO THE BORDER OF THE FIRST RIB, AN AREA WHICH IS ASSOCIATED WITH HIGHER RATES FOR PINCH-OFF. WARNINGS: THE DEVICE IS TO BE IMPLANTED, USED, MAINTAINED, AND REMOVED IN STRICT ACCORDANCE WITH INSTITUTIONAL AND OR CENTERS FOR DISEASE CONTROL (CDC) GUIDELINES OR POLICIES. REUSE OF SINGLE-USE DEVICES CREATES A POTENTIAL RISK OF PATIENT OR USER INFECTIONS. CONTAMINATION OF THE DEVICE MAY LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. DO NOT USE SYRINGES SMALLER THAN 10 ML SYRINGE WHEN ACCESSING THE PORT AS SYSTEM DAMAGE CAN OCCUR. FLUSHING OCCLUDED CATHETERS WITH SMALL SYRINGES CAN CREATE EXCESSIVE PRESSURES WITHIN THE PORT SYSTEM. ABSENCE OF A BLOOD RETURN OR A POOR BLOOD RETURN CAN BE A SIGN OF A POTENTIAL COMPLICATION SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, FIBRIN FORMATION, THROMBOSIS OR MALPOSITION. THIS SHOULD BE EVALUATED PRIOR TO DEVICE USAGE. A BLOOD RETURN SHOULD BE PRESENT PRIOR TO USAGE OF DEVICE FOR ANY THERAPY OR TESTING. DO NOT ATTEMPT TO MEASURE THE PATIENT'S BLOOD PRESSURE ON THE ARM IN WHICH A PERIPHERAL SYSTEM IS LOCATED, SINCE CATHETER OCCLUSION OR OTHER DAMAGE TO THE CATHETER COULD OCCUR. IF THE PATIENT COMPLAINS OF PAIN, OR THERE IS SWELLING WHEN THE DEVICE IS FLUSHED OR WHEN MEDICATION OR CONTRAST MEDIA IS ADMINISTERED, EVALUATE THE DEVICE FOR INFILTRATION, PROPER NEEDLE PLACEMENT, AND POTENTIAL COMPLICATIONS SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, THROMBOSIS OR MALPOSITION. FAILURE TO ASSESS THESE COMPLAINTS OR OBSERVATIONS CAN LEAD TO DEVICE FAILURE. PRECAUTIONS CAREFULLY READ AND FOLLOW ALL INSTRUCTIONS PRIOR TO USE. STRICT ASEPTIC TECHNIQUE IS OF PARAMOUNT IMPORTANCE WHEN IMPLANTING ANY DEVICE. POTENTIAL COMPLICATIONS: USE OF AN ANGIODYNAMICS PORT SYSTEM INVOLVES POTENTIAL RISKS NORMALLY ASSOCIATED WITH THE INSERTION OR USE OF ANY IMPLANTED DEVICE OR INDWELLING CATHETER INCLUDING BUT NOT LIMITED TO: EROSION OF VESSEL AND SKIN, IMPLANT REJECTION, INFECTION, NECROSIS OF SCARRING OF SKIN OVER IMPLANT AREA, VESSEL TRAUMA. POST-OPERATIVE CARE: THE INCISION SITE SHOULD BE MONITORED FOR SIGNS OF INFECTION, INFLAMMATION, HEMATOMA, DEVICE ROTATION OR EROSION. ROUTINE WOUND CARE SHOULD BE GIVEN TO THESE SITES. THE SMART PORT CT IMPLANTABLE PORT MAY BE USED IMMEDIATELY AFTER VERIFICATION OF CATHETER PLACEMENT. INSTRUCT PATIENT TO AVOID ANY HEAVY EXERTION OR STRENUOUS ACTIVITY DURING THE FIRST FEW DAYS AFTER SURGERY. GENERAL GUIDELINES: EACH ACCESS OF AN ANGIODYNAMICS SMART PORT CT IMPLANTABLE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE. FOLLOW INSTITUTIONAL UNIVERSAL PRECAUTIONS. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4). Manufacturer Narrative: THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882068&pc=DSQ | 2916596-2025-02264 | 2025-03-22 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE (MOF). THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE WOULD NOT RETURN. THE CAUSE OF DEATH WAS PROGRESSION OF MULTIORGAN FAILURE DESPITE SUPPORT AND LUNG BLEEDING. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AND BLEEDING AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882179&pc=DSQ | 2916596-2025-02435 | 2025-04-02 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882245&pc=LWS | 1028232-2025-01916 | 2025-03-11 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-04-22 04:00 | LWS | ACTICOR 7 VR-T DX DF4 PROMRI | Insufficient Information | Ventricular Fibrillation | P050023 |  | 1 | Event Description: DEVICE REPORTED EOS ON (B)(6) 2025 AFTER SHOCKS WERE DELIVERED FOR VF EPISODES SAME DAY. NO EXTERNAL SHOCKS WERE DOCUMENTED. BATTERY WAS REPORTED AS 100 PERCENT ON (B)(6) 2025. THE DEVICE WAS NOT RESET. THE PATIENT PASSED AWAY ON (B)(6) 2025 AND THE CAUSE OF DEATH WAS LISTED AS AORTIC DISSECTION, HYPERTENSION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882539&pc=FKX | 0008030665-2025-00915 | 2025-04-17 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-22 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: ON (B)(6) 2025 IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY THAT MORNING WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. ADDITIONAL INFORMATION WAS PROVIDED THROUGH FOLLOW-UP WITH THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS FOUND DECEASED BY THEIR SPOUSE ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS UNKNOWN WHAT PHASE AND CYCLE OF THE PATIENT¿S TREATMENT THEY WERE IN AT THE TIME OF DEATH. NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT. THE PATIENT¿S SPOUSE REPORTED THAT EVERYTHING WAS FINE PRIOR TO THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. IT WAS REPORTED THAT CARDIAC ISSUES ARE SUSPECTED. THE DEATH IS NOT RELATED TO THE USE OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY.Event Description: ON (B)(6) 2025 IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY THAT MORNING WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. ADDITIONAL INFORMATION WAS PROVIDED THROUGH FOLLOW-UP WITH THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS FOUND DECEASED BY THEIR SPOUSE ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS UNKNOWN WHAT PHASE AND CYCLE OF THE PATIENT¿S TREATMENT THEY WERE IN AT THE TIME OF DEATH. NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT. THE PATIENT¿S SPOUSE REPORTED THAT EVERYTHING WAS FINE PRIOR TO THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. IT WAS REPORTED THAT CARDIAC ISSUES ARE SUSPECTED. THE DEATH IS NOT RELATED TO THE USE OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. ALTHOUGH NO CAUSE OF DEATH WAS DOCUMENTED, CARDIAC ISSUES WERE SUSPECTED. CARDIOVASCULAR DISEASE IS THE LEADING CAUSE OF DEATH IN PATIENTS WITH CHRONIC KIDNEY DISEASE (CKD) WITH AN ESTIMATED 40%-50% OF ALL DEATHS IN PATIENTS WITH CKD STAGE 4 AND 5 BEING ATTRIBUTABLE TO CARDIOVASCULAR-RELATED MORTALITY. THIS PATIENT HAD SEVERAL COMORBID CARDIAC DISEASES INCLUDING CHRONIC SYSTOLIC HEART FAILURE. BASED ON THE AVAILABLE INFORMATION AND NO EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882552&pc=GEI | 3005075853-2025-02844 | 2025-03-13 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-22 04:00 | GEI | ENSEAL X1 CURVED 25CM SHAFT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K172580 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A THIN VESSEL WAS SEALED WITH COVE AND AFTER 24H MASSIVE BLEEDING OCCURRED. WHEN OPENING THE PATIENT, IT WAS IDENTIFIED THAT THE SEALED VESSEL WAS OPEN. THE PATIENT DIED. Manufacturer Narrative: (B)(4). DATE SENT: 4/22/2025. D4 BATCH #: UNKNOWN. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WERE THE INDICATIONS FOR SURGERY? LUNG NODULES WERE THE HEMOSTASIS ISSUES IDENTIFIED POST OPERATIVELY OR INTRA OPERATIVELY? POST-OPERATIVE. WERE ANY SURGICAL CLIPS USED IN THE AREA OF THE DEVICE FIRING? AFTER OPENING THE SHEARS, IT WAS NECESSARY TO PLACE A CLIP. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? BLEEDING AFTER OPENING THE SEAL IS ESTIMATED AT 2 LITERS. WAS A TRANSFUSION REQUIRED? YES. WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? NO INFORMATION. WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? NO INFORMATION WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? COAGULOPATHY PLEASE PROVIDE A TIMELINE OF EVENTS. AFTER THE CLAMPING OF THE VESSELS AND ARTERY, AFTER 1 HOUR OF THE PROCEDURE THE ARTERY CLAMPING OPENED. TWO CLIPS WERE PLACED IN THE SITE AND THERE WAS HEMOSTASIS. IN THE CASE OF THE SEAL OF THE ENSEAL, AFTER 24 HOURS BLEEDING AND OPENING OF THE SEAL WERE IDENTIFIED. IS THERE AN AUTOPSY REPORT AVAILABLE? NO. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 5/2/2025. ADDITIONAL INFORMATION WAS OBTAINED: THE ISSUE WAS WITH THE ENSEAL THAT WAS THE DEVICE WITH THE ONE HE DID HEMOSTASIS ONCE THE STAPLER DIDN¿T WORK, THE PATIENT BLED 24 HOURS POST-OP IN THE PLACE HE DID THE HEMOSTASIS WITH THE ENSEAL. THE COAGULATION TEST OF THE PATIENT WERE NORMAL BUT ONE RELATIVE ACCEPTED THAT THE PATIENT HAVE HAD A COAGULOPATHY ONCE IN HER LIFE, WHICH WAS NO NOTED IN HER MEDICAL HISTORY. Manufacturer Narrative: (B)(4). DATE SENT: 5/22/2025. ADDITIONAL INFORMATION WAS OBTAINED: AN EXTERNAL RAPID RESPONSE CALL WAS HELD ON (B)(6) 2025 TO DISCUSS THE HEMOSTASIS ISSUE AND THE PATIENT¿S UNFORTUNATE PASSING. THE SURGEON PROVIDED A SUMMARY OF THE EVENTS. WHEN STAPLING AND CUTTING THE BRONCHI DURING THE RIGHT LOBECTOMY, IT STARTED TO BLEED. HE HAD TO OPEN THE SURGERY TO PUT SOME CLIPS IN TO STOP THE BLEEDING. AFTER COMPLETING THE PROCEDURE, THE PATIENT WENT TO THE ICU. THE INTUBATION TUBE WAS REMOVED THE NEXT MORNING. IN THE AFTERNOON, THE SURGEON TOLD THE PATIENT ABOUT THE PROCEDURE. AT ABOUT 8PM THE PATIENT HAD COUGH AND SEIZERS AND BLED 1.5 LITERS OF BLOOD. HE TOOK THE PATIENT BACK INTO SURGERY. HE SAW THAT THE ARTERY WHERE THE ENSEAL WAS USED IN THE ORIGINAL PROCEDURE (NOT THE STAPLER) WAS BLEEDING. IT WAS A DIFFERENT VESSEL THAN WAS BLEEDING IN THE ORIGINAL PROCEDURE. THE SURGEON ATTEMPTED TO SUTURE THE BLEED, BUT THE PATIENT ULTIMATELY PASSED. WHEN TALKING WITH A RELATIVE, THE RELATIVE DID INDICATE THAT THE PATIENT HAD COAGULOPATHY ISSUES IN THE PAST, BUT THIS WAS NOT KNOWN DURING EITHER PROCEDURE. THE SURGEON INDICATED THAT HE WAS TAKING SMALL BRANCHES WITH THE ENSEAL. THE SEAL JUST FAILED; IT WAS CLOSED/SEAL WHEN THE ORIGINAL PROCEDURE WAS COMPLETED. HE CONFIRMED THE POST-OP BLEEDING CAME FROM EXACTLY WHERE HE USED THE ENSEAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882724&pc=LFL | 3005075853-2025-02848 | 2025-03-31 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-22 04:00 | LFL | HARMONIC 1100 SHEARS 36CM | Patient Device Interaction Problem | Failure to Anastomose; Burn(s); Cardiac Arrest | K200841 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT SURGERY FOR DEEP INFILTRATIVE ENDOMETRIOSIS WITH DISCOID RESECTION OF THE COLON. DURING THE SURGERY, A HARMONIC AND A CIRCULAR STAPLER WERE USED. LATER, THE PATIENT WAS SCHEDULED FOR EMERGENCY SURGERY DUE TO AN UNDETECTED LEAK IN THE SIGMOID COLON. THEY SUFFERED CARDIAC ARREST FOUR TIMES DURING THE PROCEDURE, BUT STILL MANAGED TO SURVIVE. THEY DIED POSTOPERATIVELY. THE CAUSE OF THE LEAKAGE WAS ATTRIBUTED TO A THERMAL LESION IN THE SIGMOID COLON THAT OCCURRED DURING THE MANIPULATION AND RELEASE OF STRUCTURES DURING THE FIRST SURGERY. Manufacturer Narrative: (B)(4). DATE SENT: 7/9/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SALES REP WAS PRESENT IN THE FIRST SURGERY THAT WAS PROGRAMED AND AS PART OF AN EVENT. THIS SURGERY WAS COMPLETED SUCCESSFULLY WITHOUT ANY ISSUES OR COMPLICATIONS. THERE WERE NO GENERATOR ALARMS. THE LEAK HAPPENED 2-3 DAYS AFTER THE SURGERY. THE SURGEON FOUND THE BURN THAT HAD NOT BEEN SEEN DURING THE FIRST PROCEDURE AND HE THINKS THAT IT COULD HAVE BEEN CAUSED BY THE HARMONIC. SHE SAID THAT SURGEONS HAVE RECEIVED TRAINING IN HARMONIC AND WERE AWARE THAT THE DEVICE HEATS UP IN THE DISTAL PART DURING ACTIVATION. Manufacturer Narrative: (B)(4). DATE SENT: 8/6/2025. ADDITIONAL INFORMATION WAS OBTAINED VIA THE GENERATOR LOG: THERE WERE NO SIGNIFICANT THERMAL EVENTS ASSOCIATED WITH A DEVICE MATCHING THE DEVICE BATCH NUMBER PROVIDED. THERE WAS ONE HAR1136 DEVICE MATCHING THE BATCH NUMBER REFERENCED. THE DEVICE WAS FROM BATCH Y7055D AND SERIAL # (B)(6). IT WAS CONNECTED TO THE GENERATOR A TOTAL OF 3 TIMES DUE TO SOME HANDSWITCH FAULTS THAT WERE OBSERVED. A TOTAL OF 104 ACTIVATIONS WERE COMPLETED ATT TONE WAS TRIGGERED SEVERAL TIMES. THERE WERE TWO ACTIVATIONS WITH ATT TONES FOR 1 SECOND. THAT WAS THE MAXIMUM TIME OBSERVED AND IS WELL WITHIN THE NORMAL REACTION TIME FOR A SURGEON. THERE WERE NO ACTIVATIONS WHERE CTM ACTIVATED. NO LONG ACTIVATIONS WERE OBSERVED. THE MAXIMUM ACTIVATION TIME OBSERVED WAS 7 SECONDS. NO HIGH ENERGY ACTIVATIONS WERE OBSERVED. Manufacturer Narrative: (B)(4). DATE SENT: 4/22/2025. D4 BATCH #: UNKNOWN. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE, NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE, A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT IS THE SURGEON¿S EXPERIENCE WITH HARMONIC DEVICE? WHAT ANATOMICAL STRUCTURES WERE THE HARMONIC USED ON DURING THE ORIGINAL PROCEDURE? WHAT HEAT MANAGEMENT TECHNIQUES WERE USED THROUGHOUT THE ORIGINAL PROCEDURE? WHAT SEGMENT OF GI TRACT WAS RESECTED DURING THE ORIGINAL PROCEDURE? HOW WAS THE DEVICE CLEANED THROUGHOUT THE PROCEDURE? WHAT POWER LEVELS WERE BEING USED DURING THE ORIGINAL PROCEDURE? WERE THERE ANY GENERATOR ALERT SCREENS DURING THE ORIGINAL PROCEDURE? HOW LONG AFTER THE INITIAL PROCEDURE WAS POST-OP LEAK IDENTIFIED? IS THE SURGEON AWARE OF THE FOLLOWING STATEMENT FROM OUR INSTRUCTION FOR USE: ¿DURING AND FOLLOWING ACTIVATION IN TISSUE, THE INSTRUMENT BLADE, CLAMP ARM, AND DISTAL 7 CM OF THE SHAFT MAY BE HOT. AVOID UNINTENDED CONTACT WITH TISSUE?¿ CAN A TIMELINE OF EVENTS BE PROVIDED? WHAT WAS THE CAUSE OF THE PATIENT¿S DEATH? AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21882745&pc=BZD | 2518422-2025-104603 | 2024-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT PASSED AWAY ONE YEAR AGO. A REPLACEMENT FOR THE REMSTAR PRO C-FLEX+ DEVICE WAS DECLINED, AND THE PATIENT'S SISTER HAS REQUESTED TO BE REMOVED FROM THE LIST. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883268&pc=NPT | 2025587-2025-02908 | 2024-03-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, CARDIOPULMONARY ARREST OCCURRED DUE TO AN UNSPECIFIED OUT-OF-HOSPITAL CONDITION. SUBSEQUENTLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: CORRECTED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883322&pc=NPT | 2025587-2025-02909 | 2024-07-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Renal Failure; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH PRE-EXISTING LOW CARDIAC FUNCTION AND A HISTORY OF OCCLUSION MYOCARDIAL INFARCTION, THE PATIENT WAS HOSPITALIZED WITH CONGESTIVE HEART FAILURE. CHRONIC KIDNEY DISEASE WITH A CREATININE OF 2-31 MG/DL WAS ALSO NOTED. DURING HOSPITALIZATION, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOTED TO BE HEART-RELATED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883699&pc=DSQ | 2916596-2025-02371 | 2025-03-21 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883733&pc=DSQ | 2916596-2025-02413 | 2025-04-02 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM AN UNKNOWN CAUSE. THIS WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE CAUSE AND DETAILS OF THE DEATH WERE NOT AVAILABLE. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883768&pc=DSQ | 2916596-2025-02449 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THE PATIENT REMAINED ON RIGHT VENTRICULAR ASSIST DEVICE (RVAD) SUPPORT FROM THEIR HEARTMATE 3 IMPLANT ON (B)(6) 2025. THE PATIENT DEVELOPED A VERY RESISTANT PSEUDOMONAS BACTEREMIA AND WAS NOT PROGRESSING OFF THE RVAD. THE FAMILY DECIDED TO REMOVE CARE AND STOP BOTH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND RVAD. Manufacturer Narrative: SECTION A - PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883799&pc=NPT | 2025587-2025-02911 | 2024-10-19 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Aneurysm | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO THE RUPTURE OF AN ABDOMINAL AORTIC ANEURYSM. Manufacturer Narrative: CORRECTION: THIS MEDWATCH WAS IDENTIFIED AS A DUPLICATE TO MEDWATCH #. PLEASE REFERENCE MEDWATCH # FOR THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {{D-EVOULTFX2329}}; PRODUCT LOT/SERIAL NUMBER {{UNKNOWN}}; PRODUCT TYPE: {{0195 HEART VALVES}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}}. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883804&pc=FOS | 2245270-2025-00035 | 2025-03-29 04:00 | Death | VYGON | 2025-04-22 04:00 | FOS | UMBILICAL CATHETER | Malposition of Device; Positioning Problem | Cardiac Tamponade; Pericardial Effusion | K921352 |  | 1 | Event Description: A PREMATURE NEWBORN 34 WEEKS + 2 DAYS, WITH A BIRTH WEIGHT OF 1480G, WAS TRANSFERRED TO NEONATAL INTENSIVE CARE AT A SECOND HOSPITAL ON 1 DAY OF LIFE FOR SURGICAL MANAGEMENT OF AN ANORECTAL MALFORMATION WITH RECTOVAGINAL FISTULA. ON ADMISSION, HER CLINICAL CONDITION WAS STABLE AND PARENTERAL NUTRITION IS ADMINISTERED THROUGH UMBILICAL VENOUS CATHETER INSERTED AT BIRTH BY THE FIRST HOSPITAL. THE DAY AFTER ADMISSION, THE PATIENT IS OPERATED ON TO DILATE THE FISTULA. THE OPERATION IS GOING WELL AND ALLOWS EVACUATION OF FAECES. ON DAY 2 OF ADMISSION, SHE SUFFERED A CARDIORESPIRATORY ARREST DUE TO TAMPONADE OF IATROGENIC ORIGIN (PERFUSO- PERICARDIUM RELATED TO MALPOSITION OF THE UMBILICAL VENOUS CATHETER). THE EVOLUTION IS UNFAVORABLE DESPITE CARDIOPULMONARY RESUSCITATION STARTED EARLY AND PERICARDIAL DRAINAGE. THE PREMATURE DIED.Event Description: A PREMATURE NEWBORN 34 WEEKS + 2 DAYS, WITH A BIRTH WEIGHT OF 1480G, WAS TRANSFERRED TO NEONATAL INTENSIVE CARE AT A SECOND HOSPITAL ON 1 DAY OF LIFE FOR SURGICAL MANAGEMENT OF AN ANORECTAL MALFORMATION WITH RECTOVAGINAL FISTULA. ON ADMISSION , HER CLINICAL CONDITION WAS STABLE AND PARENTERAL NUTRITION IS ADMINISTERED THROUGH UMBILICAL VENOUS CATHETER INSERTED AT BIRTH BY THE FIRST HOSPITAL. THE DAY AFTER ADMISSION, THE PATIENT IS OPERATED ON TO DILATE THE FISTULA. THE OPERATION IS GOING WELL AND ALLOWS EVACUATION OF FAECES . ON DAY 2 OF ADMISSION , SHE SUFFERED A CARDIORESPIRATORY ARREST DUE TO TAMPONADE OF IATROGENIC ORIGIN (PERFUSO- PERICARDIUM RELATED TO MALPOSITION OF THE UMBILICAL VENOUS CATHETER). THE EVOLUTION IS UNFAVORABLE DESPITE CARDIOPULMONARY RESUSCITATION STARTED EARLY AND PERICARDIAL DRAINAGE. THE PREMATURE DIED. Manufacturer Narrative: THIS MDR WAS INITIALLY SENT UNDER NUMBER 2245270-2025-00035 ON APRIL 22, 2025. THIS REPORT REQUIRE A CORRECTION. CORRECTION: THE PRODUCT LOT NUMBER ASSOCIATED WITH THE COMPLAINT SHOULD BE 290523PB . THE PREVIOUS REPORT LISTED 290325PB ERRONEOUSLY. THE INVOLVED UMBILICAL CATHETER IS NOT AVAILABLE. FROM THE DESCRIPTION, THIS SERIOUS ADVERSE EVENT IS CLEARLY RELATED TO THE MALPOSITION OF THE UMBILICAL CATHETER. IT IS NOT TRACED TO A DEFECT OF OUR DEVICE. THE BATCH REVIEW DOES NOT SHOW ANY NON-CONFORMITY . THE PRODUCTS ARE IN CONFORMITY WITH THE SPECIFICATIONS. ALL CONTROLS ARE COMPLIANT, INCLUDING THE 100% TESTS CARRIED OUT ON EACH CATHETER DURING THE MANUFACTURING PROCESS. THESE TESTS ARE AS FOLLOWS: THE BLUNTNESS OF THE CATHETER TIP , MARKING THE CATHETERS, TIGHTNESS AND NON-OBSTRUCTION. THE HISTORICAL DATA ANALYSIS OF COMPLAINTS THESE LAST 3 YEARS, ON THIS REFERENCE DOES NOT SHOW ANY SIMILAR INCIDENT. AT THIS TIME, NO CORRECTIVE ACTION HAS BEEN INITIATED, AS THERE IS NO EVIDENCE INDICATING A CONNECTION BETWEEN THE ISSUE AND THE DEVICE. Manufacturer Narrative: THE INVOLVED UMBILICAL CATHETER IS NOT AVAILABLE. FROM THE DESCRIPTION, THIS SERIOUS ADVERSE EVENT IS CLEARLY RELATED TO THE MALPOSITION OF THE UMBILICAL CATHETER. IT IS NOT TRACED TO A DEFECT OF OUR DEVICE. THE BATCH REVIEW DOES NOT SHOW ANY NON-CONFORMITY. THE PRODUCTS ARE IN CONFORMITY WITH THE SPECIFICATIONS. ALL CONTROLS ARE COMPLIANT, INCLUDING THE 100% TESTS CARRIED OUT ON EACH CATHETER DURING THE MANUFACTURING PROCESS. THESE TESTS ARE AS FOLLOWS: THE BLUNTNESS OF THE CATHETER TIP, MARKING THE CATHETERS, TIGHTNESS AND NON-OBSTRUCTION. THE HISTORICAL DATA ANALYSIS OF COMPLAINTS THESE LAST 3 YEARS, ON THIS REFERENCE DOES NOT SHOW ANY SIMILAR INCIDENT. AT THIS TIME, NO CORRECTIVE ACTION HAS BEEN INITIATED, AS THERE IS NO EVIDENCE INDICATING A CONNECTION BETWEEN THE ISSUE AND THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883810&pc=MIH | 3007284313-2025-03919 | 2024-11-18 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-22 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Obstruction of Flow | Thrombosis/Thrombus | P020004 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2024, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE UTILIZING THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) AND GORE® EXCLUDER® AAA ENDOPROSTHESIS (RLT AND 2 PLCS) IN ZONE 5. PRIOR TO THE PROCEDURE, THE PHYSICIAN ADDRESSED THE PATIENT'S LEFT RENAL ARTERY THAT WAS ALREADY OCCLUDED. A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE/STENT) WAS DEPLOYED INTO THE RENAL PORTAL THAT WAS FOLLOWED BY AN AMPLATZER PLUG. DURING THE PROCEDURE, THERE WAS NOT AN ADEQUATE LANDING ZONE IN THE PATIENT¿S CELIAC ARTERY. THERE WAS A SHORT BRANCH FROM THE LEFT GASTRIC ARTERY THAT BRANCHED OFF THE SPLENIC ARTERY. TO ACHIEVE A SEAL INTO THE CELIAC ARTERY, THE PHYSICIAN PLUGGED THE SPLENIC ARTERY AND COILED THE LEFT GASTRIC ARTERY. A VBX DEVICE WAS THEN IMPLANTED INTO THE PATIENT¿S HEPATIC ARTERY. ADDITIONAL VBX DEVICES WERE IMPLANTED IN THE CELIAC ARTERY AND SUPERIOR MESENTERIC ARTERY WITH NO ISSUES. ON (B)(6) I-CAST STENT WAS IMPLANTED IN THE RIGHT RENAL ARTERY (RRA) DUE TO EXISTING HIGH-GRADE STENOSIS. DURING THE PROCEDURE, TWO VBX DEVICES WERE DEPLOYED WITHIN THE RRA. AS REPORTED, PROCEDURE WENT SMOOTHLY AND THE PATIENT TOLERATED THE PROCEDURE. AFTER PROCEDURE COMPLETION, A NON-CONTRAST CT IMAGING SHOWED ALL DEVICES WERE WIDELY PATENT WITH NO COMPRESSION OR STENOSIS NOTED. ON (B)(6) 2024, THE NURSING STAFF CHANGED THE PATIENT'S BEDDING AND THE PATIENT MOVED ON HER OWN WITH NO ISSUES. HOWEVER, WHEN THE PATIENT WAS MOVED BACK INTO BED, HER BLOOD PRESSURE DROPPED INTO THE 50¿S WHEN LYING FLAT IN BED. THE PATIENT CODED AND ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS PERFORMED. AFTER RESUSCITATION, THE PATIENT WAS INTUBATED, AND IT WAS ALSO NOTED THAT SHE HAD ASPIRATED. THE PATIENT WAS TRANSFERRED FOR CTA SCAN OF THE CHEST. ABDOMEN AND PELVIS SHOWED IRREGULAR RESPONSE OF AORTA SECONDARY TO CARDIAC DYSFUNCTION. THE PATIENT HAD RIB FRACTURES NOTED ON THE CT AFTER ACLS. THE VBX STENTS IN THE RIGHT RENAL ARTERY WERE ALSO CRUSHED. AS REPORTED, THE PHYSICIAN BELIEVES THAT DUE TO CARDIAC EVENT AND EXTENDED LOW / MINIMAL BLOOD PRESSURE, THE PATIENT'S VISCERAL ARTERIES (CELIAC, HEPATIC, SMA, AND RIGHT RENAL ARTERY) ALL PRESENTED AS THROMBOSED/OCCLUDED, ALONG WITH THE TAMBE DEVICE, VBX DEVICE, AND EXCLUDER DEVICES. THE PATIENT REMAINED INTUBATED POST CT SCAN AND EXPIRED SOMETIME IN THE MORNING. THE RIGHT RENAL ARTERY VBX STENTS WERE LIKELY CRUSHED WHILE TRYING TO RESUSCITATE THE PATIENT. THE PRIMARY CAUSE OF THE PATIENT'S DEATH IS UNKNOWN. IT IS ALSO UNKNOWN IF THE PATIENT HAD SENSITIVITY TO HEPARIN AS TESTING WAS NOT DONE PRIOR TO PROCEDURE. Manufacturer Narrative: ADDED G3/G4, H1/H2, H6. Manufacturer Narrative: AS IT IS UNKNOWN WHICH DEVICE(S) IS DIRECTLY IMPLICATED IN THIS DEATH EVENT, THE FOLLOWING DEVICE(S) WILL ALSO BE INCLUDED IN THIS REPORT: CATALOG #PLC141200/ SERIAL # (B)(6)/ UDI #(B)(4). CATALOG #PLC141000/ SERIAL # (B)(6)/ UDI #(B)(4). ACCORDING TO THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE DEATH, ARTERIAL OR VENOUS THROMBOSIS, OCCLUSION OF DEVICE OR NATIVE VESSEL, SINGLE OR MULTIPLE VESSELS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883897&pc=DSQ | 2916596-2025-02450 | 2025-03-30 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Loss of consciousness; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED DUE TO A DECREASED LEVEL OF CONSCIOUSNESS. TESTS PERFORMED REVEALED THAT THE PATIENT HAD EXPERIENCED 2 LARGE EMBOLIC STROKES. CONSEQUENTLY, THE PATIENT FAMILY DECIDED TO REMOVE CARE AND STOP LEFT VENTRICULAR ASSIST DEVICE (LVAD) SUPPORT. THE PATIENT PASSED AWAY ON (B)(6) 2025. ACCORDING TO THE SITE THE LVAD HAD OPERATED AS INTENDED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21883940&pc=LWR | 2135147-2025-02119 | 2015-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-22 04:00 | LWR | TRIFECTA¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100029 |  | 1 | Event Description: THE ARTICLE, "THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY THAT COMPARED THE CLINICAL AND HEMODYNAMIC OUTCOME OF PATIENTS UNDERGOING VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT (VIV-TAVR) FOR DETERIORATED TRIFECTA AND EPIC AORTIC BIOPROSTHESES. DEVICES INCLUDED IN THE STUDY WERE EPIC AND TRIFECTA. THE ARTICLE CONCLUDED THAT COMPARED TO A STENTED PROSTHESIS WITHOUT INCREASED RISK OF CORONARY OBSTRUCTION (CO), VIV-TAVR INTO TRIFECTA PROSTHESES CAN BE PERFORMED WITH LOW RISK OF CO AND ACCEPTABLE SHORT-TERM CLINICAL OUTCOMES. AS THE RATE OF POST-VIV PPM IS SUBSTANTIAL FOR BOTH PROSTHESES, CAREFUL PATIENT SELECTION IS WARRANTED. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MATTHIAS RASCHPICHLER, LEIPZIG HEART CENTER, UNIVERSITY CLINIC OF CARDIAC SURGERY, LEIPZIG, GERMANY, WITH CORRESPONDING EMAIL: MRASCHPICHLER@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM OCTOBER 2015 TO JUNE 2020. A TOTAL OF 76 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 80 YEARS OLD AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED PRIOR STROKE, MYOCARDIAL INFARCTION, ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC RENAL FAILURE, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, PERMANENT PACEMAKER, AORTIC STENOSIS/REGURGITATION.Event Description: N/A Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRIFECTA VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR STROKE, MYOCARDIAL INFARCTION, ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC RENAL FAILURE, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, PERMANENT PACEMAKER, AORTIC STENOSIS/REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (VALVE-IN-VALVE), HOSPITALIZATION, STROKE, AORTIC STENOSIS, AORTIC REGURGITATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRIFECTA VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR STROKE, MYOCARDIAL INFARCTION, ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC RENAL FAILURE, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, PERMANENT PACEMAKER, AORTIC STENOSIS/REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (VALVE-IN-VALVE), HOSPITALIZATION, STROKE, AORTIC STENOSIS, AORTIC REGURGITATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21884191&pc=DSQ | 2916596-2025-02329 | 2025-03-27 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT LOG FILES WERE SUBMITTED FOR REVIEW. THE EVENT LOG DISPLAYED MULTIPLE CONSECUTIVE STRINGS OF LOW FLOW ALARMS WITH A RECORDED FLOW RANGING 0.0 LPM - 2.4 LPM ON (B)(6) 2025 AT 5:35 PM UNTIL 6:17 PM. THE PUMP SET SPEED (4000 RPM) WAS SET THE SAME AS THE LOW SET LIMIT (4000 RPM) PER THE LAST RECORDED EVENT IN THE EVENT LOG FILE. IT DID APPEAR THAT THE FLOW IMPROVED TOWARDS THE END OF THE LOG FILE HISTORY. THERE WERE NO OTHER UNUSUAL EVENTS SEEN WITHIN THE LOG FILES. IT WAS LATER REPORTED THAT ON (B)(6) 2025 THE PATIENT CODED ON THE FLOOR AND WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) THERAPY, IN WHICH THEY LATER PASSED AWAY THAT DAY. IT WAS REPORTED THAT RIGHT HEART VENTRICULAR FAILURE WAS THE CAUSE OF DEATH. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS; HOWEVER, A SPECIFIC CAUSE FOR THESE FINDINGS COULD NOT BE CONCLUSIVELY DETERMINED. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS ON (B)(6) 2025 FROM 17:35:08 ¿ 18:39:06. A SUSTAINED LOW FLOW ALARM WAS CAPTURED FROM THE BEGINNING OF THE FILE THROUGH 18:17:09, WITH FLOW DECREASING TO 0.0 LITERS PER MINUTE (LPM) BY 18:05:48. OF NOTE, THIS DECREASE IN FLOW APPEARED TO BE ASSOCIATED WITH A DECREASE IN THE STORED PATIENT SPEED. FLOW THEN REMAINED AT 0.0 LPM UNTIL 18:07:12, WHEN THE FLOW BEGAN TO GRADUALLY INCREASE WITH THE LOW FLOW ALARM RESOLVING BY 18:17:13 AND FLOWS RETURNING TO BASELINE BY THE END OF THE FILE. OF NOTE, THIS INCREASE IN FLOW APPEARED TO BE ASSOCIATED WITH AN INCREASE IN THE STORED PATIENT SPEED. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED AT THE STORED PATIENT SPEED. WHETHER THE PATIENT'S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. THIS SECTION ALSO NOTES THAT, IN GENERAL, THE MAGNITUDE OF THE PULSATILITY INDEX (PI) VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E., THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21884684&pc=OZD | 1220648-2025-27738 | 2025-03-29 04:00 | Death | ABIOMED, INC. | 2025-04-22 04:00 | OZD | IMPELLA CP | Failure to Pump | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND AFTER PLACEMENT, THE PHYSICIAN ATTEMPTED TO START THE PUMP ON AUTO. HOWEVER, THERE WAS AN IMMEDIATE MOTOR CURRENT HIGH ALARM NOTED, AND THE PUMP FAILED TO START. THE IMPELLA CP DEVICE WAS REMOVED AND REINSERTED. THE SECOND ATTEMPT TO START THE PUMP MET WITH THE SAME MOTOR CURRENT HIGH ALARM AND THE PUMP FAILED TO START. THE DECISION WAS MADE TO PLACE A NEW IMPELLA CP DEVICE WHICH WAS SUCCESSFULLY COMPLETED VIA THE RIGHT FEMORAL ARTERY. HOWEVER, DURING THE PERCUTANEOUS INTERVENTION TO THE 100% OCCLUDED RIGHT CORONARY ARTERY, THE PATIENT CODED. THE TEAM WAS UNABLE TO OBTAIN RETURN OF SPONTANEOUS CIRCULATION, AND THE PATIENT EXPIRED. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. DELAY IN NEEDED SUPPORT; MAY HAVE ADD TO AN ALREADY COMPLEX SITUATION WITH THE PATIENT. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21884788&pc=DSQ | 2916596-2025-02474 | 2025-04-04 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT HOME, ON HOSPICE CARE, DUE TO AN UNKNOWN CAUSE ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO HAVE BEEN DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21884981&pc=LWR | 2135147-2025-02126 | 2015-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-22 04:00 | LWR | EPIC¿ VALVE (AORTIC) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P040021 |  | 1 | Event Description: THE ARTICLE, "THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY THAT COMPARED THE CLINICAL AND HEMODYNAMIC OUTCOME OF PATIENTS UNDERGOING VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT (VIV-TAVR) FOR DETERIORATED TRIFECTA AND EPIC AORTIC BIOPROSTHESES. DEVICES INCLUDED IN THE STUDY WERE EPIC AND TRIFECTA. THE ARTICLE CONCLUDED THAT COMPARED TO A STENTED PROSTHESIS WITHOUT INCREASED RISK OF CORONARY OBSTRUCTION (CO), VIV-TAVR INTO TRIFECTA PROSTHESES CAN BE PERFORMED WITH LOW RISK OF CO AND ACCEPTABLE SHORT-TERM CLINICAL OUTCOMES. AS THE RATE OF POST-VIV PPM IS SUBSTANTIAL FOR BOTH PROSTHESES, CAREFUL PATIENT SELECTION IS WARRANTED. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MATTHIAS RASCHPICHLER, LEIPZIG HEART CENTER, UNIVERSITY CLINIC OF CARDIAC SURGERY, LEIPZIG, GERMANY, WITH CORRESPONDING EMAIL: MRASCHPICHLER@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM OCTOBER 2015 TO JUNE 2020. A TOTAL OF 76 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 80 YEARS OLD AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED PRIOR STROKE, MYOCARDIAL INFARCTION, ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC RENAL FAILURE, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, PERMANENT PACEMAKER, AORTIC STENOSIS/REGURGITATION.Event Description: N/A Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF EPIC VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR STROKE, MYOCARDIAL INFARCTION, ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC RENAL FAILURE, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, PERMANENT PACEMAKER, AORTIC STENOSIS/REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (VALVE-IN-VALVE), HOSPITALIZATION, STROKE, AORTIC STENOSIS, AORTIC REGURGITATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: THE INTERNATIONAL TRIFECTATM AND EPICTM VALVE-IN-VALVE REGISTRY: INSIGHTS INTO CLINICAL & HEMODYNAMIC OUTCOMES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885006&pc=MVK | 3008642652-2025-04110 | 2025-03-24 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 10:15:36 ON (B)(6) 2025. AT 07:15:25 ON(B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT. AT 07:16:00, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 07:16:17, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 07:16:31, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY. AT 07:30:29, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM WITH MOTION ARTIFACT. AT 07:30:29, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 180 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH BBB AND MOTION ARTIFACT. AT 07:41:47, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH UNCONDUCTED P WAVES TRANSITIONING TO VF. AT 07:42:18, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH UNCONDUCTED P WAVES. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 08:10:00 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885013&pc=DSQ | 2916596-2025-02691 | 2025-04-07 04:00 | Death | THORATEC CORPORATION | 2025-04-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE BIVENTRICULAR ASSIST DEVICE (BIVAD) PATIENT NEVER RECOVERED AFTER IMPLANTATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE CUSTOMER REPORTED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885024&pc=BZD | 2518422-2025-044232 | 2021-08-26 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | BZD | PHILIPS CPAP | Degraded | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885279&pc=CBK | 2518422-2025-104626 | 2025-03-24 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM A DEALER ON 26 MARCH 2025 THAT THEY ARE REQUESTING ANALYSIS OF A TRILOGY EVO VENTILATOR FOR DETAILS OF RECORDED DATA BETWEEN 14:00 AND 15:00 ON (B)(6) 2025 ASSOCIATED WITH "AN INCIDENT". THEY REQUESTED TO KNOW WHAT ALARM(S) OCCURRED DURING THIS RELEVANT TIME AND THE STATUS OF ANY EVENTS. ON 27 MARCH 2025, INFORMATION WAS PROVIDED BY THE SALES REPRESENTATIVE THAT THE PATIENT HAD DIED. NO PATIENT INFORMATION WAS PROVIDED. IN REGARD TO THE PATIENT DEATH, THE SALES REPRESENTATIVE STATED THERE WAS NO CAUSAL RELATIONSHIP WITH THE MEDICAL DEVICE (VENTILATOR), AND NO INFORMATION ABOUT ANY MALFUNCTION OF THE DEVICE. THE CAUSE OF THE PATIENT'S DEATH WAS REPORTED TO HAVE BEEN DETERMINED TO BE THE DELAYED RESPONSE BY THE HEALTHCARE PROVIDERS, AND THE HOSPITAL IS CURRENTLY CONSIDERING MEASURES TO ADDRESS THE HEALTHCARE PROVIDER'S RESPONSE. AT THIS TIME, THERE IS NO ALLEGATION OF A DEVICE DEFECT, MALFUNCTION OR FAILURE. IT HAS BEEN INDICATED THAT THE VENTILATOR IS AVAILABLE FOR RETURN TO THE MANUFACTURER FOR A DEVICE EVALUATION; HOWEVER, THE VENTILATOR HAS NOT YET BEEN RECEIVED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE TYPE OF REPORTABLE EVENT FIELD HAS BEEN UPDATED FROM "PRODUCT PROBLEM" TO APPROPRIATELY REFLECT "DEATH" AS DESCRIBED IN THE REPORT NARRATIVE. Manufacturer Narrative: THE SUBJECT DEVICE WAS RETURNED TO THE PHILIPS SERVICE CENTER FOR EVALUATION. THERE WERE 2 ERROR CODES INVOLVING USER SETTINGS NOTED IN THE DEVICE DOWNLOAD LOG THAT ARE OF NO CONSEQUENCE. THE DEVICE WAS EVALUATED AND TESTED WITH NO FAILURE DURING THE INITIAL EVALUATION IN THE REPAIR CENTER. DEVICE EVALUATION WAS UNABLE TO CONFIRM OR DUPLICATE A FAILURE OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885494&pc=MGB | 2024168-2025-04193 | 2025-03-06 05:00 | Death | ABBOTT VASCULAR INC. | 2025-04-22 04:00 | MGB | PERCLOSE¿ PROSTYLE¿ | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cardiac Arrest; Hemorrhage/Bleeding; Shock; Unspecified Tissue Injury; No Clinical Signs, Symptoms or Conditions | P960043 |  | 1 | Event Description: IT WAS REPORTED THAT THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) WAS PERFORMED USING TWO NON-ABBOTT THORACIC STENT GRAFT SYSTEMS FOR A DESCENDING THORACIC AORTIC ANEURYSM. THE FIRST STENT WAS ATTEMPTED TO BE ELEVATED FROM THE RIGHT FA (USING PERCLOSE), BUT IT COULD NOT BE ADVANCED AROUND THE EIA-CIA AREA. IT WAS REMOVED, AND A 24FR SHEATH WAS ELEVATED FROM THE SAME RIGHT FA. THE 24FR SHEATH ADVANCED TO THE RENAL ARTERY LEVEL. THE SECOND STENT WAS INSERTED INTO THE 24FR SHEATH AND PLACED IN THE DESCENDING AORTA AS PLANNED. THE SECOND STENT WAS LANDED IN THE DISTAL ARCH AND ANGIOGRAPHY WAS PERFORMED. DURING THIS ANGIOGRAPHY, THE FLOW OF THE CONTRAST AGENT WAS POOR AND IT STAGNATED IN THE ARCH FOR A LONG TIME. THE SECOND STENT WAS PLACED AS PLANNED, AND ANGIOGRAPHY WAS PERFORMED. THE TEVAR ITSELF WAS COMPLETED. UPON CHECKING THE ACCESS, CONTRAST AGENT WAS FOUND TO BE LEAKING OUTSIDE THE VESSEL AROUND THE RIGHT EIA-CIA AREA, CONFIRMING ACCESS DAMAGE. THE PATIENT WENT INTO SHOCK AND CARDIAC ARREST, AND CPR WAS INITIATED. AN URGENT DECISION WAS MADE TO PLACE A NON-ABBOTT STENT GRAFT. THE FIRST NON-ABBOTT STENT GRAFT (F10MM X 10CM) WAS PLACED FROM THE CIA TO THE EIA. POST-PLACEMENT ANGIOGRAPHY SHOWED REMAINING DAMAGE, SO A SECOND NON-ABBOTT STENT GRAFT (F9MM X 10CM) WAS ADDED SLIGHTLY PROXIMAL TO THE FEMORAL HEAD LEVEL. (THE NATIVE VESSEL DIAMETER WAS APPROXIMATELY 8MM.) A TYPE 3A ENDOLEAK WAS OBSERVED BETWEEN THE NON-ABBOTT STENT GRAFTS, SO A NON-ABBOTT BALLOON DILATATION CATHETER (F10MM) WAS USED FOR TOUCH-UP. THE TYPE 3A ENDOLEAK DID NOT DISAPPEAR, AND WHEN AN 8FR SHEATH WAS ADVANCED TO PLACE A THIRD NON-ABBOTT STENT GRAFT (F11MM), THE JUNCTION OF THE PREVIOUSLY PLACED NON-ABBOTT STENT GRAFT DETACHED. THE NON-ABBOTT STENT GRAFT PLACED PROXIMALLY WAS COMPLETELY PUSHED OUT OF THE VESSEL, AND THE GUIDEWIRE THAT HAD BEEN PASSING THROUGH THE TRUE LUMEN WAS ALSO DISLODGED, MAKING CANNULATION IMPOSSIBLE. A RESCUE BALLOON WAS ELEVATED FROM THE CONTRALATERAL (LEFT) SIDE FOR OCCLUSION. WHILE CONSIDERING THE NEXT STEPS, THE PATIENT REMAINED IN CARDIAC ARREST, AND CPR COULD NOT BE STOPPED. BLEEDING WAS ALSO OBSERVED FROM THE INTUBATION TUBE. AFTER 2 HOURS OF CARDIAC ARREST WITH NO RETURN OF SPONTANEOUS CIRCULATION, THE PHYSICIAN DECIDED TO CEASE RESUSCITATION EFFORTS, AND DEATH WAS CONFIRMED.Event Description: SUBSEQUENT TO THE INITIALLY FILED REPORT, THE FOLLOWING INFORMATION WAS RECEIVED: REPORTEDLY, THE SUTURES OF TWO PROSTYLE DEVICES WERE SUCCESSFULLY PRE-PLACED IN THE RIGHT COMMON FEMORAL ARTERY USING THE PRE-CLOSE TECHNIQUE VIA A 6F SHEATH HOLE PRIOR TO THE THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) PROCEDURE. THE SHEATH WAS UPSIZED TO 24F. BOTH SUTURE KNOTS WERE SUCCESSFULLY ADVANCED AND ACHIEVED HEMOSTASIS. IN THE PHYSICIAN'S OPINION, THE PROSTYLE DEVICES DID NOT CAUSE OR CONTRIBUTE TO ANY OF THE ADVERSE EVENTS OR DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. G2: POST-MARKET SAFETY MANAGEMENT AT MEDTRONIC JAPAN CO., LTD. Manufacturer Narrative: ANALYSIS WILL NOT BE REQUIRED AS ADDITIONAL INFORMATION WAS RECEIVED INDICATING THERE WAS NO DEVICE MALFUNCTION OR ADVERSE EVENT ASSOCIATED WITH THIS DEVICE. IN THE PHYSICIAN'S OPINION, THE PROSTYLE DEVICES DID NOT CAUSE OR CONTRIBUTE TO ANY OF THE ADVERSE EVENTS OR DEATH. H6- MEDICAL DEVICE PROBLEM CODE 1494 CLARIFIER: INDICATION FOR USE: REPORTEDLY, THE SHEATH WAS UPSIZED TO 24F. IT SHOULD BE NOTED THE ELECTRONIC PROSTYLE INSTRUCTIONS FOR USE (EIFU). IT SHOULD BE NOTED THAT THE ELECTRONIC PROSTYLE INSTRUCTIONS FOR USE (EIFU), STATES: FOR ACCESS SITES IN THE COMMON FEMORAL ARTERY USING 5F TO 21F SHEATHS. CORRECTION: H6- HEALTH EFFECT - CLINICAL CODE 4559, 2072, 1762,1888 WERE REMOVED. CORRECTION: H6- HEALTH EFFECT - IMPACT CODE 1802, 4641, 4604 WERE REMOVED. CORRECTION: H6- MEDICAL DEVICE PROBLEM CODE 2993 WAS REMOVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885550&pc=NPT | 2025587-2025-02920 | 2025-04-16 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT FX PLUS VALVE | Fluid/Blood Leak; Perivalvular Leak; Activation Failure; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Obstruction/Occlusion; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE OF A TRANSCATHETER BIOPROSTHETIC VALVE , THE FEMORAL ARTERIES WERE NOT ACCESSIBLE. FLUOROSCOPY REVEALED A SUCCESSFUL VALVE LOAD. ACCESS VIA THE SUBLCAVIAN ARTERY WAS ACHIEVED AND COMMISSURAL ALIGNMENT COULD NOT BE OBTAINED BUT WAS ACCEPTED. THE INITIAL IMPLANTATION DEPTH WAS GOOD, CONFIRMED FROM TWO DIRECTIONS, AND THE VALVE WAS IMPLANTED. HOWEVER, AN IMAGE SHOWED A CONSTRAINED INFLOW PART, LEADING TO POST-DILATATION WITH A 23MM BALLOON. ANGIOGRAPHY INDICATED GOOD VALVE IMPLANTATION, CORONARY PERFUSION, AND MINIMAL LEAKAGE. DURING THE CLOSURE OF THE SUBCLAVIAN ARTERY, THERE WAS A DROP IN BLOOD PRESSURE, NECESSITATING RESUSCITATION. THE WORKING DIAGNOSIS WAS CORONARY OCCLUSION. AN ATTEMPT TO PULL THE VALVE UP WITH A SNARE WAS SUCCESSFUL, BUT THERE WAS NO PERFUSION IN THE CORONARY ARTERIES, AND THE PATIENT DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE MINIMAL LEAKAGE REPORTED WAS MODERATE PARAVALVULAR LEAK. PER THE PHYSICIAN, THE CAUSE OF THE CORONARY OCCLUSION WAS UNKNOWN; HOWEVER, COULD HAVE BEEN DUE TO THE VALVE OR DUE TO THROMBUS. ADDITIONALLY, PER THE PHYSICIAN, THE VALVE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THERE WAS NO UNDERLYING PATIENT DISEASE/DISORDER/DIAGNOSIS THAT WAS THE CAUSE OF THE PATIENT OUTCOME. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A4., B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21885752&pc=LWR | 2135147-2025-02131 | 2025-03-06 05:00 | Death | ABBOTT MEDICAL | 2025-04-22 04:00 | LWR | EPIC PLUS | Nonstandard Device; Patient-Device Incompatibility | Heart Failure/Congestive Heart Failure | P040021 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 31MM EPIC VALVE WAS SELECTED FOR IMPLANT. DURING THE PROCEDURE, AFTER PARACHUTING THE VALVUE DOWN INTO THE ANNULUS, AND REMOVING THE VALVE FROM THE HOLDER, IT WAS NOTED THAT THERE WAS A BENT LEAFLET ON THE VALVE. THE FOLD WAS NOTED ON THE FREE EDGE OF THE VEIL, NEAR ITS JUNCTION WITH THE POLE. IT WAS ATTRIBUTED THAT THIS WAS DUE TO THE CALCIFICATION OF THE NATIVE VALVE. THE VALVE WAS REMOVED AND ANOTHER ABBOTT VALVE WAS USED TO COMPLETE THE PROCEDURE. THE PHYSICIAN ALLEGED THAT THERE WAS "HARDENING" ON THE ABBOTT VALVE. THE PATIENT REMAINED HEMODYNAMICALLY STABLE THROUGHOUT THE PROCEDURE. THE PHYSICIAN DOESN'T BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. ON THE SAME DAY, THE PATIENT PASSING AWAY. THE PHYSICIAN CLASSIFIES THIS DEATH AS BEING RELATED TO THE PROCEDURE AND THE PATIENT¿S COMORBIDITIES. THE IMPLANTER DOES NOT CLASSIFY THAT THE DEATH IS RELATED TO THE FUNCTION/PERFORMANCE OF THE IMPLANTED DEVICE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 31MM EPIC VALVE WAS SELECTED FOR IMPLANT. DURING THE PROCEDURE, AFTER PARACHUTING THE VALVE DOWN INTO THE ANNULUS, AND REMOVING THE VALVE FROM THE HOLDER, IT WAS NOTED THAT THERE WAS A BENT LEAFLET ON THE VALVE. THE FOLD WAS NOTED ON THE FREE EDGE OF THE VEIL, NEAR ITS JUNCTION WITH THE POLE. IT WAS ATTRIBUTED THAT THIS WAS DUE TO THE CALCIFICATION OF THE NATIVE VALVE. THE VALVE WAS REMOVED AND ANOTHER ABBOTT VALVE WAS USED TO COMPLETE THE PROCEDURE. THE PHYSICIAN ALLEGED THAT THERE WAS "HARDENING" ON THE ABBOTT VALVE. THE PATIENT REMAINED HEMODYNAMICALLY STABLE THROUGHOUT THE PROCEDURE. THE PHYSICIAN DOESN'T BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. ON THE SAME DAY, THE PATIENT PASSING AWAY. THE PHYSICIAN CLASSIFIES THIS DEATH AS BEING RELATED TO THE PROCEDURE AND THE PATIENT¿S COMORBIDITIES. THE IMPLANTER DOES NOT CLASSIFY THAT THE DEATH IS RELATED TO THE FUNCTION/PERFORMANCE OF THE IMPLANTED DEVICE. Manufacturer Narrative: THE DEVICE IS EXPECTED TO BE RETURNED FOR INVESTIGATION. IT HAS NOT YET BEEN RECEIVED. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT AFTER REMOVING THE HOLDER WHEN PARACHUTING THE EPIC MITRAL VALVE IT WAS DISCOVERED THAT THERE WAS A FOLD IN THE EPIC MITRAL VALVE LEAFLET WITH CALCIFICATION. THE VALVE WAS REPLACED WITH ANOTHER 31 MM EPIC VALVE. HOWEVER, THE PATIENT EXPIRED ON THE SAME DAY. ALSO REPORTED THAT THE PHYSICIAN DOESN'T BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. THE PHYSICIAN CLASSIFIED PATIENT DEATH AS BEING RELATED TO THE PROCEDURE AND THE PATIENT¿S COMORBIDITIES. THE DEVICE WAS RETURNED TO ABBOTT FOR INVESTIGATION AND NO ANOMALIES WERE FOUND WITH THE VALVE CUSPS. THE VALVE WAS SENT FOR PATHOLOGICAL TESTING AND THE INVESTIGATION FOUND THAT THERE WAS A POSSIBLE FOLD AND INCOMPLETE COAPTATION AT CUSP 1. THE INFLOW AND OUTFLOW SURFACES OF ALL THREE CUSPS HAVE SCATTERED NON-ADHERENT FUNGAL HYPHAE. THERE WAS NO INFLAMMATION OR SIGNIFICANT CALCIFICATIONS PRESENT. THE VALVE WAS SENT FOR FUNCTIONAL TESTING AT THE ENGINEERING TEST LAB. HYDRODYNAMIC EXAMINATION INDICATED THE VALVE MET FUNCTIONAL SPECIFICATION. A REVIEW OF FUNCTIONAL HYDRODYNAMIC TESTING OF THE RETURNED VALVE WAS PERFORMED AND SHOWED NORMAL VALVE FUNCTION WITH NORMAL SYMMETRICAL FULL OPENING OF ALL THREE CUSPS AND SYMMETRICAL COAPTATION WITH NO EVIDENCE OF A CENTRAL GAP. BASED ON THE REVIEW OF THE VIDEO DURING HYDRODYNAMIC TESTING ALL THREE LEAFLETS APPEARED NORMAL AND FUNCTIONING NORMALLY DESPITE THE APPEARANCE OF A FOLD ON PHOTOGRAPHY. DURING FUNCTIONAL TESTING THE FOLD WAS NO LONGER PRESENT. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, INCLUDING FUNCTIONAL TESTING TO ENSURE PROPER LEAFLET COAPTATION AND HEMODYNAMIC PERFORMANCE, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE MEDICAL REVIEW PERFORMED AT ABBOTT, THE EXACT CAUSE OF PATIENT DEATH CANNOT BE DETERMINED FROM THE INFORMATION PROVIDED. THE CAUSE OF THE REPORTED LEAFLET DEFORMATION AND PATIENT DEATH COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, IT WAS REPORTED THAT THE PATIENT DEATH WAS RELATED TO THE PROCEDURE AND PATIENT COMORBIDITIES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886065&pc=NPU | 2015691-2025-03135 | 2025-03-11 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-04-22 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Gradient Increase; Perivalvular Leak | Pulmonary Hypertension | P140031 |  | 1 | Event Description: AS REPORTED BY THE IMPLANT PATIENT REGISTRY, THE PATIENT'S RELATIVE NOTIFIED THAT THE PATIENT EXPIRED APPROXIMATELY THREE MONTHS AFTER A TRANSFEMORAL TMVR PROCEDURE WITH A 29 MM SAPIEN 3 VALVE INSIDE A NON-EDWARDS SURGICAL VALVE IN THE MITRAL POSITION. A MEDICAL RECORD REVIEW CONFIRMED THE PATIENT PASSED AWAY DUE TO MITRAL VALVE RESTENOSIS OF THEIR SURGICAL VALVE, REQUIRING A VIV AND EARLY FAILURE OF THE 29 MM S3 VALVE. THE TMVR PROCEDURE USING A 29 MM S3 VALVE VIA TRANSVENOUS, TRANSSEPTAL APPROACH WAS SUCCESSFUL. HOWEVER, MULTIPLE ATTEMPTS TO PLACE OCCLUDER DEVICES THROUGH THE PVL WERE UNSUCCESSFUL, AND THE DECISION WAS MADE TO MEDICALLY MANAGE THE PVL. ONE MONTH AFTER THE MITRAL VIV IMPLANT, THE S3 VALVE SHOWED DEGENERATION WITH INCREASED PVL AND GRADIENTS UP TO 13 MMHG, ASSOCIATED WITH SIGNIFICANT PULMONARY HYPERTENSION (RVSP ~ 70 MMHG), RV DILATION, INTERVENTRICULAR SEPTAL DEVIATION, AND IMPAIRED LV FILLING PARAMETERS. THE PATIENT'S STRESS TEST INDICATED LATERAL WALL ISCHEMIC CHANGES (NSTEMI). AN ECHO PERFORMED APPROXIMATELY THREE MONTHS POST-TMVR SHOWED THE MITRAL BIOPROSTHETIC VALVE WAS WELL-SEATED, WITH PEAK/MEAN GRADIENTS OF 11 AND 6 MMHG, RESPECTIVELY, AND SEVERE PARAVALVULAR MITRAL REGURGITATION. THE FOLLOWING WEEK, THE PATIENT HAD ELEVATED PROSTHETIC GRADIENTS WITH PEAK/MEAN GRADIENTS OF 22 AND 9 MMHG, RESPECTIVELY, AND MVA (CONT. EQU.) OF 2.08 CM', ALONG WITH MODERATE PARAVALVULAR MITRAL REGURGITATION. THE PATIENT PASSED AWAY FIVE DAYS LATER. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS, SECTIONS G6, H2, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAS BEEN ADDED. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, STENOSIS/INCREASED GRADIENT EVENTS FOR THE S3, S3U, S3UR VALVES POST-IMPLANTATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS (ATHEROSCLEROSIS, THROMBOSIS, ENDOCARDITIS, RHEUMATIC FEVER, PANNUS FORMATION) AND/OR PROCEDURAL FACTORS (UNDER-DEPLOYED VALVE, VALVE SIZE SELECTION, PATIENT-PROSTHESIS MISMATCH). THE COMPLAINT FOR 'POST-IMPLANTATION - INCREASED GRADIENTS' WAS CONFIRMED BASED ON PROVIDED MEDICAL RECORDS. IT IS POSSIBLE THAT ONE OR MORE PATIENT/PROCEDURAL FACTORS IDENTIFIED IN THE TECHNICAL SUMMARY MAY HAVE BEEN PRESENT AND CONTRIBUTED TO THE COMPLAINT EVENT. HOWEVER, DUE TO INSUFFICIENT INFORMATION, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST (A LARGE NATIVE MITRAL VALVE ANNULUS, A PREVIOUSLY PLACED SURGICAL VALVE IN THE MITRAL POSITION AND THE 29MM S3 VALVE PLACED WITHIN THE SURGICAL VALVE (VIV) MAY HAVE CONTRIBUTED TO THE EVENT FOR THE PVL. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886153&pc=NPT | 2025587-2025-02922 | 2023-08-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-22 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Atrial Fibrillation; Stroke/CVA; Fever; Low Blood Pressure/ Hypotension; Paralysis; Tachycardia; Ventricular Fibrillation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT PRESENTED WITH LEFT HEMIPLEGIA, WHILE ATRIAL FIBRILLATION AND ATRIAL TACHYCARDIA WERE OBSERVED. IT WAS DETERMINED THE PATIENT HAD SUFFERED A CEREBRAL INFARCTION AND WAS HOSPITALIZED. WHILE HOSPITALIZED, THE PATIENT PRESENTED WITH VENTRICULAR FIBRILLATION, AND SUBSEQUENTLY, A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS IMPLANTED. FOLLOWING THE IMPLANT OF THE CRT-D, THE PATIENT SUFFERED A SECOND CEREBRAL INFARCTION AND WAS ADMINISTERED AN UNKNOWN TREATMENT, AND THE PATIENT'S SYMPTOMS IMPROVED. FOLLOWING TREATMENT, THE PATIENT WAS TRANSFERRED TO A REHABILITATION FACILITY. SUBSEQUENTLY, THE PATIENT DEVELOPED A COVID-19 INFECTION, AND WAS HOSPITALIZED, PRESENTING WITH A FEVER AND HYPOTENSION. IT WAS REPORTED THE PATIENT THEN DIED DUE TO COVID-19. PER THE PHYSICIAN, THE ATRIAL FIBRILLATION AND ATRIAL TACHYCARDIA CONTRIBUTED TO THE FIRST CEREBRAL INFARCTION, WHILE VENTRICULAR FIBRILLATION CONTRIBUTED TO THE SECOND CEREBRAL INFARCTION.Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT PRESENTED WITH LEFT HEMIPLEGIA, WHILE ATRIAL FIBRILLATION AND ATRIAL TACHYCARDIA WERE OBSERVED. IT WAS DETERMINED THE PATIENT HAD SUFFERED A CEREBRAL INFARCTION AND WAS HOSPITALIZED. WHILE HOSPITALIZED, THE PATIENT PRESENTED WITH VENTRICULAR FIBRILLATION, AND SUBSEQUENTLY, A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS IMPLANTED. FOLLOWING THE IMPLANT OF THE CRT-D, THE PATIENT SUFFERED A SECOND CEREBRAL INFARCTION AND WAS ADMINISTERED AN UNKNOWN TREATMENT, AND THE PATIENT'S SYMPTOMS IMPROVED. FOLLOWING TREATMENT, THE PATIENT WAS TRANSFERRED TO A REHABILITATION FACILITY. SUBSEQUENTLY, THE PATIENT DEVELOPED A COVID-19 INFECTION, AND WAS HOSPITALIZED, PRESENTING WITH A FEVER AND HYPOTENSION. IT WAS REPORTED THE PATIENT THEN DIED DUE TO COVID-19. PER THE PHYSICIAN, THE ATRIAL FIBRILLATION AND ATRIAL TACHYCARDIA CONTRIBUTED TO THE FIRST CEREBRAL INFARCTION, WHILE VENTRICULAR FIBRILLATION CONTRIBUTED TO THE SECOND CEREBRAL INFARCTION. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, BOTH THE VALVE AND THE DELIVERY CATHETER SYSTEM MAY HAVE CONTRIBUTED TO THE FIRST AND SECOND STROKE. Manufacturer Narrative: UPDATED DATA: H6. H11. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0011558721); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE UPDATED DATA: B5. ADDED SECOND PARAGRAPH. D10. H6. H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886502&pc=OZD | 1220648-2025-27741 | 2025-04-03 04:00 | Death | ABIOMED, INC. | 2025-04-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS ADMITTED FOR ACUTE MYOCARDIAL INFARCTIONS/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP. THE PATIENT WAS INFUSED WITH ONE UNIT OF RED BLOOD CELLS TO ADDRESS HEMOLYSIS. AFTER EIGHT DAYS OF SUPPORT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA CP AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. INSTRUCTIONS FOR USE: SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ D4 MODEL NUMBER WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27741. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27741. H10 INSTRUCTIONS FOR USE WERE INCOMPLETE ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27741. Manufacturer Narrative: IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). THE IMPELLA CP WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886625&pc=BZD | 2518422-2025-104646 | 2024-10-13 04:00 | Death | RESPIRONICS, INC. | 2025-04-22 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. IF ADDITIONAL INFORMATION IS OBTAINED ABOUT THIS EVENT, THE PMS CLINICAL EXPERT WILL REVIEW. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING EVALUATION AN ERROR INDICATING A REBOOT OF THE CENTRAL PROCESSING UNIT (CPU). PER THE DREAMSTATION SERVICE AND TECHNICAL REFERENCE MANUAL, THIS LEVEL ERROR DOES NOT AFFECT DEVICE FUNCTIONALITY. THE SERVICE TECHNICIAN RECOMMENDED THE DEVICE BE SCRAPPED. NO FURTHER INVESTIGATION DETAILS WERE PROVIDED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886728&pc=NVY | 2017865-2025-50846 | 2025-04-02 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-22 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Respiratory Failure | P950022 |  | 1 | Event Description: IT WAS NOTED THAT DURING THE IMPLANT, THE PATIENT SUFFERED AN ARRHYTHMIA AND RESPIRATORY FAILURE WITH THE RIGHT VENTRICULAR (RV) LEAD. THERE IS NO ALLEGATION OF ABBOTT PRODUCT INVOLVEMENT IN THE PATIENT'S DEATH. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886742&pc=GEI | 3010377594-2025-00005 | 2024-11-18 05:00 | Death | BOLDER SURGICAL | 2025-04-22 04:00 | GEI | COOLSEAL¿ TRINITY | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K211579 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A COOLSEAL TRINITY PROCEDURE ON (B)(6) 2024, A COOLSEAL TRINITY 30CM AND COOLSEAL GENERATOR WERE USED DURING A THORACOSCOPIC PROCEDURE ON A CHILD (6 YEARS OLD, FEMALE), DISSECTING AND SEALING VESSELS. DURING THE PROCEDURE IT WAS IDENTIFIED THAT THE PATIENT WAS BLEEDING FROM A PREVIOUSLY DISSECTED AND SEALED BRANCH OF THE PULMONARY ARTERY, AND THERE WERE CONCERNS OF A POTENTIAL SEAL FAILURE. THE PATIENT SUFFERED SIGNIFICANT BLEEDING WHICH COULD NOT BE STOPPED, RESULTING IN PATIENT DEATH. ON A FOLLOW UP EMAIL IT WAS DETERMINED THAT THE BLEEDING STARTED 3 HOURS AFTER THE SURGERY BEGAN, MAJOR BLEEDING WAS ENCOUNTERED FROM A CORNER OF A PREVIOUSLY SEALED AND DIVIDED BASAL ARTERIAL BRANCH. THE PROCEDURE WAS A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY. NO PRE-EXISTENT CONDITIONS WERE REPORTED. MULTIPLE STEPS WERE TAKEN AFTER THE BLEEDING WAS OBSERVED TO TREAT IT, COMPRESSION, TRANEXAMIC ACID AND BLOOD TRANSFUSION. NO OTHER INFORMATION WAS RECEIVED. Manufacturer Narrative: H6: APPROPRIATE TERM/CODE NOT AVAILABLE: SUGGESTED CODE : ELECTROSURGICAL CUTTER AND COAGULATOR. LOT NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886840&pc=DYB | 3007284313-2025-03921 | 2025-03-06 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-22 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Patient Device Interaction Problem | Perforation of Vessels | K160254 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR A THORACIC AORTIC ANEURYSM USING NON-GORE STENT GRAFTS (VALIANT THORACIC STENT GRAFT, MEDTRONIC). THE FIRST VALIANT STENT GRAFT WAS INSERTED FROM THE RIGHT FEMORAL ARTERY, BUT IT COULD NOT BE ADVANCED THROUGH THE RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY. ONCE THE DEVICE WAS REMOVED, A 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS INSERTED FROM THE RIGHT FEMORAL ARTERY AND SUCCESSFULLY ADVANCED TO THE LEVEL OF THE RENAL ARTERIES. THE DEVICE WAS THEN INSERTED THROUGH THE SHEATH, AND TWO VALIANT STENT GRAFTS WERE IMPLANTED IN THE DESCENDING AORTA AS PLANNED. ACCESS ANGIOGRAPHY REVEALED VASCULAR INJURY NEAR THE RIGHT EXTERNAL ILIAC ARTERY AND RIGHT COMMON ILIAC ARTERY. THE PATIENT WENT INTO SHOCK AND A CARDIAC ARREST. CARDIAC MASSAGE WAS INITIATED, AND A GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE (VIABAHN DEVICE) WAS USED AS TREATMENT. THE FIRST VIABAHN (10 MM X 10 CM) WAS IMPLANTED AT THE SITE OF VASCULAR INJURY (RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY), AND THE SECOND VIABAHN (9 MM X 10 CM) WAS IMPLANTED SLIGHTLY PROXIMAL TO THE LEVEL OF THE FEMORAL HEAD. THE NATIVE VESSEL WAS REPORTEDLY ABOUT 8 MM IN DIAMETER. AFTER PLACEMENT, A TYPE III ENDOLEAK WAS OBSERVED AT THE OVERLAP SITE OF THE TWO VIABAHNS. BALLOON MOLDING WAS PERFORMED, BUT THE LEAK DID NOT RESOLVE. WHEN AN 8 FR OR 10 FR WAS INSERTED IN ORDER TO DELIVER THE THIRD VIABAHN TO THE LEAK SITE, THE OVERLAP OF THE TWO VIABAHNS BECAME UNINTENTIONALLY DISLODGED. THE VIABAHN IMPLANTED ON THE PROXIMAL SIDE WAS COMPLETELY PUSHED OUT OF THE VESSEL. THE GUIDEWIRE WAS ALSO PULLED OUT OF THE TRUE LUMEN, MAKING RE-CANNULATION IMPOSSIBLE. A BALLOON OCCLUSION WAS PERFORMED FROM THE CONTRALATERAL SIDE, AND THE NEXT COURSE OF TREATMENT WAS BEING CONSIDERED. THE PATIENT REMAINED IN CARDIAC ARREST, AND CARDIAC MASSAGE WAS CONTINUED. BLEEDING FROM THE INTUBATION TUBE WAS ALSO OBSERVED. AT THIS POINT, TWO HOURS HAD PASSED SINCE THE CARDIAC ARREST, AND THE HEARTBEAT HAD NOT RESUMED. THE PHYSICIAN DECIDED TO DISCONTINUE RESUSCITATION, AND THE PATIENT WAS DECEASED. <ADDITIONAL INFORMATION OBTAINED FROM THE GORE FSA ON APRIL 11, 2025> THE CAUSE OF DEATH IS UNKNOWN. THE PHYSICIAN'S COMMENT: THE CAUSAL RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND GORE DEVICES (SHEATH/VIABAHN) IS UNKNOWN. REGARDING THE VASCULAR INJURY, THERE WAS SIGNIFICANT RESISTANCE DURING INSERTION OF THE FIRST VALIANT STENT GRAFT, AND IT WAS NOT POSSIBLE TO ADVANCE THROUGH THE AREA NEAR THE RIGHT EXTERNAL AND RIGHT COMMON ILIAC ARTERIES. THERE WAS NO RESISTANCE DURING INSERTION OF THE 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH, SUGGESTING THAT THE VASCULAR INJURY LIKELY OCCURRED DURING THE INSERTION OF THE FIRST STENT GRAFT. AS THE HEARTBEAT DID NOT RESUME EVEN AFTER PLACEMENT OF THE VIABAHN, IT IS LIKELY THAT REGARDLESS OF WHETHER THE VIABAHN COULD HAVE CONTROLLED THE BLEEDING IT WAS ALREADY TOO LATE. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12:ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE (IFU), CAREFUL EVALUATION OF VESSEL SIZE, ANATOMY, TORTUOSITY, AND DISEASE STATE (INCLUDING CALCIFICATION, PLAQUE, AND THROMBUS) IN ACCORDANCE WITH THE OBSERVATION IN THE CT IMAGE OR ANGIOGRAPHY OBTAINED BY PRE-OPERATION/OPERATION OR VISUAL OBSERVATION OBTAINED BY OPERATION IS REQUIRED TO ENSURE SUCCESSFUL USE OF GORE® DRYSEAL FLEX INTRODUCER SHEATH. [IF VESSEL SIZE IS SMALLER THAN THE INTRODUCER SHEATH OUTER DIAMETER OR IF VESSEL IS NOT ADEQUATE FOR ACCESS, MAJOR BLEEDING, VESSEL DAMAGE, OR EMBOLISM, MAY RESULT ADDITIONALLY, THE IFU STATES: ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO VASCULAR TRAUMA (I.E., DISSECTION, RUPTURE, PERFORATION, TEAR, ETC.) AND BLEEDING. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886897&pc=GEI | 3010377594-2025-00006 | 2024-11-18 05:00 | Death | BOLDER SURGICAL | 2025-04-22 04:00 | GEI | COOLSEAL¿ GENERATOR | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K211579 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A COOLSEAL TRINITY PROCEDURE ON (B)(6) 2024, A COOLSEAL TRINITY 30CM AND COOLSEAL GENERATOR WERE USED DURING A THORACOSCOPIC PROCEDURE ON A CHILD (6 YEARS OLD, FEMALE), DISSECTING AND SEALING VESSELS. DURING THE PROCEDURE IT WAS IDENTIFIED THAT THE PATIENT WAS BLEEDING FROM A PREVIOUSLY DISSECTED AND SEALED BRANCH OF THE PULMONARY ARTERY, AND THERE WERE CONCERNS OF A POTENTIAL SEAL FAILURE. THE PATIENT SUFFERED SIGNIFICANT BLEEDING WHICH COULD NOT BE STOPPED, RESULTING IN PATIENT DEATH. ON A FOLLOW UP EMAIL IT WAS DETERMINED THAT THE BLEEDING STARTED 3 HOURS AFTER THE SURGERY BEGAN, MAJOR BLEEDING WAS ENCOUNTERED FROM A CORNER OF A PREVIOUSLY SEALED AND DIVIDED BASAL ARTERIAL BRANCH. THE PROCEDURE WAS A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY. NO PRE-EXISTENT CONDITIONS WERE REPORTED. MULTIPLE STEPS WERE TAKEN AFTER THE BLEEDING WAS OBSERVED TO TREAT IT, COMPRESSION, TRANEXAMIC ACID AND BLOOD TRANSFUSION. NO OTHER INFORMATION WAS RECEIVED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED LOT/SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. THE DEVICE WAS RETURNED FOR INSPECTION WHICH FOUND NO CONCERNS AND MET THE MANUFACTURER´S SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886902&pc=NIP | 2017233-2025-06099 | 2025-03-06 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-22 04:00 | NIP | GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE | Device-Device Incompatibility |  | P040037 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR A THORACIC AORTIC ANEURYSM USING NON-GORE STENT GRAFTS (VALIANT THORACIC STENT GRAFT, MEDTRONIC). THE FIRST VALIANT STENT GRAFT WAS INSERTED FROM THE RIGHT FEMORAL ARTERY, BUT IT COULD NOT BE ADVANCED THROUGH THE RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY. ONCE THE DEVICE WAS REMOVED, A 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS INSERTED FROM THE RIGHT FEMORAL ARTERY AND SUCCESSFULLY ADVANCED TO THE LEVEL OF THE RENAL ARTERIES. THE DEVICE WAS THEN INSERTED THROUGH THE SHEATH, AND TWO VALIANT STENT GRAFTS WERE IMPLANTED IN THE DESCENDING AORTA AS PLANNED. ACCESS ANGIOGRAPHY REVEALED VASCULAR INJURY NEAR THE RIGHT EXTERNAL ILIAC ARTERY AND RIGHT COMMON ILIAC ARTERY. THE PATIENT WENT INTO SHOCK AND A CARDIAC ARREST. CARDIAC MASSAGE WAS INITIATED, AND A GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE (VIABAHN DEVICE) WAS USED AS TREATMENT. THE FIRST VIABAHN (10 MM X 10 CM) WAS IMPLANTED AT THE SITE OF VASCULAR INJURY (RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY), AND THE SECOND VIABAHN (9 MM X 10 CM) WAS IMPLANTED SLIGHTLY PROXIMAL TO THE LEVEL OF THE FEMORAL HEAD. THE NATIVE VESSEL WAS REPORTEDLY ABOUT 8 MM IN DIAMETER. AFTER PLACEMENT, A TYPE III ENDOLEAK WAS OBSERVED AT THE OVERLAP SITE OF THE TWO VIABAHNS. BALLOON MOLDING WAS PERFORMED, BUT THE LEAK DID NOT RESOLVE. WHEN AN 8 FR OR 10 FR WAS INSERTED IN ORDER TO DELIVER THE THIRD VIABAHN TO THE LEAK SITE, THE OVERLAP OF THE TWO VIABAHNS BECAME UNINTENTIONALLY DISLODGED. THE VIABAHN IMPLANTED ON THE PROXIMAL SIDE WAS COMPLETELY PUSHED OUT OF THE VESSEL. THE GUIDEWIRE WAS ALSO PULLED OUT OF THE TRUE LUMEN, MAKING RE-CANNULATION IMPOSSIBLE. A BALLOON OCCLUSION WAS PERFORMED FROM THE CONTRALATERAL SIDE, AND THE NEXT COURSE OF TREATMENT WAS BEING CONSIDERED. THE PATIENT REMAINED IN CARDIAC ARREST, AND CARDIAC MASSAGE WAS CONTINUED. BLEEDING FROM THE INTUBATION TUBE WAS ALSO OBSERVED. AT THIS POINT, TWO HOURS HAD PASSED SINCE THE CARDIAC ARREST, AND THE HEARTBEAT HAD NOT RESUMED. THE PHYSICIAN DECIDED TO DISCONTINUE RESUSCITATION, AND THE PATIENT WAS DECEASED. ADDITIONAL INFORMATION OBTAINED FROM THE GORE FSA ON APRIL 11, 2025. THE CAUSE OF DEATH IS UNKNOWN. THE PHYSICIAN'S COMMENT: THE CAUSAL RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND GORE DEVICES (SHEATH/VIABAHN) IS UNKNOWN. REGARDING THE VASCULAR INJURY, THERE WAS SIGNIFICANT RESISTANCE DURING INSERTION OF THE FIRST VALIANT STENT GRAFT, AND IT WAS NOT POSSIBLE TO ADVANCE THROUGH THE AREA NEAR THE RIGHT EXTERNAL AND RIGHT COMMON ILIAC ARTERIES. THERE WAS NO RESISTANCE DURING INSERTION OF THE 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH, SUGGESTING THAT THE VASCULAR INJURY LIKELY OCCURRED DURING THE INSERTION OF THE FIRST STENT GRAFT. AS THE HEARTBEAT DID NOT RESUME EVEN AFTER PLACEMENT OF THE VIABAHN, IT IS LIKELY THAT REGARDLESS OF WHETHER THE VIABAHN COULD HAVE CONTROLLED THE BLEEDING IT WAS ALREADY TOO LATE. Manufacturer Narrative: HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: ENDOLEAK AND/OR ENDOTENSION W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21886928&pc=NIP | 2017233-2025-06100 | 2025-03-06 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-22 04:00 | NIP | GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE | Patient-Device Incompatibility; Migration |  | P040037 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR A THORACIC AORTIC ANEURYSM USING NON-GORE STENT GRAFTS (VALIANT THORACIC STENT GRAFT, MEDTRONIC). THE FIRST VALIANT STENT GRAFT WAS INSERTED FROM THE RIGHT FEMORAL ARTERY, BUT IT COULD NOT BE ADVANCED THROUGH THE RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY. ONCE THE DEVICE WAS REMOVED, A 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS INSERTED FROM THE RIGHT FEMORAL ARTERY AND SUCCESSFULLY ADVANCED TO THE LEVEL OF THE RENAL ARTERIES. THE DEVICE WAS THEN INSERTED THROUGH THE SHEATH, AND TWO VALIANT STENT GRAFTS WERE IMPLANTED IN THE DESCENDING AORTA AS PLANNED. ACCESS ANGIOGRAPHY REVEALED VASCULAR INJURY NEAR THE RIGHT EXTERNAL ILIAC ARTERY AND RIGHT COMMON ILIAC ARTERY. THE PATIENT WENT INTO SHOCK AND A CARDIAC ARREST. CARDIAC MASSAGE WAS INITIATED, AND A GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE (VIABAHN DEVICE) WAS USED AS TREATMENT. THE FIRST VIABAHN (10 MM X 10 CM) WAS IMPLANTED AT THE SITE OF VASCULAR INJURY (RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY), AND THE SECOND VIABAHN (9 MM X 10 CM) WAS IMPLANTED SLIGHTLY PROXIMAL TO THE LEVEL OF THE FEMORAL HEAD. THE NATIVE VESSEL WAS REPORTEDLY ABOUT 8 MM IN DIAMETER. AFTER PLACEMENT, A TYPE III ENDOLEAK WAS OBSERVED AT THE OVERLAP SITE OF THE TWO VIABAHNS. BALLOON MOLDING WAS PERFORMED, BUT THE LEAK DID NOT RESOLVE. WHEN AN 8 FR OR 10 FR WAS INSERTED IN ORDER TO DELIVER THE THIRD VIABAHN TO THE LEAK SITE, THE OVERLAP OF THE TWO VIABAHNS BECAME UNINTENTIONALLY DISLODGED. THE VIABAHN IMPLANTED ON THE PROXIMAL SIDE WAS COMPLETELY PUSHED OUT OF THE VESSEL. THE GUIDEWIRE WAS ALSO PULLED OUT OF THE TRUE LUMEN, MAKING RE-CANNULATION IMPOSSIBLE. A BALLOON OCCLUSION WAS PERFORMED FROM THE CONTRALATERAL SIDE, AND THE NEXT COURSE OF TREATMENT WAS BEING CONSIDERED. THE PATIENT REMAINED IN CARDIAC ARREST, AND CARDIAC MASSAGE WAS CONTINUED. BLEEDING FROM THE INTUBATION TUBE WAS ALSO OBSERVED. AT THIS POINT, TWO HOURS HAD PASSED SINCE THE CARDIAC ARREST, AND THE HEARTBEAT HAD NOT RESUMED. THE PHYSICIAN DECIDED TO DISCONTINUE RESUSCITATION, AND THE PATIENT WAS DECEASED. ADDITIONAL INFORMATION OBTAINED FROM THE GORE FSA ON APRIL 11, 2025: THE CAUSE OF DEATH IS UNKNOWN. THE PHYSICIAN'S COMMENT: THE CAUSAL RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND GORE DEVICES (SHEATH/VIABAHN) IS UNKNOWN. REGARDING THE VASCULAR INJURY, THERE WAS SIGNIFICANT RESISTANCE DURING INSERTION OF THE FIRST VALIANT STENT GRAFT, AND IT WAS NOT POSSIBLE TO ADVANCE THROUGH THE AREA NEAR THE RIGHT EXTERNAL AND RIGHT COMMON ILIAC ARTERIES. THERE WAS NO RESISTANCE DURING INSERTION OF THE 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH, SUGGESTING THAT THE VASCULAR INJURY LIKELY OCCURRED DURING THE INSERTION OF THE FIRST STENT GRAFT. AS THE HEARTBEAT DID NOT RESUME EVEN AFTER PLACEMENT OF THE VIABAHN, IT IS LIKELY THAT REGARDLESS OF WHETHER THE VIABAHN COULD HAVE CONTROLLED THE BLEEDING IT WAS ALREADY TOO LATE. Manufacturer Narrative: HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE BUT ARE NOT LIMITED TO: ENDOLEAK AND/OR ENDOTENSION, MIGRATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21872841&pc=NPT | 2015691-2025-03079 | 2025-03-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-21 04:00 | NPT | SAPIEN 3 ULTRA RESILIA VALVE | Migration or Expulsion of Device | Foreign Body In Patient | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM THE UNITED KINGDOM, IT WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE VALVE DEPLOYMENT, AN ECTOPIC HEART BEAT DUE TO A PACEMAKER CAPTURE ISSUE CAUSED THE PUSHING OF THE BALLOON FROM THE VENTRICLE INTO THE AORTIC ROOT CAUSING THE EMBOLIZATION OF THE VALVE. AFTER MANY ATTEMPTS TO REPOSITION THE VALVE IN THE DESCENDING OR ASCENDING AORTA WITHOUT SUCCESS, A SECOND 20MM NON-EDWARDS BALLOON WAS INSERTED FROM THE RADIAL ACCESS TO CREATE FURTHER OVER EXPANSION FROM A DUAL INFLATION (COMMANDER + NON-EDWARDS BALLOON). HOWEVER, THE VALVE DID NOT REMAIN STABLE AS THE AORTA DIAMETER WAS 38MM. IT WAS DECIDED TO REMOVED THE SYSTEM. DURING THE WITHDRAWAL, IT WAS DIFFICULT TO PULL THE BALLOON BACK INTO THE SHEATH AS THE USE OF TWO BALLOONS LIKELY DEFORMED THE SYSTEM. ONCE THE BALLOON WAS INSIDE THE SHEATH, THE SYSTEM WAS REMOVED AS A UNIT WITHOUT ISSUES. THEN, THE VALVE IN SITU WITHIN THE AORTA PARTIALLY FLIPPED AND THE PATIENT WAS MOVED TO SURGERY. THE PATIENT ARRESTED ON THE TABLE AND PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PROCEDURAL FACTORS (PACEMAKER CAPTURE ISSUE CAUSING THE PUSHING OF THE BALLOON FROM THE VENTRICLE INTO THE AORTIC ROOT), PATIENT FACTORS (PATIENT ANATOMY) AND/OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21872995&pc=NPT | 2025587-2025-02849 | 2024-11-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-21 04:00 | NPT | EVOLUT FX VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information; Patient Device Interaction Problem | Fever; Low Blood Pressure/ Hypotension; Septic Shock; Chills; Shaking/Tremors; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 4 MONTHS AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH FEVER AND CHILLS. INFECTIOUS ENDOCARDITIS WAS SUSPECTED DUE TO RESULTS OF AN ECHOCARDIOGRAM AND AN INCREASE IN AN INFLAMMATORY BIOMARKER. THE PATIENT WAS URGENTLY HOSPITALIZED. ANTIBACTERIAL MEDICATION WAS ADMINISTERED WITH SOME IMPROVEMENT IN FEVERS AND INFLAMMATION. THE CAUSATIVE ORGANISM WAS REPEATEDLY DETECTED, BUT THE INFECTION SITE COULD NOT BE DETERMINED VIA EXAMINATION. IT WAS NOTED THAT THE INFECTION COULD NOT BE CONTROLLED WITH THE TWO TYPES OF ANTIBIOTIC MEDICATION USED. APPROXIMATELY 1 MONTH AND 1 WEEK FOLLOWING HOSPITALIZATION, MASK-LIKE FACIAL FEATURES AND TREMORS APPEARED. DUE TO THE RAPID PROGRESSION OF THESE CONDITIONS, DECARBOXYLASE INHIBITOR BENSERAZIDE WAS ADMINISTERED. THE PROGRESSION OF NEUROLOGICAL SYMPTOMS ONLY SLIGHTLY IMPROVED AND COULD NOT BE STOPPED. AFTER 1 MONTH, THE PATIENT'S FEVER RAPIDLY WORSENED. THE PATIENT'S CONDITION COULD NO LONGER BE CONTROLLED WITH A WIDE-SPECTRUM ANTIBIOTIC MEDICATION. APPROXIMATELY 5 DAYS LATER, NORADRENALIN WAS INITIATED DUE TO A SUDDEN DROP IN BLOOD PRESSURE. ATTEMPTS TO CONTROL THE PATIENT'S BLOOD PRESSURE WERE UNSUCCESSFULLY, HOWEVER. CARDIAC ARREST WAS OBSERVED IN THE EVENING OF THE SAME DAY. SUBSEQUENTLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPTIC SHOCK. IT IS UNKNOWN WHETHER THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT'S DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21873459&pc=DSQ | 2916596-2025-02335 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY DUE TO SPONTANEOUS INTRACRANIAL BLEEDING (ICB). THE PATIENT SUFFERED FROM SPONTANEOUS ICB WHICH WAS NOT RELATED TO ANY DEVICE OR ANTICOAGULATION ISSUES. THE ICB WAS DIAGNOSED VIA A COMPUTED TOMOGRAPHY (CT) SCAN. TREATMENT WAS CONSERVATIVE. FURTHER DETAILS WERE NOT PROVIDED. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS RUNNING AS EXPECTED. THE ICB WAS RELATED TO DEATH AT LEAST. IT WAS REPORTED THAT THE DEATH WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THERE WAS NO FURTHER INFORMATION PROVIDED DUE TO HOSPITAL DATA PROTECTION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), AND THE REPORTED INTRACRANIAL BLEED AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21873620&pc=PCU | 3005099803-2025-01547 | 2016-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-21 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Use of Device Problem; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF AN EVENT INVOLVING AN AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM THROUGH THE ARTICLE "ENDOSCOPIC ULTRASOUND-GUIDED GASTROJEJUNOSTOMY WITH WIRE ENDOSCOPIC SIMPLIFIED TECHNIQUE: MOVE TOWARDS BENIGN INDICATIONS" BY JEAN-MICHEL GONZALEZ, ET AL. PER THE ARTICLE, BETWEEN 2016 TO 2023, 87 PATIENTS SUFFERING FROM BENIGN OR MALIGNANT GASTRIC OUTLET OBSTRUCTION UNDERWENT ENDOSCOPIC ULTRASOUND-GUIDED GASTROJEJUNOSTOMY. ONE PATIENT DIED AFTER THE PROCEDURE BECAUSE HE WAS IN A VERY POOR CONDITION AND DID NOT RECOVER FROM ANESTHESIA. PLEASE SEE THE REFERENCE ARTICLE FOR FULL DETAILS. NOTE: IT WAS REPORTED THAT THE AXIOS STENT WAS IMPLANTED TO TREAT A GASTRIC OUTLET OBSTRUCTION. HOWEVER, PER THE AXIOS STENT AND ELECTROCAUTERY-ENHANCED DELIVERY SYSTEM INSTRUCTIONS FOR USE (IFU), THE STENT IS INDICATED FOR USE TO FACILITATE TRANSGASTRIC OR TRANSDUODENAL ENDOSCOPIC DRAINAGE OF SYMPTOMATIC PANCREATIC PSEUDOCYSTS GREATER THAN OR EQUAL TO 6 CM IN SIZE AND WALLED-OFF NECROSIS GREATER THAN OR EQUAL TO 6 CM IN SIZE THAT ARE ADHERENT TO THE GASTRIC OR BOWEL WALL. THE DEVICE IS NOT INDICATED TO BE PLACED FOR THE TREATMENT OF GASTRIC OUTLET OBSTRUCTION. Manufacturer Narrative: BLOCK B3: THE EXACT DATE OF THE EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY DATE OF JANUARY 1, 2016, IS USED FOR THE ESTIMATED DATE OF EVENT BETWEEN 2016 AND 2023. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WERE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: JEAN-MICHEL GONZALEZ, SOHAIB OUAZZANI, SAAD M, GEOFFROY VANBIERVLIET, MOHAMED GASMI AND MARC BARTHET. "ENDOSCOPIC ULTRASOUND-GUIDED GASTROJEJUNOSTOMY WITH WIRE ENDOSCOPIC SIMPLIFIED TECHNIQUE: MOVE TOWARDS BENIGN INDICATIONS" DEPARTMENT OF GASTROENTEROLOGY, AIX-MARSEILLE UNIVERSITE, AP-HM, HOPITAL NORD, MARSEILLE, FRANCE, HTTPS://DOI.ORG/10.1111/DEN.1489. BLOCK H6: IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21874035&pc=PYX | 1220648-2025-27708 | 2025-03-26 04:00 | Death | ABIOMED, INC. | 2025-04-21 04:00 | PYX | IMPELLA RP | Increase in Pressure | Insufficient Information | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA RP SUPPORT AND AFTER THE IMPLANT, THERE WAS HIGH PURGE PRESSURE OBSERVED. THE PURGE CASSETTE WAS REPLACED WITH NO AVAIL. TISSUE PLASMA ACTIVATOR WAS ADDED TO THE PURGE. SUPPORT CONTINUED UNTIL CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE DATA LOGS CONFIRM HIGH PURGE PRESSURE ALARMS. THE LOGS SHOW A SHIFT IN MOTOR CURRENT AND PLACEMENT SIGNAL COINCIDING WITH A DROP IN PUMP FLOWS. AFTER THIS, THERE IS A 12 MINUTE RISE IN PURGE PRESSURE UNTIL HIGH PURGE PRESSURE ALARM IS TRIGGERED. CLINICAL MENTIONS TISSUE PLASMINOGEN ACTIVATOR (TPA) WAS ADDED AND LOGS SHOW PURGE PRESSURE VALUES RETURNED TO NORMAL AFTER ABOUT 4-5 HOURS. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL BUILD UP AS EVIDENCED BY THE GRADUAL RISE IN PURGE PRESSURE AFTER MOTOR CURRENT AND PLACEMENT SHIFT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP WITH SMARTASSIST SYSTEM: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21874051&pc=DSQ | 2916596-2025-02462 | 2024-11-23 05:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PNEUMONIA AND MULTI-SYSTEM ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE HOSPITAL WAS NOT AUTHORIZED TO SHARE FURTHER INFORMATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21874409&pc=NVN | 2017865-2025-50664 | 2025-03-28 04:00 | Death | ABBOTT | 2025-04-21 04:00 | NVN | ULTIPACE¿ LEAD, 52CM MRI | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Vascular Dissection; Pericardial Effusion | P960013 |  | 1 | Event Description: FOLLOWING AN IMPLANT PROCEDURE, THE PATIENT EXPERIENCED A CARDIAC EFFUSION AND CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS ATTEMPTED, ALTHOUGH NOT SUCCESSFUL. THE PATIENT'S CHEST WAS OPENED UP AND A PERFORATION OF THE RIGHT ATRIAL APPENDAGE AND DISSECTION OF THE AORTA WERE OBSERVED. THE PATIENT WAS THEN MONITORED IN THE INTENSIVE THERAPY UNIT (ITU); HOWEVER, THEY LATER PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21874669&pc=LPB | 2029046-2025-01194 | 2025-03-24 04:00 | Death | BIOSENSE WEBSTER INC | 2025-04-21 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Pulmonary Edema; Respiratory Failure | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CARDIAC ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED HYPOTENSION, CARDIAC ARREST AND RESPIRATORY FAILURE THAT REQUIRED RESUSCITATION. THE PATIENT DIED. MAPPING AND ABLATION IN A PATIENT WITH VENTRICULAR TACHYCARDIA¿S (VTS). PATIENT WAS IN NORMAL SINUS RHYTHM (SR). NO ERRORS OR OTHER WARNINGS WERE SHOWN IN CARTO OR THE NGEN GENERATOR. AFTER ABLATION, PATIENT HAD VERY LOW BLOOD PRESSURE, FALLING SPO2 AND THEY HAD TO RESUSCITATE THE PATIENT ABOUT 2 HOURS WITH ENDING IN DEATH OF THE PATIENT. THERE WAS NO PERICARDIAL EFFUSION. HE COULD NOT BE OXYGENATED BECAUSE OF "BLOODING OUT THE LUNG". MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN CLARIFICATION TO THIS COMPLAINT. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. Manufacturer Narrative: SINCE THE DATE OF DEATH WAS NOT PROVIDED, ADDITIONAL INFORMATION WAS REQUESTED. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. THEREFORE, SINCE NO EVENT DATE AND PROCEDURE DATE HAVE BEEN PROVIDED, WILL ADD THE DATE OF DEATH AS THE ALERT DATE OF (B)(6) 2025. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: IN THE 3500A INITIAL UNDER H11. ADDITIONAL MANUFACTURER NARRATIVE REPORTED, "SINCE THE DATE OF DEATH WAS NOT PROVIDED, ADDITIONAL INFORMATION WAS REQUESTED. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. THEREFORE, SINCE NO EVENT DATE AND PROCEDURE DATE HAVE BEEN PROVIDED, WILL ADD THE DATE OF DEATH AS THE ALERT DATE OF (B)(6) 2025." ADDITIONAL INFORMATION WAS RECEIVED ON 05-MAY-2025 PROVIDING THE DATE OF DEATH AS (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 05-MAY-2025. THE EVENT DATE OF THE ADVERSE EVENT WAS 24-MAR-2025. THE ADVERSE EVENT WAS DISCOVERED DURING USE OF BIOSENSE WEBSTER PRODUCTS. NO CATHETER EXCHANGES OCCURRED DURING THE PROCEDURE. NO TRANSSEPTAL PUNCTURE SITE PERFORMED DURING THE PROCEDURE. THE INTERVENTION PROVIDED WAS RESUSCITATION AND INTUBATION. THE PATIENT PASSED AWAY DURING THE CASE. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS PATIENT CONDITION, AND MOST SPECIFICALLY THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS PULMONARY EDEMA. THE PATIENT HAD PRE-EXISTING CONDITIONS (HEART ATTACK, DIABETES, LONG-TIME SMOKER). THEREFORE, UPDATED THE A3A. SEX, AND B7. MEDICAL HISTORY/PREEXISTING CONDITION, B3. DATE OF EVENT AND H6. HEALTH EFFECT - CLINICAL CODE FIELDS. IN ADDITION, PROVIDED THE PATIENT¿S DATE OF BIRTH. THEREFORE, UPDATED A2. PATIENT AGE AT THE TIME OF EVENT, A2. AGE UNIT, AND A2. DATE OF BIRTH FIELDS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT A PATIENT UNDERWENT A CARDIAC ABLATION WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED HYPOTENSION, CARDIAC ARREST AND RESPIRATORY FAILURE THAT REQUIRED RESUSCITATION. THE PATIENT DIED. THE DEVICE WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J) FOR EVALUATION 16-JUN-2025. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING J&J PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE MAGNETIC AND FORCE FEATURE WERE TESTED AND NO ERRORS WERE OBSERVED. THE ICON WAS VISUALIZED CORRECTLY AND THE FORCE VALUES AND THE VECTOR WERE OBSERVED WITHIN SPECIFICATIONS. NO MAGNETIC OR FORCE ISSUES WERE OBSERVED. IN ADDITION, DURING THE ABLATION CYCLE, THE DEFLECTION AND IRRIGATION SYSTEM WERE REVIEWED, AND THE CURVES WERE DEFLECTING WITHIN SPECIFICATIONS. NO IRRIGATION ISSUES WERE OBSERVED. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO INTERNAL ACTIONS WERE IDENTIFIED. NO MALFUNCTION WAS OBSERVED DURING THE PRODUCT ANALYSIS, OTHER ISSUES OR CIRCUMSTANCES MAY HAVE OCCURRED DURING THE USAGE OF THE DEVICE THAT COMPROMISED ITS PERFORMANCE. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS PATIENT CONDITION. THE ROOT CAUSE OF THE ADVERSE EVENT IS TRACED TO A PULMONARY EDEMA. THE INSTRUCTION FOR USE (IFU) CONTAINS THE FOLLOWING WARNING AND PRECAUTIONS: WHEN USING THE CATHETER WITH CONVENTIONAL SYSTEMS (USING FLUOROSCOPY TO DETERMINE CATHETER TIP LOCATION), OR WITH THE CARTO¿ 3 SYSTEM, CAREFUL CATHETER MANIPULATION MUST BE PERFORMED IN ORDER TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. DO NOT USE EXCESSIVE FORCE TO ADVANCE OR WITHDRAW THE CATHETER WHEN RESISTANCE IS ENCOUNTERED. THE FIRMNESS OF THE BRAIDED TIP DICTATES THAT CARE MUST BE TAKEN TO PREVENT PERFORATION OF THE HEART. AS PART OF JOHNSON & JOHNSON MEDTECH'S QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21874878&pc=OZD | 1220648-2025-27711 | 2025-01-20 05:00 | Death | ABIOMED, INC. | 2025-04-21 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding; Sepsis; Thrombocytopenia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: ON APRIL 1ST, 2025, THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT. ON (B)(6) 2025 THE PATIENT ON IMPELLA 5.5 SUPPORT DEVELOPED HEPARIN INDUCED THROMBOCYTOPENIA AND WAS SWITCHED FROM HEPARIN TO ARGATROBAN. ON (B)(6) 2025 NON-CENTRAL NERVOUS SYSTEM ARTERIAL THROMBOEMBOLISM AND MAJOR INTRA-ABDOMINAL HEMORRHAGE OCCURRED. ANTICOAGULATION WAS DISCONTINUED AND LEFT BICARBONATE IN THE PURGE. THERE WAS SPONTANEOUS BLEEDING FROM BRANCHES OF THE HEPATIC ARTERY WITH HYPOVOLEMIC SHOCK. ATTEMPTED EMBOLIZATION BY INTERVENTIONAL RADIOLOGY WAS NOTED. ON (B)(6) 2025 THE PATIENT SUFFERED AN ISCHEMIC STROKE WITH SYMPTOMS BUT THERE WAS NO SIGNIFICANT DISABILITY. ANTICOAGULATION WAS RESUMED. OCCIPITAL STROKE WITHOUT SEQUELAE. RECOVERED WITHOUT SEQUELAE. ON (B)(6) 2025 THE PATIENT HAD KLEBSIELLA PNEUMONIAE INFECTION AND SEPTIC SHOCK AND EXPIRED TWO DAYS LATER. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS, STROKE/CVA, VASCULAR DAMAGE - PERIPHERAL VESSEL, THROMBUS, AND THROMBOCYTOPENIA HAVE BEEN COMPLETED. DATA LOGS ARE NOT AVAILABLE. PRODUCT WAS RETURNED AND EVALUATED. THE PUMP'S 5S PARAMETERS WERE WITHIN SPECIFICATION. THE PUMP PASSED LASER CONTINUITY TESTING. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF INFECTION/SEPSIS, STROKE/CVA, VASCULAR DAMAGE - PERIPHERAL VESSEL, THROMBUS, AND THROMBOCYTOPENIA WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875034&pc=NGV | 2135147-2025-02096 | 2025-03-27 04:00 | Death | ABBOTT MEDICAL | 2025-04-21 04:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device; Patient Device Interaction Problem | Angina; Cardiac Arrest; Obstruction/Occlusion | P200049 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 A 28MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT USING A 14F AMPLATZER TORQVUE 45X45 DELIVERY SHEATH TO TREAT A LEFT ATRIAL APPENDAGE (LAA). THE PATIENT'S LAA WAS MEASURED UNDER THE FOLLOWING VIEWS: 24MM AT 0 DEGREES, 20MM AT 45 DEGREES, 22MM AT 90 DEGREES, AND 19MM AT 135 DEGREES. THE LANDING ZONE WAS MEASURED AT 27MM UNDER ANGIOGRAPHY IN RIGHT ANTERIOR OBLIQUE (RAO). TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS USED DURING THE PROCEDURE. THE PATIENT'S LEFT ATRIAL PRESSURE WAS 12 MMHG. DURING THE PROCEDURE THE OCCLUDER WAS PARTIALLY RE-CAPTURED MULTIPLE TIMES. THE OCCLUDER WAS REMOVED FROM THE PATIENT AND RE-SIZED TO A NEW 34MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER. THE OCCLUDER WAS DEPLOYED. IT WAS NOTED THAT THERE WAS A 30-40% MITRAL SHOULDER IN THE HIGH TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) ANGLE AND MINIMAL SEPARATION BETWEEN THE LOBE AND THE DISC. DUE TO THIS CLOSE CRITERIA WERE NOT COMPLETELY MET. THERE WAS VISIBLE COMPRESSION ON THE LOBE IN THE SHAPE OF A TIRE. THE END SCREW OFFSET AND THE OCCLUDER DID NOT MOVE OR MIGRATE DURING A TUG TEST. THE OCCLUDER WAS RELEASED FROM THE DELIVERY CABLE. A LEAK WAS OBSERVED THAT WAS LESS THAN 1MM. WITHIN AN HOUR POST-PROCEDURE THE PATIENT PRESENTED WITH CHEST DISCOMFORT. IT WAS DISCOVERED VIA CHEST X-RAY THAT THE OCCLUDER EMBOLIZED TO THE LEFT ATRIUM AND CAUSED A PARTIAL BLOCKAGE. THE OCCLUDER WAS SURGICALLY EXPLANTED. PER THE PHYSICIAN, THE OCCLUDER EMBOLIZED DUE TO POOR STABILITY. THE PATIENT STATUS WAS STABLE.Event Description: IT WAS REPORTED ON (B)(6) 2025 A 28MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT USING A 14F AMPLATZER TORQVUE 45X45 DELIVERY SHEATH TO TREAT A LEFT ATRIAL APPENDAGE (LAA). THE PATIENT'S LAA WAS MEASURED UNDER THE FOLLOWING VIEWS: 24MM AT 0 DEGREES, 20MM AT 45 DEGREES, 22MM AT 90 DEGREES, AND 19MM AT 135 DEGREES. THE LANDING ZONE WAS MEASURED AT 27MM UNDER ANGIOGRAPHY IN RIGHT ANTERIOR OBLIQUE (RAO). TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS USED DURING THE PROCEDURE. THE PATIENT'S LEFT ATRIAL PRESSURE WAS 12 MMHG. DURING THE PROCEDURE THE OCCLUDER WAS PARTIALLY RE-CAPTURED MULTIPLE TIMES. THE OCCLUDER WAS REMOVED FROM THE PATIENT AND RE-SIZED TO A NEW 34MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER. THE OCCLUDER WAS DEPLOYED. IT WAS NOTED THAT THERE WAS A 30-40% MITRAL SHOULDER IN THE HIGH TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) ANGLE AND MINIMAL SEPARATION BETWEEN THE LOBE AND THE DISC. DUE TO THIS CLOSE CRITERIA WERE NOT COMPLETELY MET. THERE WAS VISIBLE COMPRESSION ON THE LOBE IN THE SHAPE OF A TIRE. THE END SCREW OFFSET AND THE OCCLUDER DID NOT MOVE OR MIGRATE DURING A TUG TEST. THE OCCLUDER WAS RELEASED FROM THE DELIVERY CABLE. A LEAK WAS OBSERVED THAT WAS LESS THAN 1MM. WITHIN AN HOUR POST-PROCEDURE THE PATIENT PRESENTED WITH CHEST DISCOMFORT. IT WAS DISCOVERED VIA CHEST X-RAY THAT THE OCCLUDER EMBOLIZED TO THE LEFT ATRIUM AND CAUSED A PARTIAL BLOCKAGE. THE OCCLUDER WAS SURGICALLY EXPLANTED. PER THE PHYSICIAN, THE OCCLUDER EMBOLIZED DUE TO POOR STABILITY. THE PATIENT STATUS WAS HOSPITALIZATION. ON (B)(6) 2025 THE PATIENT WENT INTO ASYSTOLE. THE PATIENT HAD A DO NOT RESUSCITATE (DNR) ORDER, THEREFORE NO INTERVENTION WAS PERFORMED. THE PATIENT PASSED AWAY. Manufacturer Narrative: AN EVENT OF DEVICE MIGRATION, AN HOUR AFTER THE IMPLANT, RESIDUAL SHUNT AND PATIENT DEATH AFTER TWO MONTHS OF IMPLANT WAS REPORTED. INFORMATION FROM THE FIELD INDICATED THAT CLOSE CRITERIA WERE NOT COMPLETELY MET. ALSO NOTED THAT EMBOLIZED TO THE LEFT ATRIUM AND CAUSED A PARTIAL BLOCKAGE. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF DEVICE EMBOLIZATION IS MOST LIKELY RELATED TO DEVICE OVERSIZING AND NOT ACHIEVING A STABLE DEVICE POSITION. AND THE CAUSE OF THE REPORTED DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED SURGICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS THE DEVICE WAS SURGICALLY REMOVED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. FROM MEDICAL REVIEW: THE DEATH OCCURRING MORE THAN 2 MONTHS POST-IMPLANT IT IS UNLIKELY TO BE RELATED TO THE IMPLANT PROCEDURE OR AMULET OCCLUDER SINCE THE DEVICE WAS EXPLANTED SURGICALLY. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF DEVICE MIGRATION, AN HOUR AFTER THE IMPLANT AND RESIDUAL SHUNT WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. FROM THE IMAGING RECEIVED IT WAS NOTED THAT THERE WAS LITTLE COMPRESSION AND A LEAK ON THE MITRAL ASPECT OF THE DEVICE. THE DEVICE DID NOT MEET ALL SIGNS OF CLOSE CRITERIA. THE DEVICE HAD A MITRAL SHOULDER AND WAS NOT 2/3 PAST THE LEVEL OF THE CIRCUMFLEX AND VERY LITTLE SEPARATION NOTED ON FLUORO. THE DISC AFTER A TENSION TEST WAS PUSHED IN BY THE DELIVERY CABLE WHICH MAY HAVE DE-STABILIZED THE DEVICE AS WELL. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED EVENTS WERE DUE TO PROCEDURAL COMPLICATIONS. HOWEVER, THIS COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED SURGICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS THE DEVICE WAS SURGICALLY REMOVED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875040&pc=MIH | 3007284313-2025-03915 | 2025-03-26 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-21 04:00 | MIH | GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS | Insufficient Information | Low Blood Pressure/ Hypotension; Rupture | P200030 |  | 1 | Event Description: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT A ZONE 9 ANEURYSM, UTILIZING GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS (CXT321414, PLL161207, PLL161007, AND PLC121000). DURING THE PROCEDURE, ALL DEVICES WERE IMPLANTED WITH NO ISSUES. SOON AFTER THE PATIENT'S BLOOD PRESSURE SEEMED TO DROP, BILATERAL ILIAC RUPTURE WAS OBSERVED, AND PATIENT EXPIRED. IT WAS ALSO REPORTED PHYSICIAN DOES NOT KNOW WHAT CAUSED THIS ISSUE. PHYSICIAN ALSO REPORTED PATIENT HAD A VERY THIN ARTERY. NO FURTHER PATIENT OR CASE INFORMATION IS AVAILABLE. Manufacturer Narrative: H6: CODE B20 ¿ THE DEVICE REMAINS IMPLANTED AND, THEREFORE, WAS NOT AVAILABLE FOR DIRECT ANALYSIS BY GORE. H6: CODE C19-A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE MANUFACTURING SPECIFICATIONS. AS THE DEVICE WAS NOT ACCESSIBLE, THE PRODUCT HISTORY REVIEW (PHR) REVIEW WAS THE EXTEND OF THE INVESTIGATION. NO DEVICE PROBLEM WAS FOUND PER REVIEW OF THESE RECORDS. IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, VASCULAR SPASM OR VASCULAR TRAUMA (E.G., AORTA DISSECTION, AORTA DAMAGE, ILIO-FEMORAL VESSEL DISSECTION, BLEEDING, RUPTURE, DEATH) AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875044&pc=DSQ | 3007042319-2025-00795 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-21 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE EFFICACY OF BRIDGING WITH LOW-MOLECULAR-WEIGHT HEPARIN (LMWH) OR UNFRACTIONATED HEPARIN (UFH) IN LEFT VENTRICULAR ASSIST DEVICE (VAD) PATIENTS. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE DUE TO BLEEDING EVENTS AND OTHER UNKNOWN CAUSES. REASONS FOR BRIDGING INCLUDED SUBTHERAPEUTIC INTERNATIONAL NORMALIZED RATIO (INR), CONCERN FOR PUMP THROMBOSIS, OR PERIPROCEDURAL ANTICOAGULATION. THERE WERE PATIENTS WHO EXPERIENCED BLEEDING AND THROMBOTIC EVENTS. MOST BLEEDING EVENTS WERE GASTROINTESTINAL BLEEDS (GIBS). BLEEDING EVENTS INCLUDED BLEEDING RESULTING IN DEATH, INTRACRANIAL BLEEDING, AND BLEEDING EVENTS REQUIRING GREATER THAN TWO UNITS OF PACKED RED BLOOD CELLS (PRBCS) WITHIN A 24-HOUR PERIOD. THROMBOTIC EVENTS INCLUDED SUSPECTED OR CONFIRMED PUMP THROMBOSIS WITH HEMOLYSIS, HEART FAILURE UNEXPLAINED BY STRUCTURAL DISEASE, OR ABNORMAL PUMP PARAMETERS AND PRESENCE OF HEMOGLOBINURIA, ANEMIA, OR HYPERBILIRUBINEMIA. ISCHEMIC STROKE WITH INFARCTION ON AN IMAGING STUDY WAS ALSO DEFINED AS A THROMBOTIC EVENT. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/49 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: UNFRACTIONATED AND LOW-MOLECULAR-WEIGHT HEPARIN FOR BRIDGING PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICE: AN EVENT-BASED ANALYSIS. ASAIO JOURNAL 2021; 67;1277¿1283. DOI: 10.1097/MAT.0000000000001392. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875070&pc=DSQ | 2916596-2025-02481 | 2025-04-04 04:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO GASTROINTESTINAL (GI) BLEEDING. THE PATIENT WAS TREATED DAILY WITH PACKAGED RED BLOOD CELLS (PRBCS). GI SURGERY DID NOT HELP THE ISSUE. THE PATIENT CHOSE TO STOP ALL CARE AND THE CAUSE OF THE GI BLEED WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF GASTROINTESTINAL (GI) BLEEDING COULD NOT BE CONCLUSIVELY DETERMINED. BLEEDING IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. GI BLEEDING HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875107&pc=NGV | 2135147-2025-02098 | 2025-03-13 04:00 | Death | ABBOTT MEDICAL | 2025-04-21 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 22MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN FOR A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE USING A 14F AMPLATZER AMULET DELIVERY SHEATH. DURING THE PROCEDURE, THE MEAN LEFT ATRIAL PRESSURE WAS MORE THAN 10MMHG AND THE AMULET WAS SUCCESSFULLY IMPLANTED AFTER ONE PARTIAL RECAPTURE. FIVE HOURS AFTER THE CLOSURE, THE HEART OF THE PATIENT WAS STOPPED AND THEY MADE A MASSAGE. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS CONDUCTED AND THEY FOUND A CIRCUMFERENTIAL PERICARDIAL EFFUSION. THE SIZE OF THE EFFUSION WAS 25MM AND A CARDIAC TAMPONADE WAS ALSO PRESENT. WHEN THE PATIENT WAS STABLE, HEY TRIED TO DRAIN THE EFFUSION BUT THE PATIENT HAD A NEW CARDIAC ARREST AND REPORTED PASSED AWAY. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS TAMPONADE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 22 MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN FOR A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE USING A 14F AMPLATZER AMULET DELIVERY SHEATH. DURING THE PROCEDURE, THE MEAN LEFT ATRIAL PRESSURE WAS MORE THAN 10 MMHG, 7000 UI OF HEPARIN WAS ADMINISTRATED AND THE AMULET WAS SUCCESSFULLY IMPLANTED AFTER ONE PARTIAL RECAPTURE. FIVE HOURS AFTER THE CLOSURE, THE HEART OF THE PATIENT WAS STOPPED AND THEY MADE A MASSAGE. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS CONDUCTED AND THEY FOUND A CIRCUMFERENTIAL PERICARDIAL EFFUSION. THE SIZE OF THE EFFUSION WAS 25 MM AND A CARDIAC TAMPONADE WAS ALSO PRESENT. WHEN THE PATIENT WAS STABLE, HEY TRIED TO DRAIN THE EFFUSION BUT THE PATIENT HAD A NEW CARDIAC ARREST AND REPORTED PASSED AWAY. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS TAMPONADE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF PERICARDIAL EFFUSION, CARDIAC TAMPONADE, CARDIAC ARREST, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON THE INFORMATION AVAILABLE, THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION, CARDIAC TAMPONADE AND CARDIAC ARREST, AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS PERICARDIOCENTESIS WAS PERFORMED TO RESOLVE PERICARDIAL EFFUSION AND RESUSCITATION WAS PERFORMED DUE TO THE CARDIAC ARREST. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: COMPONENT CODE: 4755 PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4118 TYPE OF INVESTIGATION NOT YET DETERMINED. INVESTIGATION FINDINGS: 3233 RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSIONS: 11 CONCLUSION NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875157&pc=DSQ | 2916596-2025-02184 | 2025-03-13 04:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. ADDITIONAL INFORMATION WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE EXPLANTED FOR EVALUATION. IT WAS COMMUNICATED THAT DUE TO HOSPITAL POLICY, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875359&pc=OZD | 1220648-2025-27712 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-21 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemorrhage/Bleeding; Rupture | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH UNDERGOING HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED A VESSEL RUPTURE AND POSITIONING ISSUES. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. WHILE INSERTING THE 14FR PEEL AWAY SHEATH, THE VESSEL RUPTURED, AND THE GROIN WAS TREATED WITH A COVERED STENT TO CLOSE THE DAMAGED VESSEL. THE PUNCTURE WAS SEALED WITH MANTA CLOSING DEVICE. THE PHYSICIAN DECIDED TO SWITCH TO THE RIGHT SIDE FOR ACCESS, AND THE IMPELLA CP DEVICE WAS IMPLANTED. DURING THE HIGH-RISK PCI, AFTER INTERVENTION OF THE LEFT ANTERIOR DESCENDING ARTERY, THE PATIENT BECAME CRITICALLY UNSTABLE. THE IMPELLA CP DEVICE KEPT PATIENT ON A LOW-LEVEL STABLE. LEFT VENTRICULAR EJECTION FRACTION WAS 10% PRIOR TO THE PROCEDURE AND BECAME WORSE. IT WAS FURTHER NOTED THAT DURING THE IMPELLA MCS, SUCTION WAS PRESENT DUE TO LACK OF PRELOAD AND BEGINNING RIGHT HEART FAILURE. SEVERAL POSITIONING ALERTS WERE PRESENT DUE TO THE PATIENT'S MOVEMENT, AND THE PUMP HAD TO BE REPOSITIONED SEVERAL TIMES. SUBSEQUENTLY, THE WAS TRANSFERRED TO A UNIVERSITY HOSPITAL FOR TREATMENT BUT WOULD ULTIMATELY EXPIRE. Manufacturer Narrative: VASCULAR DAMAGE - PERIPHERAL VESSEL: DURING INITIAL IMPLANTATION ATTEMPT FROM LEFT GROIN, THE VESSEL RUPTURED AND THE GROIN WAS TREATED WITH A COVERED STENT TO CLOSE THE DAMAGED VESSEL. PUNCTURE WAS SEALED WITH MANTA CLOSING DEVICE. THE PUMP WAS NOT RETURNED. THE CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. POSITIONING ISSUES: AFTER PLACEMENT OF THE IMPELLA CP AND FOLLOWING HRPCI PROCEDURE, SUCTION WAS PRESENT DUE TO LACK OF PRELOAD AND BEGINNING RHF. SEVERAL POSITIONING ALERTS WERE PRESENT DUE TO PATIENT MOVEMENT. THE PUMP WAS NOT RETURNED. THE CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY PATIENT MOVEMENT DUE TO THE CLINICAL INFORMATION STATING POSITIONING ALARMS OCCURRED DURING PATIENT MOVEMENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875514&pc=DSQ | 2916596-2025-02305 | 2025-03-31 04:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Cardiac Arrest; Dyspnea; Low Blood Pressure/ Hypotension; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS SHORTNESS OF BREATH, HYPOTENSION, AND A URINE CULTURE POSITIVE FOR E. COLI. THE SOURCE OF THE INFECTION WAS BACTEREMIA. THE PATIENT EXPIRED DUE TO SEPSIS LEADING TO CARDIAC ARREST. THE DEATH WAS NOT DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE PUMP WAS EXPLANTED AND WOULD RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES. A SPECIFIC CAUSE FOR THE PATIENT¿S REPORTED INFECTION, AND SEPSIS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE DISTAL PORTION OF THE PUMP CABLE (INCLUDING THE INLINE CONNECTOR) WAS RETURNED IN ONE SEGMENT. THE MODULAR CABLE WAS RETURNED ATTACHED TO THE PUMP CABLE INLINE CONNECTOR. THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE SEALED OUTFLOW GRAFT BEND RELIEF ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE REVEALED A DECREASE IN PUMP FLOW CONSISTENT WITH THE PATIENT¿S EXPLANT PROCEDURE. THE LOG FILES APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. INCIDENTAL FINDING: REVIEW OF THE EXTRACTED LVAD EVENT LOG FILES REVEALED A PUMP RESET ON (B)(6) 2025, AND ON (B)(6) 2025. A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION, AS THE CORRESPONDING CONTROLLER EVENT LOG FILES ARE NOT AVAILABLE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), SEPSIS, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 OF THE IFU, AND SECTION 4 OF THE PATIENT HANDBOOK INSTRUCTS THE USER TO CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS AND URGES PATIENTS TO INFORM THEIR HOSPITAL CONTACT IMMEDIATELY IF THEY FIND SIGNS OF DRIVELINE DAMAGE. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, " PROVIDES ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE IN A SUB-SECTION ENTITLED "WHAT NOT TO DO: DRIVELINE AND CABLES". SECTION 8 OF THE IFU, ¿EQUIPMENT STORAGE AND CARE¿, STATES: "AS NEEDED, CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, USE WARM WATER AND MILD DISH SOAP". SECTION 4 OF THE PATIENT HANDBOOK, ¿LIVING WITH THE HEARTMATE III¿, ALSO CONTAINS INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE. THIS SECTION INSTRUCTS THE USER TO WIPE OFF ANY DIRT OR GRIME. IF THE DRIVELINE GETS DIRTY, USE A TOWEL WITH MILD DISH SOAP AND WARM WATER TO GENTLY CLEAN IT. NEVER SUBMERGE THE DRIVELINE OR OTHER SYSTEM COMPONENTS IN WATER OR LIQUID.¿ NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875798&pc=DSQ | 3007042319-2025-00799 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-21 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Drug Resistant Bacterial Infection | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING DEVICE-RELATED INFECTIONS IN LEFT VENTRICULAR ASSIST DEVICE (VAD) PATIENTS. THE AUTHORS STUDIED PATIENTS WHO UNDERWENT PUMP EXCHANGE AFTER DIAGNOSIS OF A DEVICE-RELATED INFECTION, WHICH INCLUDED DRIVELINE INFECTIONS, EXTERNAL SURFACE OF AN IMPLANTABLE COMPONENT, OR INFECTION OF THE BLOOD CONTACTING SURFACES AFTER PRIMARY IMPLANTATION OF A VAD. PATIENTS WERE INITIALLY HOSPITALIZED AND TREATED WITH MEDICATIONS INITIALLY AND RE-ADMITTED AT LEAST THREE TIMES BEFORE UNDERGOING A PUMP EXCHANGE. ORGANISMS FOUND INCLUDED STAPHYLOCOCCUS AUREUS (METHICILLIN-SENSITIVE AND METHICILLIN-RESISTANT), SERRATIA MARCESCENS, PSEUDOMONAS SPECIES, AND PROTEUS MIRABILIS. THERE WERE PATIENT DEATHS WHICH INCLUDED OPERATIVE MORTALITY (DEATH WITHIN 30-DAYS OR IN-HOSPITAL MORTALITY); HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/47 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PUMP EXCHANGE FOR MAJOR DEVICE RELATED INFECTION IN PATIENTS RECEIVING A DURABLE LEFT VENTRICULAR ASSIST DEVICE. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION. 2025; 44:686¿689. DOI: 10.1016/J.HEALUN.2024.11.032 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21875975&pc=NKM | 2135147-2025-02102 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-21 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "IMPACT OF TRANSCATHETER EDGE-TO- EDGE REPAIR IN PATIENTS WITH CARDIAC SARCOIDOSIS: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CLINICAL IMPACT OF TRANSCATHETER EDGE-TO- EDGE REPAIR IN PATIENTS WITH CS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT TRANSCATHETER EDGE-TO-EDGE REPAIR IS AN EFFECTIVE HEART FAILURE TREATMENT FOR PATIENTS WITH CS-RELATED MR; HOWEVER, THE DRUG-AND CARDIAC DEVICE-REFRACTORY CARDIOMYOPATHY IN THIS POPULATION WARRANTS CAREFUL MANAGEMENT. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS APRIL 2018 TO JUNE 2023. A TOTAL OF 1240 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 1240 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED PREVIOUS HEART FAILURE, ATRIAL FIBRILLATION, HYPERTENSION, DIABETES. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), TISSUE DAMAGE, PERICARDIAL EFFUSION, HEART FAILURE, PROLONGED HOSPITALIZATION, ARRYTHMIA. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PREVIOUS HEART FAILURE, ATRIAL FIBRILLATION, HYPERTENSION, DIABETES. COMPLICATIONS REPORTED INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), TISSUE DAMAGE, PERICARDIAL EFFUSION, HEART FAILURE, PROLONGED HOSPITALIZATION, ARRYTHMIA; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: IMPACT OF TRANSCATHETER EDGE-TO- EDGE REPAIR IN PATIENTS WITH CARDIAC SARCOIDOSIS: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21876062&pc=DSQ | 2916596-2025-02530 | 2025-04-01 04:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Encephalopathy; Hemolysis; Hemorrhage/Bleeding; Respiratory Failure; Gastrointestinal Hemorrhage; Renal Impairment; Lactate Dehydrogenase Increased | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD BEEN ADMITTED FOR BLEEDING ISSUES AND INCREASED LACTATE DEHYDROGENASE (LDH). AFTER TREATMENT DID NOT WORK, THE FAMILY DECIDED TO REMOVE CARE AND STOP LEFT VENTRICULAR ASSIST DEVICE (LVAD).Event Description: ROUTINE LABS REVEALED ELEVATED LDH IN SETTING OF PATIENT WITH HISTORY. THE PATIENT HAD A GASTROINTESTINAL BLEED (GIB) AND WAS OFF ANTICOAGULATION. THEY HAD NO HEMOLYSIS OR PUMP DYSFUNCTION. THE PATIENT WAS PLACED ON HEPARIN DRIP FOR ELEVATED LDH WHICH EVENTUALLY RESOLVED BACK TO BASELINE, UNFORTUNATELY THE USE OF HEPARIN CAUSED FURTHER GASTROINTESTINAL BLEEDING. THE SOURCE OF THE BLEEDING WAS NOT IDENTIFIED ON AN ESOPHAGOGASTRODUODENOSCOPY AND COLONOSCOPY. THE PATIENT RECEIVED MULTIPLE TRANSFUSIONS. SMALL BOWEL ENDOSCOPIES DID NOT REVEAL THE SOURCE OF GIB. THE PATIENT HAD WORSENING RENAL FUNCTION AND HYPOTENSION DUE TO ANEMIA. THEY EXPERIENCED ENCEPHALOPATHY AND EVENTUALLY REQUIRED VASOPRESSORS AND INTUBATION. THEY WERE UNABLE TO BE WEANED FROM THE VENTILATOR. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, HEMOLYSIS, RENAL FAILURE, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS. THIS SECTION (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21876417&pc=DSQ | 2916596-2025-02539 | 2025-04-04 04:00 | Death | THORATEC CORPORATION | 2025-04-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE (RVF) FOLLOWED BY MULTISYSTEM ORGAN FAILURE (MOF). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT HAD SUFFERED FROM A NON-TREATABLE RVF WHICH LED TO A PROGRESSIVE MOF WHICH WAS RELATED TO THE OUTCOME AT LEAST. THE PATIENT WAS ALREADY TREATED BY TEMPORARY ROTAFLOW RIGHT VENTRICULAR ASSIST DEVICE (RVAD) AND WAS NOT WEANABLE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS OPERATING AS EXPECTED, NO ISSUES MENTIONED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION.Event Description: THE PATIENT HAD SECONDARY RVF IN THE BEGINNING, PRIOR TO THE IMPLANT DUE TO CHRONIC HEART FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO AUTOPSY WAS PERFORMED, AND HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21876955&pc=DSQ | 3007042319-2025-00803 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-21 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Hemorrhagic Stroke; Thrombosis/Thrombus; Valvular Insufficiency/ Regurgitation | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING VENTRICULAR ASSIST DEVICE (VAD) THERAPY IN FONTAN CIRCULATORY FAILURE PATIENTS. THE AUTHORS DESCRIBED ONE PATIENT WITH AN ATRIOPULMONARY FONTAN WHO UNDERWENT A REPOSITION OF THEIR VAD FROM ATRIAL CANNULATION WITH PLACEMENT OF THE VAD IN THE LEFT VENTRICULAR (LV) APEX WITH MITRAL VALVE EXCISION AND MYECTOMY. WHILE ON VAD SUPPORT, THE PATIENT EXPERIENCED A PYLETHROMBOSIS AND A HEMORRHAGIC STROKE. THE PATIENT SUBSEQUENTLY DIED DUE TO MULTIORGAN FAILURE. THE STATUS OF THE VAD IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THE EVENT ONLY OCCURRED WITH ONE PATIENT BUT SPECIFIC DETAILS ON THE PATIENT WERE NOT PROVIDED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SINGLE CENTER EXPERIENCE WITH DURABLE CONTINUOUS FLOW SINGLE VENTRICLE ASSIST DEVICE: A VIABLE OPTION IN FONTAN CIRCULATORY FAILURE. ASAIO JOURNAL. 2023; 69;956¿961. DOI: 10.1097/MAT.0000000000001986 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21877222&pc=DSQ | 3007042319-2025-00805 | 2024-11-01 04:00 | Death | HEARTWARE, INC. | 2025-04-21 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Arrhythmia; Cardiac Arrest; Hemorrhage/Bleeding; Unspecified Infection; Tachycardia; Ventricular Fibrillation; Unspecified Nervous System Problem; Asystole | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE SAFETY OF CHEST COMPRESSIONS AS PART OF CARDIOPULMONARY RESUSCITATION (CPR) FOR PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES (VADS). PATIENTS IN THE STUDY WERE DIVIDED BETWEEN THOSE THAT UNDERWENT COMPRESSION AND THOSE THAT DID NOT. THE AUTHORS DESCRIBED THAT PATIENTS WERE ADMITTED TO THE HOSPITAL DUE TO DIAGNOSES OF INFECTION, BLEEDING, UNKNOWN VAD DYSFUNCTION, OUT-OF-HOSPITAL ARREST, AND OTHER UNKNOWN REASONS. PATIENTS EXPERIENCED CARDIAC ARREST DUE TO ARRHYTHMIAS, VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF), ASYSTOLE, PULSELESS ELECTRICAL ACTIVITY, ALTERED MENTAL STATUS, UNKNOWN VAD ALARMS/DYSFUNCTIONS, AND OTHER UNKNOWN REASONS. PATIENTS WITH AND WITHOUT COMPRESSION SUFFERED INTERMEDIATE TO SIGNIFICANT NEUROLOGIC DEFICITS. THERE WERE PATIENT DEATHS IN BOTH COMPRESSION AND NON-COMPRESSION GROUPS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/64 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SAFETY OF CHEST COMPRESSIONS IN PATIENTS WITH A DURABLE LEFT VENTRICULAR ASSIST DEVICE. JACC: HEART FAILURE. 2024. VOL. 12, NO. 11. 1928-1930. DOI: 10.1016/J.JCHF.2024.03.004 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21877357&pc=MIH | 3013164176-2025-02487 | 2025-03-26 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-21 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Insufficient Information | Low Blood Pressure/ Hypotension; Rupture | P020004 |  | 1 | Event Description: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT A ZONE 9 ANEURYSM, UTILIZING GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS (CXT321414) AND GORE® EXCLUDER® AAA ENDOPROSTHESIS (PLL161207, PLL161007, AND PLC121000). DURING THE PROCEDURE, ALL DEVICES WERE IMPLANTED WITH NO ISSUES. SOON AFTER THE PATIENT'S BLOOD PRESSURE SEEMED TO DROP, BILATERAL ILIAC RUPTURE WAS OBSERVED, AND PATIENT EXPIRED. IT WAS ALSO REPORTED PHYSICIAN DOES NOT KNOW WHAT CAUSED THIS ISSUE. PHYSICIAN ALSO REPORTED PATIENT HAD A VERY THIN ARTERY. NO FURTHER PATIENT OR CASE INFORMATION IS AVAILABLE. Manufacturer Narrative: H6: CODE C19-A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOTS MET ALL PRE-RELEASE MANUFACTURING SPECIFICATIONS. AS THE DEVICES WERE NOT ACCESSIBLE, THE PRODUCT HISTORY REVIEW (PHR) REVIEW WAS THE EXTEND OF THE INVESTIGATION. NO DEVICE PROBLEM WAS FOUND PER REVIEW OF THESE RECORDS. UPDATED SECTION G3/G4, AND H6-TYPE OF INVESTIGATION AND INVESTIGATION FINDINGS. Manufacturer Narrative: AS IT IS UNKNOWN WHICH DEVICE IS DIRECTLY IMPLICATED IN THIS PARAPLEGIA EVENT, THE FOLLOWING DEVICES WILL ALSO BE INCLUDED IN THIS REPORT: CATALOG#PLL161207/SERIAL# (B)(6)/ UDI(B)(4). CATALOG #PLL161007/ SERIAL # (B)(6)/ UDI # (B)(4). H6: CODE B20 ¿ THE DEVICE REMAINS IMPLANTED AND, THEREFORE, WAS NOT AVAILABLE FOR DIRECT ANALYSIS BY GORE. H6: CODE C21 - RESULTS PENDING COMPLETION OF PRODUCT HISTORY REVIEW. IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, VASCULAR SPASM OR VASCULAR TRAUMA (E.G., AORTA DISSECTION, AORTA DAMAGE, ILIO-FEMORAL VESSEL DISSECTION, BLEEDING, RUPTURE, DEATH) AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21877656&pc=DSQ | 3007042319-2025-00807 | 2024-09-20 04:00 | Death | HEARTWARE, INC. | 2025-04-21 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Sepsis; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEFT VENTRICULAR ASSIST DEVICE (VAD) INSERTION AFTER CORONARY ARTERY BYPASS GRAFTING (CABG) USING A LESS INVASIVE (LI) APPROACH VERSUS THE FULL STERNOTOMY (FS) TECHNIQUE. THE AUTHORS DESCRIBED PATIENT DEATHS WHICH INCLUDED IN-HOSPITAL DEATHS AND THOSE DURING THE FOLLOW UP PERIOD. THE CAUSES OF IN-HOSPITAL MORTALITY WERE MULTIORGAN FAILURE, REFRACTORY RIGHT VENTRICULAR (RV) FAILURE, SEPSIS, RESPIRATORY FAILURE, AND REFRACTORY TRACHEAL BLEEDING. DEATHS THAT OCCURRED DURING THE FOLLOW UP PERIOD WERE DUE TO UNKNOWN CAUSES. IN THE FULL STERNOTOMY GROUP, THERE WERE PATIENTS WHO EXPERIENCED RV FAILURE WHICH REQUIRED INOTROPES, EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), OR A TEMPORARY RIGHT VAD. IN BOTH LI AND FS GROUPS, ADVERSE EVENTS INCLUDED REOPERATIONS FOR BLEEDING, STROKE, HEMODIALYSIS, PROLONGED VENTILATION, INFECTION, RENAL DYSFUNCTION, UNKNOWN DEVICE MALFUNCTION, TRACHEOSTOMY, ARRHYTHMIAS, WORSENING AORTIC INSUFFICIENCY, OR GASTROINTESTINAL BLEEDS (GIB). THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/67 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: REOPERATIVE LEFT VENTRICULAR ASSIST DEVICE INSERTION FOLLOWING PREVIOUS CORONARY ARTERY BYPASS GRAFTING: LESS INVASIVE VERSUS STERNOTOMY APPROACH. ARTIFICIAL ORGANS. 2025; 0:1¿8. DOI: 10.1111/AOR.14984 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21877722&pc=KWQ | 3008583793-2025-00002 | 2025-03-11 04:00 | Death | ZAVATION MEDICAL PRODUCTS, LLC | 2025-04-21 04:00 | KWQ | LATERAL F3D 18 MM IMPLANT; COBB ELEVATOR | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | K231743 |  | 1 | Event Description: THE PRODUCT MANAGER WAS PRESENT IN THE OPERATING ROOM FOR A CASE INVOLVING A LATERAL L1-L5 LUMBAR INTERBODY FUSION (LLIF) PROCEDURE. THE PATIENT WAS POSITIONED LATERALLY WITH THE LEFT SIDE FACING UP. THE SURGEON MADE A SMALL INCISION ON THE LATERAL FLANK FOR ACCESS TO THE L1 VERTEBRAL BODY. THE PROCEDURE COMMENCED WITH THE PLACEMENT OF A ZAVATION LATERAL RETRACTOR SYSTEM, THE DISK WAS DISSECTIONED, AND THEN THE LATERAL F3D IMPLANT WAS PLACED INTO AT THE L1 LEVEL. THE OPERATION PROGRESSED TO L2, WHERE THE SURGEON UTILIZED A COBB SPINAL ELEVATOR TO EXCISE OSTEOPHYTES FOR ADEQUATE ACCESS TO THE L2 VERTEBRA. DURING THE PROCEDURE, THE PATIENT BEGAN TO EXPERIENCE SIGNIFICANT HEMORRHAGING. A GENERAL SURGEON WAS URGENTLY CALLED TO THE OPERATING ROOM FOR LIFE-SAVING INTERVENTION. UNFORTUNATELY, THE SOURCE OF THE BLEEDING COULD NOT BE IDENTIFIED IN TIME, AND THE PATIENT ULTIMATELY SUCCUMBED TO THE HEMORRHAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21878042&pc=PCU | 3005099803-2025-01783 | 2015-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-21 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation; Sepsis; Pancreatitis; Insufficient Information |  |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE: "COMBINED ENDOSCOPIC ULTRASONOGRAPHY AND ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY IN PATIENTS WITH MALIGNANT DISTAL BILIARY OBSTRUCTION IS ASSOCIATED WITH REDUCED TIME TO ONCOLOGICAL THERAPY COMPARED WITH ERCP AND SAMPLING ALONE," BY JAMES GAUCI, ET AL. PER THE ARTICLE, A RETROSPECTIVE STUDY EXAMINED PATIENTS WHO UNDERWENT ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) PROCEDURE BETWEEN JANUARY 2015 AND DECEMBER 2020. AMONG PATIENTS, THE FOLLOWING OUTCOMES WERE REPORTED: HEMORRHAGE, PANCREATITIS, PERFORATION, SEPSIS, CARDIOPULMONARY EVENTS AND DEATH. PLEASE SEE THE REFERENCED ARTICLE FOR FULL DETAILS AND FULL LISTING OF PHYSICIANS AND HEALTHCARE FACILITIES. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE JANUARY 31, 2015, BASED ON 30 DAYS FROM THE ESTIMATED EVENT DATE OF JANUARY 1, 2025. BLOCK B3: THE EXACT DATE OF THE EVENT WAS NOT REPORTED. THE ESTIMATED EVENT DATE OF JANUARY 1, 2015, IS USED, BASED ON THE START OF THE RETROSPECTIVE STUDY IN JANUARY 2015. BLOCK D4, H4: DETAILED PRODUCT INFORMATION, INCLUDING DEVICE AVAILABILITY FOR EVALUATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: JAMES GAUCI; WEI ON; BHARAT PARANANDI; MATTHEW THOMAS HUGGETT; SIMON EVERETT. "COMBINED ENDOSCOPIC ULTRASONOGRAPHY AND ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY IN PATIENTS WITH MALIGNANT DISTAL BILIARY OBSTRUCTION IS ASSOCIATED WITH REDUCED TIME TO ONCOLOGICAL THERAPY COMPARED WITH ERCP AND SAMPLING ALONE." PANCREAS 2025; 54:E101-E106. BLOCK H6: IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF HEMORRHAGE. IMDRF PATIENT CODE E1021 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF PANCREATITIS. IMDRF PATIENT CODE E2114 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF PERFORATION. IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF SEPSIS. IMDRF PATIENT CODE E2401 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIOPULMONARY EVENTS. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED FATALITIES. IMDRF IMPACT CODE F12 CAPTURES THE SERIOUS INJURY RESULTING FROM REPORTED PATIENT HARMS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21878209&pc=QZI | 2124215-2025-24887 | 2025-01-30 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-21 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Unspecified Infection; Myocardial Infarction; Pneumonia; Respiratory Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. THE PROCEDURE WAS COMPLETED ON (B)(6) 2025 AND THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO "ACUTE HYPOXIC RESPIRATORY FAILURE, BILATERAL PULMONARY INFILTRATE/ (DIFFUSE ALVEOLAR HEMMORRHAGE) DAH, (NONE-ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION) NSTEMI, PNEUMONIA". FURTHER INFORMATION REGARDING THE TIMELINE LEADING UP TO THE PATIENT'S DEATH WAS NOT PROVIDED BY THE FACILITY. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORTS TO OBTAIN THE INFORMATION ARE IN PROGRESS. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE FILED IF THE INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21878731&pc=NPU | 2015691-2025-03104 | 2024-11-09 05:00 | Death | EDWARDS LIFESCIENCES | 2025-04-21 04:00 | NPU | US S3 UR VALVE 29MM | Perivalvular Leak | Atrial Fibrillation; Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED BY IMPLANT PATIENT REGISTRY, A PATIENT UNDERWENT AN TRANSFEMORAL TMVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN MITRAL POSITION. ON POST OP DAY 9, THE PATIENT PASSED AWAY. UPON FURTHER REVIEW OF MEDICAL RECORDS, THE CASE WAS OFF-LABEL AND THE VALVE WAS IMPLANTED IN NATIVE MITRAL ANNULAR CALCIFICATION (MAC). AFTER A REVIEW OF THE MEDICAL RECORDS, THE PATIENT UNDERWENT A SUCCESSFUL TMVR USING A 29MM SAPIEN 3 ULTRA RESILIA VALVE WITH CONCURRENT MITRAL CLIPPING PROCEDURE. DURING THE TMVR PROCEDURE, THERE WERE NO EDWARDS DEVICE MALFUNCTIONS NOR PROCEDURAL COMPLICATIONS. THE THV WAS DEPLOYED WELL WITH NORMAL FUNCTION. THERE WAS SUCCESSFUL PVL PLUG IMPLANT DECREASING THE PVL TO TRACE/MILD AND ASD CLOSURE WAS ALSO PERFORMED. POD#1, THE PATIENT DEVELOPED HEMOLYSIS WITH DARK RED URINE. ASSUMING PVL WAS THE ISSUE, THE PATIENT WAS TAKEN BACK TO CATH LAB TO RE-ADDRESS THE PVL. HOWEVER, IN THE LAB THERE WAS "FAINT EVIDENCE" OF PVL AND THERE WAS NO ACTUAL PVL AMENABLE TO CLOSURE. THROUGHOUT THIS PROCEDURE TO LOOK FOR THE ROOT CAUSE OF THE HEMOLYSIS, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND REQUIRED IABP SUPPORT. THE PATIENT WAS TRANSFERRED TO CCU IN CRITICAL CONDITION. IN THE FOLLOWING DAYS, THE PATIENT DEVELOPED A-FIB WITH RVR, AKI AND WAS NEUROLOGICALLY COMPROMISED WITH CVA DUE TO EMBOLIC EVENTS. THE FAMILY ELECTED TO MAKE THE PATIENT COMFORT CARE AND THE PATIENT EXPIRED ON POD#9. Manufacturer Narrative: THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE NATIVE MITRAL MAC POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. HEMOLYSIS IS THE DESTRUCTION OF RED BLOOD CELLS. THERE ARE MULTIPLE EXTRINSIC AND INTRINSIC FACTORS INCLUDING USE OF EXTRACORPOREAL CIRCULATION, TRANSFUSION, INFECTION, TUMORS, AUTOIMMUNE DISORDERS, MEDICATION SIDE EFFECTS, A METABOLIC ABNORMALITY, AND RED BLOOD CELL MEMBRANE INSTABILITY. NORMAL RED BLOOD CELLS (ERYTHROCYTES) HAVE A LIFESPAN OF ABOUT 120 DAYS. AFTER THE LIFESPAN IS OVER THE RED BLOOD CELLS BREAK DOWN AND ARE REMOVED FROM THE CIRCULATION BY THE SPLEEN. IN SOME MEDICAL CONDITIONS, OR BECAUSE OF TAKING CERTAIN MEDICATIONS, THIS BREAKDOWN OF RED BLOOD CELLS IS INCREASED. MEDICATIONS THAT HAVE BEEN ASSOCIATED WITH HEMOLYSIS INCLUDE ACETAMINOPHEN, PENICILLIN, AND OTHER PAIN MEDICATIONS. RED CELLS MAY BE DESTROYED DUE TO DEFECTS IN THE CELLS THEMSELVES. ANOTHER CAUSE OF HEMOLYTIC ANEMIA IS BREAK DOWN DUE TO MECHANICAL DAMAGE, SUCH AS FROM ARTIFICIAL HEART VALVES OR HEART-LUNG BYPASS; OR A RESULT OF TURBULENT BLOOD FLOW PATTERN AND ERYTHROCYTE DESTRUCTION CAUSED BY PVL OR CENTRAL REGURGITATION. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF THE REPORTED EVENT WAS UNABLE TO BE DETERMINED BUT MAY HAVE BEEN DUE TO PROCEDURAL FACTORS, PATIENT FACTORS (LARGE MITRAL ANNULUS WITH MAC), AND/OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THE PVL WITH SUBSEQUENT HEMOLYSIS/HEMOLYTIC ANEMIA. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (MULTIPLE COMORBIDITIES, FEMALE, ADVANCED AGE, ATRIAL FIBRILLATION) AND/OR PROCEDURAL FACTORS (ADDITIONAL SURGERY TO ADDRESS PVL/HEMOLYSIS) OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), ARRHYTHMIAS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ATRIAL FIBRILLATION IS A COMMON PRE-EXISTING ARRHYTHMIA IN PATIENTS UNDERGOING VALVE REPLACEMENT. IT ALSO CAN BE ASSOCIATED WITH THE PROCEDURE OR CAN BE A PROGRESSION OF THE PATIENT-S DISEASE AT SOME POINT IN THE POST-OPERATIVE PERIOD. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN EW CLINICAL TECHNICAL SUMMARY FOR COMPLAINTS- ATRIAL FIBRILLATION, DESPITE ONGOING EFFORTS TO DECREASE ITS OCCURRENCE, POSTOPERATIVE ATRIAL FIBRILLATION (POAF) REMAINS A FREQUENT COMPLICATION OF CARDIAC SURGERY. ALTHOUGH THE ETIOLOGY OF POAF IS NOT COMPLETELY UNDERSTOOD, SEVERAL MECHANISMS AND RISK FACTORS HAVE BEEN IDENTIFIED. MULTIPLE STUDIES HAVE REPORTED THE MAIN FACTORS ASSOCIATED WITH AN INCREASED RISK OF POAF AS ADVANCED AGE, THE COMPLEXITY OF THE PROCEDURE, A HISTORY OF HEART FAILURE, AND LOW EF IN ADDITION TO A HIGHER EUROSCORE, WHICH IS DESCRIPTIVE OF THE CURRENT THV PATIENT POPULATION. A REVIEW OF THE LITERATURE INVOLVING THE STUDY OF THOUSANDS OF PATIENTS HAS FOUND NO CASES IN WHICH THE CAUSE OF POAF WAS DETERMINED TO BE RELATED TO THE DEVICE BEING IMPLANTED. AS A RESULT, AN IN-DEPTH INVESTIGATION INTO THESE EVENTS, BEYOND DOCUMENTATION OF THE POTENTIAL PREDISPOSING FACTORS, IS OF LIMITED VALUE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST BOTH PATIENT FACTORS AND PROCEDURAL FACTORS AND/OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21870255&pc=OZP | 2032227-2025-156892 | 2025-03-13 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-20 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOSPITAL ON (B)(6) 2025 . THE CAUSE OF DEATH WAS FATAL CAR ACCIDENT AND CUSTOMER ALSO EXPERIENCED HYPOGLYCEMIA WITH THE BLOOD GLUCOSE VALUE OF 50 MG/DL. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-7040B ,MMT-1884 AND MMT-399A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, CUSTOMER WORN INSULIN PUMP AT THE TIME OF PASSING AND SENSOR WAS IN USE AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR 332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040B. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-399A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21870397&pc=FKX | 0008030665-2025-00901 | 2025-04-06 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-20 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT'S DAUGHTER REPORTED THAT THE PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025, DUE TO UNSPECIFIED ¿HEART ISSUES.¿ NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DEATH CERTIFICATE, ESRD DEATH NOTIFICATION, AUTOPSY REPORT, TREATMENT RECORD, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE POC REPORTED THE CAUSE OF DEATH WAS DUE TO HIS CARDIAC HISTORY. THE SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21870407&pc=LKK | 3004209178-2025-06834 | 2025-04-08 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-20 04:00 | LKK | SYNCHROMED III | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM MULTIPLE SOURCES (MANUFACTURER REPRESENTATIVE, HEALTHCARE PROVIDER) REGARDING A PATIENT WHO WAS RECEIV ING MORPHINE DRUG VIA AN IMPLANTABLE PUMP FOR UNKNOWN INDICATIONS FOR USE. IT WAS REPORTED THAT THE PATIENT HAD AN INFECTION. THE CAUSE OF THE EVENT WAS UNKNOWN. THE WHOLE SYSTEM WAS EXPLANTED AND DISCARDED ON (B)(6) 2025. THE ISSUE WAS RESOLVED. THE HEALTHCARE PROVIDER (HCP) HAS NO FURTHER INFORMATION FOR MEDTRONIC. (B)(6) 2025 MPXR1291614 (HCP/REP): ADDITIONAL INFORMATION WAS PROVIDED FROM A HEALTHCARE PROVIDER (HCP) VIA A COMPANY REPRESENTATIVE (REP) STATED THAT THE PATIENT WAS SHOWING SIGNS OF SEPSIS PRIOR TO EXPLANATION THAT WAS DONE ON (B)(6) 2025. THE PATIENT WAS STILL INPATIENT AND INTUBATED IN THE INTENSIVE CARE UNIT (ICU).Event Description: INFORMATION WAS RECEIVED FROM MULTIPLE SOURCES (MANUFACTURER REPRESENTATIVE, HEALTHCARE PROVIDER) REGARDING A PATIENT WHO WAS RECEIV ING MORPHINE DRUG VIA AN IMPLANTABLE PUMP FOR UNKNOWN INDICATIONS FOR USE. IT WAS REPORTED THAT THE PATIENT HAD AN INFECTION. THE CAUSE OF THE EVENT WAS UNKNOWN. THE WHOLE SYSTEM WAS EXPLANTED AND DISCARDED ON (B)(6) 2205. THE ISSUE WAS RESOLVED. THE HEALTHCARE PROVIDER (HCP) HAS NO FURTHER INFORMATION FOR MEDTRONIC. ADDITIONAL INFORMATION WAS PROVIDED FROM A HEALTHCARE PROVIDER (HCP) VIA A COMPANY REPRESENTATIVE (REP) STATED THAT THE PATIENT WAS SHOWING SIGNS OF SEPSIS PRIOR TO EXPLANTATION THAT WAS DONE ON (B)(6) 2025. THE PATIENT WAS STILL INPATIENT AND INTUBATED IN THE INTENSIVE CARE UNIT (ICU).Event Description: ADDITIONAL INFORMATION WAS PROVIDED FROM A HEALTHCARE PROVIDER (HCP) VIA A COMPANY REPRESENTATIVE (REP) STATED THAT THE PATIENT PASSED AWAY ON SUNDAY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 8780 LOT# SERIAL# (B)(6). IMPLANTED: (B)(6) 2025. EXPLANTED: (B)(6) 2025. PRODUCT TYPE CATHETER SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 8780, SERIAL/LOT #: (B)(6), UBD: 17-DEC-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869906&pc=OZD | 1220648-2025-27695 | 2025-03-28 04:00 | Death | ABIOMED, INC. | 2025-04-19 04:00 | OZD | IMPELLA CP | Positioning Failure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AN IMPLANTED IMPELLA CP PUMP SLIPPED INTO THE PATIENT'S AORTA AT THE TIME THAT CPR WAS BEING PERFORMED ON THE PATIENT. THE PUMP WAS REPLACED IN RESPONSE TO THE POSITIONING ISSUE. THE PATIENT REPORTEDLY PASSED AWAY THE FOLLOWING DAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH IMPELLA SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. CLINICAL STATES THAT THE PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION (CPR) WHICH RESULTED IN IMPELLA SLIPPING INTO AORTA. PUMP WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS USE RELATED TO THE CPR PERFORMED WHILE PUMP STILL INSIDE THE PATIENT. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER: 1220648-2025-27695 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ A1-A5 ARE UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21863490&pc=NVY | 2017865-2025-40492 | 2025-03-01 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-18 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Deliver Shock/Stimulation; Under-Sensing | Arrhythmia; Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAD DIED DUE TO UNDER-SENSING OF VENTRICULAR FIBRILLATION. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.Event Description: NEW INFORMATION NOTED THAT THE DEVICE EXHIBITED INTERMITTENT RV UNDER-SENSING DURING A VENTRICULAR FIBRILLATION (VF) EPISODE. ALTHOUGH THE VF THERAPY ASSURANCE ALGORITHM ACTIVATED, THE DEVICE MET RETURN-TO-SINUS CRITERIA BEFORE DELIVERING THERAPY. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: CORRECTION: H6 HEALTH IMPACT - CLINICAL CODE SHOULD ALSO INCLUDE "ARRHYTHMIA."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21863496&pc=NIK | 2017865-2025-40491 | 2025-03-01 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-18 04:00 | NIK | GALLANT HF | Failure to Deliver Shock/Stimulation; Under-Sensing | Arrhythmia; Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAD DIED DUE TO UNDER-SENSING OF VENTRICULAR FIBRILLATION. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.Event Description: NEW INFORMATION NOTED THAT THE DEVICE EXHIBITED INTERMITTENT RV UNDER-SENSING DURING A VENTRICULAR FIBRILLATION (VF) EPISODE. ALTHOUGH THE VF THERAPY ASSURANCE ALGORITHM ACTIVATED, THE DEVICE MET RETURN-TO-SINUS CRITERIA BEFORE DELIVERING THERAPY. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RETURNED. Manufacturer Narrative: CORRECTION: H6 HEALTH IMPACT - CLINICAL CODE SHOULD ALSO INCLUDE "ARRHYTHMIA."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864172&pc=MNS | 2518422-2025-104422 | 2025-04-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864330&pc=DSI | 1218950-2025-000242 | 2024-12-15 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-04-18 04:00 | DSI | INTELLIVUE MX40 802.11A/B/G | Use of Device Problem; Insufficient Information | Cardiac Arrest | K113125 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT EXPERIENCED A CARDIAC ARREST TWO DAYS AFTER PCI AND PASSED AWAY ON (B)(6) 2024 AFTER UNSUCCESSFUL REANIMATION ATTEMPT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. THE MX40 GENERATED AN INOP FOR "LEAD OFF' SO THERE WAS NO ECG MEASUREMENT AVAILABLE. THE INOP WAS SENT TO THE NURSE VIA CAREEVENT BUT NO INTERVENTION WAS PERFORMED TO RESOLVE THE LEAD OFF INOP. IT WAS DETERMINED THE INOP ALARM WAS SENT TO THE PRIMARY NURSE PHONE BUT THE PHONE WAS NOT IN USE; THEREFORE, THE INOP WAS DELIVERED TO THE BUDDY NURSE'S PHONE SUCCESSFULLY. THE CONFIGURATION IS SET TO SEND THE INOP ALARMS TO THE PRIMARY NURSE AND THEN FORWARD ON TO THE BUDDY NURSE ONE TIME BUT INOP ALARMS ARE NOT FORWARDED TO THE WHOLE TEAM. IN ADDITION, DUE TO THE "LEADS OFF" INOP, THERE WAS NO RED ALARM GENERATED, WHICH WOULD HAVE ALERTED THE ENTIRE NURSING TEAM. IT WAS DETERMINED THE DEVICE WAS FUNCTIONING AS EXPECTED. THE PATIENT PASSED WAY DUE TO CARDIAC ARREST. BASED ON THE INFORMATION RECEIVED, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED PATIENT DEATH; HOWEVER, ALARM MANAGEMENT, CLINICAL WORKFLOW, OR OTHER USE ERROR RELATED TO NURSE PHONE ASSIGNMENTS MAY HAVE BEEN FACTORS. Manufacturer Narrative: THE CUSTOMER PROVIDED A BRIEF DESCRIPTION OF THE TECHNICAL FACTORS THAT CONTRIBUTED TO THE EVENT INDICATING THAT THE ALARM SYSTEM (CHAIN FROM MX40 TO MYCO ON WHICH CAREASSIST RUNS) HAS WORKED AS CONFIGURED; HOWEVER, THE CHOSEN CONFIGURATION SETTINGS DID CONTRIBUTE TO THE REPORTED EVENT. A PHILIPS PRODUCT SUPPORT ENGINEER REVIEWED THE LOGS PROVIDED AND CONFIRMED THAT THE MX40 ANNOUNCED THE ECG LEADS OFF CORRECTLY AT THE CENTRAL STATION. THE INFORMATION WAS ALSO REVIEWED BY A PHILIPS CLINICAL EXPERTS WHO PROVIDED THE FOLLOWING ANALYSIS: THE MX40 GENERATED AN INOP FOR "LEAD OFF' SO THERE WAS NO ECG MEASUREMENT AVAILABLE. THE INOP WAS SENT TO THE NURSE VIA CAREEVENT BUT NO INTERVENTION WAS PERFORMED TO RESOLVE THE LEAD OFF INOP. IT WAS DETERMINED THE INOP ALARM WAS SENT TO THE PRIMARY NURSE PHONE BUT THE PHONE WAS NOT IN USE; THEREFORE, THE INOP WAS DELIVERED TO THE BUDDY NURSE'S PHONE SUCCESSFULLY. THE CONFIGURATION IS SET TO SEND THE INOP ALARMS TO THE PRIMARY NURSE AND THEN FORWARD ON TO THE BUDDY NURSE ONE TIME BUT INOP ALARMS ARE NOT FORWARDED TO THE WHOLE TEAM. IN ADDITION, DUE TO THE "LEADS OFF" INOP, THERE WAS NO RED ALARM GENERATED, WHICH WOULD HAVE ALERTED THE ENTIRE NURSING TEAM. IT WAS DETERMINED THE DEVICE WAS FUNCTIONING AS EXPECTED. THE PATIENT PASSED WAY DUE TO CARDIAC ARREST. BASED ON THE INFORMATION RECEIVED, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED PATIENT DEATH; HOWEVER, ALARM MANAGEMENT, CLINICAL WORKFLOW, OR OTHER USE ERROR RELATED TO NURSE PHONE ASSIGNMENTS MAY HAVE BEEN FACTORS. A PHILIPS PRODUCT SUPPORT ENGINEER FOR CAREEVENT REVIEWED THE ADDITIONAL CAREEVENT LOGS THAT WERE PROVIDED AND DETERMINED THE FOLLOWING: ONLY A SCREENSHOT OF THE EVENT REPORTS / LOGS WERE PROVIDED. THE LIMITED LOGS THAT WERE SUPPLIED SHOWS THAT THE MESSAGES WERE RECEIVED IN THE CAREEVENT SYSTEM, BUT THEY WERE NOT ACKNOWLEDGED. THE ALARM ENDED (PATIENT CONDITION ENDED) WITHOUT BEING ACKNOWLEDGED¿ THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING INSTITUTION PHONE # (B)(6). THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864334&pc=MNS | 2518422-2025-104427 | 2025-04-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | MNS | DREAMSTATION ST30 | Contamination; Crack; Failure to Calibrate; Failure to Align; Gas/Air Leak; Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE TOTAL BLOWER HOURS WERE 5414.1 WITH NO ERRORS. THE BLOWER BOX ASSEMBLY WAS CONTAMINATED AND NEEDED TO BE REPLACED. THE ACCESSORY MODULE FLIP DOOR WAS MISSING FROM THE DEVICE, SO IT WAS FITTED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED AS PREVENTIVE MAINTENANCE. IN ADDITION, THE SD CARD WAS ALSO FITTED. THE DEVICE WAS CLEANED AND TESTED. THE DEVICE FAILED AT TEST STEP 114, SO THE BLOWER WAS REPLACED. THE DEVICE WAS RE-TESTED AND PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS PREVIOUSLY RETURNED TO A SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE TOTAL BLOWER HOURS WERE 5414.1 WITH NO ERRORS. THE BLOWER BOX ASSEMBLY WAS CONTAMINATED AND NEEDED TO BE REPLACED. THE ACCESSORY MODULE FLIP DOOR WAS MISSING FROM THE DEVICE, SO IT WAS FITTED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED AS PREVENTIVE MAINTENANCE. IN ADDITION, THE SD CARD WAS ALSO FITTED. THE DEVICE WAS CLEANED AND TESTED. THE DEVICE FAILED AT TEST STEP 114, SO THE BLOWER WAS REPLACED. THE DEVICE WAS RE-TESTED AND PASSED ALL TESTS. THE MANUFACTURER HAS UPDATED BOX H CODING IN THIS REPORT TO CORRECT THE CODING FROM THE PREVIOUS REPORT. THE MANUFACTURER INCLUDED THE UPDATED DEVICE PROBLEM CODES IN THIS REPORT. THE PREVIOUS REPORT HAD THE EVALUATION RESULTS CODE GRID AS CODE C92043. IN THIS REPORT, THE MANUFACTURER HAS ALSO INCLUDED THE CODE FOR MECHANICAL PROBLEM IDENTIFIED, CODE C92079. IN THE PREVIOUS REPORT THE CONCLUSION CODE GRID WAS CODED TO CAUSE TRACED TO ENVIRONMENT, CODE C139473. THE MANUFACTURER HAS INCLUDED THE CODE FOR CAUSE TRACED TO COMPONENT FAILURE, C139472 IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864402&pc=MGB | 21864402 | 2025-04-10 04:00 | Death | ABBOTT VASCULAR | 2025-04-18 04:00 | MGB | PERCLOSE¿ PROSTYLE¿ | Activation, Positioning or Separation Problem; Therapeutic or Diagnostic Output Failure | Cardiac Arrest; Exsanguination; Retroperitoneal Hemorrhage |  |  | 1 | Event Description: DURING CLOSING FROM A RIGHT FEMORAL ACCESS, TWO PERCLOSE PROSTYLE DEVICES FAILED. WHEN THE SECOND PERCLOSE PROSTYLE DEVICE FAILED, A FEMORAL PUNCTURE WAS LEFT WHICH CAUSED A RETRO-PERITONEAL BLEED. THE PRONGS ON THE PERCLOSE DID NOT DEPLOY RIGHT. THE PATIENT WAS ACTIVELY BLEEDING AND BEGAN TO DECOMPENSATE ULTIMATELY LEAD TO CARDIAC ARREST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864413&pc=BZD | 2518422-2025-104430 | 2025-04-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864459&pc=OZD | 1220648-2025-27652 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-18 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE DOCTOR WANTED TO OPTIMIZE THE PUMP POSITIONING UNDER THE PREMISE THAT IT WAS TOO DEEP, BUT THE CATHETER APPEARED TO BE STUCK. THE PHYSICIAN WAS UNABLE TO MOVE THE PUMP FORWARD OR BACKWARD INTO THE DESIRED LOCATION. THE AUTOMATED IMPELLA CONTROLLER (AIC) SUBSEQUENTLY TRIGGERED A PLACEMENT SIGNAL LOW ALARM AND THE PATIENT'S BLOOD PRESSURE BEGAN TO DROP. THE STAFF WAS ADVISED NOT TO INCREASE THE SUPPORT LEVEL UNTIL THE POSITIONING ISSUE COULD BE RESOLVED. UPON FOLLOW UP, IT WAS DISCOVERED THAT DURING THE REPOSITIONING ATTEMPTS, THE T-LOCK HAD BEEN TURNED BY THE STAFF RATHER THAN THE TOUHY-BORST VALVE, WHICH WAS WHY THE CATHETER REMAINED FIRMLY IN POSITION. REPOSITIONING WAS PERFORMED SUCCESSFULLY FOLLOWING THE CORRECTION. THE PATIENT LATER PASSED AWAY THAT SAME DAY. THERE IS NOT ENOUGH INFORMATION TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE POSITIONING ISSUE WAS USE RELATED TO IMPROPER TECHNIQUE OF USING THE TOUHY-BORST LOCK. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS. ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ CORRECT IMPELLA CATHETER POSITION: ¿FOR OPTIMAL POSITIONING OF THE IMPELLA CATHETER, THE INLET AREA OF THE CATHETER SHOULD BE 3.5 CM BELOW THE AORTIC VALVE ANNULUS AND WELL AWAY FROM PAPILLARY MUSCLE AND SUBANNULAR STRUCTURES. THE OUTLET AREA SHOULD BE WELL ABOVE THE AORTIC VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864595&pc=DTB | 2649622-2025-10555 | 2025-03-23 04:00 | Death | MPRI | 2025-04-18 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing; Capturing Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE RIGHT ATRIAL (RA) LEAD EXHIBITED INTERMITTENT UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS AT TIMES AND INAPPROPRIATE PACING OF THE ATRIA.THE RA LEAD ALSO EXHIBITED FAR FIELD R-WAVE (FFRW) OVERSENSING. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND A POSSIBLE CAPTURE ISSUE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864623&pc=DTB | 2649622-2025-10556 | 2025-03-23 04:00 | Death | MPRI | 2025-04-18 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Over-Sensing; Pacing Problem; Under-Sensing | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE RIGHT ATRIAL (RA) LEAD EXHIBITED INTERMITTENT UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS AT TIMES AND INAPPROPRIATE PACING OF THE ATRIA. THE RA LEAD ALSO EXHIBITED FAR FIELD R-WAVE (FFRW) OVERSENSING. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND A POSSIBLE CAPTURE ISSUE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864726&pc=BZD | 2518422-2025-104345 | 2025-01-08 05:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Insufficient Information | Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY DUE TO KIDNEY AND HEART FAILURE. NO MEDICAL INTERVENTION WAS MENTIONED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21864946&pc=MIH | 9612164-2025-01980 | 2025-04-03 04:00 | Death | MEDTRONIC IRELAND | 2025-04-18 04:00 | MIH | VALIANT CAPTIVIA - FF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Vascular Dissection; Pericardial Effusion | P100040 |  | 1 | Event Description: A VALIANT CAPTIVIA STENT GRAFT WAS IMPLANTED DURING THE TREATMENT OF A 65 MM THORACIC AORTIC ANEURYSM INVOLVING THE AORTIC ARCH. IT WAS REPORTED THAT DURING THE INDEX PROCEDURE, A REVERSE TEAR OF BLOOD VESSELS OCCURRED IN THE PROXIMAL ANCHORING AREA OF THE STENT GRAFT. AFTER STENT GRAFT IMPLANTATION, THE PATIENT'S BLOOD PRESSURE AND HEART RATE DECREASED. UPON ANGIOGRAPHIC EXAMINATION, CONTRAST RETENTION WAS OBSERVED IN THE BLOOD VESSELS, AND PERICARDIAL EFFUSION WAS DIAGNOSED. DESPITE CONTACTING A CARDIOLOGIST FOR CONSULTATION AND PERFORMING CARDIOPULMONARY RESUSCITATION AND OTHER RESCUE MEASURES, THE PATIENT EXPIRED ONE HOUR AFTER THE PROCEDURE. ACCORDING TO THE PHYSICIAN, IT WAS CONSIDERED THAT THE SUPPORT FORCE OF THE FREEFLO STENTS TOO STRONG OR THE FREEFLO STENTS WERE TOO SHARP, CAUSING THE REVERSE TEAR. THE PHYSICIAN SUSPECTED THAT THE CAUSE OF DEATH WAS RELATED TO THE BARE STENT OF VALIANT CAPTIVIA STENT GRAFT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: IT WAS REPORTED THAT THE CORRECT DEVICE SIZE WAS USED AND THE STENT WAS DEPLOYED IN A HEALTHY LANDING ZONE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS CONCLUSION: THE REPORTED DISSECTION AND SUBSEQUENT REPORTED ADVERSE EVENTS COULD NOT BE CONFIRMED ON THE FILMS PROVIDED; THEREFORE, THE CAUSE OF THESE EVENTS COULD NOT BE DETERMINED. LACK OF COMPLETE ANGIOGRAMS SHOWING THE DEPLOYMENT OF THE STENT AND THE EXACT MOMENT WHEN THE REVERSE TEAR WAS OBSERVED WERE NOT AVAILABLE FOR A MORE THOROUGH ASSESSMENT OF THE REPORTED EVENT. ANALYSIS OF THE RETURNED FILMS DID NOT REVEAL ANY OBVIOUS OUT-OF-SPECIFICATION STENT GRAFT INTEGRITY ISSUES. THE INSTRUCTIONS FOR USE STATES THAT " MANIPULATION OF WIRES, BALLOONS, CATHETERS, OR ENDOGRAFTS IN THE THORACIC AORTA MAY LEAD TO VASCULAR TRAUMA, INCLUDING AORTIC DISSECTION". MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE ANALYSIS: TWO STILL IMAGES RECEIVED FOR ANALYSIS. FIRST IMAGE DEPICTS A VALIANT CAPTIVIA SHELF CARTON WITH LABEL. LABELLING INFORMATION MATCHES THAT REPORTED IN THE COMPLAINT FILE. SECOND IMAGE DEPICTS THE PROXIMAL HANDLE OF A VALIANT CAPTIVIA ALONGSIDE OTHER DEVICES USED DURING THE PROCEDURE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21865210&pc=DSQ | 2916596-2025-02336 | 2025-03-31 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE DEVICE OPERATED AS EXPECTED AND THE DEATH WAS NOT CONSIDERED DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21865263&pc=NPT | 2025587-2025-02810 | 2024-12-21 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-18 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY THREE AND A HALF MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN.Event Description: IT WAS REPORTED THAT APPROXIMATELY THREE AND A HALF MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE; SERIAL #; FULL UDI # H4. DEVICE MFG DATE MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21865339&pc=DSQ | 2916596-2025-02308 | 2024-12-20 05:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). THE PATIENT EXPERIENCED HEART, KIDNEY, AND LIVER FAILURE. THE ONSET DATE OF MSOF WAS (B)(6) 2024. THE ORGAN FAILURE DID NOT EXIST PRE - LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. LVAD THERAPY DID NOT WORSEN, CAUSE OR CONTRIBUTE TO THE ORGAN FAILURE. THE PATIENT WAS HOME ON HOSPICE AFTER DISCUSSION THAT THE PATIENT WAS NOT DISCHARGEABLE. THE PATIENT SWITCHED TO HOSPICE WHILE THEY STILL HAD THE ABILITY TO MAKE DECISIONS ON THEIR OWN. THE VENTRICULAR ASSIST DEVICE (VAD) WAS OPERATING AS INTENDED AT THE TIME OF DEATH. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. Manufacturer Narrative: B7 - OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS. H6 HEALTH EFFECT - CLINICAL CODE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21865454&pc=DSQ | 2916596-2025-02311 | 2025-03-31 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT APPEARED TO HAVE A STROKE LEADING TO BRAIN DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21865636&pc=DSQ | 2916596-2025-02341 | 2025-03-28 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21865717&pc=NPT | 2025587-2025-02812 | 2025-03-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-18 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Pneumonia; Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY SIX YEARS AND ONE MONTH FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE INSIDE A PRE-EXISTING MEDTRONIC SURGICAL AORTIC VALVE, THE VALVE FAILED. THE MECHANISM OF FAILURE WAS NOT PROVIDED. SUBSEQUENTLY, IT WAS REPORTED THAT THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS DUE TO PNEUMONIA. PER THE PHYSICIAN, THE VALVE COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH. Manufacturer Narrative: UPDATED DATA: A4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2. DATE OF DEATH UNKNOWN. YEAR CONFIRMED VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866042&pc=MKJ | 1220908-2025-01416 | 2025-04-08 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-04-18 04:00 | MKJ | R SERIES | Inappropriate/Inadequate Shock/Stimulation | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A FEMALE PATIENT (AGE UNKNOWN) THE DEVICE FAILED TO DISCHARGE THE CORRECT ENERGY LEVEL DURING FOUR RESUSCITATION ATTEMPTS. THE COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CANADA EVALUATED THE DEVICE AND THE DEVICE PERFORMED TO SPECIFICATION. DEVICE LOGS CONFIRMED FOUR 200J SHOCKS WERE DELIVERED; THE FIRST DELIVERED 172.3J, NOT 72.3J AS MISTAKENLY REPORTED. ALL ENERGY DELIVERIES WERE WITHIN SPECIFICATION. THE DEVICE PASSED FUNCTIONAL TESTING WITH NO FAULT FOUND. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: JUSTIFICATION FOR NO UDI, THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866086&pc=DSQ | 3007042319-2025-00787 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-18 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE OUTCOMES OF PERICARDIAL CLOSURE USING AN EXPANDED POLYTETRAFLUOROETHYLENE (EPTFE) PATCH IN LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. PATIENTS WERE STUDIED ON OUTCOMES EITHER WITH THE USE OF A PATCH OR WITHOUT. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO DEVELOPED MEDIASTINITIS WITH PATHOGENS OF CANDIDA ALBICANS, CORYNEBACTERIUM SPECIES, ENTEROCOCCUS FAECIUM, PSEUDOMONAS AERUGINOSA, STAPHYLOCOCCUS AUREUS, STAPHYLOCOCCUS EPIDERMIDIS, STAPHYLOCOCCUS HAEMOLYTICUS, STREPTOCOCCUS ORALIS, AND SERRATIA MARCESCENS. OTHER PATIENTS EXPERIENCED DRIVELINE INFECTIONS, INFECTIONS WITH POSITIVE BLOOD CULTURES, STROKE, BLEEDING EVENTS, PUMP THROMBOSIS, AND RIGHT VENTRICULAR FAILURE WHICH REQUIRED A TEMPORARY RIGHT VAD. SOME PATIENTS REQUIRED A RETURN TO THE OPERATING ROOM WITHIN THIRTY DAYS AND UNDERWENT DELAYED STERNAL CLOSURE, REDO STERNOTOMY, AND SUBXIPHOID PERICARDIOTOMY FOR PERICARDIAL EFFUSION. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/54 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PERICARDIAL CLOSURE WITH EXPANDED POLYTETRAFLUOROETHYLENE PATCH IN LEFT VENTRICULAR ASSIST DEVICE SURGERY. ASAIO JOURNAL. 2024; 70:371¿376. DOI: 10.1097/MAT.0000000000002126. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866110&pc=NKE | 9614453-2025-01538 | 2022-11-05 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-18 04:00 | NKE | PERCEPTA QUAD CRT-P MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED UNEXPECTEDLY AT HOME SURROUNDED BY FAMILY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866116&pc=DTB | 2649622-2025-10596 | 2022-11-05 04:00 | Death | MPRI | 2025-04-18 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED UNEXPECTEDLY AT HOME SURROUNDED BY FAMILY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866126&pc=DTB | 2649622-2025-10597 | 2022-11-05 04:00 | Death | MPRI | 2025-04-18 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED UNEXPECTEDLY AT HOME SURROUNDED BY FAMILY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866127&pc=OJX | 2649622-2025-10598 | 2022-11-05 04:00 | Death | MPRI | 2025-04-18 04:00 | OJX | ATTAIN STABILITY QUAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED UNEXPECTEDLY AT HOME SURROUNDED BY FAMILY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866240&pc=DSQ | 3007042319-2025-00789 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-18 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Sepsis | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TEMPORAL TRENDS OF ROUTINELY OBTAINED BLOOD BIOMARKERS TO PREDICT BLOODSTREAM INFECTIONS IN LEFT VENTRICULAR ASSIST DEVICE (VAD) RECIPIENTS. THE AUTHORS DESCRIBED PATIENTS WHO EXPERIENCED BLOODSTREAM INFECTIONS WITH POSITIVE BLOOD CULTURES. ALBUMIN AND HEMOGLOBIN WERE SIGNIFICANT PREDICTORS FOR THE DEVELOPMENT OF BLOODSTREAM INFECTIONS DURING INITIAL HOSPITALIZATION. THERE WERE PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/60 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE UTILITY OF TEMPORAL TRENDS OF BLOOD BIOMARKERS AS PREDICTORS FOR BLOODSTREAM INFECTIONS IN LEFT VENTRICULAR ASSIST DEVICE RECIPIENTS. ARTIFICIAL ORGANS. 2024; 48:1513¿1524. DOI: 10.1111/AOR.14839 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866411&pc=DSQ | 2916596-2025-02275 | 2025-03-10 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE SEPTIC SHOCK. DURING IMPLANT DAY THE PATIENT WAS TOO MUCH IN SHOCK, SEPTIC SHOCK. THE PATIENT WAS SEDATED AND INTUBATED AND UNDERWENT HIGH DOSES OF INOTROPES. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: D5 - OPERATOR OF DEVICE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866675&pc=MCX | 2124215-2025-23387 | 2019-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-18 04:00 | MCX | ROTAWIRE AND WIRECLIP TORQUER | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS ON ROTAWIRE AND ROTAWIRE DRIVE USING THE PINC AI HEALTHCARE DATABASE FROM PREMIER. PATIENTS ARE IDENTIFIED THROUGH USE OF ROTAWIRE AND ROTAWIRE DRIVE AS IDENTIFIED IN CHARGE MASTER BILLING. ADVERSE EVENTS WERE IDENTIFIED THROUGH THE RESPECTIVE CPT/ICD-10 CODING. THE PROCEDURES WERE PERFORMED FROM (B)(6) 2019 TO (B)(6) 2023. THESE ANALYSES ARE BEING PERFORMED AS A SPECIFIC ACTIVITY TO ASSESS SAFETY AND PERFORMANCE RATES ASSOCIATED WITH THE SUBJECT DEVICE. ROTAWIRE AND ROTAWIRE DRIVE OUTCOMES WERE ASSESSED AGAINST VIPERWIRE AND VIPERWIRE ADVANCE. RATES OF THE ADVERSE EVENTS INCLUDING DEATH, VESSEL TRAUMA, MI, STROKE, ABRUPT CLOSURE, BLEEDING, EMBOLISM, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, ARRYTHMIA, AND ACUTE RENAL FAILURE ARE REPORTED BELOW. A TOTAL OF 1334 PATIENTS WHO UNDERWENT A PROCEDURE USING A ROTAWIRE. THE FOLLOWING IS A SUMMARY OF THOSE RESULTS OVER 5 YEARS (2019 -2023): MORTALITY RATES ROTA CASES: 35 OUT OF 1334 PATIENTS RESULTING IN A RATE OF 2.62%, COMPARED TO THE COMPETITOR RATE OF 0.98%-1.41%. MORTALITY RATES FOR ROTAWIRE CASES FALL WITHIN THE RANGE OF OR ARE LOWER THAN THOSE OF THE COMPETITORS. RATES FOR MORTALITY ARE THEREFORE COMPARABLE TO OR ARE LOWER THAN THOSE OF THE COMPETITIVE PRODUCTS AND ARE THEREFORE ACCEPTABLE. Manufacturer Narrative: B2 DATE OF DEATH AND B3 DATE OF EVENT: DATA WAS PROVIDED FOR EVENTS RECORDED (B)(6) 2019-(B)(6) 2023. (B)(6) 2019 SELECTED AS THE EARLIEST POSSIBLE DATE OF EVENT AND DATE OF DEATH. DATA OBTAINED FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AND NO FURTHER INFORMATION IS AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866681&pc=MCX | 2124215-2025-23390 | 2019-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-18 04:00 | MCX | ROTAWIRE DRIVE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS ON ROTAWIRE AND ROTAWIRE DRIVE USING THE PINC AI HEALTHCARE DATABASE FROM PREMIER. PATIENTS ARE IDENTIFIED THROUGH USE OF ROTAWIRE AND ROTAWIRE DRIVE AS IDENTIFIED IN CHARGE MASTER BILLING. ADVERSE EVENTS WERE IDENTIFIED THROUGH THE RESPECTIVE CPT/ICD-10 CODING. THE PROCEDURES WERE PERFORMED FROM JANUARY 2019 TO DECEMBER 2023. THESE ANALYSES ARE BEING PERFORMED AS A SPECIFIC ACTIVITY TO ASSESS SAFETY AND PERFORMANCE RATES ASSOCIATED WITH THE SUBJECT DEVICE. ROTAWIRE AND ROTAWIRE DRIVE OUTCOMES WERE ASSESSED AGAINST VIPERWIRE AND VIPERWIRE ADVANCE. RATES OF THE ADVERSE EVENTS INCLUDING DEATH, VESSEL TRAUMA, MI, STROKE, ABRUPT CLOSURE, BLEEDING, EMBOLISM, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, ARRYTHMIA, AND ACUTE RENAL FAILURE ARE REPORTED BELOW. A TOTAL OF 1214 UNDERWENT A PROCEDURE USING A ROTAWIRE DRIVE. THE FOLLOWING IS A SUMMARY OF THOSE RESULTS OVER 5 YEARS (2019 -2023): MORTALITY RATES ROTA CASES: 21 OUT OF 1214 PATIENTS RESULTING IN A RATE OF 1.73%, COMPARED TO THE COMPETITOR RATE OF 0.98%-1.41%. MORTALITY RATES FOR ROTAWIRE DRIVE CASES FALL WITHIN THE RANGE OF OR ARE LOWER THAN THOSE OF THE COMPETITORS. RATES FOR MORTALITY ARE THEREFORE COMPARABLE TO OR ARE LOWER THAN THOSE OF THE COMPETITIVE PRODUCTS AND ARE THEREFORE ACCEPTABLE. Manufacturer Narrative: B2 DATE OF DEATH AND B3 DATE OF EVENT: DATA WAS PROVIDED FOR EVENTS RECORDED (B)(6) 2019-31DEC2023. (B)(6) 2019 SELECTED AS THE EARLIEST POSSIBLE DATE OF EVENT AND DATE OF DEATH. DATA OBTAINED FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AND NO FURTHER INFORMATION IS AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866713&pc=QYU | 3012977166-2025-00001 | 2025-03-10 04:00 | Death | OWLET BABY CARE, INC. | 2025-04-18 04:00 | QYU | OWLET BABY CARE | Adverse Event Without Identified Device or Use Problem | Low Oxygen Saturation | DEN220091 |  | 1 | Event Description: OWLET WAS CONTACTED FOR A REQUEST FOR DATA REGARDING A DESATURATION NOTIFICATION. OWLET WAS INFORMED THAT THE INFANT HAD PASSED WHILE WEARING THE OWLET DREAM SOCK. Manufacturer Narrative: NO ALLEGATIONS OF DEFICIENCY WERE MADE BY THE USER. HOWEVER, OWLET CONDUCTED A THOROUGH INVESTIGATION AND REVIEWED THE HISTORICAL DATA ANALYSIS AND BACKEND DATA. THE TIMELINE AND CORRESPONDING INTERACTION WITH THE DEVICE AND DEVICE FUNCTION ALIGN WITH THE CUSTOMER NARRATIVE AND INDICATE THAT THE DEVICE FUNCTIONED AS INTENDED IN ALERTING THE CUSTOMER TO A DESATURATION EVENT. NO DEVICE DEFICIENCY OR MALFUNCTION WAS IDENTIFIED DURING THE INVESTIGATION. THEREFORE, THE CONCLUSION OF THE INVESTIGATION IS THAT THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21866911&pc=CAW | 2518422-2025-104471 | 2025-03-19 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | CAW | SIMPLYGO MINI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K111885 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT PASSED AWAY, AND HIS WIFE REACHED OUT TO INQUIRE ABOUT HOW TO DISPOSE OF THE SIMPLYGO MINI DEVICE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO DEVICE SERIAL NUMBER WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867303&pc=NIQ | 2024168-2025-04049 | 2025-03-27 04:00 | Death | ABBOTT VASCULAR INC. | 2025-04-18 04:00 | NIQ | XIENCE SKYPOINT¿ | Improper or Incorrect Procedure or Method | Cardiac Arrest; Perforation of Vessels | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A SEVERELY CALCIFIED LESION IN THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. A 3.5 X 38 MM XIENCE SKYPOINT DRUG-ELUTING STENT (DES) WAS SUCCESSFULLY IMPLANTED AT THE LESION; HOWEVER, DURING POST DILATATION, WITH DELIVERY CATHETER OF THE IMPLANTED XIENCE SKYPOINT, AT 17 ATMS [ABOVE RATED BURST PRESSURE (RBP)] A PERFORATION WAS NOTED. REMEDIAL ACTIONS WERE ATTEMPTED; HOWEVER, THE PATIENT WAS NOTED TO EXPIRE. PER THE PHYSICIAN'S OPINION, THE DEATH WAS NOT DUE TO THE USE OF THE XIENCE SKYPOINT NOR POST-DILATATION WITH THE DELIVERY CATHETER AND WAS DEEMED TO BE CAUSED BY USER ERROR. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECTS OF DEATH, PERFORATION AND CARDIAC FAILURE ARE LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (IFU) AS KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. IT SHOULD BE NOTED; PER THE PHYSICIANS OPINION THE DEATH WAS NOT DUE TO THE USE OF THE XIENCE DES NOR POST-DILATATION WITH THE DELIVERY CATHETER. THE TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. H6 MEDICAL DEVICE PROBLEM CODE: 2017 - ABOVE RBP.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867345&pc=NIK | 2124215-2025-24813 | 2025-04-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-18 04:00 | NIK | MOMENTUM CRT-D | Incorrect, Inadequate or Imprecise Result or Readings; Defibrillation/Stimulation Problem | Insufficient Information | P010012/S436 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) AFTER EXPERIENCING MULTIPLE EPISODES OF A VENTRICULAR ARRHYTHMIA ON (B)(6) 2025. THE PATIENT WAS REPORTED TO HAVE REQUIRED EXTERNAL DEFIBRILLATION IN ORDER TO RESTORE SINUS RHYTHM. INTERROGATION OF THE CRT-D REVEALED IT WAS IN SAFETY MODE. WHILE IN THE ICU, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY THE FOLLOWING DAY ON (B)(6) 2025 DUE TO COMPLICATIONS FROM THE VENTRICULAR EPISODES. THE CRT-D WAS EXPLANTED AND REMOVED FROM THE BODY AND WILL BE RETURNED BACK TO BOSTON SCIENTIFIC FOR ANALYSIS. THE PHYSICIAN WAS WANTING TO UNDERSTAND WHETHER THE CRT-D HAD DELIVERED ANY SHOCKS OR NOT, AND WHY (IF THERAPY WAS DELIVERED), SUCH SHOCKS WERE NOT EFFECTIVE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867404&pc=DSQ | 2916596-2025-02428 | 2025-03-30 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PASSED AWAY DUE TO CARDIOPULMONARY ARREST. WHETHER THE OUTCOME WAS DEVICE RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867424&pc=DSQ | 2916596-2025-02198 | 2025-03-19 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Disseminated Intravascular Coagulation (DIC); Hypovolemic Shock; Cardiogenic Shock; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT HEART FAILURE AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT RIGHT HEART FAILURE EXISTED BEFORE THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. THE DEVICE RELATED ISSUES DID NOT CAUSE/ CONTRIBUTE TO THE RIGHT HEART FAILURE. THE PATIENT EXHIBITED CARDIOGENIC AND HYPOVOLEMIC SHOCK. THE PATIENT WAS TREATED WITH FULL RIGHT HEART FAILURE MEDICATIONS. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867439&pc=MKJ | 1220908-2025-01476 | 2025-04-08 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-04-18 04:00 | MKJ | X SERIES | Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO MONITOR A FEMALE PATIENT (AGE UNKNOWN), THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. THE CREW SUBSEQUENTLY CHANGED THE PADS FOR A NEW SET AND WERE ABLE TO OBTAIN A RHYTHM ON THE SCREEN WITH THE REPLACEMENT PADS. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. HOWEVER, THE DEVICE LOGS WERE PROVIDED FOR REVIEW. REVIEW OF THE DEVICE LOGS SHOW THAT THE USER RECEIVED A PADS ON MESSAGE AT 2:45:23 WITH A CLEAR ECG SIGNAL. AT APPROXIMATELY 3:01:00, THE SIGNAL BECAME NOISY AND UNSTABLE. BY 3:04:15, REPEATED PADS OFF MESSAGES BEGAN, AND THE ECG SIGNAL WAS LOST FOR APPROXIMATELY 4 MINUTES. AT 3:10:15, A PADS ON MESSAGE AND CLEAR ECG SIGNAL RETURNED, LIKELY AFTER THE PADS WERE REPLACED. THE ISSUE APPEARS TO BE DUE TO POOR ELECTRODE TO PATIENT CONTACT (POOR COUPLING). CONTRIBUTING FACTORS COULD BE ELECTRODE PAD PLACEMENT, SKIN PREOPERATION, OR ASSOCIATED WITH PATIENT'S CLINICAL CONDITION (EXAMPLE: DIAPHORETIC). ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867551&pc=LWR | 2135147-2025-02083 | 2020-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-04-18 04:00 | LWR | TRIFECTA¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100029 |  | 1 | Event Description: THE ARTICLE, "LONGITUDINAL OUTCOMES FOLLOWING INTERNATIONAL MULTICENTER EXPERIENCE WITH ROBOTIC AORTIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY TO REPORT THE EARLY LONGITUDINAL OUTCOMES OF ROBOTIC-ASSISTED AORTIC VALVE REPLACEMENT (RAVR) AND HYPOTHESIZED THAT RAVR WOULD PROVIDE A SAFE AND EFFECTIVE ALTERNATIVE APPROACH TO TRADITIONAL SAVR THAT MAY BE SAFELY ADOPTED IN A GLOBAL MULTI-CENTER EXPERIENCE THROUGH STANDARDIZATION OF TECHNIQUE AND OUTCOME REPORTING. DEVICES INCLUDED IN THE STUDY WERE NOT SPECIFIED WITH EXCEPTION TO ONE CASE OF A TRIFECTA VALVE. THE ARTICLE CONCLUDED THAT RAVR IS SAFE AND EFFECTIVE, PROVIDING THE REPRODUCIBLE BENEFITS OF SURGICAL AVR WHILE AFFORDING A LESS INVASIVE APPROACH THAT PERMITS THE OPPORTUNITY FOR CONCOMITANT PROCEDURES. FOR LOW AND INTERMEDIATE RISK PATIENTS WITH AORTIC VALVE DISEASE, RAVR IS A POTENTIAL REPRODUCIBLE ALTERNATIVE FOR PATIENTS AND HEART TEAMS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS VINAY BADHWAR, DEPARTMENT OF CARDIOVASCULAR AND THORACIC SURGERY, WEST VIRGINIA UNIVERSITY, MORGANTOWN, WV, USA, WITH CORRESPONDING EMAIL: VINAY.BADHWAR@WVUMEDICINE.ORG]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2020 TO JULY 2024. A TOTAL OF 300 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT SPECIFIED HOW MANY RECEIVED AN ABBOTT DEVICE WITH EXCEPTION TO ONE PATIENT CASE. THE AVERAGE AGE WAS 67 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, HYPERTENSION, DIABETES MELLITUS, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, RENAL DISEASE, CEREBROVASCULAR ACCIDENT, PRIOR PERMANENT PACEMAKER, HEART FAILURE, REGURGITATION (AORTIC, MITRAL, AND/OR TRICUSPID), AORTIC STENOSIS, UNICUSPID/BICUSPID AORTIC VALVE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRIFECTA VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, HYPERTENSION, DIABETES MELLITUS, PERIPHERAL ARTERY DISEASE, CORONARY ARTERY DISEASE, RENAL DISEASE, CEREBROVASCULAR ACCIDENT, PRIOR PERMANENT PACEMAKER, HEART FAILURE, REGURGITATION (AORTIC, MITRAL, AND/OR TRICUSPID), AORTIC STENOSIS, UNICUSPID/BICUSPID AORTIC VALVE. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (REOPERATION, PACEMAKER), HOSPITALIZATION, CARDIOGENIC SHOCK, UNEXPECTED MEDICAL INTERVENTION (PROLONGED VENTILATION, DIALYSIS), RENAL FAILURE, STROKE, AORTIC STENOSIS, PARAVALVULAR LEAK. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: LONGITUDINAL OUTCOMES FOLLOWING INTERNATIONAL MULTICENTER EXPERIENCE WITH ROBOTIC AORTIC VALVE REPLACEMENT. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: LONGITUDINAL OUTCOMES FOLLOWING INTERNATIONAL MULTICENTER EXPERIENCE WITH ROBOTIC AORTIC VALVE REPLACEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867590&pc=MNS | 2518422-2025-104462 | 2025-04-16 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI. Manufacturer Narrative: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867736&pc=FRN | 9615050-2025-00187 | 2025-03-19 04:00 | Death | ICU MEDICAL COSTA RICA LTD. | 2025-04-18 04:00 | FRN | PLUM 360¿ INFUSER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K161469 |  | 1 | Event Description: IT WAS REPORTED THAT A PLUM 360 WAS INVOLVED IN A PATIENT ADVERSE EVENT. IT WAS STATED IN THE REPORT, "TWO PUMPS THAT WERE USED DURING A PATIENT CODING. THE STAFF SENT ME ¿THE FLUIDS WERE RUNNING WIDE OPEN AND THE ERROR CONTINUED SO THE NURSE SWAPPED OUT THE PUMP TO A DIFFERENT ONE AND THE SAME ERROR OCCURRED WHEN ATTEMPTING TO RUN WIDE OPEN. IT DID NOT OCCUR WHEN LEVO WAS RUNNING A CONSISTENT RATE.¿ IT ALSO STATED THAT THEY WERE UNABLE TO FIGURE OUT WHICH PUMP WAS FIRST OR SECOND. THE OTHER PUMP IS UNABLE TO SEE THE LOGS AFTER ON (B)(6) 2024 ON THE PUMP ITSELF AND ALSO ON MEDNET. THE PATIENT PASSED AWAY, BUT THEY ARE UNABLE TO TELL IF THESE PUMPS ARE A DIRECT CAUSE OF THAT. THE SECOND SERIAL NUMBER IS (B)(6). ADDITIONAL INFORMATION OBTAINED STATED THE FOLLOWING: "THE CODE RECORD IS STILL NOT SCANNED IN. WHAT I CAN GATHER FROM THE PROVIDER NOTES AND NURSING NOTES IS THAT OUR MONITOR TECH CALLED THE NURSE TO TELL HIM THAT THE PATIENT WAS SHOWING ASYSTOLE ON THE MONITOR. THE NURSE WENT TO THE ROOM AND FOUND THE PATIENT ON THE FLOOR OUTSIDE THE BATHROOM; THE PATIENT WAS PULSELESS AND NOT BREATHING. CPR WAS STARTED AND A CODE BLUE WAS CALLED. THE PATIENT WAS INTUBATED BY OUR INTENSIVIST. THE PATIENT WAS GIVEN 8 ROUNDS OF EPINEPHRINE IV. HE DID AT ONE POINT APPEAR TO HAVE A WIDE COMPLEX CARDIAC RHYTHM ON THE MONITOR WITH A RATE OF ABOUT 30 BUT HAD A BARELY PALPABLE PULSE AND NO BLOOD PRESSURE. ATROPINE WAS GIVEN IV BUT DID NOT HAVE AN EFFECT ON THE RHYTHM. THE PATIENT ONCE AGAIN WENT PULSELESS, AND CPR WAS CONTINUED. THE SAME WIDE COMPLEX RHYTHM DID RECUR BRIEFLY WITH A PULSE, BUT THE PULSE WAS THEN LOST AGAIN. CPR WAS STOPPED AFTER 31 MINUTES, AND THE PATIENT WAS PRONOUNCED DEAD. HAVING IV FLUIDS RUN AT 999CC/HR. ON THE IV PUMP HELPS TO DELIVER LIFE-SAVING MEDICATIONS SUCH AS EPINEPHRINE AND ATROPINE. CPR HELPS TO CIRCULATE THE MEDICATIONS SO THEY CAN HAVE AN EFFECT ON THE HEART. IT REALLY IS CRITICAL TO HAVE THE PUMPS FUNCTIONING APPROPRIATELY DURING THE CODE BLUE EVENT IN ORDER TO DELIVER THE MEDICATION. I'M STILL TRYING TO FIGURE OUT IF WE TOOK THE IV FLUIDS OFF THE PUMP AND TRIED TO JUST RUN THEM VIA GRAVITY." THE SECOND SERIAL NUMBER IS (B)(6). NO AUTOPSY WAS DONE AS THE CORONER DECLINED AN AUTOPSY. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR EVALUATION; HOWEVER, TESTING HAS NOT YET BEEN COMPLETED. Manufacturer Narrative: TESTS: SELF-TEST AND VISUAL INSPECTION. SELF-TEST PASSED. A VISUAL INSPECTION WAS PERFORMED AND PASSED. THE CUSTOMER COMPLAINT WASN'T CONFIRMED THAT THE DEVICE HAD AN ISSUE WITH DISTAL OCCLUSION OR THE FUNCTION OF THE PUMP. THE PROBABLE CAUSE WASN'T FOUND, NO ISSUES. THE DEVICE PASSED PERFORMANCE VERIFICATION TESTING (PVT) TESTING AND FOUND NO ISSUES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867752&pc=BZD | 2518422-2025-104480 | 2025-04-16 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867871&pc=NPT | 2015691-2025-03060 | 2025-03-25 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-18 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Degraded | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST, APPROXIMATELY 5 YEARS AND 7 MONTHS AFTER A TRANSFEMORAL TAVR PROCEDURE WITH A 29 MM SAPIEN 3 VALVE, THE VALVE FAILED. PER MEDICAL RECORDS RECEIVED FOR THE PATIENT, THE PATIENT PRESENTED WITH FATIGUE, SHORTNESS OF BREATH AND EDEMA. AN ECHOCARDIOGRAM PERFORMED REPORTED SEVERE PROSTHETIC AORTIC VALVE STENOSIS WITH PEAK VELOCITY = 4.6 M/S AND MEAN GRADIENT = 51 MMHG, LEAFLETS APPEARED MILDLY THICKENED ALTHOUGH THEY WERE NOT WELL SEEN. THERE WAS ALSO SEVERE MITRAL VALVE REGURGITATION. THE PATIENT WAS IN HOSPICE CARE DUE TO HEART FAILURE WITH PRESERVED EJECTION FRACTION, RENAL FAILURE AND AUTOIMMUNE HEPATITIS, AND PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, ALL SAPIEN THV VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. THE EXACT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED, BUT INVESTIGATION RESULTS SUGGEST THAT FACTORS/MECHANISMS MENTIONED ABOVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21867990&pc=DSQ | 2916596-2025-02375 | 2025-03-25 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO CARDIAC ARREST. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21868070&pc=FRN | 9615050-2025-00188 | 2025-03-19 04:00 | Death | ICU MEDICAL COSTA RICA LTD. | 2025-04-18 04:00 | FRN | PLUM 360¿ INFUSER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K161469 |  | 1 | Event Description: IT WAS REPORTED THAT A PLUM 360 WAS INVOLVED IN A PATIENT EVENT. IT WAS STATED IN THE REPORT, "TWO PUMPS THAT WERE USED DURING A PATIENT CODING. THE STAFF SENT ME THE FLUIDS WERE RUNNING WIDE OPEN AND THE ERROR CONTINUED SO THE NURSE SWAPPED OUT THE PUMP TO A DIFFERENT ONE AND THE SAME ERROR OCCURRED WHEN ATTEMPTING TO RUN WIDE OPEN. IT DID NOT OCCUR WHEN LEVO WAS RUNNING A CONSISTENT RATE.¿ IT ALSO STATED THAT THEY ARE UNABLE TO FIGURE OUT WHICH PUMP WAS FIRST OR SECOND. ONE OF THE PUMPS WITH THE SERIAL NUMBER (B)(6) HAD A FEW DIFFERENT ERROR CODES HAPPENING AROUND THE TIME OF THE CODE. IT HAS A DISTAL OCCLUSION ALARM AND THEN AN ¿E346 DISTAL PRESS SENSOR 2¿. THE PATIENT PASSED AWAY, BUT THEY ARE UNABLE TO TELL IF THESE PUMPS ARE A DIRECT CAUSE. "THE CODE RECORD IS STILL NOT SCANNED IN. WHAT I CAN GATHER FROM THE PROVIDER NOTES AND NURSING NOTES IS THAT OUR MONITOR TECH CALLED THE NURSE TO TELL HIM THAT THE PATIENT WAS SHOWING ASYSTOLE ON THE MONITOR. THE NURSE WENT TO THE ROOM AND FOUND THE PATIENT ON THE FLOOR OUTSIDE THE BATHROOM; THE PATIENT WAS PULSELESS AND NOT BREATHING. CPR WAS STARTED AND A CODE BLUE WAS CALLED. THE PATIENT WAS INTUBATED BY OUR INTENSIVIST. THE PATIENT WAS GIVEN 8 ROUNDS OF EPINEPHRINE IV. HE DID AT ONE POINT APPEAR TO HAVE A WIDE COMPLEX CARDIAC RHYTHM ON THE MONITOR WITH A RATE OF ABOUT 30 BUT HAD A BARELY PALPABLE PULSE AND NO BLOOD PRESSURE. ATROPINE WAS GIVEN IV BUT DID NOT HAVE AN EFFECT ON THE RHYTHM. THE PATIENT ONCE AGAIN WENT PULSELESS, AND CPR WAS CONTINUED. THE SAME WIDE COMPLEX RHYTHM DID RECUR BRIEFLY WITH A PULSE, BUT THE PULSE WAS THEN LOST AGAIN. CPR WAS STOPPED AFTER 31 MINUTES, AND THE PATIENT WAS PRONOUNCED DEAD. HAVING IV FLUIDS RUN AT 999CC/HR. ON THE IV PUMP HELPS TO DELIVER LIFE-SAVING MEDICATIONS SUCH AS EPINEPHRINE AND ATROPINE. CPR HELPS TO CIRCULATE THE MEDICATIONS SO THEY CAN HAVE AN EFFECT ON THE HEART. IT REALLY IS CRITICAL TO HAVE THE PUMPS FUNCTIONING APPROPRIATELY DURING THE CODE BLUE EVENT IN ORDER TO DELIVER THE MEDICATION. I¿M STILL TRYING TO FIGURE OUT IF WE TOOK THE IV FLUIDS OFF THE PUMP AND TRIED TO JUST RUN THEM VIA GRAVITY." Manufacturer Narrative: THE DEVICE WAS RETURNED FOR EVALUATION; HOWEVER, TESTING HAS NOT YET BEEN COMPLETED. Manufacturer Narrative: ENGINEERING DETERMINED THAT THE PROBABLE CAUSE OF THE REPORTED DISTAL PRESSURE SENSOR FAILURE WAS A MOMENTARY PRESSURE FLUCTUATION THAT OCCURRED WHILE THE DEVICE WAS INFUSING ¿ NOT DUE TO A HARDWARE MALFUNCTION. THE DISTAL OCCLUSION ALARM MAY HAVE BEEN TRIGGERED BY THE USE OF MICROBORE TUBING AT A FLOW RATE EXCEEDING 100 ML/HR, WHICH IS EXPLICITLY CAUTIONED AGAINST IN THE SYSTEM OPERATING MANUAL. WHILE ADDITIONAL ALARMS SUCH AS DISTAL AIR-IN-LINE AND ANOTHER E346 ERROR WERE OBSERVED DURING NORMAL-RATE TESTING, THESE COULD NOT BE REPRODUCED AND DID NOT INDICATE A PERSISTENT OR SYSTEMIC FAULT. OVERALL, THE DEVICE WAS FOUND TO BE OPERATING WITHIN SPECIFICATIONS, AND THE INVESTIGATION CONCLUDED THAT IT FUNCTIONED CORRECTLY. THE ALARM EVENTS WERE LINKED TO SETUP-RELATED CONDITIONS. ALTHOUGH THE PATIENT PASSED AWAY DURING THE EVENT, THE INVESTIGATION DID NOT ESTABLISH A CAUSAL LINK BETWEEN THE ALARMS AND THE OUTCOME. THEREFORE, THE DEVICE¿S CONTRIBUTION TO THE EVENT COULD NOT BE CONFIRMED, AND THIS CASE DOES NOT WARRANT ESCALATION TO CAPA
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21868326&pc=NPT | 2025587-2025-02842 | 2025-04-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-18 04:00 | NPT | EVOLUT FX DCS | Use of Device Problem; Difficult to Advance; Device Dislodged or Dislocated; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Perforation; Low Oxygen Saturation | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH A HEAVILY CALCIFIED AORTIC ARCH, NO PRE-DILATION WAS PERFORMED AND THE DELIVERY CATHETER SYSTEM (DCS) FLUSH PORT POSITIONED AT 1200 WAS UNABLE TO CROSS. THE 3, 9, 12, AND 6 O'CLOCK POSITIONS WERE ATTEMPTED AND WERE UNABLE TO CROSS AS WELL. THE ARCH WAS FINALLY ABLE TO BE CROSSED WITH NO NOSE CONE SEPARATION OBSERVED. PRESSURE QUICKLY DROPPED AS WELL AS DESATURATION. THE VALVE WAS 80% DEPLOYED AT 3 MM ON THE NON-CORONARY CUSP (NCC). ASYSTOLE WAS NOTED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. AN ECHOCARDIOGRAM REVEALED NO EFFUSION. THE PATIENT WAS PACED AS THERE WAS NO RHYTHM NOTED. CPR STOPPED AND AN AORTOGRAM WAS PERFORMED. THE VALVE MOVED AORTIC AND WAS RECAPTURED. AN AORTIC ARCH PERFORATION WAS SEEN AND WAS DEEMED INCOMPATIBLE WITH LIFE DUE TO THE POSITION OF THE TEAR AND PROLONGED HYPOTENSION. RESUSCITATION EFFORTS WERE CEASED AND THE PATIENT DIED. IT WAS REPORTED THAT THE PATIENT'S ANNULUS WAS 64.2 MM. PER THE PHYSICIAN, THE PATIENT'S HEAVILY CALCIFIED AORTA AND TIGHT BEND IN THE AORTA CONTRIBUTED TO THE EVENT.Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH A HEAVILY CALCIFIED AORTIC ARCH, NO PRE-DILATION WAS PERFORMED AND THE DELIVERY CATHETER SYSTEM (DCS) FLUSH PORT POSITIONED AT 1200 WAS UNABLE TO CROSS. THE 3, 9, 12, AND 6 O'CLOCK POSITIONS WERE ATTEMPTED AND WERE UNABLE TO CROSS AS WELL. THE ARCH WAS FINALLY ABLE TO BE CROSSED WITH NO NOSE CONE SEPARATION OBSERVED. PRESSURE QUICKLY DROPPED AS WELL AS DESATURATION. THE VALVE WAS 80% DEPLOYED AT 3 MM ON THE NON-CORONARY CUSP (NCC). ASYSTOLE WAS NOTED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. AN ECHOCARDIOGRAM REVEALED NO EFFUSION. THE PATIENT WAS PACED AS THERE WAS NO RHYTHM NOTED. CPR STOPPED AND AN AORTOGRAM WAS PERFORMED. THE VALVE MOVED AORTIC AND WAS RECAPTURED. AN AORTIC ARCH PERFORATION WAS SEEN AND WAS DEEMED INCOMPATIBLE WITH LIFE DUE TO THE POSITION OF THE TEAR AND PROLONGED HYPOTENSION. RESUSCITATION EFFORTS WERE CEASED AND THE PATIENT DIED. IT WAS REPORTED THAT THE PATIENT'S ANNULUS WAS 64.2 MM. PER THE PHYSICIAN, THE PATIENT'S HEAVILY CALCIFIED AORTA AND TIGHT BEND IN THE AORTA CONTRIBUTED TO THE EVENT. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN TO REPORT THE PATIENT'S ANATOMY WAS CALCIFIED AND INCLUDED A "HAIRPIN TURN" IN THE AORTA, THIS LED TO THE PERFORATION. ADDITIONALLY, THE PHYSICIAN INDICATED THERE WERE "NO CONCERNS REGARDING" THE VALVE OR THE DCS AND THE PATIENT'S ANATOMY WAS THE SOLE CONTRIBUTOR TO THE PATIENT'S DEATH. ADDITIONAL INFORMATION WAS PREVIOUSLY REPORTED, BUT NOT CAPTURED IN THE INITIAL NARRATIVE: THE DCS WAS INSERTED INTO THE PATIENT VIA THE RIGHT FEMORAL ARTERY AND NAVIGATED TO THE HEART. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. DEVICE CODE WAS ADDED. EVAL CODE CONCLUSION WERE UPDATED. EVAL CODE RESULT WAS UPDATED. CORRECTED DATA: B1. THE PRODUCT PROBLEM WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. B5. ADDITIONAL INFORMATION WAS PREVIOUSLY REPORTED, BUT NOT CAPTURED IN THE INITIAL NARRATIVE TO INDICATE WHERE VASCULAR ACCESS FOR DCS INSERTION WAS ACHIEVED. H6. PATIENT CODE E2203 WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. DEVICE CODE A051201 WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. ADDITIONAL CODES. APPROPRIATE ANNEX F CODES WERE INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS: UPON RECEIPT OF THE SUSPECT COMPLAINT DEVICE AT MEDTRONIC¿S QUALITY LABORATORY, VISUAL ANALYSIS SHOWED THE DELIVERY CATHETER SYSTEM (DCS) WAS RECEIVED WITHOUT A VALVE LOADED WITHIN THE CAPSULE. THE DEVICE WAS RECEIVED WITH THE CAPSULE PARTIALLY OPENED. THERE WAS A SLIGHT BEND TO THE CAPSULE. THERE WAS A SCRATCH TO THE PROXIMAL SECTION OF THE CAPSULE. DELAMINATION WAS OBSERVED OVER THE NITINOL REINFORCING FRAME AT THE DISTAL AND PROXIMAL ENDS OF THE CAPSULE. THERE WAS A BEND TO THE STABILITY SHAFT. THE HANDLE APPEARED INTACT. THE DEVICE WAS RETURNED WITH THE END CAP/SCREW GEAR SNAP FIT CONNECTED. THE HANDLE APPEARED TO RETRACT AND ADVANCE THE CAPSULE. THE TRIGGER MOVED TO FULLY ADVANCED AND RETRACTED POSITIONS AND LOCKED IN PLACE WHEN RELEASED. THE TIP-RETRIEVAL MECHANISM APPEARED INTACT. THE INNER MEMBER SHAFT AND SPINDLE HUB APPEARED INTACT WITH NO EVIDENCE OF DAMAGE. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012617231); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVOLUTFX-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE (B)(6) 2025 SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21868772&pc=DSQ | 2916596-2025-02208 | 2025-03-10 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Low Oxygen Saturation; Respiratory Failure; Aspiration Pneumonitis | P060040 |  | 1 | Event Description: THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) WITH ACUTE RESPIRATORY FAILURE, HYPOXEMIA, AND LACTIC ACIDOSIS. THE ASPIRATION PNEUMONIA INITIALLY RESPONDED TO IV ANTIBIOTICS. HOWEVER, THE PATIENT'S ORIENTATION AND RESPIRATORY STATUS CONTINUED TO DECLINE. THE FAMILY WAS CONSULTED AND THEY DECIDED TO DISCONTINUE CARE. THE PATIENT PASSED AWAY ON (B)(6) 2025 AFTER WITHDRAWAL OF CARE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, LOCAL INFECTION, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21868926&pc=MNS | 2518422-2025-104501 | 2025-04-16 04:00 | Death | RESPIRONICS, INC. | 2025-04-18 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: IN THE PREVIOUS REPORT MANUFACTURER REPORTED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. BASED ON INFORMATION AVAILABLE AT THE TIME OF THIS REVIEW, THE DEVICE WAS RETURNED FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE IT DETERMINE THAT, NO DEVICE ERROR FOUND AND THE DEVICE PASSED ALL TEST NO OTHER CONTAMINATION ALSO FOUND. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869018&pc=NIQ | 2124215-2025-23455 | 2022-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-18 04:00 | NIQ | SYNERGY? | Off-Label Use; Defective Device | Cardiac Arrest | P150003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A SYNERGY? 3.50 X 8 WAS ADVANCED FOR TREATMENT OF A PATENT DUCTUS ARTERIOSUS (PDA). HOWEVER, DURING THE PROCEDURE, THE DEVICE FAILED TO FUNCTION AS INTENDED WHICH LEAD TO CARDIAC ARREST AND THE PATIENT DIED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A SYNERGY? 3.50 X 8 WAS ADVANCED FOR TREATMENT OF A PATENT DUCTUS ARTERIOSUS (PDA). HOWEVER, DURING THE PROCEDURE, THE DEVICE FAILED TO FUNCTION AS INTENDED WHICH LED TO CARDIAC ARREST AND THE PATIENT DIED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A SYNERGY 3.50 X 8 WAS ADVANCED FOR TREATMENT OF A PATENT DUCTUS ARTERIOSUS (PDA). HOWEVER, DURING THE PROCEDURE, THE DEVICE FAILED TO FUNCTION AS INTENDED WHICH LED TO CARDIAC ARREST AND THE PATIENT DIED. IT WAS FURTHER REPORTED THAT TWO SYNERGY 3.50 X 8 MM STENTS WERE IMPLANTED. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. CORRECTED INFORMATION: A2: AGE AT TIME OF EVENT, DATE OF BIRTH AND H6: DEVICE CODES. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTED INFORMATION- A2: AGE AT TIME OF EVENT, DATE OF BIRTH AND H6: DEVICE CODES. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869144&pc=DQY | 3005168196-2025-00126 | 2025-03-21 04:00 | Death | PENUMBRA, INC. | 2025-04-18 04:00 | DQY | NEURON MAX 6F 088 LONG SHEATH | Material Fragmentation | Intracranial Hemorrhage; Brain Injury; Foreign Body Embolism | K111380 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE USING A NEURON MAX 6F 088 LONG SHEATH (NEURON MAX). WITH SOME DIFFICULTY, THE NEURON MAX WAS ADVANCED THROUGH A STENT AT THE ORIGIN OF THE LEFT COMMON CAROTID ARTERY. UPON REMOVAL OF THE NEURON MAX 6F SELECT INNER CATHETER, THE NEURON MAX DISTAL TIP WAS FOUND TO BE FRACTURED. CONTRAST INJECTION DEMONSTRATED MINIMAL FLOW INTO THE LEFT MIDDLE CEREBRAL ARTERY AS WELL AS THE LEFT ANTERIOR CEREBRAL ARTERY DUE TO THE FRACTURED NEURON MAX SEGMENT. ATTEMPTS TO RETRIEVE THE NEURON MAX FRACTURED SEGMENT WERE UNSUCCESSFUL FOLLOWED BY GROSS SUBARACHNOID HEMORRHAGE. SUBSEQUENT COIL OCCLUSION OF THE LEFT INTERNAL CAROTID ARTERY WAS PERFORMED TO CONTROL BLEEDING. THE PHYSICIAN ENDED THE PROCEDURE AT THIS POINT. POST-PROCEDURE, THE PATIENT REMAINED INTUBATED, NOT SEDATED, REQUIRING FOUR VASOPRESSORS. THE PATIENT HAD NO COUGH, GAG, OR CORNEAL REFLEX. PUPILS WERE 7MM AND NONREACTIVE BILATERALLY. DISCUSSION WITH THE SPOUSE/FAMILY GIVEN HEMODYNAMIC INSTABILITY AND UNLIKELY MEANINGFUL OUTCOME. THE FAMILY DECIDED TO TRANSITION TO COMFORT CARE, AND THE PATIENT EXPIRED ON (B)(6) 2025. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. WITHOUT THE RETURN OF THE DEVICE, THE ROOT CAUSE OF THE PROBLEM CANNOT BE DETERMINED. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869234&pc=DSQ | 2916596-2025-02447 | 2024-11-18 05:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNTREATABLE VENTRICULAR FIBRILLATION FOLLOWED BY RIGHT VENTRICULAR FAILURE (RVF). THE VENTRICULAR FIBRILLATION WAS AN ISSUE PRIOR TO LEFT VENTRICULAR ASSIST DEVICE (LVAD). NO DEVICE RELATED ISSUES WERE REPORTED AND NO DEVICE ISSUES LED TO THE OUTCOME. THE DEVICE OPERATED AS EXPECTED AND THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, RIGHT HEART FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 OF THE IFU (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869245&pc=DSQ | 2916596-2025-02491 | 2025-03-25 04:00 | Death | THORATEC CORPORATION | 2025-04-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED SEPSIS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. NO FURTHER INFORMATION WAS PROVIDED DUE TO HOSPITAL POLICY. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING SEPSIS AND DEATH. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, SEVERAL SECTIONS OF THE HEARTMATE 3 LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869479&pc=NVY | 2124215-2025-24867 | 2025-01-31 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-18 04:00 | NVY | ENDOTAK RELIANCE G | Defective Device | No Clinical Signs, Symptoms or Conditions | P910073/S077 |  | 1 | Event Description: IT WAS REPORTED VIA A FOOD AND DRUG ADMINISTRATION (FDA) MEDWATCH REPORT THAT THE PATIENT IMPLANTED WITH THIS RIGHT VENTRICULAR (RV) LEAD PASSED AWAY. IT WAS ALLEGED THAT THE LEAD WAS NOT FUNCTIONING PROPERLY. NO ADDITIONAL INFORMATION WAS ABLE TO BE OBTAINED AS THE REPORT DID NOT INCLUDE ANY CONTACT INFORMATION. ADDITIONALLY, NO LEAD SERIAL NUMBER OR PATIENT IDENTIFICATION WERE PROVIDED. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869481&pc=DTB | 2124215-2025-24869 | 2025-01-31 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-18 04:00 | DTB | FINELINE II STEROX | Defective Device | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA A FOOD AND DRUG ADMINISTRATION (FDA) MEDWATCH REPORT THAT THE PATIENT IMPLANTED WITH THIS RIGHT ATRIAL (RA) LEAD PASSED AWAY. IT WAS ALLEGED THAT THE LEAD WAS NOT FUNCTIONING PROPERLY. NO ADDITIONAL INFORMATION WAS ABLE TO BE OBTAINED AS THE REPORT DID NOT INCLUDE ANY CONTACT INFORMATION. ADDITIONALLY, NO LEAD SERIAL NUMBER OR PATIENT IDENTIFICATION WERE PROVIDED. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869622&pc=NAY | 2955842-2025-16249 | 2024-12-23 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-18 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED PYLORIC GASTRECTOMY (ROUX-EN-Y RECONSTRUCTION) PROCEDURE, THE PATIENT DIED. DURING THE PROCEDURE, WHILE CREATING THE Y-LEG ANASTOMOSIS, THE MESENTERY OF THE OTHER PORTION OF THE SMALL INTESTINE WAS CAUGHT AND ALSO SUTURED TOGETHER WITH THE Y-LEG. THIS CAUSED THE INTESTINE TO BECOME TWISTED. POST-OPERATIVELY, THE PATIENT EXPERIENCED AN INTESTINAL OBSTRUCTION DUE TO THE TWISTED INTESTINE. NO ADDITIONAL EVENT INFORMATION WAS PROVIDED. THE SURGEON STATED THIS EVENT WAS NOT CAUSED BY A DA VINCI PRODUCT. THE HOSPITAL SAFETY MANAGEMENT COMMITTEE DETERMINED THAT FUTURE CASES FOR THIS SURGERY AT THIS HOSPITAL WOULD REQUIRE THE ANASTOMOSES TO BE PERFORMED EXTRACORPOREALLY.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, LONG BIPOLAR GRASPER, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, LARGE SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND, TWO MEDIUM-LARGE CLIP APPLIERS, SMALL CLIP APPLIER, AND A SUREFORM 60 STAPLER INSTRUMENT. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A GASTRECTOMY PROCEDURE IN WHICH A SURGICAL ERROR OCCURRED IN CREATING A ROUX-EN-Y ANASTOMOSIS RESULTING IN AN OBSTRUCTION. HOW THIS TECHNICAL ERROR CONTRIBUTED TO THE DEATH IS UNKNOWN. NO ALLEGATIONS HAVE BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869635&pc=OZD | 1220648-2025-27680 | 2024-08-09 04:00 | Death | ABIOMED, INC. | 2025-04-18 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE FOLLOWING WAS REPORTED BY MEDWATCH: "IMPELLA PLACED AND DAD NEVER RECOVERED AND GOT EXTREMELY WORSE". NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869652&pc=OZD | 1220648-2025-27687 | 2025-03-26 04:00 | Death | ABIOMED, INC. | 2025-04-18 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemolysis; Hemorrhage/Bleeding; Renal Failure; Tachycardia; Cardiogenic Shock | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND APPROXIMATELY ONE DAY AFTER START OF THE SUPPORT, THE PATIENT EXPERIENCED TACHYCARDIC EVENTS FOR WHICH CARDIOVERSION WAS REQUIRED. THE PHYSICIAN ARRIVED AND REPOSITIONED THE PUMP TO OPTIMAL PLACEMENT AND DECREASED THE SUPPORT LEVEL TO ASSIST WITH MANAGEMENT OF THE TACHYCARDIA EVENTS. THE FOLLOWING DAY, THE IMPELLA WAS INADVERTENTLY PULLED BACK INTO THE AORTA WHEN THE PATIENT WAS BEING TURNED; IT WAS NOTED THE NURSE FORGOT TO CHECK THE TOUHY BORST. THE PHYSICIAN WAS UNABLE TO VISUALIZE THE ISSUE. THE FELT THE PATIENT WAS HEMODYNAMICALLY STABLE AND THE DECISION WAS MADE TO EXPLANT THE IMPELLA CP DEVICE. THE PATIENT SURVIVED THE EXPLANT AND WAS REPORTED TO BE IN STABLE CONDITION FOLLOWING THE EVENT.Event Description: ADDITIONAL NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED BLEEDING WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿PATIENT WAS INDICATED FOR RED BLOOD CELL TRANSFUSION WITH A HGB LESS THAN 9 OR A HCT OF LESS THAN 27 WITH SYMPTOMATIC ANEMIA. PATIENT'S LABS WERE DRAWN AND A HGB WAS FOUND TO BE 6.9 G/DL THE MORNING OF (B)(6) 2025. PATIENT WAS GIVEN 1 UNIT OR RBC'S AND BLOOD RE-DRAW WAS FOUND TO BE 8.3 G/DL, FURTHER RBC'S WERE NOT NEEDED AT THAT TIME.¿. THE PATIENT RECOVERED WITH NO SEQUELAE.Event Description: AN ADDITIONAL LOQI NOTIFICATION WAS RECEIVED AND INDICATED THAT THE PATIENT DEVELOPED ACUTE ANEMIA SECONDARY TO HEMOLYSIS BELIEVED TO HAVE A PROBABLE RELATIONSHIP TO IMPELLA CP SUPPORT. A BLOOD TRANSFUSION WAS PERFORMED TO TREAT THE CONDITION. IT WAS THEN NOTED THAT THE PATIENT SUFFERED FROM AN ACUTE KIDNEY INJURY WITH A PROBABLE RELATIONSHIP TO IMPELLA SUPPORT. THE CONDITION WAS BELIEVED TO HAVE STEMMED FROM CARDIOGENIC SHOCK, IMPELLA-RELATED HEMOLYSIS, AND CONTRAST NEPHROPATHY. THE PATIENT WAS STARTED ON A BUMETANIDE INFUSION AND CONTINUOUS RENAL REPLACEMENT THERAPY IN RESPONSE TO CONDITION. IT WAS REPORTED THAT THIS WAS ONGOING AT THE TIME OF THE PATIENT'S PASSING. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ B2 DEATH WAS INADVERTENTLY OMITTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27687-1. DATE OF DEATH IS UNKNOWN. B5 ADDITIONAL INFORMATION AND PATIENT OUTCOME WAS INADVERTENTLY OMITTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27687-1. H1 TYPE OF REPORT WAS ERRONEOUSLY ENTERED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27687-1. H6 HEALTH EFFECT - CLINICAL CODE 1886, 2041, 2262 AND HEALTH EFFECT - IMPACT CODES 4644 AND 1802 WERE INADVERTENTLY OMITTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27687-1. H10 INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27687-1. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND VF/VT/AF/AT HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION, AND A DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USE RELATED SINCE THE PUMP WAS PULLED BACK TO AORTA DURING TURN. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. B5 REVISED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED. THE INVESTIGATION VASCULAR DAMAGE - PERIPHERAL VESSEL HAS NOT BEEN COMPLETED. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. G2 REPORT SOURCE; STUDY WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27687. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21869661&pc=OZD | 1220648-2025-27690 | 2025-03-28 04:00 | Death | ABIOMED, INC. | 2025-04-18 04:00 | OZD | IMPELLA CP | Positioning Failure | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE IMPELLA CP WAS PLACED AND WAS TAKEN TO ICU, HOWEVER THE PATIENT WAS IN SUSTAINED VENTRICULAR TACHYCARDIA. DURING SUPPORT THERE WAS A CALL REGARDING ONGOING PLACEMENT SIGNAL NOT RELIABLE (PSNR) ALARM. THE NURSE REPORTED ACTIVE ALARM WITH AUDIO DISABLED. PLACEMENT SIGNAL WAVEFORMS REMAINED. ADDITIONALLY, THE THERE WAS SPORADIC WRONG POSITION ALARMS. MOTOR CURRENT (MC) WAS ALSO NARROW. THE PHYSICIAN WAS ADVISED TO CONFIRM POSITION. AND THEY ARE WORKING ON AN ART LINE. ULTIMATELY THE PATIENT WAS DECOMPENSATING. THE FAMILY DID NOT WANT CPR AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. AFTER A FEW HOURS ON SUPPORT, SPORADIC WRONG POSITION ALARMS OCCURRED. PREVAILING CARE TEAM THEORY IS PATIENT IS NON-PULSATILE. THE PATIENT BEGAN TO DECOMPENSATE IN ICU. PRODUCT WAS NOT RETURNED. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21852901&pc=BZD | 2518422-2025-104242 | 2022-05-14 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT DIED DUE TO ATRIAL FIBRILLATION. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DREAMSTATION AUTO BIPAP DEVICE AND THE HARM REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI. Manufacturer Narrative: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853091&pc=CAW | 2518422-2025-104245 | 2025-04-02 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | CAW | EVERFLO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K061261 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THE CAUSE OF DEATH IS NOT KNOWN. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853200&pc=NJK | 3007797756-2025-00015 | 2025-03-17 04:00 | Death | PULMONX CORPORATION | 2025-04-17 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P180002 |  | 1 | Event Description: THIS 81-YEAR-OLD MALE PATIENT SUFFERING FROM SEVERE COPD SINCE 2011 WITH COR PULMONALE DIAGNOSED IN (B)(6) 2024, WAS TREATED ON (B)(6) 2024 WITH 3 ZEPHYR VALVES PLACED IN THE RUL (RBI: 4.0 LP, RB2: 4.0, AND RB3: 4.0). THEN, THE PATIENT WAS DISCHARGED ON (B)(6) 2024 AND HAD NO PROBLEMS FOR 2 MONTHS. ON (B)(6) 2025, THE PATIENT EXPERIENCED ARTHRALGIA AND ANOREXIA WHILE BEING UNWELL SINCE A FEW DAYS. ON (B)(6) IN THE MORNING, HE WAS TRANSPORTED TO EMERGENCY CARE BECAUSE HE WAS FOUND UNABLE TO MOVE BY THE FAMILY. ON ARRIVAL THE PATIENT WAS RECEIVING ANTIBIOTIC (ERYTHROMYCIN) FOR PNEUMONIA, AND MEDICATION FOR COPD (ON SL OXYGEN) AND CARDIAC INSUFFICIENCY (FUROSEMIDE). HE WAS HYPOTENSIVE, WITH TYPE 2 RESPIRATORY FAILURE. AFTER TREATMENT WITH IV FLUIDS, A CHEST CT WAS PERFORMED AND SHOWED AN EXTENSIVE INFILTRATIVE SHADOW IN THE RLL. BLOOD TESTS SHOWED AN ELEVATED INFLAMMATORY RESPONSE. A COMMUNITY-ACQUIRED PNEUMONIA OF THE RLL WAS CONFIRMED (STREPTOCOCCUS PNEUMONIA). THUS, A SEPTIC SHOCK ASSOCIATED WITH THE ONGOING PNEUMONIA WAS CONSIDERED. THE PATIENT WENT INTO A CARDIOPULMONARY ARREST DURING IMAGING SESSION, AND - ALTHOUGH HIS OWN HEARTBEAT RESUMES AFTER ROUTINE CARDIOPULMONARY RESUSCITATION AND THE PATIENT RETURNED TO CONSCIOUSNESS - BLOOD PRESSURE REMAINED LOW AND CIRCULATION PARAMETERS VERY UNSTABLE. AFTER FAMILY CONSULTATION, THE DECISION WAS MADE NOT TO RESUSCITATE THE PATIENT AT THE TIME OF CARDIAC ARREST. DEATH WAS CONFIRMED AT THE EMERGENCY ROOM ((B)(6) 2025). IN SUMMARY, THIS SEVERE COPD PATIENT WITH CARDIAC COMPLICATION (COR PULMONALE), DIED 3 MONTHS AFTER BL VR (RUL) WITH ZEPHYR® VALVES FROM A SEPTIC SHOCK RELATED TO PNEUMONIA IN THE RLL. THE DEATH IS NOT RELATED TO THE PROCEDURE, WHICH WAS PERFORMED 3 MONTHS EARLIER, AND UNLIKELY RELATED TO THE VALVES WHICH WERE NOT IMPLANTED IN THE INFECTED LOBE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853257&pc=BZD | 2518422-2025-104246 | 2025-04-15 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853260&pc=MNS | 2518422-2025-104247 | 2025-04-15 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION WHERE IT WAS INSPECTED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS AND NO COSMETIC DAMAGE FOUND. THE DEVICE PASSED ALL TESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853656&pc=DTB | 2649622-2025-10430 | 2025-03-25 04:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | Cardiac Arrest; Ventricular Fibrillation; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING DURING STORED VENTRICULAR TACHYCARDIA (VT) EPISODES. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF), CODED AND PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853734&pc=DTB | 2649622-2025-10388 | 2025-02-26 05:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DEVELOPED AN INFECTION. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REMOVED AND THE PATIENT WAS BEING PACED BY A TEMPORARY PACING LEAD. THE PATIENT WAS THEN IMPLANTED WITH A LEADLESS IPG. TWO WEEKS FOLLOWING THIS IMPLANT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21853844&pc=NJK | 3007797756-2025-00014 | 2025-03-17 04:00 | Death | PULMONX CORPORATION | 2025-04-17 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P180002 |  | 1 | Event Description: THIS 81-YEAR-OLD MALE PATIENT SUFFERING FROM SEVERE COPD SINCE 2011 WITH COR PULMONALE DIAGNOSED IN (B)(6) 2024, WAS TREATED ON (B)(6) 2024 WITH 3 ZEPHYR VALVES PLACED IN THE RUL (RBI: 4.0 LP, RB2: 4.0, AND RB3: 4.0). THEN, THE PATIENT WAS DISCHARGED ON (B)(6) 2024 AND HAD NO PROBLEMS FOR 2 MONTHS. ON (B)(6) 2025, THE PATIENT EXPERIENCED ARTHRALGIA AND ANOREXIA WHILE BEING UNWELL SINCE A FEW DAYS. ON (B)(6) IN THE MORNING, HE WAS TRANSPORTED TO EMERGENCY CARE BECAUSE HE WAS FOUND UNABLE TO MOVE BY THE FAMILY. ON ARRIVAL THE PATIENT WAS RECEIVING ANTIBIOTIC (ERYTHROMYCIN) FOR PNEUMONIA, AND MEDICATION FOR COPD (ON SL OXYGEN) AND CARDIAC INSUFFICIENCY (FUROSEMIDE). HE WAS HYPOTENSIVE, WITH TYPE 2 RESPIRATORY FAILURE. AFTER TREATMENT WITH IV FLUIDS, A CHEST CT WAS PERFORMED AND SHOWED AN EXTENSIVE INFILTRATIVE SHADOW IN THE RLL. BLOOD TESTS SHOWED AN ELEVATED INFLAMMATORY RESPONSE. A COMMUNITY-ACQUIRED PNEUMONIA OF THE RLL WAS CONFIRMED (STREPTOCOCCUS PNEUMONIA). THUS, A SEPTIC SHOCK ASSOCIATED WITH THE ONGOING PNEUMONIA WAS CONSIDERED. THE PATIENT WENT INTO A CARDIOPULMONARY ARREST DURING IMAGING SESSION, AND - ALTHOUGH HIS OWN HEARTBEAT RESUMES AFTER ROUTINE CARDIOPULMONARY RESUSCITATION AND THE PATIENT RETURNED TO CONSCIOUSNESS - BLOOD PRESSURE REMAINED LOW AND CIRCULATION PARAMETERS VERY UNSTABLE. AFTER FAMILY CONSULTATION, THE DECISION WAS MADE NOT TO RESUSCITATE THE PATIENT AT THE TIME OF CARDIAC ARREST. DEATH WAS CONFIRMED AT THE EMERGENCY ROOM (B)(6) 2025). IN SUMMARY, THIS SEVERE COPD PATIENT WITH CARDIAC COMPLICATION (COR PULMONALE), DIED 3 MONTHS AFTER BL VR (RUL) WITH ZEPHYR® VALVES FROM A SEPTIC SHOCK RELATED TO PNEUMONIA IN THE RLL. THE DEATH IS NOT RELATED TO THE PROCEDURE, WHICH WAS PERFORMED 3 MONTHS EARLIER, AND UNLIKELY RELATED TO THE VALVES WHICH WERE NOT IMPLANTED IN THE INFECTED LOBE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21854042&pc=OLR | MW5169243 |  | Death | SOLEVA PHARMA LLC. | 2025-04-17 04:00 | OLR | ORAL WOUND DRESSING | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: THE PT WAS MARKED DECEASED IN WALGREENS RETAIL SYSTEM WHICH FLOWED TO WALGREENS SPECIALTY SYSTEM, AND WE ARE REPORTING PER CONTRACTUAL REQUIREMENTS AND STANDARD OPERATING PROCEDURE WITHIN ONE BUSINESS DAY, HOWEVER, NO CAUSE OF DEATH KNOWN. DATE OF DEATH UNK AND NO FURTHER INFO AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21854549&pc=DXY | 2182208-2025-01890 | 2025-02-26 05:00 | Death | MEDTRONIC, INC. | 2025-04-17 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DEVELOPED AN INFECTION. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REMOVED AND THE PATIENT WAS BEING PACED BY A TEMPORARY PACING LEAD. THE PATIENT WAS THEN IMPLANTED WITH A LEADLESS IPG. TWO WEEKS FOLLOWING THIS IMPLANT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21854617&pc=DSQ | 2916596-2025-02436 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL BLEEDING. THE PATIENT'S DEATH WAS RELATED TO HEALTH CONDITION. THE DEVICE WAS WORKING AS INTENDED DURING TIME OF SUPPORT. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURE'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21854717&pc=NVZ | 3004209178-2025-06662 | 2025-02-26 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-17 04:00 | NVZ | ADVISA MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DEVELOPED AN INFECTION. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REMOVED AND THE PATIENT WAS BEING PACED BY A TEMPORARY PACING LEAD. THE PATIENT WAS THEN IMPLANTED WITH A LEADLESS IPG. TWO WEEKS FOLLOWING THIS IMPLANT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21854844&pc=DTB | 2649622-2025-10392 | 2025-03-26 04:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | High impedance | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypovolemic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. POST IMPLANT OF A IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM THE RIGHT ATRIAL (RA) LEAD EXHIBITED HIGH IMPEDANCE AND OVERSENSING, THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED HIGH IMPEDANCE AND THE PATIENT'S HEART RATE WAS IN THE THIRTIES. THE FOLLOWING DAY THE PATIENT'S BLOOD PRESSURE WAS DROPPING. THE PATIENT WAS GIVEN BLOOD AND THERE WAS AN ATTEMPT TO TRANSFER THE PATIENT TO A LARGER CENTER BUT THAT SUDDENLY THERE WAS PULSELESS ELECTRICAL ACTIVITY AND HYPOVOLEMIC SHOCK WAS NOTED. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED BUT SHORTLY AFTER THE PATIENT WAS PRONOUNCED DEAD.Event Description: IT WAS NOTED THAT THE IPG SYSTEM WAS NOT RELATED TO THE PATIENTS DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21854849&pc=NVN | 2649622-2025-10400 | 2025-03-31 04:00 | Death | MPRI | 2025-04-17 04:00 | NVN | SELECTSECURE MRI SURESCAN | Under-Sensing | Septic Shock; Cardiogenic Shock; No Clinical Signs, Symptoms or Conditions | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT A SUSPECTED UNDER SENSED BEAT WAS OBSERVED ON THE RIGHT VENTRICULAR (RV) LEFT BUNDLE BRANCH (LBB) LEAD. IT WAS ALSO REPORTED THAT A DEVICE COUNTER RESET WAS OBSERVED ON THE IMPLANTABLE PULSE GENERATOR (IPG). THE RV LEAD AND IPG REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT WAS IN CARDIOGENETIC SHOCK AND SEPTIC SHOCK AND ULTIMATELY PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21855265&pc=DQY | 2124215-2025-23888 | 2025-04-12 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-17 04:00 | DQY | GUIDEZILLA? II | Difficult to Insert | No Clinical Signs, Symptoms or Conditions | K163314 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. VASCULAR ACCESS WAS OBTAINED VIA THE RIGHT RADIAL ARTERY. THE 80% STENOSED, 33MM, CONCENTRIC DE NOVO TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND MILDLY CALCIFIED LEFT ANTERIOR DESCENDING TO RIGHT CORONARY AND THE LEFT CIRCUMFLEX ARTERY. A 6F GUIDEZILLA CATHETER-GUIDE EXTENSION WAS SELECTED FOR USE. DURING THE PROCEDURE, THE GUIDEZILLA CATHETER FAILED TO PASS THROUGH BOTH THE RADIAL AND FEMORAL SHEATHS. HOWEVER, ANOTHER NON-BOSTON SCIENTIFIC GUIDE EXTENSION CATHETER WAS ADVANCED SMOOTHLY WITHOUT RESISTANCE. DUE TO THIS EVENT, THE PROCEDURE WAS CANCELLED. UNFORTUNATELY, THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21855442&pc=NVZ | 3008973940-2025-02709 | 2025-03-31 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-17 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Data Problem | Septic Shock; Cardiogenic Shock; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT A SUSPECTED UNDER SENSED BEAT WAS OBSERVED ON THE RIGHT VENTRICULAR (RV) LEFT BUNDLE BRANCH (LBB) LEAD. IT WAS ALSO REPORTED THAT A DEVICE COUNTER RESET WAS OBSERVED ON THE IMPLANTABLE PULSE GENERATOR (IPG). THE RV LEAD AND IPG REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT WAS IN CARDIOGENETIC SHOCK AND SEPTIC SHOCK AND ULTIMATELY PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21855770&pc=DTB | 2649622-2025-10387 | 2025-02-26 05:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DEVELOPED AN INFECTION. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REMOVED AND THE PATIENT WAS BEING PACED BY A TEMPORARY PACING LEAD. THE PATIENT WAS THEN IMPLANTED WITH A LEADLESS IPG. TWO WEEKS FOLLOWING THIS IMPLANT THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED IN SEGMENTS, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856347&pc=NVZ | 3008973940-2025-02706 | 2025-03-26 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-17 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Pacing Problem | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypovolemic Shock | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. POST IMPLANT OF A IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM THE RIGHT ATRIAL (RA) LEAD EXHIBITED HIGH IMPEDANCE AND OVERSENSING, THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED HIGH IMPEDANCE AND THE PATIENT'S HEART RATE WAS IN THE THIRTIES. THE FOLLOWING DAY THE PATIENT'S BLOOD PRESSURE WAS DROPPING. THE PATIENT WAS GIVEN BLOOD AND THERE WAS AN ATTEMPT TO TRANSFER THE PATIENT TO A LARGER CENTER BUT THAT SUDDENLY THERE WAS PULSELESS ELECTRICAL ACTIVITY AND HYPOVOLEMIC SHOCK WAS NOTED. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED BUT SHORTLY AFTER THE PATIENT WAS PRONOUNCED DEAD.Event Description: IT WAS NOTED THAT THE IPG SYSTEM WAS NOT RELATED TO THE PATIENTS DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856435&pc=DTB | 2649622-2025-10423 | 2021-10-28 04:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED APPROXIMATELY TWO WEEKS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG). THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856548&pc=DTB | 2649622-2025-10424 | 2021-10-28 04:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED APPROXIMATELY TWO WEEKS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG). THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856582&pc=DSQ | 2916596-2025-02442 | 2024-12-11 05:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS NOT THE CAUSE OF DEATH AND WAS FUNCTIONING WITHOUT ANY ISSUES. LVAD WOULD NOT BE RETURNED FOR ANALYSIS. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT OUTCOME, AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. THE HEARTMATE 3 LVAS IFU LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856721&pc=DSP | 2249723-2025-0001761 | 2025-04-01 04:00 | Death | DATASCOPE CORP. | 2025-04-17 04:00 | DSP | CARDIOSAVE HYBRID W/ E/F PLUG | Calibration Problem | Insufficient Information | K112372 |  | 1 | Event Description: IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) HAD POWER-UP TEST FAILS CODE # 10 AND THE HELIUM TRANSDUCER WAS NOT CALIBRATED. DEVICE WAS UNABLE TO START. PATIENT DIED ONE HOUR AFTER FAILED START OF CARDIOSAVE. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, FULL INITIAL REPORTER NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856777&pc=MFA | 3007284006-2025-00072 | 2025-04-10 04:00 | Death | SPECTRANETICS CORPORATION | 2025-04-17 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Tamponade | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEAD DUE TO CIED SYSTEM/POCKET INFECTION. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH ON THE RV LEAD, THE LEAD WAS REMOVED. THEN, SWITCHING TO THE RA LEAD, THE LEAD WAS REMOVED; HOWEVER, THE PATIENT¿S BLOOD PRESSURE DROPPED, AND CARDIAC TAMPONADE WAS DETECTED. RESCUE EFFORTS BEGAN, INCLUDING EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND THORACOTOMY. A SUPERIOR VENA CAVA PERFORATION (SVC) WAS DISCOVERED AND REPAIRED. THE PATIENT WAS TRANSFERRED TO ICU, BUT PASSED AWAY THE FOLLOWING DAY. THIS REPORT CAPTURES THE 14F GLIDELIGHT DEVICE IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: B6) RELEVANT TESTS/LABORATORY DATA - NOT AVAILABLE FROM FACILITY. D4) DEVICE SERIAL NUMBER - NOT AVAILABLE FROM FACILITY. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21856884&pc=DTB | 2649622-2025-10395 | 2025-03-26 04:00 | Death | MPRI | 2025-04-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | High impedance; Over-Sensing | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypovolemic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. POST IMPLANT OF A IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM THE RIGHT ATRIAL (RA) LEAD EXHIBITED HIGH IMPEDANCE AND OVERSENSING, THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED HIGH IMPEDANCE AND THE PATIENT'S HEART RATE WAS IN THE THIRTIES. THE FOLLOWING DAY THE PATIENT'S BLOOD PRESSURE WAS DROPPING. THE PATIENT WAS GIVEN BLOOD AND THERE WAS AN ATTEMPT TO TRANSFER THE PATIENT TO A LARGER CENTER BUT THAT SUDDENLY THERE WAS PULSELESS ELECTRICAL ACTIVITY AND HYPOVOLEMIC SHOCK WAS NOTED. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED BUT SHORTLY AFTER THE PATIENT WAS PRONOUNCED DEAD.Event Description: IT WAS NOTED THAT THE IPG SYSTEM WAS NOT RELATED TO THE PATIENTS DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857043&pc=OZP | 2032227-2025-155499 | 2025-03-15 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-17 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE PRIMARY CAUSE OF DEATH WAS HEART RELATED ISSUES AND IT WAS SUDDEN. THE PUMP WAS WORN AT THE TIME OF PASSING. THE PRODUCT INVOLVE IN THE EVENT WAS MMT-1884. TROUBLESHOOTING WAS PERFORMED AND THE REPORTING PARTY REPORT THAT OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITION INCLUDES HEART PROBLEM BESIDES DIABETES WHICH LEAD TO CUSTOMER DEATH. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857162&pc=OZD | 1220648-2025-27591 | 2025-03-21 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE 70 YEAR OLD FEMALE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, AN ECHO WAS PERFORMED WHICH REVEALED THE DEVICE HAD FLIPPED AND INLET WAS NEAR PAPILLARY BED. ECHO WAS BROUGHT BACK TO BEDSIDE AND SEVERAL ATTEMPTS WERE MADE TO REPOSITION CP. THE DOCTOR BELIEVED THE PIGTAIL OF THE CP WAS WRAPPED/WEDGED INTO CHORDAE TENDINEAE. ADDITIONALLY, THE PUMP WAS ONLY ABLE TO ACHIEVE FLOWS AT P6. THE PUMP WAS EXPLANTED AND AN IMPELLA 5.5 WAS IMPLANTED. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857280&pc=NPT | 2025587-2025-02758 | 2025-01-02 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-17 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857518&pc=PNJ | 9612164-2025-01957 | 2025-02-26 05:00 | Death | MEDTRONIC IRELAND | 2025-04-17 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT DEVELOPED AN INFECTION. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REMOVED AND THE PATIENT WAS BEING PACED BY A TEMPORARY PACING LEAD. THE PATIENT WAS THEN IMPLANTED WITH A LEADLESS IPG. TWO WEEKS FOLLOWING THIS IMPLANT THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857624&pc=FPA | 9616066-2025-00936 | 2025-02-27 05:00 | Death | SISTEMAS MEDICOS ALARIS S.A. DE C.V. | 2025-04-17 04:00 | FPA | UNSPECIFIED BD INFUSION SET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 1 | Event Description: IT WAS REPORTED THAT UNSPECIFIED BD INFUSION SET MAY OF HAD ISSUES THE FOLLOWING INFORMATION WAS RECEIVED BY THE INITIAL REPORTER WITH THE FOLLOWING. ALARIS UNKNOWN MODEL PCU WITH PCA ATTACHED AND LVP MODULE FOR LUIDS ANDINFUSION SET 3PORT PRIMARY UNKNOWN REF AND LOT. ISSUE: DURING SHIFT CHANGE RN FOUND THAT PT WAS CODING AND DEVICE WAS SENT TO BIO TO DETERMINE IF IT WAS A CONTRIBUTING FACTOR TO THE PT CODING BUT AS OF NOW UNKNOWN. PT EXPIRED UNDETERMINED CAUSES AS OF NOW, PT SUFFERED FROM SICKLE CELL. DOE: 27FEB2025. WE INITIALLY TRIED TO RUN A PUMP REPORT AND WE ARE NOT ABLE TO ACCESS THE DATA. BD WAS ABLE TO GET US THE REPORT FROM THE PUMP AND IT WAS IDENTIFIED FROM THE TIMELINE THAT THE PUMP WAS NOT A FACTOR IN THE EVENT. THE ISSUE OF NOT BEING ABLE TO ACCESS PUMP SPECIFIC REPORTS HAS NOW BEEN RESOLVED BY THE BD TEAM. Manufacturer Narrative: WE INITIALLY TRIED TO RUN A PUMP REPORT AND WE ARE NOT ABLE TO ACCESS THE DATA. BD WAS ABLE TO GET US THE REPORT FROM THE PUMP AND IT WAS IDENTIFIED FROM THE TIMELINE THAT THE PUMP WAS NOT A FACTOR IN THE EVENT. THE ISSUE OF NOT BEING ABLE TO ACCESS PUMP SPECIFIC REPORTS HAS NOW BEEN RESOLVED BY THE BD TEAM. NO PRODUCT OR PHOTO WAS RETURNED BY THE CUSTOMER. THE CUSTOMER COMPLAINT THAT THE TUBING WAS CONNECTED TO THE PATIENT AT THE TIME OF THE INCIDENT. A DEVICE HISTORY RECORD REVIEW COULD NOT BE PERFORMED BECAUSE A MODEL AND LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. DUE TO NO SAMPLE BEING RECEIVED, AN INVESTIGATION COULD NOT BE PERFORMED AND A ROOT CAUSE COULD NOT BE DETERMINED. THIS INCIDENT HAS BEEN ADDED TO OUR DATABASE OF REPORTED INCIDENTS. OUR BUSINESS TEAM REGULARLY REVIEWS THE COLLECTED DATA FOR IDENTIFICATION OF EMERGING TRENDS. YOUR ASSISTANCE IN THIS MATTER HAS BEEN HELPFUL IN TREND IDENTIFICATION AND SUPPORTING OUR COMMITMENT TO CONTINUOUS QUALITY IMPROVEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857655&pc=NJK | 3007797756-2025-00021 | 2025-02-28 05:00 | Death | PULMONX CORPORATION | 2025-04-17 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P180002 |  | 1 | Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES ON (B)(6) 2025. THE PATIENT WAS TRANSFERRED TO ICU POST-PROCEDURE AND THE VALVES WERE REMOVED IN THE FOLLOWING DAYS AFTER THE PROCEDURE. IT WAS MADE AWARE ON (B)(6) 2025 THAT THE PATIENT HAD PASSED AWAY WHILE IN RECOVERY.Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES (ONE VALVE IN THE RIGHT MIDDLE LOBE (RML) AND THREE IN THE RIGHT UPPER LOBE (RUL) ON (B)(6) 2025. PER THE PHYSICIAN, THE PROCEDURE DURATION LASTED FOR ABOUT ONE HOUR. ON (B)(6) 2025, ALL FOUR VALVES WERE EXPLANTED DUE TO A DECLINE IN THE PATIENT'S RESPIRATORY STATUS FOLLOWING THE PROCEDURE WHILE IN THE ICU. THE VALVES REMAINED IMPLANTED FOR A TOTAL OF 3 DAYS. ON (B)(6) 2025, IMMEDIATELY FOLLOWING THE BRONCHOSCOPY, THE PATIENT EXPERIENCED A FAILED EXTUBATION DUE TO CO[?] RETENTION AND HYPERSOMNOLENCE, REQUIRING RE-INTUBATION AND TRANSFER TO THE ICU. DESPITE VARIOUS VENTILATOR ADJUSTMENTS, THE PATIENT DEMONSTRATED VERY POOR AIR MOVEMENT. THE VALVES WERE REMOVED AFTER THREE DAYS; HOWEVER, THE PATIENT'S CONDITION DID NOT IMPROVE. THE PATIENT CONTINUED TO DETERIORATE AND SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025. PER THE MEDICAL CHART, THE FOLLOWING STATEMENT WAS DOCUMENTED: "DUE TO PERSISTENCE OF VENTILATOR DEPENDENCE AND INABILITY TO WEAN PATIENT SUCCESSFULLY FROM PARALYSIS AND DEEP SEDATION, COUNSELED FAMILY THAT NEXT TREATMENT OPTION WAS TO UNDERGO TRACHEOSTOMY PLACEMENT VERSUS PURSUE COMFORT FOCUSED CARE. IN SPEAKING WITH MR. (B)(6) FAMILY, THEY RELAYED THAT TRACHEOSTOMY WAS NOT IN LINE WITH HIS GOALS OF CARE. AS SUCH, PATIENT WAS TRANSITIONED TO COMFORT FOCUSED CARE ON (B)(6) 2025 AND WAS PALLIATIVELY EXTUBATED AT THIS TIME. MR. (B)(6) PASSED PEACEFULLY SHORTLY THEREAFTER ON (B)(6) 2025 AT 20:30." ACCORDING TO THE PATIENT'S CHART AND MEDICAL RECORDS, THE PRIMARY CAUSE OF DEATH WAS ACUTE HYPOXEMIC HYPERCAPNIC RESPIRATORY FAILURE ATTRIBUTED TO SEVERELY LIMITED PULMONARY RESERVE AND THE EFFECTS OF GENERAL ANESTHESIA. A SECONDARY CONTRIBUTING FACTOR WAS A COPD EXACERBATION. NO AUTOPSY WAS PERFORMED. ACCORDING TO THE PHYSICIAN, THE VALVES WERE DISCARDED AS THEY WERE NOT BELIEVED TO BE RELATED TO THE PATIENT'S CLINICAL DETERIORATION OR CAUSE OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857674&pc=NJK | 3007797756-2025-00022 | 2025-02-28 05:00 | Death | PULMONX CORPORATION | 2025-04-17 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P180002 |  | 1 | Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES ON (B)(6) 2025. THE PATIENT WAS TRANSFERRED TO ICU POST-PROCEDURE AND THE VALVES WERE REMOVED IN THE FOLLOWING DAYS AFTER THE PROCEDURE. IT WAS MADE AWARE ON (B)(6) 2025 THAT THE PATIENT HAD PASSED AWAY WHILE IN RECOVERY.Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES (ONE VALVE IN THE RIGHT MIDDLE LOBE (RML) AND THREE IN THE RIGHT UPPER LOBE (RUL) ON (B)(6) 2025. PER THE PHYSICIAN, THE PROCEDURE DURATION LASTED FOR ABOUT ONE HOUR. ON (B)(6) 2025, ALL FOUR VALVES WERE EXPLANTED DUE TO A DECLINE IN THE PATIENT'S RESPIRATORY STATUS FOLLOWING THE PROCEDURE WHILE IN THE ICU. THE VALVES REMAINED IMPLANTED FOR A TOTAL OF 3 DAYS. ON (B)(6) 2025, IMMEDIATELY FOLLOWING THE BRONCHOSCOPY, THE PATIENT EXPERIENCED A FAILED EXTUBATION DUE TO CO[?] RETENTION AND HYPERSOMNOLENCE, REQUIRING RE-INTUBATION AND TRANSFER TO THE ICU. DESPITE VARIOUS VENTILATOR ADJUSTMENTS, THE PATIENT DEMONSTRATED VERY POOR AIR MOVEMENT. THE VALVES WERE REMOVED AFTER THREE DAYS; HOWEVER, THE PATIENT'S CONDITION DID NOT IMPROVE. THE PATIENT CONTINUED TO DETERIORATE AND SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025. PER THE MEDICAL CHART, THE FOLLOWING STATEMENT WAS DOCUMENTED: "DUE TO PERSISTENCE OF VENTILATOR DEPENDENCE AND INABILITY TO WEAN PATIENT SUCCESSFULLY FROM PARALYSIS AND DEEP SEDATION, COUNSELED FAMILY THAT NEXT TREATMENT OPTION WAS TO UNDERGO TRACHEOSTOMY PLACEMENT VERSUS PURSUE COMFORT FOCUSED CARE. IN SPEAKING WITH MR. (B)(6) FAMILY, THEY RELAYED THAT TRACHEOSTOMY WAS NOT IN LINE WITH HIS GOALS OF CARE. AS SUCH, PATIENT WAS TRANSITIONED TO COMFORT FOCUSED CARE ON (B)(6) 2025 AND WAS PALLIATIVELY EXTUBATED AT THIS TIME. MR. (B)(6) PASSED PEACEFULLY SHORTLY THEREAFTER ON (B)(6) 2025 AT 20:30." ACCORDING TO THE PATIENT'S CHART AND MEDICAL RECORDS, THE PRIMARY CAUSE OF DEATH WAS ACUTE HYPOXEMIC HYPERCAPNIC RESPIRATORY FAILURE ATTRIBUTED TO SEVERELY LIMITED PULMONARY RESERVE AND THE EFFECTS OF GENERAL ANESTHESIA. A SECONDARY CONTRIBUTING FACTOR WAS A COPD EXACERBATION. NO AUTOPSY WAS PERFORMED. ACCORDING TO THE PHYSICIAN, THE VALVES WERE DISCARDED AS THEY WERE NOT BELIEVED TO BE RELATED TO THE PATIENT'S CLINICAL DETERIORATION OR CAUSE OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857792&pc=OZD | 1220648-2025-27592 | 2025-04-09 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT DUE TO HEART FAILURE. IT WAS REPORTED THAT WHILE ON SUPPORT, THE PATIENT EXPERIENCED TORSADES DE POINT AND EXPIRED. THE CAUSE OF DEATH IS UNKNOWN, BUT IT IS BELIEVED TO BE SEPTIC SHOCK, WHICH RESULTED IN VENTRICULAR FIBRILLATION. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857839&pc=GKT | 1722028-2025-00091 | 2025-01-29 05:00 | Death | TERUMO BCT | 2025-04-17 04:00 | GKT | SPECTRA OPTIA | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Abdominal Pain; Headache; Nausea; Vomiting; Jaundice; Melena | K183081 |  | 1 | Event Description: PER JOURNAL ARTICLE "INTENSIVE THERAPEUTIC PLASMA EXCHANGE¿NEW APPROACH TO TREAT AND RESCUE PATIENTS WITH SEVERE FORM OF YELLOW FEVER" BY HO, Y., NUKUI, Y., VILLAÇA, P. R., OKAZAKI, E., TATSUI, N. H., NETTO, L. C., JOELSONS, D., DA ROCHA, T. R. F., DE MELLO MALTA, F., PINHO, J. R. R., SEGURADO, A. A. C., & ROCHA, V. THE STUDY DISCUSSES THE USE OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN PATIENTS WITH SEVERE YELLOW FEVER, UTILIZING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. WHILE THE STUDY REPORTS A SIGNIFICANT REDUCTION IN MORTALITY WITH INTENSIVE TPE, IT ALSO NOTES ADVERSE EVENTS SUCH AS HYPOCALCEMIA AND ALLERGIC REACTIONS DURING THE PROCEDURES. SPECIFICALLY, ONE PATIENT EXPERIENCED AN ALLERGIC REACTION DURING TPE, WHICH IS A CONCERN FOR PATIENT SAFETY. THE USE OF TERUMO PRODUCTS IN THIS CONTEXT RAISES QUESTIONS ABOUT THE MANAGEMENT OF SUCH ADVERSE EVENTS AND THE NEED FOR MONITORING DURING PROCEDURES. THE FINDINGS SUGGEST THAT WHILE THE DEVICES ARE EFFECTIVE IN MANAGING SEVERE YELLOW FEVER, THERE ARE POTENTIAL RISKS ASSOCIATED WITH THEIR USE THAT NEED TO BE ADDRESSED TO ENSURE PATIENT SAFETY. THIS OBSERVATIONAL CASE-SERIES STUDY EVALUATES THE EFFECTIVENESS OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN TREATING SEVERE YELLOW FEVER, PARTICULARLY IN PATIENTS WITH ACUTE LIVER FAILURE. THE STUDY INVOLVED THREE GROUPS OF PATIENTS RECEIVING DIFFERENT TREATMENT REGIMENS, WITH A FOCUS ON THE OUTCOMES ASSOCIATED WITH INTENSIVE TPE USING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. THE RESULTS INDICATED A SIGNIFICANT REDUCTION IN MORTALITY RATES AMONG PATIENTS RECEIVING INTENSIVE TPE COMPARED TO THOSE RECEIVING STANDARD INTENSIVE CARE SUPPORT OR HIGH-VOLUME TPE. HOWEVER, THE STUDY ALSO HIGHLIGHTED ADVERSE EVENTS, INCLUDING ALLERGIC REACTIONS AND HYPOCALCEMIA, WHICH OCCURRED DURING THE PROCEDURES. THE AUTHORS CONCLUDED THAT INTENSIVE TPE IS A FEASIBLE AND EFFECTIVE INTERVENTION FOR SEVERE YELLOW FEVER, BUT FURTHER STUDIES ARE NEEDED TO CONFIRM THESE FINDINGS AND ADDRESS THE ASSOCIATED RISKS. ONE WAS A 48-YO MALE PATIENT (VPP) WHO REPORTED LONG-TERM USE OF ALCOHOL, WITH FOUR DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (403 AND 148 TIMES GREATER THAN THE UPPER LIMIT OF NORMAL [ULN], RESPECTIVELY), WITH HIGH AMMONIA LEVELS (3.7 TIMES THE ULN) AND LOW LEVELS OF FV (REDUCTION OF 90% FROM NORMAL). FOUR DIFFERENT BACTERIA WERE IDENTIFIED 48 HS LATER IN ADMISSION BLOOD CULTURE (E. COLI, S. VIRIDIANS, S. PNEUMONIAE AND F. NUCLEATUM). DESPITE ICS, ANTIBIOTICS, AND INTENSIVE TPE, THE PATIENT PRESENTED WITH A MASSIVE ALVEOLAR HEMORRHAGE AND DIED 72 HS AFTER ADMISSION (TABLE 2). THE SECOND WAS ALSO A 48-YO MALE PATIENT (VSS), PREVIOUSLY VACCINATED AGAINST YFV (10 YEARS BEFORE), WITH SIX DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (163 AND 103 TIMES THE ULN, RESPECTIVELY), HIGH AMMONIA LEVELS (2.3 TIMES THE ULN), AND NORMAL LEVELS OF FV. AFTER FIVE DAYS OF INTENSIVE TPE, ALL LABORATORY RESULTS WERE NORMALIZED. HOWEVER, ON THE SIXTH DAY OF ADMISSION, S. AUREUS WAS IDENTIFIED IN A BLOOD CULTURE AND THE PATIENT DIED SEVEN DAYS LATER. ONLY MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. HYPOCALCEMIA (IONIC CALCIUM < 4.0 MG/DL) WAS MORE FREQUENT IN G2 (P < 0.001) THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. SPECIFIC DETAILS, SUCH AS PATIENT INFORMATION AND OUTCOME, WERE NOT INCLUDED IN THE ARTICLE FOR THESE EVENTS, THEREFORE THIS REPORT IS BEING PROVIDED AS A SUMMARY OF THE EVENTS. IT IS UNKNOWN AT THIS TIME IF MEDICAL INTERVENTION WAS REQUIRED FOR THESE EVENTS THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION.Event Description: PER JOURNAL ARTICLE "INTENSIVE THERAPEUTIC PLASMA EXCHANGE¿NEW APPROACH TO TREAT AND RESCUE PATIENTS WITH SEVERE FORM OF YELLOW FEVER" BY HO, Y., NUKUI, Y., VILLAÇA, P. R., OKAZAKI, E., TATSUI, N. H., NETTO, L. C., JOELSONS, D., DA ROCHA, T. R. F., DE MELLO MALTA, F., PINHO, J. R. R., SEGURADO, A. A. C., & ROCHA, V. THE STUDY DISCUSSES THE USE OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN PATIENTS WITH SEVERE YELLOW FEVER, UTILIZING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. WHILE THE STUDY REPORTS A SIGNIFICANT REDUCTION IN MORTALITY WITH INTENSIVE TPE, IT ALSO NOTES ADVERSE EVENTS SUCH AS HYPOCALCEMIA AND ALLERGIC REACTIONS DURING THE PROCEDURES. SPECIFICALLY, ONE PATIENT EXPERIENCED AN ALLERGIC REACTION DURING TPE, WHICH IS A CONCERN FOR PATIENT SAFETY. THE USE OF TERUMO PRODUCTS IN THIS CONTEXT RAISES QUESTIONS ABOUT THE MANAGEMENT OF SUCH ADVERSE EVENTS AND THE NEED FOR MONITORING DURING PROCEDURES. THE FINDINGS SUGGEST THAT WHILE THE DEVICES ARE EFFECTIVE IN MANAGING SEVERE YELLOW FEVER, THERE ARE POTENTIAL RISKS ASSOCIATED WITH THEIR USE THAT NEED TO BE ADDRESSED TO ENSURE PATIENT SAFETY. THIS OBSERVATIONAL CASE-SERIES STUDY EVALUATES THE EFFECTIVENESS OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN TREATING SEVERE YELLOW FEVER, PARTICULARLY IN PATIENTS WITH ACUTE LIVER FAILURE. THE STUDY INVOLVED THREE GROUPS OF PATIENTS RECEIVING DIFFERENT TREATMENT REGIMENS, WITH A FOCUS ON THE OUTCOMES ASSOCIATED WITH INTENSIVE TPE USING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. THE RESULTS INDICATED A SIGNIFICANT REDUCTION IN MORTALITY RATES AMONG PATIENTS RECEIVING INTENSIVE TPE COMPARED TO THOSE RECEIVING STANDARD INTENSIVE CARE SUPPORT OR HIGH-VOLUME TPE. HOWEVER, THE STUDY ALSO HIGHLIGHTED ADVERSE EVENTS, INCLUDING ALLERGIC REACTIONS AND HYPOCALCEMIA, WHICH OCCURRED DURING THE PROCEDURES. THE AUTHORS CONCLUDED THAT INTENSIVE TPE IS A FEASIBLE AND EFFECTIVE INTERVENTION FOR SEVERE YELLOW FEVER, BUT FURTHER STUDIES ARE NEEDED TO CONFIRM THESE FINDINGS AND ADDRESS THE ASSOCIATED RISKS. THE SECOND WAS ALSO A 48-YO MALE PATIENT (VSS), PREVIOUSLY VACCINATED AGAINST YFV (10 YEARS BEFORE), WITH SIX DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (163 AND 103 TIMES THE ULN, RESPECTIVELY), HIGH AMMONIA LEVELS (2.3 TIMES THE ULN), AND NORMAL LEVELS OF FV. AFTER FIVE DAYS OF INTENSIVE TPE, ALL LABORATORY RESULTS WERE NORMALIZED. HOWEVER, ON THE SIXTH DAY OF ADMISSION, S. AUREUS WAS IDENTIFIED IN A BLOOD CULTURE AND THE PATIENT DIED SEVEN DAYS LATER. ONLY MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. HYPOCALCEMIA (IONIC CALCIUM < 4.0 MG/DL) WAS MORE FREQUENT IN G2 (P < 0.001) THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. SPECIFIC DETAILS, SUCH AS FURTHER PATIENT INFORMATION AND OUTCOME, WERE NOT INCLUDED IN THE ARTICLE FOR THESE EVENTS, THEREFORE THIS REPORT IS BEING PROVIDED AS A SUMMARY OF THE EVENTS. IT IS UNKNOWN AT THIS TIME IF MEDICAL INTERVENTION WAS REQUIRED FOR THESE EVENTS THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: INVESTIGATION: THIS IS AN OBSERVATIONAL CASE-SERIES UNICENTRIC STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK. IT WAS CONDUCTED AT HOSPITAL DAS CLINICAS (HCFMUSP), A 2200-BED TERTIARY CARE TEACHING HOSPITAL AFFILIATED TO THE UNIVERSITY OF SAO PAULO WITH 110 INTENSIVE CARE BEDS AND A 7-BED INFECTIOUS DISEASE INTENSIVE CARE UNIT (ID-ICU). ON 9 JANUARY 2018, THE ID-ICU HAD BECOME THE CENTER OF EXPERTISE FOR SEVERE FORMS OF YF IN SP. SEVERITY CRITERIA FOR ID-ICU ADMISSION WERE: ASPARTATE AMINOTRANSFERASE (AST) OR ALANINE AMINOTRANSFERASE (ALT) > 3000 U/L; INR > 1.5; PLATELET COUNT < 90 × 109/L; KIDNEY DYSFUNCTION; HEMORRHAGIC MANIFESTATIONS; ENCEPHALOPATHY; OR HEMODYNAMIC INSTABILITY. HOWEVER, IN 2019, A YELLOW FEVER OUTBREAK OCCURRED IN A REGION OF THE STATE OF SAO PAULO WHERE TRANSPORTATION TO HCFMUSP TAKES AT LEAST 6 H. FOR THIS REASON, ALL PATIENTS WHO HAD ALT/AST > 500 WERE REFERRED TO HCFMUSP. FROM 9 JAN TO 15 FEB 2018, TREATMENT CONSISTED OF STANDARD INTENSIVE CARE SUPPORT (ICS). IN THE SECOND HALF OF FEBRUARY 2018, HIGH-VOLUME TPE (HV-TPE) WAS USED IN SOME PATIENTS WITH SEVERE YF IN OUR SERVICE. TRANSIENT IMPROVEMENT OF LABORATORY RESULTS WAS OBSERVED, HOWEVER HEMORRHAGIC PHENOMENA OCCURRED BETWEEN SESSIONS OF HV-TPE WITHOUT CHANGING THE OUTCOME. SINCE APRIL 2018, THIS STRATEGY WAS MODIFIED TO MORE INTENSIVE TPE TWICE A DAY AND SUPPORTIVE TRANSFUSION, IF NECESSARY. THIS INTENSIVE PROTOCOL WAS USED ON ONE PATIENT IN 2018 SUCCESSFULLY, AND SINCE THEN IT HAS BEEN USED IN THE STATE OF SAO PAULO¿S 2019 YF OUTBREAK. TPE WAS ADMINISTERED TO PATIENTS WITH LEVELS OF AMMONIA > 70 MCMOL/L AND/OR FACTOR V (FV) < 50%. THE DEFINITION OF THE CUT-OFF OF AMMONIA WAS BASED ON THE 2018 EXPERIENCE WHEN PATIENTS WITH AMMONIA > 70 MCMOL/L HAD A HIGHER RELATIVE RISK OF DEATH (RR 2.0) (DATA NOT PUBLISHED). THE CUT-OFF OF FV < 50% WAS DEFINED BASED ON CLICHY¿S CRITERIA, CONSIDERING A REDUCTION OF 30% FROM THE REFERENCE VALUE OF 70¿150% FOR ALL PATIENTS. IN ORDER TO COMPARE TO PREVIOUS CASES, ALL PATIENTS WITH AMMONIA > 70 MCMOL/L AND/OR FV < 50% WERE SELECTED. FOR THIS ANALYSIS, PATIENTS WERE CLASSIFIED INTO THREE GROUPS ACCORDING TO TYPE OF TREATMENT GIVEN: GROUP ONE (G1)¿ICS; GROUP TWO (G2)¿ICS AND HV-TPE; AND GROUP THREE (G3)¿ICS AND INTENSIVE TPE. CLINICAL AND LABORATORY DATA WERE EXTRACTED RETROSPECTIVELY AND STORED IN THE REDCAP WEB PLATFORM. HV-TPE STRATEGY (G2): PROPOSED TREATMENT CONSISTED OF THREE CONSECUTIVE DAYS OF PLASMA EXCHANGE. THE VOLUME CHANGED WAS 10% OF BODY WEIGHT (BW) WITH FRESH FROZEN PLASMA (FFP). ONE TO THREE ADDITIONAL SESSIONS OF ONE PLASMA VOLUME EACH WERE IMPLEMENTED ACCORDING TO AMMONIA LEVELS. SUPPORTIVE TRANSFUSION WAS GUIDED MAINLY BY HEMORRHAGIC EPISODES. INTENSIVE-TPE STRATEGY (G3): VOLUME OF EXCHANGE WITH FFP WAS CALCULATED ACCORDING TO BW, HEMATOCRIT (HT), AND PLASMA VOLUME (BW × 0.7 × [1 - HT]), CONSIDERING ONE PLASMA VOLUME PER SESSION. PROCEDURE WAS PERFORMED TWICE A DAY, WITH ADDITIONAL FFP INFUSION (10 ML/KG) AT NIGHT. PLATELET CONCENTRATE WAS TRANSFUSED TO KEEP PLATELET COUNT > 30 × 109/L OR, IN THE PRESENCE OF BLEEDING, >50 × 109/L. IN SEVERE OR LIFE-THREATENING BLEEDS, TARGET WAS >80¿100 × 109/L. CRYOPRECIPITATE TRANSFUSION WAS INDICATED IF FIBRINOGEN LEVEL WAS <100 MG/DL. BASED ON BLEEDING AS WELL AS FV AND/OR AMMONIA LEVELS, THE FREQUENCY OF TPE WAS TAPERED TO ONCE A DAY, EVERY OTHER DAY AND DISCONTINUATION. DURING THE DE-ESCALATION PERIOD, FFP WAS INFUSED (5 ML/KG) 1¿2 TIMES/DAY, FOLLOWING THE SAME CRITERIA ABOVE. TRANEXAMIC ACID (10 MG/KG/DOSE, IV, EVERY 6¿8 H) WAS PRESCRIBED UNTIL DISCONTINUATION OF TPE AND TOPICAL USE WAS IMPLEMENTED IN CASES OF MUCOCUTANEOUS BLEEDING. TPE WAS PERFORMED USING FFP REPLACEMENT, WITHOUT ANTICOAGULANT (CITRATE), ONCE ALL PATIENTS HAD ALF. CALCIUM CHLORIDE WAS ADMINISTERED AT THE FOLLOWING RATES: FOR HV-TPE GROUP, 5 MMOL PER EACH LITER OF PLASMA EXCHANGE; FOR INTENSIVE-TPE GROUP, FIXED 6.8 MMOL OF CALCIUM PER SESSION. AFTER EACH SESSION, SERUM CALCIUM LEVEL WAS ASSESSED, AND ADDITIONAL SUPPLEMENTATION WAS CONSIDERED. COBE SPECTRA AND OPTIA (TERUMOBCT, TOKYO, JAPAN) AND COM.TECR (FRESENIUS KABI, BAD HOMBURG, GERMANY) TPE DEVICES WERE EMPLOYED, USING A WITHDRAWAL FLOW OF 70 ML/MINUTE ON AVERAGE. IN CASES WHERE CONTINUOUS VENOVENOUS HEMODIALYSIS WAS INDICATED, THE PLASMA EXCHANGE SYSTEM WAS CONNECTED TO THE RETURN CIRCUIT OF THE DIALYSIS SYSTEM. YFV-RNA WAS ISOLATED FROM 250 ¿L OF SERUM SAMPLES AND 20 ¿G OF RNASE-FREE GLYCOGEN WERE ADDED TO THE AQUEOUS PHASE AS AN RNA CARRIER. REAL-TIME QUANTITATIVE YFV REVERSE TRANSCRIPTION (RT-QPCR) WAS CARRIED OUT USING SUPERSCRIPTR III ONE-STEP RT-QPCR SYSTEM WITH PLATINUMR TAQ DNA POLYMERASE (INVITROGEN, THERMO FISHER SCIENTIFIC BRAND, CARLSBAD, CA, USA). THE YFV 5'NON-CODING REGION WAS TARGETED USING SPECIFIC PRIMERS AND PROBE, AS PREVIOUSLY DESCRIBED. ALL SAMPLES WERE TESTED IN TRIPLICATE. THE PRIMARY END POINT WAS MORTALITY. BASELINE CLINICAL AND LABORATORY ASPECTS WERE COMPARED AMONG GROUPS. SURVIVAL CURVES WERE CONSTRUCTED USING THE KAPLAN¿MEIER METHOD, AND THE LOG-RANK TEST WAS USED TO ASSESS DIFFERENCES BETWEEN CURVES. PROPORTIONAL HAZARD ASSUMPTION WAS CHECKED GRAPHICALLY BY LOG-MINUS-LOG VS. LOG (TIME) PLOT. COMPARISONS BETWEEN TWO GROUPS WERE ASSESSED USING THE STUDENT T-TEST, OR THE MANN¿WHITNEY TEST IF APPROPRIATE FOR QUANTITATIVE VARIABLES, AND THE FISHER EXACT TEST FOR CATEGORICAL VARIABLES. ASSUMPTIONS OF NORMALITY AND HOMOGENEITY OF VARIANCE FOR T-TEST WERE EVALUATED. DIFFERENCES BETWEEN THE 3 GROUPS IN QUANTITATIVE VARIABLES WERE EVALUATED BY THE KRUSKAL¿WALLIS TEST. COMPARISON OF PROPORTIONS BETWEEN THE 3 GROUPS WAS PERFORMED BY USING THE FISHER-EXACT TEST. QUANTITATIVE VARIABLES WERE PRESENTED AS MEDIAN (INTERQUARTILE RANGE [IQR]). CATEGORICAL VARIABLES WERE EXPRESSED AS COUNTS AND PERCENTAGES. NORMALITY WAS ASSESSED BY THE SHAPIRO¿WILK NORMALITY TEST, THE D¿AGOSTINO¿PEARSON OMNIBUS NORMALITY TEST, AND VISUAL INSPECTION OF HISTOGRAMS. ALL P-VALUES WERE TWO-SIDED. RESULTS WERE CONSIDERED SIGNIFICANT IF P < 0.05. STATISTICAL ANALYSES WERE PERFORMED USING SAS VERSION 9.4 (SAS INSTITUTE INC., CARY, NC, USA). DURING THE 2018/2019 YELLOW FEVER OUTBREAK IN SÃO PAULO, 114 CONFIRMED CASES OF YELLOW FEVER WERE REFERRED TO THE ID-ICU OF HCFMUSP. SIXTY-SIX PATIENTS HAD LEVELS OF AMMONIA > 70 MCMOL/L AND/OR AN FV OF <50% AND WERE INCLUDED FOR ANALYSIS: 41 (62%) WERE TREATED WITH ICS (G1), 11 (17%) RECEIVED ICS PLUS HV-TPE (G2) AND 14 (21%) RECEIVED ICS PLUS INTENSIVE TPE (G3). THERE WERE NO DIFFERENCES AMONG THE THREE GROUPS RELATED TO SEX AND CLINICAL PRESENTATION AT ADMISSION. ALL GROUPS HAD VALUES OF AST/ALT > 3000 U/L, HOWEVER PATIENTS OF G3 HAD LOWER LEVELS OF AST/ALT COMPARED TO THE OTHER GROUPS, PROBABLY AS A CONSEQUENCE OF THE CHANGE IN THE TRANSFER CRITERIA, WHICH ALLOWED EARLIER ASSISTANCE, PREVENTING THE PROGRESSION OF THE DISEASE. NO DIFFERENCES WERE OBSERVED AMONG THE GROUPS WITH RESPECT TO THE OTHER LABORATORY TESTS. CONCERNING G3, THE PATIENTS WERE ADMITTED WITH A MEDIAN OF SIX DAYS OF SYMPTOMS AND FEVER WAS NOTED IN 86% OF THEM. SURPRISINGLY, JAUNDICE WAS PRESENT IN ONLY 43% OF THE PATIENTS AT ADMISSION. HEMORRHAGIC PHENOMENA OCCURRED IN 11 PATIENTS (79%), THE VAST MAJORITY BEING GASTROINTESTINAL BLEEDING. TWELVE PATIENTS (86%) HAD SEIZURES AND/OR HEADACHES AS NEUROLOGIC MANIFESTATIONS. ONE OF THE FEMALE PATIENTS (LO) WAS PREGNANT AND WAS DIAGNOSED WITH REACTIVE HEMOPHAGOCYTIC SYNDROME, WHICH WAS SUCCESSFULLY TREATED WITH HUMAN INTRAVENOUS IMMUNOGLOBULIN. THE MEDIAN VOLUME TREATED AND TIME PER PROCEDURE WERE HIGHER IN G2 (P = 0.005 AND P < 0.001, RESPECTIVELY) ALTHOUGH G3 HAD A HIGHER MEDIAN NUMBER OF PROCEDURES PER PATIENT (P < 0.001) (TABLE 1). HOWEVER, 45% OF THE G2 PATIENTS DIED BEFORE CONCLUDING THE PROPOSED TREATMENT (THREE CONSECUTIVE DAYS). TRANSFUSION REACTIONS SUCH AS TRANSFUSION-RELATED ACUTE LUNG INJURY, TRANSFUSION ASSOCIATED CIRCULATORY OVERLOAD OR ABO INCOMPATIBILITY WERE NOT OBSERVED. ONLY MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. HYPOCALCEMIA (IONIC CALCIUM < 4.0 MG/DL) WAS MORE FREQUENT IN G2 (P < 0.001). THE MORTALITY RATES IN G1 AND G2 WERE SIMILAR, AT 85% AND 82%, RESPECTIVELY. THE MORTALITY RATE FOR PATIENTS IN G3 WAS SIGNIFICANTLY INFERIOR, AS DEMONSTRATED IN THE KAPLAN¿MEIER SURVIVAL CURVE. ONLY TWO PATIENTS DIED IN G3. ONE WAS A 48-YO MALE PATIENT (VPP) WHO REPORTED LONG-TERM USE OF ALCOHOL, WITH FOUR DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (403 AND 148 TIMES GREATER THAN THE UPPER LIMIT OF NORMAL [ULN], RESPECTIVELY), WITH HIGH AMMONIA LEVELS (3.7 TIMES THE ULN) AND LOW LEVELS OF FV (REDUCTION OF 90% FROM NORMAL). FOUR DIFFERENT BACTERIA WERE IDENTIFIED 48 HS LATER IN ADMISSION BLOOD CULTURE (E. COLI, S. VIRIDIANS, S. PNEUMONIAE AND F. NUCLEATUM). DESPITE ICS, ANTIBIOTICS, AND INTENSIVE TPE, THE PATIENT PRESENTED WITH A MASSIVE ALVEOLAR HEMORRHAGE AND DIED 72 HS AFTER ADMISSION (TABLE 2). THE SECOND WAS ALSO A 48-YO MALE PATIENT (VSS), PREVIOUSLY VACCINATED AGAINST YFV (10 YEARS BEFORE), WITH SIX DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (163 AND 103 TIMES THE ULN, RESPECTIVELY), HIGH AMMONIA LEVELS (2.3 TIMES THE ULN), AND NORMAL LEVELS OF FV. AFTER FIVE DAYS OF INTENSIVE TPE, ALL LABORATORY RESULTS WERE NORMALIZED. HOWEVER, ON THE SIXTH DAY OF ADMISSION, S. AUREUS WAS IDENTIFIED IN A BLOOD CULTURE AND THE PATIENT DIED SEVEN DAYS LATER. COMPARING G2 AND G3, THE EVOLUTION OF THE MEDIAN LEVELS OF AMMONIA EXHIBITED OPPOSITE DIRECTIONS: AN INCREASING TENDENCY FOR G2 AND A DECREASING TENDENCY FOR G3 (P = 0.0002). ON THE OTHER HAND, THE LEVELS OF FV SHOWED A SIMILAR INCREASING TENDENCY IN G2 AND G3, HOWEVER, WITH STATISTICAL SIGNIFICANCY (P = 0.014). (FIGURE 2A,B). THREE DAYS AFTER TPE (DAY 4), THE SERUM YF VIRAL LOAD BECAME UNDETECTABLE IN 91.7% OF G3 COMPARED TO 28.6% OF G2. THE HIGH MORTALITY RATES OF SEVERE YF FORMS DURING THE YF OUTBREAK IN SP DESPITE INTENSIVE CARE SUPPORT PROMPTED US TO USE HV-TPE FOR YF-ALF. IN OUR CASUISTIC, USING A SIMILAR LARGE-VOLUME EXCHANGE STRATEGY, WE COULD NOT OBSERVE ANY BENEFIT IN MORTALITY. G1 AND G2 HAD SIMILAR MORTALITY RATES DESPITE THE USE OF TPE. WE IMPUTE THIS FINDING TO THE FULMINANT ONSET AND AGGRESSIVE COURSE OF ALF-YF, FASTER THAN THAT OBSERVED IN OTHER ETIOLOGIES. THIS HYPOTHESIS WAS CORROBORATED BY THE SHORT SURVIVAL OF G1 PATIENTS AND BY THE TRANSIENT IMPROVEMENT OF LABORATORY RESULTS OBSERVED AFTER HV-TPE IN G2. THE CLOSE RELATIONSHIP BETWEEN ELEVATED AMMONIA LEVELS AND THE DEVELOPMENT OF ENCEPHALOPATHY IS WELL ESTABLISHED. DURING THE FIRST MONTH OF THE OUTBREAK, WE OBSERVED THAT A PRESENCE OF AMMONIA > 70 MCMOL/L AND/OR FV < 50% WERE ASSOCIATED WITH HIGHER MORTALITY. AMONG 59 PATIENTS WITH THIS LABORATORY PROFILE, THE FIRST 48 WERE TREATED WITH ICS OR LT, AND THIRTY-NINE (81%) OF THEM DIED. CLEMMESEN ET AL., USING HV-TPE IN 11 PATIENTS WITH ALF AND HEPATIC ENCEPHALOPATHY GRADE III-IV, DEMONSTRATED A REDUCTION OF ARTERIAL AMMONIA. NEVERTHELESS, IN OUR CASUISTIC USING A SIMILAR STRATEGY, WE OBSERVED AN INCREASING TREND OF AMMONIA LEVELS INSTEAD. IN LARSEN¿S STUDY, AN INITIAL SIGNIFICANT DECREASE OF AMMONIA LEVELS WAS OBSERVED ALTHOUGH IT WAS NOT SUSTAINED AFTER DAY THREE. IN ORDER TO OVERCOME THE AGGRESSIVE BEHAVIOR OF ALF-YF AND AIMING TO OBTAIN SAFE LEVELS OF AMMONIA, A LOWER VOLUME OF TPE PER SESSION AND A REDUCTION OF THE INTERVAL BETWEEN THE PROCEDURES WERE IMPLEMENTED (INTENSIVE TPE). THIS APPROACH ALLOWED A SUSTAINED DECREASE OF AMMONIA (FIGURE 2A, G3), RESULTING IN A HIGHER SURVIVAL RATE. THE STUDIES ON HV-TPE IN HEPATOTROPIC VIRAL DISEASES ARE LIMITED. LARSEN ET AL. INCLUDED ONLY 11 PATIENTS WITH ACUTE VIRAL HEPATITIS, AND SIX OF THEM SUBMITTED TO HV-TPE [13]. ONLY TWO AMONG 11 PATIENTS TREATED BY CLEMMESEN ET AL. HAD VIRAL HEPATITIS. LIKEWISE, AKDOGAN ET AL., USING A STANDARD PLASMA EXCHANGE PROTOCOL, HAD ONLY 20% (8/39) CASES OF VIRAL HEPATITIS IN THEIR CASUISTIC. UNLIKE DRUG-INDUCED OR TOXIC HEPATITIS, IN VIRAL ETIOLOGIES THE CAUSAL AGENT IS STILL PRESENT AND THE LIVER INJURY PERSISTS. DESPITE THE FACT THAT COMMON KNOWLEDGE PRESUMES A VIREMIA OF FIVE DAYS IN YF, RECENT STUDIES DEMONSTRATED LONGER PERSISTENCE OF THE VIRUS IN THE SERUM AND OTHER ORGANS. ADDITIONALLY, IN A COHORT OF THE 2018 SP OUTBREAK, A HIGHER VIRAL LOAD WAS DEMONSTRATED TO BE AN INDEPENDENT PREDICTOR FACTOR ASSOCIATED WITH DEATH. AND, AS WE HAVE SHOWN, INTENSIVE TPE WAS ABLE TO REMOVE THE VIRUS, THUS REDUCING LIVER DAMAGE, ANOTHER POSSIBLE REASON TO JUSTIFY THE SUCCESS OF INTENSIVE TPE IN SEVERE YELLOW FEVER. ALTHOUGH TRANSAMINASE LEVELS IN G3 PATIENTS WERE LOWER THAN IN G2, THEY REMAINED MORE THAN 100 TIMES ABOVE THE UPPER LIMIT OF NORMAL, UNDERSCORING THE SEVERITY OF THE DISEASE. ADDITIONALLY, STUDIES HAVE INCREASINGLY DEMONSTRATED THAT THE YELLOW FEVER VIRUS EXHIBITS TROPISM FOR ORGANS BEYOND THE LIVER, WITH RNA OR ANTIGENS DETECTABLE IN THE MYOCARDIUM, LUNGS, KIDNEYS, PANCREAS, AND SPLEEN. FURTHERMORE, COAGULOPATHY HAS BEEN SHOWN TO OCCUR INDEPENDENTLY OF HEPATOCELLULAR TROPISM. THESE FINDINGS HIGHLIGHT THAT THE INVOLVEMENT OF MULTIPLE ORGANS SIGNIFICANTLY CONTRIBUTES TO THE SEVERITY OF THE DISEASE, AND THEREFORE SHOW THE ROLE OF TPE IN RESTORING METABOLIC AND IMMUNOLOGICAL BALANCE AND REPLENISHING COAGULATION FACTORS. INTENSIVE TPE APPARENTLY WAS SAFER WITH A LOWER INCIDENCE OF HYPOCALCEMIA COMPARED TO HV-TPE. IT WAS ALSO ASSOCIATED WITH A REDUCED MEDIAN VOLUME TREATED AND SHORTER TIME PER SESSION, ALLOWING SEQUENTIAL TREATMENT FOR A MAJOR NUMBER OF PATIENTS. THE POSSIBILITY OF SIMULTANEOUS TREATMENT FOR A LARGE NUMBER OF PATIENTS IS PARTICULARLY IMPORTANT CONSIDERING THE DRAMATIC INCREASE OF YF IN REGIONS WITH HEALTH RESOURCE CONSTRAINTS, AND THE SEASONAL PATTERN OF THE DISEASE IN OUTBREAKS. LT IS A RECOGNIZED APPROACH FOR THE MANAGEMENT OF ALF, HOWEVER THERE IS VERY RESTRICTED EXPERIENCE IN ITS USE AS A TREATMENT FOR ALF-YF. MOREOVER, THE HIGH COST OF LT AND THE SHORTAGE OF LIVER DONORS CURTAIL THE FEASIBILITY OF THE PROCEDURE. SOME LIMITATIONS OF OUR STUDY INCLUDE THE SMALL SAMPLE SIZE AND THE NEED FOR RANDOMIZED TRIALS TO CONFIRM THE EFFICACY OF THIS STRATEGY. HOWEVER, GIVEN THE FULMINANT PROGRESSION AND HIGH LETHALITY OF THE DISEASE, AN ADAPTIVE PRAGMATIC TRIAL COULD REPRESENT A FEASIBLE APPROACH TO DEMONSTRATE THIS TREATMENT¿S EFFECTIVENESS. WE DEMONSTRATED THAT IN PATIENTS WITH ALF SECONDARY TO YF, INTENSIVE TPE ASSOCIATED WITH GUIDED TRANSFUSION IS A FEASIBLE AND EFFECTIVE ARTIFICIAL SUPPORT TO KEEP PATIENTS ALIVE UNTIL HEPATIC REGENERATION AND RECOVERY OCCURS. THIS WAS DEMONSTRATED BY A REDUCTION OF 84% IN MORTALITY. THEREFORE, IT IS IMPORTANT TO SHARE OUR SUCCESSFUL EXPERIENCE, ALTHOUGH LIMITED TO A SMALL NUMBER OF CASES. UNDOUBTEDLY THERE ARE MANY KNOWLEDGE GAPS RELATED TO THE EFFECTS OF THERAPEUTIC PLASMA EXCHANGE IN THE COURSE OF YF AND OUR LIMITATIONS OF A SMALL SAMPLE SIZE, AND THE OBSERVATIONAL AND SINGLE-CENTER NATURE OF THIS STUDY. FURTHER STUDIES ARE REQUIRED TO BETTER UNDERSTAND THE ROLE OF TPE IN SEVERE YF. SINCE THIS WAS A JOURNAL PUBLICATION AN OBSERVATIONAL STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK, THE LOT NUMBERS WERE NOT REQUESTED; THEREFORE, A DISPOSABLE LOT HISTORY SEARCH COULD NOT BE CONDUCTED. ACCORDING TO THERAPEUTIC APHERESIS: A PHYSICIAN'S HANDBOOK, ADVERSE EVENTS OCCUR DURING THERAPEUTIC PROCEDURES WITH A FREQUENCY OF 4.8%. SYMPTOMS OF THESE ALLERGIC REACTIONS MAY INCLUDE HIVES, DYSPNEA, WHEEZING, BURNING EYES, TACHYCARDIA, HYPOTENSION, AND OR FACIAL SWELLING AND FLUSHING. MILD REACTIONS CAN BE TREATED WITH DIPHENHYDRAMINE ADMINISTERED THROUGH AN IV. ACCORDING TO THERAPEUTIC APHERESIS: A PHYSICIAN'S HANDBOOK, ADVERSE EVENTS OCCUR DURING THERAPEUTIC PROCEDURES WITH A FREQUENCY OF 4.8%. SOME OF THE MOST COMMON REACTIONS INCLUDE FEVER, URTICARIA, HYPOCALCEMIC SYMPTOMS, PRURITUS, DYSPNEA, TACHYCARDIA, AND MILD HYPOTENSION. TRANSIENT HYPOCALCEMIA ASSOCIATED WITH APHERESIS IS USUALLY WELL TOLERATED. SYMPTOMS OFTEN SHOW AS PARESTHESIA (TINGLING) BUT PATIENTS MAY ALSO EXPERIENCE UNUSUAL TASTE, NAUSEA, LIGHTHEADEDNESS, SHIVERING, AND TREMORS. SEVERE HYPOCALCEMIA MAY ALSO CAUSE MUSCLE CONTRACTIONS AND CAN PROGRESS TO TETANY AND SEIZURES IF HYPOCALCEMIA ESCALATES AND IS NOT CORRECTED. SINCE THIS WAS A JOURNAL PUBLICATION AN OBSERVATIONAL STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK, THE LOT NUMBERS WERE NOT REQUESTED; THEREFORE, A DHR SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA FOR RELEASE. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED. HO, Y., NUKUI, Y., VILLAÇA, P. R., OKAZAKI, E., TATSUI, N. H., NETTO, L. C., JOELSONS, D., DA ROCHA, T. R. F., DE MELLO MALTA, F., PINHO, J. R. R., SEGURADO, A. A. C., & ROCHA, V. (2025). INTENSIVE THERAPEUTIC PLASMA EXCHANGE¿NEW APPROACH TO TREAT AND RESCUE PATIENTS WITH SEVERE FORM OF YELLOW FEVER. TROPICAL MEDICINE AND INFECTIOUS DISEASE, 10(2), 39. HTTPS://DOI.ORG/10.3390/TROPICALMED10020039. Manufacturer Narrative: THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION IN B.5, H.1, H.6 AND H.11 INVESTIGATION: THIS IS AN OBSERVATIONAL CASE-SERIES UNICENTRIC STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK. IT WAS CONDUCTED AT HOSPITAL DAS CLINICAS (HCFMUSP), A 2200-BED TERTIARY CARE TEACHING HOSPITAL AFFILIATED TO THE UNIVERSITY OF SAO PAULO WITH 110 INTENSIVE CARE BEDS AND A 7-BED INFECTIOUS DISEASE INTENSIVE CARE UNIT (ID-ICU). ON 9 JANUARY 2018, THE ID-ICU HAD BECOME THE CENTER OF EXPERTISE FOR SEVERE FORMS OF YF IN SP. SEVERITY CRITERIA FOR ID-ICU ADMISSION WERE: ASPARTATE AMINOTRANSFERASE (AST) OR ALANINE AMINOTRANSFERASE (ALT) > 3000 U/L; INR > 1.5; PLATELET COUNT < 90 × 109/L; KIDNEY DYSFUNCTION; HEMORRHAGIC MANIFESTATIONS; ENCEPHALOPATHY; OR HEMODYNAMIC INSTABILITY. HOWEVER, IN 2019, A YELLOW FEVER OUTBREAK OCCURRED IN A REGION OF THE STATE OF SAO PAULO WHERE TRANSPORTATION TO HCFMUSP TAKES AT LEAST 6 H. FOR THIS REASON, ALL PATIENTS WHO HAD ALT/AST > 500 WERE REFERRED TO HCFMUSP. FROM 9 JAN TO 15 FEB 2018, TREATMENT CONSISTED OF STANDARD INTENSIVE CARE SUPPORT (ICS). IN THE SECOND HALF OF FEBRUARY 2018, HIGH-VOLUME TPE (HV-TPE) WAS USED IN SOME PATIENTS WITH SEVERE YF IN OUR SERVICE. TRANSIENT IMPROVEMENT OF LABORATORY RESULTS WAS OBSERVED, HOWEVER HEMORRHAGIC PHENOMENA OCCURRED BETWEEN SESSIONS OF HV-TPE WITHOUT CHANGING THE OUTCOME. SINCE APRIL 2018, THIS STRATEGY WAS MODIFIED TO MORE INTENSIVE TPE TWICE A DAY AND SUPPORTIVE TRANSFUSION, IF NECESSARY. THIS INTENSIVE PROTOCOL WAS USED ON ONE PATIENT IN 2018 SUCCESSFULLY, AND SINCE THEN IT HAS BEEN USED IN THE STATE OF SAO PAULO¿S 2019 YF OUTBREAK. TPE WAS ADMINISTERED TO PATIENTS WITH LEVELS OF AMMONIA > 70 MCMOL/L AND/OR FACTOR V (FV) < 50%. THE DEFINITION OF THE CUT-OFF OF AMMONIA WAS BASED ON THE 2018 EXPERIENCE WHEN PATIENTS WITH AMMONIA > 70 MCMOL/L HAD A HIGHER RELATIVE RISK OF DEATH (RR 2.0) (DATA NOT PUBLISHED). THE CUT-OFF OF FV < 50% WAS DEFINED BASED ON CLICHY¿S CRITERIA, CONSIDERING A REDUCTION OF 30% FROM THE REFERENCE VALUE OF 70¿150% FOR ALL PATIENTS. IN ORDER TO COMPARE TO PREVIOUS CASES, ALL PATIENTS WITH AMMONIA > 70 MCMOL/L AND/OR FV < 50% WERE SELECTED. FOR THIS ANALYSIS, PATIENTS WERE CLASSIFIED INTO THREE GROUPS ACCORDING TO TYPE OF TREATMENT GIVEN: GROUP ONE (G1)¿ICS; GROUP TWO (G2)¿ICS AND HV-TPE; AND GROUP THREE (G3)¿ICS AND INTENSIVE TPE. CLINICAL AND LABORATORY DATA WERE EXTRACTED RETROSPECTIVELY AND STORED IN THE REDCAP WEB PLATFORM. HV-TPE STRATEGY (G2): PROPOSED TREATMENT CONSISTED OF THREE CONSECUTIVE DAYS OF PLASMA EXCHANGE. THE VOLUME CHANGED WAS 10% OF BODY WEIGHT (BW) WITH FRESH FROZEN PLASMA (FFP). ONE TO THREE ADDITIONAL SESSIONS OF ONE PLASMA VOLUME EACH WERE IMPLEMENTED ACCORDING TO AMMONIA LEVELS. SUPPORTIVE TRANSFUSION WAS GUIDED MAINLY BY HEMORRHAGIC EPISODES. INTENSIVE-TPE STRATEGY (G3): VOLUME OF EXCHANGE WITH FFP WAS CALCULATED ACCORDING TO BW, HEMATOCRIT (HT), AND PLASMA VOLUME (BW × 0.7 × [1 - HT]), CONSIDERING ONE PLASMA VOLUME PER SESSION. PROCEDURE WAS PERFORMED TWICE A DAY, WITH ADDITIONAL FFP INFUSION (10 ML/KG) AT NIGHT. PLATELET CONCENTRATE WAS TRANSFUSED TO KEEP PLATELET COUNT > 30 × 109/L OR, IN THE PRESENCE OF BLEEDING, >50 × 109/L. IN SEVERE OR LIFE-THREATENING BLEEDS, TARGET WAS >80¿100 × 109/L. CRYOPRECIPITATE TRANSFUSION WAS INDICATED IF FIBRINOGEN LEVEL WAS <100 MG/DL. BASED ON BLEEDING AS WELL AS FV AND/OR AMMONIA LEVELS, THE FREQUENCY OF TPE WAS TAPERED TO ONCE A DAY, EVERY OTHER DAY AND DISCONTINUATION. DURING THE DE-ESCALATION PERIOD, FFP WAS INFUSED (5 ML/KG) 1¿2 TIMES/DAY, FOLLOWING THE SAME CRITERIA ABOVE. TRANEXAMIC ACID (10 MG/KG/DOSE, IV, EVERY 6¿8 H) WAS PRESCRIBED UNTIL DISCONTINUATION OF TPE AND TOPICAL USE WAS IMPLEMENTED IN CASES OF MUCOCUTANEOUS BLEEDING. TPE WAS PERFORMED USING FFP REPLACEMENT, WITHOUT ANTICOAGULANT (CITRATE), ONCE ALL PATIENTS HAD ALF. CALCIUM CHLORIDE WAS ADMINISTERED AT THE FOLLOWING RATES: FOR HV-TPE GROUP, 5 MMOL PER EACH LITER OF PLASMA EXCHANGE; FOR INTENSIVE-TPE GROUP, FIXED 6.8 MMOL OF CALCIUM PER SESSION. AFTER EACH SESSION, SERUM CALCIUM LEVEL WAS ASSESSED, AND ADDITIONAL SUPPLEMENTATION WAS CONSIDERED. COBE SPECTRA AND OPTIA (TERUMOBCT, TOKYO, JAPAN) AND COM.TECR (FRESENIUS KABI, BAD HOMBURG, GERMANY) TPE DEVICES WERE EMPLOYED, USING A WITHDRAWAL FLOW OF 70 ML/MINUTE ON AVERAGE. IN CASES WHERE CONTINUOUS VENOVENOUS HEMODIALYSIS WAS INDICATED, THE PLASMA EXCHANGE SYSTEM WAS CONNECTED TO THE RETURN CIRCUIT OF THE DIALYSIS SYSTEM. YFV-RNA WAS ISOLATED FROM 250 ¿L OF SERUM SAMPLES AND 20 ¿G OF RNASE-FREE GLYCOGEN WERE ADDED TO THE AQUEOUS PHASE AS AN RNA CARRIER. REAL-TIME QUANTITATIVE YFV REVERSE TRANSCRIPTION (RT-QPCR) WAS CARRIED OUT USING SUPERSCRIPTR III ONE-STEP RT-QPCR SYSTEM WITH PLATINUMR TAQ DNA POLYMERASE (INVITROGEN, THERMO FISHER SCIENTIFIC BRAND, CARLSBAD, CA, USA). THE YFV 5'NON-CODING REGION WAS TARGETED USING SPECIFIC PRIMERS AND PROBE, AS PREVIOUSLY DESCRIBED. ALL SAMPLES WERE TESTED IN TRIPLICATE. THE PRIMARY END POINT WAS MORTALITY. BASELINE CLINICAL AND LABORATORY ASPECTS WERE COMPARED AMONG GROUPS. SURVIVAL CURVES WERE CONSTRUCTED USING THE KAPLAN¿MEIER METHOD, AND THE LOG-RANK TEST WAS USED TO ASSESS DIFFERENCES BETWEEN CURVES. PROPORTIONAL HAZARD ASSUMPTION WAS CHECKED GRAPHICALLY BY LOG-MINUS-LOG VS. LOG (TIME) PLOT. COMPARISONS BETWEEN TWO GROUPS WERE ASSESSED USING THE STUDENT T-TEST, OR THE MANN¿WHITNEY TEST IF APPROPRIATE FOR QUANTITATIVE VARIABLES, AND THE FISHER EXACT TEST FOR CATEGORICAL VARIABLES. ASSUMPTIONS OF NORMALITY AND HOMOGENEITY OF VARIANCE FOR T-TEST WERE EVALUATED. DIFFERENCES BETWEEN THE 3 GROUPS IN QUANTITATIVE VARIABLES WERE EVALUATED BY THE KRUSKAL¿WALLIS TEST. COMPARISON OF PROPORTIONS BETWEEN THE 3 GROUPS WAS PERFORMED BY USING THE FISHER-EXACT TEST. QUANTITATIVE VARIABLES WERE PRESENTED AS MEDIAN (INTERQUARTILE RANGE [IQR]). CATEGORICAL VARIABLES WERE EXPRESSED AS COUNTS AND PERCENTAGES. NORMALITY WAS ASSESSED BY THE SHAPIRO¿WILK NORMALITY TEST, THE D¿AGOSTINO¿PEARSON OMNIBUS NORMALITY TEST, AND VISUAL INSPECTION OF HISTOGRAMS. ALL P-VALUES WERE TWO-SIDED. RESULTS WERE CONSIDERED SIGNIFICANT IF P < 0.05. STATISTICAL ANALYSES WERE PERFORMED USING SAS VERSION 9.4 (SAS INSTITUTE INC., CARY, NC, USA). DURING THE 2018/2019 YELLOW FEVER OUTBREAK IN SÃO PAULO, 114 CONFIRMED CASES OF YELLOW FEVER WERE REFERRED TO THE ID-ICU OF HCFMUSP. SIXTY-SIX PATIENTS HAD LEVELS OF AMMONIA > 70 MCMOL/L AND/OR AN FV OF <50% AND WERE INCLUDED FOR ANALYSIS: 41 (62%) WERE TREATED WITH ICS (G1), 11 (17%) RECEIVED ICS PLUS HV-TPE (G2) AND 14 (21%) RECEIVED ICS PLUS INTENSIVE TPE (G3). THERE WERE NO DIFFERENCES AMONG THE THREE GROUPS RELATED TO SEX AND CLINICAL PRESENTATION AT ADMISSION. ALL GROUPS HAD VALUES OF AST/ALT > 3000 U/L, HOWEVER PATIENTS OF G3 HAD LOWER LEVELS OF AST/ALT COMPARED TO THE OTHER GROUPS, PROBABLY AS A CONSEQUENCE OF THE CHANGE IN THE TRANSFER CRITERIA, WHICH ALLOWED EARLIER ASSISTANCE, PREVENTING THE PROGRESSION OF THE DISEASE. NO DIFFERENCES WERE OBSERVED AMONG THE GROUPS WITH RESPECT TO THE OTHER LABORATORY TESTS. CONCERNING G3, THE PATIENTS WERE ADMITTED WITH A MEDIAN OF SIX DAYS OF SYMPTOMS AND FEVER WAS NOTED IN 86% OF THEM. SURPRISINGLY, JAUNDICE WAS PRESENT IN ONLY 43% OF THE PATIENTS AT ADMISSION. HEMORRHAGIC PHENOMENA OCCURRED IN 11 PATIENTS (79%), THE VAST MAJORITY BEING GASTROINTESTINAL BLEEDING. TWELVE PATIENTS (86%) HAD SEIZURES AND/OR HEADACHES AS NEUROLOGIC MANIFESTATIONS. ONE OF THE FEMALE PATIENTS (LO) WAS PREGNANT AND WAS DIAGNOSED WITH REACTIVE HEMOPHAGOCYTIC SYNDROME, WHICH WAS SUCCESSFULLY TREATED WITH HUMAN INTRAVENOUS IMMUNOGLOBULIN. THE MEDIAN VOLUME TREATED AND TIME PER PROCEDURE WERE HIGHER IN G2 (P = 0.005 AND P < 0.001, RESPECTIVELY) ALTHOUGH G3 HAD A HIGHER MEDIAN NUMBER OF PROCEDURES PER PATIENT (P < 0.001) (TABLE 1). HOWEVER, 45% OF THE G2 PATIENTS DIED BEFORE CONCLUDING THE PROPOSED TREATMENT (THREE CONSECUTIVE DAYS). TRANSFUSION REACTIONS SUCH AS TRANSFUSION-RELATED ACUTE LUNG INJURY, TRANSFUSION ASSOCIATED CIRCULATORY OVERLOAD OR ABO INCOMPAT... Edited for length. See MAUDE Adverse Event Report for complete event narratives.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857888&pc=MNS | 2518422-2025-043986 | 2022-10-26 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | MNS | DREAMSTATION BIPAP AUTOSV | Degraded | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRMED FOAM PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21857915&pc=NJK | 3007797756-2025-00019 | 2025-02-28 05:00 | Death | PULMONX CORPORATION | 2025-04-17 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P180002 |  | 1 | Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES ON (B)(6) 2025. THE PATIENT WAS TRANSFERRED TO ICU POST-PROCEDURE AND THE VALVES WERE REMOVED IN THE FOLLOWING DAYS AFTER THE PROCEDURE. IT WAS MADE AWARE ON (B)(6) 2025 THAT THE PATIENT HAD PASSED AWAY WHILE IN RECOVERY.Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES (ONE VALVE IN THE RIGHT MIDDLE LOBE (RML) AND THREE IN THE RIGHT UPPER LOBE (RUL) ON (B)(6) 2025. PER THE PHYSICIAN, THE PROCEDURE DURATION LASTED FOR ABOUT ONE HOUR. ON (B)(6) 2025, ALL FOUR VALVES WERE EXPLANTED DUE TO A DECLINE IN THE PATIENT'S RESPIRATORY STATUS FOLLOWING THE PROCEDURE WHILE IN THE ICU. THE VALVES REMAINED IMPLANTED FOR A TOTAL OF 3 DAYS. ON (B)(6) 2025, IMMEDIATELY FOLLOWING THE BRONCHOSCOPY, THE PATIENT EXPERIENCED A FAILED EXTUBATION DUE TO CO[?] RETENTION AND HYPERSOMNOLENCE, REQUIRING RE-INTUBATION AND TRANSFER TO THE ICU. DESPITE VARIOUS VENTILATOR ADJUSTMENTS, THE PATIENT DEMONSTRATED VERY POOR AIR MOVEMENT. THE VALVES WERE REMOVED AFTER THREE DAYS; HOWEVER, THE PATIENT'S CONDITION DID NOT IMPROVE. THE PATIENT CONTINUED TO DETERIORATE AND SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025. PER THE MEDICAL CHART, THE FOLLOWING STATEMENT WAS DOCUMENTED: "DUE TO PERSISTENCE OF VENTILATOR DEPENDENCE AND INABILITY TO WEAN PATIENT SUCCESSFULLY FROM PARALYSIS AND DEEP SEDATION, COUNSELED FAMILY THAT NEXT TREATMENT OPTION WAS TO UNDERGO TRACHEOSTOMY PLACEMENT VERSUS PURSUE COMFORT FOCUSED CARE. IN SPEAKING WITH MR. ***'S FAMILY, THEY RELAYED THAT TRACHEOSTOMY WAS NOT IN LINE WITH HIS GOALS OF CARE. AS SUCH, PATIENT WAS TRANSITIONED TO COMFORT FOCUSED CARE ON (B)(6) 2025 AND WAS PALLIATIVELY EXTUBATED AT THIS TIME. MR.*** PASSED PEACEFULLY SHORTLY THEREAFTER ON (B)(6) 2025 AT 20:30." ACCORDING TO THE PATIENT'S CHART AND MEDICAL RECORDS, THE PRIMARY CAUSE OF DEATH WAS ACUTE HYPOXEMIC HYPERCAPNIC RESPIRATORY FAILURE ATTRIBUTED TO SEVERELY LIMITED PULMONARY RESERVE AND THE EFFECTS OF GENERAL ANESTHESIA. A SECONDARY CONTRIBUTING FACTOR WAS A COPD EXACERBATION. NO AUTOPSY WAS PERFORMED. ACCORDING TO THE PHYSICIAN, THE VALVES WERE DISCARDED AS THEY WERE NOT BELIEVED TO BE RELATED TO THE PATIENT'S CLINICAL DETERIORATION OR CAUSE OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858041&pc=LYJ | 1644487-2025-00420 | 2025-03-19 04:00 | Death | CYBERONICS - HOUSTON | 2025-04-17 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: A LINKEDIN POST WAS FOUND WHICH NOTED THAT THE PATIENT PASSED AWAY DUE TO SUDEP. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858052&pc=MKJ | 1220908-2025-01449 | 2025-04-08 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-04-17 04:00 | MKJ | R SERIES | Patient Device Interaction Problem | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED WHILE DEFIBRILLATING A MALE PATIENT (AGE UNKNOWN), AFTER REMOVING THE ELECTRODE PADS, BURNS WERE FOUND ON THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL UNITED KINGDOM FOR EVALUATION. THE DEVICE AND MULTIFUNCTION CABLE (MFC) WERE TESTED AND FOUND TO BE FULLY FUNCTIONAL WITH NO DEVICE MALFUNCTIONS. LOG DATA FROM THE EVENT SHOWED LARGE DIFFERENCES BETWEEN INITIAL AND PRE-SHOCK IMPEDANCE READINGS, WHICH INDICATES POOR ELECTRODE CONTACT WITH THE PATIENT'S SKIN. THIS IS THE LIKELY CAUSE FOR THE ENERGY OUTPUT DELIVERED BASED ON MEASURED IMPEDANCE AT THE TIME. THE RETURNED PADS SHOWED NO SIGNS OF BURNS. BASED ON TESTING AND DATA REVIEW, THE ISSUE IS DUE TO POOR PAD COUPLING. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858067&pc=NPT | 2135147-2025-02008 | 2025-03-25 04:00 | Death | ABBOTT MEDICAL | 2025-04-17 04:00 | NPT | FLEXNAV DELIVERY SYSTEM | Difficult to Advance | Ventricular Fibrillation; Heart Block; Aortic Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VISON VALVE WAS CHOSEN FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE PERIMETER OF THE NATIVE VALVE WAS 76.6MM AND LEFT VENTRICULAR OUTFLOW TRACT (LVOT) WAS 20.1MM. DURING PROCEDURE, THE RIGHT FEMORAL ACCESS IN PATIENT WITH SEVERE PERIPHERAL VASCULAR DISEASE AND PERCUTANEOUS TRANSLUMINAL PERIPHERAL ANGIOPLASTY (PTA) OF AN ILIAC. THERE WAS EXTREMELY DIFFICULT TO ADVANCE THE DELIVERY SYSTEM THROUGH THE NON-ABBOTT INTRODUCER, VERY DIFFICULT TO NAVIGATE THE FLEXNAV TROUGH THE TORTUOUS AORTA AND TO PASS THE AORTIC ARCH AND VALVE. THE PATIENT DEVELOPED A TOTAL ATRIOVENTRICULAR (AV) WITH EVOLUTION TO A VENTRICULAR FIBRILLATION (VFIBRILLATION). A TEMPORARY PACEMAKER WAS INSERTED. THERE WAS NO PERICARDIAL EFFUSION OR AORTIC (AO) DISSECTION WERE NOTED BY TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). AFTER PRE-DILATATION A MASSIVE AORTIC INSUFFICIENCY (AI) WAS NOTED. AFTER 45 MINUTES OF REANIMATION THE PATIENT WAS REPORTED PASSED AWAY. THE VALVE WAS NOT IMPLANTED. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS PROCEDURE AND COMORBIDITIES.Event Description: N/A. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. INFORMATION FROM THE FIELD INDICATED THAT THERE WAS EXTREMELY DIFFICULT TO ADVANCE THE DELIVERY SYSTEM THROUGH THE NON-ABBOTT INTRODUCER, VERY DIFFICULT TO NAVIGATE THE FLEXNAV TROUGH THE TORTUOUS AORTA AND TO PASS THE AORTIC ARCH AND VALVE. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON ALL AVAILABLE INFORMATION, THE REPORTED DIFFICULT ADVANCEMENT AND INSERTION DIFFICULTIES INTO THE PATIENT APPEAR TO BE RELATED TO CHALLENGING PATIENT ANATOMY (TORTUOUS ANATOMY). THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FROM MEDICAL REVIEW: "THE CAUSE OF DEATH THEREFORE DOES NOT APPEAR TO BE RELATED TO THE NAVITOR VALVE BUT IS A PROCEDURAL COMPLICATION RELATED TO RHYTHM CHANGE IN A MULTIMORBID PATIENT. IF ADDITIONAL IMAGING BECOMES AVAILABLE AN AMENDED MEDICAL REVIEW WILL BE PROVIDED."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858245&pc=MVK | 3008642652-2025-03984 | 2025-03-17 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-17 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED NINE APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 23:14:17 ON (B)(6) 2025. AT 12:48:49 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 150 BPM. AT 12:50:04, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 160 BPM. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 50 BPM TRANSITIONING TO VT @ 130 BPM. AT 12:51:21, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 150 BPM DEGRADING TO VF WITH CPR/MOTION ARTIFACT. AT 12:51:53, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 50 BPM DEGRADING TO VF. AT 12:52:24, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM. AT 12:53:00, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 40 BPM DEGRADING TO VF. AT 12:53:32, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 12:54:06, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 13:04:32, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 13:06:52, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 13:07:33, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 13:08:00, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 30 BPM. THE DEVICE WAS SHUT DOWN AT 14:21:27 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858334&pc=GDW | 3005075853-2025-02733 | 2024-08-16 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-17 04:00 | GDW | ECHELON 3000 45 STANDARD | Patient Device Interaction Problem | Cardiac Arrest | K213633 |  | 1 | Event Description: IT WAS REPORTED THAT, DURING A THORACOSCOPIC RIGHT UPPER PARTIAL PNEUMONECTOMY, THERE WERE NO PROBLEMS DURING THE OPERATION ON (B)(6) 2024, BUT THE PATIENT DIED TWO DAYS AFTER THE OPERATION (B)(6) 2024. THE DEVICE WAS USED 4-5 FIRINGS ON RIGHT UPPER LOBE OF LUNG. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO CARDIAC TAMPONADE ON LEFT HEMOTHORAX. THE RESULTS OF PATHOLOGICAL AUTOPSY ON (B)(6) 2025, THERE WAS A 1MM HOLE ON THE OUTSIDE OF THE AORTA. NO DEVICE WILL BE RETURNING. NO FURTHER INFORMATION IS AVAILABLE. AFFILIATE REFERENCE NUMBER: (B)(4). Manufacturer Narrative: (B)(4). DATE SENT: 5/13/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: IS THERE AN AUTOPSY REPORT AVAILABLE? NOT CLEAR. IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? NOT CLEAR. WAS THERE BLOOD IN PERICARDIUM POST MORTEM EXAM? YES. HOW MUCH BLOOD WAS IN THE LEFT HEMOTHORAX ON POST MORTEM EXAM? N/A WAS THERE ANY EVIDENCE OF ANY BLOODING IN THE 1MM HOLE IN THE AORTA DURING THE POST MORTEM EXAM? BY A DEATH AUTOPSY, THE WHOLE WAS CONFIRMED. WHERE THERE ANY OTHER INSTRUMENTS USED IN THE AREA? N/A DOES THE SURGEON BELIEVE THERE IS A DESIGN OR MANUFACTURING DEFICIENCY WITH THE DEVICE THAT WOULD OF CAUSED OR CONTRIBUTED TO THE DAMAGE ON THE AORTA? THE THOUGHT IS THAT IF THERE WAS EXTERNAL DAMAGE, IT WOULD LIKELY BE DUE TO A STAPLE. IF SO PLEASE PROVIDE DETAILS. ARE ANY VIDEOS FROM THE PROCEDURE AVAILABLE? YES, ARE THERE ANY PICTURES FROM THE POST MORTEM OF THE HOLE? NOT CLEAR. IS A COPY OF THE AUTOPSY AVAILABLE? NOT CLEAR YET. Manufacturer Narrative: (B)(4). DATE SENT: 4/17/2025. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE PRIMARY DISEASE IS LUNG CANCER. WHERE WAS THE STAPLE LINE FINALLY CREATED? IT SEEMS THAT PARTIAL RESECTION OF S3 WAS PERFORMED. THE SURGEON SAID THAT THERE WERE NO PROBLEMS SUCH AS MALFORMATION OF THE STAPLES OR BLEEDING AT THE AREA WHERE THE PARTIAL RESECTION WAS PERFORMED. IT SEEMS THAT THE STAPLER MADE A HOLE ON THE OUTSIDE OF THE AORTA, BUT A DIRECT CAUSAL RELATIONSHIP HAS NOT BEEN PROVEN, AS A RESULT OF THE AUTOPSY, THERE WERE TWO HOLES WHICH ABOUT 1 MM IN SIZE ON THE RIGHT SIDE OF THE AORTA CARDIAC TAMPONADE MADE BY THE BLOOD POOLED BETWEEN THE PERICARDIUM AND THE HEART, AND THE PRESSURE WAS APPLIED TO THE PERICARDIUM. IT WAS ASSERTED FROM THE WAY THE HOLE WAS CREATED THAT IT WAS CAUSED BY DAMAGE FROM THE OUTSIDE OF THE BLOOD VESSEL. (SAFETY MANAGEMENT DESCRIBED IT AS "ABRASION".) IT IS HIGHLY POSSIBLE THAT SOME KIND OF DAMAGE WAS INFLICTED AFTER THE SURGERY, AND IT IS SPECULATED THAT THE ONLY THING THAT COULD HAVE CAUSED THE DAMAGE WAS THE METAL STAPLER (METAL). THE HOLE OPENED IN THE WEAK STRENGTH PART OF THE PERICARDIUM (=ON THE LEFT SIDE), LEADING TO LEFT HEMOTHORAX. IN ADDITION, ACCORDING TO THE PATHOLOGICAL OPINION, THE EVIDENCE THAT THE HOLE WHICH ABOUT 1 MM IN SIZE ON THE RIGHT SIDE OF THE AORTA WAS CAUSED BY AN EXTERNAL DAMAGE IS THAT THE ADVENTITIA OF THE BLOOD VESSEL WAS TORN, THE HOLE WAS LARGE, AND THE HOLE BECAME SMALLER AS IT APPROACHED THE INTIMA, SO IT IS THINKING THAT IT IS HIGHLY POSSIBLE THE DAMAGE WAS INFLICTED FROM THE OUTSIDE OF THE AORTA. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858377&pc=DSQ | 2916596-2025-02382 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO PROGRESSIVE LIVER FAILURE UNDER MULTIPLE ORGAN FAILURE. THE PATIENT SUFFERED FROM A PROGRESSIVE MULTIPLE ORGAN FAILURE. SEVERE LIVER FAILURE WAS RELATED TO THE OUTCOME AT LEAST. NO DEVICE ISSUES PRIOR OUTCOME WAS MENTIONED. THERE WERE NO DEVICE RELATED ISSUES LEAD TO OUTCOME. THE DEVICE WAS RUNNING AS EXPECTED. THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) WAS NOT RETURNED FOR EVALUATION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE (INCLUDING HEPATIC FAILURE) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858600&pc=NVZ | 3008973940-2025-02717 | 2021-10-28 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-17 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED APPROXIMATELY TWO WEEKS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG). THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858625&pc=DSQ | 2916596-2025-02219 | 2025-02-15 05:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO THE PATIENT BEING ON PALLIATIVE CARE FROM RESPIRATORY FAILURE. IT WAS NOT PROVIDED IF THIS WAS DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED AND RETURNED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT DEATH INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RESPIRATORY FAILURE COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT PROVIDE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858745&pc=OZD | 1220648-2025-27616 | 2025-03-25 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Fluid/Blood Leak | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WAS A PURGE LEAK FROM THE SIDEARM AT THE LUER-CONNECTOR. A PURGE BYPASS WAS CREATED AND SUPPORT WAS CONTINUED. IT WAS FURTHER REPORTED THAT BASED UPON CONVERSATION WITH THE FAMILY, CARE WAS WITHDRAWN AND THE PATIENT SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF PURGE LEAK ¿ PUMP HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INVESTIGATION. SIDEARM BYPASS WAS OBSERVED ON THE RETURNED PRODUCT. THE SEVERED SIDEARM WAS INSPECTED, AND A CRACK ALONG THE LENGTH OF THE YELLOW LUER WAS NOTED. THE SIDEARM WAS CONNECTED TO THE PURGE FLUID VIA THE PURGE CASSETTE, AND THE LEAK WAS REPRODUCED. DATA LOG WAS NOT PROVIDED FOR CASE RUN. THE ROOT CAUSE OF THE PURGE LEAK - PUMP WAS DETERMINED TO BE DUE TO A DAMAGED SIDEARM YELLOW LUER DUE TO HIGH TENSION IN COMBINATION WITH DISINFECTANT AND CERTAIN DRUG INGREDIENTS SUCH AS OIL, ALCOHOL, LIPIDS ETC. Manufacturer Narrative: THE PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: CLEANING: ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858768&pc=PYX | 1220648-2025-27595 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | PYX | IMPELLA RP | Restricted Flow rate | Hemolysis; Thrombosis/Thrombus | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE MOTOR CURRENT ON THE IMPLANTED IMPELLA RP PUMP SUDDENLY SPIKED COINCIDING WITH A DROP IN FLOWS. AT THAT POINT IN TIME, THE PATIENT'S LABS WERE NOTED TO HAVE HEMOLYZED. THE FLOW ISSUES CONTINUED INTO THE FOLLOWING DAY AT WHICH POINT THE FLOW RATES FINALLY BEGAN TO NORMALIZE. DESPITE THE IMPROVEMENT, THE PATIENT'S PATHOLOGY RESULTS PROMPTED THE DECISION TO WITHDRAW CARE. THE PATIENT EXPIRED FOLLOWING REMOVAL OF THE PUMP. THERE ARE NO NOTES THAT THE PATIENT'S HEMOLYSIS HAD RESOLVED PRIOR TO THEIR PASSING. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND HEMOLYSIS HAS BEEN COMPLETED. THE DATA LOGS CONFIRMED FAILURE MODE WITH MOTOR CURRENT (MC) SPIKED FOLLOWED LOW PUMP FLOW THAT TRIGGERED AUTO SUCTION RESPONSE AND SUCTION ALARMS. THIS EVENT ONLY OCCURRED ABOUT 5 MINUTES ON 3/27 AND 15 MINUTES ON 03/28 THEN BACK TO NORMAL RANGE. NO OTHER ISSUE WAS FOUND ON THE LOGS. PRODUCT WAS RETURNED FOR REVIEW. VISUAL INSPECTION FOUND BIOMATERIAL INSIDE THE OF CAGE. NO OTHER ISSUES WERE FOUND ON THE REST OF THE PUMP. PUMP PASSED HEMOLYSIS TEST. THE ROOT CAUSE OF LOW FLOW WAS MOST LIKELY DUE TO BIOMATERIAL INGESTION SINCE THE BIOMATERIAL WAS FOUND ON THE PUMP, THE DATA LOGS MEET THE PATTERN OF LOW FLOW AND HEMOLYSIS AND THE PUMP HAD NO ISSUE UNDER INSPECTION VISUALLY. BASED ON THE INVESTIGATION FINDINGS THE FAILURE MODE THROMBUS WAS ADDED AND AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. D9 DEVICE RETURNED TO MANUFACTURER DATE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858770&pc=MVK | 3008642652-2025-03983 | 2025-02-27 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-17 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:40:40 ON (B)(6) 2025. AT 15:19:12 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION ARTIFACT. AT 15:19:47, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT. AT 15:20:14, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 20:18:36 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858862&pc=DSQ | 2916596-2025-02246 | 2025-03-24 04:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THERE WERE NO FURTHER DETAILS AND THE CAUSE WAS BEING DETERMINED. THE PUMP WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858968&pc=DSQ | 2916596-2025-02186 | 2024-09-12 04:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Intracranial Hemorrhage; Respiratory Failure; Ruptured Aneurysm; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025 WITH A SUBARACHNOID HEMORRHAGE (B)(6) AND ALTERED MENTAL STATUS AND WAS TREATED WITH KEPPRA (LEVETIRACETAM). WHILE ADMITTED THE PATIENT ALSO HAD MULTIPLE EPISODES OF BACTEREMIA DUE TO AN ONGOING INFECTION THAT STARTED (B)(6) 2024, WITH THE PATIENT HAVING FEVERS AND LEUKOCYTOSIS, EACH TREATED WITH INTRAVENOUS ANTIBIOTICS, WHICH RESOLVED THE INFECTION. THE PATIENT'S (B)(6) WAS FOUND TO HAVE BEEN CAUSED BY A THROMBOSED INFECTED INTRACRANIAL ANEURYSM, FOR WHICH NO NEUROLOGICAL INTERVENTION WAS RECOMMENDED. THEY LATER WENT INTO VENTILATOR-DEPENDENT RESPIRATORY FAILURE (VDRF), WAS TRACHEOSTOMY DEPENDENT, AND ALSO HAD RIGHT VENTRICULAR FAILURE (RVF), WHICH WAS TREATED WITH DOBUTAMINE. THE CAUSE OF THE RESPIRATORY FAILURE WAS UNKNOWN. IT WAS NOTED THAT THE PATIENT'S RVF EXISTED BEFORE LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION, AND THE RVF WAS NOT ATTRIBUTED TO A DEVICE ISSUE. THE PATIENT PASSED AWAY DUE TO MULTIPLE EPISODES OF CARDIOPULMONARY ARREST ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: DUE TO SYSTEM LIMITATIONS, THE FOLLOWING WERE NOT INCLUDED IN SECTION H6: 2551 - COGNITIVE CHANGES, 1858 - FEVER, 4694 - HIGH WHITE BLOOD CELL COUNT. CORRECTED DATA: SECTION H6: HEALTH EFFECT - IMPACT CODES AND HEALTH EFFECT - CLINICAL CODES CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS RESPIRATORY FAILURE, THROMBOEMBOLISM, INFECTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. SECTION 6 ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, AND OUTLINES INDICATIONS OF THROMBUS AS WELL AS HOW TO RESPOND TO SUCH EVENTS. THIS SECTION ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6 - ADDITIONAL CODES: 4868 - HEMODYNAMIC INSTABILITY, 1735 - BACTERIAL INFECTION, 4846 - BACTEREMIA, 4426 - UNSPECIFIED NERVOUS SYSTEM PROBLEM, 1858 - FEVER. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21858997&pc=DSQ | 2916596-2025-02248 | 2025-03-22 04:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Adult Respiratory Distress Syndrome | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO RESPIRATORY DISTRESS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS DEATH AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859091&pc=DSQ | 3007042319-2025-00779 | 2025-04-11 04:00 | Death | HEARTWARE, INC. | 2025-04-17 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Muscle Weakness; Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) PRESENTED TO THE EMERGENCY DEPARTMENT WITH DECREASED RESPONSIVENESS AND DIFFICULTY WALKING. ON EVALUATION, THE PATIENT WAS FOUND TO BE DIFFICULT TO AROUSE. A COMPUTED TOMOGRAPHY (CT) SCAN REVEALED AN INTRACRANIAL HEMORRHAGE. INITIAL LABORATORY RESULTS SHOWED AN INTERNATIONAL NORMALIZED RATIO (INR) OF 1.8, INDICATING SUBTHERAPEUTIC ANTICOAGULATION. ANTICOAGULATION WAS REVERSED USING VITAMIN K AND PROTHROMBIN COMPLEX CONCENTRATE (PCC). THE PATIENT SUBSEQUENTLY UNDERWENT A DECOMPRESSIVE HEMICRANIECTOMY TO RELIEVE INTRACRANIAL PRESSURE. DESPITE THESE INTERVENTIONS, THE PATIENT¿S CONDITION CONTINUED TO DETERIORATE. ULTIMATELY, THE DECISION WAS MADE TO TRANSITION TO COMFORT MEASURES, AND VAD SUPPORT WAS WITHDRAWN. THE PATIENT SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED NO ANOMALIES WITHIN THE ANALYZED PERIOD. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING AN INTRACRANIAL HEMORRHAGE AND THE VAD SUPPORT WAS WITHDRAWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION, BLEEDING AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF BLEEDING EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859098&pc=OZD | 1220648-2025-27627 | 2025-03-25 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT. DURING SUPPORT IT WAS REPORTED THAT THE RIGHT LEG WAS MOTTLED. HYPOTENSION CONTINUED THE NEXT DAY AND THE RIGHT LEG WAS VERY MOTTLED. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT'S LACTATE CONTINUED TO RISE. FULMINANT ORGAN FAILURE, AND ACIDOSIS INHIBITING MECHANISM OF ACTION OF MEDICATIONS. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS THE PATIENT EXPIRED ON SUPPORT. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859127&pc=MGB | 2024168-2025-03971 | 2025-03-26 04:00 | Death | ABBOTT VASCULAR INC. | 2025-04-17 04:00 | MGB | PERCLOSE PROGLIDE¿ | Device Damaged by Another Device | Hemorrhage/Bleeding; Unspecified Tissue Injury | P960043 |  | 1 | Event Description: IT WAS REPORTED THAT THIS WAS AN ARTERIOTOMY CLOSURE OF A COMMON FEMORAL ARTERY USING A PROGLIDE DEVICE RELATIVE TO A THROMBECTOMY PROCEDURE. REPORTEDLY, THE PROGLIDE DEVICE WAS USED WITHIN DAYS OF A COMMON FEMORAL ARTERY ENDARTERECTOMY AT SAME ACCESS SITE THAT WAS CLOSED VIA SURGICAL SUTURING. AFTER THE FOOT WAS DEPLOYED IN STEP 1, THE PROGLIDE DEVICE WAS RETRACTED FOR TACTILE SENSATION OF THE FOOT AGAINST THE ANTERIOR WALL OF THE ARTERY BUT IN DOING SO THE FOOT BROKE THE PREVIOUSLY PLACED SURGICAL SUTURES. THE PATIENT BEGAN TO BLEED AND HEMOSTASIS WAS UNABLE TO BE ACHIEVED. THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR SURGICAL REPAIR. THE PATIENT EXPIRED AT A LATER DATE. THERE WAS A REPORTED CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE OR THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED. MEDICAL REVIEW CONCLUDED THE PATIENT¿S CAUSE OF DEATH IS MOST LIKELY DUE TO HEMORRHAGIC SHOCK AND THE PROGLIDE DEVICE WAS ONLY INDIRECTLY INVOLVED AS IT WAS USED FOR THE PROCEDURE RELATED TO INTERACTION WITH THE PREVIOUSLY PLACED SUTURE. THE REPORTED PATIENT EFFECTS OF EFFECTS OF DEATH, HEMORRHAGE AND VASCULAR INJURY ARE KNOWN INHERENT RISKS OF THE PROCEDURE. THE SUBSEQUENT TREATMENTS APPEAR TO BE RELATED TO PROCEDURE CIRCUMSTANCE. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2, B3: DATE ESTIMATED. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859240&pc=DSQ | 3007042319-2025-00775 | 2025-04-01 04:00 | Death | HEARTWARE, INC. | 2025-04-17 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE PASSED AWAY DUE TO CHRONIC ILLNESS. THE PATIENT HAD ENTERED HOSPICE CARE BECAUSE OF THE CHRONIC ILLNESS AND UNDERWENT A COMPASSIONATE VAD DEACTIVATION DURING A HOSPITAL STAY, WHICH WAS THE TERMINAL EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS IN HOSPICE CARE FOR A PELVIC FRACTURE WHICH RESULTED IN AN INFECTED DECUBITUS ULCER AND OSTEOMYELITIS WITH SIGNIFICANT PAIN. THE VAD WAS DECOMMISSIONED COMPASSIONATELY TO WITHDRAW CARE ALLOWING FOR A NATURAL DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION TO B5. ADDITIONAL REVIEW OF THE EVENT SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. NO ADDITIONAL REPORTS WILL BE SUBMITTED FOR THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859340&pc=OZD | 1220648-2025-27632 | 2025-03-25 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A THROMBUS DEVELOPED IN THE LEFT VENTRICLE DURING IMPELLA 5.5 SUPPORT. THE APICAL THROMBUS WAS FOUND NEAR THE INLET OF THE IMPELLA 5.5. CARDIOLOGY DISCUSSED OPTIONS AND THE DECISION WAS ULTIMATELY MADE TO EXPLANT THE PUMP. THE PATIENT'S STATUS WAS UPDATED TO EXPIRED FOLLOWING EXPLANT. PER MEDICAL SAFETY OFFICER REVIEW, THE PUMP WAS PREMATURELY EXPLANTED BECAUSE OF THE THROMBUS WHICH MAY OR MAY NOT BE RELATED TO THE USE OF THE IMPELLA DEVICE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: PER MEDICAL SAFETY OFFICER REVIEW, THE PUMP WAS PREMATURELY EXPLANTED BECAUSE OF THE THROMBUS WHICH MAY OR MAY NOT BE RELATED TO THE USE OF THE IMPELLA DEVICE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ Manufacturer Narrative: THE INVESTIGATION OF THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. D4 MODEL NUMBER WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27632. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27632.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859451&pc=LWP | 2017865-2025-40439 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-17 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PULSE GENERATOR WAS REMOVED DUE TO PATIENT DEATH. THE CAUSE OF DEATH PROVIDED WAS, "CARDIAC DYSRHYTHMIA DUE TO DILATED CARDIOMYOPATHY IN THE SETTING OF REMOVE TRICUSPID VALVE ENDOCARDITIS". THE PRODUCT WAS RETURNED WITHOUT CLEAR INFORMATION AS TO WHETHER THE DEATH WAS DEVICE RELATED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR ANALYSIS DUE TO PATIENT DEATH. ANALYSIS PERFORMED, INCLUDED ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS, AND INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE PRODUCT WAS RETURNED AND VISUAL INSPECTION WAS NORMAL. THE CAUSE OF INFECTION COULD NOT BE TRACED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859628&pc=NJK | 3007797756-2025-00020 | 2025-02-28 05:00 | Death | PULMONX CORPORATION | 2025-04-17 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P180002 |  | 1 | Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES ON (B)(6) 2025. THE PATIENT WAS TRANSFERRED TO ICU POST-PROCEDURE AND THE VALVES WERE REMOVED IN THE FOLLOWING DAYS AFTER THE PROCEDURE. IT WAS MADE AWARE ON (B)(6) 2025 THAT THE PATIENT HAD PASSED AWAY WHILE IN RECOVERY.Event Description: A BLVR PROCEDURE WAS COMPLETED INCLUDING IMPLANTATION OF FOUR VALVES (ONE VALVE IN THE RIGHT MIDDLE LOBE (RML) AND THREE IN THE RIGHT UPPER LOBE (RUL) ON (B)(6) 2025. PER THE PHYSICIAN, THE PROCEDURE DURATION LASTED FOR ABOUT ONE HOUR. ON (B)(6) 2025, ALL FOUR VALVES WERE EXPLANTED DUE TO A DECLINE IN THE PATIENT'S RESPIRATORY STATUS FOLLOWING THE PROCEDURE WHILE IN THE ICU. THE VALVES REMAINED IMPLANTED FOR A TOTAL OF 3 DAYS. ON (B)(6) 2025, IMMEDIATELY FOLLOWING THE BRONCHOSCOPY, THE PATIENT EXPERIENCED A FAILED EXTUBATION DUE TO CO[?] RETENTION AND HYPERSOMNOLENCE, REQUIRING RE-INTUBATION AND TRANSFER TO THE ICU. DESPITE VARIOUS VENTILATOR ADJUSTMENTS, THE PATIENT DEMONSTRATED VERY POOR AIR MOVEMENT. THE VALVES WERE REMOVED AFTER THREE DAYS; HOWEVER, THE PATIENT'S CONDITION DID NOT IMPROVE. THE PATIENT CONTINUED TO DETERIORATE AND SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025. PER THE MEDICAL CHART, THE FOLLOWING STATEMENT WAS DOCUMENTED: "DUE TO PERSISTENCE OF VENTILATOR DEPENDENCE AND INABILITY TO WEAN PATIENT SUCCESSFULLY FROM PARALYSIS AND DEEP SEDATION, COUNSELED FAMILY THAT NEXT TREATMENT OPTION WAS TO UNDERGO TRACHEOSTOMY PLACEMENT VERSUS PURSUE COMFORT FOCUSED CARE. IN SPEAKING WITH MR.(B)(6) FAMILY, THEY RELAYED THAT TRACHEOSTOMY WAS NOT IN LINE WITH HIS GOALS OF CARE. AS SUCH, PATIENT WAS TRANSITIONED TO COMFORT FOCUSED CARE ON (B)(6) 2025 AND WAS PALLIATIVELY EXTUBATED AT THIS TIME. MR.(B)(6) PASSED PEACEFULLY SHORTLY THEREAFTER ON (B)(6) 2025 AT 20:30." ACCORDING TO THE PATIENT'S CHART AND MEDICAL RECORDS, THE PRIMARY CAUSE OF DEATH WAS ACUTE HYPOXEMIC HYPERCAPNIC RESPIRATORY FAILURE ATTRIBUTED TO SEVERELY LIMITED PULMONARY RESERVE AND THE EFFECTS OF GENERAL ANESTHESIA. A SECONDARY CONTRIBUTING FACTOR WAS A COPD EXACERBATION. NO AUTOPSY WAS PERFORMED. ACCORDING TO THE PHYSICIAN, THE VALVES WERE DISCARDED AS THEY WERE NOT BELIEVED TO BE RELATED TO THE PATIENT'S CLINICAL DETERIORATION OR CAUSE OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859669&pc=NPT | 2025587-2025-02789 | 2024-05-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-17 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Hemorrhage/Bleeding; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A NON-MEDTRONIC TRANSCATHETER VALVE WAS ATTEMPTED IN THE PATIENT; HOWEVER, THIS VALVE COULD NOT BE ADVANCED. THE NON-MEDTRONIC VALVE WAS WITHDRAWN FROM THE PATIENT AND THE EVOLUTFX-26 WAS PLACED. THREE DAYS LATER, A HEMORRHAGE OF CRITERIA 3 OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM CLASSIFICATION SCALE WAS NOTED AT THE ACCESS SITE. APPROXIMATELY 2 WEEKS LATER, THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION DUE TO HEART FAILURE. THE HEART FAILURE DID NOT IMPROVE AND THE PATIENT ULTIMATELY PASSED AWAY. THE CAUSE OF DEATH WAS HEART FAILURE.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE A NON-MEDTRONIC TRANSCATHETER VALVE WAS ATTEMPTED IN THE PATIENT; HOWEVER, THIS VALVE COULD NOT BE ADVANCED. THE NON-MEDTRONIC VALVE WAS WITHDRAWN FROM THE PATIENT AND THE EVOLUTFX-26 WAS PLACED. THREE DAYS LATER, A HEMORRHAGE OF CRITERIA 3 OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM CLASSIFICATION SCALE WAS NOTED AT THE ACCESS SITE. APPROXIMATELY 2 WEEKS LATER, THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) DUE TO HEART FAILURE. THE HEART FAILURE DID NOT IMPROVE AND THE PATIENT ULTIMATELY PASSED AWAY. THE CAUSE OF DEATH WAS HEART FAILURE. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY THE PATIENT WAS IN CARDIOGENIC SHOCK AND UNDERWENT THE TRANSCATHETER AORTIC VALVE IMP LANTATION (TAVI) WHILE ECMO WAS INSERTED. VASCULAR ACCESS FOR THE TAVI PROCEDURE WAS ACHIEVED VIA THE RIGHT "SIDE" WHICH WAS THE SAME SIDE AS THE VASCULAR INJURY. THE INJURY WAS NOTED FOLLOWING THE TAVI PROCEDURE. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM DID NOT CONTRIBUTE TO THE VASCULAR INJURY, INSTEAD THE PATIENT'S ANATOMY INCLUDED TORTUOSITY IN THE CHEST WHICH DID CONTRIBUTE TO THE INJURY. ALTHOUGH THE ECMO WAS REMOVED AFTER SURGERY, THE HEART FAILURE DID NOT RECOVER. THE PATIENT DIED APPROXIMATELY TWO WEEKS AFTER THE TAVI PROCEDURE. NO MALFUNCTIONS WERE OBSERVED IN THE IMPLANTED DEVICE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE; SERIAL #; AND FULL UDI # H4. DEVICE MFG DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859980&pc=NIK | 9614453-2025-01522 | 2025-04-01 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-17 04:00 | NIK | CLARIA MRI CRT-D SURESCAN | Pacing Problem | Bradycardia; Dizziness; Cardiogenic Shock | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS SYMPTOMATIC WITH DIZZINESS. IT WAS NOTED THAT THE PATIENT HAD A HEART RATE AROUND FORTY-SEVEN BEATS PER MINUTE WHICH WAS LOWER THAN THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PROGRAMMED LOWER RATE. THE DEVICE REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DIED SHORTLY AFTER DUE TO CARDIOGENIC SHOCK WITH SECONDARY CAUSES NOTED AS CARDIAC ARREST, VENTRICULAR FIBRILLATION (VF) AND VIRAL NON-ISCHEMIC CARDIOMYOPATHY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 0181 LEAD IMPLANTED (B)(6) 2008. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859990&pc=DSQ | 3007042319-2025-00783 | 2023-07-12 04:00 | Death | HEARTWARE, INC. | 2025-04-17 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A REVIEW OF PEDIATRIC PATIENTS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE (VAD). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THEY DESCRIBE ADVERSE EVENTS IN THESE PATIENTS WHICH INCLUDED RIGHT VENTRICULAR (RV) FAILURE, BLEEDING EVENTS, WHICH INCLUDED GASTROINTESTINAL BLEED (GIB). INFECTIONS WHICH INCLUDE DRIVELINE INFECTIONS, STROKES, RESPIRATORY FAILURE, CEREBRAL HEMORRHAGE, AND PUMP THROMBOSIS. SIGNS OF RV FAILURE INCLUDED DECREASED VAD FLOW, SUCTION EVENTS, HEPATIC CONGESTION, ACUTE KIDNEY INJURY, AND DECREASED CARDIAC OUTPUT. INTERVENTIONS PERFORMED FOR THE VARIOUS ADVERSE EVENTS INCLUDE SURGICAL, MEDICAL MANAGEMENT, AND THE IMPLANT OF A RIGHT VAD. DEVICES EXHIBITED MALFUNCTIONS OF UNKNOWN ORIGIN. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS UNKNOWN. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. REFERENCED ARTICLE: COMING SOON TO AN OR NEAR YOU: A CHILD WITH A VENTRICULAR ASSIST DEVICE. CURRENT ANESTHESIOLOGY REPORTS. 2023. 1 3:172¿180. DOI: 10.1007/S40140-023-00565-1 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21859999&pc=OZD | 1220648-2025-27612 | 2025-03-23 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION AND WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT EXPERIENCED ACCESS SITE BLEEDING, LOW PUMP FLOW, AND DEVICE MALPOSITIONING WITH REPOSITIONING. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED TO THE HOSPITAL FROM JAIL IN A DECLINING CONDITION. THE PATIENT WAS NOTED TO HAVE STOPPED EATING TWO DAYS PRIOR. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE IN THE SETTING OF CARDIOGENIC SHOCK, WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT WAS PREPARED AND TRANSFERRED OUT TO A HIGH-LEVEL CARE FACILITY THE FOLLOWING DAY. ONCE ON THE UNIT, THE PATIENT WAS BLEEDING AT THE IMPELLA ACCESS SITE AND ALL OTHER ACCESS SITES. THE PATIENT WAS INTUBATED AND SEDATED AND PLACED ON A TRIPLE LUMEN CENTRAL VENOUS CATHETER AND ARTERIAL LINE. ARTERIAL BLOOD GAS SHOWED COMBINED ACIDOSIS, BUT WAS IMPROVING. THE PATIENT WAS PACED VIA TRANSVENOUS PACEMAKER AND RESTING ON IMPELLA CP SUPPORT AT PERFORMANCE LEVEL P-2 TO P-4. HOWEVER, THERE WAS INTERMITTENT SUCTION ALARMS. THE PERFORMANCE LEVEL WAS UNABLE TO GO ABOVE P-4 WITHOUT SUCTION. THE PATIENT WAS TRANSFUSED WITH ONE UNIT OF PACKED RED BLOOD CELLS (PRBCS). THE IMPELLA SITE STILL WAS NOTED WITH CONSTANT BLEEDING. PRESSURE WAS HELD FOR 30 MINUTES WITH NO CHANGE. ANGLE WAS SUPPORTED BUT APPROXIMATELY TWO INCHES OF THE REPOSITIONING SHEATH WAS EXPOSED. AT BEDSIDE, THE PHYSICIAN RE-SUTURED THE SITE AND REPOSITIONED THE IMPELLA CP PUMP. THE BLEED WAS NOTED TO HAVE STOPPED AT ALL SITES. IT WAS NOTED THAT THE PATIENT WAS NOT A CANDIDATE FOR EXTRACORPOREAL MEMBRANE OXYGENATION. THE PATIENT REMAINED IN CRITICAL CONDITION BUT STABLE AFTER REWARMING, REPOSITIONING AND VOLUME ADMINISTRATION. THE FOLLOWING DAY THE PATIENT REMAINED IN CRITICAL CONDITION. OVERNIGHT THE PATIENT EXPERIENCED BLEEDING FROM THE ENDOTRACHEAL TUBE. THE VOLUME WAS RESUSCITATED NOW WITH TWO UNITS OF PRBCS. ANTICOAGULATION WAS PUT ON HOLD. THE PATIENT WAS TOO UNSTABLE FOR PAN SCAN. A DISCUSSION WITH THE FAMILY WAS PLANNED FOR PLAN OF CARE THAT MAY INCLUDE PALLIATIVE CARE. THE FOLLOWING NIGHT THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR A WASHOUT. STERNAL WIRES WAS THE CULPRIT OF THE CHEST BLEEDING. THE PATIENT RECEIVED "LOTS" OF BLOOD PRODUCTS. THE BLEEDING SUBSIDED, AND IT WAS NOTED THE PATIENT LOOKED MARKEDLY BETTER. THE PATIENT WAS NOTED TO BE HEMODYNAMICALLY STABLE. THE PLAN WAS TO WEAN THE IMPELLA WITH HOPES THE PATIENT WAKES UP FOR EXTUBATING. HOWEVER, LATER IN THE EVENING, THE PATIENT WOULD EXPIRE. AND CARE WAS WITHDRAWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, LOW PUMP FLOW, AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING - MAJOR: THE CAUSE OF THE ACCESS SITE BLEEDING - MAJOR MOST LIKELY WAS PATIENT CONDITION RELATED AS THE PATIENT WAS BLEEDING AT ALL ACCESS SITES. LOW PUMP FLOW: COMPLAINT DEVICE NOT RETURNED FOR ANALYZING. NO DATA LOG AVAILABLE TO REVIEW. NOT ENOUGH EVIDENCE PROVIDED TO CONFIRM CAUSE OF LOW FLOW. THE CAUSE OF THE LOW PUMP FLOW ISSUE CANNOT BE DETERMINED SINCE COMPLAINT DEVICE AND DATA LOGS WERE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. POSITIONING ISSUES: THE CAUSE OF THE POSITIONING ISSUES MOST LIKELY WAS PATIENT CONDITION RELATED TO THE PATIENT'S SMALL LEFT VENTRICLE. Manufacturer Narrative: THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR. THE PATIENT WAS VERY ILL AND THE IMPACT TO THE OUTCOME AS A RESULT OF THE IMPELLA RELATED EVENT IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860089&pc=DSQ | 2916596-2025-02125 | 2025-03-15 04:00 | Death | THORATEC CORPORATION | 2025-04-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DOWN AT HOME, DECEASED, ON (B)(6) 2025. IT APPEARED TO THE CUSTOMER THAT THE PATIENT UNPLUGGED THEIR OWN MODULAR CABLE. LOG FILES WERE SUBMITTED FOR REVIEW, AND CAPTURED A DRIVELINE DISCONNECT AT 10:21:43. THE DETAILS LEADING UP TO THE DEATH WERE UNKNOWN, AND WHICH END THE MODULAR CABLE DISCONNECTED FROM WAS UNKNOWN. THE PUMP WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED A PUMP STOP ASSOCIATED WITH A DRIVELINE DISCONNECT. A SPECIFIC CAUSE FOR THE DISCONNECTION OF THE DRIVELINE COULD NOT BE CONCLUSIVELY DETERMINED. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. AT 10:21:43 ON (B)(6) 2025, THE DRIVELINE WAS DISCONNECTED, CAUSING THE PUMP TO STOP. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED WHILE THE DRIVELINE WAS CONNECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 2 OF THE IFU, "SYSTEM OPERATIONS" (UNDER "SYSTEM CONTROLLER WARNINGS AND CAUTIONS"), AND SECTION 2 OF THE PATIENT HANDBOOK, ¿HOW YOUR HEART PUMP WORKS¿, INSTRUCT THE USER TO CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK FURTHER EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. SECTION 2 OF THE IFU ALSO CONTAINS A SECTION TITLED "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", WHICH PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "EDUCATING AND TRAINING PATIENTS, FAMILIES, AND CAREGIVERS"), EXPLAINS THAT DURING THE PATIENT SELECTION, PREIMPLANT, AND POSTOPERATIVE PERIOD, THE PATIENT MUST RECEIVE INSTRUCTIONS REGARDING THE OPERATION AND CARE OF EVERY SYSTEM COMPONENT (INCLUDING THE DRIVELINE AND WHAT TO DO IN AN EMERGENCY). SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, DESCRIBE ALL ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. THE PATIENT HANDBOOK ALSO CONTAINS A SECTION ON HANDLING EMERGENCIES AND FURTHER INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE USER THINKS, FOR ANY REASON, ANY PORTION OF THE EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THE USER IS UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860216&pc=OZD | 1220648-2025-27633 | 2025-03-26 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHEN THE PATIENT SELF-EXPLANTED THE 5.5 AND PLACED BACK ON VA ECMO. THE PATIENT WAS THEN IMPLANTED WITH A NEW IMPELLA 5.5. WHILE ON SUPPORT, AFTER 10 DAYS OF THE 5.5 SUPPORT A CVA BLEED WAS OBSERVED. THE TEAM REMOVED HEPARIN FROM THE PURGE AND USED SODIUM BICARB. THE PUMP REMAINED ON FOR 5 MORE DAYS AND THEN THE CARE WAS WITHDRAWN AND ORGANS HARVESTED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860385&pc=DSQ | 3007042319-2025-00781 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-17 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A META-ANALYSIS REVIEW OF CONCOMITANT TRICUSPID VALVE SURGERY IN PATIENTS WITH SIGNIFICANT TRICUSPID REGURGITATION UNDERGOING LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. PATIENTS WERE STUDIED IN GROUPS OF A VAD ONLY OR A VAD PLUS A TRICUSPID VALVE SURGERY (TVS). THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND CATEGORIZED AS EARLY, OCCURRING IN-HOSPITAL OR LATE, OCCURRING AFTER THE INDEX HOSPITALIZATION. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED EARLY AND LATE RIGHT HEART FAILURE WITH THE NEED FOR A RIGHT VAD, POST-OPERATIVE STROKES, RENAL FAILURE, AND EXTENDED STAYS IN THE INTENSIVE CARE UNIT (ICU). THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER CHARACTERISTICS IS MALE. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. REFERENCED ARTICLE: CONCOMITANT TRICUSPID VALVE SURGERY IN PATIENTS WITH SIGNIFICANT TRICUSPID REGURGITATION UNDERGOING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION: A SYSTEMATIC REVIEW AND META-ANALYSIS. ARTIFICIAL ORGANS. 2024; 48:1392¿1403. DOI: 10.1111/AOR.14819 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860487&pc=BZD | 2518422-2025-104358 | 2025-02-18 05:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Cancer | K131982 |  | 1 | Event Description: A DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT DUE TO CANCER. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860491&pc=MNT | 2518422-2025-043988 | 2025-03-24 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | MNT | RESPIRONICS | Adverse Event Without Identified Device or Use Problem |  | K102985 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT FROM THE CUSTOMER, REPORTING THAT A PATIENT HAD EXPIRED WHILE THE DEVICE WAS IN USE. THE CUSTOMER HAD REQUESTED ASSISTANCE ON READING THE EVENT LOG THAT WAS CREATED AT THE TIME OF THE EVENT. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, IT WAS REPORTED THAT DURING CLINICAL AND THERAPEUTIC USE OF THE V60 VENTILATOR, THE DEVICE WAS ALARMING, BUT FUNCTIONAL AND A PATIENT WAS NOTED TO HAVE EXPIRED. THE REPORTER HAS REQUESTED INTERPRETATION OF THE DEVICE DRPT (DIAGNOSTIC REPORT). DEVICE EVENT LOGS WERE ANALYZED BY A PHILIPS PMS (POST MARKET SURVEILLANCE) CLINICAL EXPERT. THE ALLEGED TIME AND DATE OF EVENT (MARCH 24, 2025 AT APPROXIMATELY 2000) REVEALS THAT THE DEVICE HAD EXHIBITED MULTIPLE PATIENT THERAPY ALARMS (LOW TIDAL VOLUME, LOW INHALATION PRESSURE, LOW MINUTE VENTILATION, HIGH RATE, AND PATIENT DISCONNECT ALARMS) FROM 2000 UNTIL 2009. DURING THIS TIME, END USER INTERACTION WAS NOTED AT 2003 (END USER ADJUSTED DEVICE DELIVERED FIO2 FROM 50% TO 100%). NO FURTHER ALARMS WERE NOTED AND THE DEVICE WAS POWERED OFF AT 2049. NO ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED WITHIN THE COMPLAINT RECORD AGAINST THE DEVICE. FURTHER GOOD FAITH EFFORTS ARE REQUIRED TO DETERMINE CAUSE AND/OR CONTRIBUTION. FURTHER GOOD FAITH EFFORTS YIELDED INFORMATION STATING THAT THE CONTACT ESTABLISHED BY THE REPORTER WAS AN INQUIRY AS TO INTERPRETATION OF THE DATA PROVIDED IN THE DRPT AND NOT INTENDED TO LODGE A COMPLAINT OF DEVICE CAUSE AND/OR CONTRIBUTION TO THE PATIENT OUTCOME. DESPITE THE INQUIRY, NO FURTHER INFORMATION HAS YIELDED ADDITIONAL INFORMATION THAT WOULD AMEND THE PREVIOUS DISPOSITIONED CLINICAL ASSESSMENT. A SERVICE ENGINEER (SE) EVALUATED THE DEVICE AND REPORTED THAT THE DEVICE POWERED ON AND OPERATED NORMALLY IN VENTILATION MODE. THE SE NOTED THAT NO ISSUES COULD BE DUPLICATED WHILE TESTING THE DEVICE. THE DIAGNOSTIC LOGS WERE PULLED AND REVIEWED AND SHOWED NO ERROR CODES. A COMPLETE PERFORMANCE VERIFICATION TEST WAS PERFORMED, AND THE DEVICE PASSED ALL TEST. THE DEVICE WAS RETURNED TO SERVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860530&pc=FKX | 0008030665-2025-00887 | 2025-04-07 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-17 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: ON 8/APR/2025, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT HAD CONTRACTED INFLUENZA DURING THE SAME TIME FRAME. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS DISCOVERED BY A FAMILY MEMBER WHEN THEY ENTERED THE BEDROOM TO ASSIST THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HER EXTENSIVE CARDIAC HISTORY AND INFLUENZA. PER THE PDRN, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION.Event Description: ON 8/APR/2025, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT HAD CONTRACTED INFLUENZA DURING THE SAME TIME FRAME. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS DISCOVERED BY A FAMILY MEMBER WHEN THEY ENTERED THE BEDROOM TO ASSIST THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HER EXTENSIVE CARDIAC HISTORY AND INFLUENZA. PER THE PDRN, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND SUBSEQUENT DEATH, AS THE PATIENT WAS UNDERGOING HD THERAPY WHEN HE EXPIRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT¿S EXTENSIVE CARDIAC HISTORY AND INFLUENZA. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE: HEATER TRAY DISCOLORED THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860542&pc=FMI | 3006260740-2025-02429 | 2025-03-20 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-04-17 04:00 | FMI | INTRAOSSEOUS, INTRAOSSEOUS NEEDLE KITS | Material Deformation; Material Twisted/Bent | No Clinical Signs, Symptoms or Conditions | K203193 |  | 1 | Event Description: IT WAS REPORTED BY THE CUSTOMER, "DURING A CARDIAC ARREST LAST NIGHT, BENT NEEDLE THAT MADE IT IMPOSSIBLE TO PLACE AN IO DEVICE. PARAMEDIC WAS UNABLE TO PLACE A IO DURING A CARDIAC ARREST DUE TO NEEDLES ET BENDING DURING PLACEMENT.. PARAMEDIC DID PLACE AN IV, HOWEVER, PATIENT EXPIRED." Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF AN INTRAOSSEOUS (IO) NEEDLE BENDING DURING ATTEMPTED INSERTION WAS CONFIRMED AND THE CAUSE APPEARED TO BE USE-RELATED. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 25 MM INTRAOSSEOUS (IO) NEEDLE. USAGE RESIDUES WERE OBSERVED ON THE SAMPLE. THE OBTURATOR WAS INLAID THROUGH THE NEEDLE SHAFT. A BEND WAS OBSERVED IN THE NEEDLE SHAFT NEAR THE LUER ADAPTER. THE NEEDLE TIP EXHIBITED DEFORMATION. MICROSCOPIC INSPECTION OF THE NEEDLE TIP CONFIRMED DEFORMATION. THE NEEDLE TIP WAS TWISTED AND PEELED AWAY FROM THE OBTURATOR TIP. THE OBTURATOR TIP WAS FOLDED TOWARD THE NEEDLE BEVEL. THE BEND IN THE NEEDLE SHAFT WAS CONSISTENT WITH SHEER, OR LATERAL STRESS PLACED ON THE NEEDLE SHAFT. THE TIP DEFORMATION WAS CONSISTENT WITH EXCESSIVE FORCE PLACED ON THE NEEDLE TIP WHILE THE NEEDLE WAS ROTATING. THE NEEDLE BEND FURTHER SUGGESTED THAT FORCE WAS APPLIED WHILE THE NEEDLE WAS NOT PLACED AT A 90 DEGREE ANGLE TO THE PATIENT¿S SKIN. THE PRODUCT INSTRUCTIONS FOR USE (IFU) INFORM USERS TO USE MODERATE STEADY PRESSURE WHEN INSERTING THE NEEDLE INTO THE BONE AT A 90 DEGREE ANGLE TO THE SKIN. POSITIONING THE NEEDLE IN THIS FASHION MINIMIZES THE AMOUNT OF LATERAL PRESSURE, REDUCING THE RISK OF NEEDLE DEFORMATION/DAMAGE. THE OBSERVED NEEDLE DEFORMATION APPEARED TO BE THE RESULT OF INSERTION TECHNIQUES. IT IS UNKNOWN TO WHAT EXTENT INITIAL NEEDLE PLACEMENT FAILURE ULTIMATELY CONTRIBUTED TO DECLINING PATIENT CONDITIONS. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE CUSTOMER REPORTED THE PATIENT PASSED AWAY BUT DID NOT PROVIDE THE PATIENT'S DATE OF DEATH. THE DATE IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE PATIENT'S DEATH. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IT WILL BE EVALUATED. A SUPPLEMENTAL WILL BE SUBMITTED WITH EVALUATION RESULTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21860769&pc=NPT | 2015691-2025-03007 | 2025-03-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE SYSTEM | Migration or Expulsion of Device | Vascular Dissection; Foreign Body Embolism | P140031 |  | 1 | Event Description: AS REPORTED BY EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, REGARDING A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION VIA TRANSFEMORAL APPROACH. DURING DEPLOYMENT OF THE FIRST 29MM SAPIEN 3 ULTRA RESILIA VALVE, THE 29MM COMMANDER DELIVERY SYSTEM BALLOON WAS 80% INFLATED WHEN THERE WAS A PVC WHICH EMBOLIZED THE VALVE. THE VALVE WAS PULLED BACK TO THE DESCENDING AORTA AND DEPLOYED. A SECOND 29MM SAPIEN 3 ULTRA RESILIA VALVE AND 29MM COMMANDER DELIVERY SYSTEM WERE USED SUCCESSFULLY WITH 3 EXTRA CC OF VOLUME. AFTER THE CASE WAS OVER, 2 ANGIOS OF THE ARCH AND THE DESCENDING AORTA WERE PERFORMED, AND NOTHING WAS OBSERVED. THE PATIENT LEFT THE ROOM STABLE. THE NEXT DAY A SURVEILLANCE CT WAS PERFORMED AND A DISSECTION DISTAL TO THE LEFT SUBCLAVIAN WHICH CONTINUED TO PROPAGATE WAS OBSERVED. THE PATIENT WAS NOT A SURGICAL BAILOUT AND COMFORT MEASURES ONLY WERE TAKEN. THE PATIENT EXPIRED TWO DAYS LATER. PER THE FIELD CLINICAL SPECIALIST, THE EMBOLISM WAS DUE TO THE PVC. THE DISSECTION WAS LIKELY DUE TO WHEN THE BALLOON AND EXPANDED VALVE WERE PULLED BACK OVER THE ARCH. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (PVC, INTERVENTION FOR NON-TARGET DEPLOYMENT) CAUSED OR CONTRIBUTED TO THE EMBOLIZATION AND THE DISSECTION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21861006&pc=GKT | 1722028-2025-00092 | 2025-01-29 05:00 | Death | TERUMO BCT | 2025-04-17 04:00 | GKT | SPECTRA OPTIA | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Abdominal Pain; Fever; Headache; Nausea; Vomiting; Jaundice; Hematemesis | K183081 |  | 1 | Event Description: PER JOURNAL ARTICLE ""INTENSIVE THERAPEUTIC PLASMA EXCHANGE¿NEW APPROACH TO TREAT AND RESCUE PATIENTS WITH SEVERE FORM OF YELLOW FEVER"" BY HO, Y., NUKUI, Y., VILLAÇA, P. R., OKAZAKI, E., TATSUI, N. H., NETTO, L. C., JOELSONS, D., DA ROCHA, T. R. F., DE MELLO MALTA, F., PINHO, J. R. R., SEGURADO, A. A. C., & ROCHA, V. THE STUDY DISCUSSES THE USE OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN PATIENTS WITH SEVERE YELLOW FEVER, UTILIZING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. WHILE THE STUDY REPORTS A SIGNIFICANT REDUCTION IN MORTALITY WITH INTENSIVE TPE, IT ALSO NOTES ADVERSE EVENTS SUCH AS HYPOCALCEMIA AND ALLERGIC REACTIONS DURING THE PROCEDURES. SPECIFICALLY, ONE PATIENT EXPERIENCED AN ALLERGIC REACTION DURING TPE, WHICH IS A CONCERN FOR PATIENT SAFETY. THE USE OF TERUMO PRODUCTS IN THIS CONTEXT RAISES QUESTIONS ABOUT THE MANAGEMENT OF SUCH ADVERSE EVENTS AND THE NEED FOR MONITORING DURING PROCEDURES. THE FINDINGS SUGGEST THAT WHILE THE DEVICES ARE EFFECTIVE IN MANAGING SEVERE YELLOW FEVER, THERE ARE POTENTIAL RISKS ASSOCIATED WITH THEIR USE THAT NEED TO BE ADDRESSED TO ENSURE PATIENT SAFETY. THIS OBSERVATIONAL CASE-SERIES STUDY EVALUATES THE EFFECTIVENESS OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN TREATING SEVERE YELLOW FEVER, PARTICULARLY IN PATIENTS WITH ACUTE LIVER FAILURE. THE STUDY INVOLVED THREE GROUPS OF PATIENTS RECEIVING DIFFERENT TREATMENT REGIMENS, WITH A FOCUS ON THE OUTCOMES ASSOCIATED WITH INTENSIVE TPE USING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. THE RESULTS INDICATED A SIGNIFICANT REDUCTION IN MORTALITY RATES AMONG PATIENTS RECEIVING INTENSIVE TPE COMPARED TO THOSE RECEIVING STANDARD INTENSIVE CARE SUPPORT OR HIGH-VOLUME TPE. HOWEVER, THE STUDY ALSO HIGHLIGHTED ADVERSE EVENTS, INCLUDING ALLERGIC REACTIONS AND HYPOCALCEMIA, WHICH OCCURRED DURING THE PROCEDURES. THE AUTHORS CONCLUDED THAT INTENSIVE TPE IS A FEASIBLE AND EFFECTIVE INTERVENTION FOR SEVERE YELLOW FEVER, BUT FURTHER STUDIES ARE NEEDED TO CONFIRM THESE FINDINGS AND ADDRESS THE ASSOCIATED RISKS. ONE WAS A 48-YO MALE PATIENT (VPP) WHO REPORTED LONG-TERM USE OF ALCOHOL, WITH FOUR DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (403 AND 148 TIMES GREATER THAN THE UPPER LIMIT OF NORMAL [ULN], RESPECTIVELY), WITH HIGH AMMONIA LEVELS (3.7 TIMES THE ULN) AND LOW LEVELS OF FV (REDUCTION OF 90% FROM NORMAL). FOUR DIFFERENT BACTERIA WERE IDENTIFIED 48 HS LATER IN ADMISSION BLOOD CULTURE (E. COLI, S. VIRIDIANS, S. PNEUMONIAE AND F. NUCLEATUM). DESPITE ICS, ANTIBIOTICS, AND INTENSIVE TPE, THE PATIENT PRESENTED WITH A MASSIVE ALVEOLAR HEMORRHAGE AND DIED 72 HS AFTER ADMISSION (TABLE 2). THE SECOND WAS ALSO A 48-YO MALE PATIENT (VSS), PREVIOUSLY VACCINATED AGAINST YFV (10 YEARS BEFORE), WITH SIX DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (163 AND 103 TIMES THE ULN, RESPECTIVELY), HIGH AMMONIA LEVELS (2.3 TIMES THE ULN), AND NORMAL LEVELS OF FV. AFTER FIVE DAYS OF INTENSIVE TPE, ALL LABORATORY RESULTS WERE NORMALIZED. HOWEVER, ON THE SIXTH DAY OF ADMISSION, S. AUREUS WAS IDENTIFIED IN A BLOOD CULTURE AND THE PATIENT DIED SEVEN DAYS LATER. "ONLY MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. HYPOCALCEMIA (IONIC CALCIUM < 4.0 MG/DL) WAS MORE FREQUENT IN G2 (P < 0.001)" THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. SPECIFIC DETAILS, SUCH AS PATIENT INFORMATION AND OUTCOME, WERE NOT INCLUDED IN THE ARTICLE FOR THESE EVENTS, THEREFORE THIS REPORT IS BEING PROVIDED AS A SUMMARY OF THE EVENTS. IT IS UNKNOWN AT THIS TIME IF MEDICAL INTERVENTION WAS REQUIRED FOR THESE EVENTS THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION.Event Description: PER JOURNAL ARTICLE "INTENSIVE THERAPEUTIC PLASMA EXCHANGE¿NEW APPROACH TO TREAT AND RESCUE PATIENTS WITH SEVERE FORM OF YELLOW FEVER"" BY HO, Y., NUKUI, Y., VILLAÇA, P. R., OKAZAKI, E., TATSUI, N. H., NETTO, L. C., JOELSONS, D., DA ROCHA, T. R. F., DE MELLO MALTA, F., PINHO, J. R. R., SEGURADO, A. A. C., & ROCHA, V. THE STUDY DISCUSSES THE USE OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN PATIENTS WITH SEVERE YELLOW FEVER, UTILIZING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. WHILE THE STUDY REPORTS A SIGNIFICANT REDUCTION IN MORTALITY WITH INTENSIVE TPE, IT ALSO NOTES ADVERSE EVENTS SUCH AS HYPOCALCEMIA AND ALLERGIC REACTIONS DURING THE PROCEDURES. SPECIFICALLY, ONE PATIENT EXPERIENCED AN ALLERGIC REACTION DURING TPE, WHICH IS A CONCERN FOR PATIENT SAFETY. THE USE OF TERUMO PRODUCTS IN THIS CONTEXT RAISES QUESTIONS ABOUT THE MANAGEMENT OF SUCH ADVERSE EVENTS AND THE NEED FOR MONITORING DURING PROCEDURES. THE FINDINGS SUGGEST THAT WHILE THE DEVICES ARE EFFECTIVE IN MANAGING SEVERE YELLOW FEVER, THERE ARE POTENTIAL RISKS ASSOCIATED WITH THEIR USE THAT NEED TO BE ADDRESSED TO ENSURE PATIENT SAFETY. THIS OBSERVATIONAL CASE-SERIES STUDY EVALUATES THE EFFECTIVENESS OF INTENSIVE THERAPEUTIC PLASMA EXCHANGE (TPE) IN TREATING SEVERE YELLOW FEVER, PARTICULARLY IN PATIENTS WITH ACUTE LIVER FAILURE. THE STUDY INVOLVED THREE GROUPS OF PATIENTS RECEIVING DIFFERENT TREATMENT REGIMENS, WITH A FOCUS ON THE OUTCOMES ASSOCIATED WITH INTENSIVE TPE USING TERUMO BCT'S COBE SPECTRA AND OPTIA DEVICES. THE RESULTS INDICATED A SIGNIFICANT REDUCTION IN MORTALITY RATES AMONG PATIENTS RECEIVING INTENSIVE TPE COMPARED TO THOSE RECEIVING STANDARD INTENSIVE CARE SUPPORT OR HIGH-VOLUME TPE. HOWEVER, THE STUDY ALSO HIGHLIGHTED ADVERSE EVENTS, INCLUDING ALLERGIC REACTIONS AND HYPOCALCEMIA, WHICH OCCURRED DURING THE PROCEDURES. THE AUTHORS CONCLUDED THAT INTENSIVE TPE IS A FEASIBLE AND EFFECTIVE INTERVENTION FOR SEVERE YELLOW FEVER, BUT FURTHER STUDIES ARE NEEDED TO CONFIRM THESE FINDINGS AND ADDRESS THE ASSOCIATED RISKS. ONE WAS A 48-YO MALE PATIENT (VPP) WHO REPORTED LONG-TERM USE OF ALCOHOL, WITH FOUR DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (403 AND 148 TIMES GREATER THAN THE UPPER LIMIT OF NORMAL [ULN], RESPECTIVELY), WITH HIGH AMMONIA LEVELS (3.7 TIMES THE ULN) AND LOW LEVELS OF FV (REDUCTION OF 90% FROM NORMAL). FOUR DIFFERENT BACTERIA WERE IDENTIFIED 48 HS LATER IN ADMISSION BLOOD CULTURE (E. COLI, S. VIRIDIANS, S. PNEUMONIAE AND F. NUCLEATUM). DESPITE ICS, ANTIBIOTICS, AND INTENSIVE TPE, THE PATIENT PRESENTED WITH A MASSIVE ALVEOLAR HEMORRHAGE AND DIED 72 HS AFTER ADMISSION (TABLE 2). "ONLY MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. HYPOCALCEMIA (IONIC CALCIUM < 4.0 MG/DL) WAS MORE FREQUENT IN G2 (P < 0.001)" THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. SPECIFIC DETAILS, SUCH AS FURTHER PATIENT INFORMATION AND OUTCOME, WERE NOT INCLUDED IN THE ARTICLE FOR THESE EVENTS, THEREFORE THIS REPORT IS BEING PROVIDED AS A SUMMARY OF THE EVENTS. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: INVESTIGATION: THIS IS AN OBSERVATIONAL CASE-SERIES UNICENTRIC STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK. IT WAS CONDUCTED AT HOSPITAL DAS CLINICAS (HCFMUSP), A 2200-BED TERTIARY CARE TEACHING HOSPITAL AFFILIATED TO THE UNIVERSITY OF SAO PAULO WITH 110 INTENSIVE CARE BEDS AND A 7-BED INFECTIOUS DISEASE INTENSIVE CARE UNIT (ID-ICU). ON 9 JANUARY 2018, THE ID-ICU HAD BECOME THE CENTER OF EXPERTISE FOR SEVERE FORMS OF YF IN SP. SEVERITY CRITERIA FOR ID-ICU ADMISSION WERE: ASPARTATE AMINOTRANSFERASE (AST) OR ALANINE AMINOTRANSFERASE (ALT) > 3000 U/L; INR > 1.5; PLATELET COUNT < 90 × 109/L; KIDNEY DYSFUNCTION; HEMORRHAGIC MANIFESTATIONS; ENCEPHALOPATHY; OR HEMODYNAMIC INSTABILITY. HOWEVER, IN 2019, A YELLOW FEVER OUTBREAK OCCURRED IN A REGION OF THE STATE OF SAO PAULO WHERE TRANSPORTATION TO HCFMUSP TAKES AT LEAST 6 H. FOR THIS REASON, ALL PATIENTS WHO HAD ALT/AST > 500 WERE REFERRED TO HCFMUSP. FROM 9 JAN TO 15 FEB 2018, TREATMENT CONSISTED OF STANDARD INTENSIVE CARE SUPPORT (ICS). IN THE SECOND HALF OF FEBRUARY 2018, HIGH-VOLUME TPE (HV-TPE) WAS USED IN SOME PATIENTS WITH SEVERE YF IN OUR SERVICE. TRANSIENT IMPROVEMENT OF LABORATORY RESULTS WAS OBSERVED, HOWEVER HEMORRHAGIC PHENOMENA OCCURRED BETWEEN SESSIONS OF HV-TPE WITHOUT CHANGING THE OUTCOME. SINCE APRIL 2018, THIS STRATEGY WAS MODIFIED TO MORE INTENSIVE TPE TWICE A DAY AND SUPPORTIVE TRANSFUSION, IF NECESSARY. THIS INTENSIVE PROTOCOL WAS USED ON ONE PATIENT IN 2018 SUCCESSFULLY, AND SINCE THEN IT HAS BEEN USED IN THE STATE OF SAO PAULO¿S 2019 YF OUTBREAK. TPE WAS ADMINISTERED TO PATIENTS WITH LEVELS OF AMMONIA > 70 MCMOL/L AND/OR FACTOR V (FV) < 50%. THE DEFINITION OF THE CUT-OFF OF AMMONIA WAS BASED ON THE 2018 EXPERIENCE WHEN PATIENTS WITH AMMONIA > 70 MCMOL/L HAD A HIGHER RELATIVE RISK OF DEATH (RR 2.0) (DATA NOT PUBLISHED). THE CUT-OFF OF FV < 50% WAS DEFINED BASED ON CLICHY¿S CRITERIA, CONSIDERING A REDUCTION OF 30% FROM THE REFERENCE VALUE OF 70¿150% FOR ALL PATIENTS. IN ORDER TO COMPARE TO PREVIOUS CASES, ALL PATIENTS WITH AMMONIA > 70 MCMOL/L AND/OR FV < 50% WERE SELECTED. FOR THIS ANALYSIS, PATIENTS WERE CLASSIFIED INTO THREE GROUPS ACCORDING TO TYPE OF TREATMENT GIVEN: GROUP ONE (G1)¿ICS; GROUP TWO (G2)¿ICS AND HV-TPE; AND GROUP THREE (G3)¿ICS AND INTENSIVE TPE. CLINICAL AND LABORATORY DATA WERE EXTRACTED RETROSPECTIVELY AND STORED IN THE REDCAP WEB PLATFORM. HV-TPE STRATEGY (G2): PROPOSED TREATMENT CONSISTED OF THREE CONSECUTIVE DAYS OF PLASMA EXCHANGE. THE VOLUME CHANGED WAS 10% OF BODY WEIGHT (BW) WITH FRESH FROZEN PLASMA (FFP). ONE TO THREE ADDITIONAL SESSIONS OF ONE PLASMA VOLUME EACH WERE IMPLEMENTED ACCORDING TO AMMONIA LEVELS. SUPPORTIVE TRANSFUSION WAS GUIDED MAINLY BY HEMORRHAGIC EPISODES. INTENSIVE-TPE STRATEGY (G3): VOLUME OF EXCHANGE WITH FFP WAS CALCULATED ACCORDING TO BW, HEMATOCRIT (HT), AND PLASMA VOLUME (BW × 0.7 × [1 - HT]), CONSIDERING ONE PLASMA VOLUME PER SESSION. PROCEDURE WAS PERFORMED TWICE A DAY, WITH ADDITIONAL FFP INFUSION (10 ML/KG) AT NIGHT. PLATELET CONCENTRATE WAS TRANSFUSED TO KEEP PLATELET COUNT > 30 × 109/L OR, IN THE PRESENCE OF BLEEDING, >50 × 109/L. IN SEVERE OR LIFE-THREATENING BLEEDS, TARGET WAS >80¿100 × 109/L. CRYOPRECIPITATE TRANSFUSION WAS INDICATED IF FIBRINOGEN LEVEL WAS <100 MG/DL. BASED ON BLEEDING AS WELL AS FV AND/OR AMMONIA LEVELS, THE FREQUENCY OF TPE WAS TAPERED TO ONCE A DAY, EVERY OTHER DAY AND DISCONTINUATION. DURING THE DE-ESCALATION PERIOD, FFP WAS INFUSED (5 ML/KG) 1¿2 TIMES/DAY, FOLLOWING THE SAME CRITERIA ABOVE. TRANEXAMIC ACID (10 MG/KG/DOSE, IV, EVERY 6¿8 H) WAS PRESCRIBED UNTIL DISCONTINUATION OF TPE AND TOPICAL USE WAS IMPLEMENTED IN CASES OF MUCOCUTANEOUS BLEEDING. TPE WAS PERFORMED USING FFP REPLACEMENT, WITHOUT ANTICOAGULANT (CITRATE), ONCE ALL PATIENTS HAD ALF. CALCIUM CHLORIDE WAS ADMINISTERED AT THE FOLLOWING RATES: FOR HV-TPE GROUP, 5 MMOL PER EACH LITER OF PLASMA EXCHANGE; FOR INTENSIVE-TPE GROUP, FIXED 6.8 MMOL OF CALCIUM PER SESSION. AFTER EACH SESSION, SERUM CALCIUM LEVEL WAS ASSESSED, AND ADDITIONAL SUPPLEMENTATION WAS CONSIDERED. COBE SPECTRA AND OPTIA (TERUMOBCT, TOKYO, JAPAN) AND COM.TECR (FRESENIUS KABI, BAD HOMBURG, GERMANY) TPE DEVICES WERE EMPLOYED, USING A WITHDRAWAL FLOW OF 70 ML/MINUTE ON AVERAGE. IN CASES WHERE CONTINUOUS VENOVENOUS HEMODIALYSIS WAS INDICATED, THE PLASMA EXCHANGE SYSTEM WAS CONNECTED TO THE RETURN CIRCUIT OF THE DIALYSIS SYSTEM. YFV-RNA WAS ISOLATED FROM 250 ¿L OF SERUM SAMPLES AND 20 ¿G OF RNASE-FREE GLYCOGEN WERE ADDED TO THE AQUEOUS PHASE AS AN RNA CARRIER. REAL-TIME QUANTITATIVE YFV REVERSE TRANSCRIPTION (RT-QPCR) WAS CARRIED OUT USING SUPERSCRIPTR III ONE-STEP RT-QPCR SYSTEM WITH PLATINUMR TAQ DNA POLYMERASE (INVITROGEN, THERMO FISHER SCIENTIFIC BRAND, CARLSBAD, CA, USA). THE YFV 5'NON-CODING REGION WAS TARGETED USING SPECIFIC PRIMERS AND PROBE, AS PREVIOUSLY DESCRIBED. ALL SAMPLES WERE TESTED IN TRIPLICATE. THE PRIMARY END POINT WAS MORTALITY. BASELINE CLINICAL AND LABORATORY ASPECTS WERE COMPARED AMONG GROUPS. SURVIVAL CURVES WERE CONSTRUCTED USING THE KAPLAN¿MEIER METHOD, AND THE LOG-RANK TEST WAS USED TO ASSESS DIFFERENCES BETWEEN CURVES. PROPORTIONAL HAZARD ASSUMPTION WAS CHECKED GRAPHICALLY BY LOG-MINUS-LOG VS. LOG (TIME) PLOT. COMPARISONS BETWEEN TWO GROUPS WERE ASSESSED USING THE STUDENT T-TEST, OR THE MANN¿WHITNEY TEST IF APPROPRIATE FOR QUANTITATIVE VARIABLES, AND THE FISHER EXACT TEST FOR CATEGORICAL VARIABLES. ASSUMPTIONS OF NORMALITY AND HOMOGENEITY OF VARIANCE FOR T-TEST WERE EVALUATED. DIFFERENCES BETWEEN THE 3 GROUPS IN QUANTITATIVE VARIABLES WERE EVALUATED BY THE KRUSKAL¿WALLIS TEST. COMPARISON OF PROPORTIONS BETWEEN THE 3 GROUPS WAS PERFORMED BY USING THE FISHER-EXACT TEST. QUANTITATIVE VARIABLES WERE PRESENTED AS MEDIAN (INTERQUARTILE RANGE [IQR]). CATEGORICAL VARIABLES WERE EXPRESSED AS COUNTS AND PERCENTAGES. NORMALITY WAS ASSESSED BY THE SHAPIRO¿WILK NORMALITY TEST, THE D¿AGOSTINO¿PEARSON OMNIBUS NORMALITY TEST, AND VISUAL INSPECTION OF HISTOGRAMS. ALL P-VALUES WERE TWO-SIDED. RESULTS WERE CONSIDERED SIGNIFICANT IF P < 0.05. STATISTICAL ANALYSES WERE PERFORMED USING SAS VERSION 9.4 (SAS INSTITUTE INC., CARY, NC, USA). DURING THE 2018/2019 YELLOW FEVER OUTBREAK IN SÃO PAULO, 114 CONFIRMED CASES OF YELLOW FEVER WERE REFERRED TO THE ID-ICU OF HCFMUSP. SIXTY-SIX PATIENTS HAD LEVELS OF AMMONIA > 70 MCMOL/L AND/OR AN FV OF <50% AND WERE INCLUDED FOR ANALYSIS: 41 (62%) WERE TREATED WITH ICS (G1), 11 (17%) RECEIVED ICS PLUS HV-TPE (G2) AND 14 (21%) RECEIVED ICS PLUS INTENSIVE TPE (G3). THERE WERE NO DIFFERENCES AMONG THE THREE GROUPS RELATED TO SEX AND CLINICAL PRESENTATION AT ADMISSION. ALL GROUPS HAD VALUES OF AST/ALT > 3000 U/L, HOWEVER PATIENTS OF G3 HAD LOWER LEVELS OF AST/ALT COMPARED TO THE OTHER GROUPS, PROBABLY AS A CONSEQUENCE OF THE CHANGE IN THE TRANSFER CRITERIA, WHICH ALLOWED EARLIER ASSISTANCE, PREVENTING THE PROGRESSION OF THE DISEASE. NO DIFFERENCES WERE OBSERVED AMONG THE GROUPS WITH RESPECT TO THE OTHER LABORATORY TESTS. CONCERNING G3, THE PATIENTS WERE ADMITTED WITH A MEDIAN OF SIX DAYS OF SYMPTOMS AND FEVER WAS NOTED IN 86% OF THEM. SURPRISINGLY, JAUNDICE WAS PRESENT IN ONLY 43% OF THE PATIENTS AT ADMISSION. HEMORRHAGIC PHENOMENA OCCURRED IN 11 PATIENTS (79%), THE VAST MAJORITY BEING GASTROINTESTINAL BLEEDING. TWELVE PATIENTS (86%) HAD SEIZURES AND/OR HEADACHES AS NEUROLOGIC MANIFESTATIONS. ONE OF THE FEMALE PATIENTS (LO) WAS PREGNANT AND WAS DIAGNOSED WITH REACTIVE HEMOPHAGOCYTIC SYNDROME, WHICH WAS SUCCESSFULLY TREATED WITH HUMAN INTRAVENOUS IMMUNOGLOBULIN. THE MEDIAN VOLUME TREATED AND TIME PER PROCEDURE WERE HIGHER IN G2 (P = 0.005 AND P < 0.001, RESPECTIVELY) ALTHOUGH G3 HAD A HIGHER MEDIAN NUMBER OF PROCEDURES PER PATIENT (P < 0.001) (TABLE 1). HOWEVER, 45% OF THE G2 PATIENTS DIED BEFORE CONCLUDING THE PROPOSED TREATMENT (THREE CONSECUTIVE DAYS). TRANSFUSION REACTIONS SUCH AS TRANSFUSION-RELATED ACUTE LUNG INJURY, TRANSFUSION ASSOCIATED CIRCULATORY OVERLOAD OR ABO INCOMPATIBILITY WERE NOT OBSERVED. ONLY MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. HYPOCALCEMIA (IONIC CALCIUM < 4.0 MG/DL) WAS MORE FREQUENT IN G2 (P < 0.001). THE MORTALITY RATES IN G1 AND G2 WERE SIMILAR, AT 85% AND 82%, RESPECTIVELY (TABLE 1). THE MORTALITY RATE FOR PATIENTS IN G3 WAS SIGNIFICANTLY INFERIOR, AS DEMONSTRATED IN THE KAPLAN¿MEIER SURVIVAL CURVE (FIGURE 1). ONLY TWO PATIENTS DIED IN G3 (TABLE 2). ONE WAS A 48-YO MALE PATIENT (VPP) WHO REPORTED LONG-TERM USE OF ALCOHOL, WITH FOUR DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (403 AND 148 TIMES GREATER THAN THE UPPER LIMIT OF NORMAL [ULN], RESPECTIVELY), WITH HIGH AMMONIA LEVELS (3.7 TIMES THE ULN) AND LOW LEVELS OF FV (REDUCTION OF 90% FROM NORMAL). FOUR DIFFERENT BACTERIA WERE IDENTIFIED 48 HS LATER IN ADMISSION BLOOD CULTURE (E. COLI, S. VIRIDIANS, S. PNEUMONIAE AND F. NUCLEATUM). DESPITE ICS, ANTIBIOTICS, AND INTENSIVE TPE, THE PATIENT PRESENTED WITH A MASSIVE ALVEOLAR HEMORRHAGE AND DIED 72 HS AFTER ADMISSION (TABLE 2). THE SECOND WAS ALSO A 48-YO MALE PATIENT (VSS), PREVIOUSLY VACCINATED AGAINST YFV (10 YEARS BEFORE), WITH SIX DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (163 AND 103 TIMES THE ULN, RESPECTIVELY), HIGH AMMONIA LEVELS (2.3 TIMES THE ULN), AND NORMAL LEVELS OF FV. AFTER FIVE DAYS OF INTENSIVE TPE, ALL LABORATORY RESULTS WERE NORMALIZED. HOWEVER, ON THE SIXTH DAY OF ADMISSION, S. AUREUS WAS IDENTIFIED IN A BLOOD CULTURE AND THE PATIENT DIED SEVEN DAYS LATER. COMPARING G2 AND G3, THE EVOLUTION OF THE MEDIAN LEVELS OF AMMONIA EXHIBITED OPPOSITE DIRECTIONS: AN INCREASING TENDENCY FOR G2 AND A DECREASING TENDENCY FOR G3 (P = 0.0002). ON THE OTHER HAND, THE LEVELS OF FV SHOWED A SIMILAR INCREASING TENDENCY IN G2 AND G3, HOWEVER, WITH STATISTICAL SIGNIFICANCY (P = 0.014). (FIGURE 2A,B). THREE DAYS AFTER TPE (DAY 4), THE SERUM YF VIRAL LOAD BECAME UNDETECTABLE IN 91.7% OF G3 COMPARED TO 28.6% OF G2. THE HIGH MORTALITY RATES OF SEVERE YF FORMS DURING THE YF OUTBREAK IN SP DESPITE INTENSIVE CARE SUPPORT PROMPTED US TO USE HV-TPE FOR YF-ALF. IN OUR CASUISTIC, USING A SIMILAR LARGE-VOLUME EXCHANGE STRATEGY, WE COULD NOT OBSERVE ANY BENEFIT IN MORTALITY. G1 AND G2 HAD SIMILAR MORTALITY RATES DESPITE THE USE OF TPE. WE IMPUTE THIS FINDING TO THE FULMINANT ONSET AND AGGRESSIVE COURSE OF ALF-YF, FASTER THAN THAT OBSERVED IN OTHER ETIOLOGIES. THIS HYPOTHESIS WAS CORROBORATED BY THE SHORT SURVIVAL OF G1 PATIENTS AND BY THE TRANSIENT IMPROVEMENT OF LABORATORY RESULTS OBSERVED AFTER HV-TPE IN G2. THE CLOSE RELATIONSHIP BETWEEN ELEVATED AMMONIA LEVELS AND THE DEVELOPMENT OF ENCEPHALOPATHY IS WELL ESTABLISHED. DURING THE FIRST MONTH OF THE OUTBREAK, WE OBSERVED THAT A PRESENCE OF AMMONIA > 70 MCMOL/L AND/OR FV < 50% WERE ASSOCIATED WITH HIGHER MORTALITY. AMONG 59 PATIENTS WITH THIS LABORATORY PROFILE, THE FIRST 48 WERE TREATED WITH ICS OR LT, AND THIRTY-NINE (81%) OF THEM DIED. CLEMMESEN ET AL., USING HV-TPE IN 11 PATIENTS WITH ALF AND HEPATIC ENCEPHALOPATHY GRADE III-IV, DEMONSTRATED A REDUCTION OF ARTERIAL AMMONIA. NEVERTHELESS, IN OUR CASUISTIC USING A SIMILAR STRATEGY, WE OBSERVED AN INCREASING TREND OF AMMONIA LEVELS INSTEAD. IN LARSEN¿S STUDY, AN INITIAL SIGNIFICANT DECREASE OF AMMONIA LEVELS WAS OBSERVED ALTHOUGH IT WAS NOT SUSTAINED AFTER DAY THREE. IN ORDER TO OVERCOME THE AGGRESSIVE BEHAVIOR OF ALF-YF AND AIMING TO OBTAIN SAFE LEVELS OF AMMONIA, A LOWER VOLUME OF TPE PER SESSION AND A REDUCTION OF THE INTERVAL BETWEEN THE PROCEDURES WERE IMPLEMENTED (INTENSIVE TPE). THIS APPROACH ALLOWED A SUSTAINED DECREASE OF AMMONIA (FIGURE 2A, G3), RESULTING IN A HIGHER SURVIVAL RATE. THE STUDIES ON HV-TPE IN HEPATOTROPIC VIRAL DISEASES ARE LIMITED. LARSEN ET AL. INCLUDED ONLY 11 PATIENTS WITH ACUTE VIRAL HEPATITIS, AND SIX OF THEM SUBMITTED TO HV-TPE. ONLY TWO AMONG 11 PATIENTS TREATED BY CLEMMESEN ET AL. HAD VIRAL HEPATITIS. LIKEWISE, AKDOGAN ET AL., USING A STANDARD PLASMA EXCHANGE PROTOCOL, HAD ONLY 20% (8/39) CASES OF VIRAL HEPATITIS IN THEIR CASUISTIC. UNLIKE DRUG-INDUCED OR TOXIC HEPATITIS, IN VIRAL ETIOLOGIES THE CAUSAL AGENT IS STILL PRESENT AND THE LIVER INJURY PERSISTS. DESPITE THE FACT THAT COMMON KNOWLEDGE PRESUMES A VIREMIA OF FIVE DAYS IN YF, RECENT STUDIES DEMONSTRATED LONGER PERSISTENCE OF THE VIRUS IN THE SERUM AND OTHER ORGANS. ADDITIONALLY, IN A COHORT OF THE 2018 SP OUTBREAK, A HIGHER VIRAL LOAD WAS DEMONSTRATED TO BE AN INDEPENDENT PREDICTOR FACTOR ASSOCIATED WITH DEATH. AND, AS WE HAVE SHOWN, INTENSIVE TPE WAS ABLE TO REMOVE THE VIRUS, THUS REDUCING LIVER DAMAGE, ANOTHER POSSIBLE REASON TO JUSTIFY THE SUCCESS OF INTENSIVE TPE IN SEVERE YELLOW FEVER. ALTHOUGH TRANSAMINASE LEVELS IN G3 PATIENTS WERE LOWER THAN IN G2, THEY REMAINED MORE THAN 100 TIMES ABOVE THE UPPER LIMIT OF NORMAL, UNDERSCORING THE SEVERITY OF THE DISEASE. ADDITIONALLY, STUDIES HAVE INCREASINGLY DEMONSTRATED THAT THE YELLOW FEVER VIRUS EXHIBITS TROPISM FOR ORGANS BEYOND THE LIVER, WITH RNA OR ANTIGENS DETECTABLE IN THE MYOCARDIUM, LUNGS, KIDNEYS, PANCREAS, AND SPLEEN. FURTHERMORE, COAGULOPATHY HAS BEEN SHOWN TO OCCUR INDEPENDENTLY OF HEPATOCELLULAR TROPISM. THESE FINDINGS HIGHLIGHT THAT THE INVOLVEMENT OF MULTIPLE ORGANS SIGNIFICANTLY CONTRIBUTES TO THE SEVERITY OF THE DISEASE, AND THEREFORE SHOW THE ROLE OF TPE IN RESTORING METABOLIC AND IMMUNOLOGICAL BALANCE AND REPLENISHING COAGULATION FACTORS. INTENSIVE TPE APPARENTLY WAS SAFER WITH A LOWER INCIDENCE OF HYPOCALCEMIA COMPARED TO HV-TPE. IT WAS ALSO ASSOCIATED WITH A REDUCED MEDIAN VOLUME TREATED AND SHORTER TIME PER SESSION, ALLOWING SEQUENTIAL TREATMENT FOR A MAJOR NUMBER OF PATIENTS. THE POSSIBILITY OF SIMULTANEOUS TREATMENT FOR A LARGE NUMBER OF PATIENTS IS PARTICULARLY IMPORTANT CONSIDERING THE DRAMATIC INCREASE OF YF IN REGIONS WITH HEALTH RESOURCE CONSTRAINTS, AND THE SEASONAL PATTERN OF THE DISEASE IN OUTBREAKS. LT IS A RECOGNIZED APPROACH FOR THE MANAGEMENT OF ALF, HOWEVER THERE IS VERY RESTRICTED EXPERIENCE IN ITS USE AS A TREATMENT FOR ALF-YF. MOREOVER, THE HIGH COST OF LT AND THE SHORTAGE OF LIVER DONORS CURTAIL THE FEASIBILITY OF THE PROCEDURE. SOME LIMITATIONS OF OUR STUDY INCLUDE THE SMALL SAMPLE SIZE AND THE NEED FOR RANDOMIZED TRIALS TO CONFIRM THE EFFICACY OF THIS STRATEGY. HOWEVER, GIVEN THE FULMINANT PROGRESSION AND HIGH LETHALITY OF THE DISEASE, AN ADAPTIVE PRAGMATIC TRIAL COULD REPRESENT A FEASIBLE APPROACH TO DEMONSTRATE THIS TREATMENT¿S EFFECTIVENESS. WE DEMONSTRATED THAT IN PATIENTS WITH ALF SECONDARY TO YF, INTENSIVE TPE ASSOCIATED WITH GUIDED TRANSFUSION IS A FEASIBLE AND EFFECTIVE ARTIFICIAL SUPPORT TO KEEP PATIENTS ALIVE UNTIL HEPATIC REGENERATION AND RECOVERY OCCURS. THIS WAS DEMONSTRATED BY A REDUCTION OF 84% IN MORTALITY. THEREFORE, IT IS IMPORTANT TO SHARE OUR SUCCESSFUL EXPERIENCE, ALTHOUGH LIMITED TO A SMALL NUMBER OF CASES. UNDOUBTEDLY THERE ARE MANY KNOWLEDGE GAPS RELATED TO THE EFFECTS OF THERAPEUTIC PLASMA EXCHANGE IN THE COURSE OF YF AND OUR LIMITATIONS OF A SMALL SAMPLE SIZE, AND THE OBSERVATIONAL AND SINGLE-CENTER NATURE OF THIS STUDY. FURTHER STUDIES ARE REQUIRED TO BETTER UNDERSTAND THE ROLE OF TPE IN SEVERE YF. SINCE THIS WAS A JOURNAL PUBLICATION AN OBSERVATIONAL STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK, THE LOT NUMBERS WERE NOT REQUESTED; THEREFORE, A DISPOSABLE LOT HISTORY SEARCH COULD NOT BE CONDUCTED. ACCORDING TO THERAPEUTIC APHERESIS: A PHYSICIAN'S HANDBOOK, ADVERSE EVENTS OCCUR DURING THERAPEUTIC PROCEDURES WITH A FREQUENCY OF 4.8%. SYMPTOMS OF THESE ALLERGIC REACTIONS MAY INCLUDE HIVES, DYSPNEA, WHEEZING, BURNING EYES, TACHYCARDIA, HYPOTENSION, AND OR FACIAL SWELLING AND FLUSHING. MILD REACTIONS CAN BE TREATED WITH DIPHENHYDRAMINE ADMINISTERED THROUGH AN IV. ACCORDING TO THERAPEUTIC APHERESIS: A PHYSICIAN'S HANDBOOK, ADVERSE EVENTS OCCUR DURING THERAPEUTIC PROCEDURES WITH A FREQUENCY OF 4.8%. SOME OF THE MOST COMMON REACTIONS INCLUDE FEVER, URTICARIA, HYPOCALCEMIC SYMPTOMS, PRURITUS, DYSPNEA, TACHYCARDIA, AND MILD HYPOTENSION. TRANSIENT HYPOCALCEMIA ASSOCIATED WITH APHERESIS IS USUALLY WELL TOLERATED. SYMPTOMS OFTEN SHOW AS PARESTHESIA (TINGLING) BUT PATIENTS MAY ALSO EXPERIENCE UNUSUAL TASTE, NAUSEA, LIGHTHEADEDNESS, SHIVERING, AND TREMORS. SEVERE HYPOCALCEMIA MAY ALSO CAUSE MUSCLE CONTRACTIONS AND CAN PROGRESS TO TETANY AND SEIZURES IF HYPOCALCEMIA ESCALATES AND IS NOT CORRECTED. SINCE THIS WAS A JOURNAL PUBLICATION AN OBSERVATIONAL STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK, THE LOT NUMBERS WERE NOT REQUESTED; THEREFORE, A DHR SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA FOR RELEASE. BASED ON THE CLINICAL INFORMATION PROVIDED, THE SPECTRA OPTIA PERFORMED AS INTENDED, THERE WERE NO SUGGESTED DEVICE FAILURES, MALFUNCTIONS, MISLABELING, IMPROPER OR INADEQUATE DESIGN OR MANUFACTURING ERRORS, NOR WAS THERE ANY REPORTED USER ERRORS THAT CONTRIBUTED TO OR CAUSED ADVERSE EVENTS. ROOT CAUSE: A ROOT CAUSE ASSESSMENT WAS PERFORMED FOR THE FIRST PATIENT DEATH. ACCORDING TO THE AUTHORS, ¿ONE WAS A 48-YO MALE PATIENT (VPP) WHO REPORTED LONG-TERM USE OF ALCOHOL, WITH FOUR DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (403 AND 148 TIMES GREATER THAN THE UPPER LIMIT OF NORMAL [ULN], RESPECTIVELY), WITH HIGH AMMONIA LEVELS (3.7 TIMES THE ULN) AND LOW LEVELS OF FV (REDUCTION OF 90% FROM NORMAL). FOUR DIFFERENT BACTERIA WERE IDENTIFIED 48 HRS LATER IN ADMISSION BLOOD CULTURE (E. COLI, S. VIRIDIANS, S. PNEUMONIAE AND F. NUCLEATUM). DESPITE ICS, ANTIBIOTICS, AND INTENSIVE TPE, THE PATIENT PRESENTED WITH A MASSIVE ALVEOLAR HEMORRHAGE AND DIED 72 HS AFTER ADMISSION.¿ THE PATIENT IN QUESTION HAD UNDERLYING LIVER DISEASE DUE TO HIS LONG-TERM ALCOHOL USE WHICH PREDISPOSED HIM TO RISKS OF ADDITIONAL COMPLICATIONS. THIS COMBINED WITH YELLOW FEVER WOULD HAVE WORSENED HIS ALREADY COMPROMISED LIVER FUNCTION. A ROOT CAUSE ASSESSMENT WAS PERFORMED FOR THE SECOND PATIENT DEATH. ACCORDING TO THE AUTHORS, ¿THE SECOND WAS ALSO A 48-YO MALE PATIENT (VSS), PREVIOUSLY VACCINATED AGAINST YFV (10 YEARS BEFORE), WITH SIX DAYS OF SYMPTOMS. AST AND ALT LEVELS WERE MARKEDLY ELEVATED (163 AND 103 TIMES THE ULN, RESPECTIVELY), HIGH AMMONIA LEVELS (2.3 TIMES THE ULN), AND NORMAL LEVELS OF FV. AFTER FIVE DAYS OF INTENSIVE TPE, ALL LABORATORY RESULTS WERE NORMALIZED. HOWEVER, ON THE SIXTH DAY OF ADMISSION, S. AUREUS WAS IDENTIFIED IN A BLOOD CULTURE AND THE PATIENT DIED SEVEN DAYS LATER. BASED ON THE AVAILABLE INFORMATION A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BUT IT IS LIKELY DUE TO PROGRESSION OF THE PATIENT¿S DISEASE AND/OR COMPLICATIONS ASSOCIATED WITH YELLOW FEVER. A ROOT CAUSE ASSESSMENT WAS PERFORMED FOR THE REPORTED MILD ALLERGIC SYMPTOMS OCCURRED DURING OR SOON AFTER THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED BUT IT IS LIKELY DUE TO ONE OR A COMBINATION OF THE POSSIBLE CAUSES LISTED BELOW: HYPERSENSITIVITY TO THE ETHYLENE OXIDE USED TO STERILIZE THE DISPOSABLE SET. HYPERSENSITIVITY TO THE COMPONENTS INSIDE THE DISPOSABLE SET. ALLERGIC REACTIONS TO THE REPLACEMENT FRESH FROZEN PLASMA. ANAPHYLACTIC REACTIONS HAVE BEEN REPORTED IN IGA-DEFICIENT PATIENTS WITH ANTI-IGA ANTIBODIES AND PATIENTS WITH HAPTOGLOBIN DEFICIENCY, MOST REACTIONS ARE IDIOSYNCRATIC AND NOT ASSOCIATED WITH A SPECIFIC SERUM PROTEIN DEFICIENCY, POLYMORPHISM, OR IDENTIFIABLE CAUSE. IMMEDIATE HYPERSENSITIVITY REACTION (PRODUCTION OF IGE ANTIBODIES IN RESPONSE TO ALLERGENS (PLASMA PROTEINS IN FRESH FROZEN PLASMA). PASSIVE SENSITIZATION (PASSIVE TRANSFER OF DONOR ALLERGENS OR DONOR ANTIBODIES (E.G. ALLERGEN-SPECIFIC IGE); TRANSIENT ALLERGIC SENSITIVITY TO ALLERGENS. A ROOT CAUSE ASSESSMENT WAS PERFORMED FOR THE REPORTED HYPOCALCEMIA. THESE REACTIONS OCCUR DUE TO DECREASED IONIZED CALCIUM IN CIRCULATION AS A RESULT OF EXOGENOUS CITRATE ADMINISTERED DURING THE APHERESIS PROCEDURE AND ARE INFLUENCED BY PATIENT PHYSIOLOGY, THE PATIENT'S DISEASE STATE, THE RATE OF AC INFUSION, THE CITRATE CONTENTS IN THE REPLACEMENT FLUID, AND/OR THE LENGTH OF THE PROCEDURE. THESE SYMPTOMS MAY BE TREATED WITH ORAL OR INTRAVENOUS CALCIUM SUPPLEMENTS OR BY ADJUSTING THE AC INFUSION RATE. HO, Y., NUKUI, Y., VILLAÇA, P. R., OKAZAKI, E., TATSUI, N. H., NETTO, L. C., JOELSONS, D., DA ROCHA, T. R. F., DE MELLO MALTA, F., PINHO, J. R. R., SEGURADO, A. A. C., & ROCHA, V. (2025). INTENSIVE THERAPEUTIC PLASMA EXCHANGE¿NEW APPROACH TO TREAT AND RESCUE PATIENTS WITH SEVERE FORM OF YELLOW FEVER. TROPICAL MEDICINE AND INFECTIOUS DISEASE, 10(2), 39. HTTPS://DOI.ORG/10.3390/TROPICALMED10020039. Manufacturer Narrative: INVESTIGATION: THIS IS AN OBSERVATIONAL CASE-SERIES UNICENTRIC STUDY EVALUATING TWO DIFFERENT TPE STRATEGIES FOR PATIENTS WITH A SEVERE FORM OF YELLOW FEVER DURING THE SAO PAULO, BRAZIL 2018¿2019 YELLOW FEVER OUTBREAK. IT WAS CONDUCTED AT HOSPITAL DAS CLINICAS (HCFMUSP), A 2200-BED TERTIARY CARE TEACHING HOSPITAL AFFILIATED TO THE (B)(6) WITH 110 INTENSIVE CARE BEDS AND A 7-BED INFECTIOUS DISEASE INTENSIVE CARE UNIT (ID-ICU). ON (B)(6) 2018, THE ID-ICU HAD BECOME THE CENTER OF EXPERTISE FOR SEVERE FORMS OF YF IN SP. SEVERITY CRITERIA FOR ID-ICU ADMISSION WERE: ASPARTATE AMINOTRANSFERASE (AST) OR ALANINE AMINOTRANSFERASE (ALT) > 3000 U/L; INR > 1.5; PLATELET COUNT < 90 × 109/L; KIDNEY DYSFUNCTION; HEMORRHAGIC MANIFESTATIONS; ENCEPHALOPATHY; OR HEMODYNAMIC INSTABILITY. HOWEVER, IN 2019, A YELLOW FEVER OUTBREAK OCCURRED IN A REGION OF THE STATE OF SAO PAULO WHERE TRANSPORTATION TO HCFMUSP TAKES AT LEAST 6 H. FOR THIS REASON, ALL PATIENTS WHO HAD ALT/AST > 500 WERE REFERRED TO HCFMUSP. FROM (B)(6) 2018, TREATMENT CONSISTED OF STANDARD INTENSIVE CARE SUPPORT (ICS). IN THE SECOND HALF OF (B)(6) 2018, HIGH-VOLUME TPE (HV-TPE) WAS USED IN SOME PATIENTS WITH SEVERE YF IN OUR SERVICE. TRANSIENT IMPROVEMENT OF LABORATORY RESULTS WAS OBSERVED; HOWEVER HEMORRHAGIC PHENOMENA OCCURRED BETWEEN SESSIONS OF HV-TPE WITHOUT CHANGING THE OUTCOME. SINCE (B)(6) 2018, THIS STRATEGY WAS MODIFIED TO MORE INTENSIVE TPE TWICE A DAY AND SUPPORTIVE TRANSFUSION, IF NECESSARY. THIS INTENSIVE PROTOCOL WAS USED ON ONE PATIENT IN 2018 SUCCESSFULLY, AND SINCE THEN IT HAS BEEN USED IN THE (B)(6) 2019 YF OUTBREAK. TPE WAS ADMINISTERED TO PATIENTS WITH LEVELS OF AMMONIA > 70 MCMOL/L AND/OR FACTOR V (FV) < 50%. THE DEFINITION OF THE CUT-OFF OF AMMONIA WAS BASED ON THE 2018 EXPERIENCE WHEN PATIENTS WITH AMMONIA > 70 MCMOL/L HAD A HIGHER RELATIVE RISK OF DEATH (RR 2.0) (DATA NOT PUBLISHED). THE CUT-OFF OF FV < 50% WAS DEFINED BASED ON CLICHY¿S CRITERIA, CONSIDERING A REDUCTION OF 30% FROM THE REFERENCE VALUE OF 70¿150% FOR ALL PATIENTS. IN ORDER TO COMPARE TO PREVIOUS CASES, ALL PATIENTS WITH AMMONIA > 70 MCMOL/L AND/OR FV < 50% WERE SELECTED. FOR THIS ANALYSIS, PATIENTS WERE CLASSIFIED INTO THREE GROUPS ACCORDING TO TYPE OF TREATMENT GIVEN: GROUP ONE (G1)¿ICS; GROUP TWO (G2)¿ICS AND HV-TPE; AND GROUP THREE (G3)¿ICS AND INTENSIVE TPE. CLINICAL AND LABORATORY DATA WERE EXTRACTED RETROSPECTIVELY AND STORED IN THE REDCAP WEB PLATFORM. HV-TPE STRATEGY (G2): PROPOSED TREATMENT CONSISTED OF THREE CONSECUTIVE DAYS OF PLASMA EXCHANGE. THE VOLUME CHANGED WAS 10% OF BODY WEIGHT (BW) WITH FRESH FROZEN PLASMA (FFP). ONE TO THREE ADDITIONAL SESSIONS OF ONE PLASMA VOLUME EACH WERE IMPLEMENTED ACCORDING TO AMMONIA LEVELS. SUPPORTIVE TRANSFUSION WAS GUIDED MAINLY BY HEMORRHAGIC EPISODES. INTENSIVE-TPE STRATEGY (G3): VOLUME OF EXCHANGE WITH FFP WAS CALCULATED ACCORDING TO BW, HEMATOCRIT (HT), AND PLASMA VOLUME (BW × 0.7 × [1 - HT]), CONSIDERING ONE PLASMA VOLUME PER SESSION. PROCEDURE WAS PERFORMED TWICE A DAY, WITH ADDITIONAL FFP INFUSION (10 ML/KG) AT NIGHT. PLATELET CONCENTRATE WAS TRANSFUSED TO KEEP PLATELET COUNT > 30 × 109/L OR, IN THE PRESENCE OF BLEEDING, >50 × 109/L. IN SEVERE OR LIFE-THREATENING BLEEDS, TARGET WAS >80¿100 × 109/L. CRYOPRECIPITATE TRANSFUSION WAS INDICATED IF FIBRINOGEN LEVEL WAS <100 MG/DL. BASED ON BLEEDING AS WELL AS FV AND/OR AMMONIA LEVELS, THE FREQUENCY OF TPE WAS TAPERED TO ONCE A DAY, EVERY OTHER DAY AND DISCONTINUATION. DURING THE DE-ESCALATION PERIOD, FFP WAS INFUSED (5 ML/KG) 1¿2 TIMES/DAY, FOLLOWING THE SAME CRITERIA ABOVE. TRANEXAMIC ACID (10 MG/KG/DOSE, IV, EVERY 6¿8 H) WAS PRESCRIBED UNTIL DISCONTINUATION OF TPE AND TOPICAL USE WAS IMPLEMENTED IN CASES OF MUCOCUTANEOUS BLEEDING. TPE WAS PERFORMED USING FFP REPLACEMENT, WITHOUT ANTICOAGULANT (CITRATE), ONCE ALL PATIENTS HAD ALF. CALCIUM CHLORIDE WAS ADMINISTERED AT THE FOLLOWING RATES: FOR HV-TPE GROUP, 5 MMOL PER EACH LITER OF PLASMA EXCHANGE; FOR INTENSIVE-TPE GROUP, FIXED 6.8 MMOL OF CALCIUM PER SESSION. AFTER EACH SESSION, SERUM CALCIUM LEVEL WAS ASSESSED, AND ADDITIONAL SUPPLEMENTATION WAS CONSIDERED. COBE SPECTRA AND OPTIA (TERUMOBCT, TOKYO, JAPAN) AND COM.TECR (FRESENIUS KABI, BAD HOMBURG, GERMANY) TPE DEVICES WERE EMPLOYED, USING A WITHDRAWAL FLOW OF 70 ML/MINUTE ON AVERAGE. IN CASES WHERE CONTINUOUS VENOVENOUS HEMODIALYSIS WAS INDICATED, THE PLASMA EXCHANGE SYSTEM WAS CONNECTED TO THE RETURN CIRCUIT OF THE DIALYSIS SYSTEM. YFV-RNA WAS ISOLATED FROM 250 L OF SERUM SAMPLES AND 20 G OF RNASE-FREE GLYCOGEN WERE ADDED TO THE AQUEOUS PHASE AS AN RNA CARRIER. R... Edited for length. See MAUDE Adverse Event Report for complete event narratives.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21861343&pc=MNS | 2518422-2025-104346 | 2025-03-07 05:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | MNS | DREAMSTATION BIPAP AUTOSV | Degraded; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH DREAMSTATION BIPAP AUTOSV DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: IN THE PREVIOUS REPORT, PRODUCT BRAND NAME, PRODUCT UDI, DEVICE SERVICED BY 3RDP?, DEVICE AVAIL. FOR EVAL?, DEVICE MANUFACTURE DATE, REASON DEVICE NOT EVALUATED, (DEVICE) PROBLEM CODE GRID, REMEDIAL ACTION INITIATED AND RECALL (Z) NUMBER SECTION WHICH WERE SELECTED INCORRECTLY, HAS BEEN UPDATED/ CORRECTED IN THIS FOLLOW UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21861468&pc=QFG | 3013756811-2025-68465 | 2025-02-24 05:00 | Death | TANDEM DIABETES CARE | 2025-04-17 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DUE TO DIABETIC KETOACIDOSIS. A REVIEW OF THE PUMP DATA LOGS IS EXPECTED; HOWEVER, THE PUMP DATA LOG REVIEW HAS NOT YET BEEN COMPLETED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21861858&pc=BZD | 2518422-2025-104387 | 2025-03-21 04:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21862394&pc=BZD | 2518422-2025-104395 | 2024-11-15 05:00 | Death | RESPIRONICS, INC. | 2025-04-17 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21862567&pc=QEW | 3005168196-2025-00125 | 2025-03-22 04:00 | Death | PENUMBRA, INC. | 2025-04-17 04:00 | QEW | INDIGO SYSTEM FLASH ASPIRATION CATHETER | Adverse Event Without Identified Device or Use Problem | Hemoptysis | K240030 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE TO TREAT AN ACUTE PULMONARY EMBOLISM IN THE LEFT PULMONARY ARTERY (LPA) AND RIGHT SEGMENTAL BRANCHES USING AN INDIGO SYSTEM LIGHTNING FLASH ASPIRATION TUBING (FLASH TUBING) AND AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER). IT WAS NOTED THAT THE PATIENT HAD A RIGHT VENTRICULAR (RV) STRAIN. DURING THE PROCEDURE, THE FLASH CATHETER WAS ADVANCED INTO THE LPA WITHOUT ISSUES AND ASPIRATION WAS INITIATED. AFTER APPROXIMATELY THREE MINUTES OF ASPIRATION, THE PATIENT BEGAN TO COUGH UP BLOOD AND THE PROCEDURE WAS ABORTED. THE PHYSICIAN THEN ADVANCED A BALLOON CATHETER TO STABILIZE THE PATIENT AND PERFORMED AN ANGIOGRAM. HOWEVER, THE PHYSICIAN COULD NOT FIND AN AREA OF PERFORATION. IT WAS REPORTED THAT THE PATIENT DEVELOPED AIRWAY OBSTRUCTION, AND CARDIOPULMONARY ARREST. RESUSCITATION EFFORTS THAT INCLUDED ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) INTERVENTION, CRICOTHYROTOMY, AND MECHANICAL VENTILATION WERE PERFORMED. THE PATIENT WAS ALSO ADMINISTERED PROTAMINE SULFATE, TWO UNITS OF FRESH RED BLOOD CELLS (PRBCS), AND TWO UNITS OF FRESH FROZEN PLASMA (FFP). THE PATIENT WAS ABLE TO REGAIN A STRONG PULSE FOR A SHORT TIME. AFTER APPROXIMATELY 37 MINUTES, THE PATIENT EXPIRED. THE DEATH WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM, AND A DEFINITE RELATIONSHIP TO THE PROCEDURE.Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE TO TREAT AN ACUTE PULMONARY EMBOLISM IN THE LEFT PULMONARY ARTERY (LPA) AND RIGHT SEGMENTAL BRANCHES USING AN INDIGO SYSTEM LIGHTNING FLASH ASPIRATION TUBING (FLASH TUBING) AND AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER). IT WAS NOTED THAT THE PATIENT HAD A RIGHT VENTRICULAR (RV) STRAIN. DURING THE PROCEDURE, THE FLASH CATHETER WAS ADVANCED INTO THE LPA WITHOUT ISSUES AND ASPIRATION WAS INITIATED. AFTER APPROXIMATELY THREE MINUTES OF ASPIRATION, THE PATIENT BEGAN TO COUGH UP BLOOD AND THE PROCEDURE WAS ABORTED. THE PHYSICIAN THEN ADVANCED A BALLOON CATHETER TO STABILIZE THE PATIENT AND PERFORMED AN ANGIOGRAM. HOWEVER, THE PHYSICIAN COULD NOT FIND AN AREA OF PERFORATION. IT WAS REPORTED THAT THE PATIENT DEVELOPED AIRWAY OBSTRUCTION, AND CARDIOPULMONARY ARREST. RESUSCITATION EFFORTS THAT INCLUDED ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) INTERVENTION, CRICOTHYROTOMY, AND MECHANICAL VENTILATION WERE PERFORMED. THE PATIENT WAS ALSO ADMINISTERED PROTAMINE SULFATE, TWO UNITS OF FRESH RED BLOOD CELLS (PRBCS), AND TWO UNITS OF FRESH FROZEN PLASMA (FFP). THE PATIENT WAS ABLE TO REGAIN A STRONG PULSE FOR A SHORT TIME. AFTER APPROXIMATELY 37 MINUTES, THE PATIENT EXPIRED. THE DEATH WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM, AND A DEFINITE RELATIONSHIP TO THE PROCEDURE. ON 27-MAY-2025 AN UPDATE WAS RECEIVED INDICATING THAT THE ADVERSE EVENT WAS UPDATED TO HEMOPTYSIS AND THAT THE EVENT WAS UNRELATED TO THE INDIGO ASPIRATION SYSTEM.Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE TO TREAT AN ACUTE PULMONARY EMBOLISM IN THE LEFT PULMONARY ARTERY (LPA) AND RIGHT SEGMENTAL BRANCHES USING AN INDIGO SYSTEM LIGHTNING FLASH ASPIRATION TUBING (FLASH TUBING) AND AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER). IT WAS NOTED THAT THE PATIENT HAD A RIGHT VENTRICULAR (RV) STRAIN. DURING THE PROCEDURE, THE FLASH CATHETER WAS ADVANCED INTO THE LPA WITHOUT ISSUES AND ASPIRATION WAS INITIATED. AFTER APPROXIMATELY THREE MINUTES OF ASPIRATION, THE PATIENT BEGAN TO COUGH UP BLOOD AND THE PROCEDURE WAS ABORTED. THE PHYSICIAN THEN ADVANCED A BALLOON CATHETER TO STABILIZE THE PATIENT AND PERFORMED AN ANGIOGRAM. HOWEVER, THE PHYSICIAN COULD NOT FIND AN AREA OF PERFORATION. IT WAS REPORTED THAT THE PATIENT DEVELOPED AIRWAY OBSTRUCTION, AND CARDIOPULMONARY ARREST. RESUSCITATION EFFORTS THAT INCLUDED ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) INTERVENTION, CRICOTHYROTOMY, AND MECHANICAL VENTILATION WERE PERFORMED. THE PATIENT WAS ALSO ADMINISTERED PROTAMINE SULFATE, TWO UNITS OF FRESH RED BLOOD CELLS (PRBCS), AND TWO UNITS OF FRESH FROZEN PLASMA (FFP). THE PATIENT WAS ABLE TO REGAIN A STRONG PULSE FOR A SHORT TIME. AFTER APPROXIMATELY 37 MINUTES, THE PATIENT EXPIRED. THE DEATH WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM, AND A DEFINITE RELATIONSHIP TO THE PROCEDURE. ON 27-MAY-2025 AN UPDATE WAS RECEIVED INDICATING THAT THE ADVERSE EVENT WAS UPDATED TO HEMOPTYSIS AND THAT THE EVENT WAS UNRELATED TO THE INDIGO ASPIRATION SYSTEM. ON 19-JUNE-2025, ADDITIONAL INFORMATION INDICATED THAT THE CLINICAL EVENTS COMMITTEE (CEC) ADJUDICATED THE EVENT TO HAVE A PROBABLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND THE INDEX PROCEDURE. ON 9-JULY-2025, THE CEC UPDATED THE ADJUDICATION FROM PROBABLE TO POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND THE INDEX PROCEDURE. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTIONS ARE BEING UPDATED BASED ON ADDITIONAL INFORMATION PROVIDED BY PENUMBRA CLINICAL TEAM ON 19-JUN-25: 1. SECTION B. BOX 2. OUTCOMES ATTRIBUTED TO ADVERSE EVENT. 2. SECTION B. BOX 5. DESCRIBE EVENT OR PROBLEM. 3. SECTION H. BOX 6. HEALTH EFFECT IMPACT CODE 3. NOTE: THE PATIENT'S DEATH WAS PREVIOUSLY REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND A DEFINITE RELATIONSHIP TO THE PROCEDURE. ON 27-MAY-25, THE ADVERSE EVENT WAS UPDATED FROM DEATH TO HEMOPTYSIS. THE HEMOPTYSIS WAS REPORTED TO BE UNRELATED TO THE INDIGO ASPIRATION SYSTEM. HOWEVER, ON 19-JUNE-2025, THE CLINICAL EVENTS COMMITTEE (CEC) ADJUDICATED THE ADVERSE EVENT OF HEMOPTYSIS TO HAVE A PROBABLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND THE INDEX PROCEDURE. ON 09-JULY-2025, THE CEC UPDATED THE ADJUDICATION FROM PROBABLE TO POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND THE INDEX PROCEDURE. PER THE AUTOPSY REPORT, THE PATIENT'S CAUSE OF DEATH WAS DUE TO ONGOING HEMORRHAGE AS A RESULT OF REPERFUSION INJURY AFTER THE PULMONARY EMBOLISM WAS EVACUATED FROM THE PULMONARY VASCULATURE. THERE WAS NO EVIDENCE OF MECHANICAL DISRUPTION TO THE VESSELS OR THE AIRWAYS. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTION WAS INADVERTENTLY MISSED ON THE INITIAL MFR REPORT AND IS BEING UPDATED ON THIS FOLLOW-UP MFR REPORT: 3004168196-2025-00125. SECTION H. BOX 6. HEALTH EFFECT IMPACT CODE 3. PLEASE NOTE THE FOLLOWING UPDATED INFORMATION: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTION IS BEING UPDATED BASED ON ADDITIONAL INFORMATION PROVIDED BY PENUMBRA CLINICAL TEAM ON 27-MAY-2025: 1. SECTION B. BOX 5. DESCRIBE EVENT OR PROBLEM. NOTE: THE DEATH WAS PREVIOUSLY REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND A DEFINITE RELATIONSHIP TO THE PROCEDURE. ON 27-MAY-2025, AN UPDATE WAS RECEIVED INDICATING THAT THE CAUSE OF DEATH WAS HEMOPTYSIS. HEMOPTYSIS WAS UNRELATED TO THE INDIGO ASPIRATION SYSTEM. THEREFORE, THIS EVENT IS NOT CONSIDERED REPORTABLE AGAINST THE INDIGO ASPIRATION SYSTEM. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. HEMOPTYSIS AND DEATH ARE INCLUDED AS POSSIBLE COMPLICATIONS IN THE INSTRUCTIONS FOR USE (IFU) FOR THE INDIGO ASPIRATION SYSTEM. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21862655&pc=OZD | 1220648-2025-27639 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND SOME TIME THEREAFTER IMPLANT, SAME DAY, THE PATIENT HAD ACCESS SITE BLEEDING. HEPARIN WAS WITHHELD AND TWO UNITS OF BLOOD PRODUCTS WERE ADMINISTERED. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: THERE ARE LIMITED DETAILS AND THE IMPACT FROM THE IMPELLA-RELATED EVENT IS UNKNOWN. ALTHOUGH CARE WAS WITHDRAWN, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21862817&pc=OZD | 1220648-2025-27648 | 2025-03-20 04:00 | Death | ABIOMED, INC. | 2025-04-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED LOW PUMP FLOW RESULTING IN DEVICE REPLACEMENT. HOWEVER, THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT CAME IN SEPTIC AND FOUND TO HAVE ENDOCARDITIS FROM LIVER DISEASE. ECHOCARDIOGRAM SHOWED AORTIC INSUFFICIENCY AND POSSIBLE VENTRICULAR SEPTAL DEFECT. CORONARIES WERE NOTED AS CLEAN. EJECTION FRACTION WAS 25% AND LEFT VENTRICULAR END-DIASTOLIC PRESSURE WAS 32. THE DECISION WAS MADE TO PLACE AN AXILLARY IMPELLA CP DEVICE WHICH WAS SUCCESSFULLY PLACED. FLOWS WERE 3.6 LITERS PER MINUTE AT PERFORMANCE LEVEL P-9. THE PEEL-AWAY SHEATH WAS REMOVED, AND REPOSITIONING SHEATH WAS PLACED WITH DEVICE MOVED IN VENTRICLE. THE DEVICE WAS REPOSITIONED, AND THE FLOWS WOULD NOT GO ABOVE 1.6 LITERS PER MINUTE CONSEQUENTLY, THE DEVICE WAS REMOVED, CLEANED AND RAN IN FLUIDS WITH FLOWS STILL NOT GOING ABOVE 1.6 LITERS PER MINUTE. CONSEQUENTLY, THE DEVICE WAS REPLACED WITH A NEW IMPELLA CP DEVICE. IT WAS NOTED THERE WAS NO CLOT OBSERVED ON THE IMPELLA CP DEVICE (PUMP 1) AFTER REMOVAL. THE PATIENT WOULD EXPIRE APPROXIMATELY TWO DAYS LATER. THE SURGICAL TEAM AND FAMILY DECIDED THAT AORTIC VALVE SURGERY WOULD BE "EXTREMELY" RISKY, AND THE DECISION WAS MADE TO REMOVE LIFE-SUSTAINING MEASURES INCLUDING IMPELLA CP SUPPORT. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE WAS A DELAY IN THE NEEDED MECHANICAL CIRCULATORY SUPPORT WHICH MAY HAVE ADDED TO THE ALREADY POOR PROGNOSIS SITUATION. ALTHOUGH THE IMPELLA IS NOT THE CAUSE OF DEATH, THE IMPACT TO THE OUTCOME IS UNKNOWN. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE TH IMPELLA CP DEVICE PUMP 1 AS AN ASSOCIATED FACTOR TO THE EVENT OUTCOME. THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE RETURNED COMPLAINT DEVICE VISUALLY INSPECTED. DAMAGED IMPELLER BLADES FOUND. NO BIOMATERIAL, NO CANNULA KINK OBSERVED. LOW PUMP FLOW: THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS DAMAGED IMPELLER BLADES DUE TO INTERACTION WITH OTHER DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21846523&pc=PNJ | 9612164-2025-01930 | 2025-04-15 04:00 | Death | MEDTRONIC IRELAND | 2025-04-16 04:00 | PNJ | MICRA¿ AV2 | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure | P150033 |  | 1 | Event Description: THE PATIENT`S REP REPORTED THAT APPROXIMATELY ONE POST-IMPLANT, THE PATIENT DIED FROM HEART FAILURE, AND THE DEVICE WAS RELATED TO THE PATIENT`S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847202&pc=MVK | 3008642652-2025-03935 | 2025-03-09 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-16 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 20:10:35 ON (B)(6) 2025. AT 13:42:11 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 13:42:59, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 13:44:53, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 10 BPM WITH CPR/MOTION ARTIFACT. AT 13:45:47, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 10 BPM WITH CPR/MOTION ARTIFACT. AT 13:46:41, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS IDIOVENTRICULAR RHYTHM @ 10 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 10 BPM WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 14:43:16 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847454&pc=OZD | 1220648-2025-27522 | 2025-03-18 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT STILL HAD SIGNIFICANT HEMOLYSIS DUE TO IMPELLA PIGTAIL BEING WRAPPED IN MITRAL APPARATUS. PLASMA FREE HEMOGLOBIN REMAINED ELEVATED. THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5.Event Description: ADDITIONAL INFORMATION WAS FURTHER REPORTED THAT STATED THAT WHILE THE PATIENT HAD CONTINUED HEMOLYSIS, THE PATIENT WAS ESCALATED TO THE IMPELLA 5.5. LATER ON, THE PATIENT DEVELOPED A MASSIVE GASTROINTESTINAL BLEED WHICH REQUIRED MULTIPLE UNITS OF BLOOD. HOWEVER, DUE TO THE PATIENT'S CONDITION, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER, ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE, ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS, ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS, INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES, MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION, ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE REGARDING THE GASTROINTESTINAL HEMORRHAGE A SUPPLEMENTAL REPORT WILL BE SUBMITTED HEMOLYSIS: CLINICAL STATES THAT OLD CP WAS SWAPPED OUT AND THE NEW CP HAS SIGNIFICANT HEMOLYSIS DUE TO THE PIGTAIL BEING WRAPPED IN THE MITRAL APPARATUS. NO OTHER INFORMATION WAS PROVIDED. PUMP WAS NOT RETURNED. PARTIAL DATA LOGS WERE PROVIDED. THEREFORE, THE ROOT CAUSE OF THE HEMOLYSIS ISSUE WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE AS MENTIONED IN THE CLINICAL INFORMATION. POSITIONING ISSUE: CLINICAL STATES THAT THE IMAGING DONE FOR INITIAL PLACEMENT, CONFIRMED THAT THE PUMP WAS NOT IN THE CORRECT POSITION AND WAS FACING THE MITRAL APPARATUS. REPOSITIONING WAS UNSUCCESSFUL. PUMP WAS NOT RETURNED FOR INVESTIGATION. THEREFORE, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND THE REASON FOR THE IMPROPER POSITIONING IS UNKNOWN. ADDITIONAL INFORMATION: B2: SELECTED DEATH AND ADDED DATE OF DEATH. B5-ADDITIONAL EVENT DETAILS LEADING TO PATIENT DEATH. H6-ADDED CLINICAL CODE 4476, IMPACT CODE 1802, 4643. MDR MANUFACTURER REPORT #1220648-2025-27846 WAS FILED ON THE IMPELLA 5.5 FOR THE GASTROINTESTINAL BLEED. MDR MANUFACTURER REPORT # 1220648-2025-27850 WAS FILED ON THE IMPELLA RP FOR THE GASTROINTESTINAL BLEED. BOTH THE IMPELLA 5.5 AND IMPELLA RP WERE IN-SITU AT THE TIME OF GI BLEED. MEDICAL SAFETY OFFICER REVIEW: THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE PATIENT'S SIGNIFICANT HEMOLYSIS WHILE USING THE IMPELLA CP DEVICE (FIRST IMPLANTED IMPELLA) AND CONTINUED WITH THE ESCALATION TO THE IMPELLA 5.5, COULD HAVE BEEN ONE OF MULTIPLE REASONS THAT THE MASSIVE BLEED OCCURRED DAYS LATER WHEN ON BIPELLA (IMPELLA 5.5 AND RP FLEX). THE PATIENT'S DEMISE CASCADED FROM THE CP MALFUNCTION CAUSING SIGNIFICANT HEMOLYSIS AND COULD NOT BE DISSOCIATED FROM THE OUTCOME OF DEATH. THERE WAS ADDITIONAL INTERVENTION AND THE PATIENT REQUIRED AN ESCALATION BECAUSE OF HIGHER THERAPY. THE PATIENT WAS IN POOR CONDITION AND SUBSEQUENTLY MET HIS/HER DEMISE DUE TO SIGNIFICANT GI BLEED IN WHICH IT CANNOT BE SAID IT IS NOT ASSOCIATED WITH THE IMPELLA DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847464&pc=BZD | 2518422-2025-104169 | 2025-04-14 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847489&pc=OZD | 1220648-2025-27523 | 2025-03-18 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A CRITICAL PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. BLEEDING WAS SUSPECTED, BUT STAFF COULD NOT SEE ANY ACTIVE BLEED UNDER FLUORO, CONSIDERING DISSECTION. GIVING EPI PUSHES, FLUID, BLOOD PRODUCT AND INTERMITTENTLY CODING PATIENT FOR LOW BLOOD PRESSURE. BLEEDING WAS FROM THE RIGHT FEMORAL ARTERY (WHERE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) SHEATH WAS PREVIOUSLY). THE 14FR PEEL AWAY WAS REMOVED AND REPO SHEATH WAS ADVANCED. STAFF PLACED SUTURES AND MATCHED ANGLE ENTRY, SIGNIFICANTLY SLOWING THE BLEEDING. THE PATIENT WAS STABILIZED AND SENT TO ICU. THE PATIENT WAS STILL HAVING SUDDEN DROPS IN BLOOD PRESSURE, HAVING TO INTERMITTENTLY CODE, GIVE BLOOD PRODUCTS AND FLUID. RIGHT FLANK HEMATOMA WAS NOTED IN ICU AND IT KEPT GETTING BIGGER AS TIME WENT ON. THE PATIENT WAS SENT TO GET CT SCAN - SHOWED RETROPERITONEAL BLEED. GEN SURG CONSULTED TO FIX THE BLEED (HE WAS WILLING TO FIX). THE PATIENTS FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND VASCULAR DAMAGE: PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE THE BLEEDING WAS RESOLVED BY MATCHING THE ANGLE AND PLACING SUTURES. VASCULAR DAMAGE: THE BLEED FROM RIGHT FEMORAL ARTERY (RFA) HAS BEEN MENTIONED BUT NO DETAILS RELATED TO THE CAUSE OF DAMAGE. HENCE, THE ROOT CAUSE OF THE VASCULAR DAMAGE WAS UNABLE TO BE DETERMINED AS INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847550&pc=MNS | 2518422-2025-104152 | 2025-04-11 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS OR FAULTS FOUND. THE DEVICE IS CURRENTLY AWAITING CONFIRMATION FOR RETURN TO LOAN OR TO BE SCRAPPED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. IN THE PREVIOUS FILE THE REPORT WAS SUBMITTED FOR AN INITIAL WHICH IS UPDATED TO THE FINAL IN THIS REPORT. ADDITIONALLY, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WERE ALSO UPDATED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847645&pc=MHX | 9610816-2025-000308 | 2025-02-16 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-04-16 04:00 | MHX | PATIENT INFORMATION CENTER IX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K211900 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED WAY WHILE CONNECTED TO THE MONITOR IN THE TIMEFRAME OF (B)(6) 2025. WHEN THE CUSTOMER WAS ATTEMPTING TO COLLECT THE LOGS ON (B)(6) 2025, THE APPLICATION SHUT DOWN; THEREFORE, THE CUSTOMER WAS REQUESTING ASSISTANCE IN DOWNLOADING THE LOGS. IN ADDITION, IT WAS INDICATED THE CUSTOMER ALLEGED NO FAILURE OR FAULT OF THE DEVICE AS IT RELATES TO THE PATIENT DEATH. A REVIEW OF THE LOGS BY PRODUCT SUPPORT ENGINEERING CONFIRMED THE DEVICE WAS FUNCTIONING AS EXPECTED DURING THE EVENT TIMEFRAME.
Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847736&pc=BZD | 2518422-2025-104170 | 2025-04-15 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21847964&pc=MNS | 2518422-2025-104174 | 2025-04-15 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848047&pc=OZD | 1220648-2025-27530 | 2025-03-18 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND APPROXIMATELY FOUR DAYS AFTER START OF THE SUPPORT, THE PATIENT EXPERIENCED RUNS OF VENTRICULAR TACHYCARDIA. THE PUMP WAS REPOSITIONED IN RESPONSE TO THE CONDITION. SUPPORT WAS ABLE TO CONTINUE WITHOUT INTERRUPTION. THE IMPELLA 5.5 DEVICE WAS SUCCESSFULLY WEANED AND EXPLANTED WITH A SURVIVED PATIENT OUTCOME.Event Description: ADDITIONAL INFORMATION RECEIVED THAT EIGHTEEN DAYS AFTER INITIATING OF SUPPORT, CARE WAS WITHDRAWN AND THE PATIENT'S STATUS WAS UPDATED TO EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: B2 OUTCOME DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. 1220648-2025-27530. B5 PATIENT OUTCOME OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. 1220648-2025-27530. G3 DATE RECEIVED BY MANUFACTURE WAS 4/6/2025. H1 TYPE OF REPORT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. 1220648-2025-27530. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. 1220648-2025-27530.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848103&pc=QZI | 2124215-2025-23736 | 2022-07-04 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-16 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Unspecified Heart Problem |  |  | 1 | Event Description: A PULSE FIELD ABLATION PROCEDURE WAS PERFORMED WITH A FARAWAVE CATHETER. IT WAS REPORTED THAT THE PATIENT EXPERIENCED LEFT VENTRICULAR DYSFUNCTION, CARDIOGENIC SHOCK AND DEATH.Event Description: A PULSE FIELD ABLATION PROCEDURE WAS PERFORMED WITH A FARAWAVE CATHETER. IT WAS REPORTED THAT THE PATIENT EXPERIENCED LEFT VENTRICULAR DYSFUNCTION, CARDIOGENIC SHOCK AND DEATH. Manufacturer Narrative: CORRECTION ADDED A3A: SEX: FEMALE; A3B: GENDER: FEMALE; A4: WEIGHT: 54; A4: UNITE OF MEASURE - WEIGHT: KG. THE FOLLOWING EVENT IS A RETROSPECTIVE CHART REVIEW AND IS BEING REPORTED AS AN EVENT THAT COULD HAVE POTENTIALLY BEEN PREVIOUSLY REPORTED TO BOSTON SCIENTIFIC. AS THE DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BOSTON SCIENTIFIC, THERE ARE SIGNIFICANT LIMITATIONS TO BOSTON SCIENTIFIC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BOSTON SCIENTIFIC. AS DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BOSTON SCIENTIFIC AND THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: THE FOLLOWING EVENT IS A RETROSPECTIVE CHART REVIEW AND IS BEING REPORTED AS AN EVENT THAT COULD HAVE POTENTIALLY BEEN PREVIOUSLY REPORTED TO BOSTON SCIENTIFIC. AS THE DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BOSTON SCIENTIFIC, THERE ARE SIGNIFICANT LIMITATIONS TO BOSTON SCIENTIFIC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BOSTON SCIENTIFIC. AS DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BOSTON SCIENTIFIC AND THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848153&pc=OZD | 1220648-2025-27531 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiac Tamponade; Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT HAD AN ATTEMPTED IMPELLA CP IMPLANT AND AFTER THE SHEATH AND GUIDEWIRE WERE INSERTED, THE PATIENT HAD A CARDIAC TAMPONADE WITH A LEFT VENTRICLE (LV) RUPTURE. THE IMPLANT WAS ABORTED. THE IMPELLA PACKAGE WAS OPENED, BUT DUE TO LV RUPTURE IMPELLA WAS NOT INSERTED. INSTEAD, IABP+ECMO WAS INSERTED AND A SURGICAL PROCEDURE WAS PERFORMED. IT IS UNKNOWN IF THE GUIDEWIRE OR ATTEMPTED IMPLANT CONTRIBUTED TO THE LV RUPTURE. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES AND VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VENTRICLE PERFORATION: DURING INSERTION OF 16FR INSERTION SHEATH, A NON ABIOMED PRODUCT, AND OPERATION OF THE GUIDEWIRE, BLOOD PRESSURE DECREASED, AND CARDIAC TAMPONADE AND LEFT VENTRICLE (LV) RUPTURE (FREE WALL RUPTURE) WERE CONFIRMED BY ANGIOGRAPHY. THE PUMP WAS NOT RETURNED. IT IS UNCLEAR WHETHER THE 16 FR SHEATH, GUIDEWIRE, OR PUMP CONTRIBUTED TO THE PERFORATION AND HOW. THE CAUSE OF THE VENTRICLE PERFORATION WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. DELIVERY ISSUES: THE IMPELLA PACKAGE WAS OPENED, BUT DUE TO CARDIAC TAMPONADE AND LV RUPTURE, IMPELLA WAS NOT INSERTED. THE CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY USE RELATED AS PER THE CLINICAL DETAILS, THE PUMP WAS NOT IMPLANTED DUE TO CARDIAC TAMPONADE AND LV PERFORATION. INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM FOR USE DURING CARDIOGENIC SHOCK, SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531. B2 DATE OF DEATH WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531. D2 COMMON DEVICE NAME WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 HEALTH EFFECT - CLINICAL CODE 2135 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531. H6 MEDICAL DEVICE PROBLEM CODE 2682 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531 H10 INSTRUCTIONS FOR USE WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27531. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848171&pc=NPT | 2025587-2025-02705 | 2023-05-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-16 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Dysphagia/ Odynophagia; Hematoma; Necrosis; Pressure Sores | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD BEEN HOSPITALIZED DUE TO A TROCHANTER FRACTURE. NECROSIS AND MULTIPLE BEDSORES WERE OBSERVED IN THE LOWER LIMBS. 6 DAYS LATER, THE PATIENT DEVELOPED DYSARTHRIA AND APPROXIMATELY 2 WEEKS AFTER THIS, A COMPUTED TOMOGRAPHY SCAN REVEALED AN ACUTE SUBDURAL HEMATOMA. THE PATIENT PASSED AWAY APPROXIMATELY 4 DAYS LATER. IT WAS NOTED THAT, DUE TO THE PATIENT'S AGE OF (B)(6) YEARS AND HIGH FRAILTY, THE CAUSE OF DEATH WAS WORSENING OF GENERAL CONDITION PRIOR TO OPERATION FOR THE FRACTURE. PER THE PHYSICIAN, IT WAS A CAUSE OF LONG-TERM DEATH CLOSE TO SENILITY. PER THE PHYSICIAN, NO ADVERSE EVENTS WERE OBSERVED IN THE POST-PROCEDURE ECHOCARDIOGRAM. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848333&pc=MNS | 2518422-2025-104179 | 2025-04-15 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848346&pc=NVZ | 2017865-2025-39838 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: THE PRODUCT WAS RETURNED DUE TO PATIENT DEATH WITHOUT A COMPLAINT ASSOCIATED TO THE PRODUCT. CAUSE OF DEATH: CARDIAC ARREST, MALIGNANT NEOPLASM OF UNSPECIFIED BRONCHUS OR LUNG AND DEMENTIA. Manufacturer Narrative: THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND. CORRECTION - G2 - REPORT SOURCE - USER FACILITY SHOULD HAVE BEEN SELECTED IN INITIAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848348&pc=NVN | 2017865-2025-39840 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: THE PRODUCT WAS RETURNED DUE TO PATIENT DEATH WITHOUT A COMPLAINT ASSOCIATED TO THE PRODUCT. CAUSE OF DEATH: CARDIAC ARREST, MALIGNANT NEOPLASM OF UNSPECIFIED BRONCHUS OR LUNG AND DEMENTIA. Manufacturer Narrative: CORRECTION - G2 - REPORT SOURCE - USER FACILITY SHOULD HAVE BEEN SELECTED IN INITIAL REPORT. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL AND X-RAY EXAMINATIONS OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848349&pc=NVN | 2017865-2025-39841 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: THE PRODUCT WAS RETURNED DUE TO PATIENT DEATH WITHOUT A COMPLAINT ASSOCIATED TO THE PRODUCT. CAUSE OF DEATH: CARDIAC ARREST, MALIGNANT NEOPLASM OF UNSPECIFIED BRONCHUS OR LUNG AND DEMENTIA. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. Manufacturer Narrative: CORRECTION - G2 - REPORT SOURCE - USER FACILITY SHOULD HAVE BEEN SELECTED IN INITIAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848370&pc=DSQ | 3007042319-2025-00768 | 2024-12-26 05:00 | Death | HEARTWARE, INC. | 2025-04-16 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Sepsis | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING VENTRICULAR ASSIST DEVICE (VAD) IMPLANTS IN SMALL PATIENTS. THE AUTHORS DESCRIBED TWO PATIENT DEATHS. ONE DEATH WAS DUE TO SEPSIS ON DAY 373 POST VAD IMPLANT AND THE OTHER WAS DUE TO A CEREBRAL HEMORRHAGE ON DAY 52 POST IMPLANT. THERE WERE PATIENTS WHO EXPERIENCED DRIVELINE INFECTION, CEREBROVASCULAR ACCIDENT (CVA), CARDIAC TAMPONADE, AND THE NEED FOR RIGHT VAD SUPPORT. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: MID-TERM OUTCOME OF HEARTWARE HVAD IN SMALL PATIENTS IN JAPAN. DOKKYO MEDICAL JOURNAL. 2024. 3(4): 289-297. DOI: 10.51040/DKMJ.2024-019 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848408&pc=NGV | 2124215-2025-23550 | 2025-04-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-16 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. ONE DAY POST PROCEDURE WHILE THE PATIENT WAS IN RECOVERY, THEY EXPERIENCED A CEREBRAL VASCULAR ACCIDENT. THE PATIENT DID NOT RECOVER, AND THEY DIED FROM THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848441&pc=DQA | 2936999-2025-00208 |  | Death | MEDIANA CO. LTD. | 2025-04-16 04:00 | DQA | NELLCOR | Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K120773 |  | 1 | Event Description: IT WAS REPORTED THAT THE MONITOR'S ALARM HISTORY WAS REQUESTED FOR INVESTIGATION. THE ALARM HISTORY COULD NOT BE VIEWED AT THAT TIME, SO ONLY THE ALARM SETTINGS WAS RECORDED. THE SENSOR WAS DETACHED AND THE POWER WAS ON. THE PATIENT PASSED AWAY AT THE MEDICAL FACILITY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848563&pc=MNT | 2518422-2025-043789 | 2025-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | MNT | RESPIRONICS | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | K102985 |  | 1 | Event Description: AN AUTHORIZED SERVICE PROVIDER (ASP) REPORTED THAT A PATIENT PASSED AWAY WHILE THE V60 VENTILATOR WAS IN USE ON THE PATIENT AND THAT THE DEVICE WAS FOUND ALARMING FOR A DEPLETED BACKUP BATTERY. THE REPORTER IS UNSURE IF THE DEVICE HAD ANYTHING TO DO WITH THE PATIENT¿S OUTCOME, AS IT WAS COMMUNICATED THAT THE PATIENT WAS ALREADY IN THE PROCESS OF PASSING AWAY WHEN BEING PLACED ONTO THE VENTILATOR. ON (B)(6) 2025, THE MANAGER OF RESPIRATORY CARE OR NEURODIAGNOSTIC SERVICES AT THE HOSPITAL CONFIRMED IN A GOOD FAITH EFFORT RESPONSE THAT THE DEVICE ALARMED AS INTENDED WHEN THE DEVICE WAS CHECKED, AND THAT THE BACKUP BATTERY WAS FOUND TO BE DEPLETED. THE MANAGER ALSO STATED THAT THE PATIENT DEATH WAS NOT DIRECTLY ASSOCIATED WITH THE EVENT. THE V60 VENTILATOR WAS PLUGGED INTO AN ELECTRICAL OUTLET, AND THE V60 VENTILATOR HAD BEEN FUNCTIONING PROPERLY. THE MANAGER STATED THAT THEY JUST WANTED TO CONFIRM THAT THE BACKUP BATTERY WORKED PROPERLY AND REQUESTED AN EVALUATION OF THE DEVICE TO DETERMINE HOW LONG THE DEVICE MAY HAVE BEEN USING THE BACKUP BATTERY. THE ASP AGREED TO TRAVEL TO THE CUSTOMER¿S SITE TO PERFORM AN EVALUATION OF THE DEVICE TO CONFIRM THAT THE DEVICE AND ITS BACKUP BATTERY ARE FULLY FUNCTIONAL. A CLINICAL ASSESSMENT WAS PERFORMED BY A PHILIPS CLINICAL EXPERT ON THIS CASE. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND CONTRIBUTION TO THE EVENT CANNOT CURRENTLY BE CONFIRMED, NOR REFUTED, BASED ON THE EVIDENCE PRESENT. ACCORDING TO THE V60 USER MANUAL ABOUT THE BACKUP BATTERY, ¿TO REDUCE THE RISK OF POWER FAILURE TO THE VENTILATOR, PAY CLOSE ATTENTION TO THE BATTERY'S CHARGE LEVEL. THE BATTERY'S OPERATION TIME IS APPROXIMATE AND IS AFFECTED BY VENTILATOR SETTINGS, DISCHARGE AND RECHARGE CYCLES, BATTERY AGE, AND AMBIENT TEMPERATURE. THE BACKUP BATTERIES ARE INTENDED FOR SHORT-TERM USE ONLY. THEY ARE NOT INTENDED TO BE A PRIMARY POWER SOURCE. AS A SAFEGUARD, THE VENTILATOR PROVIDES A LOW BATTERY ALARM. IT ALSO HAS A CAPACITOR-DRIVEN BACKUP ALARM THAT SOUNDS FOR AT LEAST 2 MINUTES WHEN BATTERY POWER IS COMPLETELY LOST. THE VENTILATOR CHARGES THE BATTERY WHENEVER THE VENTILATOR IS CONNECTED TO AC, WITH OR WITHOUT THE VENTILATOR SWITCHED ON. THE BATTERY (CHARGED) LED FLASHES TO SHOW THAT THE BATTERY IS BEING CHARGED. CHECK THE BATTERY CHARGE LEVEL BEFORE PUTTING A PATIENT ON THE VENTILATOR AND BEFORE UNPLUGGING THE VENTILATOR FOR TRANSPORT OR OTHER PURPOSES. THE POWER SOURCE SYMBOL AT THE BOTTOM OF THE RIGHT-HAND CORNER OF THE SCREEN SHOWS THE POWER SOURCE IN USE AND, IF THE VENTILATOR IS RUNNING ON BATTERY, THE LEVEL OF BATTERY CHARGE. IF THE BATTERY IS NOT FULLY CHARGED, RECHARGE IT BY CONNECTING THE VENTILATOR TO AC POWER FOR A MINIMUM OF 5 HOURS. PRESSING THE HELP BUTTON SHOWS YOU THE APPROXIMATE TIME REMAINING UNTIL THE BATTERY IS FULLY CHARGED. IF THE BATTERY IS NOT FULLY CHARGED WITHIN 16 HOURS, OR THE VENTILATOR DISPLAYS A CHECK VENT: BATTERY FAILED ALARM, REPLACE THE BATTERY.¿ IN ADDITION, THE V60 USER MANUAL¿S WARNINGS, CAUTIONS, AND NOTES ON ALARMS AND MESSAGES STATES THAT ¿IF AC POWER FAILS AND THE BACKUP BATTERY IS NOT INSTALLED OR IS DEPLETED, AN AUDIBLE AND VISUAL ALARM ANNUNCIATES FOR AT LEAST 2 MINUTES. IMMEDIATELY DISCONTINUE VENTILATOR USE AND SECURE AN ALTERNATIVE MEANS OF VENTILATION. AS IN MOST VENTILATORS WITH PASSIVE EXHALATION PORTS, WHEN POWER IS LOST, SUFFICIENT AIR IS NOT PROVIDED THROUGH THE CIRCUIT, AND EXHALED AIR MAY BE REBREATHED.¿. FURTHER GOOD FAITH EFFORTS ARE BEING COMPLETED TO THE CUSTOMER TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ON 16APR2025, A 3RD PARTY AUTHORIZED SERVICE PROVIDER (ASP) WENT TO THE CUSTOMER SITE AND EVALUATED THE DEVICE. THE ASP COMPLETED A PERFORMANCE VERIFICATION TEST (PVT) WHICH THE DEVICE PASSED, CONFIRMING THAT THE DEVICE MET THE MANUFACTURER¿S SPECIFICATIONS FOR USE AND WAS NOT EXPERIENCING ANY ISSUES. THE DEVICE WAS CONFIRMED TO BE FULLY FUNCTIONAL. A COPY OF A PORTION OF THE DEVICE DRPT (DIAGNOSTIC REPORT) WAS PROVIDED AND IT CONFIRMED THAT, ACROSS APPROXIMATELY 7 AND A HALF MINUTES, THE DEVICE PRODUCED A LOW INTERNAL BATTERY ALARM 4 TIMES, CONFIRMING THAT THE DEVICE ALARMED AND NOTIFIED THE USER APPROPRIATELY, AS INTENDED. WITH THE DEVICE HAVING ALARMED AS EXPECTED AND PASSED THE PVT, IT IS CONFIRMED THAT THERE WAS NO DEVICE ISSUE AT THE TIME OF THE PATIENTS PASSING. IF THE DEVICE DID SHUT DOWN DUE TO BATTERY DEPLETION, IT IS CONCLUDED THAT IT WAS DUE TO A LACK OF CORRECTIVE ACTION BY THE USER FOLLOWING THE INITIAL LOW INTERNAL BATTERY ALARM. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848738&pc=OZD | 1220648-2025-27538 | 2025-03-23 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE 72-YEAR-OLD MALE PATIENT ON IMPELLA CP SUPPORT DEVELOPED LIMB ISCHEMIA. ABSENT PULSES DISTAL TO CP, BUT PRESENT VIA DOPPLER ON CONTRALATERAL SIDE. IT WAS FURTHER REPORTED THAT CARE WAS WITHDREW AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21848773&pc=MKJ | 2112020-2025-00244 | 2025-03-25 04:00 | Death | CARDIAC SCIENCE CORPORATION | 2025-04-16 04:00 | MKJ | POWERHEART G5 KIT, AUTO, US ENG/ESP LA | Incorrect Interpretation of Signal | Insufficient Information | P160033 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT AN 18-YEAR-OLD FEMALE PATIENT, THE DEVICE ISSUED A "SHOCK ADVISED" PROMPT FOR A HEART RHYTHM THEY BELIEVE WAS NON-SHOCKABLE. THE SHOCK WAS NOT DELIVERED TO THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. HOWEVER, THEY BELIEVE IT WAS NOT A RESULT OF THE REPORTED MALFUNCTION. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE DEVICE WAS FOUND TO BE FUNCTIONING PROPERLY, WITH NO ERRORS BEFORE THE EVENT. DURING USE, IT INITIALLY DETECTED A SHOCKABLE RHYTHM AND BEGAN CHARGING, BUT AS PART OF THE SAME ANALYSIS REVERSED THE DECISION TO A NON-SHOCKABLE RHYTHM. AS A RESULT, THE DEVICE CANCELED THE SHOCK, AND NO SHOCK WAS DELIVERED. THIS BEHAVIOR IS CONSISTENT WITH THE DEVICE'S DESIGN AND SAFETY FEATURES, AS DESCRIBED IN THE USER GUIDE. THE DEVICE FUNCTIONS AS INTENDED. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849050&pc=JXG | 3013886523-2025-00075 |  | Death | INTEGRA LIFESCIENCES MANSFIELD | 2025-04-16 04:00 | JXG | CHPV RT ANG RESERVOIR UNITIZED | Infusion or Flow Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K221840 |  | 1 | Event Description: IT WAS REPORTED: ON AN UNSPECIFIED DATE, A 58-YEAR-OLD MALE PATIENT HAD A TRAFFIC ACCIDENT, WHICH LED TO POSTTRAUMATIC HYDROCEPHALUS. HE HAD UNDERGONE DECOMPRESSIVE CRANIOTOMY, AND A PIECE OF SKULL BONE REMOVED ON THE LEFT SIDE OF HIS SKULL. ON (B)(6) 2024, A HAKIM PROGRAMMABLE VALVE WAS IMPLANTED (ID: 823148) WITH AN INITIAL PRESSURE OF 160MMHG, WHICH WAS LATER INCREASED TO 180 MMHG, THEN TO 190 MMHG. THE PATIENT WAS DISCHARGED ON (B)(6) 2024. AT HOME, ON (B)(6) 2024, THE SHUNT PRESSURE WAS ADJUSTED TO 200 MM HG TO FACILITATE SURGERY WHILE PLACING THE BONE BACK. WHILE THIS MAINTAINED THE REQUIRED PRESSURE FOR A WEEK, THE VENTRICLES SUBSEQUENTLY COLLAPSED, LEADING TO CEREBRAL VENTRICLE COLLAPSE. THE AREA OF THE BRAIN SWELLED BACK FOR SOME TIME DUE TO HIGH PRESSURE; THEREFORE, SURGERY WAS NOT DONE. ON (B)(6) 2025, THE SHUNT PRESSURE WAS ADJUSTED TO 200 MM HG. SIMILAR TO EARLIER ATTEMPTS, THE PRESSURE REMAINED STABLE FOR A WEEK BEFORE THE VENTRICLES ¿SUNK¿ AGAIN, INDICATING MALFUNCTIONING OF THE VP SHUNT. ON (B)(6) 2025, THE PATIENT HAD X-RAY OF SKULL AP, LAT, AND ON (B)(6) 2025, A SENIOR REPRESENTATIVE CONFIRMED THAT THE VP SHUNT WAS MALFUNCTIONING AND ADVISED ANOTHER SURGERY FOR REPLACEMENT. THE PATIENT WAS CRITICALLY ILL AND IN URGENT NEED OF SHUNT REPLACEMENT SURGERY, HOWEVER, BUT THE PATIENT'S CONDITION WAS NOT STABLE AT THAT TIME. THE PATIENT'S CONDITION DETERIORATED AND THE PATIENT HAD ¿EXPIRED¿.Event Description: N/A Manufacturer Narrative: UPDATED FIELDS: D4, D9, G3, G6, H2, H3, H4, H6, H11. THE HAKIM VALVE (ID 823148) WAS NOT RETURNED FOR EVALUATION; THEREFORE, AN EVALUATION OF THE DEVICE COULD NOT BE PERFORMED. LOT NUMBER INFORMATION HAS BEEN PROVIDED; THEREFORE, DEVICE HISTORY RECORD (DHR) WAS REVIEWED, AND NO ANOMALIES WERE FOUND. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED. Manufacturer Narrative: UPDATED FIELDS: G3, G6, H1, H2, H11. ADDITIONAL INFORMATION PROVIDED BY INTEGRA'S MEDICAL AFFAIRS DIRECTOR: "I HAD A BRIEF TELEPHONIC CONVERSATION WITH DR. (B)(6). FROM (B)(6) HOSPITAL, REGARDING THE PATIENT¿S CASE. DR. (B)(6). MENTIONED THAT THE PATIENT HAD SUSTAINED A SEVERE HEAD INJURY AND WAS INITIALLY MANAGED WITH A FIXED PRESSURE SHUNT. DUE TO COMPLICATIONS RELATED TO OVER DRAINAGE, THE PATIENT WAS LATER SWITCHED TO A PROGRAMMABLE VALVE." "HE NOTED THAT THE OVER DRAINAGE OCCURRED AT LOWER PRESSURE SETTINGS, AND THE VALVE PRESSURE WAS SUBSEQUENTLY INCREASED TO 200 MM H2O. DESPITE THIS ADJUSTMENT, A SUNKEN FLAP WAS STILL OBSERVED." "AS DR. (B)(6). WAS OCCUPIED AT THE TIME OF THE CALL, WE COULD NOT DISCUSS THE CASE IN FURTHER DETAIL. HE CONVEYED THE KEY POINTS BUT DID NOT EXPRESS INTEREST IN AN EXTENDED DISCUSSION." Manufacturer Narrative: UPDATED FIELDS: G3, G6, H1, H2, H11. CORRECTED FIELD: H1. ADDITIONAL INFORMATION PROVIDED BY INTEGRA'S MEDICAL AFFAIRS DIRECTOR: "I HAD A DISCUSSION WITH DR. (B)(6) (SURGEON) CONCERNING THE CONSENSUS ON SUNKEN FLAP SYNDROME - BASED RECOMMENDATIONS FOR THE DIAGNOSIS AND SURGICAL MANAGEMENT OF CRANIOPLASTY AND POST-TRAUMATIC HYDROCEPHALUS AS OUTLINED BY THE EUROPEAN PANEL." "DR. (B)(6) HAS UNDERSTOOD THE CONSENSUS AND ITS RELEVANCE TO CRANIOPLASTY IN THE CONTEXT OF POST-TRAUMATIC HYDROCEPHALUS. SHOULD THERE BE ANY NEED FOR FURTHER CLARIFICATION FROM DR. (B)(6), I WILL PROVIDE ADDITIONAL UPDATES ACCORDINGLY." Manufacturer Narrative: AN INVESTIGATION HAS BEEN INITIATED BASED ON THE REPORTED INFORMATION. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW-UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849091&pc=OWB | 3003768277-2025-003511 | 2025-03-15 04:00 | Death | PHILIPS MEDICAL SYSTEMS NEDERLAND B.V. | 2025-04-16 04:00 | OWB | AZURION | Radiation Output Failure |  | K200917 |  | 1 | Event Description: IT WAS REPORTED TO PHILIPS THAT DURING AN EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE, INVOLVING A PATIENT WITH ST-ELEVATION MYOCARDIAL INFARCTION (STEMI), THE SYSTEM DISPLAYED AN X-RAY TUBE CURRENT OUT-OF-RANGE WARNING MESSAGE. THE PROCEDURE WAS ABORTED. THE PATIENT BEGAN TO DECOMPENSATE, AND THE PHYSICIAN INSERTED A BALLOON PUMP. THE PATIENT WAS THEN TRANSFERRED TO ANOTHER FACILITY, WHERE THEY WERE NOTED TO BE ON A VENTILATOR AND UNRESPONSIVE. ON (B)(6) 2025, THE CUSTOMER INFORMED PHILIPS THAT THE PATIENT HAD PASSED AWAY. NO OTHER INFORMATION WAS PROVIDED. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849172&pc=DSQ | 2916596-2025-02144 | 2025-03-21 04:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Respiratory Acidosis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AT AN OUTSIDE FACILITY RELATED TO VENTRICULAR TACHYCARDIA/ARRHYTHMIAS AND PROFOUND ACIDOSIS. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS CARDIAC ARRHYTHMIA, AND DEATH AS ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849411&pc=OZD | 1220648-2025-27544 | 2025-03-21 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A 64-YEAR-OLD MALE PATIENT ON IMPELLA 5.5 AND VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION FOR SUPPORT DURING HIGH-RISK SURGERY. AFTER THE CHEST WAS CLOSED AND THE SITE WAS SECURED, THE PATIENT HAD TO BE TAKEN BACK TO THE OPERATING ROOM DUE TO BLEEDING IN THE CHEST CAVITY. ULTIMATELY THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849479&pc=DSQ | 2916596-2025-02180 | 2025-03-25 04:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Asystole; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT WAS FOUND DOWN BY HIS FAMILY, CARDIOPULMONARY RESUSCITATION (CPR) STARTED. WHEN EMERGENCY MEDICAL SERVICES (EMS) ARRIVED, THE PATIENT WAS IN ASYSTOLE AND THE PUMP ALARMING LOW FLOWS. PER EMS, THE PATIENT'S POWER OF ATTORNEY DIDN¿T WANT ANY FURTHER RESUSCITATION EFFORTS. PER EMS, THE DRIVELINE WAS DISCONNECTED AND PUMP TURNED OFF. IT WAS REPORTED THAT THE PATIENT HAD JUST BEEN DISCHARGED FROM THE HOSPITAL DUE TO URINARY RETENTION ON (B)(6) 2025. THE PATIENT DID A HOME INTERNATIONAL NORMALIZED RATIO ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE PROVIDED FOR REVIEW. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 OF THE IFU ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849598&pc=BZD | 2518422-2025-104204 | 2025-03-20 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CALLED AND WANTED TO KNOW IF SHE NEEDED TO RETURN THE REPLACEMENT DEVICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S WIFE WAS INFORMED SHE COULD RETURN THE DEVICE IF SHE WANTED TO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAD PASSED AWAY. THE PATIENT'S WIFE CALLED AND WANTED TO KNOW IF SHE NEEDED TO RETURN THE REPLACEMENT DEVICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S WIFE WAS INFORMED SHE COULD RETURN THE DEVICE IF SHE WANTED TO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H CODING IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849790&pc=OZD | 1220648-2025-27553 | 2025-03-21 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A 47-YEAR-OLD FEMALE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD ARRHYTHMIAS. IMPELLA¿S POSITIONING WAS CHECKED TWICE AND APPEARED TO BE IN THE CORRECT POSITION. PATIENT HAD CONSISTENT VENTRICULAR FIBRILLATION FOR HOURS DESPITE INTERVENTIONS. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849807&pc=FRN | 9615050-2025-00174 | 2025-03-19 04:00 | Death | ICU MEDICAL COSTA RICA LTD. | 2025-04-16 04:00 | FRN | PLUM 360 (V 15.X) | Unexpected Shutdown | No Clinical Signs, Symptoms or Conditions | K141789 |  | 1 | Event Description: IT WAS REPORTED THAT A PLUM 360 SHUT DOWN UNEXPECTEDLY DURING USE. THE REPORTER STATED THAT "THE PUMP HAD A REPLACE BATTERY ALERT FROM 2 DAGO AGO BUT NO ONE PULL THE PUMP FROM SERVICE NOR SEND IN A WORK ORDER. A LEVO INFUSION WAS STARTED ON THE PATIENT AND THE PUMP WAS UNPLUGGED TO TRANSPORT THE PATIENT. THE PUMP SHUT DOWN AND THE PATIENT DIED." NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR EVALUATION; HOWEVER, TESTING HAS NOT YET BEEN COMPLETED. Manufacturer Narrative: EVALUATION: ENGINEERING EVALUATES THAT THE BATTERY DRAINAGE FROM LEVEL 4 TO LEVEL 1 IN LESS THAN 4 HOURS INDICATES A BAD BATTERY CONDITION. THE PROBABLE CAUSE OF THE PUMP ABRUPTLY SHUTTING DOWN WAS DUE TO THE SUDDEN BATTERY LEVEL DROP CONSEQUENTLY CAUSING THE DIAGNOSTIC RECORDS TO LOG A DIRTY_BIT_DETECTED/SHUTDOWN-FAILURE-DETECTED EVENTS COUPLED WITH A BOOTUP SIGNAL. ALL INSTANCES OF A REPLACE BATTERY ALARM, THE BATTERY SHOULD HAVE BEEN REPLACED AT THAT TIME. FURTHERMORE, ENGINEERING EVALUATES THE PUMP IS PERFORMING CORRECTLY IN DETECTING AND ALARMING OF A LOW AND REPLACE BATTERY CONDITION AND INFUSING AT THE EXPECTED RATE WITHIN THE DELIVERY ACCURACY TOLERANCE. ACCORDING TO THE SYSTEM OPERATING MANUAL DOCUMENT, N56-REPLACE BATTERY ALARM OCCURS WHEN THE BATTERY OR BATTERY CHARGE CIRCUITRY NEEDS SERVICING. THIS IS A LOW PRIORITY ALARM, AND IT DOES NOT STOP THE INFUSION. ACCORDING TO THE PLUM A+ AND PLUM 360 ACTIVE MALFUNCTIONS AND ALARM DOCUMENTATION, REPLACE BATTERY ALARM OCCURS WHEN A QUESTIONABLE BATTERY EVENT IS DETECTED 3 CONSECUTIVE TIMES. (ASSERTED TO INDICATE A FULL BATTERY DISCHARGE FASTER THAN EXPECTED, I.E. APPROACHING THE END OF LIFE AND THEREFORE NEEDS TO BE REPLACED). THIS ALARM PERSISTS AND REASSERTS OVER, POWER CYCLES UNTIL IT IS CLEARED BY PRESSING THE ¿CHANGE BATTERY¿ SOFT KEY IN BIOMED MODE AFTER INSTALLING A NEW BATTERY. ACCORDING TO THE SYSTEM OPERATING MANUAL DOCUMENT, N252-DEPLETED BATTERY ALARM OCCURS WHEN THE INFUSER IS RUNNING ON BATTERY POWER, AND THE BATTERY VOLTAGE DROPS BELOW DEPLETED VOLTAGE LIMIT (5.45V). THIS IS A HIGH PRIORITY ALARM WHICH STOPS THE INFUSION AND PREPARES THE PUMP FOR SHUTDOWN IF NOT PLUGGED INTO AC WITHIN 3 MINUTES AFTER IT IS TRIGGERED. THE CLEAR CONDITION IS TO CONNECT THE DEVICE TO AC POWER SOURCE. CONCLUSION: ENGINEERING CONFIRMS THE CUSTOMER¿S COMPLAINT OF THE PUMP DISPLAYING A REPLACE BATTERY ALARM AND ABRUPTLY SHUTTING DOWN. THE CUSTOMER DOES NOT APPEAR TO HAVE TAKEN REMEDIAL ACTION TO REPLACE THE BATTERY, AND ENGINEERING EVALUATES THE PROBABLE CAUSE TO BE A COMBINATION OF BATTERY FAILURE COUPLED WITH USER ERROR I.E., FAILURE TO ADDRESS ALARMS IN A TIMELY MANNER. ENGINEERING RECOMMENDS SERVICE CENTER REPLACE THE BATTERY, DOCUMENT THE BATTERY AGE AND HISTORY (SO FAR AS KNOWN), PERFORM PREVENTATIVE MAINTENANCE TEST, ENSURE THE PIEZO ALARM IS AUDIBLE AND PERFORM VERIFICATION TESTING FOR THE DEVICE POWER SYSTEM (BATTERY, POWER SUPPLY, ETC.) TO ENSURE PUMP IS OPERATIONAL. 1A. THE PUMP SHUT DOWN DURING PATIENT TRANSPORT BECAUSE THE CSB BATTERY WAS CRITICALLY DEPLETED WHEN THE DEVICE WAS UNPLUGGED FROM AC POWER. B. THE BATTERY DEPLETION WAS A DIRECT RESULT OF THE CUSTOMER¿S FAILURE TO RESPOND TO MULTIPLE REPLACE BATTERY ALARMS OVER A PERIOD OF AT LEAST TWO DAYS. THE PUMP REMAINED IN SERVICE WITHOUT ACTION, ALLOWING THE BATTERY TO DRAIN COMPLETELY. C. WHILE THE BATTERY WAS CONFIRMED TO BE DEFECTIVE, THE EVENT COULD HAVE BEEN PREVENTED IF THE ALARMS HAD BEEN ADDRESSED IN A TIMELY MANNER. D. THIS EVENT DOES NOT REQUIRE ESCALATION TO A NEW CAPA, AS THE KEY CONTRIBUTING FACTOR WAS USER ERROR. HOWEVER, CAPA IS ALREADY ADDRESSING BATTERY PERFORMANCE ISSUES, WHICH INCLUDES THE HARDWARE ASPECT RELATED TO THIS CASE. 2. ¿ALTHOUGH THE REPORTED DEVICE MALFUNCTION WAS CONFIRMED, IT COULD NOT BE CONFIRMED WHETHER THIS MALFUNCTION CONTRIBUTED TO THE REPORTED PATIENT HARM BASED ON THE INFORMATION PROVIDED TO THE MEDICAL ASSESSMENT TEAM. (SEE MEDICAL ASSESSMENT RELATED TO (B)(4) ATTACHED IN TWD.) BASED ON THIS CONCLUSION, EXISTING CAPA WILL BE LEVERAGED AND NO ADDITIONAL CAPA ESCALATION IS REQUIRED.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21849883&pc=DSQ | 2916596-2025-02640 | 2025-04-05 04:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI ORGAN FAILURE IN THE INTENSIVE CARE UNIT (ICU). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTI ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850074&pc=OZD | 1220648-2025-27556 | 2025-03-22 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD ACCESS SITE BLEEDING. SEVERAL BLOOD PRODUCTS WERE GIVEN. SUPPORT WAS CONTINUED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE PATIENT CONDITION BECAUSE THE BLEEDING IS OCCURRING FROM ALL THE ACCESS SITES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850085&pc=OZD | 1220648-2025-27557 | 2025-03-24 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Decrease in Pressure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING PREPARATION OF THE PURGE SYSTEM FOR IMPELLA CP PLACEMENT, AN ERROR OCCURRED ON THE AUTOMATED IMPELLA CONTROLLER (AIC) SAYING NO PURGE FLUID. AFTER TROUBLESHOOTING IT SEEMS THAT THE DRIVE SHAFT IN THE AIC WAS BLOCKED AND THEREFORE NO PURGE FLUID COULD BE PUMPED. THE AIC WAS EXCHANGED AND IMPELLA WAS SUCCESSFULLY PLACED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED SOON AFTER ARRIVING AT ICU. THE MINIMAL INTERRUPTION IN THERAPY CANNOT BE EXCLUDED FROM THE PATIENT¿S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - PURGE ISSUE ¿ OTHER HAS BEEN COMPLETED. THE CONSOLE LOGS WERE CONSISTENT WITH WHAT WAS REPORTED IN THE COMPLAINT. FURTHER INSPECTION OF THE CONSOLE REVEALED THAT THERE WAS DRIED GLUCOSE ON THE PURGE MOTOR. THE OBSERVED DRIED GLUCOSE LIKELY CAUSED THE PURGE MOTOR TO GET STUCK TO THE PURGE HOUSING. BECAUSE THE PURGE MOTOR WAS STUCK, IT WAS UNABLE TO ROTATE AND PURGE THE LINE. THE ROOT CAUSE OF THE PURGE FLUID DETECTION ISSUE WAS A SEIZED PURGE MOTOR DUE TO DRIED GLUCOSE INGRESS. G1 REPORTING CONTACT ADDRESS WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27557. G1 MANUFACTURING SITE NAME AND ADDRESS WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27557. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850091&pc=NPT | 2025587-2025-02721 | 2024-02-19 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-16 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, IT WAS UNKNOWN IF THE DEATH WAS RELATED TO THE VALVE AND THE VALVE IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850330&pc=NPT | 2015691-2025-02936 | 2025-03-24 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-16 04:00 | NPT | EDWARDS COMMANDER DELIVERY | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Pericardial Effusion; Insufficient Information |  |  | 1 | Event Description: AS REPORTED BY AN EDWARDS BRAZIL AFFILIATE, DURING A TAVR PROCEDURE WITH AN UNKNOWN EDWARDS VALVE, THE PATIENT DIED AFTER THE VALVE WAS IMPLANTED. THERE WAS NO ADDITIONAL INFORMATION PROVIDED.Event Description: IT WAS INITIALLY REPORTED BY AN EDWARDS BRAZIL AFFILIATE, THAT DURING A TAVR PROCEDURE WITH AN UNKNOWN EDWARDS VALVE, THE PATIENT DIED AFTER THE VALVE WAS IMPLANTED. THERE WAS NO ADDITIONAL INFORMATION PROVIDED AT THE TIME. HOWEVER, PER ADDITIONAL INFORMATION RECEIVED, AFTER THE VALVE WAS IMPLANTED, THERE WAS PARAVALVULAR LEAK NOTED, SO THE TEAM OPTED FOR POST-DILATION. AT THE TIME THE BALLOON WAS BEING DEFLATED, THE PACEMAKER WAS LOST AND IT IS BELIEVED THAT THE BALLOON MOVED INJURING THE INTIMA LAYER OF THE AORTA. THE PATIENT HAD A PERICARDIAL EFFUSION, AND DRAINAGE WAS PERFORMED, BUT THE HEART HAD NO RESERVE. DESPITE RESUSCITATION MANEUVERS, THE PATIENT DID NOT RETURN AND PASSED AWAY. Manufacturer Narrative: CORRECTION TO H6. Manufacturer Narrative: CORRECTION TO B5, D1, D4, H4, H6. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (LOSS OF PACING) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURES OFTEN HAVE COMPLEX MEDICAL HISTORIES AND MULTIPLE CO-MORBIDITIES. THEY MAY ALSO HAVE LIMITED CARDIAC RESERVE THAT CAN IMPAIR RECOVERY FROM THE PROCEDURE. AFTER TVR PROCEDURES, PATIENTS ARE CLOSELY MONITORED, WHICH INCLUDES ASSESSMENT OF DEVICE IMPLANTS. DESPITE MULTIPLE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THE REPORTED EVENT. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE (I.E., CAUSE OF DEATH) THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. NO FURTHER FOLLOW-UP FOR ADDITIONAL INFORMATION IS REQUIRED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850432&pc=PYX | 1220648-2025-27568 | 2025-03-24 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Hemorrhage/Bleeding; Perforation of Vessels | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS OOZING FROM ALL PUNCTURE SITES. HEPARIN WAS REMOVED FROM THE PURGE AND BICARB WAS PLACED. ADDITIONALLY, THE IMPELLA FLOWS DROPPED AND PLACEMENT SIGNAL WAS OFF THE SCREEN. UNABLE TO RE-ZERO PAS. THE PUMP WAS EXPLANTED AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO DISSOCIATE THE IMPELLA PUMP AS AN ASSOCIATED FACTOR TO THE OUTCOME GIVEN BLEEDING EVENT AND PUMP FLOW THERAPY SUBOPTIMAL. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA RP WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS AND CAUTIONS INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND. ¿BE SURE THAT THE STOPCOCK ON THE REPOSITIONING SHEATH IS ALWAYS KEPT IN THE CLOSED POSITION. SIGNIFICANT BLEED BACK CAN RESULT IF THE STOPCOCK IS OPEN.¿ ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE FEMORAL VEIN. CLOSE AND DRESS THE WOUND.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER, USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER: ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿ IMPELLA RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER, SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿ RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER, SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿ Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, ACCESS SITE BLEEDING, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VASCULAR DAMAGE - PERIPHERAL VESSEL: CLINICAL DETAILS NOTED THAT BLEEDING WAS OCCURRING FROM MULTIPLE SITES. IT WAS NOT NOTED IF ANY REPAIRS WERE PERFORMED FOR BLEEDING. THE ROOT CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL COULD NOT BE DEFINITIVELY DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED. ACCESS SITE BLEEDING THE CLINICAL DETAILS NOTED THAT PATIENT WAS OOZING FROM ALL PUNCTURE SITES. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY PATIENT CONDITION RELATED AS PATIENT WAS BLEEDING FROM MULTIPLE SITES PER CLINICAL DETAILS. LOW PUMP FLOW: DATA LOGS WERE REVIEWED AND LT LOG AT TIME OF ISSUE SHOWED A LARGE MOTOR CURRENT SPIKE WITH A SPEED DEVIATION AND SUDDEN DROP IN FLOWS. ADDITIONALLY, A STEP UP IN PLACEMENT SIGNAL WAS OBSERVED AT TIME OF MOTOR CURRENT SPIKE. PUMP FLOWS REMAINED LOWER THAN P-LEVEL RANGE FOR REMAINDER OF CASE. CLINICAL DETAILS NOTED THAT PATIENT WAS NOT ON ANY FORMS OF ANTICOAGULATION AT TIME OF REPORTED ISSUE. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY DUE TO BIOMATERIAL INGESTION BASED ON RETURNED DATA LOG ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850433&pc=NPT | 2025587-2025-02727 | 2025-02-20 05:00 | Death | MEDTRONIC IRELAND | 2025-04-16 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Bradycardia; Low Blood Pressure/ Hypotension; Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVFXPLUS-29, THERE WAS A DROP IN BLOOD PRESSURE AND HEART RATE AS THE DELIVERY CATHETER SYSTEM (DCS) WAS ADVANCED THROUGH THE DESCENDING AORTA AND AROUND THE AORTIC ARCH. THE VALVE WAS POSITIONED ACROSS THE NATIVE VALVE, AND THE BLOOD PRESSURE AND HEART RATE CONTINUED TO DECREASE, LEADING TO THE VALVE BEING PULLED BACK INTO THE ASCENDING AORTA. CARDIOPULMONARY RESUSCITATION WAS INITIATED, AND THE PATIENT WAS PREPPED FOR OPEN HEART SURGERY AND PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION. IT WAS NOTED THAT THE GUIDEWIRE PERFORATED THE LEFT VENTRICLE (LV), RESULTING IN A LV PERFORATION. THE PROCEDURE WAS ULTIMATELY ABORTED. THE PATIENT PASSED AWAY THE SAME DAY. THE CAUSE OF DEATH WAS LV PERFORATION. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN TO NOTE THE DCS CANNOT BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {EVFXPLUS-29}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} PRODUCT {CONFIDA GUIDEWIRE}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850602&pc=OZD | 1220648-2025-27569 | 2025-03-23 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ARRESTED AT HOME, REMAINED PULSELESS UNTIL ARRIVAL TO EMERGENCY ROOM. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED, AND THE PATIENT WAS TAKEN TO CATHETERIZATION LAB. LEFT ANTERIOR DESCENDING ARTERY WAS FOUND OCCLUDED. IT WAS OPENED AND STENTED. THE IMPELLA CP WAS PLACED POST THIS INTERVENTION. THE PATIENT THEN HAD PULSELESS ELECTRICAL ACTIVITY ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR 20 MINUTES WITH MULTIPLE ROUNDS OF EPINEPHRINE. RETURN OF SYSTEMIC CIRCULATION WAS ATTAINED AGAIN. ECHOCARDIOGRAM SHOWED THE IMPELLA CP WAS PLACED ACROSS THE MITRAL VALVE. THE PHYSICIAN ATTEMPTED TO REPOSITION THE IMPELLA CP AWAY FROM THE MITRAL VALVE, BUT IT WAS PULLED INTO THE AORTA. THE PHYSICIAN WAS UNABLE TO GET BACK THE IMPELLA CP INTO THE VENTRICLE. IMPELLA CP WAS REMOVED FROM THE BODY. THE VENTRICLE WAS REASSESSED AND THE SAME IMPELLA CP WAS INSERTED USING A V18 WIRE. AFTER RESTARTING THE IMPELLA CP, THE PHYSICIAN WAS CONCERNED ABOUT THE PLACEMENT SIGNAL AND MOTOR CURRENT WAVEFORMS AND LOW FLOW. THE IMPELLA CP WAS REMOVED AND A NEW IMPELLA CP WAS INSERTED. THE PATIENT CODED AGAIN IMMEDIATELY AFTER PLACING THE NEW PUMP AND EXPIRED IN THE PROCEDURAL AREA. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS THE PUMP STAYED WITH THE EXPIRED PATIENT. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850666&pc=NVZ | 2017865-2025-40309 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVZ | FORTIFY ASSURA DR ICD, US | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, AND RIGHT ATRIAL (RA) LEAD WERE REMOVED DUE TO PATIENT DEATH. THE CAUSE OF DEATH PROVIDED WAS, "PROBABLE CARDIAC DYSRHYTHMIA IN A PERSON WITH HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE CHRONIC OBSTRUCTIVE DISEASE AND DIABETES MELLITUS". THE PRODUCTS WERE RETURNED WITHOUT CLEAR INFORMATION AS TO WHETHER THE DEATH WAS DEVICE RELATED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL. Manufacturer Narrative: CORRECTION: D6B.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850669&pc=DSQ | 2916596-2025-02464 | 2020-06-01 04:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿CLINICAL OUTCOMES OF CARDIAC TRANSPLANTATION IN HEART FAILURE PATIENTS WITH PREVIOUS MECHANICAL CARDIOCIRCULATORY SUPPORT?¿ IDENTIFYING THAT HM3 MAY BE RELATED TO READMISSION FOR HF, GI BLEEDING, CEREBROVASCULAR EVENTS, ARRHYTHMIAS, AND LVAD-RELATED INFECTION, AND TRANSPLANT. THIS IS A RETROSPECTIVE, OBSERVATIONAL, SINGLE-CENTER STUDY THAT EVALUATED CLINICAL OUTCOMES OF 105 PATIENTS WHO UNDERWENT CARDIAC TRANSPLANTATION BETWEEN JANUARY 2010 AND JUNE 2020 AT HOSPITAL ¿PAPA GIOVANNI XIII¿ IN BERGAMO (ITALY). ADULT PATIENTS (AGED = 18 YEARS) WITH ISCHEMIC OR DILATED CARDIOMYOPATHY WHO UNDERWENT ISOLATED HEART TRANSPLANTATION, AND PATIENTS WHO RECEIVED AN LVAD, EITHER AS A BRIDGE TO TRANSPLANT (BTT) OR BRIDGE TO CANDIDACY (BTC), COLLECTIVELY CATEGORIZED UNDER THE BTT GROUP, WERE INCLUDED IN THIS STUDY. SEVENTY-SEVEN (N=77) PATIENTS WERE IN THE DTT GROUP, WHILE TWENTY-EIGHT PATIENTS WERE IN THE BTT GROUP. THE BRIDGE-TO-TRANSPLANT DURATION WAS 554±346 DAYS, DURING WHICH THE COMPLICATIONS WERE RECORDED. IT WAS NOT POSSIBLE TO TRACE PATIENTS WHO HAD LVAD IMPLANTATION BUT WERE NOT TRANSPLANTED FOR VARIOUS REASONS (DEATH, REFUSAL TO TRANSPLANT, TRANSPLANTATION TO ANOTHER HOSPITAL CENTER, OFF TRANSPLANT LIST). READMISSION FOR HF, N (%) = 5 (17.9); GI BLEEDING, N (%) = 8 (28.6); CEREBROVASCULAR EVENTS, N (%) = 11 (39.3), ARRHYTHMIAS, N (%) = 6 (21.4), LVAD-RELATED INFECTION, N (%) = 17 (60.7). Manufacturer Narrative: SECTION D5: OPERATOR OF DEVICE CORRECTED. SECTION E: INITIAL REPORTER INFORMATION CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01JUN2020, SINCE THE CARDIAC TRANSPLANTATION OCCURRED BETWEEN JANUARY 2010 AND JUNE 2020. D¿ALONZO, MICHELE ET AL. J. CLIN. MED. 2025, 14, 275. HTTPS:// DOI.ORG/10.3390/JCM14010275. CARDIAC SURGERY UNIT, SPEDALI CIVILI, UNIVERSITY OF BRESCIA, 25124 BRESCIA, ITALY. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850670&pc=NVN | 2017865-2025-40310 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, AND RIGHT ATRIAL (RA) LEAD WERE REMOVED DUE TO PATIENT DEATH. THE CAUSE OF DEATH PROVIDED WAS, "PROBABLE CARDIAC DYSRHYTHMIA IN A PERSON WITH HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE CHRONIC OBSTRUCTIVE DISEASE AND DIABETES MELLITUS". THE PRODUCTS WERE RETURNED WITHOUT CLEAR INFORMATION AS TO WHETHER THE DEATH WAS DEVICE RELATED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850686&pc=NVY | 2017865-2025-40311 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, AND RIGHT ATRIAL (RA) LEAD WERE REMOVED DUE TO PATIENT DEATH. THE CAUSE OF DEATH PROVIDED WAS, "PROBABLE CARDIAC DYSRHYTHMIA IN A PERSON WITH HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE CHRONIC OBSTRUCTIVE DISEASE AND DIABETES MELLITUS". THE PRODUCTS WERE RETURNED WITHOUT CLEAR INFORMATION AS TO WHETHER THE DEATH WAS DEVICE RELATED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850725&pc=OZD | 1220648-2025-27572 | 2025-03-24 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS ONGOING SUCTION ALARMS. ECHO SHOWED 4.6CM INLET TO AORTIC VALVE ANNULUS. THE ECHO ALSO SHOWED INLET CLOSE TO MITRAL VALVE: NOT IMPEDING MOVEMENT, BUT LEAFLET MAY BE CLOSE TO INLET WHILE THE VALVE IS OPEN. THE DOCTOR DECLINED TO REPOSITION. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE PUMP WAS NOT RETURNED FOR INVESTIGATION. CLINICAL OBSERVATIONS INDICATED THAT THE PUMP APPEARED TO BE POSITIONED DEEPLY, WITH THE INLET LOCATED CLOSE TO THE MITRAL VALVE. DESPITE THIS POSITIONING, THE PUMP WAS KEPT IN PLACE. THERE WERE INTERMITTENT SUCTION ALARMS DURING THE CASE, WHICH PROMPTED A REDUCTION IN THE PUMP'S P-LEVEL TO MANAGE THE SUCTION CONCERNS. THE DATA LOG DID NOT RECORD ANY POSITION-RELATED ALARMS OR MOTOR CURRENT SPIKE. THE PUMP HAD SEVERAL SUCTION ALARMS AND LOW PUMP FLOW WHILE RUNNING ON A HIGHER P-LEVEL (P9). REDUCING THE P-LEVEL TO P7 RESOLVED THE LOW PUMP FLOW, BUT THE SUCTION ALARM PERSISTED FOR SOME TIME. LATER DURING THE CASE, THE P-LEVEL WAS FURTHER REDUCED TO RESOLVE THE SUCTION CONDITION. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, BASED ON DATA LOG REVIEW. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS: ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ CORRECT IMPELLA CATHETER POSITION: ¿FOR OPTIMAL POSITIONING OF THE IMPELLA CATHETER, THE INLET AREA OF THE CATHETER SHOULD BE 3.5 CM BELOW THE AORTIC VALVE ANNULUS AND WELL AWAY FROM PAPILLARY MUSCLE AND SUBANNULAR STRUCTURES. THE OUTLET AREA SHOULD BE WELL ABOVE THE AORTIC VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850752&pc=OZD | 1220648-2025-27570 | 2025-03-24 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD A MASSIVE STROKE. CARE WAS LATER WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE USE OF THE IMPELLA AND THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850843&pc=DSQ | 2916596-2025-02343 | 2025-03-27 04:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY, AND THE REASON WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850903&pc=OZD | 1220648-2025-27576 | 2025-03-24 04:00 | Death | ABIOMED, INC. | 2025-04-16 04:00 | OZD | IMPELLA CP | Positioning Failure; Increase in Pressure | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PUMP LEFT DEEP AT 4 CM AFTER REPOSITIONING FROM 5 CM. SUPPORT CONTINUED AND STAFF MONITORED. ADDITIONALLY, LOW PURGE FLOWS AND HIGH PURGE PRESSURES WERE NOTED. TISSUE PLASMA ACTIVATOR WAS ADDED TO THE PURGE AND SUPPORT CONTINUED. CARE WAS WITHDRAWN ON SUPPORT AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850971&pc=BZD | 2518422-2025-104183 | 2025-02-28 05:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | BZD | DREAMSTATION 2 ADVANCED AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE DAUGHTER OF THE PATIENT REPORTED THE PATIENT EXPIRED. THE REPORTED INQUIRED ABOUT THE REGISTRATION INFORMATION FOR THE DREAMSTATION 2 ADVANCED AUTO CPAP DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850975&pc=NPT | 2025587-2025-02736 | 2025-03-20 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-16 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE ATTEMPTED IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOP LASTY (BAV) WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON AS STANDARD FOR THE IMPLANTING PHYSICIAN. IT APPEARED THAT THE BALLOON MOVED DURING RAPID PACING WITH THE TEMPORARY PACEMAKER; HOWEVER, THE BALLOON INFLATED FULLY. SUBSEQUENTLY, THE BALLOON WAS DEFLATED AND THE RAPID PACING WAS STOPPED. THE PATIENT BECAME HYPOTENSIVE. THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT AND ADVANCED TO THE ANNULUS. UPON INITIAL DEPLOYMENT, IT WAS NOTED THAT THE VALVE WAS TOO DEEP. THE VALVE WAS RECAPTURED. AN ANGIOGRAM WAS PERFORMED WHILE THE RECAPTURED VALVE REMAINED IN THE ASCENDING AORTA. THE ANGIOGRAM REVEALED AN EFFUSION IN THE LEFT VENTRICULAR OUTFLOW TRACT. THE VALVE AND DCS WERE WITHDRAWN FROM THE PATIENT. LIFE SAVING MEASURES WERE PERFORMED TO TREAT THE EFFUSION; HOWEVER, THESE WERE UNSUCCESSFUL. THE PROCEDURE WAS DELAYED BY 2 HOURS AND THE PATIENT PASSED AWAY. UPON REVIEW OF THE PROCEDURAL FILMS IT WAS DETERMINED BY THE PHYSICIAN THAT THE CAUSE OF THE EFFUSION WAS LIKELY THE PRE-IMPLANT BAV. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVOLUTFX-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012480189); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21850977&pc=NPT | 2025587-2025-02737 | 2025-04-09 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-16 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH A BICUSPID ANATOMY WITH SOME CALCIUM TO THE RIGHT CORONARY CUSP (RCC)/NON-CORONARY CUSP (NCC) RAPHE, A PRE-BALLOON DILATION WAS PERFORMED WITH AN 18 MM NON-MEDTRONIC BALLOON. THE VALVE WAS 80% DEPLOYED AND WAS DETERMINED TO BE TOO HIGH. THE VALVE WAS RECAPTURED. AS THE VALVE WAS PULLED BACK, THE PATIENT'S PRESSURES BEGAN TO DROP AND IT WAS BELIEVED THAT THE VALVE WAS OBSTRUCTING OUTFLOW. THE VALVE WAS FULLY RECAPTURED AND THE PRESSURES CONTINUED TO DROP. THE VALVE WAS DEPLOYED RAPIDLY INTO A NEW POSITION. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED AND THE PATIENT SUFFERED A PULSELESS ELECTRICAL ACTIVITY (PEA). A LUCAS MACHINE WAS USED TO SUPPORT THE CPR. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION AND PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT NEVER REGAINED AN OUTPUT AND AFTER ATTEMPTS OF RESUSCITATION, THE PATIENT DIED. PER THE PHYSICIAN, THE CAUSE IS UNKNOWN HOWEVER IT WAS LIKELY A WIRE PERFORATION OR AN AORTIC INJURY SUCH AS AN ANNULAR RUPTURE OR DISSECTION. THERE WAS NO FLOW OF EXTRAVASATION NOTED OF FLUOROSCOPY FOLLOWING THE PROCEDURE.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH A BICUSPID ANATOMY WITH SOME CALCIUM TO THE RIGHT CORONARY CUSP (RCC)/NON-CORONARY CUSP (NCC) RAPHE, A PRE-BALLOON DILATION WAS PERFORMED WITH AN 18 MM NON-MEDTRONIC BALLOON. THE VALVE WAS 80% DEPLOYED AND WAS DETERMINED TO BE TOO HIGH. THE VALVE WAS RECAPTURED. AS THE VALVE WAS PULLED BACK, THE PATIENT'S PRESSURES BEGAN TO DROP AND IT WAS BELIEVED THAT THE VALVE WAS OBSTRUCTING OUTFLOW. THE VALVE WAS FULLY RECAPTURED AND THE PRESSURES CONTINUED TO DROP. THE VALVE WAS DEPLOYED RAPIDLY INTO A NEW POSITION. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED AND THE PATIENT SUFFERED A PULSELESS ELECTRICAL ACTIVITY (PEA). AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION AND PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT NEVER REGAINED AN OUTPUT AND AFTER ATTEMPTS OF RESUSCITATION, THE PATIENT DIED. PER THE PHYSICIAN, THE CAUSE IS UNKNOWN HOWEVER IT WAS LIKELY A WIRE PERFORATION OR AN AORTIC INJURY SUCH AS AN ANNULAR RUPTURE OR DISSECTION. THERE WAS NO FLOW OF EXTRAVASATION NOTED OF FLUOROSCOPY FOLLOWING THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY THE CAUSE OF DEATH WAS THE PERICARDIAL EFFUSION. THE CAUSE OF THE PERICARDIAL EFFUSION IS UNKNOWN. Manufacturer Narrative: UPDATED DATA: D4 H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. CORRECTED DATA: H6. THIS PATIENT CODE WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012631054); PRODUCT TYPE: 0195-HEART VALVES; SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX PLUS VALVE; PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21851040&pc=DSQ | 2916596-2025-02137 | 2025-02-28 05:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY SECONDARY TO MULTI-SYSTEM ORGAN FAILURE, IN A PERIPHERAL HOSPITAL. NO PUMP RELATED ISSUES WERE REPORTED. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. PUMP WAS WORKING AS EXPECTED. Manufacturer Narrative: E3 - OCCUPATION: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21851045&pc=DYB | 1820334-2025-00399 | 2025-03-19 04:00 | Death | COOK INC | 2025-04-16 04:00 | DYB | FLEXOR RAABE GUIDING SHEATH | Break | Cardiac Arrest | K142829 |  | 1 | Event Description: AS REPORTED, DURING AN ANGIOGRAM WITH REVASCULARIZATION ATTEMPT, A FLEXOR RAABE GUIDING SHEATH UNRAVELED AND SEPARATED. THE PATIENT HAD SIGNIFICANT ATHEROSCLEROTIC PERIPHERAL ARTERY DISEASE, AND THE REVASCULARIZATION ATTEMPT WAS FOR IMPROVEMENT OF WOUND HEALING ON THE RIGHT FOOT. ACCESS WAS REPORTEDLY OBTAINED IN THE POPLITEAL AND LEFT COMMON FEMORAL ARTERIES (LCFA). THE COMPLAINT DEVICE WAS ADVANCED VIA A CONTRALATERAL APPROACH FROM THE LCFA TO THE MID- SUPERFICIAL FEMORAL ARTERY, WITHOUT ENCOUNTERING RESISTANCE. THE ACCESS SITE WAS NOT SCARRED; HOWEVER, THE ANATOMY WAS CALCIFIED AND MILDLY TORTUOUS. THE BIFURCATION WAS ALSO CALCIFIED. VARIOUS 0.018-INCH WIRE GUIDES WERE USED THROUGH THE SHEATH DURING THE PROCEDURE, AS WELL AS A COOK MICROCATHETER, OTHER MANUFACTURERS' BALLOON CATHETERS, AND ANOTHER MANUFACTURER¿S GUIDING CATHETER. RESISTANCE WAS NOT ENCOUNTERED DURING INSERTION OR REMOVAL OF ANY DEVICE THROUGH THE SHEATH. DURING THE PROCEDURE, WHILE ATTEMPTING TO EXCHANGE THE COMPLAINT DEVICE WITH A 6-FRENCH SHEATH OVER ANOTHER MANUFACTURER'S WIRE IN ANTICIPATION OF STENTING, APPROXIMATELY TWENTY CENTIMETERS OF THE SHEATH "SHEARED" IN THE VESSEL UPON ATTEMPTED GRADUAL REMOVAL. THE DILATOR WAS NOT REINSERTED PRIOR TO ATTEMPTED REMOVAL OF THE SHEATH. RESISTANCE WAS REPORTEDLY NOT ENCOUNTERED DURING SHEATH REMOVAL. VASCULAR SURGERY WAS CONSULTED FOR EMERGENT EVALUATION OF THE RETAINED SHEATH. THE USER CONFIRMED THAT A WIRE DID NOT SEPARATE DURING THE PROCEDURE. PER THE CUSTOMER, THE SHEATH SEPARATION DID NOT RESULT IN A LIFE-THREATENING EVENT, AND THE PATIENT DID NOT BECOME HEMODYNAMICALLY UNSTABLE AFTER THE SHEATH SEPARATED. BILATERAL SURGICAL CUTDOWNS WERE PERFORMED, AND THE RETAINED FRACTURED SHEATH WAS REMOVED. THE PATIENT WAS SENT TO RECOVERY WITH GOOD PERFUSION. TWELVE DAYS AFTER THE PROCEDURE, THE PATIENT, WHO HAD A HISTORY OF CORONARY ARTERY DISEASE, DEVELOPED CHEST PAIN, EXPERIENCED CARDIAC ARREST AND DIED. AN AUTOPSY WILL NOT BE PROVIDED TO COOK TO AID IN THE INVESTIGATION OF THE EVENT. Manufacturer Narrative: H3: DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED, THAT A DEATH OR SERIOUS INJURY OCCURRED, NOR THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: BLANK FIELDS ON THIS FORM INDICATE THE INFORMATION IS UNKNOWN, UNCHANGED, OR UNAVAILABLE. SUMMARY OF EVENT: AS REPORTED, DURING AN ANGIOGRAM WITH REVASCULARIZATION ATTEMPT, A FLEXOR RAABE GUIDING SHEATH UNRAVELED AND SEPARATED. THE PATIENT HAD SIGNIFICANT ATHEROSCLEROTIC PERIPHERAL ARTERY DISEASE, AND THE REVASCULARIZATION ATTEMPT WAS FOR IMPROVEMENT OF WOUND HEALING ON THE RIGHT FOOT. ACCESS WAS REPORTEDLY OBTAINED IN THE POPLITEAL AND LEFT COMMON FEMORAL ARTERIES (LCFA). THE COMPLAINT DEVICE WAS ADVANCED VIA A CONTRALATERAL APPROACH FROM THE LCFA TO THE MID- SUPERFICIAL FEMORAL ARTERY, WITHOUT ENCOUNTERING RESISTANCE. THE ACCESS SITE WAS NOT SCARRED; HOWEVER, THE ANATOMY WAS CALCIFIED AND MILDLY TORTUOUS. THE BIFURCATION WAS ALSO CALCIFIED. VARIOUS 0.018-INCH WIRE GUIDES WERE USED THROUGH THE SHEATH DURING THE PROCEDURE, AS WELL AS A COOK MICROCATHETER, OTHER MANUFACTURERS¿ BALLOON CATHETERS, AND ANOTHER MANUFACTURER¿S GUIDING CATHETER. RESISTANCE WAS NOT ENCOUNTERED DURING INSERTION OR REMOVAL OF ANY DEVICE THROUGH THE SHEATH. DURING THE PROCEDURE, WHILE ATTEMPTING TO EXCHANGE THE COMPLAINT DEVICE WITH A 6-FRENCH SHEATH OVER ANOTHER MANUFACTURER¿S WIRE IN ANTICIPATION OF STENTING, APPROXIMATELY TWENTY CENTIMETERS OF THE SHEATH ¿SHEARED¿ IN THE VESSEL UPON ATTEMPTED GRADUAL REMOVAL. THE DILATOR WAS NOT REINSERTED PRIOR TO ATTEMPTED REMOVAL OF THE SHEATH. RESISTANCE WAS REPORTEDLY NOT ENCOUNTERED DURING SHEATH REMOVAL. VASCULAR SURGERY WAS CONSULTED FOR EMERGENT EVALUATION OF THE RETAINED SHEATH. THE USER CONFIRMED THAT A WIRE DID NOT SEPARATE DURING THE PROCEDURE. PER THE CUSTOMER, THE SHEATH SEPARATION DID NOT RESULT IN A LIFE-THREATENING EVENT, AND THE PATIENT DID NOT BECOME HEMODYNAMICALLY UNSTABLE AFTER THE SHEATH SEPARATED. BILATERAL SURGICAL CUTDOWNS WERE PERFORMED, AND THE RETAINED FRACTURED SHEATH WAS REMOVED. THE PATIENT WAS SENT TO RECOVERY WITH GOOD PERFUSION. TWELVE DAYS AFTER THE PROCEDURE, THE PATIENT, WHO HAD A HISTORY OF CORONARY ARTERY DISEASE, DEVELOPED CHEST PAIN, EXPERIENCED CARDIAC ARREST AND DIED. AN AUTOPSY WILL NOT BE PROVIDED TO COOK TO AID IN THE INVESTIGATION OF THE EVENT. INVESTIGATION EVALUATION: REVIEWS OF THE COMPLAINT HISTORY, DEVICE HISTORY RECORD (DHR), INSTRUCTIONS FOR USE (IFU), MANUFACTURING INSTRUCTIONS, AND QUALITY CONTROL PROCEDURES WERE CONDUCTED DURING THE INVESTIGATION. A VISUAL INSPECTION OF THE COMPLAINT DEVICE WAS ALSO CONDUCTED. THE COMPLAINT DEVICE WAS RETURNED TO COOK FOR INVESTIGATION IN TWO SECTIONS. THE PROXIMAL SECTION WAS UNRAVELED; HOWEVER, THE SEPARATED SHEATH MATERIAL WAS CONNECTED BY THE UNRAVELED METAL COIL. THE DISTAL SECTION WAS SEPARATED, MEASURING APPROXIMATELY 48-CENTIMETERS IN LENGTH. A DOCUMENT-BASED INVESTIGATION EVALUATION WAS PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORD FOUND NO RELEVANT NON-CONFORMANCES ON THE LOT. A REVIEW OF COMPLAINT HISTORY FOUND NO ADDITIONAL COMPLAINTS FOR THIS LOT NUMBER. THE PRODUCT IFU STATES THAT REINSERTION OF THE DILATOR PRIOR TO REMOVAL OF THE SHEATH ¿INCREASES THE STRENGTH OF THE SHEATH AND LESSENS THE RISK OF DEVICE SEPARATION¿ AND SUGGESTS INSERTION OF THE DILATOR PRIOR TO REMOVAL IF RESISTANCE IS ENCOUNTERED OR ANTICIPATED DURING WITHDRAWAL OF THE DEVICE. A REVIEW OF THE DEVICE MASTER RECORD (DMR) CONCLUDED THAT SUFFICIENT INSPECTION ACTIVITIES ARE IN PLACE TO IDENTIFY THIS FAILURE MODE PRIOR TO DISTRIBUTION. THE INFORMATION PROVIDED UPON REVIEW OF THE DMR, DHR, IFU, AND INVESTIGATION OF THE RETURNED DEVICE SUGGESTS THAT THERE IS EVIDENCE THE DEVICE WAS MANUFACTURED TO SPECIFICATION. THERE IS NO EVIDENCE OF NON-CONFORMING DEVICES IN-HOUSE OR IN THE FIELD. BASED ON THE INFORMATION PROVIDED AND THE RESULTS OF THE INVESTIGATION, COOK HAS CONCLUDED THAT THE PATIENT¿S ANATOMY LIKELY CONTRIBUTED TO THIS EVENT. THE ANATOMY WAS MILDLY TORTUOUS AND THE VESSELS, INCLUDING THE BIFURCATION, WERE CALCIFIED. THE DILATOR WAS NOT REINSERTED PRIOR TO ATTEMPTED REMOVAL OF THE SHEATH. ALTHOUGH RESISTANCE WAS REPORTEDLY NOT ENCOUNTERED DURING SHEATH ADVANCEMENT OR REMOVAL, THE SHEATH MATERIAL WAS BUNCHED/TWISTED/DEFORMED, UNRAVELED, AND SEPARATED UPON RETURN TO COOK. IT IS LIKELY THAT THE SHEATH CAUGHT WITHIN THE TORTUOUS AND CALCIFIED ANATOMY UPON REMOVAL OF THE DEVICE. THE RISK ANALYSIS FOR THIS FAILURE MODE WAS REVIEWED AND NO ADDITIONAL ESCALATION WAS REQUIRED. THE APPROPRIATE PERSONNEL HAVE BEEN NOTIFIED, AND COOK WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED, THAT A DEATH OR SERIOUS INJURY OCCURRED, NOR THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21851212&pc=DSQ | 2916596-2025-02258 | 2025-03-27 04:00 | Death | THORATEC CORPORATION | 2025-04-16 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN HOSPICE CARE WHILE INPATIENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A3: PATIENT GENDER WAS NOT PROVIDED. SECTION D4: LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21851228&pc=DSQ | 3007042319-2025-00771 | 2025-04-13 04:00 | Death | HEARTWARE, INC. | 2025-04-16 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke; Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT WITH A HISTORY OF DILATED CARDIOMYOPATHY EXPERIENCED A STROKE AND WAS SUBSEQUENTLY HOSPITALIZED. THE PATIENT WAS TRANSFERRED TO A PALLIATIVE CARE WARD AT AN EXTERNAL HOSPITAL NEAR THEIR HOME. THE ADVERSE EVENTS ASSOCIATED WITH THIS CASE INCLUDED NEUROLOGICAL DYSFUNCTION, SPECIFICALLY A STROKE, AND DEATH. THE PATIENT RECEIVED PALLIATIVE CARE AS A MEDICAL INTERVENTION. IT WAS UNKNOWN IF THE EVENT WAS RELATED TO THE DEVICE OR PROCEDURE, AND THERE WERE NO ALLEGATIONS AGAINST THE VENTRICULAR ASSIST DEVICE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IN USE. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. REVIEW OF THE DEVICE HISTORY RECORD CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING A STROKE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF NEUROLOGICAL DYSFUNCTION EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21851596&pc=NKM | 2135147-2025-02002 | 2025-03-20 04:00 | Death | ABBOTT MEDICAL | 2025-04-16 04:00 | NKM | MITRACLIP | Migration | Mitral Valve Insufficiency/ Regurgitation; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4+, THIN LEAFLETS, CALCIFIED ANNULUS, SEVERE AORTIC REGURGITATION AND TRICUSPID REGURGITATION. A MITRACLIP XTW WAS INSERTED AND DEPLOYED ON THE MITRAL VALVE. A SECOND MITRACLIP XTW WAS INSERTED AND DEPLOYED ON THE MITRAL VALVE, REDUCING MR TO A GRADE OF 2. HOWEVER, POST PROCEDURE, IT WAS OBSERVED THAT THE SECOND CLIP WAS SLIGHTLY TILTED LATERALLY. IT WAS NOTED THAT THE CLIPS WERE STABLE ON THE LEAFLETS. ON (B)(6) 2025, A CARDIAC SURGERY WAS PERFORMED TO REPAIR AORTIC VALVE, MITRAL VALVE, AND TRICUSPID VALVE. AN ECHOCARDIOGRAM WAS PERFORMED AND REVEALED THAT THE MR HAD INCREASED TO A GRADE OF 4. IT WAS NOTED THAT THE CLIPS REMAINED ATTACHED TO BOTH LEAFLETS. THE PROCEDURE WAS NOT SUCCESSFUL, AND THE PATIENT PASSED AWAY DURING THE INTERVENTION.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: IT WAS REPORTED THAT THE PATIENT'S STATUS PRIOR TO IMPLANT WAS SICK AND HAD AORTIC INSUFFICIENCY (AI) AND TRICUSPID INSUFFICIENCY (TI). DURING THE PROCEDURE, LEAFLET DAMAGE WAS OBSERVED. POST MITRACLIP PROCEDURE, A DECISION TO PERFORM A MULTI VALVE SURGICAL PROCEDURE WAS MADE BASED ON THE PATIENT STATUS POST MITRACLIP PROCEDURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON AVAILABLE INFORMATION AND PER THE PHYSICIAN, THE REPORTED PARTIAL CLIP MOVEMENT ASSOCIATED WITH THE CLIP TILTING SLIGHTLY WAS DUE TO PRE-EXISTING PATIENT ANATOMY. THE REPORTED RECURRENT MR SEEMS TO BE A CASCADING EFFECT OF THE REPORTED PARTIAL CLIP MOVEMENT. THE CAUSE OF THE REPORTED DEATH AND TISSUE INJURY WERE UNABLE TO BE DETERMINED. THE REPORTED PATIENT EFFECTS OF MITRAL REGURGITATION, DEATH, AND TISSUE INJURY, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED HOSPITALIZATION AND SURGICAL INTERVENTION WERE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21851938&pc=DRE | 3007284006-2025-00069 | 2025-03-28 04:00 | Death | SPECTRANETICS CORPORATION | 2025-04-16 04:00 | DRE | TIGHTRAIL ROTATING DILATOR SHEATH | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Great Vessel Perforation | K150360 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND A RIGHT ATRIAL (RA) LEAD DUE TO CIED SYSTEM/POCKET INFECTION AND BACTEREMIA. SPECTRANETICS LEAD LOCKING DEVICES (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH ON THE RV LEAD, THE LEAD WAS REMOVED SUCCESSFULLY. NEXT, TARGETING THE RA LEAD WITH THE 16F GLIDELIGHT, ADVANCEMENT WAS MADE TO THE INNOMINATE/SUPERIOR VENA CAVA (SVC) JUNCTION, WHERE PROGRESS STALLED. THEN, SWITCHING TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, THE RA LEAD WAS REMOVED. UPON REMOVAL OF THE LEAD, THE PATIENT'S BLOOD PRESSURE DROPPED, AND RESCUE EFFORTS BEGAN, INCLUDING STERNOTOMY. A PERFORATION FROM THE SVC/INNOMINATE JUNCTION AND THE SVC WAS DISCOVERED AND REPAIRED. THE PATIENT SURVIVED THE PROCEDURE; HOWEVER, HE PASSED AWAY THE FOLLOWING DAY. THIS REPORT CAPTURES THE 13F TIGHTRAIL DEVICE IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: D4/H4) THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE TIGHTRAIL WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS, GREAT VESSEL PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21852146&pc=BZD | 2518422-2025-104241 | 2025-02-20 05:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | K113068 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT DIED DUE TO COMPLICATIONS FROM A STROKE. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE REMSTAR AUTO A-FLEX DEVICE AND THE HARM REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI. Manufacturer Narrative: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21852187&pc=DSQ | 3007042319-2025-00772 | 2024-10-26 04:00 | Death | HEARTWARE, INC. | 2025-04-16 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Pneumonia; Cardiogenic Shock; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRANSCATHETER AORTIC VALVE REPLACEMENT IN PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE (VAD). INDICATIONS FOR AORTIC VALVE REPLACEMENT INCLUDED WORSENING HEART FAILURE AND CARDIOGENIC SHOCK. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE PROGRESSION OF HEART FAILURE, PNEUMONIA, AND OTHER UNKNOWN CAUSES. THERE WERE PATIENTS WHO EXPERIENCED STROKE, TRANSIENT ISCHEMIC ATTACK (TIA), MYOCARDIAL INFARCTION (MI), PACEMAKER IMPLANTATION, VASCULAR COMPLICATIONS, CARDIAC TAMPONADE, ACUTE KIDNEY INJURY (AKI), BLEEDING EVENTS WHICH REQUIRED TRANSFUSIONS, AND REINTERVENTIONS FOR UNKNOWN CAUSES. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/UNKNOWN. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. REFERENCED ARTICLE: POOLED COMPARATIVE ANALYSIS OF TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICE. GENERAL THORACIC AND CARDIOVASCULAR SURGERY. 2025. 73:125¿129. DOI: 10.1007/S11748-024-02100-3 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21852403&pc=BZD | 2518422-2025-104244 | 2025-03-31 04:00 | Death | RESPIRONICS, INC. | 2025-04-16 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A PATIENT'S SON STATING THE PATIENT DID NOT RECEIVE THE DEVICE. THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21852508&pc=NIK | 2017865-2025-40217 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NIK | GALLANT VR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN THE EMERGENCY ROOM EXPERIENCING A CHEST CONTUSION ON (B)(6) 2025, AND PASSED AWAY ON (B)(6) 2025. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21852521&pc=NVY | 2017865-2025-40218 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-16 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN THE EMERGENCY ROOM EXPERIENCING A CHEST CONTUSION ON (B)(6) 2025, AND PASSED AWAY ON (B)(6) 2025. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837991&pc=DTB | 2649622-2025-09980 | 2025-03-16 04:00 | Death | MPRI | 2025-04-15 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REPROGRAMMED AND BECAUSE OF THE IPG THE PATIENT PASSED AWAY THREE DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21838001&pc=DTB | 2649622-2025-09979 | 2025-03-16 04:00 | Death | MPRI | 2025-04-15 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REPROGRAMMED AND BECAUSE OF THE IPG THE PATIENT PASSED AWAY THREE DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21838007&pc=NVZ | 3008973940-2025-02617 | 2025-03-16 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-15 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS REPROGRAMMED AND BECAUSE OF THE IPG THE PATIENT PASSED AWAY THREE DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21838492&pc=NPT | 2025587-2025-02624 | 2024-06-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Respiratory Insufficiency | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 18 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS) OCCURRED AND A POLICY OF RELAXATION WAS CHOSEN. THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS ARDS. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21838769&pc=DYB | 9612164-2025-01913 | 2025-03-25 04:00 | Death | MEDTRONIC IRELAND | 2025-04-15 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED NON OPTIMAL TH RESHOLDS WHICH WERE ADJUDGED AS A MYOCARDIAL PROBLEM AND PLACEMENT COULD NOT BE PERFORMED. THE LEADLESS IPG WAS COLLECTED, NOT IMPLANTED AND IMPLANTATION WAS DELAYED AND PLANNED FOR A LATER DATE. DEATH WAS CONFIRMED IN THE STATE OF RESPIRATORY ARREST BEFORE THE PLANNED IMPLANT. AIRWAY OBSTRUCTION DUE TO AGING IS CONSIDERED A MAJOR FACTOR. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21838770&pc=PNJ | 9612164-2025-01912 | 2025-03-25 04:00 | Death | MEDTRONIC IRELAND | 2025-04-15 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED NON-OPTIMAL THRESHOLDS WHICH WERE ADJUDGED AS A MYOCARDIAL PROBLEM AND PLACEMENT COULD NOT BE PERFORMED. THE LEADLESS IPG WAS COLLECTED, NOT IMPLANTED AND IMPLANTATION WAS DELAYED AND PLANNED FOR A LATER DATE. DEATH WAS CONFIRMED IN THE STATE OF RESPIRATORY ARREST BEFORE THE PLANNED IMPLANT. AIRWAY OBSTRUCTION DUE TO AGING IS CONSIDERED A MAJOR FACTOR. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839131&pc=QZI | MW5169017 | 2025-03-06 05:00 | Death | MEDTRONIC, INC. | 2025-04-15 04:00 | QZI | PULSED FIELD ABLATION (PFA) CATHETER | Adverse Event Without Identified Device or Use Problem | Malaise |  |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE AFIB PFA PROCEDURE ON (B)(6) 2025, THE PATIENT EXPIRED. THE CALLER WAS MADE AWARE OF THE PATIENT DEATH BY HOSPITAL STAFF TODAY, (B)(6) 2025. IT WAS REPORTED TO THE CALLER THAT THE PATIENT EXPIRED "OVER THE WEEKEND," WITHIN THREE TO FOUR DAYS AFTER THE PROCEDURE, THE EXACT DATE IS UNKNOWN. STAFF REPORTED TO THE CALLER THAT THE PATIENT INFORMED THE FAMILY THAT THEY WERE NOT FEELING WELL BUT DID NOT GO TO THE HOSPITAL. NO FURTHER DETAILS WERE KNOWN. THE CALLER STATED THAT THE PROCEDURE ON (B)(6) 2025 WAS UNREMARKABLE, THERE WERE NO COMPLICATIONS. CALLER REPORTED THAT THE PATIENT HAD A POOR EJECTION FRACTION PRIOR TO THE PROCEDURE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839149&pc=LWS | 9614453-2025-01472 | 2025-03-27 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-15 04:00 | LWS | EVERA¿ S VR | Insufficient Information | Cardiac Arrest | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT 16 DAYS POST-IMPLANT OF AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WITH NO REPORTED PROCEDURAL COMPLICATIONS, THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO DIABETES-RELATED ORTHOPEDIC DISEASE AND TRANSFERRED TO THE CARDIOLOGY WARD THE FOLLOWING DAY. THE PATIENT'S CONDITION WAS REPORTED TO HAVE BEEN POOR DURING THE STAY IN THE HOSPITAL, WITH ELECTROLYTE IMBALANCE, FREQUENT HYPOKALEMIA, AND MANY UNDERLYING DISEASES. AFTER 17 DAYS IN-HOSPITAL, THE PATIENT UNDERWENT A JOINT REPLACEMENT PROCEDURE WHERE THEY SUFFERED A SUDDEN CARDIAC ARREST AFTER THE SURGERY. RESUSCITATION EFFORTS WERE PERFORMED FOR 30 MINUTES, AND PER HOSPITAL STAFF, THE PATIENT DID NOT RECEIVE ANY HIGH-VOLTAGE THERAPY. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) AND DISCHARGED HOME TWO DAYS LATER AT THE REQUEST OF THE PATIENT'S FAMILY. THE PATIENT SUBSEQUENTLY DIED TWO DAYS POST-DISCHARGE FOLLOWING FAILED RESCUE EFFORTS. A PATIENT REPRESENTATIVE STATED THAT THEY WERE INFORMED BY A CARDIOLOGIST FROM THE HOSPITAL THAT THE PATIENT'S ICD DID NOT WORK DURING THE CARDIAC ARREST. THE PATIENT REPRESENTATIVE STATED THAT BECAUSE OF THIS INFORMATION, THEY BELIEVE THERE WAS A PROBLEM WITH THE PRODUCT QUALITY WHICH LED TO THE PATIENT'S DEATH AFTER SURGERY. THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839237&pc=OZP | 2032227-2024-282988 | 2024-11-15 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-15 04:00 | OZP | PUMP MMT-1884L MM780G BLE 3.0 MG/ML | Appropriate Term/Code Not Available | Insufficient Information | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AND THE CUSTOMER'S NEXT OF KIN SUSPECTED THAT ITS THE PUMP THAT CAUSE THE PATIENTS DEATH. CAUSE OF DEATH WAS UNKNOWN CAUSE OF DEATH. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884L. TROUBLESHOOTING WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE SMARTGUARD/AUTO MODE FEATURE OF THE INSULIN PUMP. NO FURTHER COMPLICATIONS WERE REPORTED. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE THE USE OF THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884L. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839534&pc=DSI | 1218950-2025-000236 | 2025-04-07 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-04-15 04:00 | DSI | MX40 1.4 GHZ SMART HOPPING | Defective Alarm | Insufficient Information | K113125 |  | 1 | Event Description: THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS FIELD SERVICE PERSONNEL AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON AN MX40 1.4 GHZ SMART HOPPING INDICATING THAT AN ASYSTOLE ALARM WAS NOT GENERATED AT THE CENTRAL STATION FOR BED 284 AND THE PATIENT PASSED AWAY. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. Manufacturer Narrative: A PHILIPS TECHNICAL CONSULTANT (TC) ONSITE INVESTIGATED THE REPORTED ISSUE. IT WAS NOTED THAT THE NURSE FOUND THE PATIENT IN DISTRESS, BECAUSE THE PATIENT HAD BLUISH COLOR TO THE SKIN. THE CUSTOMER PULLED UP A STRIP DISPLAY ON THE CENTRAL STATION THAT SHOWED THE INCIDENT IN QUESTION OCCURRED AT 10:29:40AM ON (B)(6) 2025 WITH A RED ASYSTOLE BANNER SHOWING ABOVE THE STRIP DISPLAY. THE TC REQUESTED MX 40 BE RELEASED TO THEM FOR SEQUESTRATION. THE TC WAS INFORMED AT THAT TIME THAT THE PATIENT HAD EXPIRED AND THAT THE MX40 WAS ON ANOTHER PATIENT. THE CUSTOMER SWAPPED OUT THE MX40 AND GAVE TO THE TC. THE TC WENT TO CENTRAL AND VERIFIED WITH STAFF THAT AUDIBLE ASYSTOLE ALARMS WERE AVAILABLE UTILIZING A SEPARATE DEPARTMENT MX40 DEVICE. THE TC CONTACTED ANOTHER TC WHO WAS ONSITE LATER IN THE DAY ON (B)(6) 2025. THIS TC WAS ABLE TO REMOTELY LOG INTO THE CENTRAL STATION AND PULL LOGS FOR THIS INCIDENT. THE LOGS SHOW THAT THE DEVICE HAD ALARMED FOR AN ASYSTOLE AND HAD BEEN ACKNOWLEDGED. (B)(6) 2025 10:30:19 (B)(6) 2-284 ASYSTOLE ENDED. (B)(6) 2025,10:30:19, (B)(6), 2-284, ASYSTOLE ENDED. (B)(6) 2025,10:30:19, (B)(6), 2-284, ASYSTOLE ENDED. (B)(6) 2025,10:30:01, (B)(6), 2-284, ACKNOWLEDGE. (B)(6) 2025,10:30:01, (B)(6), 2-284, ACKNOWLEDGE. (B)(6) 2025,10:29:57, (B)(6), 2-284, ASYSTOLE GENERATED AT 10:29:46. (B)(6) 2025,10:29:57, (B)(6), 2-284, ASYSTOLE GENERATED AT 10:29:46. (B)(6) 2025,10:29:57, (B)(6), 2-284, ASYSTOLE GENERATED AT 10:29:46. THE TC USED A SIMULATOR TO SHOW A HEART RATE OF 70/BPM AND MONITORED THE CENTRAL STATION. THE TC PLACED THE SIMULATOR INTO AN ASYSTOLE; THE HEART RATE TRANSMITTED PROPERLY, AND THE CENTRAL ALARMED FOR ASYSTOLE. THE TC THEN RETURNED THE SIMULATOR TO 70/BPM; THE CENTRAL STATION MONITORED PROPERLY, AND THE ASYSTOLE CLEARED. THIS TEST WAS REPEATED AND ALSO PASSED. THE TC WAS UNABLE TO DUPLICATE THE ORIGINAL COMPLAINT OF THE DEVICE NOT ALARMING. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839659&pc=DWF | 21839659 | 2025-01-17 05:00 | Death | CARDIAC ASSIST, INC. | 2025-04-15 04:00 | DWF | PROTEK DUO | Adverse Event Without Identified Device or Use Problem | Dyspnea; Exsanguination; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Hypovolemic Shock; Mitral Valve Stenosis; Electrolyte Imbalance; Cardiogenic Shock; Respiratory Acidosis; Thrombocytopenia |  |  | 1 | Event Description: PATIENT DEVELOPED PROGRESSIVE DYSPNEA FOLLOWING COVID INFECTION [DATE REDACTED]. RECENT SURFACE ECHOCARDIOGRAM REVEALED PROGRESSION TO SEVERE MITRAL STENOSIS. TRANSESOPHAGEAL ECHOCARDIOGRAM WAS CONFIRMATORY FOR SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 MMHG WITH SEVERE THICKENING OF THE POSTERIOR GREATER THAN ANTERIOR MITRAL VALVE LEAFLETS AND MITRAL ANNULAR CALCIFICATION. CARDIAC CATHETERIZATION CONFIRMED MODERATE AORTIC AND SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 AND 35 MMHG, LEFT MAIN DISEASE EXTENDING INTO THE PROXIMAL LAD [LEFT ANTERIOR DESCENDING]. PATIENT MET IN CONSULTATION WITH DR. AND HAS ELECTED TO UNDERGO SURGICAL INTERVENTION. PATIENT WAS ADMITTED PREOPERATIVELY ON [DATE REDACTED] FOR COMPLETION OF PREOPERATIVE TESTING. ON [DATE REDACTED], PATIENT UNDERWENT REPLACEMENT OF IMPELLA WITH PROTEK DEVICE IN THE RIGHT HEART. APPROX. 1 HOUR AFTER RETURN TO THE CVICU [CARDIOVASCULAR INTENSIVE CARE UNIT] AFTER DEVICE PLACEMENT, THE CHEST TUBE OUTPUT INCREASED DRAMATICALLY, AND PRESSOR REQUIREMENTS ALSO INCREASED. THE PATIENT WAS GIVEN MULTIPLE MASSIVE TRANSFUSION PACKS WHILE THE FAMILY AND SURGICAL TEAM DETERMINED THE BEST PATH FORWARD, AND ULTIMATELY ELECTED TO HAVE THE PATIENT PROCEED WITH COMFORT MEASURES. TRANSFUSIONS WERE STOPPED, AND THE PATIENT PROMPTLY PASSED SECONDARY TO HYPOVOLEMIC SHOCK. CVICU INTENSIVIST NOTE: ANESTHESIOLOGY CRITICAL CARE: WAS CALLED TO BEDSIDE WITH INCREASED CHEST TUBE OUTPUT, FIRST PRIMARILY FROM LEFT PLEURAL TUBE AT APPROXIMATELY 200CC/HR AND SUBSEQUENTLY FROM THE MEDIASTINAL TUBES AS WELL. MASSIVE BLOOD TRANSFUSION WITH EQUAL PRBC [PACKED RED BLOOD CELLS], FFP [FRESH FROZEN PLASMA], PLATELETS WERE STARTED AND ESCALATION OF PRESSORS. TEE [TRANSESOPHAGEAL ECHOCARDIOGRAM] WAS DONE WHICH SHOWED THAT THE NEWLY PLACED PROTEK DUO CATH WAS IN THE RV [RIGHT VENTRICLE] WITH WHAT LOOKED LIKE AN RV INJURY PRESENT. SURGICAL TEAM WAS MADE AWARE AND PRESENT DURING DISCUSSIONS. WE SPOKE WITH FAMILY WHO AGREED THAT FURTHER SURGERY WOULD LIKELY NOT RESULT IN AN OUTCOME AND A QUALITY OF LIFE THAT THE PATIENT WOULD WANT. WE PROCEEDED TO CORRECT ANY COAGULOPATHY AND DESPITE THIS, THE BLEEDING CONTINUED. WE THEN SPOKE WITH FAMILY AND AGREED TO PROCEED WITH COMFORT CARE MEASURES. I PROVIDED 184 MINUTES OF CRITICAL CARE SERVICES WHICH WERE NECESSARY DUE TO THE PATIENT'S FOLLOWING CONDITION NECESSITATING CRITICAL CARE MANAGEMENT: ACIDOSIS, ACUTE BLOOD LOSS ANEMIA, ACUTE KIDNEY INJURY, ARRHYTHMIA, CARDIOGENIC SHOCK, COAGULOPATHY, ELECTROLYTE ABNORMALITIES, HYPOTENSION, HYPOVOLEMIC SHOCK, LIVER FAILURE, MULTISYSTEM ORGAN FAILURE, RENAL FAILURE, THROMBOCYTOPENIA. THE CRITICAL CARE TREATMENT AND MANAGEMENT INCLUDED ALKALI THERAPY, ANTIBIOTICS, BLOOD PRODUCT ADMINISTRATION, CONTINUOUS RENAL REPLACEMENT THERAPY, DISCUSSION OF PATIENT'S CONDITION AND THERAPEUTIC OPTIONS WITH MEDICAL DECISION MAKERS, ELECTROLYTE AND NUTRITION MANAGEMENT, FLUID ADMINISTRATION, INVASIVE HEMODYNAMIC MONITORING, MANAGEMENT OF MECHANICAL VENTILATION, REVIEW OF MEDICAL RECORDS, SEDATION MANAGEMENT, TITRATION OF INOTROPE INFUSIONS, TITRATION OF VASOPRESSOR INFUSIONS, DISCUSSION WITH CONSULTING SERVICES, TITRATION OF ANTICOAGULANTS.Event Description: PATIENT DEVELOPED PROGRESSIVE DYSPNEA FOLLOWING COVID INFECTION [DATE REDACTED]. RECENT SURFACE ECHOCARDIOGRAM REVEALED PROGRESSION TO SEVERE MITRAL STENOSIS. TRANSESOPHAGEAL ECHOCARDIOGRAM WAS CONFIRMATORY FOR SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 MMHG WITH SEVERE THICKENING OF THE POSTERIOR GREATER THAN ANTERIOR MITRAL VALVE LEAFLETS AND MITRAL ANNULAR CALCIFICATION. CARDIAC CATHETERIZATION CONFIRMED MODERATE AORTIC AND SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 AND 35 MMHG, LEFT MAIN DISEASE EXTENDING INTO THE PROXIMAL LAD [LEFT ANTERIOR DESCENDING]. PATIENT MET IN CONSULTATION WITH DR. AND HAS ELECTED TO UNDERGO SURGICAL INTERVENTION. PATIENT WAS ADMITTED PREOPERATIVELY ON [DATE REDACTED] FOR COMPLETION OF PREOPERATIVE TESTING. ON [DATE REDACTED], PATIENT UNDERWENT REPLACEMENT OF IMPELLA WITH PROTEK DEVICE IN THE RIGHT HEART. APPROX. 1 HOUR AFTER RETURN TO THE CVICU [CARDIOVASCULAR INTENSIVE CARE UNIT] AFTER DEVICE PLACEMENT, THE CHEST TUBE OUTPUT INCREASED DRAMATICALLY, AND PRESSOR REQUIREMENTS ALSO INCREASED. THE PATIENT WAS GIVEN MULTIPLE MASSIVE TRANSFUSION PACKS WHILE THE FAMILY AND SURGICAL TEAM DETERMINED THE BEST PATH FORWARD, AND ULTIMATELY ELECTED TO HAVE THE PATIENT PROCEED WITH COMFORT MEASURES. TRANSFUSIONS WERE STOPPED, AND THE PATIENT PROMPTLY PASSED SECONDARY TO HYPOVOLEMIC SHOCK. CVICU INTENSIVIST NOTE: ANESTHESIOLOGY CRITICAL CARE: WAS CALLED TO BEDSIDE WITH INCREASED CHEST TUBE OUTPUT, FIRST PRIMARILY FROM LEFT PLEURAL TUBE AT APPROXIMATELY 200CC/HR AND SUBSEQUENTLY FROM THE MEDIASTINAL TUBES AS WELL. MASSIVE BLOOD TRANSFUSION WITH EQUAL PRBC [PACKED RED BLOOD CELLS], FFP [FRESH FROZEN PLASMA], PLATELETS WERE STARTED AND ESCALATION OF PRESSORS. TEE [TRANSESOPHAGEAL ECHOCARDIOGRAM] WAS DONE WHICH SHOWED THAT THE NEWLY PLACED PROTEK DUO CATH WAS IN THE RV [RIGHT VENTRICLE] WITH WHAT LOOKED LIKE AN RV INJURY PRESENT. SURGICAL TEAM WAS MADE AWARE AND PRESENT DURING DISCUSSIONS. WE SPOKE WITH FAMILY WHO AGREED THAT FURTHER SURGERY WOULD LIKELY NOT RESULT IN AN OUTCOME AND A QUALITY OF LIFE THAT THE PATIENT WOULD WANT. WE PROCEEDED TO CORRECT ANY COAGULOPATHY AND DESPITE THIS, THE BLEEDING CONTINUED. WE THEN SPOKE WITH FAMILY AND AGREED TO PROCEED WITH COMFORT CARE MEASURES. I PROVIDED 184 MINUTES OF CRITICAL CARE SERVICES WHICH WERE NECESSARY DUE TO THE PATIENT'S FOLLOWING CONDITION NECESSITATING CRITICAL CARE MANAGEMENT: ACIDOSIS, ACUTE BLOOD LOSS ANEMIA, ACUTE KIDNEY INJURY, ARRHYTHMIA, CARDIOGENIC SHOCK, COAGULOPATHY, ELECTROLYTE ABNORMALITIES, HYPOTENSION, HYPOVOLEMIC SHOCK, LIVER FAILURE, MULTISYSTEM ORGAN FAILURE, RENAL FAILURE, THROMBOCYTOPENIA. THE CRITICAL CARE TREATMENT AND MANAGEMENT INCLUDED ALKALI THERAPY, ANTIBIOTICS, BLOOD PRODUCT ADMINISTRATION, CONTINUOUS RENAL REPLACEMENT THERAPY, DISCUSSION OF PATIENT'S CONDITION AND THERAPEUTIC OPTIONS WITH MEDICAL DECISION MAKERS, ELECTROLYTE AND NUTRITION MANAGEMENT, FLUID ADMINISTRATION, INVASIVE HEMODYNAMIC MONITORING, MANAGEMENT OF MECHANICAL VENTILATION, REVIEW OF MEDICAL RECORDS, SEDATION MANAGEMENT, TITRATION OF INOTROPE INFUSIONS, TITRATION OF VASOPRESSOR INFUSIONS, DISCUSSION WITH CONSULTING SERVICES, TITRATION OF ANTICOAGULANTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839835&pc=PYX | 1220648-2025-27471 | 2025-03-18 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Stroke/CVA; Hemolysis | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WITH RIGHT HEART FAILURE ARRIVED AT THE CARDIAC CATHETERIZATION LAB FROM THE INTENSIVE CARE UNIT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND VENTILATOR. THE IMPELLA RP WAS INSERTED. ECMO WAS DECREASED AND WORKING IN CONJUNCTION WITH THE IMPELLA RP FLEX. THE NURSE STATED THAT THE IMPELLA RP PLACEMENT SIGNAL AND PA SIGNAL WERE MISSING AND THAT THERE COULD BE A THROMBUS IN THE CANNULA. HEMOGLOBIN HAD INCREASED AND DECREASES OF FLOWS WERE NOTED. ADDITIONALLY, THE PATIENT DEVELOPED HEMOLYSIS AND DUE TO INTRACRANIAL HEMORRHAGE, THE PATIENTS CARE WAS WITHDRAWN. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS THE PATIENT PASSED AWAY AND THE PUMP WAS NOT REMOVED. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839853&pc=FBN | 3005099803-2025-01652 | 2025-03-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-15 04:00 | FBN | SPYSCOPE DS II | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K183636 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT SPYSCOPE DS II ACCESS AND DELIVERY CATHETER WAS USED IN THE DUODENUM FOR THE TREATMENT OF STONES ON (B)(6) 2025. DURING THE SPYGLASS PROCEDURE, WHILE PERFORMING IRRIGATION THE ANESTHETIST WARNED OF A RAPID DETERIORATION IN THE PATIENTS VITALS. THE DEVICE WAS SAFELY REMOVED, AND DOCTORS ATTEMPTED CPR. HOWEVER, THE PATIENT PASSED AWAY SHORTLY AFTER. AIR WAS USED FOR INSUFFLATION DURING THE PROCEDURE. IN THE PHYSICIANS ASSESSMENT THE DEVICE DID NOT CONTRIBUTE TO THE PATIENTS COMPLICATIONS. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. Manufacturer Narrative: BLOCK H6: IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST. IMDRF IMPACT CODE F2306 CAPTURES THE REPORTABLE EVENT OF CARDIOPULMONARY RESUSCITATION. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21839869&pc=DSQ | MW5169018 | 2025-02-22 05:00 | Death | BERLIN HEART, INC. | 2025-04-15 04:00 | DSQ | BERLIN HEART | Material Separation | Air Embolism |  |  | 1 | Event Description: BERLIN CANNULA SEPARATED FROM THE CANNULAS PLASTIC SUTURED SEWING RING ON THE DAY OF HEART TRANSPLANT CAUSING AIR EMBOLISM/ DEATH OF MY 2 YEAR OLD SON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21840049&pc=LWS | 2649622-2025-10132 | 2025-03-16 04:00 | Death | MPRI | 2025-04-15 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing; Low impedance | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED INTERMITTENT UNDERSENSING AND TRIGGERED OUT OF RANGE (OOR) ALERTS FOR LOW PACING IMPEDANCE. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21840294&pc=NIQ | 2124215-2025-21909 | 2025-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-15 04:00 | NIQ | SYNERGY? XD | Break; Difficult to Remove | Cardiac Arrest; Chest Pain; Respiratory Arrest; Insufficient Information | P150003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH THE DEPLOYMENT OF A 2.75 X 24 MM SYNERGY XD DRUG-ELUTING STENT. DURING THE PROCEDURE, THE BALLOON SHAFT FRACTURED PROXIMAL TO THE MID-LEFT ANTERIOR DESCENDING (LAD) ARTERY POST-DEPLOYMENT. ALTHOUGH THE STENT WAS SUCCESSFULLY DEPLOYED, THE BALLOON COULD NOT BE REMOVED. THE FRACTURED DEVICE WAS EXTRACTED USING A SNARE. AFTER THE REMOVAL, THE PATIENTS CONDITION DETERIORATED RAPIDLY. A FULL CODE WAS INITIATED IN AN ATTEMPT TO STABILIZE THE PATIENT, BUT DESPITE ALL RESUSCITATIVE EFFORTS, THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) WITH THE DEPLOYMENT OF A 2.75 X 24MM SYNERGY XD DRUG-ELUTING STENT. DURING THE PROCEDURE, THE BALLOON SHAFT FRACTURED PROXIMAL TO THE MID-LEFT ANTERIOR DESCENDING (LAD) ARTERY POST-DEPLOYMENT. ALTHOUGH THE STENT WAS SUCCESSFULLY DEPLOYED, THE BALLOON COULD NOT BE REMOVED. THE FRACTURED DEVICE WAS EXTRACTED USING A SNARE. AFTER THE REMOVAL, THE PATIENTS CONDITION DETERIORATED RAPIDLY. A FULL CODE WAS INITIATED IN AN ATTEMPT TO STABILIZE THE PATIENT, BUT DESPITE ALL RESUSCITATIVE EFFORTS, THE PATIENT PASSED AWAY. IT WAS FURTHER REPORTED VIA MEDWATCH MW5153914 THAT THE PROCEDURE BEGAN WITH VASCULAR ACCESS THROUGH THE RIGHT GROIN. DIAGNOSTIC IMAGING IDENTIFIED A 95% STENOSED TARGET LESION LOCATED IN THE MILDLY TORTUOUS AND SEVERELY CALCIFIED PROXIMAL TO MID SEGMENT OF THE LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE LESION WAS PRE-DILATED USING A 2.0 X 20 MM NON-BOSTON SCIENTIFIC BALLOON CATHETER. A STENT WAS THEN DEPLOYED AT 20 ATMOSPHERES. DURING DEPLOYMENT, THE STENT FAILED TO FULLY EXPAND, AND THE BALLOON COULD NOT BE WITHDRAWN. ANGIOGRAPHY REVEALED THAT THE BALLOON MARKERS HAD BEEN LEFT WITHIN THE VESSEL. THE PATIENT BEGAN EXPERIENCING CHEST DISCOMFORT AND WAS INTUBATED. ALTHOUGH THE RETAINED BALLOON MARKERS WERE SUCCESSFULLY RETRIEVED, THE PATIENTS CONDITION CONTINUED TO DETERIORATE, LEADING TO CARDIAC AND RESPIRATORY ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND AN ATTEMPT WAS MADE TO PLACE AN IMPELLA DEVICE, BUT IT WAS UNSUCCESSFUL. THE PATIENT DIED. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, SUCH AS HOW MANY SECONDS DID THE USER ALLOW FOR THE BALLOON TO DEFLATE PRIOR TO ATTEMPTING TO WITHDRAW BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: B5 - DESCRIBE EVENT OR PROBLEM: UPDATED TO INCLUDE NEW INFORMATION OBTAINED. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, SUCH AS HOW MANY SECONDS DID THE USER ALLOW FOR THE BALLOON TO DEFLATE PRIOR TO ATTEMPTING TO WITHDRAW BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21840302&pc=QAU | 1037905-2025-00284 | 2025-01-16 05:00 | Death | WILSON-COOK MEDICAL INC | 2025-04-15 04:00 | QAU | HEMOSPRAY ENDOSCOPIC HEMOSTAT | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | K200972 |  | 1 | Event Description: DURING AN ENDOSCOPIC PROCEDURE TO TREAT A BLEEDING ULCER IN THE STOMACH, THE PHYSICIAN USED A COOK HEMOSPRAY ENDOSCOPIC HEMOSTAT. IT WAS REPORTED THAT THE PATIENT HAD BEEN REPEATEDLY BLEEDING FROM A DUODENAL ULCER FOR THREE DAYS, AND PRIOR TO THIS TREATMENT, HE HAD BEEN TREATED WITH CLIPS TO STOP THE BLEEDING TWICE. DUE TO ANEMIA, LOW BLOOD PRESSURE, AND TARRY STOOLS, THE PHYSICIAN UNDERWENT ANOTHER ENDOSCOPY. THERE WAS A LARGE AMOUNT OF BLOOD CLOTS AND BLEEDING FROM THE DUODENAL ULCER. THEY USED TWO CLIPS AND SPRAYED WITH HEMOSPRAY. AT THAT POINT, THE BLEEDING HAD STOPPED, BUT ABOUT 9 HOURS LATER, BLOOD PRESSURE DROPPED DUE TO RE-BLEEDING, AND AN ENDOSCOPY WAS PERFORMED. DUE TO GUSHING BLEEDING, IT WAS NOT POSSIBLE TO STOP THE BLEEDING, AND THE SCOPE WAS REMOVED BECAUSE THE PATIENT'S VITALS COLLAPSED, AND THE PATIENT DIED WITHOUT BEING ABLE TO UNDERGO SURGERY OR AG TREATMENT. PER THE PHYSICIAN, "I BELIEVE THAT THE CAUSE OF REBLEEDING WAS THE EXCESSIVE STIMULATION OF EXPOSED BLOOD VESSELS BY APPLYING CLIPS, AND THE FACT THAT THE HEMOSTATIC AGENT WAS NOT DISTRIBUTED PROPERLY DUE TO THE PRESENCE OF CLIPS AT THE BLEEDING SITE." Manufacturer Narrative: ADDITIONAL INFORMATION: SECTION E PHONE: (B)(6). THE INVESTIGATION IS ON-GOING. A FOLLOW-UP EMDR WILL BE PROVIDED WITHIN 30 DAYS OF SUBMISSION OF THIS REPORT. Manufacturer Narrative: ADDITIONAL INFORMATION: SECTION E PHONE: (B)(6). INVESTIGATION EVALUATION: A PRODUCT EVALUATION WAS NOT PERFORMED IN RESPONSE TO THIS REPORT BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT PROVIDED TO COOK FOR EVALUATION. THE REPORT COULD NOT BE CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE CONDUCTED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. INVESTIGATION CONCLUSION: WE COULD NOT CONDUCT A COMPLETE INVESTIGATION BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT RETURNED FOR EVALUATION. A DEFINITIVE CAUSE FOR THE REPORTED OBSERVATION COULD NOT BE DETERMINED. HOWEVER, A DEVICE FAILURE WAS NOT REPORTED. BASED ON THE RECEIVED INFORMATION, THE PATIENTS SUFFERED A REBLEED AFTER SUCCESSFUL HEMOSTASIS. THE IFU STATES, "ALL ENDOSCOPIC HEMOSTATIC THERAPIES, INCLUDING HEMOSPRAY, HAVE AN ASSOCIATED RISK OF REBLEEDING, PARTICULARLY IN SITUATIONS WHERE THE CAUSE OF BLEEDING IS AN UNRESOLVED UNDERLYING DISEASE. AFTER HEMOSTASIS HAS BEEN ACHIEVED, MONITOR PATIENTS FOR REBLEEDING PER THE RELEVANT CLINICAL GUIDELINES." THE IFU PRECAUTIONS, "LIKE OTHER MODALITIES, HEMOSPRAY MAY NOT BE EFFECTIVE FOR ALL TYPES OF BLEEDS. GASTROINTESTINAL BLEEDING MAY EXACERBATE EXISTING COMORBIDITIES, INCREASING THE POTENTIAL FOR ADVERSE EVENTS INCLUDING PATIENT MORTALITY." PRIOR TO DISTRIBUTION, ALL HEMOSPRAY ENDOSCOPIC HEMOSTATS ARE SUBJECTED TO A VISUAL INSPECTION TO ENSURE DEVICE INTEGRITY. CORRECTIVE ACTION: BASED ON THE QUALITY ENGINEERING RISK ASSESSMENT NO CORRECTIVE ACTION IS WARRANTED AT THIS TIME. A REVIEW OF THE COMPLAINT HISTORY WAS CONDUCTED. BASED ON THIS REVIEW, THE LIKELIHOOD OF THIS TYPE OF REPORT IS RARE. QUALITY ASSURANCE WILL CONTINUE TO MONITOR FOR COMPLAINT TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21840343&pc=DSQ | 2916596-2025-02187 | 2024-06-12 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Nervous System Injury | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS A DEVASTATING NEUROLOGICAL INJURY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING NEUROLOGICAL EVENTS (NOT STROKE RELATED), AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGIC DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21840594&pc=DSQ | 2916596-2025-02086 | 2025-03-11 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Mitral Valve Insufficiency/ Regurgitation; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED DUE TO DETERIORATION OF CLINICAL CONDITIONS INCLUDING SHORTNESS OF BREATH AT REST AND FEELING LACK OF AIR. THE PATIENT WAS DIAGNOSED WITH MASSIVE MITRAL VALVE REGURGITATION (MR) AND WAS NOT REACTING ON LEFT VENTRICULAR ASSIST DEVICE (LVAD) SPEED INCREASE WHILE THE SIZE OF THE LEFT VENTRICLE (LV) GOT SMALLER. THE PATIENT WAS ON INOTROPES AND VENTILATION IN THE INTENSIVE CARE UNIT (ICU). COMPUTED TOMOGRAPHY ANGIOGRAPHY (CT) SHOWED NO OBSTRUCTION AT THE HEARTMATE 3 OUTFLOW GRAFT OR AROUND THE LVAD INFLOW. ECHOCARDIOGRAM RAMP TEST WAS SUCCESSFUL, AND THE LV SIZE GOT SMALLER WITH HIGHER ROTATIONS PER MINUTE (RPM), THE AORTIC VALVE WAS CLOSED AT HIGHER SPEED RANGE. BLOOD LABS SHOWED NO HEMOLYSIS. THE PATIENT WAS STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) DUE TO ANURIA. INOTROPES WERE ALSO STARTED, AND FURTHER STEPS WERE AWAITING CONFIRMATION. THE SUBMITTED LOG FILES IT APPEARED THAT THE SPEED WAS TURNED DOWN BRIEFLY TO 4000RPM CAUSING LOW SPEED ADVISORIES AND ALSO THE FLOWS DROPPED BELOW 2.5LPM. OTHER THAN THAT, THE LVAD APPEARED TO FUNCTION AS INTENDED. IT WAS NOTED THAT THE PULSATILITY INDEX (PI) VALUES WERE ON THE LOW SIDE, BUT IT COULD BE DUE TO THE VALVE CONCERN. ADDITIONAL INFORMATION WAS RECEIVED THAT THE RENAL DYSFUNCTION AND MR WERE NOT RELATED TO OR CAUSED BY THE DEVICE OR DEVICE THERAPY. THE PATIENT DEVELOPED IMPAIRED RENAL FUNCTION SINCE (B)(6) 2025 WITH CREATININE INCREASE FROM 121 TO 265. SEVERAL TRANSTHORACIC ECHOCARDIOGRAMS (TTE) AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SHOWED SEVERE MR WITH SEVERE CENTRAL AND ECCENTRIC REGURGITATION. FLOW WAS OCCUPYING 1/3 OF THE LEFT ATRIAL CAVITY AND THE LEFT VENTRICULAR END-DIASTOLIC DIAMETER (LVEDD) WAS 9.6CM. RIGHT HEART CATHETERIZATION PERFORMED ON (B)(6) 2025 SHOWED INCREASED PRESSURE IN THE RIGHT CHAMBERS AND THE PULMONARY CAPILLARY WEDGE PRESSURE (PCWP) WAS 41MMHG. THE PATIENT WAS TREATED WITH DIURETIC THERAPY, BASIS CONGESTIVE HEART FAILURE THERAPY. DESPITE THE ONGOING THERAPY, THE CREATININE LEVEL INCREASED TO 335 MMOL/L AND HEMODIALYSIS WAS STARTED ON (B)(6) 2025 WITH NO EFFECT. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PUMP WAS NOT EXPLANTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. THE ONSET OF THE MITRAL REGURGITATION IS REPORTED UNDER 2916596-2025-02619. Manufacturer Narrative: SECTION A: DUE TO PRIVACY POLICY PATIENT INFORMATION COULD NOT BE PROVIDED BY THE SITE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 06MAR2025 THROUGH 20MAR2025. NO NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE FIXED SPEED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSRTUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", WARNS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A: PATIENT INFORMATION REQUESTED. SECTION E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21840896&pc=NPT | 2025587-2025-02670 | 2025-03-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPT | EVOLUT FX DCS | Device Dislodged or Dislocated; Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING DEPLOYMENT, THE FIRST VALVE DISLODGED. A SECOND VALVE WAS LOADED INTO A DELIVERY CATHETER SYSTEM; HOWEVER, THE PATIENT BEGAN TO DECOMPENSATE BEFORE DELIVERY COULD BEGIN. AN ASCENDING AORTIC DISSECTION WAS IDENTIFIED. INTERVENTIONAL RADIOLOGY AND VASCULAR SURGERY WERE CALLED. AN ATTEMPT WAS MADE TO SNARE THE DISLODGED VALVE AND PLACE A STENT. DESPITE THESE EFFORTS, THE PATIENT DIED. THE MEDTRONIC DEVICE WAS NOTED TO HAVE CONTRIBUTED TO THE ADVERSE OUTCOME. OTHER CONTRIBUTING FACTORS INCLUDED THE PATIENT'S DILATED ASCENDING AORTA.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING DEPLOYMENT, THE FIRST VALVE DISLODGED. A SECOND VALVE WAS LOADED INTO A DELIVERY CATHETER SYSTEM; HOWEVER, THE PATIENT BEGAN TO DECOMPENSATE BEFORE DELIVERY COULD BEGIN. AN ASCENDING AORTIC DISSECTION WAS IDENTIFIED. INTERVENTIONAL RADIOLOGY AND VASCULAR SURGERY WERE CALLED. AN ATTEMPT WAS MADE TO SNARE THE DISLODGED VALVE AND PLACE A STENT. DESPITE THESE EFFORTS, THE PATIENT DIED. THE MEDTRONIC DEVICE WAS NOTED TO HAVE CONTRIBUTED TO THE ADVERSE OUTCOME. OTHER CONTRIBUTING FACTORS INCLUDED THE PATIENT'S DILATED ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THE PATIENT'S UNDERLYING CONDITION OF SEVERE AORTIC STENOSIS WHICH REQUIRED THE TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE CONTRIBUTED TO THE DEATH.Event Description: IT WAS FURTHER REPORTED THAT POCKET CULTURES WERE NEGATIVE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC DISSECTION OCCURRED WHEN THE SECOND SYSTEM WAS ADVANCING OVER THE THE AORTIC ARCH, MOVING THE FIRST VALVE THAT HAD DISLODGED. PER THE PHYSICIAN, THE SECOND DCS DID NOT CAUSE OR CONTRIBUTE TO THE DISSECTION. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VFXPLUS-29, SERIAL/LOT #: (B)(6), UBD: 16-JUL-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21841224&pc=DSQ | 2916596-2025-02149 | 2025-03-21 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Obstruction/Occlusion; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM (ER) ON (B)(6) 2025 WITH RECTAL BLEEDING. THE PATIENT WAS FOUND TO HAVE A SMALL BOWEL OBSTRUCTION. THE PATIENT BEGAN TO DECOMPENSATE IN THE ER WITH MEAN ARTERIAL PRESSURE (MAPS) IN THE 50S AND ULTIMATELY CODED BEFORE THEY WERE ABLE TO BE TRANSPORTED TO THE INTENSIVE CARE UNIT (ICU). RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS OBTAINED. THE PATIENT'S FAMILY DECIDED TO PURSUE COMFORT CARE ON (B)(6) 2025, AND THE PATIENT PASSED AWAY LATER THAT DAY. Manufacturer Narrative: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE PATIENT¿S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21841510&pc=NPT | 2025587-2025-02674 | 2025-04-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPT | EVOLUT FX DCS | Device Dislodged or Dislocated; Patient Device Interaction Problem; Central Regurgitation | Stroke/CVA; Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Hypoesthesia; Loss of consciousness; Decreased Sensitivity; Movement Disorder; Ischemia Stroke; Unspecified Nervous System Problem; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE, A TEMPORARY PACEMAKER WAS INSERTED. THE PATIENT UNDERWENT A LEFT CAROTID CUTDOWN UNDER GENERAL ANESTHETIC, AND PRE-DILATATION WAS PERFORMED WITH A 22 MILLIMETER NON-MEDTRONIC BALLOON. FOLLOWING PRE-DILATATION, THE PATIENT WENT INTO COMPLETE HEART BLOCK AND WAS SUPPORTED BY THE TEMPORARY PACEMAKER. THE PATIENT REMAINED HYPOTENSIVE WITH AORTIC INSUFFICIENCY (AI). THE DELIVERY CATHETER SYSTEM (DCS) WAS INSERTED THROUGH AN 18 FR NON-MEDTRONIC SHEATH OVER A NON-MEDTRONIC GUIDEWIRE. PRIOR TO THE FIRST DEPLOYMENT, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION, NECESSITATING CARDIOPULMONARY RESUSCITATION (CPR) AND EXTERNAL SHOCKS MULTIPLE TIMES. THE VALVE WAS INITIALLY DEPLOYED TO 80%, BUT THE PATIENT CONTINUED TO EXPERIENCE VENTRICULAR FIBRILLATION, REQUIRING FURTHER CPR AND SHOCKS. FLUOROSCOPY REVEALED VALVE DISLODGEMENT INTO THE ASCENDING AORTA, PROMPTING RECAPTURE AND REDEPLOYMENT TO 80%. AGAIN VENTRICULAR FIBRILLATION WAS NOTED AND REQUIRED MORE CPR AND SHOCKING. DESPITE STABILIZATION EFFORTS, THE VALVE DISLODGED AGAIN, REQUIRING ANOTHER RECAPTURE AND REDEPLOYMENT. ONCE THE PATIENT BEGAN TO STABILIZE, THE VALVE WAS FULLY DEPLOYED WITH FINAL DEPTH MEASUREMENTS OF 4MM NON-CORONARY CUSP (NCC) AND 8 MILLIMETER ON LEFT CORONARY CUSP (LCC), ACHIEVING COMMISSURAL ALIGNMENT WITH NO PARAVALVULAR LEAK (PVL) OR AI. THE PATIENT WAS INTUBATED WITH A TEMPORARY PACEMAKER AND SENT TO RECOVERY. DESPITE STABILIZATION, THE PATIENT BECAME NEUROLOGICALLY UNRESPONSIVE AFTER SEDATION WAS WEANED, AND A BRAIN SCAN SHOWED MULTIPLE AREAS OF INFARCT. LIFE SUPPORT WAS DISCONTINUED, AND THE PATIENT PASSED AWAY TWO DAYS POST VALVE IMPLANT. PER THE PHYSICIAN, THE VALVE DISLODGEMENT WAS LIKELY DUE TO CPR AND POSSIBLE MOVEMENT OF THE EXTERNAL SHEATH DURING THE CPR.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE, A TEMPORARY PACEMAKER WAS INSERTED. THE PATIENT UNDERWENT A LEFT CAROTID CUTDOWN UNDER GENERAL ANESTHETIC, AND PRE-DILATATION WAS PERFORMED WITH A 22 MILLIMETER NON-MEDTRONIC BALLOON. FOLLOWING PRE-DILATATION, THE PATIENT WENT INTO COMPLETE HEART BLOCK AND WAS SUPPORTED BY THE TEMPORARY PACEMAKER. THE PATIENT REMAINED HYPOTENSIVE WITH AORTIC INSUFFICIENCY (AI). THE DELIVERY CATHETER SYSTEM (DCS) WAS INSERTED THROUGH AN 18 FR NON-MEDTRONIC SHEATH OVER A NON-MEDTRONIC GUIDEWIRE. PRIOR TO THE FIRST DEPLOYMENT, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION, NECESSITATING CARDIOPULMONARY RESUSCITATION (CPR) AND EXTERNAL SHOCKS MULTIPLE TIMES. THE VALVE WAS INITIALLY DEPLOYED TO 80%, BUT THE PATIENT CONTINUED TO EXPERIENCE VENTRICULAR FIBRILLATION, REQUIRING FURTHER CPR AND SHOCKS. FLUOROSCOPY REVEALED VALVE DISLODGEMENT INTO THE ASCENDING AORTA, PROMPTING RECAPTURE AND REDEPLOYMENT TO 80%. AGAIN VENTRICULAR FIBRILLATION WAS NOTED AND REQUIRED MORE CPR AND SHOCKING. DESPITE STABILIZATION EFFORTS, THE VALVE DISLODGED AGAIN, REQUIRING ANOTHER RECAPTURE AND REDEPLOYMENT. ONCE THE PATIENT BEGAN TO STABILIZE, THE VALVE WAS FULLY DEPLOYED WITH FINAL DEPTH MEASUREMENTS OF 4MM NON-CORONARY CUSP (NCC) AND 8 MILLIMETER ON LEFT CORONARY CUSP (LCC), ACHIEVING COMMISSURAL ALIGNMENT WITH NO PARAVALVULAR LEAK (PVL) OR AI. THE PATIENT WAS INTUBATED WITH A TEMPORARY PACEMAKER AND SENT TO RECOVERY. DESPITE STABILIZATION, THE PATIENT BECAME NEUROLOGICALLY UNRESPONSIVE AFTER SEDATION WAS WEANED, AND A BRAIN SCAN SHOWED MULTIPLE AREAS OF INFARCT. LIFE SUPPORT WAS DISCONTINUED, AND THE PATIENT PASSED AWAY TWO DAYS POST VALVE IMPLANT. PER THE PHYSICIAN, THE VALVE DISLODGEMENT WAS LIKELY DUE TO CPR AND POSSIBLE MOVEMENT OF THE EXTERNAL SHEATH DURING THE CPR. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PATIENT HAD SEVERE AORTIC STENOSIS. CLARIFICATION WAS RECEIVED: THE SEVERE AI OCCURRED AFTER PRE-IMPLANT BALLOON DILATION AND PRIOR TO THE VALVE IMPLANT; THEREFORE, IT WAS CENTRAL REGURGITATION. ON THE FIRST POST-OPERATIVE DAY, THE PATIENT WAS WEANED OFF OF SEDATION. HOWEVER, UPON ASSESSMENT THE PATIENT DID NOT OBEY COMMANDS; THE PUPILS WERE EQUAL AND REACTIVE TO LIGHT, BUT SLUGGISH. THE PATIENT DID NOT HAVE VISIBLE MEANINGFUL MOTOR RESPONSES AND THERE WAS NO MOTOR RESPONSE TO DEEP CENTRAL PAIN. THE NEUROLOGICAL REPORT STATED: ¿EXTENSIVE AREA OF DIFFUSION RESTRICTION INVOLVING ALL LOBES OF THE SUPRATENTORIAL BRAIN AS WELL AS THE BILATERAL BASAL GANGLIA AND THALAMI AND THROUGHOUT THE CEREBELLUM IN KEEPING WITH EVOLVING INFARCTS. MANY OF THE DIFFUSION-WEIGHTED RESTRICTION LESIONS LOOK LIKE CORTICAL INFARCTS ESPECIALLY THE CEREBELLAR AND THALAMIC INFARCTIONS. TYPICALLY WITH HYPOXIC ISCHEMIC INJURY AND OTHER STRUCTURES SUCH AS THE DENTATE NUCLEUS AND BASAL GANGLIA ARE AFFECTED AS WELL. IT CERTAINLY POSSIBLE THAT AN ELEMENT OF THESE LESION MAY BE HYPOXIC ISCHEMIC." PER THE PHYSICIAN, THE STROKE WAS NOT RELATED TO EITHER THE VALVE OR THE DCS. IT WAS NOTED THE PATIENT'S CALCIFIED ANNULUS, WITH A CALCIUM SCORE 1800, WAS A CONTRIBUTING FACTOR TO THE STROKE. THE PHYSICIAN REPORTED THE PATIENT'S DEATH WAS RELATED TO THE PATIENT'S UNDERLYING CONDITION IN COMBINATION WITH THE SEVERE AI THAT PRESENTED AFTER THE PRE-IMPLANT DILATION. THE PHYSICIAN NOTED THE VALVE AND DCS CAN BE EXCLUDED AS CONTRIBUTING TO THE PATIENT'S DEATH. PER THE PHYSICIAN, THE PATIENT "PASSED WITHOUT DISTRESS WITH CESSATION OF LIFE SUPPORT." Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. PATIENT CODES/ANNEX E WERE ADDED CORRECTED DATA: B2. OUTCOME ATTRIBUTED TO ADVERSE EVENT. DISABILITY AND HOSPITALIZATION WERE OMITTED FROM THE INITIAL MEDWATCH REPORT. D. SUSPECT MEDICAL DEVICE FULL UDI# D10. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: MEDTRONIC PRODUCT BRAND NAME: EVOLUT FX VALVE; MANUFACTURED DATE: 2025-03-11; USE BY DATE: 2027-03-11; FULL UDI#: (B)(4) H6. PATIENT CODES/ANNEX E CODES ADDED THAT WERE NOT INCLUDED IN THE INITIAL MEDWATCH REPORT. ADDITIONAL CODES. HEALTH IMPACT/ANNEX F CODES ADDED THAT WERE NOT INCLUDED IN THE INITIAL MEDWATCH REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX VALVE; PRODUCT ID EVOLUTFX-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21841943&pc=OZD | 1220648-2025-27490 | 2025-03-19 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND SAME DAY SOME TIME POST IMPLANT THE PATIENT WAS NOTED TO HAVE HEMOLYSIS. IT WAS NOTED THE POSITION OF THE PUMP AND THE VOLUME STATUS WERE "GOOD." THE FOLLOWING DAY CARE WAS WITHDRAWN FROM THE PATIENT, AND THE PATIENT EXPIRED. ADDITIONAL DETAILS SURROUNDING THE DEMISE OUTCOME IS UNKNOWN AND THE IMPACT AS A RESULT OF THE HEMOLYSIS IS UNKNOWN AS WELL. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. CLINICAL STATES THAT HEMOLYSIS WAS OBSERVED AFTER PUMP IMPLANT WITH PUMP IN GOOD POSITION AND GOOD VOLUME. PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOGS WERE INVESTIGATED AND THERE WAS NO MOTOR CURRENT SPIKE OR POSITIONING ISSUE OBSERVED. THE DATA LOGS WERE INCONCLUSIVE. THEREFORE THE ROOT CAUSE OF THE HEMOLYSIS ISSUE WAS NOT DETERMINED SINCE THE PUMP WAS NOT RETURNED AND DATA LOGS WERE INCONCLUSIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842162&pc=DSQ | 2916596-2025-02237 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Embolism/Embolus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE PATIENT WAS WEANING FROM BYPASS, THE LEFT VENTRICULAR ASSIST DEVIC (LVAD) WAS STARTED AND THE SPEED INCREASED. THE FLOW WAS VERY LOW AT THE BEGINNING, NOT HIGHER THAN 0,4 LITER PER MINUTE (LPM) BUT DECREASED TO ZERO. IT REMAINED ZERO ALTHOUGH THE BYPASS WAS STOPPED ALREADY. THE HOSPITAL CALLED THE COMPANY FOR ANY PRESUMPTIONS, WHY THE FLOW REMAINED ZERO. PUMP PARAMETERS EXCEPT THE FLOW WERE NORMAL. PULSATILITY INDEX (PI) WAS VERY LOW, POWER WAS ALSO LOW AND WAS WHAT WAS EXPECTED IN THIS ZERO-FLOW MOMENT. ECHOCARDIOGRAM COULD NOT SHOW ANY INFLOW STREAM INTO THE INFLOW CANNULA; THE OUTFLOW IN THE AORTA COULD BE VISUALIZED DUE TO ARTEFACTS AND TURBULENCES FROM THE 2 BYPASSES. THE LEFT VENTRICLE (LV) SEEMED TO BE EMPTY BUT NOT COLLAPSED COMPLETELY. THE RIGHT VENTRICULAR FUNCTION WAS ALREADY BAD PRIOR LVAD AND REMAINS IN A BAD CONDITION AFTERWARDS. TO EXCLUDE ANY OCCLUSIONS OF THE INFLOW OR AIR LOCKS THE PUMP WAS AGAIN DISCONNECTED FROM THE RING AND FLUSHED RETROGRADE PLUS RINSED WITH WATER. NO VISUAL ISSUES COULD BE SEEN, AND THE PUMP WAS AGAIN CONNECTED TO THE PATIENT. THE FLOW REMAINED ZERO. ACCORDING TO THE VAD COORDINATOR AFTER EVERYTHING WAS DOUBLE CHECKED THEY DECIDED TO GO FOR A PUMP EXCHANGE DIRECTLY TO EXCLUDE ANY TECHNICAL ISSUES WITH THE DEVICE. AFTER THE EXCHANGE, THERE WAS 5 MINUTES FLOW ON THE PUMP BUT THE PICTURE FROM BEFORE WAS RECURRENT AFTERWARDS ALSO WITH THE NEW PUMP, AGAIN ZERO FLOW. ACCORDING TO THE ANESTHETIST THE RIGHT VENTRICLE WAS DOING NOTHING AT THIS MOMENT AND THE SURGEONS DECIDED TO GO FOR A TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) (CARDIOHELP) TO SUPPORT THE RIGHT VENTRICULAR FAILURE (RVF). AFTER STARTING THE RVAD, THE FLOW OF THE LVAD INCREASED DIRECTLY AND BOTH DEVICES WERE ADJUSTED ACCORDINGLY AND ENDED UP WITH NEARLY 4 LPM ON BOTH SIDES. THE HOSPITAL ALSO CHECKED THE EXPLANTED PUMP IN A BUCKET OF WATER WITH A TEST CIRCUIT AND THE PUMP WAS BEHAVING AS EXPECTED, NO ISSUE OCCURRED, NO TECHNICAL ALARM OCCURRED AND ALSO NO THROMBUS FORMATIONS WERE WASHED OUT. THEREFORE, THE CLINICIANS SUSPECTED THAT THE PATIENT DEVELOPED A FULMINANT RVF WHICH WAS NOT DIRECTLY DIAGNOSED, AND THE LV WAS COLLAPSED COMPLETELY AND OCCLUDED THE INFLOW ITSELF. THE HOSPITAL WAS DISCUSSING IF THEY WOULD RETURN THE EXPLANTED PUMP ALTHOUGH THEY DO NOT ASSUME A DEVICE MALFUNCTION. LOG FILES ALSO CONFIRMED THAT THE PUMP/ROTOR DID NOT SHOW ANY MALFUNCTIONING OR ISSUES WHICH LEAD TO THIS SITUATION. THE PATIENT REMAINED RECOVERING ON THE INTENSIVE CARE UNIT (ICU).Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY DUE TO EMBOLIC STROKE ON 28MAR2025. THE PATIENT SUFFERED FROM A SEVERE EMBOLIC STROKE (COMPUTED TOMOGRAPHY SCAN) 2 DAYS AFTER AN ULTIMA RATIO IMPLANTATION FROM ECMO (NO ABBOTT DEVICE). ROOT CAUSE WAS NOT CONFIRMED AS ALSO NO AUTOPSY WAS PERFORMED. THE DEVICE OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT UNDERWENT A PUMP EXCHANGE ON THE DAY OF IMPLANT AFTER EXPERIENCING FLOW ISSUES ON THE FIRST IMPLANTED PUMP. AFTER THE EXCHANGE, THERE WAS 5 MINUTES OF FLOW ON THE PUMP UNTIL THE SAME FLOW ISSUES WERE OBSERVED AGAIN WITH FLOW DROPPING TO ZERO. ACCORDING TO THE ANESTHETIST, THE RIGHT VENTRICLE WAS DOING NOTHING AT THIS TIME AND THE SURGEONS DECIDED TO TRY ADDING A TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WITH EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) TO SUPPORT THE RIGHT VENTRICULAR FAILURE (RVF). AFTER STARTING THE RVAD, THE FLOW OF THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) INCREASED DIRECTLY, AND BOTH DEVICES WERE ADJUSTED ACCORDINGLY LEADING TO SUFFICIENT FLOWS ON BOTH SIDES. THE CLINICIANS SUSPECTED THAT THE PATIENT DEVELOPED A FULMINANT RVF WHICH WAS NOT DIRECTLY DIAGNOSED, AND THE LEFT VENTRICLE (LV) HAD COLLAPSED COMPLETELY AND OCCLUDED THE INFLOW ITSELF. THE RIGHT HEART FAILURE EXISTED PRE-LVAD IMPLANT AND WAS NOTED TO BE MILD TO MODERATE UNDER THE IMPELLA DEVICE. A DEVICE RELATED ISSUE WAS NOT SUSPECTED TO HAVE CONTRIBUTED TO THE RIGHT HEART FAILURE. THE PATIENT REMAINED RECOVERING IN THE INTENSIVE CARE UNIT (ICU). IT WAS LATER REPORTED THAT THE PATIENT ULTIMATELY EXPIRED DUE TO AN EMBOLIC STROKE. THE PATIENT SUFFERED FROM A SEVERE EMBOLIC STROKE (NOTED ON CARDIAC COMPUTED TOMOGRAPHY (CCT)) 2 DAYS AFTER AN ULTIMA RATIO IMPLANTATION FROM ECMO. A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE CONFIRMED, AND NO AUTOPSY WAS PERFORMED. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. G, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, RIGHT HEART FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER MLP-051915, AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PATIENT REMAINS ONGOING ON HEARTMATE 3 LVAS, SERIAL NUMBER MLP-051915, WITH NO FURTHER RELATED EVENTS REPORTED AT THIS TIME. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. G, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR MLP-051915 WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842171&pc=NPT | 2025587-2025-02681 | 2024-07-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: UPDATED: D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842172&pc=NPT | 2025587-2025-02682 | 2025-03-19 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPT | EVOLUT FX PLUS VALVE | Difficult to Advance; Patient Device Interaction Problem | Stroke/CVA; Dysphasia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH CALCIFIED ANATOMY, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS IT WAS DIFFICULT TO CROSS THE AORTIC VALVE. SUBSEQUENTLY, THE VALVE WAS IMPLANTED IN THE PATIENT AND THE PROCEDURE WAS COMPLETED. THE FOLLOWING NIGHT, THE PATIENT EXPERIENCED A CEREBRAL VASCULAR ACCIDENT AND PASSED AWAY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT DEVELOPED DIMINISHED MENTATION AND EXPRESSIVE APHASIA AS THEY WERE WAKING UP FROM ANESTHESIA WITHIN A FEW HOURS OF THE PROCEDURE. SUBSEQUENTLY, A MULTIFOCAL ISCHEMIC STROKE WAS CONFIRMED. PER THE PHYSICIAN, THE STROKE SEEMED EMBOLIC AND RELATED TO THE PROCEDURE. IT WAS REPORTED THAT THE STROKE WAS LIKELY RELATED TO THE PATIENT'S EXTENSIVE ASCENDING AORTA/AORTIC ARCH CALCIFICATION. THE PATIENT THEN HAD RAPID NEUROLOGICAL DECLINE AND IT WAS DECIDED TO SWITCH THE PATIENT OVER TO COMFORT MEASURES. LATER THAT NIGHT, THE PATIENT DIED. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329, PRODUCT ID: D-EVOLUTFX-2329, SERIAL/LOT: UNKNOWN, USE BY DATE: UNKNOWN, UDI: UNKNOWN. UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842218&pc=OZD | 1220648-2025-27491 | 2025-03-16 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA 5.5 | Material Split, Cut or Torn | Atrial Fibrillation; Atrial Flutter; Low Blood Pressure/ Hypotension | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THERE WAS PRODUCT DAMAGE. THE STERILE SLEEVE RIPPED AND DETACHED ON THE DISTAL SEGMENT OF THE CATHETER BY THE RED IMPELLA PLUG. THE MIDDLE AND PROXIMAL SEGMENT OF THE STERILE SLEEVE WAS STILL ATTACHED AND IN GOOD CONDITION. THE NURSE CLEANED WITH BETADINE AND REINFORCED WITH TEGADERMS. IT WAS LATER REPORTED THAT THE PATIENT DECOMPENSATED AND CODED. THE IMPELLA POSITION WAS CHECKED DUE TO CONCERNS OF PLACEMENT SIGNAL. ADDITIONALLY, THE PATIENT EXPERIENCED ATRIAL FIBRILLATION AND ATRIAL FLUTTER FOR WHICH THE P-LEVEL WAS ADJUSTED FOR. HOWEVER, THE PATIENT REMAINED HYPOTENSIVE AND RESUSCITATION WAS CONTINUED. THE PATIENT SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF PRODUCT DAMAGE AND VF/VT/AF/AT HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION, AND A DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE. THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT) DURING PUMP USE. ADDITIONALLY, THE SLEEVE HAD RIPPED AND DETACHED, BUT THIS WAS MANAGED WITH TEGADERM. THE CAUSE OF THE STERILE SLEEVE DAMAGE ISSUE WAS MOST LIKELY USE RELATED, BASED ON GIVEN CLINICAL DETAILS. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842281&pc=OZD | 1220648-2025-27492 | 2025-03-23 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT (PT) CRITICALLY ILL, IN CARDIOGENIC SHOCK, REPORTEDLY DECOMPENSATED FURTHER ONCE ARRIVING TO THE ICU. DECISION WAS MADE FOR IMPELLA CP. SUCCESSFUL IMPELLA IMPLANT. INITIAL INTRINSIC PULSATILITY VERY WEAK, HOWEVER DID IMPROVE AFTER INITIATING SUPPORT. PT¿S PRESSURES ALSO IMPROVED. DURING SUPPORT IT WAS REPORTED THAT THE PATIENT WAS MECHANICALLY VENTILATED. THE PT'S CLOTTING FACTORS HAD BEEN MARKEDLY ABNORMAL AND NURSING STAFF STATED HE WAS ¿BLEEDING FROM EVERY LINE.¿ WITH THE DIC, PATIENT REQUIRED PACKED RED BLOOD CELLS, CRYO, AND FRESH FROZEN PLASMA. THIS WAS NOTED TO NOT BE IMPELLA RELATED. ADDITIONALLY, THE PATIENT STATUS HAD IMPROVED SLIGHTLY; HOWEVER, STAFF WERE CONCERNED ABOUT HIS NEUROLOGIC STATUS-GCS OF 5 (SUSPECT POSSIBLE CVA). IT WAS NOTED THAT IN THE LAST 24 HOURS, PATIENT HAD RECEIVED 17 UNITS OF BLOOD PRODUCTS. THEY STARTED HEMODIALYSIS AS WELL. THE NURSE NOTED SIGNIFICANT BLEEDING FROM THE IMPELLA GROIN SITE. AN INTENSIVIST ADJUSTED THE ANGLE OF INSERTION SLIGHTLY AND THEN RESECURED IT WITH SUTURES. THIS DID APPEAR TO ALLEVIATE THE BLEEDING ISSUE. ULTIMATELY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED, AN INTENSIVIST ADJUSTED THE ANGLE OF INSERTION SLIGHTLY AND THEN RESECURED IT WITH SUTURES TO ACHIEVE HEMOSTASIS. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842665&pc=PCU | 3005099803-2025-01671 | 2018-07-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-15 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Use of Device Problem; Positioning Problem | Peritonitis |  |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE: "EFFECTIVENESS OF ENDOSCOPIC ULTRASOUND (EUS)-GUIDED CHOLEDOCHODUODENOSTOMY VS. EUS-GUIDED GALLBLADDER DRAINAGE FOR JAUNDICE IN PATIENTS WITH MALIGNANT DISTAL BILIARY OBSTRUCTION AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY: RETROSPECTIVE, MULTICENTER STUDY (GALLBLADEUS STUDY)," BY ANTOINE DEBOURDEAU, ET AL. PER THE ARTICLE, A MULTICENTER RETROSPECTIVE STUDY INCLUDED 78 PATIENTS WITH OBSTRUCTIVE JAUNDICE SECONDARY TO MALIGNANT DISTAL BILIARY OBSTRUCTION (MDBO) WHO UNDERWENT ENDOSCOPIC ULTRASOUND-GUIDED GALLBLADDER DRAINAGE (EUS-GBD) (41 PATIENTS) OR ENDOSCOPIC ULTRASOUND-GUIDED CHOLEDOCHODUODENOSTOMY (EUS-CDS) (37 PATIENTS) WITH LUMEN-APPOSING METAL STENTS AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) BETWEEN JULY 2018 AND JULY 2022. A FATALITY OCCURRED DUE TO THE LUMEN-APPOSING METAL STENT (LAMS) ENTERING THE PERITONEUM, WHICH COULD NOT BE MANAGED ENDOSCOPICALLY. PERCUTANEOUS DRAINAGE AND EXTERNAL DRAIN PLACEMENT DID NOT IMPROVE THE RESULTING SEPTIC CONDITION FROM BILIARY PERITONITIS. Event Description: BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE: "EFFECTIVENESS OF ENDOSCOPIC ULTRASOUND (EUS)-GUIDED CHOLEDOCHODUODENOSTOMY VS. EUS-GUIDED GALLBLADDER DRAINAGE FOR JAUNDICE IN PATIENTS WITH MALIGNANT DISTAL BILIARY OBSTRUCTION AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY: RETROSPECTIVE, MULTICENTER STUDY (GALLBLADEUS STUDY)," BY ANTOINE DEBOURDEAU, ET AL. PER THE ARTICLE, A MULTICENTER RETROSPECTIVE STUDY INCLUDED 78 PATIENTS WITH OBSTRUCTIVE JAUNDICE SECONDARY TO MALIGNANT DISTAL BILIARY OBSTRUCTION (MDBO) WHO UNDERWENT ENDOSCOPIC ULTRASOUND-GUIDED GALLBLADDER DRAINAGE (EUS-GBD) (41 PATIENTS) OR ENDOSCOPIC ULTRASOUND-GUIDED CHOLEDOCHODUODENOSTOMY (EUS-CDS) (37 PATIENTS) WITH LUMEN-APPOSING METAL STENTS AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) BETWEEN JULY 2018 AND JULY 2022. A FATALITY OCCURRED DUE TO THE LUMEN-APPOSING METAL STENT (LAMS) ENTERING THE PERITONEUM, WHICH COULD NOT BE MANAGED ENDOSCOPICALLY. PERCUTANEOUS DRAINAGE AND EXTERNAL DRAIN PLACEMENT DID NOT IMPROVE THE RESULTING SEPTIC CONDITION FROM BILIARY PERITONITIS. PLEASE SEE THE REFERENCED ARTICLE FOR FULL DETAILS AND FULL LISTING OF PHYSICIANS AND HEALTHCARE FACILITIES. NOTE: IT WAS REPORTED THAT THE AXIOS STENT WAS IMPLANTED DURING AN ENDOSCOPIC ULTRASOUND (EUS)-GUIDED CHOLEDOCHODUODENOSTOMY PROCEDURE. HOWEVER, PER THE AXIOS STENT AND ELECTROCAUTERY-ENHANCED DELIVERY SYSTEM INSTRUCTIONS FOR USE (IFU), THE STENT IS INDICATED FOR USE TO FACILITATE TRANSGASTRIC OR TRANSDUODENAL ENDOSCOPIC DRAINAGE OF A PANCREATIC PSEUDOCYST OR WALLED-OFF NECROSIS WITH GREATER THAN OR EQUAL TO 70% FLUID CONTENT, GALLBLADDER IN PATIENTS WITH ACUTE CHOLECYSTITIS WHO ARE AT HIGH RISK OR UNSUITABLE FOR SURGERY, AND THE BILE DUCT AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) IN PATIENTS WITH BILIARY OBSTRUCTION DUE TO A MALIGNANT STRICTURE. THE DEVICE IS NOT INDICATED TO BE USED FOR AN ENDOSCOPIC ULTRASOUND (EUS)-GUIDED CHOLEDOCHODUODENOSTOMY PROCEDURE. Manufacturer Narrative: BLOCK B3: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED DATE OF PATIENT'S DEATH OF (B)(6) 2018, WAS CHOSEN AS THE BEST ESTIMATE BASED ON 30 DAYS FROM THE RETROSPECTIVE STUDY START DATE OF (B)(6) 2018. BLOCK D4, H4: DETAILED PRODUCT INFORMATION, INCLUDING DEVICE AVAILABILITY FOR EVALUATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1 (INITIAL REPORTER ADDRESS 1): THE INITIAL REPORTER ADDRESS IS (B)(6). BLOCK G2: LITERATURE SOURCE: ANTOINE DEBOURDEAU; JULES DANIEL; LUDOVIC CAILLO; ERIC ASSENAT; PHILIPPE POUDEROUX; JEAN-FRANCOIS BOURGAUX; MARTIN BERTRAND; THOMAS BARDOL. "EFFECTIVENESS OF ENDOSCOPIC ULTRASOUND (EUS)-GUIDED CHOLEDOCHODUODENOSTOMY VS. EUS-GUIDED GALLBLADDER DRAINAGE FOR JAUNDICE IN PATIENTS WITH MALIGNANT DISTAL BILIARY OBSTRUCTION AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY: RETROSPECTIVE, MULTICENTER STUDY (GALLBLADEUS STUDY)." DIGESTIVE ENDOSCOPY 2025; 37: 103-114. BLOCK H6: IMDRF DEVICE CODE A1502 CAPTURES THE REPORTABLE EVENT OF LAMS ENTERING THE PERITONEUM. IMDRF PATIENT CODE E1024 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF SEPTIC CONDITION FROM BILIARY PERITONITIS. IMDRF IMPACT CODE F2301 CAPTURES THE PERCUTANEOUS DRAINAGE AND EXTERNAL DRAIN PLACEMENT. IMDRF IMPACT CODE F02 CAPTURES THE PATIENT DEATH. Manufacturer Narrative: BLOCKS B5 AND H6 (DEVICE CODES) HAVE BEEN CORRECTED. BLOCK B3: THE EXACT DATE OF THE PATIENT'S DEATH IS UNKNOWN. THE PROVIDED DATE OF PATIENT'S DEATH OF (B)(6) 2018, WAS CHOSEN AS THE BEST ESTIMATE BASED ON 30 DAYS FROM THE RETROSPECTIVE STUDY START DATE OF JULY 2018. BLOCK D4, H4: DETAILED PRODUCT INFORMATION, INCLUDING DEVICE AVAILABILITY FOR EVALUATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1 (INITIAL REPORTER ADDRESS 1): THE INITIAL REPORTER ADDRESS (B)(6). BLOCK G2: LITERATURE SOURCE: ANTOINE DEBOURDEAU; JULES DANIEL; LUDOVIC CAILLO; ERIC ASSENAT; PHILIPPE POUDEROUX; JEAN-FRANCOIS BOURGAUX; MARTIN BERTRAND; THOMAS BARDOL. "EFFECTIVENESS OF ENDOSCOPIC ULTRASOUND (EUS)-GUIDED CHOLEDOCHODUODENOSTOMY VS. EUS-GUIDED GALLBLADDER DRAINAGE FOR JAUNDICE IN PATIENTS WITH MALIGNANT DISTAL BILIARY OBSTRUCTION AFTER FAILED ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY: RETROSPECTIVE, MULTICENTER STUDY (GALLBLADEUS STUDY)." DIGESTIVE ENDOSCOPY 2025; 37: 103-114. BLOCK H6: IMDRF DEVICE CODE A1502 CAPTURES THE REPORTABLE EVENT OF LAMS ENTERING THE PERITONEUM. IMDRF PATIENT CODE E1024 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF SEPTIC CONDITION FROM BILIARY PERITONITIS. IMDRF IMPACT CODE F2301 CAPTURES THE PERCUTANEOUS DRAINAGE AND EXTERNAL DRAIN PLACEMENT. IMDRF IMPACT CODE F02 CAPTURES THE PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842764&pc=DSQ | 2916596-2025-02135 | 2025-03-24 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Fall; Obstruction/Occlusion; Ischemia Stroke; Unspecified Mental, Emotional or Behavioural Problem; Thrombosis/Thrombus | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT (ED) WITH ALTERED MENTAL STATUS, LEFT HEMINEGLECT, AND MILD FACIAL DROOP IN THE SETTING OF KNOWN PUMP THROMBOSIS. THE PATIENT WAS IN HOSPICE CARE PRIOR TO GRADUATION FROM THEIR SERVICE TWO WEEKS PRIOR. THERE WERE NO UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS ADMITTED IN (B)(6) 2024 WITH STROKE SYMPTOMS AND WAS FOUND TO HAVE A LEFT SUBDURAL HEMATOMA. THE PATIENT ALSO SHOWED SMALL INCREASES IN THE PUMP FLOW AND POWER, BUT NO SPEED CHANGES HAD BEEN MADE. THE THROMBUS WAS SUSPECTED TO HAVE DEVELOPED AROUND THAT TIME. THE PATIENT WAS TAKEN OFF ANTICOAGULATION MEDICATIONS WARFARIN AND ACETYLSALICYLIC ACID IN (B)(6) 2024 DUE TO RECURRING FALLS AND HEAD BLEEDS. THE PATIENT HAD COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) OF THE HEAD AND NECK DONE ON (B)(6) 2025 WHICH SHOWED SUBACUTE TO CHRONIC APPEARING INFRACT OF THE MEDIAL LEFT PARIETAL LOBE, WHICH WAS NEW FROM (B)(6) 2024. THERE WAS ALSO OCCLUSION OF THE PROXIMAL M2 BRANCH. AFTER CONVERSATIONS WITH THE HEART FAILURE TEAM, NEUROLOGY, PALLIATIVE CARE, AND ED STAFF, NO TREATMENTS WERE PERFORMED. THE PLAN OF CARE WAS ONGOING. THE PATIENT'S MENTAL STATUS VARIED FROM AGITATION REQUIRING MEDICATIONS TO HELP CALM THEM TO MINIMAL ENGAGEMENT TO THEIR SURROUNDINGS. THEIR SPOUSE WAS TALKING WITH THE TEAM DAILY TO DETERMINE THE NEXT STEPS FOR THE PATIENT. THE OPTIONS RANGED FROM HOSPICE WITH THE DEACTIVATION OF THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) TO PERUSING REHABILITATION EFFORTS DEPENDING ON IF THE PATIENT WAS TO SHOW SOME IMPROVEMENT IN THEIR CONDITION. THE PATIENT BEING TAKEN OFF ANTICOAGULATION IN (B)(6) 2024 WAS REPORTED UNDER MFR # 2916596-2024-52794. THE PATIENT'S ADMISSION IN (B)(6) 2024 WAS REPORTED UNDER MFR # 2916596-2024-52767.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 SECONDARY TO CEREBROVASCULAR ACCIDENT (CVA), PUMP THROMBOSIS, AND WAS NOT A SURGICAL CANDIDATE. Manufacturer Narrative: CORRECTED DATA: SECTION A2: PATIENT AGE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED PUMP THROMBOSIS COULD NOT BE CONFIRMED AS NO PRODUCT WAS RETURNED FOR EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE (HM) II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILE CAPTURED THE DATA FROM (B)(6) 2025 THROUGH (B)(6) 2025. THE SYSTEM APPEARED TO BE OPERATING AS INTENDED AT THE FIXED SPEED, AND THERE WERE NO NOTABLE ALARMS ACTIVE IN THE LOG FILE. REVIEW OF THE SUBMITTED COMPUTED TOMOGRAPHY (CT) SCAN VIDEOS COULD NOT CONFIRM THE REPORTED THE REPORTED STROKE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS DEVICE THROMBOSIS, STROKE, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 AND SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, OUTLINE INDICATIONS OF PUMP THROMBOSIS AND HOW TO RESPOND TO SUCH EVENTS. ADDITIONALLY, SECTION 6, UNDER ¿ANTICOAGULATION¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D4: DEVICE EXPIRATION DATE CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS UPDATED WITH ALL OF THE CURRENT INFORMATION AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842771&pc=DSQ | 2916596-2025-02055 | 2025-03-17 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension; Cardiovascular Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM COMPLICATIONS OF HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED, THE PUMP WAS NOT EXPLANTED. LOG FILES WERE SUBMITTED FOR ANALYSIS. THE EVENT LOG CAPTURED LOW FLOW ALARMS ON (B)(6) 2025 FROM 04:26 UNTIL 05:18. POWER CABLE DISCONNECT ALARMS WERE ALSO NOTED STARTING AT 04:50 TO 04:51. THE DRIVELINE WAS THEN DISCONNECTED AT 05:18. THERE WAS NO PUMP INTERRUPTION DURING THESE EVENTS AS THE SYSTEM CONTROLLER¿S EMERGENCY BACKUP BATTERY FUNCTIONED AS INTENDED. THE ALARMS RESOLVED UPON CONNECTING TO THE POWER MODULE. THERE WERE NO OTHER NOTABLE ALARM CONDITIONS OR PARAMETER CHANGES. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6 - MEDICAL DEVICE PROBLEM CODE: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW HAZARD ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 14MAR2025 THROUGH 17MAR2025. CONTINUOUS AND TRANSIENT LOW FLOW HAZARD ALARMS WERE CAPTURED, WHICH SOME WERE ASSOCIATED WITH ELEVATED PUISATILITY INDEX (PI) VALUES. A DRIVELINE DISCONNECT ALARM WAS CAPTURED, CONSISTENT WITH THE REMOVAL OF DEVICE SUPPORT FOLLOWING THE PATIENT'S DEATH. THE PUMP OPERATED AT THE SET SPEED THROUGHOUT THE ENTIRETY OF THE LOG FILE WHILE THE DRIVELINE WAS CONNECTED. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO COMPLICATIONS OF HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING HYPERTENSION, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW AND PI. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS, SUCH AS HYPERTENSION, CAN RESULT IN LOW FLOW. SECTIONS 1 AND 4 OF THE IFU ALSO STATE THAT THE PI CALCULATION REPRESENTS CARDIAC PULSATILITY, WITH VALUES TYPICALLY RANGING FROM 1 TO 10. GENERALLY, THE MAGNITUDE OF THE PI VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E. THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT", STATES THAT POST-IMPLANTATION HYPERTENSION MAY BE TREATED AT THE DISCRETION OF THE ATTENDING PHYSICIAN, AND ANY THERAPY THAT CONSISTENTLY MAINTAINS MEAN ARTERIAL BLOOD PRESSURE LESS THAN 90 MMHG SHOULD BE CONSIDERED ADEQUATE. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21842801&pc=FKX | 0008030665-2025-00865 | 2025-04-10 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-15 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT¿S SON HAD ARISEN TO USE THE RESTROOM AT APPROXIMATELY 1:30 AM, AND NOTED HIS MOTHER WAS REPORTEDLY SLEEPING COMFORTABLY. HOWEVER, AT APPROXIMATELY 6:30 AM THE PATIENT¿S SON DISCOVERED THE PATIENT HAD EXPIRED WHEN HE WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 THROUGH (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HER SIGNIFICANT CARDIAC HISTORY (CORONARY ARTERY DISEASE, STENOSED AORTIC VALVE NEEDING REPLACEMENT). THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK.Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT¿S SON HAD ARISEN TO USE THE RESTROOM AT APPROXIMATELY 1:30 AM, AND NOTED HIS MOTHER WAS REPORTEDLY SLEEPING COMFORTABLY. HOWEVER, AT APPROXIMATELY 6:30 AM THE PATIENT¿S SON DISCOVERED THE PATIENT HAD EXPIRED WHEN HE WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 THROUGH (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HER SIGNIFICANT CARDIAC HISTORY (CORONARY ARTERY DISEASE, STENOSED AORTIC VALVE NEEDING REPLACEMENT). THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE PDRN REPORTED THE CAUSE OF THE SERIOUS ADVERSE EVENTS WAS THE PATIENT¿S SIGNIFICANT CARDIAC COMORBIDITIES. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. THE PDRN SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR NO SHOWED SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY ON THE BOTTOM COVER. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21843079&pc=OZD | 1220648-2025-27495 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA CP | Difficult to Remove | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. ONCE IMPELLA WAS EXPLANTED, REMOVAL ISSUES WERE NOTED ONE OF THE PERCLOSE SUTURES RIPPED. THE PHYSICIAN DECIDED TO PLACE ANGIOSEAL AND ONCE ANGIOSEAL WAS DEPLOYED, GROIN BLEEDING DID NOT CEASE. FEMORAL ANGIOGRAM WAS PERFORMED AND A STENT WAS DEPLOYED, BUT DID NOT RESOLVE ISSUE. THE PATIENT WAS TAKEN TO OPERATING ROOM, WHERE HE CODED. A FEM-FEM BYPASS WAS PERFORMED. THE PATIENT WAS THEN TAKEN BACK TO THE ICU AND ULTIMATELY EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND REMOVAL ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. REMOVAL ISSUES: PUMP WAS NOT RETURNED. DATA LOGS ARE NOT APPLICABLE AS THE INCIDENT TOOK PLACE AFTER IMPELLA REMOVAL. THE CAUSE OF THE REMOVAL ISSUE WAS USE RELATED DUE TO THE TEARING OF THE PERCLOSE SUTURE THAT WAS PLACED FOLLOWING THE PROCEDURE WHERE THE IMPELLA WAS REMOVED. ACCESS SITE BLEEDING: PUMP WAS NOT RETURNED. DATA LOGS ARE NOT APPLICABLE AS THE INCIDENT TOOK PLACE AFTER IMPELLA REMOVAL. THE CAUSE OF THE ACCESS SITE BLEEDING WAS USE RELATED DUE TO THE TEARING OF THE PERCLOSE SUTURE THAT WAS PLACED FOLLOWING THE PROCEDURE WHERE THE IMPELLA WAS REMOVED. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27495. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿AVOID MANUAL COMPRESSION OF THE INLET AND OUTLET AREAS OF THE CANNULA ASSEMBLY.¿ SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS: ¿HANDLE WITH CARE. THE IMPELLA CATHETER CAN BE DAMAGED DURING REMOVAL FROM PACKAGING, PREPARATION, INSERTION, AND REMOVAL. DO NOT BEND, PULL, OR PLACE EXCESS PRESSURE ON THE CATHETER OR MECHANICAL COMPONENTS AT ANY TIME.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21843677&pc=DSQ | 2916596-2025-02185 | 2024-04-14 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON HOSPICE CARE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HM 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21843681&pc=DSQ | 2916596-2025-02303 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO SEPSIS ON (B)(6) 2025. NO PRODUCTS WERE TO BE RETURNED. THE PUMP WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED IN THE HOSPITAL. Manufacturer Narrative: SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21843698&pc=DSQ | 2916596-2025-02216 | 2025-03-23 04:00 | Death | THORATEC CORPORATION | 2025-04-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THIS WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21844240&pc=NPT | 2025587-2025-02693 | 2025-03-21 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO CALCIFICATION. FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, A POST-IMPLANT BAV USING A 20 MILLIMETER (MM) NON-MEDTRONIC BALLOON WAS PERFORMED DUE TO MILD-MODERATE PARAVALVULAR LEAK (PVL); HOWEVER, THE PATIENT'S BLOOD PRESSURE DID NOT FULLY RECOVER FOLLOWING THE POST-IMPLANT BAV, AND MEASURES WERE TAKEN TO RECOVER THE BLOOD PRESSURE. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, BUT THE PATIENT DID NOT RECOVER. PRIOR TO ENDING THE RECOVERY EFFORTS, A POST-IMPLANT ANGIOGRAPHY WAS PERFORMED THAT SHOWED NO PERICARDIAL EFFUSION AND CORONARIES WERE PERFUSED. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO CALCIFICATION. FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, A POST-IMPLANT BAV USING A 20 MILLIMETER (MM) NON-MEDTRONIC (TRUE) BALLOON WAS PERFORMED DUE TO MILD-MODERATE PARAVALVULAR LEAK (PVL); HOWEVER, THE PATIENT'S BLOOD PRESSURE DID NOT FULLY RECOVER FOLLOWING THE POST-IMPLANT BAV, AND MEASURES WERE TAKEN TO RECOVER THE BLOOD PRESSURE. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, BUT THE PATIENT DID NOT RECOVER. PRIOR TO ENDING THE RECOVERY EFFORTS, A POST-IMPLANT ANGIOGRAPHY WAS PERFORMED THAT SHOWED NO PERICARDIAL EFFUSION AND CORONARIES WERE PERFUSED. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE POST-IMPLANT BALLOON DILATION COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH BUT THE DELIVERY CATHETER SYSTEM (DCS) COULD BE EXCLUDED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT TYPE: {0195-HEARTVALVES}; PRODUCT ID {NON-MEDTRONIC POST-IMPLANT BAV}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845068&pc=IYN | 3019216-2025-000176 | 2025-03-18 04:00 | Death | PHILIPS ULTRASOUND, INC | 2025-04-15 04:00 | IYN | EPIQ CVX | Application Program Freezes, Becomes Nonfunctional | Stroke/CVA | K132304 |  | 1 | Event Description: IT WAS REPORTED THAT THE EPIQ CVX ULTRASOUND SYSTEM DISPLAYED AN ERROR CODE DURING A TEE (TRANSESOPHAGEAL ECHOCARDIOGRAM) EXAMINATION ON A STROKE PATIENT. THE ISSUE REMAINED AFTER THE SYSTEM WAS RESTARTED TWICE. THE PROCEDURE WAS ABLE TO BE COMPLETED AFTER AN EXCHANGE OF THE ULTRASOUND SYSTEM. THE PATIENT ULTIMATELY EXPIRED; HOWEVER, THERE IS CURRENTLY NO EVIDENCE TO SUGGEST THE PATIENT EXPIRED AS A RESULT OF THE DEVICE MALFUNCTION AND ADDITIONAL INFORMATION IS BEING OBTAINED. A PHILIPS SERVICE ENGINEER REPLACED THE SYSTEM¿S ACQUISITION MODULE TO ADDRESS THE CUSTOMER¿S IMMEDIATE CONCERN. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845116&pc=NPV | 2025587-2025-02697 | 2025-02-04 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-15 04:00 | NPV | MELODY TRANSCATHETER PULMONARY VALVE | Insufficient Information; Patient Device Interaction Problem; Device Stenosis | Hemorrhage/Bleeding; High Blood Pressure/ Hypertension; Pulmonary Valve Stenosis; Unspecified Heart Problem | P140017 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A 33-YEAR-OLD MALE PATIENT WITH SEVERE PULMONARY STENOSIS OF A PREVIOUSLY IMPLANTED HOMOGRAFT WITH A MELODY BIOPROSTHETIC VALVE. THE PATIENT PRESENTED WITH RESPIRATORY FAILURE WHICH REQUIRED EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT AS WELL AS SEVERE PULMONARY HYPERTENSION WITH SEVERE RIGHT VENTRICULAR DYSFUNCTION WHICH REQUIRED HEMODIAFILTRATION. DUE TO THE PATIENT¿S CONDITION, HE WAS INCLUDED ON THE HEART-LUNG TRANSPLANT LIST. ON DAY 5 THE PATIENT WAS SWITCHED TO A NON-MEDTRONIC DOUBLE-LUMEN CANNULA FOR RIGHT HEART SUPPORT WITH AN OXYGENATOR AS A BRIDGE THERAPY TO HEART-LUNG TRANSPLANT. HEMODIAFILTRATION WAS DISCONTINUED ONE WEEK LATER. DURING REHABILITATION WHILE ON THE TRANSPLANT WAIT LIST, THE PATIENT EXPERIENCED UNCONTROLLABLE GASTRO-INTESTINAL BLEEDING AND DIED TWO WEEKS LATER. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO THIS PATIENT OR MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: CARLOS DOMINGUEZ-MASSA, ET AL. PROTEKDUO CANNULA THROUGH MELODY TRANSCATHETER PULMONARY VALVE PROSTHESIS BRIDGE TO HEART-LUNG TRANSPLANTATION. MEDICINA INTENSIVA FEB 4:502119 2025. 10.1016/J.MEDINE.2025.502119. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845367&pc=OZD | 1220648-2025-27509 | 2025-03-19 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA CP | Positioning Failure | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PUMP FELL OUT THE LEFT VENTRICLE. STAFF WERE UNABLE TO CONFIRM THE POSITION OF THE PIGTAIL, AND THE PUMP WAS EXPLANTED AND EXTRACORPOREAL MEMBRANE OXYGENATION WAS IMPLANTED. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. ACCORDING TO THE CLINICAL REPORT, THE IMPELLA WAS PULLED INTO THE AORTA DURING THE PATIENT'S TRANSFER FROM THE TABLE TO THE BED. THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY RELATED TO USE, AS THE PUMP POSITION WAS COMPROMISED DURING THE PATIENT'S TRANSFER. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H6 CODES 4582 AND 4628 WERE REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT 1220648-2025-27509. NEW CODES WERE ADDED TO HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845453&pc=OZD | 1220648-2025-27512 | 2025-03-20 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE LEFT ACCESS SIDE WAS USED FOR IMPELLA CP PLACEMENT AND PCI PROCEDURE. AT THE POST 14FR SHEATH EXPLANT THE FEMORAL VESSEL WAS PERFORATED REQUIRING COVERED STENT AND/OR BALLOON TAMPONADE AND THE PATIENT EXPIRED IN THE PROCEDURAL AREA. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845476&pc=OZD | 1220648-2025-27515 | 2025-03-20 04:00 | Death | ABIOMED, INC. | 2025-04-15 04:00 | OZD | IMPELLA CP | Positioning Failure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS A SUCTION ALARM AND REQUESTED FOR SUPPORT; RECOMMENDATIONS WERE GIVEN TO DECREASE THE P-LEVEL AND VERIFY THE POSITION VIA ECHO. THE PHYSICIAN REPORTED THAT THE PUMP WAS MUCH TOO FAR IN THE VENTRICLE, SO PULLED THE IMPELLA 10 CM OUT OF THE VENTRICLE UNTIL IT WAS COMPLETELY IN THE AORTA; DUE TO DECREASE OF RR, RECOMMENDATION GIVEN TO START THE CPR AND IMPLANT A NEW IMPELLA. IT WAS NOTED THAT DURING THE COMPLETE EVENT THE IMPELLA SHOWED THE RED FAILURE MODE "IMPELLA IN VENTRICLE", DESPITE THE IMPELLA BEING IN THE AORTA. ULTIMATELY THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ CORRECT IMPELLA CATHETER POSITION ¿FOR OPTIMAL POSITIONING OF THE IMPELLA CATHETER, THE INLET AREA OF THE CATHETER SHOULD BE 3.5 CM BELOW THE AORTIC VALVE ANNULUS AND WELL AWAY FROM PAPILLARY MUSCLE AND SUBANNULAR STRUCTURES. THE OUTLET AREA SHOULD BE WELL ABOVE THE AORTIC VALVE.¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE CAUSE OF THE POSITIONING ISSUES MOST LIKELY WAS USE RELATED DUE THE PUMP WAS MUCH TOO FAR IN THE VENTRICLE AND PULL THE PUMP ABOUT 10CM OUT OF THE VENTRICLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845714&pc=MHX | 3008729547-2025-00031 | 2025-03-16 04:00 | Death | CRITIKON DE MEXICO S. DE R.L. DE C.V. | 2025-04-15 04:00 | MHX | APEXPRO TELEMETRY SYSTEM V6 | Display or Visual Feedback Problem; Appropriate Term/Code Not Available | Insufficient Information | K080251 |  | 1 | Event Description: IT WAS ALLEGED THAT THE PATIENT STOPPED BEING MONITORED BY THE GEHC MONITORING SYSTEM. THE PATIENT EXPIRED. Manufacturer Narrative: GE HEALTHCARE'S INVESTIGATION FOUND THAT THE PATIENT WAS CONNECTED TO A BEDSIDE MONITOR IN COMBINATION MONITORING MODE. IN COMBINATION MODE, ECG DATA IS PROVIDED BY THE TELEMETRY SERVER AND DISPLAYED ON THE BEDSIDE MONITOR AND CENTRAL STATION. THE INVESTIGATION CONCLUDED THAT THE PATIENT DROPPED OFF THE CENTRAL STATION DISPLAY DUE TO AN UNUSUAL SEQUENCE OF EVENTS, UNEXPECTED SYSTEM PERFORMANCE, AND USER INTERACTIONS TAKEN AT THE CENTRAL STATION AND BEDSIDE MONITOR. THE CENTRAL STATION AND TELEMETRY SERVER PERFORMED TO SPECIFICATIONS. THE BEDSIDE MONITOR EXPERIENCED TWO SPONTANEOUS REBOOTS. THE CAUSE OF THE REBOOTS WAS NOT DEFINITIVELY DETERMINED. FULL MONITORING, INCLUDING PATIENT ALARMS, WAS MAINTAINED AT THE BESIDE MONITOR DURING THE INCIDENT. THE CUSTOMER WAS PROVIDED WITH THE INVESTIGATION FINDINGS. THERE IS NO INDICATION THEY EXPRESSED ANY FURTHER CONCERNS. GE HEALTHCARE HAS DETERMINED THAT NO FURTHER ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: LEGAL MANUFACTURER: HCS RESEARCH PARK - 9900 INNOVATION DRIVE USA WAUWATOSA, WI 53226 A1-6: NOT PROVIDED BY THE CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845759&pc=EOQ | 2955842-2025-14931 | 2025-03-17 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-15 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE THE PATIENT EXPIRED. THE ENDOLUMINAL SALES ASSOCIATE (ESA) REPORTED THAT UPON ARRIVAL AT THE HOSPITAL PROCEDURE ROOM, THEY SAW THAT THE ION SYSTEM WAS UNDOCKED AND BLOOD WAS COMING OUT OF THE ENDOTRACHEAL TUBE. THE PHYSICIAN HAD PLACED A BRONCHOSCOPE, WAS ADMINISTERING SALINE AND EPINEPHRINE VIA THE ENDOTRACHEAL TUBE, AND WAS ATTEMPTING TO STOP THE BLEEDING VIA SUCTIONING THE AIRWAY. THE ESA WAS LATER INFORMED THAT THE PATIENT HAD EXPIRED. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: AN ION SYSTEM LOG REVIEW COULD NOT BE PERFORMED AS THE LOGS ARE NOT AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE WITH UNKNOWN MEDICAL CO-MORBIDITIES UNDERWENT AN ION BRONCHOSCOPY FOR BIOPSY. DURING THE PROCEDURE BLEEDING WAS NOTED IN THE ET TUBE PROMPTING TRANSITION TO MANUAL BRONCHOSCOPY TO INSTILL EPINEPHRINE, SALINE AND SUCTIONING. THE PATIENT DIED DURING THE PROCEDURE. THERE WAS NO ALLEGATION OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE REPORTED EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. ATTEMPTS AT OBTAINING FURTHER INFORMATION HAS BEEN UNSUCCESSFUL. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845804&pc=NAY | 2955842-2025-14023 | 2024-05-01 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-15 04:00 | NAY | DAVINCI SINGLE PORT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Respiratory Failure; Embolism/Embolus | K171294 |  | 1 | Event Description: A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE STUDY TO COMPARE THE OUTCOMES OF 96 PATIENTS UNDERGOING ROBOTIC-ASSISTED RADICAL CYSTECTOMY (RARC) WITH URINARY DIVERSION FOR BLADDER CANCER. THE STUDY AIMED TO EVALUATE THE POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF 47 SINGLE-PORT (SP) VS 49 MULTIPORT ROBOTIC-ASSISTED RADICAL CYSTECTOMY (MP RARC). THE STUDY REPORTED THAT IN THE SP COHORT, TWO MORTALITIES WERE NOTED WITH ONE PATIENT EXPIRING ON POSTOPERATIVE DAY 1 SECONDARY TO ACUTE HEMORRHAGIC SHOCK AND RESPIRATORY FAILURE FROM PULMONARY EMBOLISM AND A SECOND PATIENT EXPIRING SECONDARY TO HEMORRHAGIC SHOCK AND RESPIRATORY DISTRESS. THE STUDY CONCLUDED THAT THE SP ROBOTIC PLATFORM OFFERS A SAFE ALTERNATIVE TO THE MP ROBOTIC PLATFORM IN PERFORMING RARC WITH URINARY DIVERSION. NO CLINICALLY SIGNIFICANT DIFFERENCE WAS DEMONSTRATED IN PERIOPERATIVE OR ONCOLOGIC OUTCOMES. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS REPORTED, NOR DID THE AUTHORS ALLEGED THAT ANY INTUITIVE PRODUCTS CAUSED OR CONTRIBUTED TO THE DEATHS. THE DESIGNATED AUTHOR CONTACT HAS NOT RESPONDED TO A REQUEST FOR ADDITIONAL INFORMATION. Manufacturer Narrative: THIS MEDWATCH REPORT REPRESENTS THE DA VINCI SINGLE PORT (SP) DEATH REPORTS. SEE SECTION H10 FOR THE ASSOCIATED DA VINCI X MULTIPORT (MP) REPORTED DEATHS. AN INVESTIGATION COULD NOT BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATES OR DA VINCI SYSTEM IDENTIFIERS. THERE WERE NO DEVICE ISSUES DOCUMENTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENTS IN THIS LITERATURE ARTICLE DIED FOLLOWING A RADICAL CYSTECTOMY, ONE FROM ACUTE HEMORRHAGIC SHOCK WITH PULMONARY EMBOLISM AND THE OTHER FROM HEMORRHAGIC SHOCK AND RESPIRATORY DISTRESS. THE DETAILS OF HOW THEY DIED ARE NOT REPORTED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: FANG, A. M., HAYEK, O., KAYLOR, J. M., PEYTON, C. C., FERGUSON, J. E., NIX, J. W., & RAIS BAHRAMI, S. (2024). POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF SINGLE-PORT VS MULTIPORT ROBOTIC-ASSISFTED RADICAL CYSTECTOMY. JOURNAL OF ENDOUROLOGY, 38(5), 438¿443. HTTPS://DOI.ORG/10.1089/END.2023.0553 SECTION A2 AGE: THE STUDY POPULATION AGE RANGE FOR DA VINCI SP WAS 67.1 +/-11.1 YEARS. MEAN 67 YEARS. SECTION A3 GENDER: THE STUDY POPULATION FOR DA VINCI SP WAS MALE - 39; FEMALE - 8. SECTION A4 WEIGHT: THE STUDY POPULATION FOR DA VINCIE SP WAS BMI OF 28.8+/-4.9 KG/M2. SECTION B3 EVENT DATE IS UNKNOWN. THE DATE REFLECTS THE DATE THE ARTICLE WAS PUBLISHED - MAY 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845822&pc=NAY | 2955842-2025-14026 | 2024-05-01 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-15 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Respiratory Acidosis | K131861 |  | 1 | Event Description: A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE STUDY TO COMPARE THE OUTCOMES OF 96 PATIENTS UNDERGOING ROBOTIC-ASSISTED RADICAL CYSTECTOMY (RARC) WITH URINARY DIVERSION FOR BLADDER CANCER. THE STUDY AIMED TO EVALUATE THE POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF 47 SINGLE-PORT (SP) VS 49 MULTIPORT (MP) ROBOTIC-ASSISTED RADICAL CYSTECTOMY. THE STUDY REPORTED THAT IN THE MP COHORT, A MORTALITY WAS NOTED WITH THE PATIENT EXPIRING ON POSTOPERATIVE DAY 2 SECONDARY TO SEVERE ACIDOSIS AND HYPOTENSION. THE STUDY CONCLUDED THAT THE SP ROBOTIC PLATFORM OFFERS A SAFE ALTERNATIVE TO THE MP ROBOTIC PLATFORM IN PERFORMING RARC WITH URINARY DIVERSION. NO CLINICALLY SIGNIFICANT DIFFERENCE WAS DEMONSTRATED IN PERIOPERATIVE OR ONCOLOGIC OUTCOMES. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS REPORTED, NOR DID THE AUTHORS ALLEGED THAT ANY INTUITIVE PRODUCTS CAUSED OR CONTRIBUTED TO THE DEATHS. THE DESIGNATED AUTHOR CONTACT HAS NOT RESPONDED TO A REQUEST FOR ADDITIONAL INFORMATION. Manufacturer Narrative: THIS MEDWATCH REPORT REPRESENTS THE DA VINCI MULTIPORT (MP) REPORTED DEATHS. SEE SECTION H10 FOR THE ASSOCIATED DA VINCI SINGLE PORT (SP) DEATH REPORTS. AN INVESTIGATION COULD NOT BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATES OR DA VINCI SYSTEM IDENTIFIERS. THERE WERE NO DEVICE ISSUES DOCUMENTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THIS ARTICLE DISCUSSED A SERIES OF PATIENTS COMPARING MULTIPORT AND SINGLE PORT ROBOTIC SURGERY WHO UNDERWENT RADICAL CYSTECTOMY. THIS REPORT IS REGARDING A MULTIPORT PATIENT WHO DIED ON POST OPERATIVE DAY 2 SECONDARY TO SEVERE ACIDOSIS AND HYPOTENSION. NO INTRAOPERATIVE DETAILS ARE PROVIDED. HOW THEY DIED AND THE CIRCUMSTANCES WHICH LED TO THEIR DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: FANG, A. M., HAYEK, O., KAYLOR, J. M., PEYTON, C. C., FERGUSON, J. E., NIX, J. W., & RAIS BAHRAMI, S. (2024). POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF SINGLE-PORT VS MULTIPORT ROBOTIC-ASSISTED RADICAL CYSTECTOMY. JOURNAL OF ENDOUROLOGY, 38(5), 438¿443. HTTPS://DOI.ORG/10.1089/END.2023.0553. SECTION A2 AGE: THE STUDY POPULATION AGE RANGE FOR DA VINCI MP WAS 67.0 +/- 9.4 YEARS. MEAN AGE 67 YEARS. SECTION A3 GENDER: THE STUDY POPULATION FOR DA VINCI MP WAS MALE - 41; FEMALE - 8. SECTION A4 WEIGHT: THE STUDY POPULATION FOR DA VINCIE MP WAS BMI OF 30.3+/-6.5 KG/M2. SECTION B3 EVENT DATE IS UNKNOWN. THE DATE REFLECTS THE DATE THE ARTICLE WAS PUBLISHED - MAY 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21845823&pc=NAY | 2955842-2025-15264 | 2025-03-13 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-15 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Myocardial Infarction | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY WEDGE RESECTION, THE PROCEDURE WAS CONVERTED TO OPEN AFTER APPROXIMATELY 4 HOURS OF DISSECTION FOR LYSIS OF ADHESIONS. THE SURGEON INFORMED THE CLINICAL SALES REPRESENTATIVE THAT THE CONVERSION WAS DUE TO THE SURGEON NOT BEING ABLE TO REMOVE ALL OF THE ADHESIONS, THUS WAS UNABLE TO VISUALIZE THE PATIENT ANATOMY, AND IN ORDER TO PREVENT FURTHER BLEEDING FROM THE EXTENDED DISSECTION. THE PATIENT EXPIRED TWO HOURS POST CONVERSION DUE TO CARDIAC SHOCK DURING THE OPEN PROCEDURE. PRIOR TO THE PROCEDURE, THE SURGEON ACKNOWLEDGED THAT THEY MIGHT HAVE TO CONVERT TO OPEN AS THE PATIENT WAS KNOWN TO BE UNSTABLE DUE TO A HISTORY OF CARDIAC ISSUES AND AN UNSPECIFIED INFECTION LIKELY CAUSING ADHESIONS. THERE WAS NO ORGAN DAMAGE REPORTED. THE PATIENT REQUIRED AN UNSPECIFIED AMOUNT OF BLOOD TRANSFUSIONS DURING THE OPEN SURGERY. IT WAS REPORTED THAT THE DEATH WAS NOT RELATED TO DA VINCI PRODUCTS.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, TIP-UP FENESTRATED GRASPER, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD KNOWN CARDIAC DISEASE AND PREDICTED TO HAVE CHALLENGING ANATOMY GOING INTO THE PROCEDURE. THEY CONVERTED TO OPEN AND THE PATIENT HAD AN EVENT 2 HOURS INTO THE OPEN PORTION OF THE OPERATION AND EXPIRED. NO ALLEGATION HAS BEEN MADE REGARDING ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONVERTED TO THIS EVENT. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833136&pc=MKJ | 3030677-2025-001172 | 2025-04-05 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-04-14 04:00 | MKJ | FRX (W B) DEFIB, US ENGLISH, EXCHG | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE DID NOT DELIVER A SHOCK WHILE IN USE ON A PATIENT. THE PATIENT WAS NOT RESPONSIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833517&pc=OZD | MW5168975 | 2024-12-11 05:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA CATH DEVICE | Material Frayed | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IMPELLA CONTROLLER BEGAN TO ALARM, LINE TRACED AND DID FIND WIRE FRAYING. DEVICE PULLED; PT EXPIRED BUT NOT ATTRIBUTED TO DEVICE. PT HAD SIGNS OF HEART FAILURE BEFORE ALARMS. NO EVENT ISOLATED THAT CAUSED DAMAGE TO LINE. LINE WAS TAKEN BACK TO ABIOMED FOR RESEARCH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833566&pc=NVN | MW5168976 |  | Death | ABBOTT MEDICAL/ST. JUDE MEDICAL | 2025-04-14 04:00 | NVN | TENDRIL STS 2088TC LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: CLINICAL STUDY REPORTED THE PATIENT PASSED AWAY IN THEIR SLEEP AT HOME. CAUSE OF DEATH IS NOT NOTED OR AVAILABLE., THE DATE OF DEATH IS CONFIRMED THROUGH PUBLIC OBITUARY. THE (B)(6) 2023, DEVICE INTERROGATION WAS REVIEWED BY A PHYSICIAN AND DEATH WAS RULED NOT TO BE DEVICE RELATED. AS-VS PACING WAS 99%, RECOVERY FROM THE DEVICE EXPLANT/IMPLANT ON (B)(6) 2023, WENT WELL., DEVICE INTERROGATION FROM (B)(6) 2023, SAYS: "DOWNLOAD SHOWS VENT NOISE REVERSION, NO EGM TO VIEW; EXTENDED DIAGNOSTICS SHOW IT OCCURRED ON DAY DEVICE REPLACED. 1 MODE SWITCH ALSO JUST SEEN ON, EXTENDED DIAGNOSTIC; EPISODE LASTED 8 SECONDS ON DEVICE REPLACEMENT DAY, AGAIN NO EGM TO VIEW. NORMALLY FUNCTIONING ICD IN DDD MODE." VENT NOISE REVERSION WAS DUE TO THE PROCEDURE ON (B)(6) 2023. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833609&pc=LOD | MW5168977 | 2025-04-01 04:00 | Death | ZIMMER BIOMET INC. | 2025-04-14 04:00 | LOD | BONE CEMENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: SUBJECT ID: (B)(6). STUDY NO: DSJ-2016-03. CLINICAL ADVERSE EVENT RECEIVED FOR DEATH - CAUSE UNKNOWN. DEVICE AND PROCEDURE (RELATEDNESS). DEVICE LOCATION: LEFT. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833668&pc=NPT | 2025587-2025-02628 | 2025-03-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-14 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Non specific EKG/ECG Changes; Pleural Effusion; Cardiac Tamponade; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AN UNSPECIFIED HEART BLOCK WAS OBSERVED. ADDITIONALLY, A COMPUTED TOMOGRAPHY SCAN REVEALED A LARGE AMOUNT OF PERICARDIAL EFFUSION AND PLEURAL EFFUSION. CARDIAC TAMPONADE OF DELAY WAS SUSPECTED. SUBSEQUENTLY, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AN UNSPECIFIED HEART BLOCK WAS OBSERVED. ADDITIONALLY, A COMPUTED TOMOGRAPHY (CT) SCAN REVEALED A LARGE AMOUNT OF PERICARDIAL EFFUSION AND PLEURAL EFFUSION. CARDIAC TAMPONADE OF DELAY WAS SUSPECTED. SUBSEQUENTLY, THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT EMERGENCY TRANSPORT WAS PROVIDED DUE TO THE PATIENT EXPERIENCING CARDIAC ARREST AT HOME. POSTMORTEM CT IMAGING SHOWED AN AORTIC DISSECTION. PER THE PHYSICIAN, IT WAS UNKNOWN WHETHER THE EVENT WAS RELATED TO THE DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUTFX DELIVERY CATHETER SYSTEM (DCS), PRODUCT ID: D-EVOLUTFX-2329, LOT: 0012416828, USE BY DATE: 2026-08-27, UDI: (B)(4). UPDATED: B3, B5, H6, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012416828); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012399238); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833768&pc=NPT | 2025587-2025-02631 | 2025-03-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-14 04:00 | NPT | EVOLUT FX DCS | Fluid/Blood Leak; Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER PROCEDURE INVOLVING THE EVFXPLUS-26, THE PATIENT, BECAME UNSTABLE AND HYPOTENSIVE UPON PLACEMENT OF THE WORKING WIRE IN THE VENTRICLE AND AORTIC INSUFFICIENCY WAS NOTED. OF NOTE, THE DELIVERY CATHETER SYSTEM HAD BEEN INSERTED INTO THE GROIN BUT WAS NOT YET ADVANCED TO THE ANNULUS. DESPITE CHEST COMPRESSIONS AND ADMINISTRATION OF EPINEPHRINE, RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT DIED. PER THE PHYSICIAN, THE PATIENT WAS HIGH-RISK HAVING HAD A PRE-OPERATIVE ECHOCARDIOGRAM REVEAL AN EJECTION FRACTION OF 20-25%. THE CAUSE OF DEATH WAS VENTRICULAR OVERLOAD POSSIBLY CAUSED BY THE PLACEMENT OF THE GUIDEWIRE.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER PROCEDURE INVOLVING THE EVFXPLUS-26, THE PATIENT, BECAME UNSTABLE AND HYPOTENSIVE UPON PLACEMENT OF THE WORKING WIRE IN THE VENTRICLE AND AORTIC INSUFFICIENCY WAS NOTED. OF NOTE, THE DELIVERY CATHETER SYSTEM HAD BEEN INSERTED INTO THE GROIN BUT WAS NOT YET ADVANCED TO THE ANNULUS. DESPITE CHEST COMPRESSIONS AND ADMINISTRATION OF EPINEPHRINE, RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT DIED. PER THE PHYSICIAN, THE PATIENT WAS HIGH-RISK HAVING HAD A PRE-OPERATIVE ECHOCARDIOGRAM REVEAL AN EJECTION FRACTION OF 20-25%. THE CAUSE OF DEATH WAS VENTRICULAR OVERLOAD POSSIBLY CAUSED BY THE PLACEMENT OF THE GUIDEWIRE. ADDITIONAL INFORMATION WAS RECEIVED THAT A NON-MEDTRONIC (SAFARI) GUIDEWIRE WAS USED FOR THE PROCEDURE. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {EVFXPLUS-26}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} PRODUCT ID {L-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} PRODUCT {GUIDEWIRE}; MANUFACTURER {UNKNOWN}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21833904&pc=NPT | 2025587-2025-02635 | 2024-12-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-14 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH OF WAS UNKNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834004&pc=BZD | 2518422-2025-104074 | 2025-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | REMSTAR AUTO A-FLEX W/HUMID, SD CARD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834007&pc=BZD | 2518422-2025-104066 | 2024-10-19 04:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834013&pc=BZD | 2518422-2025-104073 | 2021-12-28 05:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Degraded | Cancer | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834027&pc=BZD | 2518422-2025-104070 | 2025-04-07 04:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE HAS PASSED AWAY THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834031&pc=BZD | 2518422-2025-104067 | 2022-08-19 04:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Atrial Fibrillation; Pulmonary Emphysema; High Blood Pressure/ Hypertension; Chronic Obstructive Pulmonary Disease (COPD); Unspecified Immune System Problem; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SEVERE CHRONIC OBSTRUCTIVE PULMONARY DISEASE, EMPHYSEMA, INSULIN TYPE 2 W/ANGIOPATHY, NEPHROPATHY, DYSLIPIDEMIA, HYPERTENSIVE HEAT DISEASE WITH HEART FAILURE, CHRONIC PAROXYSMAL ATRIAL FIBRILLATION, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION IS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834039&pc=DSQ | 2916596-2025-02249 | 2025-03-27 04:00 | Death | THORATEC CORPORATION | 2025-04-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Abscess; Bacterial Infection; Purulent Discharge; Fall; Muscle Weakness | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE RIGHT VENTRICLE FAILURE. PER VENTRICULAR ASSIST DEVICE (VAD) TEAM, THE PATIENT HAD A DRIVELINE INFECTION AND RIGHT VENTRICLE DYSFUNCTION. RIGHT HEART FAILURE EXISTED PRE-LVAD IMPLANT, AS THE PATIENT WAS FIRST DIAGNOSED WITH HEART FAILURE IN 2000, AND RIGHT VENTRICLE DYSFUNCTION WAS FIRST NOTED IN 2021 PRE-VAD IMPLANTATION. A DEVICE RELATED ISSUE DID NOT CONTRIBUTE OR CAUSE TO RIGHT HEART FAILURE. THE PATIENT EXPERIENCED ANASARCA, GENERALIZED FATIGUE AND WEAKNESS, INCREASING IN SEVERITY, SHORTNESS OF BREATH, LIGHTHEADEDNESS, AND SEVERE MALNUTRITION. THE EVENT WAS TREATED WITH COMFORT AND HOSPICE CARE. THE PATIENT HAD MULTIPLE READMISSIONS FOR CHRONIC DRIVELINE INFECTION FROM 2023-2025. THE PATIENT WAS ADMITTED AFTER A FALL FROM (B)(6) 2025 TO (B)(6) 2025 FOR FAILURE TO THRIVE, AND GENERALIZED WEAKNESS. CLINICAL SYMPTOMS AND CULTURES INCLUDED PURULENT GREEN AND YELLOW DISCHARGE AROUND THE DRIVELINE, AND LEUKOCYTOSIS. MYOBACTERIUM FORTUITUM DRIVELINE ABSCESS AND VAD INFECTION WITH POSITIVE CULTURES FOR ACID-FAST BACILLI (AFB) ON (B)(6) 2023 AND (B)(6) 2023. THE PATIENT WAS MAINTAINED ON OMADACYCLINE AND LEVAQUIN. MOST RECENTLY, AN ABDOMINAL WALL ABSCESS INCISION AND DRAINAGE WAS PERFORMED, AND THE PATIENT WAS STARTED ON IMIPENEM. A WOUND VAC WAS ALSO PLACED. THE DEVICE WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS ALSO NOTED THAT THE PATIENT HAD A DRIVELINE INFECTION. THE PATIENT WAS ADMITTED AFTER A FALL FOR FAILURE TO THRIVE AND GENERALIZED WEAKNESS. PURULENT GREEN AND YELLOW DRAINAGE WAS NOTED AROUND THE DRIVELINE, AS WELL AS LEUKOCYTOSIS. AN ABDOMINAL WALL ABSCESS INCISION AND DRAINAGE WAS PERFORMED, AND THE PATIENT WAS STARTED ON IMIPENEM. A WOUND VAC WAS ALSO PLACED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING INFECTION AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834151&pc=NIQ | 2024168-2025-03819 | 2010-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-04-14 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN AN ARTICLE SUMMARY THAT THE PURPOSE OF THE STUDY WAS TO COMPARE 10-YEAR CLINICAL OUTCOMES OF TITANIUM-NITRIDE-OXIDE-COATED (TINO)-STENT VERSUS PERMANENT POLYMER EVEROLIMUS-ELUTING (EES), FOR THE PRIMARY ENDPOINT OF MAJOR ADVERSE CARDIAC EVENTS (MACE) IN PATIENTS WITH ACUTE CORONARY SYNDROME (ACS). FROM JANUARY 2009 TO SEPTEMBER 2010, 827 ACS PATIENTS WERE RANDOMIZED WHO UNDERWENT EARLY PCI TO RECEIVE EITHER A NON-ABBOTT TINO-STENT OR EES XIENCE V STENTS. 410 XIENCE V STENTS WERE IMPLANTED AND 417 TINO STENTS WERE IMPLANTED. AT ONE YEAR FOLLOW UP, BOTH GROUPS HAD CARDIAC DEATH, NON-FATAL MYOCARDIAL INFARCTION, ISCHEMIA DRIVEN TARGET LESION REVASCULARIZATION (TLR), NON-CARDIAC DEATH, AND DEFINITE STENT THROMBOSIS. ADDITIONAL FOLLOW UP WAS PERFORMED YEARLY THROUGH 5 YEARS AND THEN AGAIN AT 10 YEARS. IT WAS FOUND THAT THERE WAS A TREND FAVORING THE EES GROUP FOR TLR AT 1 YEAR FOLLOW-UP. AFTER 5 YEARS OF FOLLOW-UP THE RATE OF ST WAS LOW AND COMPARABLE IN BOTH STUDY GROUPS. AT 10-YEAR FOLLOW-UP, THE RATE OF MACE WAS SIGNIFICANTLY LOWER IN ACS PATIENTS TREATED WITH TINO-STENTS COMPARED TO PATIENTS TREATED WITH EES. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "10-YEAR FOLLOW-UP OF PATIENTS WITH TITANIUM-NITRIDE-OXIDE-COATED STENTS VERSUS EVEROLIMUS-ELUTING STENTS IN ACUTE CORONARY SYNDROME". Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. C4: TREATMENT/THERAPY START DATE ESTIMATED. D6A: DATE OF IMPLANT ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT ARTICLE TITLED, 10-YEAR FOLLOW-UP OF PATIENTS WITH TITANIUM-NITRIDE-OXIDE-COATED STENTS VERSUS EVEROLIMUS-ELUTING STENTS IN ACUTE CORONARY SYNDROME. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH AND RELATIONSHIP TO PRODUCT IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834168&pc=NPT | 2025587-2025-02639 | 2025-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-14 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS HOSPITALIZED DUE TO CONGESTIVE HEART FAILURE (CHF). SUBSEQUENTLY, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE CHF AND DEVICE OR IMPLANT PROCEDURE. PER THE PHYSICIAN, THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CLARIFIED PREVIOUSLY REPORTED INFORMATION INDICATING THAT THE PATIENT DIED APPROXIMATELY 9 MONTHS FOLLOWING THE IMPLANT OF THE VALVE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS RELATED TO CONGESTIVE HEART FAILURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834294&pc=QBJ | 3004753838-2025-087292 | 2025-04-09 04:00 | Death | DEXCOM, INC. | 2025-04-14 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | Coma | K213919 |  | 1 | Event Description: IT WAS REPORTED BY INSULET THAT AN UNKNOWN PATIENT EXPERIENCED BRAIN DEATH FOLLOWING COMA WITH ALLEGATION OF UNSPECIFIED DEXCOM MALFUNCTION. THE REPORT WAS MADE BY AN UNIDENTIFIED INDIVIDUAL ON A SOCIAL MEDIA POST TO INSULET. INSULET AND DEXCOM WERE UNABLE TO DETERMINE THE PATIENT REFERENCED THE COMPLAINT. THERE WAS NO DOCUMENTATION OF FURTHER MEDICAL EVALUATION OR INTERVENTION. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. PROBLEM COULD NOT BE CONFIRMED AND PROBABLE CAUSE CANNOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834326&pc=DSQ | 2916596-2025-02189 | 2025-03-24 04:00 | Death | THORATEC CORPORATION | 2025-04-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS SEPSIS. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION IS UNKNOWN. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834554&pc=NPT | 2025587-2025-02644 | 2025-03-15 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-14 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure; Unspecified Heart Problem | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE OUTCOMES OF TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) WITH SELF-EXPANDING VALVES IN JAPANESE PATIENTS ON DIALYSIS. THE STUDY POPULATION CONSISTED OF TEN PATIENTS WHO UNDERWENT TAVR WITH A SELF-EXPANDING VALVE FROM THE MEDTRONIC EVOLUT PORTFOLIO (EVOLUT R, PRO, OR PRO+). ADVERSE OUTCOMES RECOUNTED IN THE ARTICLE: CARDIOPULMONARY RESUSCITATION, VASOPRESSOR, AND PERCUTANEOUS CARDIOPULMONARY SUPPORT REQUIRED DUE TO THE OCCURRENCE OF HYPOTENSION AND CARDIAC ARREST DURING THE TAVR PROCEDURE (ONE CASE); MYOCARDIAL INFARCTION (ONE CASE); MAJOR BLEEDING VASCULAR COMPLICATION (ARTERY HEMORRHAGE) ATTRIBUTED TO THE TAVR PROCEDURE (ONE CASE); AND AORTIC REGURGITATION (TRACE TO MILD). ADDITIONALLY, BETWEEN THE 6-MONTH AND 12-MONTH FOLLOW-UP POINTS, ONE PATIENT EXPERIENCED SUDDEN CARDIAC DEATH DURING SLEEP, AND ANOTHER PATIENT DIED IN THE HOSPITAL DUE TO HEART FAILURE, WITH NO FURTHER DETAILS KNOWN. Manufacturer Narrative: CITATION: SHISHIDO K, YAMANAKA F, MORIYAMA N, ET AL. SAFETY AND EFFECTIVENESS OF SELF-EXPANDING TAVR IN JAPANESE DIALYSIS PATIENTS WITH SEVERE AORTIC STENOSIS: 1-YEAR OUTCOMES. J CARDIOL. PUBLISHED ONLINE MARCH 15, 2025. DOI:10.1016/J.JJCC.2025.03.008. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. EARLIEST APPROVED EVOLUT PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834601&pc=JXG | 9612501-2025-01043 | 2025-03-14 04:00 | Death | MEDTRONIC DOMINICANA | 2025-04-14 04:00 | JXG | UNKNOWN VALVE/SHUNT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: MICHAEL J. STUART, ALISON WRAY, MARK DEXTER, ROBERT AJ. CAMPBELL. VALIDATION OF DATA CAPTURE IN THE AUSTRALASIAN SHUNT REGISTRY WITH A PROSPECTIVELY MAINTAINED INSTITUTIONAL DATABASE. JOURNAL OF CLINICAL NEUROSCIENCE 135 (2025). DOI: 10.1016/J.JOCN.2025.111179 A B S T R A C T BACKGROUND: THE AUSTRALASIAN SHUNT REGISTRY WAS ESTABLISHED IN 2016 WITH THE AIM OF PROVIDING SAFETY AND QUALITY DATA IN ADDITION TO ENHANCING CAPACITY FOR RESEARCH TO IMPROVE HEALTH OUTCOMES FOR PATIENTS WITH CEREBROSPINAL FLUID SHUNTS. THE QUEENSLAND CHILDREN¿S HOSPITAL IS THE LARGEST SINGLE CONTRIBUTOR OF PAEDIATRIC PATIENT DATA TO THE REGISTRY AND MAINTAINS AN INDEPENDENT INSTITUTIONAL SHUNT DATABASE. THE AIM OF THIS STUDY WAS TO VALIDATE THE DEMOGRAPHIC DATA AND OUTCOMES CAPTURED BY THE REGISTRY AGAINST THAT INSTITUTIONAL DATABASE. METHODS: DATA FROM THE INSTITUTIONAL DATABASE AND A SAMPLE OF DATA FROM THE SAME INSTITUTION HELD BY THE AUSTRALASIAN SHUNT REGISTRY WERE ACQUIRED FROM THE PERIOD OF 1 JANUARY 2017 TO 1 OCTOBER 2024. PATIENTS = 18 YEARS OF AGE UNDERGOING PLACEMENT OF A NEW VENTRICULOPERITONEAL SHUNT SYSTEM WERE INCLUDED. REVISION AND MORTALITY DATA WERE ACQUIRED FROM THE AUSTRALASIAN SHUNT REGISTRY AND QUEENSLAND STATEWIDE ELECTRONIC MEDICAL RECORDS INDEPENDENTLY. COMPARISON WAS MADE BETWEEN DEMOGRAPHIC AND MORTALITY DATA CAPTURED BY EACH DATABASE. THE PRIMARY OUTCOME OF TIME TO SHUNT REVISION (SHUNT SURVIVAL) IN EACH SAMPLE WAS ASSESSED THROUGH THE GENERATION OF KAPLAN-MEIER CURVES AND ANALYSIS BY BOTH WILCOXON AND LOG-RANK TESTS. RESULTS: OVER THE STUDY PERIOD THE SHUNT REGISTRY REPORTED A FULL OR PARTIAL OPT-OUT RATE OF 9.5 %. 344 PATIENTS WERE IDENTIFIED AND INCLUDED FROM THE INSTITUTIONAL DATABASE, AND 294 PATIENTS WERE INCLUDED FROM THE REGISTRY. THE IDENTIFIED PATIENT SAMPLES WERE DEMOGRAPHICALLY SIMILAR WITH A MEANAGE OF 5 YEARS, AND 44 % FEMALE IN BOTH GROUPS. THE MOST COMMON AETIOLOGIES OF HYDROCEPHALUS IN BOTH COHORTS WERE CONGENITAL (28 % VS 28 %, P = 0.49), TUMOUR (33 % VS 27 %, P = 0.08) AND HAEMORRHAGE (25 % VS 18 %, P = 0.02). MORTALITY DURING FOLLOW UP WAS CONSISTENT ACROSS BOTH SAMPLES (13 % VS 11 %, P = 0.27). SIMILARLY, BINARY REVISION STATUS DURING THE FOLLOW UP PERIOD WAS SIMILAR (34 % VS 32 %, P = 0.3). KAPLAN-MEIER ANALYSIS OF TIME TO REVISION (SHUNT SURVIVAL) ESTIMATED SHUNT SURVIVAL TO BE 4.82 YEARS (95 % CI 4.42¿¿5.22 YEARS) IN THE INSTITUTIONAL DATABASE AND 5.25 (95 % CI 4.81¿5.69 YEARS) IN THE REGISTRY WITH NO SIGNIFICANT DIFFERENCES BETWEEN THE SAMPLES ON WILCOXON, P = 0.3 OR LATE LOG-RANK, P = 0.36 TESTS. CONCLUSIONS: THE AUSTRALASIAN SHUNT REGISTRY APPEARS TO CAPTURE A VALID SAMPLE WHICH IS REPRESENTATIVE OF THE DEMOGRAPHICS AND CLINICAL OUTCOMES OF PATIENTS TREATED AT ONE LARGE CONTRIBUTING INSTITUTION. ONGOING EFFORTS TO ENSURE COMPREHENSIVE DATA CAPTURE AT ALL PARTICIPATING SITES ARE JUSTIFIED TO ENSURE THAT FUTURE FINDINGS DERIVED FROM REGISTRY DATA ARE REPRESENTATIVE OF THE STUDIED POPULATION. REPORTED EVENTS: - THERE WERE 76 MORTALITIES DURING FOLLOW UP. THE LITERATURE ARTICLE INDICATED THAT DELTA AND STRATA VALVES WERE USED, HOWEVER, IT DID NOT INDICATE WHETHER THE VALVES WERE INVOLVED WITH THE REPORTED EVENTS. SEE ATTACHED LITERATURE ARTICLE. Manufacturer Narrative: B2. INTERVENTION REQUIRED CHECKBOX WAS CHECKED, HOWEVER, IT SHOULD NOT HAVE BEEN CHECKED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834685&pc=DSQ | 3007042319-2025-00760 | 2025-04-04 04:00 | Death | HEARTWARE, INC. | 2025-04-14 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPIRED DUE TO BRAIN BLEEDING. NO ADDITIONAL INFORMATION WILL BE FORTHCOMING. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IN USE. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO A BRAIN BLEED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE BLEEDING AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21834725&pc=NIK | 9614453-2025-01462 | 2025-02-02 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-04-14 04:00 | NIK | COBALT¿ HF CRT-D MRI SURESCAN¿ | Insufficient Information | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) IS DECEASED. THE PATIENT HAD A HISTORY OF MULTIPLE VENTRICULAR TACHYCARDIA (VT) ABLATIONS AND HEART FAILURE. ADDITIONAL INFORMATION FOR A CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21835041&pc=HSB | 0009613350-2025-00289 | 2025-03-18 04:00 | Death | ZIMMER GMBH | 2025-04-14 04:00 | HSB | ZNNÂ¿¢, CMN LAG SCREW, Ã¸ 10.5 MM, 95 MM, INCLUDING SET SCREW | Adverse Event Without Identified Device or Use Problem | Bone Fracture(s); Hemorrhage/Bleeding; Muscle/Tendon Damage | K231114 |  | 1 | Event Description: IT WAS REPORTED THAT INITIAL OPERATION WAS PERFORMED WITH CM NAIL SYSTEM, PATIENT UNDERWENT EXPLANTED SURGERY APPROXIMATELY 10 MONTHS POST IMPLANTATION. DURING THE SURGERY SURGEON HAD DIFFICULTY REMOVING THE LAG SCREW AND NAIL AND SHOWED A DROP IN BLOOD PRESSURE INTRAOPERATIVELY BUT WAS ABLE TO PULL IT OUT SUCCESSFULLY. POSTOPERATIVE ECHO IMAGING SHOWED BLEEDING IN THE ABDOMINAL CAVITY, AND CT IMAGING REVEALED EVIDENCE THAT THE LAG SCREW HAD PENETRATED THE BONY HEAD. THE OUTER PELVIS WAS PUNCTURED, AND THE INNER PELVIS WAS SLIGHTLY FRACTURED. THE PATIENT WAS RUSHED TO ANOTHER EMERGENCY FACILITY, BUT THE PATIENT DIED ONE DAY AFTER SURGERY. THE CAUSE OF THE HEMORRHAGE WAS ATTRIBUTED TO THE SURGEON'S DIFFICULTY IN REMOVING THE LAG SCREW, WHICH FOR SOME REASON WAS PUSHED IN THE DIRECTION OF THE EPIPHYSIS, DAMAGING THE ILIOPSOAS MUSCLE. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4). D10: ITEM # 47249318110, CMN FEMORAL NAIL, CCD 125°, LEFT, 10 MM, 18 CM, LOT # 3123884. ITEM # 00249000450, CM LAG SCREW INSERTER- ASIA, LOT # UNKOWN. ITEM # 47249318110, CM LAG SCREW RET SHAFT ASIA, LOT # UNKOWN. ITEM # 47248403050, 5.0 MM DIAMETER CORTICAL SCREW - RED FIXED ANGLE 3.5 MM HEX HEAD, LOT # 66260877. ITEM # 47248700200, CEPHALOMEDULLARY NAIL CAP 0 MM HEIGHT, LOT # 66435604. G2: REPORT SOURCE JAPAN. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21835042&pc=HSB | 0009613350-2025-00290 | 2025-03-18 04:00 | Death | ZIMMER GMBH | 2025-04-14 04:00 | HSB | CMN FEMORAL NAIL, CCD 125Â°, LEFT, Ã¸ 10 MM, 18 CM | Adverse Event Without Identified Device or Use Problem | Bone Fracture(s); Hemorrhage/Bleeding; Muscle/Tendon Damage | K231114 |  | 1 | Event Description: IT WAS REPORTED THAT INITIAL OPERATION WAS PERFORMED WITH CM NAIL SYSTEM, PATIENT UNDERWENT EXPLANTED SURGERY APPROXIMATELY 10 MONTHS POST IMPLANTATION. DURING THE SURGERY SURGEON HAD DIFFICULTY REMOVING THE LAG SCREW AND NAIL AND SHOWED A DROP IN BLOOD PRESSURE INTRAOPERATIVELY BUT WAS ABLE TO PULL IT OUT SUCCESSFULLY. POSTOPERATIVE ECHO IMAGING SHOWED BLEEDING IN THE ABDOMINAL CAVITY, AND CT IMAGING REVEALED EVIDENCE THAT THE LAG SCREW HAD PENETRATED THE BONY HEAD. THE OUTER PELVIS WAS PUNCTURED, AND THE INNER PELVIS WAS SLIGHTLY FRACTURED. THE PATIENT WAS RUSHED TO ANOTHER EMERGENCY FACILITY, BUT THE PATIENT DIED ONE DAY AFTER SURGERY. THE CAUSE OF THE HEMORRHAGE WAS ATTRIBUTED TO THE SURGEON'S DIFFICULTY IN REMOVING THE LAG SCREW, WHICH FOR SOME REASON WAS PUSHED IN THE DIRECTION OF THE EPIPHYSIS, DAMAGING THE ILIOPSOAS MUSCLE. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4): D10: ITEM # 47248509510, ZNNÂ¿¢, CMN LAG SCREW, Ã¸ 10.5 MM, 95 MM, INCLUDING SET SCREW, LOT # 3182509. ITEM # 00249000450, CM LAG SCREW INSERTER- ASIA, LOT # UNKOWN. ITEM # 47249318110, CM LAG SCREW RET SHAFT ASIA-, LOT # UNKOWN. ITEM # 47248403050, 5.0 MM DIAMETER CORTICAL SCREW - RED FIXED ANGLE 3.5 MM HEX HEAD, LOT # 66260877. ITEM # 47248700200, CEPHALOMEDULLARY NAIL CAP 0 MM HEIGHT, LOT # 66435604. G2: REPORT SOURCE JAPAN THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21835043&pc=HSB | 0009613350-2025-00291 | 2025-03-18 04:00 | Death | ZIMMER GMBH | 2025-04-14 04:00 | HSB | CM LAG SCREW INSERTER- ASIA | Adverse Event Without Identified Device or Use Problem | Bone Fracture(s); Hemorrhage/Bleeding; Muscle/Tendon Damage | K231114 |  | 1 | Event Description: IT WAS REPORTED THAT INITIAL OPERATION WAS PERFORMED WITH CM NAIL SYSTEM, PATIENT UNDERWENT EXPLANTED SURGERY APPROXIMATELY 10 MONTHS POST IMPLANTATION. DURING THE SURGERY SURGEON HAD DIFFICULTY REMOVING THE LAG SCREW AND NAIL AND SHOWED A DROP IN BLOOD PRESSURE INTRAOPERATIVELY BUT WAS ABLE TO PULL IT OUT SUCCESSFULLY. POSTOPERATIVE ECHO IMAGING SHOWED BLEEDING IN THE ABDOMINAL CAVITY, AND CT IMAGING REVEALED EVIDENCE THAT THE LAG SCREW HAD PENETRATED THE BONY HEAD. THE OUTER PELVIS WAS PUNCTURED, AND THE INNER PELVIS WAS SLIGHTLY FRACTURED. THE PATIENT WAS RUSHED TO ANOTHER EMERGENCY FACILITY, BUT THE PATIENT DIED ONE DAY AFTER SURGERY. THE CAUSE OF THE HEMORRHAGE WAS ATTRIBUTED TO THE SURGEON'S DIFFICULTY IN REMOVING THE LAG SCREW, WHICH FOR SOME REASON WAS PUSHED IN THE DIRECTION OF THE EPIPHYSIS, DAMAGING THE ILIOPSOAS MUSCLE. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4). D10: ITEM # 47249318110, CMN FEMORAL NAIL, CCD 125Â°, LEFT, Ã¸ 10 MM, 18 CM, LOT # 3123884. ITEM # 47248509510, ZNNÂ¿¢, CMN LAG SCREW, Ã¸ 10.5 MM, 95 MM, INCLUDING SET SCREW, LOT # 3182509. ITEM # 47249318110, CM LAG SCREW RET SHAFT ASIA-, LOT # UNKOWN. ITEM # 47248403050, 5.0 MM DIAMETER CORTICAL SCREW - RED FIXED ANGLE 3.5 MM HEX HEAD, LOT # 66260877. ITEM # 47248700200, CEPHALOMEDULLARY NAIL CAP 0 MM HEIGHT, LOT # 66435604. G2: REPORT SOURCE JAPAN. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21835046&pc=HSB | 0009613350-2025-00292 | 2025-03-18 04:00 | Death | ZIMMER GMBH | 2025-04-14 04:00 | HSB | CM LAG SCREW RET SHAFT ASIA- | Adverse Event Without Identified Device or Use Problem | Bone Fracture(s); Hemorrhage/Bleeding; Muscle/Tendon Damage | K231114 |  | 1 | Event Description: IT WAS REPORTED THAT INITIAL OPERATION WAS PERFORMED WITH CM NAIL SYSTEM, PATIENT UNDERWENT EXPLANTED SURGERY APPROXIMATELY 10 MONTHS POST IMPLANTATION. DURING THE SURGERY SURGEON HAD DIFFICULTY REMOVING THE LAG SCREW AND NAIL AND SHOWED A DROP IN BLOOD PRESSURE INTRAOPERATIVELY BUT WAS ABLE TO PULL IT OUT SUCCESSFULLY. POSTOPERATIVE ECHO IMAGING SHOWED BLEEDING IN THE ABDOMINAL CAVITY, AND CT IMAGING REVEALED EVIDENCE THAT THE LAG SCREW HAD PENETRATED THE BONY HEAD. THE OUTER PELVIS WAS PUNCTURED, AND THE INNER PELVIS WAS SLIGHTLY FRACTURED. THE PATIENT WAS RUSHED TO ANOTHER EMERGENCY FACILITY, BUT THE PATIENT DIED ONE DAY AFTER SURGERY. THE CAUSE OF THE HEMORRHAGE WAS ATTRIBUTED TO THE SURGEON'S DIFFICULTY IN REMOVING THE LAG SCREW, WHICH FOR SOME REASON WAS PUSHED IN THE DIRECTION OF THE EPIPHYSIS, DAMAGING THE ILIOPSOAS MUSCLE. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4): D10: ITEM # 47249318110, CMN FEMORAL NAIL, CCD 125Â°, LEFT, Ã¸ 10 MM, 18 CM, LOT # 3123884. ITEM # 47248509510, ZNNÂ¿¢, CMN LAG SCREW, Ã¸ 10.5 MM, 95 MM, INCLUDING SET SCREW, LOT # 3182509. ITEM # 00249000450, CM LAG SCREW INSERTER- ASIA, LOT # UNKOWN. ITEM # 47248403050, 5.0 MM DIAMETER CORTICAL SCREW - RED FIXED ANGLE 3.5 MM HEX HEAD, LOT # 66260877. ITEM # 47248700200, CEPHALOMEDULLARY NAIL CAP 0 MM HEIGHT, LOT # 66435604. G2: REPORT SOURCE JAPAN THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21835151&pc=DSQ | 2916596-2025-02334 | 2024-09-12 04:00 | Death | THORATEC CORPORATION | 2025-04-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS EXCHANGED ON (B)(6) 2024 FOR AN EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO) EXTENDING THE LENGTH OF THE OUTFLOW GRAFT (MFR# 2916596-2024-05237). THE PATIENT HAD A COMPLICATED POST-OPERATIVE COURSE DUE TO BLEEDING (MFR# 2916596-2025-02452). THE PATIENT HAD A SIGNIFICANT STROKE ON (B)(6) 2024 AND THE FAMILY DECIDED TO WITHDRAW CARE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21835188&pc=NPT | 2025587-2025-02647 | 2025-03-19 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-14 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Cardiogenic Shock; Low Cardiac Output | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, A 26 MM NON-MEDTRONIC VALVE (SAPIEN 3) WAS PLANNED TO BE IMPLANTED VIA LEFT SUBCLAVIAN ACCESS BUT COULD NOT BE ADVANCED INTO THE AORTA AND A DISSECTION OF THE LEFT SUBCLAVIAN ARTERY OCCURRED. ACCORDING TO THE REPORTER, IT WAS BELIEVED THAT THE NON-MEDTRONIC VALVE OR DELIVERY CATHETER SYSTEM CAUSED THE DISSECTION. SUBSEQUENTLY, THE DISSECTED ARTERY WAS STENTED, AND A 29 MM EVOLUT FX+ VALVE WAS SUCCESSFULLY IMPLANTED. HOWEVER, ONCE THE VALVE BECAME "OCCLUSIVE," THE PATIENT DECOMPENSATED AND A COMBINATION OF HYPOTENSION, HEMODYNAMIC INSTABILITY, LOW CARDIAC OUTPUT, DECREASED EJECTION FRACTION, AND CARDIOGENIC SHOCK ENSUED. THEN THE PATIENT PROGRESSED INTO VENTRICULAR FIBRILLATION, AND ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) WAS INITIATED. ULTIMATELY, THE IMPLANT TEAM PLACED THE PATIENT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). AT THE TIME OF THE REPORT, THE PATIENT WAS LISTED AS ALIVE AND STILL ON ECMO.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, A 26 MM NON-MEDTRONIC VALVE WAS PLANNED TO BE IMPLANTED VIA LEFT SUBCLAVIAN ACCESS BUT COULD NOT BE ADVANCED INTO THE AORTA AND A DISSECTION OF THE LEFT SUBCLAVIAN ARTERY OCCURRED. ACCORDING TO THE REPORTER, IT WAS BELIEVED THAT THE NON-MEDTRONIC VALVE OR DELIVERY CATHETER SYSTEM CAUSED THE DISSECTION. SUBSEQUENTLY, THE DISSECTED ARTERY WAS STENTED, AND A 29 MM EVOLUT FX+ VALVE WAS SUCCESSFULLY IMPLANTED. HOWEVER, ONCE THE VALVE BECAME "OCCLUSIVE," THE PATIENT DECOMPENSATED AND A COMBINATION OF HYPOTENSION, HEMODYNAMIC INSTABILITY, LOW CARDIAC OUTPUT, DECREASED EJECTION FRACTION, AND CARDIOGENIC SHOCK ENSUED. THEN THE PATIENT PROGRESSED INTO VENTRICULAR FIBRILLATION, AND ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) WAS INITIATED. ULTIMATELY, THE IMPLANT TEAM PLACED THE PATIENT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). AT THE TIME OF THE REPORT, THE PATIENT WAS LISTED AS ALIVE AND STILL ON ECMO. ADDITIONAL INFORMATION WAS RECEIVED WHICH CLARIFIED THAT THE VALVE BECOMING "OCCLUSIVE" REFERRED TO ANNULAR CONTACT MADE WITH NO FORWARD OUTPUT UNTIL THE VALVE FRAME WAS OPEN ENOUGH TO ALLOW BLOOD FLOW. NINE DAYS LATER, THE PATIENT WAS TAKEN OFF ECMO AND SUBSEQUENTLY DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS DUE TO THE PATIENT'S WEAK HEART MUSCLE AND INABILITY TO RECOVER. PER THE PHYSICIAN, THE VALVE COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH. ADDITIONALLY, THE PATIENT'S UNDERLYING MEDICAL HISTORY OF POOR EJECTION FRACTION AND HEART FAILURE CAUSED/CONTRIBUTED TO THE DEATH. Manufacturer Narrative: UPDATED DATA: B2. B5. SECOND PARAGRAPH ADDED H1. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836242&pc=NVZ | 2017865-2025-40102 | 2025-01-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NVZ | FORTIFY ASSURA DR ICD, US | Insufficient Information | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836357&pc=NIK | 2017865-2025-40104 | 2025-01-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NIK | GALLANT HF | Insufficient Information | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836382&pc=NVN | 2017865-2025-40103 | 2025-01-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem; Insufficient Information | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: H6. THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN TWO PIECES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES HOWEVER AN INTERNAL SHORT WAS NOTED BETWEEN CONDUCTOR PATHS AT THE DISTAL REGION CONSISTENT WITH PROCEDURAL DAMAGE VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836425&pc=NVY | 2017865-2025-40105 | 2025-01-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem; Insufficient Information | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: H6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836431&pc=NVN | 2017865-2025-40107 | 2025-01-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NVN | ULTIPACE¿ LEAD, 52CM MRI | Adverse Event Without Identified Device or Use Problem; Insufficient Information | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION H6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836457&pc=OJX | 2017865-2025-40106 | 2025-01-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | OJX | QUARTET LEAD DOUBLE BEND, 86 CM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: H6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836530&pc=OZD | 1220648-2025-27418 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Shock | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, A HEMOTHORAX OF UNKNOWN CAUSE WAS CONFIRMED. THE BLEEDING COULD NOT BE CONTROLLED, AND THE PATIENT DIED OF HEMORRHAGIC SHOCK BEFORE BEING REMOVED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836547&pc=OZD | 1220648-2025-27421 | 2025-04-01 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT PASSED AWAY FROM AN INFECTION DURING IMPELLA 5.5 SUPPORT. THE SOURCE OF THE INFECTION IS UNKNOWN. AS SUCH, THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836559&pc=OZD | 1220648-2025-27425 | 2025-02-23 05:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THERE WAS AN UNSPECIFIED DEATH. A NEW IMPELLA WAS INSERTED AFTER THE ORIGINAL ONE WAS ACCIDENTLY REMOVED DURING TRANSPORTATION. AFTER IMPELLA CP INSERTION, VENTRICULAR SEPTAL PERFORATION (VSP) WAS FOUND BY ECHOCARDIOGRAM AND SURGERY WAS PLANNED, BUT THE PATIENT'S OVERALL CONDITION DETERIORATED AND THEY PASSED AWAY FROM SEVERE HEART FAILURE. THE PATIENT WAS ALREADY IN A SERIOUS CONDITION AND THE RELATIONSHIP TO IMPELLA IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836567&pc=OZD | 1220648-2025-27423 | 2025-03-16 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND THE PATIENT REPORTED LOSING FEELING IN HER IMPELLA LEG. THE MEDICAL DOCTOR INSERTED REPERFUSION SHEATH AND SYMPTOMS IMPROVED. IT WAS FURTHER REPORTED THAT THE PATIENT CODED AND EXPIRED. NO FURTHER INFORMATION PROVIDED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) - ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836576&pc=OZD | 1220648-2025-27434 | 2025-03-18 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Hemorrhage/Bleeding; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE DECISION WAS MADE TO PLACE AN IMPELLA 5.5 ON A PATIENT WITH MULTI VESSEL DISEASE. THE IMPELLA 5.5 WAS PLACED PER PROTOCOL WITHOUT ISSUE. DURING SUPPORT, THE PATIENT EXPERIENCED SIGNIFICANT ABDOMINAL PAIN FOLLOWED BY A LACTATE OF 16 REQUIRING INCREASES IN CHEMICAL SUPPORT. COMPUTED TOMOGRAPHY IMAGING REVEALED A THROMBUS IN ABDOMINAL AORTA THAT WAS OCCLUDING THE SUPERIOR MESENTERIC ARTERY (SMA) AND PARTIALLY OCCLUDING THE CELIAC ARTERY. ADDITIONALLY, THERE WAS EXCESSIVE BLEEDING FROM THE IMPELLA SITE WITH COAGULOPATHY/SHOCK LIVER REQUIRING BLOOD PRODUCTS. THE PATIENT WAS TAKEN TO OPERATING ROOM EMERGENTLY FOR EXPLORATORY LAPAROTOMY. NECROTIC ASCENDING AND TRANSVERSE COLON RESECTED FOLLOWED BY SMA THROMBECTOMY. ADDITIONAL AREA OF BOWEL WAS REMOVED. THE BELLY WAS LEFT OPEN WITH WOUND VAC IN PLACE. URINE OUTPUT HAD PICKED BACK UP SINCE VOLUME RESUSCITATION. THE IMPELLA POCKET WAS WASHED OUT AND RE-CLOSED. IT WAS ALSO NOTED THAT THE PATIENT HAD SUCTION ALARMS THAT WERE VOLUME RELATED. MORE BLOOD PRODUCTS AND VOLUME WERE ADMINISTERED AND SUCTION ALARMS RESOLVED. ADDITIONAL PACKED RED BLOOD CELLS AND PLASMA WERE ALSO BEING ADMINISTERED. ULTIMATELY, THE FAMILY WITHDREW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836581&pc=OZD | 1220648-2025-27436 | 2025-03-15 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA 5.5 WITH A HISTORY OF CEREBROVASCULAR ACCIDENT (CVA) WAS TAKEN FOR A HEAD COMPUTED TOMOGRAPHY AND FOUND A NEW ISCHEMIC CVA. HEPARIN WAS HELD. ADDITIONALLY, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION ARREST THAT REQUIRED 33 SHOCKS. CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. MEDICAL SAFETY REVIEW OF THE EVENT WAS COMPLETED, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. ALTHOUGH THE DEMISE IS UNLIKELY FOR THE IMPELLA PUMP GIVEN ANTICIPATED MISUSE: PATIENT POST-SURGERY WITH ACTIVATED CLOTTING TIME SUBTHERAPEUTIC FOR IMPELLA USE DUE TO THE BLEED POTENTIAL. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. THE ROOT CAUSE OF STROKE AND VENTRICULAR FIBRILLATION CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836617&pc=OZD | 1220648-2025-27420 | 2025-04-01 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: ABIOMED RECEIVED INFORMATION THROUGH SUBMITTED LITERATURE "PROCEEDINGS OF THE JAPANESE SOCIETY OF INTENSIVE CARE MEDICINE (2024; VOL.31. NOSUPPL.1,S1051-)" THAT A PATIENT EXPERIENCED MULTIPLE ADVERSE EVENTS DURING IMPELLA SUPPORT. IT WAS INITIALLY NOTED THAT A THROMBUS WAS DISCOVERED VIA COMPUTED TOMOGRAPHY SCAN IN THE DESCENDING AORTA AROUND THE IMPELLA SHAFT. ANTICOAGULATION THERAPY WAS STRENGTHENED TO TREAT THE CONDITION. THE PATIENT WAS THEN FOUND TO HAVE SUFFERED A CEREBRAL HEMORRHAGE. WHEN THE ORIGINAL THROMBUS WAS FOUND TO HAVE SHRUNK, THE DECISION WAS MADE TO REMOVE THE PUMP. ON A LATER DATE, A BARE STENT WAS PLACED IN THE PROXIMAL SUPERIOR MESENTERIC ARTERY AND THE PROCEDURE WAS COMPLETED, BUT THE PATIENT'S MULTIPLE ORGAN FAILURE PROGRESSED AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING AND IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. D.4 CATALOG NUMBER, SERIAL NUMBER, LOT NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN. D.6A AND D.6B IMPLANT AND EXPLANT DATES ARE UNKNOWN. E.1 NAME PREFIX/TITLE, FIRST NAME AND LAST NAME ARE UNKNOWN. H.4 DEVICE MANUFACTURE DATE IS UNKNOWN. LITERATURE WAS NOT AVAILABLE TO ATTACH TO THIS REPORT. Manufacturer Narrative: AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION. SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ H.10 ADDED INSTRUCTIONS FOR USE AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27420.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836689&pc=OZD | 1220648-2025-27456 | 2025-03-17 04:00 | Death | ABIOMED, INC. | 2025-04-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, WHILE PUSHING IN A COMPANION SHEATH, THE IMPELLA PUMP WAS NOT STABILIZED AND SUBSEQUENTLY PERFORATED THE LEFT VENTRICLE. AS SUPPORT CONTINUED, THE PATIENT'S BLOOD PRESSURE BEGAN TO CONSISTENTLY DROP. SUCTION WAS NOTED AND THE PUMP WAS PULLED BACK. AT THAT POINT IN TIME THE PATIENT'S BLOOD PRESSURE DROPPED DRAMATICALLY. AN ECHOCARDIOGRAM SHOWED A HOLE IN THE APEX. PERICARDIOCENTESIS WAS PERFORMED, BUT THE PATIENT BECAME HYPOVOLEMIC. VOLUME RESUSCITATION WAS NOT SUCCESSFUL, AND CPR WAS PERFORMED FOR THIRTY MINUTES WITHOUT RECOVERY. THE PATIENT PASSED AWAY. FOLLOWING THE CASE, IT WAS NOTED THAT THE LEFT VENTRICLE MAY HAVE BEEN FRIABLE DUE TO A RECENT INFARCTION. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION." ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836786&pc=QZI | 2124215-2025-23164 | 2025-03-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-14 04:00 | QZI | FARAWAVE NAV | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation |  |  | 1 | Event Description: DURING A PULSED FIELD ABLATION (PFA) PROCEDURE A FARAWAVE WAS SELECTED FOR USE. THE LEFT PULMONARY VEINS WERE MAPPED AND ABLATED SUCCESSFULLY. THEN, THE PHYSICIAN MOVED TO THE RIGHT SUPERIOR PULMONARY VEIN (RSPV) AND HAD SOME DIFFICULTIES, THE GUIDEWIRE GOT INTO THE PERICARDIUM AND IT WAS OBSERVED THAT THE BLOOD PRESSURE DROPPED. TRANSOPHAGEAL ECHO WAS PERFORMED AND BLOOD IN THE PERICARDIUM WAS OBSERVED. A PERICARDIOCENTESIS WAS PERFORMED, BUT DRAINING WAS NOT ENOUGH SO THE PATIENT WAS TAKEN TO CARDIAC SURGERY TO SOLVE THE ISSUE. A WOUND WAS OBSERVED ON THE ANTERIOR ROOF, CLOSE TO THE RSPV. THE PATIENT WAS PUT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT WAS KEPT FOR OBSERVATION, AND IT'S EXPECTED TO RECOVER. THE DEVICES WILL NOT BE RETURNED AS THEY WERE DISPOSED. IT WAS FURTHER REPORTED THAT TWO NON-BOSTON SCIENTIFIC DEVICES WERE USED FOR TRANSEPTAL PUNCTURE. IT WAS OBSERVED THAT THE CATHETER WAS IN AN UNUSUAL POSITION/ORIENTATION ON THE MAPPING SYSTEM. A CARDIAC MASSAGE WAS PERFORMED BY THE ENTIRE MEDICAL TEAM BEFORE THE DRAINAGE WAS PERFORMED, WHILE A TRANSTHORACIC ECHO CONFIRMED A SIGNIFICANT PERICARDIAL EFFUSION. THE PATIENT HAS EXPIRED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. THIS PRODUCT IS NOT APPROVED IN THE UNITED STATES, HOWEVER, IT HAS BEEN ASSESSED AS REPORTABLE AS THERE IS A SIMILAR PRODUCT APPROVED IN THE UNITED STATES. WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION RELATED TO UNITED STATES REGISTRATION AS THE SPECIFIC PRODUCT IS ONLY COMMERCIALLY AVAILABLE OUTSIDE OF THE UNITED STATES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836793&pc=BZD | 2518422-2025-104120 | 2022-01-01 05:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21836861&pc=BZD | 2518422-2025-104121 | 2025-04-10 04:00 | Death | RESPIRONICS, INC. | 2025-04-14 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE DEVICE WAS RETURNED FOR EVALUATION. DURING THE EVALUATION, NO FAULTS WERE REPORTED. THE ACCESSORY MODULE FLIP DOOR WAS FOUND MISSING FROM THE DEVICE. THE POLLEN FILTER AND FINE FILTER WERE REPLACED DUE TO PREVENTATIVE MAINTENANCE. THE DEVICE WAS CLEANED AND TESTED. THE DEVICE PASSED ALL TESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837118&pc=MGB | 2024168-2025-03848 | 2022-03-24 04:00 | Death | ABBOTT VASCULAR | 2025-04-14 04:00 | MGB | UNK PERCLOSE | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Unspecified Tissue Injury | P960043 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WHO UNDERWENT HEMOSTASIS WITH [PERCLOSE] DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE THREE YEARS AGO DIED FROM AN INFECTION (SEPSIS). THE TAVI PROCEDURE WAS PERFORMED AT (B)(6) HOSPITAL AND THE PATIENT WAS LATER TRANSFERRED TO A HOSPITAL WHERE A FAMILY MEMBER WORKS AS A NURSE (A FACILITY SIMILAR TO A NURSING HOME), WHERE THE PATIENT PASSED AWAY. THE TIME OF DEATH WAS NOT RECENT; IT OCCURRED AFTER THE TRANSFER, AND THERE HAVE BEEN VARIOUS EXCHANGES SINCE THEN. DURING THESE EXCHANGES, A PLASTIC SURGEON EXAMINED THE SUTURE SITE AND DIAGNOSED LYMPHATIC LEAKAGE, COMPLICATING THE SITUATION. A CARDIOVASCULAR SURGEON INDICATED THAT THE [PERCLOSE] SUTURE WAS THE CAUSE. NO FURTHER INFORMATION COULD BE OBTAINED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE PATIENT EFFECTS OF DEATH, SEPSIS, INFECTION, AND TISSUE INJURY ARE LISTED IN THE PROSTYLE INSTRUCTIONS FOR USE AS POTENTIAL ADVERSE EVENTS ASSOCIATED WITH USE OF THE SUTURE MEDIATED CLOSURE DEVICES. BASED ON THE CASE INFORMATION, THERE WAS NO CONCLUSIVE CAUSE(S) FOR THE REPORTED DIFFICULTY, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY THAT CAN BE DETERMINED. THERE IS THEREFORE, NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO CORRECTIVE ACTION IS REQUIRED.. B2, B3: ESTIMATED DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837150&pc=DSQ | 2916596-2025-02068 | 2025-02-24 05:00 | Death | THORATEC CORPORATION | 2025-04-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED RIGHT VENTRICULAR (RV) FAILURE AFTER THE PUMP WAS EXCHANGED. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MULTISYSTEM ORGAN FAILURE. THE PUMP WORKED AS ANTICIPATED. THIS WAS ALL THOUGHT NOT TO BE RELATED TO THE VENTRICULAR ASSIST DEVICE (VAD). IT WAS UNKNOWN WHETHER PUMP WOULD BE EXPLANTED OR RETURNED. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS, AS WELL AS THE SUBSEQUENT PATIENT OUTCOME, COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS MULTIPLE ORGAN FAILURES/DYSFUNCTIONS (INCLUDING RIGHT HEART FAILURE, RESPIRATORY FAILURE, AND RENAL DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CHAPTER 6, ¿PATIENT CARE AND MANAGEMENT¿, STATES THAT SUDDEN RIGHT VENTRICULAR FAILURE MAY DEVELOP DURING OR SHORTLY AFTER PUMP IMPLANTATION AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837151&pc=DQX | 3003775027-2025-00103 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | PROWATERFLEX | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K163426 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI PROWATERFLEX GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS: METHODS [STUDY DESIGN AND POPULATION] THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS] MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE] THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS] FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME] IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME] TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837292&pc=DQX | 3003775027-2025-00104 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | FIELDER XT | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K171933 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI FIELDER XT GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS: METHODS. [STUDY DESIGN AND POPULATION] THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS]: MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE]: THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS] FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME] IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME] TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837293&pc=DQX | 3003775027-2025-00105 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | FIELDER | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K163426 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI FIELDER GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479 TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS: METHODS [STUDY DESIGN AND POPULATION] THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS]: MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON-Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE]: THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS]: FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME]: IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME]: TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS' CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837361&pc=DQX | 3003775027-2025-00111 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | CONFIANZA | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K171933 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI CONFIANZA GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS: METHODS: [STUDY DESIGN AND POPULATION]. THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS]: MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE]: THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS: [POPULATION CHARACTERISTICS] FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME]: IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME]: TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837362&pc=DQX | 3003775027-2025-00108 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | MIRACLE 4.5 | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K171933 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI MIRACLE 4.5 GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS: METHODS [STUDY DESIGN AND POPULATION] THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS] MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE] THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS] FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME] IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME] TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837363&pc=DQX | 3003775027-2025-00106 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | FIELDER FC | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K163426 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI FIELDER FC GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479 TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY EXCERPTS: METHODS [STUDY DESIGN AND POPULATION] THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS] MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE] THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS] FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME] IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME] TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837364&pc=DQX | 3003775027-2025-00110 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | MIRACLE 12 | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K171933 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI MIRACLE 12 GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY EXCERPTS: METHODS [STUDY DESIGN AND POPULATION] THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS] MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE] THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS] FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME] IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME] TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837365&pc=DQX | 3003775027-2025-00107 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | MIRACLE 3 | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K163426 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI MIRACLE 3 GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS: METHODS: [STUDY DESIGN AND POPULATION]: THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS]: MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE]: THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS: [POPULATION CHARACTERISTICS]: FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME]: IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME]: TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS CHARACTERISTICS. AGE, (MEAN
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837366&pc=DQX | 3003775027-2025-00109 |  | Death | ASAHI INTECC CO., LTD. | 2025-04-14 04:00 | DQX | MIRACLE 6 | Insufficient Information | Stroke/CVA; Myocardial Infarction; Perforation of Vessels; Cardiac Tamponade; Thromboembolism; Renal Impairment | K171933 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITERATURE THAT AN ASAHI MIRACLE 6 GUIDE WIRE MIGHT HAVE CAUSED OR CONTRIBUTED TO MAJOR ADVERSE CARDIAC OR CEREBROVASCULAR EVENTS (MACES). PUBLICATION: EUROINTERVENTION 2011;7:472-479. TITLE: IN-HOSPITAL OUTCOMES OF PERCUTANEOUS CORONARY INTERVENTION IN PATIENTS WITH CHRONIC TOTAL OCCLUSION: INSIGHTS FROM THE ERCTO (EUROPEAN REGISTRY OF CHRONIC TOTAL OCCLUSION) REGISTRY. EXCERPTS METHODS [STUDY DESIGN AND POPULATION]: THE ERCTO REGISTRY IS A PROSPECTIVE, REAL-WORLD REGISTRY IN WHICH ALL PATIENTS TREATED PERCUTANEOUSLY FOR ONE OR MORE CTO LESIONS INVOLVING A MAIN CORONARY ARTERY (>-2.5MM) OR A SAPHENOUS BYPASS CONDUIT WERE ENROLLED. DATA COLLECTION WAS CARRIED OUT IN 16 CENTRES ACROSS EUROPE. THE ENROLMENT STARTED ON THE 1ST OF JANUARY 2008 AND CONTINUES TO DATE. THE SCHEDULED DURATION OF CLINICAL FOLLOW-UP IS THREE YEARS AFTER CTO PROCEDURE. [DEFINITIONS]: MAJOR ADVERSE CARDIAC EVENTS (MACE) WERE DEFINED AS CARDIAC DEATH, ACUTE MYOCARDIAL INFARCTION (AMI) OR TARGET VESSEL REVASCULARISATION (TVR). ALL DEATHS WERE CONSIDERED OF CARDIAC ORIGIN UNLESS OTHERWISE DOCUMENTED. NON Q-WAVE AMI WAS DEFINED AS A CREATINE KINASE-MB (CK-MB) ENZYME ELEVATION > 3 TIMES THE UPPER LIMIT OF THE NORMAL VALUE. WHEN NEW PATHOLOGICAL Q WAVES, IN ADDITION TO ENZYME ELEVATION, WERE OBSERVED IN THE ELECTROCARDIOGRAM, THE EVENT WAS DEFINED AS A Q-WAVE AMI. TVR WAS DEFINED AS REPEATED REVASCULARISATION WITHIN THE TREATED VESSEL. STENT THROMBOSIS WAS DEFINED, ACCORDING WITH THE DEFINITION OF THE ACADEMIC RESEARCH CONSORTIUM. A PROCEDURE-RELATED CONTRAST-INDUCED NEPHROPATHY (CIN) WAS DEFINED AS AN INCREASE OF 25% OR 0.5MG/DL IN SERUM CREATININE AT 48 HOURS AFTER PCI COMPARING BASELINE VALUES. [ANGIOGRAPHIC PROCEDURE]: THE STRATEGIES USED INCLUDED THE SINGLE GUIDEWIRE TECHNIQUE, THE PARALLEL WIRE TECHNIQUE, INTRAVASCULAR ULTRASOUND (IVUS)-GUIDED WIRING TECHNIQUE, SUBINTIMAL TRACKING AND RE-ENTRY TECHNIQUES, AS WELL AS RETROGRADE TECHNIQUES DESCRIBED ELSEWHERE. DEDICATED GUIDEWIRES WERE MOSTLY USED IN ORDER TO PERFORM THESE TECHNIQUES, AND RANGED FROM SOFT POLYMERIC WIRES, INCLUDING FIELDER XT (ASAHI INTECC), FIELDER AND FIELDER FC (ASAHI INTECC), TO STIFF FLAT OR TAPERED SPRING ONES SUCH AS MIRACLE 3-12 (ASAHI INTECC), CONFIANZA FAMILY (ASAHI INTECC). ALL PATIENTS WERE CONTROLLED FOR ANY POST-PROCEDURAL EVENT SUCH AS EVIDENCE OF PERICARDIAL TAMPONADE, CHEST PAIN, HEART FAILURE, BLEEDING, OR ANY ISCHAEMIC EVENT. CK-MB, AND TROPONIN I WERE MEASURED AT 6-8 AND 12-24 HOURS AFTER THE PROCEDURE (TOTAL RATE OF SAMPLE FOR CK-MB AT EACH TIME POINTS 87%, FOR TROPONIN I 24%), WHILE SERUM CREATININE WAS MEASURED AT 24 AND 48 HOURS AFTER PROCEDURE (TOTAL RATE OF SAMPLE 83%). RESULTS [POPULATION CHARACTERISTICS]: FROM JANUARY 2008 TO THE END OF DECEMBER 2009 A TOTAL OF 1,914 PATIENTS WITH 1,983 CTO LESIONS WERE REGISTERED INTO A WEBSITE DATABASE. BASELINE PATIENTS' CHARACTERISTICS ARE SHOWN IN TABLE 1. IN ACCORDANCE WITH PREVIOUS OTHER LARGE REGISTRIES, THE MAJORITY OF PATIENTS WERE MALES. [ANGIOGRAPHIC OUTCOME]: IN THE MAJORITY OF CASES, THE WIRES WHICH FINALLY REACHED THE DISTAL TRUE LUMEN WERE STIFF TAPERED GUIDEWIRES SUCH AS THE CONQUEST GUIDEWIRES SERIES, CROSS-IT GUIDEWIRES SERIES AND PERSUADER 9 (38.8%). THE DETAILS OF THE WIRES THAT CROSSED THE LESIONS ARE REPORTED IN FIGURE 1. NOTABLY, THE FIELDER XT GUIDEWIRE WAS ABLE TO REACH THE DISTAL TRUE LUMEN IN 10.5% OF LESIONS. [IN-HOSPITAL OUTCOME]: TABLE 4 SUMMARISES COMPLICATIONS AND IN-HOSPITAL OUTCOMES, STRATIFIED ACCORDING TO THE TYPE OF APPROACH CARRIED OUT. POST-PROCEDURE ASSESSMENT OF CK-MB WAS PERFORMED IN 97% OF PATIENTS. AMONG THESE, A MERE 26 PATIENTS SATISFIED THE DEFINITION OF MI. THERE WAS AN UNFAVOURABLE TREND TOWARDS NON-Q WAVE MI IN PATIENTS TREATED WITH THE RETROGRADE APPROACH, WHILE NO DIFFERENCES WERE FOUND CONCERNING THE Q-WAVE MI. THERE WAS ONE CASE OF DEFINITE STENT THROMBOSIS AND ONE CASE OF STROKE, BOTH IN PATIENTS TREATED WITH THE ANTEROGRADE APPROACH. CORONARY PERFORATION OCCURRED IN 2.6% OF CASES, BUT MORE FREQUENTLY IN PATIENTS WHO UNDERWENT THE RETROGRADE APPROACH. CORONARY PERFORATION CAUSED CARDIAC TAMPONADE IN 0.5% OF CASES (11/1,983); IN ALL OF THESE CASES PERICARDIOCENTESIS WAS PERFORMED. ACCESS SITE VASCULAR COMPLICATIONS WERE OBSERVED RARELY (0.7%), AND THEY REQUIRED VASCULAR SURGERY IN 2/14 (14.3%) OF THESE CASES. IN-HOSPITAL DEATH OCCURRED IN SEVEN PATIENTS (0.3%); IN TWO CASES, DEATH WAS NOT RELATED TO CARDIAC CAUSE (REACUTISATION OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE); IN THE OTHER TWO CASES DEATH WAS CAUSED BY DAMAGE TO THE DONOR ARTERY (ONE WAS A FLOW-LIMITING DISSECTION OF THE LEFT MAIN DURING A RETROGRADE APPROACH; THE SECOND, A THROMBOSIS OF THE DONOR ARTERY RELATED TO THE SIMULTANEOUS CORONARY CANNULATION DURING AN ANTEGRADE PROCEDURE); TWO PATIENTS DIED AFTER EMERGENCY CARDIAC SURGERY, UNDERTAKEN BECAUSE OF CORONARY PERFORATION; AND ONE PATIENT DIED FROM SUBACUTE STENT THROMBOSIS. NO DIFFERENCES WERE OBSERVED IN THE DISTRIBUTION OF ACCESS SITE COMPLICATIONS BETWEEN PATIENTS TREATED BY FEMORAL VS. RADIAL APPROACHES (0.7% VERSUS 0.4%, P<0.6). PATIENTS WITH PROCEDURAL FAILURE SHOWED A HIGHER RATE OF 30-DAY DEATH AND NON-FATAL MI IN COMPARISON WITH PATIENTS SUCCESSFULLY REVASCULARISED (1.9% VERSUS 0, P<0.001 AND 3.6% VERSUS 0.8%, P<0.001, RESPECTIVELY). TABLE 1. BASELINE PATIENTS' CHARACTERISTICS. AGE, (MEAN}SD) --- 64.1}11.6. MALE, (%) --- 1627 (85). DIABETES, (%) --- 563 (28.8). HYPERTENSION, (%) --- 1,412 (72.3). DYSLIPIDAEMIA, (%) --- 1,463 (74.9). FAMILY HISTORY OF CAD, (%) --- 609 (31.2). SMOKE, (%) --- 851 (43.6). PREVIOUS STROKE, (%) --- 46 (2.4). PERIPHERAL VASCULAR DISEASE, (%) --- 194 (9.9). PREVIOUS MI, (%) --- 847 (43.4). PREVIOUS PCI, (%) --- 1,002 (51.2). PREVIOUS CABG, (%) --- 285 (14.6). LOW LVEF (<35%), (%) --- 146 (7.6). EGFR, (MEAN+-SD) --- 92.7+-70.4. FIGURE 1. THIS GRAPH SHOWS THE PREVALENCE OF WIRES ABLE TO SUCCESSFULLY CROSS THE CTO LESIONS IN THE STUDY. SOFT SPRING WIRES INCLUDE: BMW, BMW UNIVERSAL, PROWATERFLEX, RUNTHROUGH. SOFT PLASTIC WIRES INCLUDE: FIELDER, WHISPER, PT GRAPHIC AS FLAT TIP TYPE AND FIELDER XT AS TAPERED TIP TYPE. MEDIUM SPRING WIRES INCLUDE: MEDIUM AND MIRACLE 3 AS FLAT TIP TYPE AND CROSSIT-100 AS TAPERED TIP TYPE. MEDIUM PLASTIC WIRES INCLUDE ALL PILOT WIRE SERIES. STIFF SPRING WIRES INCLUDE: MIRACLE 4.5, 6, 9, 12, PERSUADER 6 AS FLAT TIP TYPE AND CONQUEST WIRE SERIES, CROSSIT-200, 300, 400 AND PERSUADER 9 AS TAPERED TIP TYPE. SUCCESS RATE (%): SOFT SPRING WIRE: FLAT TIP ---2.2, TAPERED TIP --- 0. SOFT PLASTIC WIRE: FLAT TIP ---16.5, TAPERED TIP --- 10.5. MEDIUM SPRING WIRE: FLAT TIP ---11.4, TAPERED TIP --- 3.4. MEDIUM PLASTIC WIRE: FLAT TIP ---17.2, TAPERED TIP --- 0. STIFF SPRING WIRE: FLAT TIP ---11.7, TAPERED TIP --- 27.1. TABLE 4. PROCEDURAL COMPLICATION AND IN-HOSPITAL OUTCOME. ALL PROCEDURES (N=1983) /ANTEROGRADE (N=1749) /RETROGRADE (N=234) /P. [PROCEDURAL COMPLICATION]: CONTRAST INDUCED NEPHROPATHY, (%) --- 18 (0.9) /15 (0.8) /3 (1.2) /0.5. CORONARY PERFORATION, (%) --- 53 (2.6) /42 (2.1) /11 (4.7) /0.04. CARDIAC TAMPONADE, (%) --- 11 (0.5) /9 (0.5) /2 (0.8) 0.5. EMERGENT CABG, (%) --- 3 (0.1) /3 (0.2) /0 /0.5. VASCULAR ACCESS SITE COMPLICATION, (%) --- 14 (0.7) /12 (0.6) /2 (0.8) /0.5. [IN-HOSPITAL OUTCOME]: DEATH, (%) --- 7 (0.3) /6 (0.3) /1 (0.4) /0.8. NON-Q WAVE MYOCARDIAL INFARCTION, (%) --- 23 (1.2) /18 (1) /5 (2.1) /0.08. Q WAVE MYOCARDIAL INFARCTION, (%) --- 3 (0.1) /2 (0.1) /1 (0.4) /0.1. STENT THROMBOSIS, (%) --- 1 (0.05) /1 (0.05) /0 /0.8. TVR --- 1 (0.05) /1 (0.05) /0 /0.8. STROKE, (%) --- 1 (0.05) /1 (0.05) /0 /0.8. Manufacturer Narrative: MANUFACTURING SITE: MANUFACTURING SITE COULD NOT BE IDENTIFIED BECAUSE THE PRODUCT LOT NUMBER INFORMATION WAS NOT AVAILABLE. THE BRAND NAME OF THE SUBJECT DEVICE IS MIRACLE 6, WHICH IS DISTRIBUTED AS MIRACLEBROS 6 IN THE US. THEREFORE, THE BRAND NAME IN D1 IS MIRACLE 6 AS INDICATED IN THE PACKAGE LABEL OF THE SUBJECT DEVICE. DEVICE EVALUATION COULD NOT BE PERFORMED BECAUSE THE AFFECTED DEVICES WERE DISCARDED AND NOT RETURNED FROM THE USER FACILITY. LOT HISTORY RECORDS REVIEW COULD NOT BE CONDUCTED BECAUSE LOT INFORMATION WAS UNAVAILABLE. ALL THE SHIPPED PRODUCTS WERE INSPECTED IN THE PRODUCTION PROCESS AND SATISFIED THE PRODUCT SPECIFICATIONS AND RELEASE CRITERIA; THEREFORE, IT WAS CONCLUDED THAT THERE WAS NO ANOMALY IN PRODUCT QUALITY. DUE TO LIMITED WRITTEN INFORMATION ON THE ARTICLE WHICH DEVICES HAD BEEN USED FOR WHICH PATIENTS AND WHAT CAUSED THE REPORTED MACES, IT WAS UNABLE TO DETERMINE WHETHER OR HOW MIRACLE 6 GUIDE WIRES HAD CAUSED OR CONTRIBUTED TO THE EVENTS. REFERRING TO KNOWN SIMILAR EVENTS, IT WAS PRESUMED THAT PATIENTS' ANATOMY AND/OR PROCEDURAL CONTENTS WERE LIKELY ASSOCIATED WITH THE EVENTS. IT WAS CONCLUDED THAT THE REPORTED EVENTS WERE NOT ATTRIBUTED TO PRODUCT QUALITY. NO CAPA WILL BE TAKEN. INSTRUCTIONS FOR USE (IFU) STATES: [WARNINGS] ~ OBSERVE MOVEMENT OF THIS GUIDE WIRE IN THE VESSELS. BEFORE THIS GUIDE WIRE IS MOVED OR TORQUED, THE TIP MOVEMENT SHOULD BE EXAMINED AND MONITORED UNDER FLUOROSCOPY. DO NOT MOVE OR TORQUE THE GUIDE WIRE WITHOUT OBSERVING CORRESPONDING MOVEMENT OF THE TIP; OTHERWISE, THE GUIDE WIRE MAY BE DAMAGED AND/OR TRAUMA MAY OCCUR. IN ADDITION, ENSURE THAT THE DISTAL TIP OF THIS GUIDE WIRE AND ITS LOCATION IN THE VESSEL ARE VISIBLE DURING MANIPULATIONS OF THE GUIDE WIRE. ~ NEVER PUSH, AUGER, WITHDRAW, OR TORQUE THIS GUIDE WIRE THAT MEETS RESISTANCE. TORQUING OR PUSHING THIS GUIDE WIRE AGAINST RESISTANCE MAY CAUSE DAMAGE AND/OR TIP SEPARATION OF THIS GUIDE WIRE OR DIRECT DAMAGE TO A VESSEL. RESISTANCE MAY BE FELT AND/OR OBSERVED UNDER FLUOROSCOPY BY NOTING ANY BUCKLING OF THE GUIDE WIRE. IF THE PROLAPSE OF THE GUIDE WIRE TIP IS OBSERVED, DO NOT ALLOW THE TIP TO REMAIN IN A PROLAPSED POSITION; OTHERWISE, DAMAGE TO THE GUIDE WIRE MAY OCCUR. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE ANY NECESSARY REMEDIAL ACTION. ~ IF RESISTANCE IS FELT DUE TO SPASM, BENDING OF THE GUIDE WIRE, OR DUE TO TRAP WHILE OPERATING THIS GUIDE WIRE IN THE BLOOD VESSEL OR REMOVING IT, DO NOT TORQUE AND/OR PULL THE GUIDE WIRE ITSELF. STOP THE PROCEDURE. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE APPROPRIATE REMEDIAL ACTION. IF THE GUIDE WIRE IS MOVED EXCESSIVELY, IT MAY BREAK OR BECOME DAMAGED, WHICH MAY CAUSE BLOOD VESSEL INJURY OR RESULT IN FRAGMENTS BEING LEFT INSIDE THE VESSEL. ~ USE THIS GUIDE WIRE CAREFULLY AS THE GUIDE WIRE MAY PENETRATE THE BLOOD VESSEL. OTHERWISE, IT MAY CAUSE ADVERSE EVENTS SUCH AS BLOOD VESSEL PERFORATION AND CORONARY ARTERY DISSECTION. THE HIGHER TORQUE PERFORMANCE, STIFFER DISTAL END, AND/OR HIGHER ADVANCEMENT FORCE MAY PRESENT A HIGHER RISK OF PERFORATION OR INJURY THAN IF USING A MORE FLEXIBLE GUIDE WIRE. THEREFORE, USE THE MOST FLEXIBLE GUIDE WIRE THAT WILL TREAT THE LESION (I.E., THE GUIDE WIRE WITH THE SMALLEST TIP LOAD THAT WILL TREAT THE LESION), AND TAKE DUE CARE TO MINIMIZE THE RISK OF PERFORATION OR OTHER DAMAGE TO BLOOD VESSELS. [MALFUNCTION AND ADVERSE EFFECTS]: ~ DAMAGE TO A VESSEL, INCLUDING POSSIBLE VESSEL PERFORATION. ~ CARDIAC TAMPONADE DUE TO VESSEL PERFORATION. ~ VESSEL THROMBUS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837510&pc=NVY | 2017865-2025-40141 | 2025-03-25 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Under-Sensing | Arrhythmia | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD CPR AND EXTERNAL SHOCKS BY THE EMS. INTERROGATION NOTED THAT THE PATIENT EXPERIENCED EPISODES OF ARRHYTHMIA, WHICH THE DEVICE DID NOT TREAT THE PATIENT¿S RHYTHM APPROPRIATELY, AND UNDER-SENSING WAS NOTED ON THE RIGHT VENTRICULAR (RV) LEAD. PROGRAMMING CHANGES WERE MADE. THE PATIENT WAS NOTED TO BE UNSTABLE AND INTUBATED. IT WAS LATER STATED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21837511&pc=NIK | 2017865-2025-40140 | 2025-03-25 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-14 04:00 | NIK | GALLANT DR | Failure to Convert Rhythm | Arrhythmia | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD CPR AND EXTERNAL SHOCKS BY THE EMS. INTERROGATION NOTED THAT THE PATIENT EXPERIENCED EPISODES OF ARRHYTHMIA, WHICH THE DEVICE DID NOT TREAT THE PATIENT¿S RHYTHM APPROPRIATELY, AND UNDER-SENSING WAS NOTED ON THE RIGHT VENTRICULAR (RV) LEAD. PROGRAMMING CHANGES WERE MADE. THE PATIENT WAS NOTED TO BE UNSTABLE AND INTUBATED. IT WAS LATER STATED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE PROVIDED SESSION RECORD WAS REVIEWED BY TECHNICAL SERVICES, AND IT WAS OBSERVED THAT THERE WAS NO INDICATION OF DEVICE MALFUNCTION. THE DEVICE WAS PERFORMING PER PROGRAMMED SETTINGS. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21823557&pc=DTB | 2649622-2025-09756 | 2025-03-24 04:00 | Death | MPRI | 2025-04-11 04:00 | DTB | CAPSUREFIX LEAD | Adverse Event Without Identified Device or Use Problem | Erythema; Unspecified Infection; Sepsis; Swelling/ Edema | P930039 |  | 1 | Event Description: IT WAS FURTHER REPORTED THAT THE IPG WAS USED AS AN EXTERNAL PULSE GENERATOR.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED DUE TO SEPSIS OF UNKNOWN ORIGIN.Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED AN INFECTION WITH REDNESS AND SWELLING. IT WAS ALSO REPORTED THAT THE PATIENT REQUIRED ANTIBIOTIC MEDICATION. THE LEADS WERE EXPLANTED AND REPLACED. THE IMPLANTABLE PULSE GENERATOR (IPG) REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT POCKET CULTURES WERE NEGATIVE. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2011; W1DR01 IPG IMPLANTED (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21823572&pc=DTB | 2649622-2025-09758 | 2025-03-24 04:00 | Death | MPRI | 2025-04-11 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Erythema; Unspecified Infection; Sepsis; Swelling/ Edema | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED AN INFECTION WITH REDNESS AND SWELLING. IT WAS ALSO REPORTED THAT THE PATIENT REQUIRED ANTIBIOTIC MEDICATION. THE LEADS WERE EXPLANTED AND REPLACED. THE IMPLANTABLE PULSE GENERATOR (IPG) REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE IPG WAS USED AS AN EXTERNAL PULSE GENERATOR.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED DUE TO SEPSIS OF UNKNOWN ORIGIN.Event Description: IT WAS FURTHER REPORTED THAT POCKET CULTURES WERE NEGATIVE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5568-45 LEAD IMPLANTED (B)(6) 2011; 5076-52 LEAD; IMPLANTED (B)(6) 2011. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21823583&pc=PNJ | 9612164-2025-01878 | 2025-04-09 04:00 | Death | MEDTRONIC IRELAND | 2025-04-11 04:00 | PNJ | MICRA¿ AV2 | Electrical /Electronic Property Problem; High impedance; Adverse Event Without Identified Device or Use Problem; High Capture Threshold | Cardiac Tamponade; Respiratory Failure | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED UNSATISFACTORY ELECTRICAL PARAMETERS. THE LEADLESS IPG WAS REPOSITIONED AND THE ELECTRICAL PARAMETERS WERE NOW SATISFACTORY. AT THIS POINT THE PATIENT EXPERIENCED RESPIRATORY FAILURE AND TAMPONADE. RESUSCITATION WAS STARTED AND THE PATIENT RECEIVED ADRENALINE. PERICARDIOCENTESIS WAS ALSO PERFORMED. HOWEVER, THE PATIENT DIEDEvent Description: IT WAS FURTHER REPORTED THAT THE LEADLESS IPG EXHIBITED HIGH IMPEDANCE AND THRESHOLDS. Manufacturer Narrative: CORRECTION TO H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA AV] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21823606&pc=NVZ | 3008973940-2025-02564 | 2025-03-24 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-11 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Erythema; Unspecified Infection; Sepsis; Swelling/ Edema | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED AN INFECTION WITH REDNESS AND SWELLING. IT WAS ALSO REPORTED THAT THE PATIENT REQUIRED ANTIBIOTIC MEDICATION. THE LEADS WERE EXPLANTED AND REPLACED. THE IMPLANTABLE PULSE GENERATOR (IPG) REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE IPG WAS USED AS AN EXTERNAL PULSE GENERATOR.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED DUE TO SEPSIS OF UNKNOWN ORIGIN.Event Description: IT WAS FURTHER REPORTED THAT POCKET CULTURES WERE NEGATIVE. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED (B)(6) 2011; 5568-45 LEAD; IMPLANTED (B)(6) 2011. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21823642&pc=OZO | 2032227-2025-153116 | 2019-05-02 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-11 04:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Break | Heart Block | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2019 AND THE CAUSE OF DEATH WAS HEART ILLNESS. THE LAST KNOWN PRODUCT(S) FOR THE CUSTOMER WERE AS FOLLOWS: MMT-1715K, MMT-332A, UNOMEDICAL. THE PUMP WAS WORN AT THE TIME OF PASSING. MMT-1715K WAS RETURNED FOR ANALYSIS. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. THE PUMP PASSED THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST AND FORCE SENSOR TEST. UNABLE TO PERFORM THE OCCLUSION TEST, DISPLACEMENT TEST, AND THE DAT DUE TO DETACHED RETAINER RING. THE PUMP WAS RECEIVED WITH A FADED SERIAL NUMBER LABEL. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: MINOR SCRATCHED DISPLAY WINDOW, SCRATCHED DISPLAY WINDOW COVER, SCRATCHED CASE, FADED END CAP ADDRESS LABEL, PILLOWING KEYPAD OVERLAY, CRACKED KEYPAD OVERLAY AT THE SELECT BUTTON, KEYPAD OVERLAY TEXTURE DAMAGE AND DETACHED RETAINER RING. TEST P-CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. HISTORY DOWNLOAD WAS SUCCESSFUL USING THUS AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP DID NOT HAVE A BATTERY INSTALLED WHEN RECEIVED. THE PUMP WAS RECEIVED WITHOUT THE ORIGINAL BATTERY CAP. UNABLE TO LIST THE DAILY TOTAL OF ALL INSULIN DELIVERED ON THE EVENT DATE (B)(6) 2019 AND THE 7 DAYS PRIOR TO THAT DATE DUE TO NO DATA AVAILABLE IN THE PUMP HISTORY FILE. UNABLE TO VERIFY BOLUS DELIVERY, SOURCE OF BOLUSES DELIVERED, DAILY TOTAL OF ALL INSULIN DELIVERED AND ANY ALARMS/SUSPENDS FOR EVENT DATE (B)(6) 2019 DUE TO NO DATA AVAILABLE IN THE PUMP HISTORY FILE. UNABLE TO PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATE (B)(6) 2019 IN THE PUMP HISTORY FILE DUE TO NO DATA AVAILABLE. UNABLE TO COMPLETE THE FUNCTIONAL TESTING DUE TO DETACHED RETAINER RING. THE PUMP PASSED THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST AND FORCE SENSOR TEST. UNABLE TO PERFORM THE OCCLUSION TEST, DISPLACEMENT TEST, AND THE DAT DUE TO DETACHED RETAINER RING. COSMETIC DAMAGE WAS CONFIRMED AT THE BACK OF THE PUMP (FADED SERIAL NUMBER LABEL). DAMAGE- LABEL DAMAGE OR MISSING CONFIRMED. DAMAGE-RETAINER RING DAMAGE CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21824716&pc=NPT | 2015691-2025-02819 | 2024-03-24 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-11 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE JAPANESE TAVI REGISTRY THAT ON POST OPERATIVE DAY (POD) 3 OF A TRANSAORTIC TAVR PROCEDURE WITH A 26MM SAPIEN 3 VALVE CONSCIOUS IMPAIRMENT WAS OBSERVED. A COMPUTED TOMOGRAPHY (CT) REVEALED AN OCCLUSION OF THE LEFT MIDDLE CEREBRAL ARTERY AND A THROMBECTOMY WAS PERFORMED. ON POD 5, SEDATION WAS DISCONTINUED, BUT DUE TO SEIZURES, SEDATION WAS READMINISTERED. ON POD 80, THE PATIENT EXPIRED DUE TO THE CEREBRAL INFARCT. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE STROKE CANNOT BE DETERMINED, IT IS POSSIBLE THE MECHANISMS OF THE TAVR PROCEDURE DESCRIBED ABOVE OR PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21824729&pc=OZD | 1220648-2025-27395 | 2025-03-15 04:00 | Death | ABIOMED, INC. | 2025-04-11 04:00 | OZD | IMPELLA 5.5 | Pumping Stopped | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ON IMPELLA 5.5 SUPPORT WHEN AN IMPELLA DEFECTIVE ALARM OCCURRED, AND THE PUMP STOPPED. THE IMPELLA WAS EXCHANGED. THE PATIENT EXPIRED FIVE DAYS LATER FROM MULTIPLE ORGAN FAILURE. PER ABIOMED'S MEDICAL SAFETY OFFICER REVIEW, THE IMPELLA 5.5 DID NOT WORK/STOPPED AND IT CANNOT BE SAID THAT THE IMPELLA 5.5 PUMP DID NOT CONTRIBUTE TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF PUMP STOP HAS BEEN COMPLETED. AFTER 5 DAYS ON IMPELLA , SUDDEN IMPELLA FAILURE ALARM OCCURRED AND PUMP STOPPED. IT WAS REPLACED BY A SECOND PUMP. DATA LOGS WERE NOT RECOVERED. PUMP WAS RETURNED AND INSPECTED. NO PHYSICAL ABNORMALITIES WERE NOTED ON THE PUMP HOWEVER IT WAS UNABLE TO CONNECT TO ANY CONSOLE DURING INVESTIGATION. CONNECTION TO A CONSOLE RESULTED IN IMPELLA DEFECTIVE ALARM BEING TRIGGERED. NO BROKEN PATIENT CABLE PINS OR BLOOD IN THE RED HANDLE WERE OBSERVED ON THE RETURNED PUMP. THE CAUSE OF THE PUMP STOP WAS NOT DETERMINED SINCE DATA LOGS WERE NOT RECOVERED AND THE RETURNED PUMP COULD NOT BE DETECTED ON ANY CONSOLE DURING INVESTIGATION. Manufacturer Narrative: PER ABIOMED'S MEDICAL SAFETY OFFICER REVIEW, THE IMPELLA 5.5 DID NOT WORK/STOPPED AND IT CANNOT BE SAID THAT THE IMPELLA 5.5 PUMP DID NOT CONTRIBUTE TO THE PATIENTS OUTCOME. THE IMPELLA PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: DIRECT AORTIC INSERTION ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA 5.5 WITH SMARTASSIST SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21824909&pc=OZD | 1220648-2025-27400 | 2025-03-13 04:00 | Death | ABIOMED, INC. | 2025-04-11 04:00 | OZD | IMPELLA 5.5 | Pumping Stopped | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) NOTED WITH POST-CARDIOTOMY CARDIOGENIC SHOCK (PCCS)/LOW CARDIAC OUTPUT SYNDROME (LCOS) WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PUMP STOPPED DUE TO MOTOR CURRENT SPIKE. THE PUMP WAS EVENTUALLY RESTARTED ON P-1 SETTING AND WAS WORKED UP TO P-3. THERE WAS NO REPORT OF PATIENT HARM.Event Description: ADDITIONAL INFORMATION RECEIVED THAT THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE DATE OF EXPLANT IS UNKNOWN THE FIELD WAS LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF THE ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. ENTERING CARDIAC OUTPUT (IMPELLA CP WITH SMART ASSIST), SECTION: DIRECT AORTIC INSERTION: ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED: ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF TEMPORARY PUMP STOP HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INVESTIGATION. DATA LOGS WERE NOT PROVIDED FOR THE CASE. NO PHYSICAL DAMAGE WAS NOTED ON THE RETURNED PRODUCT. UPON FURTHER INVESTIGATION, A BIOMETER WAS FOUND WRAPPED AROUND THE IMPELLER. THE PUMP WAS SOCKED IN A BUCKET OF WATER FOR 15 MINUTES AND WAS THEN ABLE TO START, WITH OBSERVED SATURATED PUMP FLOW. THE CAUSE OF THE TEMPORARY PUMP STOP ISSUE WAS NOT DETERMINED SINCE ISSUE WAS NOT REPRODUCE ON RETURNED PRODUCT AND DATA LOGS WERE NOT AVAILABLE TO REVIEW. B1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM, BASED ON ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B2 REVISED TO REFLECT DEATH AND DATE OF DEATH BASED ON ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B5 REVISED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. D6B EXPLANT DATE ADDED. H1 REVISED TO REFLECT DEATH. H6 REVISED BASED ON ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21824953&pc=DSQ | 2916596-2025-02138 | 2025-03-21 04:00 | Death | THORATEC CORPORATION | 2025-04-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Malaise; Respiratory Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPERIENCED SHORTNESS OF BREATH, MALAISE, AND FATIGUE. THE PATIENT WAS TRANSITIONED TO COMFORT CARE. THE PATIENT PASSED AWAY SECONDARY TO RESPIRATORY FAILURE. THE CAUSE OF RESPIRATORY FAILURE WAS DECOMPENSATED HEART FAILURE. NO AUTOPSY WAS PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH AND RESPIRATORY FAILURE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825035&pc=OZD | 1220648-2025-27402 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-04-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT A PATIENT EXPIRED AND IT WAS NOTED THAT THE CAUSE OF DEATH WAS INFECTION AND INTESTINAL ISCHEMIA. THE RELATIONSHIP BETWEEN THE CAUSE OF DEATH AND THE IMPELLA IS UNKNOWN. THE PHYSICIAN NOTED IT IS UNLIKELY THE INFECTION IS RELATED TO THE IMPELLA INSERTION AS IT OCCURRED TO EARLY. AFTER ESCALATION TO THE IMPELLA 5.5 AND UPON RETURNING TO THE ROOM, THE PATIENT DEVELOPED A FEVER. THE CAUSE OF THE INFECTION IS UNKNOWN. INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT DIED AND THE PUMP WAS REMOVED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825047&pc=NVZ | 3004209178-2025-06298 | 2025-01-16 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-11 04:00 | NVZ | ADVISA MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Twitching | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE NURSE NOTICED TWITCHING AROUND THE PACEMAKER AREA OF THE PATIENT AND WONDERED IF THE DEVICE WAS MALFUNCTIONING. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825073&pc=DTB | 2649622-2025-09815 | 2025-01-16 05:00 | Death | MPRI | 2025-04-11 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE NURSE NOTICED TWITCHING AROUND THE PACEMAKER AREA OF THE PATIENT AND WONDERED IF THE DEVICE WAS MALFUNCTIONING. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825082&pc=DTB | 2649622-2025-09817 | 2025-01-16 05:00 | Death | MPRI | 2025-04-11 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE NURSE NOTICED TWITCHING AROUND THE PACEMAKER AREA OF THE PATIENT AND WONDERED IF THE DEVICE WAS MALFUNCTIONING. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825160&pc=DSQ | 2916596-2025-02015 | 2025-03-10 04:00 | Death | THORATEC CORPORATION | 2025-04-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fatigue; Muscle Weakness; Sepsis; Septic Shock; Weight Changes; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 FROM BEING TOTALLY SEPTIC, DUE TO A DRIVELINE INFECTION. THE PATIENT WAS PLACED ON COMFORT MEASURES ONLY (CMO). THE ONSET OF THE INFECTION WAS ON (B)(6) 2025, WHICH WAS IDENTIFIED BY FATIGUE, WEAKNESS AND WEIGHT LOSS. THE DEVICE OPERATED AS EXPECTED AND WOULD NOT RETURN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825199&pc=NPT | 2025587-2025-02572 | 2023-02-01 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-11 04:00 | NPT | EVOLUT PRO PLUS VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information; Patient Device Interaction Problem | Bacterial Infection; Disseminated Intravascular Coagulation (DIC); Endocarditis; Septic Shock; Heart Failure/Congestive Heart Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THROMBOCYTOPENIA WAS NOTED. APPROXIMATELY 2 MONTHS LATER, THE PATIENT WAS HOSPITALIZED AT A DIFFERENT HOSPITAL FOR DETAILED EXAMINATION. A DIAGNOSIS OF IMMUNE THROMBOCYTOPENIC PURPURA WAS CONFIRMED; HOWEVER, WHEN INITIATION OF TREATMENT WAS ATTEMPTED, THE PATIENT DEVELOPED SEPTIC SHOCK DUE TO STREPTOCOCCUS BOVIS. HEART FAILURE AND DISSEMINATED INTRAVASCULAR COAGULATION WERE ALSO NOTED. THE PATIENT PASSED AWAY NEARLY ONE MONTH FOLLOWING THE HOSPITALIZATION. IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. PER THE PHYSICIAN, COMPLICATIONS OF INFECTIOUS ENDOCARDITIS COULD NOT BE RULED OUT AS CONTRIBUTING FACTORS.Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THROMBOCYTOPENIA WAS NOTED. APPROXIMATELY 2 MONTHS LATER, THE PATIENT WAS HOSPITALIZED AT A DIFFERENT HOSPITAL FOR DETAILED EXAMINATION. A DIAGNOSIS OF IMMUNE THROMBOCYTOPENIC PURPURA WAS CONFIRMED; HOWEVER, WHEN INITIATION OF TREATMENT WAS ATTEMPTED, THE PATIENT DEVELOPED SEPTIC SHOCK DUE TO STREPTOCOCCUS BOVIS. HEART FAILURE AND DISSEMINATED INTRAVASCULAR COAGULATION WERE ALSO NOTED. THE PATIENT PASSED AWAY NEARLY ONE MONTH FOLLOWING THE HOSPITALIZATION. IT IS UNKNOWN WHETHER THE DEVICE OR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. PER THE PHYSICIAN, COMPLICATIONS OF INFECTIOUS ENDOCARDITIS COULD NOT BE RULED OUT AS CONTRIBUTING FACTORS. Manufacturer Narrative: B3: DATE IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B2. OUTCOME ATTRIBUTED TO ADVERSE EVENT D. SUSPECT MEDICAL DEVICE FULL UNIQUE IDENTIFIER (UDI) # MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825279&pc=LYJ | 1644487-2025-00399 | 2025-02-25 05:00 | Death | CYBERONICS - HOUSTON | 2025-04-11 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DURING THEIR SLEEP DUE TO A SEIZURE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY; DEFECTS¿ OR MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825535&pc=NPT | 2025587-2025-02573 | 2024-04-28 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-11 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiogenic Shock | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WENT DIED FROM CARDIOGENIC SHOCK AND ORGAN FAILURE THAT DID NOT IMPROVE AFTER SURGERY. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825644&pc=DSQ | 2916596-2025-02062 | 2025-03-16 04:00 | Death | THORATEC CORPORATION | 2025-04-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOPULMONARY FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REVISION A IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825758&pc=DSQ | 2916596-2025-02063 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-04-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO DRIVELINE INFECTION. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825890&pc=CBI | 3012307300-2025-04199 | 2025-03-01 05:00 | Death | SMITHS HEALTHCARE MANUFACTURING S.A. DE C.V. | 2025-04-11 04:00 | CBI | ENDOBRONCHIAL TUBE | No Flow | Hypoventilation | K953483 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS UNABLE TO BE VENTILATED, AND THE PATIENT EXPIRED. THE EVENT OCCURRED WHILE AT A FACILITY. NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: B3: DATE OF EVENT IS UNKNOWN. B2: DATE OF DEATH IS UNKNOWN. NO DEVICE WAS RECEIVED FOR INVESTIGATION. THEREFORE, THE REPORTED COMPLAINT IS UNABLE TO BE CONFIRMED. IF THE DEVICE IS RECEIVED, THE INVESTIGATION WILL BE RE-OPENED FOR FURTHER EVALUATIONS. DEVICE HISTORY RECORD (DHR) REVIEW COULD NOT BE COMPLETED AS NO LOT NUMBER WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21825952&pc=NPT | 2025587-2025-02576 | 2025-03-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-11 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Respiratory Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE RESPIRATORY FAILURE; HOWEVER, THE CAUSE OF THE RESPIRATORY FAILURE WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826114&pc=OUT | 2029214-2025-00959 | 2024-04-18 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-11 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Aneurysm; Nervous System Injury; Ruptured Aneurysm |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TREATMENT OF PEDIATRIC INTRACRANIAL ANEURYSMS: INSTITUTIONAL CASE SERIES AND SYSTEMATIC LITERATURE REVIEW. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE 5 PATIENTS DIED DUE TO SEQUELAE OF THEIR ANEURYSMS; 2 PATIENTS DIED DUE TO FRANK RE-RUPTURE. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: ANEURYSMAL RUPTURE OCCURRED IN 19 (57.5%) PATIENTS, OF WHICH 8 (24.2%) WERE CATEGORIZED AS HUNT-HESS GRADES IV OR V. ANEURYSM RECURRENCE OR RE-RUPTURE OCCURRED IN FIVE (15.2%) PATIENTS, AND 5 PATIENTS (15.2%) DIED DUE TO SEQUELAE OF THEIR ANEURYSMS. FRANK RE-RUPTURE OCCURRED IN FOUR (12.1%) PATIENTS (TWO CLIP LIGATIONS, ONE PED, AND ONE STENT-ASSISTED COIL EMBOLIZATION), WITH TWO OF THESE EVENTS RESULTING IN PATIENT DEATH. TWENTY-ONE PATIENTS (63.6%) HAD A GOOD OUTCOME (MODIFIED RANKIN SCALE SCORE 0¿2) ON LAST FOLLOW UP. TWO PATIENTS (16.7%) WHO UNDERWENT ENDOVASCULAR TREATMENT HAD UNFAVORABLE OUTCOME. AT THE TIME OF LATEST FOLLOW-UP, 12 (36.4%) PATIENTS HAD A MODIFIED RANKIN SCORE (MRS) OF 0, SIX (18.2%) HAD A SCORE OF 1, THREE (9.1%) HAD A SCORE OF 2, THREE (9.1%) HAD A SCORE OF 3, FOUR (9.1%) HAD A SCORE OF 4, AND SIX (18.2%) PATIENTS WERE DECEASED. TEN (30.3%) PATIENTS REQUIRED MORE THAN ONE SURGERY, OF WHICH THREE (30%) WERE UNPLANNED (ONE REPEAT CLIPPING DUE TO ANEURYSMAL RECURRENCE, ONE COIL ASSISTED VESSEL SACRIFICE DUE TO ANEURYSM RE RUPTURE AFTER STENT-ASSISTED COILING, AND ONE BALLOON ANGIOPLASTY DUE TO AN ENDOLEAK AFTER PLACEMENT OF A PED). TWO (6.1%) PATIENTS REQUIRED A THIRD INTERVENTION TO ADEQUATELY TREAT THEIR LESIONS. POSTOPERATIVELY, THERE WERE SIX (18.2%) PATIENTS WITH NEW DEFICITS, INCLUDING VISUAL FIELD IMPAIRMENT, VOCAL CORD PARALYSIS, AND EXT REMITY/HEMIBODY WEAKNESS. DEFICITS WERE TEMPORARY IN FOUR (12.1%) PATIENTS AND PERMANENT IN TWO (6.0%) PATIENTS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THERE WERE NO MEDTRONIC PRODUCTS INVOLVED. Manufacturer Narrative: G2: CITATION: BRANDEL, M. G., PLONSKER, J. H., RENNERT, R. C., PRODUTURI, G., SARIPELLA, M., WALI, A. R., MCCANN, C., RAVINDRA, V. M., SANTIAGO-DIEPPA, D. R., PANNELL, J. S., STEINBERG, J. A., KHALESSI, A. A., <(>&<)> LEVY, M. L. (2024). TREATMENT OF PEDIATRIC INTRACRANIAL ANEURYSMS: INSTITUTIONAL CASE SERIES AND SYSTEMATIC LITERATURE REVIEW. CHILD¿S NERVOUS SYSTEM, 40(8), 2419¿2429. HTTPS://D OI.ORG/10.1007/S00381-024-06384-X. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826564&pc=DQY | 2029214-2025-00962 | 2023-11-02 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-11 04:00 | DQY | REFLEX CATHETER | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ischemia Stroke | K110055 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: 5F OR 6F NAVIEN INTERMEDIATE CATHETER, MARKSMAN, PHENOM 27, PIPELINE EMBOLIZATION DEVICE. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE: TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION, 2 PATIENTS DIED OF I NTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION, AND 1 DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. ANOTHER 2 PATIENTS DIED OF UNCORRELATED HEART FAILURE AND PNEUMONIA. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION. TWO PATIENTS DIED OF INTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION. ONE PATIENT DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. FOUR SURVIVING PATIENTS WERE LOST TO RADIOLOGIC IMAGING FOLLOW-UP DUE TO POOR STATUS. SIGNIFICANTLY HIGHER RATES OF COMPLETE OCCLUSION OF VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS (VBDA) WERE SHOWN IN THE FLOW DIVERTER (FD) GROUP THAN IN THE LOW-PROFILE BRAIDED STENT (LPBS) GROUP (LPBS, 10%, VERSUS FD, 41.2%; P ¼ .028). 12 PATIENTS (60%) IN THE LPBS GROUP AND 7 PATIENTS (41.2) IN THE FD GROUP EXPERIENCED NEAR-COMPLETE OCCLUSION. 6 PATIENTS (30%) IN THE LPBS GROUP AND 3 PATIENTS (17.6%) IN THE FD GROUP EXPERIENCED INCOMPLETE OCCLUSION. 11 PATIENTS (44%) IN THE LPBS GROUP AND 9 PATIENTS (40.9%) IN THE FD GROUP EXPERIENCED POOR CLINICAL OUTCOME (MRS 2). 12 PATIENTS (48%) IN THE LPBS GROUP EXPERIENCED AGGRAVATED SYMPTOMS. 8 PATIENTS (36.4%) IN THE FD GROUP EXPERIENCED AGGRAVATED SYMPTOMS. PROCEDURAL COMPLICATIONS INCLUDED ISCHEMIC COMPLICATIONS, WHICH ACCOUNTED FOR MOST DEFICITS, AND HEMORRHAGIC COMPLICATIONS. THERE WERE NEW NEUROLOGIC DEFICITS DUE TO PROCEDURAL COMPLICATIONS: HEMORRHAGIC; ONE PATIENT (4%) IN THE LPBS GROUP, 2 PATIENTS (9.1%) IN THE FD GROUP. ISCHEMIC: 5 PATIENTS (20%) IN THE LPBS GROUP, 9 PATIENTS (40.9%) IN THE FD GROUP. 6 (24%) OF PATIENTS IN THE LBS GROUP EXPERIENCED MILD-TO-SEVERE-DISABILITY, 6 (27.3%) IN THE FD GROUP. 10 PATIENTS (50%) IN THE LPBS AND 5 PATIENTS (29.4%) IN THE FD GROUP EXPERIENCED PROGRESSION OF MASS EFFECT. ADDITIONALLY, THE STUDY NOTED THAT BRAINSTEM COMPRESSION RESULTING FROM VBDAS POSED A SIGNIFICANT THERAPEUTIC CHALLENGE, AND WHILE NO SIGNIFICANT DIFFERENCES WERE SHOWN, THE LPBS PROVED TO BE LESS EFFECTIVE IN PREVENTING PROGRESSION OF MASS EFFECT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION: JI, Z., HE, C., LI, J., GENG, J., HU, P., LI, G., & ZHANG, H. (2023). SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. ORIGINAL RESEARCH NEUROINTERVENTION, 45(2), 176¿182. HTTPS://DOI.ORG/10.3174/AJNR.A8091. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826586&pc=KRA | 2029214-2025-00963 | 2023-11-02 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-11 04:00 | KRA | MARKSMAN | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ischemia Stroke | K091559 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: 5F OR 6F NAVIEN INTERMEDIATE CATHETER, MARKSMAN, PHENOM 27, PIPELINE EMBOLIZATION DEVICE. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE: TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION, 2 PATIENTS DIED OF INTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION, AND 1 DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. ANOTHER 2 PATIENTS DIED OF UNCORRELATED HEART FAILURE AND PNEUMONIA. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION. TWO PATIENTS DIED OF INTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION. ONE PATIENT DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. FOUR SURVIVING PATIENTS WERE LOST TO RADIOLOGIC IMAGING FOLLOW-UP DUE TO POOR STATUS. SIGNIFICANTLY HIGHER RATES OF COMPLETE OCCLUSION OF VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS (VBDA) WERE SHOWN IN THE FLOW DIVERTER (FD) GROUP THAN IN THE LOW-PROFILE BRAIDED STENT (LPBS) GROUP (LPBS, 10%, VERSUS FD, 41.2%; P ¼ .028). 12 PATIENTS (60%) IN THE LPBS GROUP AND 7 PATIENTS (41.2) IN THE FD GROUP EXPERIENCED NEAR-COMPLETE OCCLUSION. 6 PATIENTS (30%) IN THE LPBS GROUP AND 3 PATIENTS (17.6%) IN THE FD GROUP EXPERIENCED INCOMPLETE OCCLUSION. 11 PATIENTS (44%) IN THE LPBS GROUP AND 9 PATIENTS (40.9%) IN THE FD GROUP EXPERIENCED POOR CLINICAL OUTCOME (MRS 2). 12 PATIENTS (48%) IN THE LPBS GROUP EXPERIENCED AGGRAVATED SYMPTOMS. 8 PATIENTS (36.4%) IN THE FD GROUP EXPERIENCED AGGRAVATED SYMPTOMS. PROCEDURAL COMPLICATIONS INCLUDED ISCHEMIC COMPLICATIONS, WHICH ACCOUNTED FOR MOST DEFICITS, AND HEMORRHAGIC COMPLICATIONS. THERE WERE NEW NEUROLOGIC DEFICITS DUE TO PROCEDURAL COMPLICATIONS: HEMORRHAGIC; ONE PATIENT (4%) IN THE LPBS GROUP, 2 PATIENTS (9.1%) IN THE FD GROUP. ISCHEMIC: 5 PATIENTS (20%) IN THE LPBS GROUP, 9 PATIENTS (40.9%) IN THE FD GROUP. 6 (24%) OF PATIENTS IN THE LBS GROUP EXPERIENCED MILD-TO-SEVERE-DISABILITY, 6 (27.3%) IN THE FD GROUP. 10 PATIENTS (50%) IN THE LPBS AND 5 PATIENTS (29.4%) IN THE FD GROUP EXPERIENCED PROGRESSION OF MASS EFFECT. ADDITIONALLY, THE STUDY NOTED THAT BRAINSTEM COMPRESSION RESULTING FROM VBDAS POSED A SIGNIFICANT THERAPEUTIC CHALLENGE, AND WHILE NO SIGNIFICANT DIFFERENCES WERE SHOWN, THE LPBS PROVED TO BE LESS EFFECTIVE IN PREVENTING PROGRESSION OF MASS EFFECT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION:JI, Z., HE, C., LI, J., GENG, J., HU, P., LI, G., <(>&<)> ZHANG, H. (2023). SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. ORIGINAL RESEARCH NEUROINTERVENTION, 45(2), 176¿182. HTTPS://DOI.ORG/10.3174/AJNR.A8091. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826599&pc=OUT | 2029214-2025-00965 | 2023-11-02 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-11 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ischemia Stroke |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: 5F OR 6F NAVIEN INTERMEDIATE CATHETER, MARKSMAN, PHENOM 27, PIPELINE EMBOLIZATION DEVICE. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE : TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION, 2 PATIENTS DIED OF I NTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION, AND 1 DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. ANOTHER 2 PATIENTS DIED OF UNCORRELATED HEART FAILURE AND PNEUMONIA. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION. TWO PATIENTS DIED OF INTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION. ONE PATIENT DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. FOUR SURVIVING PATIENTS WERE LOST TO RADIOLOGIC IMAGING FOLLOW-UP DUE TO POOR STATUS. SIGNIFICANTLY HIGHER RATES OF COMPLETE OCCLUSION OF VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS (VBDA) WERE SHOWN IN THE FLOW DIVERTER (FD) GROUP THAN IN THE LOW-PROFILE BRAIDED STENT (LPBS) GROUP (LPBS, 10%, VERSUS FD, 41.2%; P ¼ .028). 12 PATIENTS (60%) IN THE LPBS GROUP AND 7 PATIENTS (41.2) IN THE FD GROUP EXPERIENCED NEAR-COMPLETE OCCLUSION. 6 PATIENTS (30%) IN THE LPBS GROUP AND 3 PATIENTS (17.6%) IN THE FD GROUP EXPERIENCED INCOMPLETE OCCLUSION. 11 PATIENTS (44%) IN THE LPBS GROUP AND 9 PATIENTS (40.9%) IN THE FD GROUP EXPERIENCED POOR CLINICAL OUTCOME (MRS 2). 12 PATIENTS (48%) IN THE LPBS GROUP EXPERIENCED AGGRAVATED SYMPTOMS. 8 PATIENTS (36.4%) IN THE FD GROUP EXPERIENCED AGGRAVATED SYMPTOMS. PROCEDURAL COMPLICATIONS INCLUDED ISCHEMIC COMPLICATIONS, WHICH ACCOUNTED FOR MOST DEFICITS, AND HEMORRHAGIC COMPLICATIONS. THERE WERE NEW NEUROLOGIC DEFICITS DUE TO PROCEDURAL COMPLICATIONS: HEMORRHAGIC; ONE PATIENT (4%) IN THE LPBS GROUP, 2 PATIENTS (9.1%) IN THE FD GROUP. ISCHEMIC: 5 PATIENTS (20%) IN THE LPBS GROUP, 9 PATIENTS (40.9%) IN THE FD GROUP. 6 (24%) OF PATIENTS IN THE LBS GROUP EXPERIENCED MILD-TO-SEVERE-DISABILITY, 6 (27.3%) IN THE FD GROUP. 10 PATIENTS (50%) IN THE LPBS AND 5 PATIENTS (29.4%) IN THE FD GROUP EXPERIENCED PROGRESSION OF MASS EFFECT. ADDITIONALLY, THE STUDY NOTED THAT BRAINSTEM COMPRESSION RESULTING FROM VBDAS POSED A SIGNIFICANT THERAPEUTIC CHALLENGE, AND WHILE NO SIGNIFICANT DIFFERENCES WERE SHOWN, THE LPBS PROVED TO BE LESS EFFECTIVE IN PREVENTING PROGRESSION OF MASS EFFECT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION: JI, Z., HE, C., LI, J., GENG, J., HU, P., LI, G., <(>&<)> ZHANG, H. (2023). SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. ORIGINAL RESEARCH NEUROINTERVENTION, 45(2), 176¿182. HTTPS://DOI.ORG/10.3174/AJNR.A8091. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826600&pc=KRA | 2029214-2025-00964 | 2023-11-02 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-11 04:00 | KRA | CATHERA | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ischemia Stroke | K151638 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: 5F OR 6F NAVIEN INTERMEDIATE CATHETER, MARKSMAN, PHENOM 27, PIPELINE EMBOLIZATION DEVICE. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSES OF DEATH WERE : TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION, 2 PATIENTS DIED OF I NTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION, AND 1 DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. ANOTHER 2 PATIENTS DIED OF UNCORRELATED HEART FAILURE AND PNEUMONIA. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: TWO PATIENTS DIED OF BRAINSTEM FUNCTION FAILURE RESULTING FROM BRAINSTEM INFARCTION. TWO PATIENTS DIED OF INTRACEREBRAL HEMORRHAGE RESULTING FROM THROMBOLYSIS FROM BRAINSTEM INFARCTION. ONE PATIENT DIED DIRECTLY FROM INTRACRANIAL HEMORRHAGE. FOUR SURVIVING PATIENTS WERE LOST TO RADIOLOGIC IMAGING FOLLOW-UP DUE TO POOR STATUS. SIGNIFICANTLY HIGHER RATES OF COMPLETE OCCLUSION OF VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS (VBDA) WERE SHOWN IN THE FLOW DIVERTER (FD) GROUP THAN IN THE LOW-PROFILE BRAIDED STENT (LPBS) GROUP (LPBS, 10%, VERSUS FD, 41.2%; P ¼ .028). 12 PATIENTS (60%) IN THE LPBS GROUP AND 7 PATIENTS (41.2) IN THE FD GROUP EXPERIENCED NEAR-COMPLETE OCCLUSION. 6 PATIENTS (30%) IN THE LPBS GROUP AND 3 PATIENTS (17.6%) IN THE FD GROUP EXPERIENCED INCOMPLETE OCCLUSION. 11 PATIENTS (44%) IN THE LPBS GROUP AND 9 PATIENTS (40.9%) IN THE FD GROUP EXPERIENCED POOR CLINICAL OUTCOME (MRS 2). 12 PATIENTS (48%) IN THE LPBS GROUP EXPERIENCED AGGRAVATED SYMPTOMS. 8 PATIENTS (36.4%) IN THE FD GROUP EXPERIENCED AGGRAVATED SYMPTOMS. PROCEDURAL COMPLICATIONS INCLUDED ISCHEMIC COMPLICATIONS, WHICH ACCOUNTED FOR MOST DEFICITS, AND HEMORRHAGIC COMPLICATIONS. THERE WERE NEW NEUROLOGIC DEFICITS DUE TO PROCEDURAL COMPLICATIONS: HEMORRHAGIC; ONE PATIENT (4%) IN THE LPBS GROUP, 2 PATIENTS (9.1%) IN THE FD GROUP. ISCHEMIC: 5 PATIENTS (20%) IN THE LPBS GROUP, 9 PATIENTS (40.9%) IN THE FD GROUP. 6 (24%) OF PATIENTS IN THE LBS GROUP EXPERIENCED MILD-TO-SEVERE-DISABILITY, 6 (27.3%) IN THE FD GROUP. 10 PATIENTS (50%) IN THE LPBS AND 5 PATIENTS (29.4%) IN THE FD GROUP EXPERIENCED PROGRESSION OF MASS EFFECT. ADDITIONALLY, THE STUDY NOTED THAT BRAINSTEM COMPRESSION RESULTING FROM VBDAS POSED A SIGNIFICANT THERAPEUTIC CHALLENGE, AND WHILE NO SIGNIFICANT DIFFERENCES WERE SHOWN, THE LPBS PROVED TO BE LESS EFFECTIVE IN PREVENTING PROGRESSION OF MASS EFFECT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION:JI, Z., HE, C., LI, J., GENG, J., HU, P., LI, G., & ZHANG, H. (2023). SAFETY AND EFFICACY OF LOW-PROFILE BRAIDED STENTS VERSUS FLOW DIVERTERS IN THE RECONSTRUCTIVE TECHNIQUE IN THE TREATMENT OF PATIENTS WITH VERTEBROBASILAR DOLICHOECTASIA ANEURYSMS: A COHORT OF 47 PATIENTS WITH LONG-TERM FOLLOW-UP. ORIGINAL RESEARCH NEUROINTERVENTION, 45(2), 176¿182. HTTPS://DOI.ORG/10.3174/AJNR.A8091. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826826&pc=DSQ | 2916596-2025-02016 | 2025-03-18 04:00 | Death | THORATEC CORPORATION | 2025-04-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem | Chest Pain; Ischemia Stroke; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR MESSAGED TO REPORT THAT THE PATIENT WAS "GOING TO THE OPERATING ROOM (OR) TO REMOVE LEFT CORONARY CUSP CLOT". AFTER THE VAD COORDINATOR STATED THE PATIENT HAD A CLOT AND WAS GOING FOR A COMPUTED TOMOGRAPHY (CT) SCAN. NO FURTHER COMMUNICATION WAS RECEIVED WHEN IT WAS ASKED HOW THE PATIENT WAS DOING. THERE WERE NO CHANGES TO PATIENT CONDITION THAT MAY HAVE CONTRIBUTED. THE PATIENT'S INTERNATIONAL NORMALISED RATIO (INR) WAS 1.5 AND WAS BEING BRIDGED WITH ENOXAPARIN. THERE WERE OCCASIONAL LOW FLOWS THAT OCCURRED. THE PATIENT HAD ONE SIDED DEFICITS AND SLURRED SPEECH. THE PATIENT PLAN OF CARE WAS TO WATCH AND WAIT. THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR STATED THAT THE PATIENT EXPERIENCED CHEST PAIN AND WAS TAKEN TO THE CATH LAB. THE CATH SHOWED LEFT MAIN OCCLUSION AND THE TEAM WENT TO OPEN IT UP. STROKE SYMPTOMS STARTED TO APPEAR WHILE IN THE CATH LAB. IT WAS THEN NOTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT HAD A CEREBROVASCULAR ACCIDENT (CVA) THAT WAS CATASTROPHIC. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. AN ISCHEMIC STROKE OCCURRED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORT OF LOW FLOW ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE SUBMITTED FOR EVALUATION. A SPECIFIC CAUSE FOR THE REPORTED ALARMS COULD NOT BE CONCLUSIVELY DETERMINED HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, THROMBOEMBOLISM, AND DEATH. SECTION 6 OF THE IFU ENTITLED ¿PATIENT CARE AND MANAGEMENT¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 ALSO PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR VALUES. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21826905&pc=DSQ | 2916596-2025-02267 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-04-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Loss of consciousness; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN AORTIC ROOT THROMBUS WITH AN ABRUPT LOSS OF CONSCIOUSNESS AND LOSS OF ALL LEFT VENTRICULAR ASSIST DEVICE (LVAD) FLOW. THE PUMP SPEED WAS SET TO 5400 REVOLUTIONS PER MINUTE (RPMS) WHEN FLOW WAS LOST AND THE PATIENT LOST CONSCIOUSNESS. DURING ARREST, THE PUMP SPEED WAS LOWERED TO 4000 RPMS IN AN ATTEMPT TO INCREASE LEFT VENTRICLE (LV) FILLING TIME AND REGAIN FLOW, WHICH WAS UNSUCCESSFUL, AND FLOW REMAINED AT 0 LITERS PER MINUTE (LPM) WITH POWER 2.4 WATTS. ON AN ECHOCARDIOGRAM OF THE FULL LEFT VENTRICLE, THE AORTIC ROOT THROMBUS WAS UNABLE TO BE VISUALIZED. EXTRACORPOREAL CARDIOPULMONARY RESUSCITATION (ECPR) VA-ECMO (VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION) CANNULATION WAS DONE HOWEVER LVAD FLOW NEVER RECOVERED. THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE PUMP WOULD NOT RETURN. THE DEVICE OPERATED AS EXPECTED PRIOR TO THE EVENT. IT WAS NOTED THAT ANTICOAGULATION HAD BEEN WITHHELD ON (B)(6) 2025 DUE TO STROKE AND HEMORRHAGE (MANUFACTURER REPORT NUMBER 2916596-2025-02441).Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD A KNOWN CHRONIC DRIVELINE INFECTION WITH MULTIPLE EPISODES OF NON-COMPLIANCE WITH CHRONIC SUPPRESSIVE ANTIBIOTICS AND ANTICOAGULATION (MFR #: 2916596-2021-02164). Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. THE REPORT OF PERIPHERAL THROMBOEMBOLISM COULD NOT BE CONFIRMED THROUGH THIS INVESTIGATION AS NO IMAGES WERE PROVIDED, AND THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE REPORT OF LOW FLOW EVENTS WAS CONFIRMED VIA THE SUBMITTED LOG FILES; HOWEVER, A SPECIFIC CAUSE FOR THE LOW FLOW EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE CONTROLLER EVENT LOG FILES CONTAINED EVENTS FROM 26MAR2025. LOW FLOW ALARMS WERE ACTIVE THROUGHOUT THE DURATION OF THE LOG FILE DUE TO THE ESTIMATED FLOW BEING 0.0 LPM. ADDITIONALLY, LOW SPEED ADVISORY ALARMS WERE ACTIVE DUE TO THE PATIENT SPEED BEING SET BELOW THE LOW SPEED LIMIT. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO OPERATE AS INTENDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING THROMBOEMBOLISM, NEUROLOGIC DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. IT ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AND NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATIONS. THIS SECTION ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINES ALL SYSTEM ALARM CONDITIONS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827014&pc=OZD | 1220648-2025-27411 | 2025-03-17 04:00 | Death | ABIOMED, INC. | 2025-04-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PUMP POSITION BECAME LOW, AND THE PATIENT EXPERIENCED EXTREMELY BLOODY URINE. THE NEXT DAY, CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED. AFTERWARDS, THE PUMP WAS REPOSITIONED AND PULLED BACK WHICH HELPED SLOWLY RESOLVE THE HEMOLYSIS. ADDITIONALLY, THE PATIENT SUFFERED A MASSIVE STROKE. DISCUSSIONS ABOUT LIMITED FURTHER OPTIONS WERE MADE DUE TO PATIENT SUFFERING SEVERE STROKE, CRRT, AND CARDIAC REQUIREMENTS NEEDED.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. PER THE MEDICAL SAFETY OFFICER THE USE OF THE IMPELLA CP DEVICE CANNOT BE DISSOCIATED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE AS AN ASSOCIATED FACTOR. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF STOKE/CVA AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE CAUSE OF THE HEMOLYSIS ISSUE WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE, BASED ON GIVEN CLINICAL DETAILS. AS THIS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED. B5 MEDICAL SAFETY OFFICER STATEMENT ADDED B2 DEATH AND DATE OF DEATH WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27411. B5 NARRATIVE ERRONEOUSLY STATED THE PATIENT OUTCOME ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27411. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27411. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27411.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827053&pc=NGV | 2135147-2025-01907 |  | Death | ABBOTT MEDICAL | 2025-04-11 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON AN UNKNOWN DATE, AN UNKNOWN SIZE AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS IMPLANTED USING AN UNKNOWN DELIVERY SYSTEM. ON AN UNKNOWN DATE, IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: N/A Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2 - DATE OF DEATH IS AN ESTIMATED DATE. Manufacturer Narrative: AN EVENT OF PATIENT DEATH WAS REPORTED. A RETURNED DEVICE INSPECTION, TO RULE OUT ANY DEVICE-RELATED CAUSES, COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE BATCH/LOT NUMBER WAS NOT PROVIDED, AND THEREFORE THE DEVICE HISTORY RECORD AND LOT-SPECIFIC SIMILAR COMPLAINT COULD NOT BE REVIEWED. BASED ON THE LIMITED INFORMATION RECEIVED, A CAUSE OF THE REPORTED DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827280&pc=NPT | 2025587-2025-02596 | 2024-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-11 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS SEPSIS. Manufacturer Narrative: UPDATED: D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827286&pc=OUT | 2029214-2025-00972 | 2025-02-28 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-11 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Neuropathy; Loss of consciousness; Obstruction/Occlusion; Hemorrhagic Stroke |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG-TERM SAFETY AND EFFICACY OF THE FRED X FLOW DIVERTER FOR INTRACRANIAL ANEURYSMS: A MULTICENTER STUDY OF 154 PATIENTS. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: PIPELINE EMBOLIZATION DEVICE (PED). DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. AMONG ALL PED PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: THE PATIENT PRESENTED WITH A PARTIALLY THROMBOSED GIANT ANEURYSM IN THE POSTERIOR CAROTID WALL. SHE UNDERWENT TELESCOPING FLOW DIVERSION USING 2 FRED X STENTS AND A PED SHIELD WITH ADJUNCTIVE COIL EMBOLIZATION. DURING THE PROCEDURE, THERE WAS OCCLUSION OF THE MIDDLE CEREBRAL ARTERY, WHICH REQUIRED BALLOON ANGIOPLASTY AND STENTING. SHE RECEIVED A BOLUS OF AGGRASTAT (MEDICURE PHARMA) IN ADDITION TO BEING PRELOADED WITH BRILINTA (ASTRAZENECA) SINCE SHE DID NOT RESPOND TO PLAVIX (SANOFI). SHE BECAME NONRESPONSIVE AND DEVELOPED A THIRD CRANIAL NERVE PALSY DURING THE POSTOPERATIVE PERIOD. HEAD CT SHOWED AN INTRACEREBRAL HEMORRHAGE (ICH), AND SHE WAS TAKEN TO THE OPERATING ROOM FOR DECOMPRESSIVE HEMICRANIECTOMY. HER NEUROLOGICAL EXAMINATION DECLINED OVER THE COURSE OF HER HOSPITAL STAY, AND SHE WAS DECLARED BRAIN-DEAD ON POSTOPERATIVE DAY 8. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION: AUTHORS: ROY, J.M., EL NAAMANI, K., AMARAVADI, C., MAJMUNDAR, S., MOUCHTOURIS, N., PAUL, A.R., FIELD, N.C., ZHANG, Y., BURKHARDT, J.K., KUHN, A.L., PURI, A., LEVY, E., GOOCH, R., JABBOUR, P., ROSENWASSER, R.H. LONG-TERM SAFETY AND EFFICACY OF THE FRED X FLOW DIVERTER FOR INTRACRANIAL ANEURYSMS: A MULTICENTER STUDY OF 154 PATIENTS. JOURNAL OF NEUROSURGERY 2025. DOI: 10.3171/2024.10.JNS241233 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827301&pc=MGB | 2024168-2025-03755 | 2025-03-19 04:00 | Death | ABBOTT VASCULAR INC. | 2025-04-11 04:00 | MGB | PERCLOSE¿ PROSTYLE¿ | Improper or Incorrect Procedure or Method; Obstruction of Flow; Deformation Due to Compressive Stress; Difficult to Advance | Hemorrhage/Bleeding; Pain; Shock; Perforation of Vessels; Vascular Dissection; Multiple Organ Failure | P960043 |  | 1 | Event Description: IT WAS REPORTED THAT THIS WAS AN ARTERIOTOMY CLOSURE OF THE SEVERELY CALCIFIED RIGHT COMMON FEMORAL ARTERY USING A PROSTYLE DEVICE AFTER A CORONARY ANGIOGRAM INTERVENTIONAL PROCEDURE USING A 6F SHEATH. REPORTEDLY, STRONG RESISTANCE WAS NOTED DURING INSERTING/ADVANCEMENT OF THE PROSTYLE DEVICE AND THE PATIENT COMPLAINED OF PAIN. NO BACKFLOW WAS OBSERVED FROM THE MARKER TUBE OF THE PROSTYLE, SUGGESTING IT WAS NOT INSERTED UP TO THE MARKER PORT. THE DEVICE WAS REMOVED, AND MANUAL COMPRESSION WAS APPLIED TO ACHIEVE HEMOSTASIS. TWO HOURS LATER, THE PATIENT WENT INTO SHOCK, AND THE PATIENT DIED OF MULTIPLE ORGAN FAILURE. A CT SCAN REVEALED BLEEDING NEAR THE PUNCTURE SITE. AFTER REVIEWING THE SITUATION AT THE HOSPITAL, IT WAS FOUND THAT THE PUNCTURE SITE HAD SEVERE CALCIFICATION AND WAS LOCATED ABOVE THE INGUINAL LIGAMENT. IT WAS FOUND THAT THE PUNCTURE NEEDLE WAS USED ABOUT 10 TIMES DURING THE PROCEDURE, WHICH MAY HAVE CAUSED SOME DAMAGE (TEARS) TO THE BLOOD VESSEL. THE ANGIOGRAPHIC IMAGES DURING THE PROCEDURE SHOWED THAT THE GUIDEWIRE WAS STILL INSERTED WHEN THE PROSTYLE WAS INSERTED WITH EXCESS FORCE, AND THE MARKER PORT AREA WAS PUSHED TO THE POINT OF FORMING Z-SHAPE, INDICATING THAT THE PUNCTURE SITE WAS TORN AND BLED DUE TO DEVICE DEFORMATION. ALTHOUGH THE CASE WAS NOT APPROPRIATE FOR THE PROSTYLE AND PROPER USAGE WAS NOT FOLLOWED, SUCH AS ACCESS SITE AND EXCESSIVE PUSHING, THE USE OF THE PROSTYLE MAY HAVE INDUCED A VASCULAR TEAR, WHICH LED TO HEMORRHAGIC SHOCK. ALTHOUGH THE PHYSICIAN INDICATED THE INCIDENT WAS NOT CAUSED BY THE PROSTYLE DEVICE, THE PHYSICIAN BELIEVES THAT THE PROSTYLE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: SUBSEQUENT TO THE INITIALLY FILED REPORT, THE FOLLOWING INFORMATION WAS RECEIVED: THE PHYSICIAN INDICATED THE INCIDENT WAS NOT CAUSED BY THE PROSTYLE DEVICE BUT RATHER INAPPROPRIATE USE OF THE DEVICE AGAINST THE INSTRUCTIONS FOR USE. REVIEW OF THE REPORTED DETAILS INDICATED THE PROSTYLE CONTRIBUTED TO THE VASCULAR TEAR/ BLEED LEADING TO HEMORRHAGIC SHOCK AND PATIENT DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: H6: MEDICAL DEVICE PROBLEM CODE: 2017 - AGAINST RESISTANCE, 2017 AND EXCESSIVE FORCE, 2017 - FAILURE TO FOLLOW STEPS / INSTRUCTIONS. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT OF SPECIFIC ISSUE. THE PATIENT EFFECTS OF HEMORRHAGE, DISSECTION, PAIN, AND DEATH ARE LISTED IN THE PROSTYLE INSTRUCTIONS FOR USE (IFU), AS POTENTIAL ADVERSE EVENT ASSOCIATED WITH USE OF THE VESSEL CLOSURE DEVICES. REPORTEDLY, STRONG RESISTANCE WAS NOTED DURING INSERTING/ADVANCEMENT OF THE PROSTYLE DEVICE AND THE PATIENT COMPLAINED OF PAIN. IT WAS FOUND THAT THE PUNCTURE SITE HAD SEVERE CALCIFICATION AND WAS LOCATED ABOVE THE INGUINAL LIGAMENT. IT SHOULD BE NOTED THAT THE PROSTYLE INSTRUCTIONS FOR USE (IFU), STATES: DO NOT ADVANCE OR WITHDRAW THE PERCLOSE PROGLIDE SMC DEVICE AGAINST RESISTANCE UNTIL THE CAUSE OF THAT RESISTANCE HAS BEEN DETERMINED. EXCESSIVE FORCE USED TO ADVANCE OR TORQUE THE PERCLOSE PROGLIDE SMC DEVICE SHOULD BE AVOIDED, AS THIS MAY LED TO SIGNIFICANT VESSEL DAMAGE AND / OR BREAKAGE OF THE DEVICE, WHICH MAY NECESSITATE INTERVENTION AND / OR SURGICAL REMOVAL OF THE DEVICE AND VESSEL REPAIR. SECTION 6.0 STATES: DO NOT USE THE PERCLOSE PROSTYLE SMCR SYSTEM IF THE PUNCTURE SITE IS LOCATED ABOVE THE MOST INFERIOR BORDER OF THE INFERIOR EPIGASTRIC ARTERY (IEA) AND / OR ABOVE THE INGUINAL LIGAMENT BASED UPON BONY LANDMARKS, SINCE SUCH A PUNCTURE SITE MAY RESULT IN A RETROPERITONEAL HEMATOMA. PERFORM A FEMORAL ANGIOGRAM TO VERIFY THE LOCATION OF THE PUNCTURE SITE. MEDICAL REVIEW WAS PERFORMED. THE REVIEW DETAILS CONCLUDED THERE ARE MULTIPLE USE ERRORS AND NOT FOLLOWING THE IFU IN THIS CASE INCLUDING USING THE PROSTYLE IN A SEVERELY CALCIFIED ARTERY, THE PUNCTURE SITE WAS LOCATED ABOVE THE INGUINAL LIGAMENT, EXCESSIVE FORCE AND MANIPULATION DUE TO SEVERE RESISTANCE UPON ADVANCEMENT WHICH PROBABLY LED TO VESSEL TEAR AND PERFORATION LEADING TO SEVERE HEMORRHAGE DEATH DUE TO MULTIPLE ORGAN FAILURE CAUSED BY HEMORRHAGIC SHOCK. THIS EVENT WAS NOT DIRECTLY DUE TO THE USE OF PROSTYLE BUT THE MULTIPLE USE ERRORS, CONTRIBUTED/CAUSED THE VESSEL TEAR OR PERFORATION LEADING TO SEVERE HEMORRHAGE AND SHOCK WHICH LED TO AN OUTCOME OF DEATH. THE PHYSICIAN ALSO CONCLUDED THAT THE INCIDENT WAS NOT CAUSED BY THE PROSTYLE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED EVENT APPEARS TO BE RELATED TO USER ERROR. MEDICAL REVIEW WAS PERFORMED AND CONCLUDED THERE ARE MULTIPLE USE ERRORS AND NOT FOLLOWING THE IFU IN THIS CASE INCLUDING USING THE PROSTYLE IN A SEVERELY CALCIFIED ARTERY, THE PUNCTURE SITE WAS LOCATED ABOVE THE INGUINAL LIGAMENT, EXCESSIVE FORCE AND MANIPULATION DUE TO SEVERE RESISTANCE UPON ADVANCEMENT WHICH PROBABLY LED TO VESSEL TEAR AND PERFORATION LEADING TO SEVERE HEMORRHAGE DEATH DUE TO MULTIPLE ORGAN FAILURE CAUSED BY HEMORRHAGIC SHOCK. THE REPORTED PATIENT EFFECTS OF EFFECTS OF HEMORRHAGE, DISSECTION, PAIN, AND DEATH ARE KNOWN INHERENT RISK OF THE PROCEDURE. THE SUBSEQUENT TREATMENTS APPEAR TO BE RELATED TO PROCEDURE CIRCUMSTANCE. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827513&pc=KNS | 2124215-2025-22699 | 2025-04-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-11 04:00 | KNS | REZUM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K191505 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED BLEEDING AND A DECLINE IN ACTIVITIES OF DAILY LIVING (ADL) AFTER UNDERGOING A WATER VAPOR THERAPY PROCEDURE BACK IN 2023, COMPOUNDED BY PRE-EXISTING MEDICAL HISTORY, AND SUBSEQUENTLY PASSED AWAY. THE DIRECT CAUSAL RELATIONSHIP IS NOT CLEAR.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED BLEEDING AND A DECLINE IN ACTIVITIES OF DAILY LIVING (ADL) AFTER UNDERGOING A WATER VAPOR THERAPY PROCEDURE BACK IN 2023, COMPOUNDED BY PRE-EXISTING MEDICAL HISTORY, AND SUBSEQUENTLY PASSED AWAY. THE DIRECT CAUSAL RELATIONSHIP IS NOT CLEAR. Manufacturer Narrative: THERE WAS NO DEVICE AVAILABLE FOR ANALYSIS; THEREFORE, NO PHYSICAL OR VISUAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. THE REPORTED PATIENT SYMPTOMS ARE A KNOWN RISK ASSOCIATED WITH THESE DEVICES AS INDICATED IN THE INSTRUCTIONS FOR USE. BASED ON THE INFORMATION AVAILABLE, A CONCLUSION CODE OF KNOWN INHERENT RISK OF DEVICE WAS ASSIGNED TO THIS INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827526&pc=NGV | 2135147-2025-01913 |  | Death | ABBOTT MEDICAL | 2025-04-11 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON AN UNKNOWN DATE, AN UNKNOWN SIZE AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS IMPLANTED USING AN UNKNOWN DELIVERY SYSTEM. ON AN UNKNOWN DATE, IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: N/A. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2 - DATE OF DEATH IS AN ESTIMATED DATE. Manufacturer Narrative: AN EVENT OF PATIENT DEATH WAS REPORTED. A RETURNED DEVICE INSPECTION, TO RULE OUT ANY DEVICE-RELATED CAUSES, COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE BATCH/LOT NUMBER WAS NOT PROVIDED, AND THEREFORE THE DEVICE HISTORY RECORD AND LOT-SPECIFIC SIMILAR COMPLAINT COULD NOT BE REVIEWED. BASED ON THE LIMITED INFORMATION RECEIVED, A CAUSE OF THE REPORTED DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827534&pc=KRA | 3010909876-2025-01386 | 2025-03-04 05:00 | Death | TRISALUS LIFE SCIENCES | 2025-04-11 04:00 | KRA | TRINAV LV INFUSION SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K230957 |  | 1 | Event Description: IT WAS REPORTED THAT A NEUROENDOCRINE TUMOR WAS TREATED WITH 3 SYRINGES OF BLAND EMBO BEADS, SIZES 100-500. THE TRINAV LV 120 LENGTH DEVICE WAS USED FOR INFUSION. NO COMPLICATIONS OBSERVED DURING CASE AND PATIENT DISPLAYED NO COMPLICATIONS 2 DAYS POST PROCEDURE. PATIENT DIED 3 DAYS POST PROCEDURE UNEXPECTEDLY AT HOME. CAUSE OF DEATH UNKNOWN. PHYSICIAN UNSURE WHY OR WHAT HAPPENED. THE PHYSICIAN WAS EXPECTING TO TREAT PATIENT MULTIPLE TIMES WITH BLAND EMBO AND WAS NOT EXPECTING SUDDEN DEATH. UPON SUBSEQUENT FOLLOW-UP INVESTIGATION WITH THE USER FACILITY, IT WAS REPORTED THAT ON POST-PROCEDURE DAY NUMBER 3, THE PATIENT BEGAN TO FEEL WEIRD AND SHAKY, SO HE WENT TO THE EMERGENCY DEPARTMENT. HE WAS NOTED TO BE HYPOTENSIVE (BP 90/60) AND HAD ELEVATED LIVER FUNCTION TESTS. TROPONIN AND OTHER CARDIAC LABS WERE NOT ORDERED. THE PATIENT WAS DISCHARGED WITH PAIN MEDICATION. PATIENT HAD A VERY HIGH DISEASE BURDEN AND CARCINOID SYNDROME PRIOR TO TREATMENT. ALTHOUGH HE HAD TAKEN A LONG-ACTING SOMATOSTATIN ANALOGUE (LANREOTIDE), HE WAS AT HIGH RISK FOR CARCINOID CRISIS DUE TO HIS HIGH TUMOR BURDEN. OCTREOTIDE WAS NOT USED FOR PROPHYLAXIS. HIS PRESENTATION TO THE EMERGENCY DEPARTMENT IS VERY CONCERNING FOR CARCINOID CRISIS AND THIS WAS NOT TREATED. IT IS THEREFORE SUSPECT THE PATIENT DIED OF CARCINOID CRISIS AT HOME. Manufacturer Narrative: USER FACILITY REPORTED THERE WAS NO EVIDENCE OR SUSPICION OF DEVICE MALFUNCTION. PHYSICIAN COMMENTED THE DEVICE FUNCTIONED WELL AND THERE WAS NO REFLUX OR ANY OTHER ALLEGED PERFORMANCE ISSUE WITH THE CATHETER. DEVICE UDI AND LOT NUMBER NOT REPORTED BY USER FACILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827573&pc=QZI | 2124215-2025-22827 | 2025-03-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-11 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY DUE TO CARDIAC TAMPONADE. AFTER ABLATIONS WERE APPLIED TO THE LEFT INFERIOR PULMONARY VEIN (LIPV) A NURSE COMMENTED THAT THE PATIENT LOOKED "CYANOTIC". THE FLUOROSCOPY AND ECHO WERE CHECKED AND IT WAS SEEN THAT THE HEART WAS NOT MOVING. RESUSCITATION EFFORTS WERE STARTED AND PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION WHICH WERE INITIAL RESOLVED WITH EXTERNAL SHOCKS, HOWEVER, AFTER PERICARDIAL ISSUES HAD BEEN RESOLVED THE HEART DID NOT RESUME NORMAL MOVEMENT. THE PATIENT WAS DECLARED DEAD AT THIS TIME. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827666&pc=MNS | 2518422-2025-104024 | 2025-04-09 04:00 | Death | RESPIRONICS, INC. | 2025-04-11 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.Event Description: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE DEVICE EVALUATION RESULTS AND CODING. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE DEVICE EVALUATION RESULTS AND CODING. THE RENTAL DREAMSTATION ST30 DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION, NO ERRORS WERE FOUND, AND THE DEVICE WAS SUCCESSFULLY TESTED. IT HAS NOW BEEN RETURNED TO STOCK. ADDITIONALLY, THREE GOOD FAITH EFFORT ATTEMPTS WERE MADE TO GATHER FURTHER INFORMATION, BUT NO ADDITIONAL DETAILS WERE RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827692&pc=NPT | 2135147-2025-01916 | 2025-03-17 04:00 | Death | ABBOTT MEDICAL | 2025-04-11 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. DURING DEPLOYMENT OF THE DEVICE, THE PATIENTS PRESSURE DROPPED AT AN ACCEPTABLE LEVEL (80/40). JUST AFTER THE FINAL VALVE RELEASE, THE PRESSURE DROPPED DRAMATICALLY. IMMEDIATELY THE ANESTHETIST REACTED BY GIVING PRESSURE SUSTAINING DRUGS. THE ORIGIN OF THE LOW PRESSURE WAS UNCLEAR, THE VALVE WAS IDEALLY PLACED (4 MM DEPTH) AND THE FUNCTIONALITY WAS OK BASED ON ANGIO. THE PHYSICIAN CHECKED THE PUNCTURE SITE WITH SOME DOUBTS CONCERNING A POTENTIAL EFFUSION AT THIS LEVEL. THE FEMORAL AND ILIAC WAS NOTED TO BE OK THROUGH. THE PERICARDIAC SPACE WAS CHECKED TO EXCLUDE A POTENTIAL TAMPONADE. CORONARY PATENCY WAS ALSO CONTROLLED WITH SOME FLOW AT THE RIGHT AND LEFT LEVEL, BUT IT WAS NOT EASY TO VISUALIZE THE CORONARIES DISTALITY DUE TO THE LOW PRESSURE AND FLOW. PATIENT MANAGEMENT RECOVERY WAS DONE WITH CARDIAC MASSAGE FOR +/- 40 MINUTES. THE PATIENT ENDED UP PASSING AWAY DURING THE PROCEDURE DUE TO CARDIAC ARREST. TIMING BETWEEN THE BEGINNING OF THE TAVI PROCEDURE AND THE DECISION TO STOP THE PROCEDURE WAS +/- 90 MINUTES. THE IMPLANTER CLASSIFIED THIS DEATH RELATED TO THE PROCEDURE AND NOT TO THE IMPLANTED DEVICE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. DURING DEPLOYMENT OF THE DEVICE, THE PATIENTS PRESSURE DROPPED AT AN ACCEPTABLE LEVEL (80/40). JUST AFTER THE FINAL VALVE RELEASE, THE PRESSURE DROPPED DRAMATICALLY. IMMEDIATELY THE ANESTHETIST REACTED BY GIVING PRESSURE SUSTAINING DRUGS. THE ORIGIN OF THE LOW PRESSURE WAS UNCLEAR, THE VALVE WAS IDEALLY PLACED (4 MM DEPTH) AND THE FUNCTIONALITY WAS OK BASED ON ANGIO. THE PHYSICIAN CHECKED THE PUNCTURE SITE WITH SOME DOUBTS CONCERNING A POTENTIAL EFFUSION AT THIS LEVEL. THE FEMORAL AND ILIAC WAS NOTED TO BE OK THROUGH. THE PERICARDIAC SPACE WAS CHECKED TO EXCLUDE A POTENTIAL TAMPONADE. CORONARY PATENCY WAS ALSO CONTROLLED WITH SOME FLOW AT THE RIGHT AND LEFT LEVEL, BUT IT WAS NOT EASY TO VISUALIZE THE CORONARIES DISTALITY DUE TO THE LOW PRESSURE AND FLOW. PATIENT MANAGEMENT RECOVERY WAS DONE WITH CARDIAC MASSAGE FOR +/- 40 MINUTES. THE PATIENT ENDED UP PASSING AWAY DURING THE PROCEDURE DUE TO CARDIAC ARREST. TIMING BETWEEN THE BEGINNING OF THE TAVI PROCEDURE AND THE DECISION TO STOP THE PROCEDURE WAS +/- 90 MINUTES. THE IMPLANTER CLASSIFIED THIS DEATH RELATED TO THE PROCEDURE AND NOT TO THE IMPLANTED DEVICE. Manufacturer Narrative: AN EVENT OF CARDIAC ARREST, HYPOTENSION, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION AND MEDICAL REVIEW, THE CAUSE OF THE REPORTED CARDIAC ARREST, HYPOTENSION, AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE THIS IS RELATED TO PROCEDURAL CONDITION. HOWEVER, THIS CANNOT BE CONFIRMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: NA. THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827831&pc=BZD | 2518422-2025-104022 | 2025-03-28 04:00 | Death | RESPIRONICS, INC. | 2025-04-11 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT NEVER RECEIVED THE REPLACEMENT RECERTIFIED DREAMSTATION ST30 DEVICE, AND DID NOT NEED THE REPLACEMENT SINCE THE USER HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI. Manufacturer Narrative: CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827872&pc=MIH | 2017233-2025-06066 | 2025-03-19 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-11 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P040043 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT WAS IMPLANTED WITH GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE AORTIC COMPONENT AND SIDE BRANCH) AND A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM TO TREAT A THORACIC ANEURYSM. IT WAS REPORTED THAT THE PATIENT¿S ANEURYSM RUPTURED AFTER THE FINAL DEVICE WAS IMPLANTED. COMPRESSIONS WERE STARTED, BLOOD PRODUCTS WERE GIVEN (AMOUNT IS UNKNOWN). THE PATIENT EXPIRED IN THE OPERATING ROOM. Manufacturer Narrative: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU), POTENTIAL DEVICE OR PROCEDURE RELATED ADVERSE EVENTS SECTION INCLUDES BUT IS NOT LIMITED TO ANEURYSM RUPTURE AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827907&pc=MIH | 2017233-2025-06067 | 2025-03-19 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-11 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P210032 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT WAS IMPLANTED WITH GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE AORTIC COMPONENT AND SIDE BRANCH) AND A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM TO TREAT A THORACIC ANEURYSM. IT WAS REPORTED THAT THE PATIENT¿S ANEURYSM RUPTURED AFTER THE FINAL DEVICE WAS IMPLANTED. COMPRESSIONS WERE STARTED; BLOOD PRODUCTS WERE GIVEN (AMOUNT IS UNKNOWN). THE PATIENT EXPIRED IN THE OPERATING ROOM. Manufacturer Narrative: ACCORDING TO THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE AORTIC COMPONENT AND SIDE BRANCH) INSTRUCTIONS FOR USE (IFU), POTENTIAL DEVICE OR PROCEDURE RELATED ADVERSE EVENTS SECTION INCLUDES BUT IS NOT LIMITED TO ANEURYSM RUPTURE AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21827933&pc=MIH | 2017233-2025-06068 | 2025-03-19 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-04-11 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P210032 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT WAS IMPLANTED WITH GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE AORTIC COMPONENT AND SIDE BRANCH) AND A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM TO TREAT A THORACIC ANEURYSM. IT WAS REPORTED THAT THE PATIENT¿S ANEURYSM RUPTURED AFTER THE FINAL DEVICE WAS IMPLANTED. COMPRESSIONS WERE STARTED, BLOOD PRODUCTS WERE GIVEN (AMOUNT IS UNKNOWN). THE PATIENT EXPIRED IN THE OPERATING ROOM. Manufacturer Narrative: ACCORDING TO THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE AORTIC COMPONENT AND SIDE BRANCH) INSTRUCTIONS FOR USE (IFU), POTENTIAL DEVICE OR PROCEDURE RELATED ADVERSE EVENTS SECTION INCLUDES BUT IS NOT LIMITED TO ANEURYSM RUPTURE AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21828421&pc=MGB | 2024168-2025-03789 | 2024-01-02 05:00 | Death | ABBOTT VASCULAR INC. | 2025-04-11 04:00 | MGB | PERCLOSE PROGLIDE¿ | Improper or Incorrect Procedure or Method; Malposition of Device | Ischemia; Obstruction/Occlusion | P960043 |  | 1 | Event Description: DURING A TRAINING SESSION WITH DR. (B)(6)'S TEAM, DR. (B)(6) REPORTED THAT A PRODUCT EXPERIENCE OCCURRED IN 2024, WITH THE CARDIOLOGY TEAM AT (B)(6) HOSPITAL. REPORTEDLY, THE TEAM PERFORMED A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) INTERVENTIONAL PROCEDURE, WHICH WENT SMOOTHLY, USING A PERCLOSE DEVICE. THE PATIENT WAS TRANSFERRED TO THE ICU, BUT THE NEXT DAY, THE VASCULAR TEAM WAS CALLED TO RE-EVALUATE THE PATIENT DUE TO LIMB ISCHEMIA. UNFORTUNATELY, WHEN RE-EVALUATING THE PATIENT, THEY DID NOT SURVIVE AND PASSED AWAY. THE PHYSICIAN STATED IT IS VERY LIKELY THAT THE ARTERY WAS DISEASED AND MAY HAVE COLLAPSED WHEN THE DEVICE WAS INSERTED. THE CARDIOLOGY TEAM DID NOT PERFORM A PRIOR ULTRASOUND BEFORE USING THE PERCLOSE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED. FACTORS THAT CONTRIBUTE TO BACK WALL STITCH, INCLUDE, BUT NOT LIMITED TO, USER TECHNIQUE (LATERAL PUNCTURE, USING THE DEVICE IN A FEMORAL VESSEL < 5MM). THE REPORTED PATIENT EFFECTS OF EFFECTS OF OCCLUSION , ISCHEMIA, AND DEATH ARE KNOWN INHERENT RISKS OF THE PROCEDURE. THE SUBSEQUENT TREATMENTS APPEAR TO BE RELATED TO PROCEDURE CIRCUMSTANCE. THE PATIENT EFFECTS OF OCCLUSION, ISCHEMIA, AND DEATH ARE LISTED IN THE PROGLIDE INSTRUCTIONS FOR USE (IFU), AS POTENTIAL ADVERSE EVENT ASSOCIATED WITH USE OF THE VESSEL CLOSURE DEVICES. MEDICAL REVIEW WAS PERFORMED. THE REVIEW DETAILS CONCLUDED, LIMB ISCHEMIA, A CONDITION WHERE THERE'S INADEQUATE BLOOD FLOW TO THE LIMBS, CAN BE A COMPLICATION AFTER USING THE PERCLOSE PROGLIDE SUTURE-MEDIATED CLOSURE DEVICE (SMCR). THIS CAN OCCUR DUE TO VARIOUS FACTORS, INCLUDING COMMON FEMORAL ARTERY OCCLUSION, ESPECIALLY IF THE SUTURE CAPTURES THE POSTERIOR WALL OF THE ARTERY, THROMBOSIS WHEN BLOOD CLOTS CAN FORM AT THE CLOSURE SITE OR IN THE AFFECTED ARTERY, REDUCING BLOOD FLOW, OR DISSECTION WHEN THE DEVICE POTENTIALLY CAUSE A TEAR IN THE ARTERY WALL, LEADING TO A DISSECTION AND COMPROMISED BLOOD FLOW. ACUTE LIMB ISCHEMIA IS A MEDICAL EMERGENCY AND REQUIRES IMMEDIATE TREATMENT TO PREVENT TISSUE DAMAGE AND POTENTIAL AMPUTATION. IT WAS REPORTED THAT FEMORAL ANGIOGRAM WAS NOT PERFORMED. IT SHOULD BE NOTED THAT THE PROGLIDE INSTRUCTIONS FOR USE (IFU), STATES: PRIOR TO DEPLOYMENT OF THE PERCLOSE PROGLIDE SMC DEVICE, PERFORM A FEMORAL ANGIOGRAM TO EVALUATE THE ACCESS SITE FOR VESSEL SIZE, CALCIUM DEPOSITS, TORTUOSITY, AND FOR DISEASE OR DISSECTIONS OF THE WALL TO AVOID DEVICE CUFF MISSES (DEVICE NEEDLES NOT ENGAGING WITH THE CUFFS) AND / OR POSTERIOR WALL SUTURE PLACEMENT AND POSSIBLE LIGATION OF THE ANTERIOR AND POSTERIOR WALLS OF THE VESSEL. IT IS UNKNOWN IF THE IFU DEVIATION OF NO FEMORAL ANGIOGRAM CONTRIBUTED TO THE REPORTED EVENT. IN THIS CASE, MEDICAL REVIEW CONCLUDED IT IS UNKNOWN IF THE PROGLIDE DIRECTLY CAUSED THE PATIENT¿S DEATH ESPECIALLY THAT THE VESSEL WAS NOT PREPPED OR IMAGED FOR ANY PRE-EXISTING DISEASE AND NO OTHER INFORMATION WAS PROVIDED EXCEPT THE PROCEDURE WENT SMOOTHLY. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. Manufacturer Narrative: B5: THE DATE OF EVENT IS ESTIMATED. D4: A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN. H6: MEDICAL DEVICE PROBLEM CODE: 2017 - FAILURE TO FOLLOW STEPS / INSTRUCTIONS. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21829644&pc=LYJ | 1644487-2025-00400 | 2025-03-15 04:00 | Death | CYBERONICS - HOUSTON | 2025-04-11 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S OFFICIAL CAUSE OF DEATH WAS HYPOXIC CEREBRAL DAMAGE. THE PATIENT'S DEVICE IS NOT AVAILABLE FOR RETURN. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21829812&pc=BZD | 2518422-2025-043768 | 2022-08-03 04:00 | Death | RESPIRONICS, INC. | 2025-04-11 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGES DEATH. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21829814&pc=BZD | 2518422-2025-043746 | 2024-05-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-11 04:00 | BZD | REMSTAR AUTO | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED, AND SECTION H 11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WERE 26 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21829834&pc=LWS | 2124215-2025-23135 | 2025-03-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-11 04:00 | LWS | RESONATE? HF ICD DR | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS IMPLANTED SUCCESSFULLY. HOWEVER, THE FOLLOWING DAY, JUST PRIOR TO DISCHARGE, THE PATIENT'S VITALS WERE ABNORMAL. A RAPID RESPONSE TEAM (RRT) WAS CALLED. THE PATIENT WAS FOUND TO HAVE A SMALL PERICARDIAL EFFUSION AND IT WAS DECIDED TO PERFORM A PERICARDIOCENTESIS. THE PATIENT EXPIRED DURING THIS PROCEDURE. NO X-RAYS WERE OBTAINED, SO IT WAS UNKNOWN IF A LEAD HAD CAUSED A PERFORATION. THE CAUSE OF THE PERICARDIAL EFFUSION WAS NOT KNOWN. AVAILABLE INFORMATION SUGGESTS THE DEVICE AND LEADS WERE NOT EXPLANTED POST-MORTEM. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21830401&pc=FKX | 0008030665-2025-00844 | 2025-03-07 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-11 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED WHILE UNDERGOING TREATMENT ON A LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INITIAL REPORTING. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CONTINUOUS CYCLIC PD (CCPD) THERAPY IN THE EARLY MORNING OF (B)(6) 2025. THE PATIENT'S POINT-OF-CONTACT (POC) DISCOVERED THE PATIENT HAD EXPIRED WHEN THEY WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT'S TREATMENT DATA WAS REVIEWED AND IT INDICATED THE PATIENT COMPLETED TREATMENT, WITH NO ALARMS NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT'S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO THE PATIENT¿S SIGNIFICANT CARDIAC HISTORY. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WAS SCHEDULED TO BE PICKED-UP/RETURNED.Event Description: IT WAS REPORTED TO FRESENIUS THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED WHILE UNDERGOING TREATMENT ON A LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INITIAL REPORTING. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CONTINUOUS CYCLIC PD (CCPD) THERAPY IN THE EARLY MORNING OF (B)(6) 2025. THE PATIENT¿S POINT-OF-CONTACT (POC) DISCOVERED THE PATIENT HAD EXPIRED WHEN THEY WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT'S TREATMENT DATA WAS REVIEWED AND IT INDICATED THE PATIENT COMPLETED TREATMENT, WITH NO ALARMS NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO THE PATIENT¿S SIGNIFICANT CARDIAC HISTORY. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WAS SCHEDULED TO BE PICKED-UP/RETURNED. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, DUE TO THEIR SIGNIFICANT CARDIAC HISTORY. THE SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE END-STAGE RENAL DISEASE (ESRD) POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G4, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED A CASSETTE MEMBRANE. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, A VALVE ACTUATION TEST, AND A TEACH PUMPS TEST. A SIMULATED TREATMENT USING AS-RECEIVED TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THERE WERE NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21830443&pc=BZD | 2518422-2025-043777 | 2021-08-25 04:00 | Death | RESPIRONICS, INC. | 2025-04-11 04:00 | BZD | RENTAL, DREAMSTATION ST30, GB | Degraded | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGES DEATH. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGES DEATH. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THIS REPORT IS BEING SENT TO CORRECT THE PREVIOUSLY FILED REPORT. THE MANUFACTURER RECEIVED INFORMATION THAT A DEVICE WAS BEING RETURNED SINCE A PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION AGAINST THE DEVICE AT THIS TIME. THE DEVICE WAS RETURNED FOR EVALUATION. DURING SERVICING OF THE DEVICE NO DEVICE PROBLEM HAS BEEN DETECTED. THE DEVICE IS NOT A PART OF RECALL. SECTION G2, G4 AND H6 HAVE BEEN UPDATED/CORRECTED IN THIS FOLLOW-UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21830646&pc=BZD | 2518422-2025-043776 | 2024-01-12 05:00 | Death | RESPIRONICS, INC. | 2025-04-11 04:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION AUTO CPAP DEVICE STATING THAT THE PATIENT PASSED AWAY AS PER PATIENT'S WIFE. CALLER RECEIVED THE REPLACEMENT DEVICE AND WANTS TO RETURN IT TO PHILIPS AS IT IS NOT NEEDED. CALLER IS REQUESTING FOR A RETURN LABEL VIA US MAIL AND STATED THAT CALLER IS NOT INTERESTED IN CLASS ACTION LAWSUIT. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER¿S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST-LIKE CONTAMINATION, ADHESIVE RESIDUE, BLACK DUST LIKE CONTAMINATION IN THE DEVICE. THE MANUFACTURER CONFIRMED THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN IN THE RETURNED MATERIALS. THE MANUFACTURER CONCLUDES CONTAMINATION WERE EXTERNAL TO THE DEVICE. THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. IN THIS REPORT, BOX D, G, H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21831190&pc=QFG | 3019004087-2025-00305 | 2025-03-11 04:00 | Death | BETA BIONICS | 2025-04-11 04:00 | QFG | ILET BIONIC PANCREAS | Pumping Stopped | Hyperglycemia | K231485 |  | 1 | Event Description: IT WAS REPORTED THAT AN ILET USER WAS FOUND UNRESPONSIVE AND PASSED AWAY ON (B)(6) 2025. REVIEW OF THE DEVICE DATA SHOWED THAT INSULIN DOSING WAS MANUALLY PAUSED BY THE USER FOR 2 HOURS AT 7:20 PM ON(B)(6) 2025. RESUME INSULIN ALARMS WERE TRIGGERED AS INTENDED AT 9:20 PM, BUT INSULIN DOSING WAS NEVER RESUMED. CGM GLUCOSE ROSE AFTER INSULIN WAS PAUSED AND REMAINED ELEVATED UNTIL THE DEVICE LOST POWER AT 9:26 AM ON (B)(6) 2025, THE DAY BEFORE THEIR DEATH. DEVICE LOGS INDICATE NO FURTHER INTERACTIONS WITH THE DEVICE AFTER 7:20 PM ON (B)(6) 2025. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. AVAILABLE ILET LOGS WERE REVIEWED. THE ILET LOGS INCLUDED DATA FROM THE REPORTED EVENT DATE. NO EVIDENCE OF DEVICE MALFUNCTION WAS FOUND IN THE ENGINEERING LOGS. DEVICE DATA CONFIRMED HYPERGLYCEMIA BEGINNING IN THE EVENING OF (B)(6) 2025 AND CONTINUING UNTIL THE DEVICE LOST POWER. THE ILET DOSED INSULIN WHEN IT WAS ABLE IN RESPONSE TO AVAILABLE CGM GLUCOSE VALUES AS INTENDED. THE ILET TRIGGERED DEVICE AND CGM GLUCOSE ALERTS AS INTENDED. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. WHILE THE EXACT CAUSE OF DEATH IS UNKNOWN, THE USER WAS FOUND UNRESPONSIVE, AND CLINICAL LOGS SHOW THE LAST KNOWN CGM GLUCOSE VALUES WERE EXCEEDING 400 MG/DL ON THE DAY BEFORE THEIR DEATH. INSULIN DELIVERY WAS MANUALLY PAUSED BY THE USER AT 7:20 PM ON (B)(6) 2025 AND WAS NOT RESUMED, RESULTING IN SUSTAINED HYPERGLYCEMIA. THE USER DID NOT INTERACT WITH THE DEVICE AGAIN AFTER PAUSING INSULIN. IF THE PRODUCT IS RECEIVED AT A LATER DATE, AND NEW RELEVANT INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21831193&pc=NKE | 2017865-2025-39902 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-11 04:00 | NKE | QUADRA ALLURE MP CRT-P | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA AND CARDIOGENIC SHOCK. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE VOLTAGE WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) LEVEL UPON RECEIPT. DEVICE IMAGE AND SESSION RECORDS INDICATED THE DEVICE WAS WITHIN NORMAL RANGE OF OPERATION. FURTHER ANALYSIS PERFORMED FOUND NO ANOMALY. ASSESSMENT OF TOTAL PROJECTED LONGEVITY WAS PERFORMED, AND THE DEVICE WAS FOUND TO HAVE APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21831196&pc=NVN | 2017865-2025-39903 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-11 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA AND CARDIOGENIC SHOCK. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL AND X-RAY INSPECTIONS OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21831198&pc=NIK | 2017865-2025-39906 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-11 04:00 | NIK | QUARTET LEAD SMALL-S, 86 CM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA AND CARDIOGENIC SHOCK. Manufacturer Narrative: THE REPORTED EVENT WAS ¿PATIENT DEATH DUE TO VENTRICULAR TACHYCARDIA AND CARDIOGENIC SHOCK.¿ AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL EXAMINATION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING OF THE LEAD DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. UNABLE TO MEASURE S-CURVE DUE TO THE CONDITION OF THE LEAD AS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21831199&pc=NVN | 2017865-2025-39905 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-11 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA AND CARDIOGENIC SHOCK. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21831906&pc=LPB | 2029046-2025-01133 | 2025-03-18 04:00 | Death | BIOSENSE WEBSTER INC | 2025-04-11 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE DIAGNOSTIC PORTION OF AN ATRIAL FLUTTER ABLATION PROCEDURE, THE PATIENT EXPERIENCED ARRHYTHMIAS AND EXPIRED. DURING ENTRAINMENT WITH ABLATION CATHETER, THE PHYSICIAN CAME OFF PACING AND THE PATIENT WAS IN AN UNKNOWN RHYTHM. ALSO, THEY SAW MULTIPLE FLUTTER WAVES AND A DISASSOCIATED QRS WAVES. THE PHYSICIAN PUT THE CATHETER IN THE VENTRICLE AND TRIED PACING FROM THE VENTRICLE WITH THE CATHETER. DURING THE EVENT, THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY (PEA). THE RHYTHM SIMULATED VENTRICULAR ESCAPE, AND THE PATIENT FLAT-LINED. THEY ATTEMPTED TO SHOCK THE PATIENT TWO TIMES AND CHEST COMPRESSIONS STARTED. COMPRESSIONS WERE CONDUCTED FOR OVER 40 MINUTES, AND ULTIMATELY, THE PATIENT EXPIRED. THE PATIENT WAS IN FLUTTER RHYTHM OF 2 TO 1 AND 3 TO 1. DURING ADVANCEMENT OF THE CATHETERS, THE PATIENT CONTINUED TO BE IN FLUTTER. ONE CATHETER (DECANAV) PLACED INTO THE CORONARY SINUS AND THE ABLATION CATHETER (THERMOCOOL SMARTTOUCH) PLACED IN THE CAVO-TRICUSPID ISTHMUS (CTI) LINE. NO ABLATION OR ANY ELECTRICAL MAPPING DONE DURING THIS CASE. THE NIGHT PRIOR TO THE PROCEDURE, THE PATIENT SUFFERED A STROKE, BUT THE PATIENT WAS CLEARED AFTER CONFIRMATION WITH TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THE PATIENT HAD MITRAL VALVE REGURGITATION. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 23-APR-2025. IT WAS REPORTED THAT THE PHYSICIAN THOUGHT THE CAUSE OF DEATH WAS CARDIOVASCULAR RELATED DUE TO THE PREVIOUS STROKE, AS WELL AS MITRAL AND TRICUSPID VALVE REGURGITATION. THE PHYSICIAN DID NOT THINK IT WAS ANYTHING RELATED TO BIOSENSE WEBSTER INC. (BWI) PRODUCTS. THE PATIENT HAD MULTIPLE HEALTH COMPLICATIONS. IT WAS ALSO REPORTED THAT THE LOT NUMBER FOR THE PRODUCT IS NOT AVAILABLE BECAUSE THE CATHETER BOX WAS THROWN OUT. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21932253&pc=BZD | 2518422-2025-044489 | 2023-12-12 05:00 | Death | RESPIRONICS, INC. | 2025-04-30 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGES DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX A: PATIENT IDENTIFIER WAS CORRECTED. BOX E: REPORTER FIRST NAME AND REPORTER LAST NAME WAS CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21932309&pc=OZP | 2032227-2025-163986 | 2025-03-28 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-30 04:00 | OZP | MINIMED 770G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE PRIMARY CAUSE OF DEATH WAS ALCOHOL CONSUMPTION AND DIABETES. THE CUSTOMER WAS HOSPITALIZED ON (B)(6) 2025 WITH A BLOOD GLUCOSE VALUE OF ABOVE 900 MG/DL. THE PUMP WAS NOT WORN AT THE TIME OF PASSING AS THE CUSTOMER WAS KNOWN TO BE WEARING THE PUMP ON (B)(6) 2025 FOR THE LAST TIME. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1880, MMT-431AG, MMT-342, MMT-7020LA. TROUBLESHOOTING WAS PERFORMED AND IT WAS NOTICED THAT THE CUSTOMER WAS USING THE SENSOR BUT IT WAS NOT WORN DURING THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1880. NO PRODUCT RETURN IS REQUIRED FOR MMT-431AG. NO PRODUCT RETURN IS REQUIRED FOR MMT-342. NO PRODUCT RETURN IS REQUIRED FOR MMT-7020LA. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21932867&pc=MIH | 9612164-2025-02133 | 2022-11-09 05:00 | Death | MEDTRONIC IRELAND | 2025-04-30 04:00 | MIH | ENDURANT STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿EDITOR'S CHOICE - COMPARISON OF OUTCOMES FOR MAJOR CONTEMPORARY ENDOGRAFT DEVICES USED FOR ENDOVASCULAR REPAIR OF INTACT ABDOMINAL AORTIC ANEURYSMS¿. THE TIME FRAME FOR THIS STUDY WAS OVER A NINE-YEAR PERIOD. 2874 PATIENTS WERE INCLUDED IN THE STUDY. ENDURANT AND NON-MDT STENT GRAFTS WERE IMPLANTED IN THE PATIENT POPULATION. PATIENT MORTALITY WAS REPORTED IN THE GROUP OF PATIENTS THAT RECEIVED ENDURANT STENT GRAFTS AND SOME OF THE DEATHS WERE DEEMED AORTIC RELATED. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿COMPARISON OF OUTCOMES FOR MAJOR CONTEMPORARY ENDOGRAFT DEVICES USED FOR ENDOVASCULAR REPAIR OF INTACT ABDOMINAL AORTIC ANEURYSMS¿. FALSTER MO, GARLAND SK, JORM LR, BEILES CB, FREEMAN AJ, SEDRAKYAN A, SOTADE OT, VARCOE RL EUR J VASC ENDOVASC SURG. 2023 FEB;65(2):272-280. DOI: 10.1016/J.EJVS.2022.11.005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21932942&pc=BZD | 2518422-2025-044494 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-04-30 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT DUE TO THE ISSUE WITH THE DEVICE. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973, AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21933428&pc=OZO | 2032227-2025-164412 | 2024-09-30 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-30 04:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE PRIMARY CAUSE OF DEATH WAS UNKNOWN. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNK_SENSOR, MMT-332A, UNOMEDICAL, MMT-1715K. TROUBLESHOOTING WAS PERFORMED AND THE REPORTING PARTY SUSPECTED THAT THE DEATH WAS CAUSED DUE TO PUMP MALFUNCTION. NO PRODUCT RETURN IS REQUIRED FOR UNK_SENSOR. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. PRODUCT RETURN IS REQUIRED FOR MMT-1715K.Event Description: UPDATED SUMMARY: NO PRODUCT RETURN IS REQUIRED FOR MMT-1715K.Event Description: UPDATED SUMMARY: THE EVENT IS NOT REPORTABLE ON THE INSULIN PUMP AS THE CUSTOMER HAD NOT ALLEGED ANY MALFUNCTION ON THE PUMP AND ALSO THE REASON OF DEATH WAS NOT DUE TO DIABETES. ALSO, THE INSULIN PUMP WAS WORN AT THE TIME OF DEATH. HENCE THE EVENT SUBMITTED UNDER REGULATORY REPORT NUMBER 2032227-2025-164412 IS NOT REPORTABLE. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT AND/OR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN RECEIVED REGARDING PRODUCT RETURN WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. THE INFORMATION HAS BEEN PROVIDED IN SECTIONS B5 (UPDATED THE SUMMARY), H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) AND H11 (PRODUCT RETURN STATUS) WITH THIS REPORT. ALSO, ADDITIONAL INFORMATION HAS BEEN RECEIVED REGARDING THE AWARE DATE CHANGE WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. SO, THE AWARE DATE HAS BEEN CHANGED TO 03-APR-2025 AS PER NEW ADDITIONAL INFORMATION. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21933470&pc=DSQ | 2916596-2025-02644 | 2025-04-12 04:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21933514&pc=DSQ | 2916596-2025-02860 | 2025-02-20 05:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Sepsis; Fungal Infection; Respiratory Failure; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT HAD A PROTRACTED POST OPERATIVE COURSE WITH DEBILITY, VASOPLEGIA, ILEUS, AND SEPSIS. THEY ALSO HAD CONTINUED DEPENDENCE ON CONTINUOUS RENAL REPLACEMENT THERAPY FOR RENAL FAILURE, BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) FOR ACUTE RESPIRATORY FAILURE WITH CARBON DIOXIDE RETENTION AND ALTERATION IN MENTAL STATUS. THE PATIENT WAS FOLLOWED BY INFECTIOUS DISEASES SERVICE FOR METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) BACTEREMIA AND POSITIVE FUNGITELL. THE PATIENT COMPLETED RIFAMPIN AND DAPTOMYCIN THERAPY FOLLOWED ALSO BY PALLIATIVE CARE AND PSYCHOLOGY SERVICES FOR SUPPORT. THE PATIENT WAS INTUBATED FOR ACUTE RESPIRATORY FAILURE. BRONCHOSCOPY WAS RULED OUT MUCOUS PLUGGING. THE PATIENT REQUIRED ESCALATION OF PRESSORS, EPINEPHRINE, NOREPINEPHRINE, VASOPRESSIN IV FLUID BOLUSES, STRESS DOSE STEROIDS WITH NO RESPONSE. THE DEATH WAS NOT DEVICE RELATED AND NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, INFECTION, SEPSIS, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21933760&pc=OZD | 1220648-2025-27958 | 2025-04-03 04:00 | Death | ABIOMED, INC. | 2025-04-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT PRESENTED WITH LETHARGY AND EXTREME FATIGUE AND DISCOMFORT. RADIAL ACCESS WAS GAINED AND THE RIGHT CORONARY ARTERY WAS 100% PROXIMALLY OCCLUDED. THE FOLLOWING DAY AFTER START OF IMPELLA MCS, IT WAS NOTED THAT THE PATIENT CONTINUED TO HAVE LIMITATIONS TO IMPELLA FLOW DESPITE ADEQUATE POSITION AND PRELOAD. THE NEXT DAY IT WAS NOTED THAT THE PATIENT'S HEMODYNAMICS LOOKED "WORSE" SUGGESTING RIGHT VENTRICULAR DYSFUNCTION. APPROXIMATELY FOUR DAYS LATER, THE PATIENT EXPERIENCED HEMOLYSIS AND WAS FOLLOWED BY THROMBOCYTOPENIA DURING THE IMPELLA MCS. THE PATIENT WAS GIVEN THREE UNITS OF PLATELETS AND ONE UNIT OF RED BLOOD CELLS IN RESPONSE TO THE CONDITIONS. SUPPORT WAS CONTINUED WITH THE IMPLANTED PUMP. HOWEVER, THE PATIENT CONTINUED TO BECOME WORSE. THE PHYSICIAN DISCUSSED STARTING CONTINUED RENAL REPLACEMENT THERAPY THROUGH THE EXTRACORPOREAL MEMBRANE OXYGENATION (EMCO) CIRCUIT. APPROXIMATELY SIX DAYS LATER, DAY 11 OF SUPPORT, EXTREMITIES WERE MOTTLED, OOZING YELLOWISH BLOOD FROM PREVIOUS FEMORAL ECMO SITES. THE PATIENT WAS NOTED TO BE IN IRREVERSIBLE ORGAN FAILURE AND SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND THROMBOCYTOPENIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21934104&pc=DSQ | 2916596-2025-02583 | 2025-04-09 04:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Decreased Pump Speed; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Syncope/Fainting; Unspecified Eye / Vision Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT CAME INTO THE EMERGENCY DEPARTMENT (ED) UNRESPONSIVE, WITH NO FLOW. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED WITHIN ONE MINUTE OF THE PATIENT'S ARRIVAL, AND THE PATIENT WAS MADE RESPONSIVE WITH CPR AND DEFIBRILLATOR SHOCKS. THE CAUSE OF THE ISSUE WAS UNKNOWN. THE PATIENT'S FLOWS RETURNED TO NORMAL. THE CAUSE OF THE EVENT WAS UNCLEAR. A COMPUTED TOMOGRAPHY (CT) SCAN WAS TAKEN OF THE HEAD AND FOUND NO ACUTE INTRACRANIAL HEMORRHAGE, AND A COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) OF THE CHEST FOUND NO PLEURAL EFFUSION. NO CT/CTA IMAGES WERE AVAILABLE. AN ECHOCARDIOGRAM WAS ALSO TAKEN AND FOUND NO ACUTE ABNORMALITIES. THE PATIENT REMAINED INTUBATED, WAS OFF SEDATION, AND WAS NOTED TO HAVE LEFT GAZE DEVIATION. A REPEAT ELECTROENCEPHALOGRAM (EEG) WAS TAKEN, WHICH WAS NEGATIVE FOR ANY SEIZURE-LIKE ACTIVITY. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED THE DROP IN FLOW AND POWER ON (B)(6) 2025 AT 12:18:38, WHICH DID NOT APPEAR RELATED TO A PRESSURE OR VOLUME INCREASE, AND AS PUMP TEMPERATURES WERE ALSO NORMAL NO INGESTION WAS EXPECTED. THE PATIENT PASSED AWAY ON (B)(6) 2025 FOLLOWING WITHDRAWAL OF CARE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED INTERMITTENT LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THIS FINDING COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED INTERMITTENT LOW FLOW ALARMS ON (B)(6) 2025 WITH FLOW VALUES AS LOW AS 0.9 LITERS PER MINUTE (LPM). THE SYSTEM OTHERWISE APPEARED TO BE OPERATING AS INTENDED AT THE SET SPEED, AND THERE WERE NO OTHER NOTABLE ALARMS ACTIVE IN THE LOG FILES. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WOULD NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH AND RESPIRATORY FAILURE. SECTION 1 ALSO PROVIDES AN EXPLANATION OF ALL PUMP PARAMETERS, INCLUDING FLOW AND POWER. THIS SECTION EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿ STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM (LITERS PER MINUTE) AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 ¿ALARMS AND TROUBLESHOOTING¿ DESCRIBES ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, IS CURRENTLY AVAILABLE. SECTION 5 "ALARMS AND TROUBLESHOOTING" OUTLINES SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THIS DOCUMENT INSTRUCTS THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21934359&pc=DQY | 2029214-2025-01098 | 2025-03-15 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-04-30 04:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhagic Stroke | K180715 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THE PATIENT UNDERWENT HEAD COMPUTER TOMOGRAPHY (CT) WITHOUT CONTRAST WHICH REVEALED HEMORRHAGIC TRANSFORMATION, CLINICALLY SIGNIFICANT. TEST DATE WAS (B)(6) 2025. THE RELATIVE TIME FROM THE PROCEDURE WAS POST-PROCEDURE. THIS WAS NOT A RE-CURRENT OR NEW STROKE. RATIONALE FOR RELATING ADVERSE EVENT (AE) TO SYSTEM OR PROCEDURE WAS MASSIVE STROKE WITH HEMORRHAGIC TRANSFORMATION, USUAL CONSEQUENCE OF STROKE. THE AE HAD A CAUSAL RELATIONSHIP WITH THE DISEASE UNDER STUDY AND WAS NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. AE DID NOT RESULT FROM A DEVICE DEFICIENCY. PATIENT EXPERIENCED VOLUMINOUS HEMORRHAGIC TRANSFORMATION WITH A LEFT HEMISPHERIC HEMATOMA, RUPTURED IN THE VENTRICLES, RESPONSIBLE FOR SUBFALCORAL AND TEMPORAL INVOLVEMENT. BASELINE MODIFIED RANKIN SCALE (MRS) SCORE 0, NATIONAL INSTITUTES OF HEALTH STROKE SCALE (NIHSS) 27. THIS EVENT RESULTED IN DEATH; THE PATIENT'S DATE OF DEATH WAS
(B)(6) 2025. THIS EVENT DID NOT RESULT IN CONGENITAL ANOMALY, DISABILITY, HOSPITALIZATION, OR MEDICAL INTERVENTION, AND WAS NOT CONSIDERED LIFE-THREATENING. THE SITE ASSESSED THIS EVENT AS NOT RELATED TO THE DEVICE AND POSSIBLY RELATED TO THE STUDY PROCEDURE. FINAL POST-PROCEDURE MTICI SCORE WAS 2C.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT HAD BEEN HOSPITALIZED AND TREATED FOR A LEFT MIDDLE CEREBRAL ARTERY (MCA) M1 STROKE. THE PATIENT WAS NEVER DISCHARGED AND DIED IN HOSPITAL DUE TO HEMORRHAGIC TRANSFORMATION OF THE INDEX STROKE. CAUSALITY ASSESSMENT PROVIDED AS POSSIBLY RELATED TO THE PROCEDURE, NOT RELATED TO DEVICES. THE EVENT WAS DIRECTLY RELATED TO DISEASE UNDER STUDY/INDEX STROKE. RATIONALE FOR THE RELATIONSHIP TO PROCEDURE: "MASSIVE STROKE WITH HEMORRHAGIC TRANSFORMATION, USUAL CONSEQUENCE OF STROKE."Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PRE-PROCEDURE MTICI SCORE WAS 0 AND POST PROCEDURE WAS 2C. DISCHARGE MRS SCORE WAS 6, AND NO ADDITIONAL MRS OR NIHSS SCORE WAS PROVIDED PRIOR TO DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5. UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H6. CODING UPDATED BASED ON AVAILABLE INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21934903&pc=DSQ | 2916596-2025-02567 | 2025-03-31 04:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935110&pc=OZD | 1220648-2025-27966 | 2025-04-10 04:00 | Death | ABIOMED, INC. | 2025-04-30 04:00 | OZD | IMPELLA CP | Positioning Failure | Perforation of Vessels | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 71-YEAR-OLD FEMALE PATIENT ON AN IMPELLA CP FOR SUPPORT DURING A HIGH-RISK PROCEDURE. IT WAS REPORTED THAT DURING REPOSITIONING, THE IMPELLA WAS INADVERTENTLY PULLED BACK INTO THE AORTA. PRESSOR REQUIREMENTS WERE INCREASED AS A RESULT OF THE ISSUE. ATTEMPTS TO READVANCE THE PUMP WERE UNSUCCESSFUL AS A RESULT OF THE PATIENT'S AORTIC STENOSIS. IT WAS THEN NOTED THAT THE IMPELLA SITE WAS HEMORRHAGING. THE PATIENT EXPERIENCED MIXED CARDIOGENIC/HEMORRHAGIC SHOCKED. MANUAL PRESSURE WAS HELD AND OVER FIVE UNITS OF BLOOD WERE GIVEN TO TRY TO MANAGE THE CONDITION. HEMOSTASIS WAS NOT ACHIEVED. THE PATIENT WAS MADE DNR BY THE FAMILY. VASCULAR SURGERY WAS MADE AWARE OF THE ARTERIAL INJURY AT THE IMPELLA SITE, BUT THE PATIENT ULTIMATELY PASSED AWAY DUE TO THE HEMORRHAGIC EVENT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. ACCESS SITE BLEEDING - MAJOR: CLINICAL DETAILS NOTED THAT THE SURGICAL GUIDEWIRE PRESENT IN THE FEMORAL ARTERY NEXT TO IMPELLA SITE AND WHEN REPOSITIONED, BLEEDING AT IMPELLA SITE WAS OBSERVED. IT WAS NOTED THAT ANGLE MATCHING AND FORWARD SUTURING WERE APPLIED WHEN IMPELLA WAS INSERTED. MULTIPLE UNITS OF BLOOD WERE GIVEN AND MANUAL PRESSURE WAS IN ATTEMPT TO ACHIEVE HEMOSTASIS WITHOUT SUCCESS. AS THE CLINICAL DETAILS NOTED THAT EXCESSIVE BLEEDING BUT NO VASCULAR DAMAGE OCCURRED, THE FAILURE MODE WAS CHANGED FROM "VASCULAR DAMAGE - PERIPHERAL VESSEL" TO "ACCESS SITE BLEEDING - MAJOR". THE ROOT CAUSE OF THE ACCESS SURE BLEEDING COULD NOT BE DEFINITIVELY DETERMINED AS PUMP WAS ANGLE MATCHED AND FORWARD SUTURED, BLOOD WAS GIVEN, AND MANUAL PRESSURE WAS HELD IN ATTEMPT TO RESOLVE BLEEDING. POSITIONING ISSUES: CLINICAL DETAILS NOTED THAT WHEN GUIDEWIRE WAS BEING REPOSITIONED AT IMPELLA ACCESS SITE, "IMPELLA IN AORTA" ALARMS WERE TRIGGERED AND ECHO SHOWED PUMP IN AORTA WITH PIGTAIL ACROSS AV AND IN AORTA. THE ROOT CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY AN INTERACTION WITH OTHER DEVICES AS THE PUMP CAME OUT OF POSITION WHEN REPOSITIONING OF A GUIDEWIRE AT THE ACCESS SITE BASED ON CLINICAL DETAILS. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. ACCESS SITE BLEEDING - MAJOR: CLINICAL DETAILS NOTED THAT THE SURGICAL GUIDEWIRE PRESENT IN THE FEMORAL ARTERY NEXT TO IMPELLA SITE AND WHEN REPOSITIONED, BLEEDING AT IMPELLA SITE WAS OBSERVED. IT WAS NOTED THAT ANGLE MATCHING AND FORWARD SUTURING WERE APPLIED WHEN IMPELLA WAS INSERTED. MULTIPLE UNITS OF BLOOD WERE GIVEN AND MANUAL PRESSURE WAS IN ATTEMPT TO ACHIEVE HEMOSTASIS WITHOUT SUCCESS. AS THE CLINICAL DETAILS NOTED THAT EXCESSIVE BLEEDING BUT NO VASCULAR DAMAGE OCCURRED, THE FAILURE MODE WAS CHANGED FROM "VASCULAR DAMAGE - PERIPHERAL VESSEL" TO "ACCESS SITE BLEEDING - MAJOR". THE ROOT CAUSE OF THE ACCESS SURE BLEEDING COULD NOT BE DEFINITIVELY DETERMINED AS PUMP WAS ANGLE MATCHED AND FORWARD SUTURED, BLOOD WAS GIVEN, AND MANUAL PRESSURE WAS HELD IN ATTEMPT TO RESOLVE BLEEDING. POSITIONING ISSUES: CLINICAL DETAILS NOTED THAT WHEN GUIDEWIRE WAS BEING REPOSITIONED AT IMPELLA ACCESS SITE, "IMPELLA IN AORTA" ALARMS WERE TRIGGERED AND ECHO SHOWED PUMP IN AORTA WITH PIGTAIL ACROSS AV AND IN AORTA. THE ROOT CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY AN INTERACTION WITH OTHER DEVICES AS THE PUMP CAME OUT OF POSITION WHEN REPOSITIONING OF A GUIDEWIRE AT THE ACCESS SITE BASED ON CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935239&pc=MIH | 2247858-2025-00122 | 2025-04-16 04:00 | Death | BOLTON MEDICAL, INC. | 2025-04-30 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Vascular Dissection | P200045 |  | 1 | Event Description: "DEATH: ON (B)(6), AN EMERGENCY TEVAR WAS PERFORMED FOR AN AORTIC DISSECTION. THE RELAY PRO STENT GRAFT CONCERNED (28-N4-32-209-28U, LOT#: 2404170130) WAS IMPLANTED CENTRALLY AND A RELAY PRO STENT GRAFT (28-N4-32-164-28U, LOT#: 2407130037) WAS ADDITIONALLY IMPLANTED PERIPHERALLY. THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEMS WITH THE PROCEDURE OR THE DEVICES THEMSELVES. ON (B)(6), THE PATIENT'S CONDITION SUDDENLY CHANGED, AND THE PATIENT DIED. AFTER DEATH, A CT SCAN REVEALED THAT THE INTERNAL PRESSURE IN THE FALSE LUMEN HAD INCREASED DUE TO THE EXPANSION OF THE TRUE LUMEN CAUSED BY THE IMPLANTATION OF THE STENT GRAFTS, AND THE FALSE LUMEN HAD RUPTURED AROUND THE ABDOMEN. PHYSICIAN'S COMMENT: THE PHYSICIAN DID THE BEST THAT HE COULD WITH THE MEANS AVAILABLE TO THE HOSPITAL. THERE WAS NO OTHER CHOICE. PHYSICIAN'S COMMENT ON CAUSAL RELATIONSHIP: THE MEANS TAKEN FOR THE AORTIC DISSECTION WERE THE BEST, GIVEN THE EMERGENCY TREATMENT ON APRIL 14, BUT THE TRUE LUMEN EXPANDED MORE THAN EXPECTED. OPERATION TYPE: EMERGENCY TEVAR FOR AN AORTIC DISSECTION, NO BLOOD LOSS, ANCILLARY DEVICE USED: 28-N4-32-164-28U (LOT#: 2407130037), NO IMAGE AVAILABLE DUE TO HOSPITAL POLICY, NO PRE-CASE PLAN AVAILABLE DUE TO HOSPITAL POLICY, NO ADDITIONAL INFORMATION AVAILABLE DUE TO HOSPITAL POLICY, (TC#: (B)(4))." PATIENT OUTCOME: "THE PATIENT DIED." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR 2247858-2025-00121. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935240&pc=MIH | 2247858-2025-00121 | 2025-04-16 04:00 | Death | BOLTON MEDICAL, INC. | 2025-04-30 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Vascular Dissection | P200045 |  | 1 | Event Description: "DEATH: ON (B)(6), AN EMERGENCY TEVAR WAS PERFORMED FOR AN AORTIC DISSECTION. THE RELAY PRO STENT GRAFT CONCERNED (28-N4-32-209-28U, LOT#: 2404170130) WAS IMPLANTED CENTRALLY AND A RELAY PRO STENT GRAFT (28-N4-32-164-28U, LOT#: 2407130037) WAS ADDITIONALLY IMPLANTED PERIPHERALLY. THE PROCEDURE WAS COMPLETED WITHOUT ANY PROBLEMS WITH THE PROCEDURE OR THE DEVICES THEMSELVES. ON (B)(6), THE PATIENT'S CONDITION SUDDENLY CHANGED, AND THE PATIENT DIED. AFTER DEATH, A CT SCAN REVEALED THAT THE INTERNAL PRESSURE IN THE FALSE LUMEN HAD INCREASED DUE TO THE EXPANSION OF THE TRUE LUMEN CAUSED BY THE IMPLANTATION OF THE STENT GRAFTS, AND THE FALSE LUMEN HAD RUPTURED AROUND THE ABDOMEN. PHYSICIAN'S COMMENT: THE PHYSICIAN DID THE BEST THAT HE COULD WITH THE MEANS AVAILABLE TO THE HOSPITAL. THERE WAS NO OTHER CHOICE. PHYSICIAN'S COMMENT ON CAUSAL RELATIONSHIP: THE MEANS TAKEN FOR THE AORTIC DISSECTION WERE THE BEST, GIVEN THE EMERGENCY TREATMENT ON (B)(6), BUT THE TRUE LUMEN EXPANDED MORE THAN EXPECTED. OPERATION TYPE: EMERGENCY TEVAR FOR AN AORTIC DISSECTION, NO BLOOD LOSS, ANCILLARY DEVICE USED: 28-N4-32-164-28U (LOT#: 2407130037), NO IMAGE AVAILABLE DUE TO HOSPITAL POLICY, NO PRE-CASE PLAN AVAILABLE DUE TO HOSPITAL POLICY, NO ADDITIONAL INFORMATION AVAILABLE DUE TO HOSPITAL POLICY, (TC#: (B)(4))." PATIENT OUTCOME: "THE PATIENT DIED." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 2 IS BEING REPORTED UNDER MDR 2247858-2025-00122. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935322&pc=BZD | 2518422-2025-104977 | 2021-01-01 05:00 | Death | RESPIRONICS, INC. | 2025-04-30 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Degraded | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A USER OF A REMSTAR PRO C-FLEX+ DEVICE PASSED AWAY. THE PATIENT'S WIFE HAS THE DEVICE AND REQUESTED RETURN LABELS IN ORDER TO RETURN THE DEVICE. THE PATIENT PASSED AWAY IN 2021 AND THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935374&pc=OZD | 1220648-2025-27967 | 2025-04-09 04:00 | Death | ABIOMED, INC. | 2025-04-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation of Vessels | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 79-YEAR-OLD MALE WITH AN IMPELLA CP DUE TO ACUTE MYOCARDIAL INFARCTION. IT WAS REPORTED THAT DURING THE IMPELLA CP PROCEDURE ONE PERCLOSE ATTEMPTED AND FAILED DUE TO CALCIUM IN THE VESSEL. THE IMPELLA STARTED SUCCESSFULLY TO SUPPORT THE PATIENT. THE NEXT DAY THE PATIENT HAD A LARGE HEMATOMA AT IMPELLA SITE. ANTICOAGULATION WAS HELD. IT WAS ALSO REPORTED THAT THE PATIENT STARTED TO COUGH UP CLOTTED BLOOD, WHICH TURNED INTO FRANK BLOOD. PATIENT EMERGENTLY INTUBATED AND TWO UNITS OF BLOOD WERE ADMINISTERED. BLEEDING CONTINUED TO WORSEN IN ET TUBE, PATIENT WAS BRONCHED THREE MORE TIMES TO IDENTIFY BLEEDING, BUT NO BLEEDING WAS FOUND. PATIENT CONTINUED TO DETERIORATE, AND DECISION WAS MADE FOR COMFORT CARE. PATIENT PASSED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935894&pc=CAW | 3013095415-2025-00394 | 2025-01-13 05:00 | Death | INVACARE FLORIDA | 2025-04-30 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200890 |  | 1 | Event Description: IT WAS REPORTED TO VENTEC THAT A HOSPICE PATIENT WAS ON AN INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR WHEN HE PASSED AWAY. THE REPORTER, A DISTRIBUTOR, ADVISED THAT THEY DID NOT HAVE ANY FURTHER INFORMATION ABOUT THE PATIENT'S HISTORY/TIME IN HOSPICE. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN DECEMBER OF 2013, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE "NONE". VENTEC CONTINUES TO INVESTIGATE THE REPORTED ISSUE. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21935998&pc=LWS | 3008973940-2025-03011 | 2025-04-18 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-30 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Signal Artifact/Noise; Break; Fracture; Over-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT A LEAD INTEGRITY ALERT (LIA) HAD TRIGGERED FOR THE RIGHT VENTRICULAR (RV) LEAD DUE TO ELEVATED SENSING INTEGRITY COUNTER (SIC) AND NON-SUSTAINED VENTRICULAR TACHYCARDIA NOISE EPISODES. OVERSENSING OF A NON-PHYSIOLOGIC MORPHOLOGY WAS ALSO OBSERVED IN SEVERAL EPISODES ON THE RVTIP TO RVRING EGM (ELECTROGRAM) FOR THE LEAD. ADDITIONALLY, A CHEST X-RAY SEEMS TO SHOW BREAKING OF THE LEAD TO THE SIDE OF THE DEVICE, THUS A FRACTURE WAS SUSPECTED. THE RV LEAD WAS DEACTIVATED AND REMAINS IN-SITU. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT IT WAS THE THERAPIES THAT WERE DEACTIVATED AND THEREFORE THE LEAD REMAINS IN USE. IT WAS NOTED THAT THE PATIENT THEN SELF DISCHARGED, AGAINST MEDICAL ADVICE, DUE TO GOING TRAVELLING.Event Description: IT WAS FURTHER REPORTED, FROM ADDITIONAL INFORMATION OBTAINED FROM THE HOSPITAL, THAT APPROXIMATELY 12 DAYS LATER THE PATIENT HAD DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. ANALYSIS OF THE DEVICE MEMORY ALSO INDICATED NOISE, THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET, AND OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21936014&pc=NPT | 2025587-2025-03249 | 2024-12-25 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-30 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Renal Failure; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, COMPLETE HEART BLOCK (CHB) WAS OBSERVED. ADDITIONALLY, AN ESCAPE RHYTHM WAS NOT REACHED AND CARDIOPULMONARY ARREST (CPA) OCCURRED. SUBSEQUENTLY, RESUSCITATION WAS PERFORMED, HOWEVER, THE PATIENT'S LEVEL OF CONSCIOUSNESS DID NOT IMPROVE. THE PATIENT WAS GIVEN A 4 MILLIGRAM (MG) TRANSDERMAL BETA BLOCKER PATCH AND THERE WAS NO RECURRENCE OF CHB. APPROXIMATELY 2 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, THE PATIENT'S BLOOD PRESSURE DECREASED DURING DIALYSIS. THE PATIENT WAS ADMINISTERED ALBUMIN AND RED BLOOD CELLS (RBC) WHILE DIALYSIS WAS BEING PERFORMED FOR 6 HOURS, BUT IT BECAME MORE DIFFICULT. IT WAS DETERMINED THAT DIALYSIS WOULD NOT BE PERFORMED ANY FURTHER AND WOULD THEREFORE BE DISCONTINUED. APPROXIMATELY 3 MONTHS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT'S BLOOD PRESSURE DECREASED AND THE PATIENT DIED. THE CAUSE OF DEATH WAS RENAL RELATED. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND VALVE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. B3: DATE IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21936152&pc=CBK | 2518422-2025-044508 | 2021-11-23 05:00 | Death | RESPIRONICS, INC. | 2025-04-30 04:00 | CBK | TRILOGY 100 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A TRILOGY 100 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. LIMITED INFORMATION WAS PROVIDED BY THE CUSTOMER. HOWEVER, THE CUSTOMER DID INDICATE THAT THE PATIENT EXPIRED DUE TO A NON-SPECIFIC NATURAL DISEASE PROGRESSION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THE DEVICE POWERED ON AS IT SHOULD. NO ADDITIONAL REPAIRS WERE PERFORMED. IT WAS NOTED THAT THE DEVICE WAS NOT NEEDED FOR INVENTORY AND THE UNIT WILL BE SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21936441&pc=DSQ | 2916596-2025-02827 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN HOSPICE AFTER ELECTIVELY TURNING THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) OFF. THE PATIENT WAS ADMITTED AFTER MULTIPLE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SHOCKS. IT WAS DECIDED TO TURN THE ICD OFF AND THEN THE PATIENT WAS TRANSITIONED TO HOSPICE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. THE CAUSE OF THE ICD SHOCKS WAS VENTRICULAR TACHYCARDIA WITH NO SYMPTOMS. IT WAS UNKNOWN IF THE PATIENT HAD A HISTORY OF ARRHYTHMIA PRIOR TO LVAD IMPLANT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21936601&pc=DSQ | 2916596-2025-02678 | 2025-04-08 04:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REVISION A IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21936630&pc=DXF | 2124215-2025-27594 | 2025-02-06 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-04-30 04:00 | DXF | VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING AN ATRIAL FIBRILLATION WATCHMAN PROCEDURE, A VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE KIT WAS SELECTED FOR USE. THE PHYSICIAN COMPLETED TRANSEPTAL PUNCTURE USING ONLY INTRACARDIAC ECHOCARDIOGRAPHY (ICE), THEN THE BLOOD PRESSURE DROPPED. AN EFFUSION WAS NOTICED. THE PATIENT WAS TAKEN TO SURGERY AS THE DRAIN WAS UNABLE TO CONTROL THE SITUATION. AN OPEN CHEST SURGERY WAS PERFORMED, BUT A HOLE IN THE LEFT ATRIAL APPENDAGE WAS NOTICED AND THE BLOOD COULD NOT BE CONTROLLED AND THE PATIENT DID NOT MAKE IT. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). IT WAS HARD TO MAKE A TRANSEPTAL DUE TO ANATOMY AND ICE ONLY. NO MALFUNCTIONS WITH THE VERSACROSS DEVICES WERE REPORTED. ACT WAS THERAPEUTIC. IT WAS BELIEVED THAT THE FARADRIVE MAY HAVE CAUSE THE PERFORATION IN THE LAA. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21936902&pc=KDI | 0008030665-2025-00984 | 2025-04-10 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-30 04:00 | KDI | 2008T HD SYS. CDX W/BIBAG BLUE STAR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K173972 |  | 1 | Event Description: IT WAS REPORTED THAT A HEMODIALYSIS (HD) PATIENT WAS IN TREATMENT MODE UTILIZING A FRESENIUS 2008T HEMODIALYSIS (HD) MACHINE WHEN THEY REQUESTED TO BE TAKEN OFF THE DEVICE DUE TO NOT FEELING WELL (UNSPECIFIED). THE NURSE PROCEEDED TO RINSE BACK THE PATIENT'S BLOOD TO END THE TREATMENT. AT THIS TIME, THE NURSE CALLED FOR THE RAPID RESPONSE TEAM TO ASSIST AS THE PATIENT STILL WAS NOT FEELING WELL. THE RAPID RESPONSE TEAM WAS ATTEMPTING TO HELP THE PATIENT, AND IN THE PROCESS, THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THERE WERE NO ISSUES OR PROBLEMS WITH THE 2008T MACHINE DURING THE TREATMENT. NO DEVICE EVALUATION WAS REQUESTED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE DIALYSIS UNIT MANAGER. THE PATIENT WAS IN THE HOSPITAL INTENSIVE CARE UNIT (ICU) RECEIVING HD TREATMENT. THE PATIENT WAS AT THE END OF THE TREATMENT ON (B)(6) 2025 WHEN THEY REPORTED NOT FEELING WELL. THE TREATMENT WAS ENDED, AND THE PATIENT'S BLOOD WAS RINSED BACK. THE PATIENT SUBSEQUENTLY PASSED AWAY, BUT NO CAUSE OF DEATH WAS PROVIDED. IT WAS NOT REPORTED IF THE PATIENT RECEIVED ANY RESUSCITATIVE EFFORTS. THE MANAGER STATED THIS PATIENT WAS AN ICU PATIENT AND VERY ILL. THERE WERE NO MACHINE OR TREATMENT ISSUES. THE EVENT WAS REPORTED AS PER FACILITY PROTOCOL. THE MACHINE WAS SEQUESTERED PER PROTOCOL AND HAS UNDERGONE ALL EVALUATION AND TESTING WITH NO ISSUES FOUND. THE MACHINE WILL BE PUT BACK INTO SERVICE PENDING SIGN-OFF FROM THE MEDICAL DIRECTOR. NO FURTHER INFORMATION WAS PROVIDED RELATED TO THIS ADVERSE EVENT.Event Description: IT WAS REPORTED THAT A HEMODIALYSIS (HD) PATIENT WAS IN TREATMENT MODE UTILIZING A FRESENIUS 2008T HEMODIALYSIS (HD) MACHINE WHEN THEY REQUESTED TO BE TAKEN OFF THE DEVICE DUE TO NOT FEELING WELL (UNSPECIFIED). THE NURSE PROCEEDED TO RINSE BACK THE PATIENT¿S BLOOD TO END THE TREATMENT. AT THIS TIME, THE NURSE CALLED FOR THE RAPID RESPONSE TEAM TO ASSIST AS THE PATIENT STILL WAS NOT FEELING WELL. THE RAPID RESPONSE TEAM WAS ATTEMPTING TO HELP THE PATIENT, AND IN THE PROCESS, THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THERE WERE NO ISSUES OR PROBLEMS WITH THE 2008T MACHINE DURING THE TREATMENT. NO DEVICE EVALUATION WAS REQUESTED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE DIALYSIS UNIT MANAGER. THE PATIENT WAS IN THE HOSPITAL INTENSIVE CARE UNIT (ICU) RECEIVING HD TREATMENT. THE PATIENT WAS AT THE END OF THE TREATMENT ON (B)(6) 2025 WHEN THEY REPORTED NOT FEELING WELL. THE TREATMENT WAS ENDED, AND THE PATIENT¿S BLOOD WAS RINSED BACK. THE PATIENT SUBSEQUENTLY PASSED AWAY, BUT NO CAUSE OF DEATH WAS PROVIDED. IT WAS NOT REPORTED IF THE PATIENT RECEIVED ANY RESUSCITATIVE EFFORTS. THE MANAGER STATED THIS PATIENT WAS AN ICU PATIENT AND VERY ILL. THERE WERE NO MACHINE OR TREATMENT ISSUES. THE EVENT WAS REPORTED AS PER FACILITY PROTOCOL. THE MACHINE WAS SEQUESTERED PER PROTOCOL AND HAS UNDERGONE ALL EVALUATION AND TESTING WITH NO ISSUES FOUND. THE MACHINE WILL BE PUT BACK INTO SERVICE PENDING SIGN-OFF FROM THE MEDICAL DIRECTOR. NO FURTHER INFORMATION WAS PROVIDED RELATED TO THIS ADVERSE EVENT. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN BEING ON HEMODIALYSIS (HD) UTILIZING THE 2008T HD MACHINE AND THE PATIENT EVENT OF NOT FEELING WELL WITH SUBSEQUENT DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE 2008T MACHINE. ADDITIONALLY, THERE IS NO ALLEGATION OF A MACHINE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. IT WAS REPORTED THAT THIS PATIENT WAS VERY ILL AND ADMITTED TO THE ICU OF THE HOSPITAL AT THE TIME OF TREATMENT. THERE WERE NO REPORTED ISSUES WITH THE MACHINE AND NO ISSUES FOUND DURING POST EVALUATION AND TESTING. ALTHOUGH A CAUSE OF DEATH WAS NOT PROVIDED, THE PATIENT WAS REPORTEDLY VERY ILL. ALTHOUGH MAINTENANCE DIALYSIS PREVENTS DEATH FROM UREMIA, MORTALITY AMONG PATIENTS WITH END-STAGE KIDNEY DISEASE (ESKD) REMAINS HIGH. AMONG PATIENTS WITH ESKD, CARDIOVASCULAR DISEASE ACCOUNTS FOR APPROXIMATELY 50 PERCENT OF DEATHS. INFECTIONS ARE THE SECOND MOST COMMON CAUSE OF DEATH, FOLLOWED BY WITHDRAWAL FROM DIALYSIS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE FRESENIUS 2008T HEMODIALYSIS MACHINE CAN BE EXCLUDED AS CAUSING THE PATIENT'S DEATH. Manufacturer Narrative: PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. ADDITIONALLY, NO ON-SITE EVALUATION WAS PERFORMED BY A FRESENIUS FIELD SERVICE TECHNICIAN (FST). A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21937277&pc=LZO | 3010536692-2025-00203 |  | Death | MICROPORT ORTHOPEDICS INC. | 2025-04-30 04:00 | LZO | PROFEMUR® TL CLASSIC STEM | Insufficient Information | Unspecified Infection |  |  | 1 | Event Description: ALLEGEDLY, A LITERATURE DOCUMENT WAS RECEIVED DETAILING A STUDY CONDUCTED IN GREECE (TOTTAS (2) 2020 - MINIMA SHORT STEM VERSUS STANDARD PROFEMUR TL STEM IN PRIMARY TOTAL HIP REPLACEMENT A COMPARATIVE STUDY). THE DOCUMENT MENTIONS A REVISION SURGERY DUE TO INFECTION.Event Description: ALLEGEDLY, A LITERATURE DOCUMENT WAS RECEIVED DETAILING A STUDY CONDUCTED IN GREECE (TOTTAS (2) 2020 - MINIMA SHORT. STEM VERSUS STANDARD PRO-FEMUR TL STEM IN PRIMARY TOTAL HIP REPLACEMENT A COMPARATIVE STUDY). THE DOCUMENT. MENTIONS A REVISION SURGERY DUE TO INFECTION. Manufacturer Narrative: THIS EVENT WILL BE UPDATED ONCE THE INVESTIGATION IS COMPLETE. TRENDS WILL BE EVALUATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21937328&pc=NPT | 2025587-2025-03258 | 2024-09-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-30 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO MALNUTRITION AND HYPONATREMIA. SUBSEQUENTLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPSIS. THE CAUSE OF THE SEPSIS WAS UNSPECIFIED. OF NOTE, AN ECHOCARDIOGRAM REVEALED NO ABNORMALITIES IN THE VALVE. Manufacturer Narrative: CORRECTED DATA: H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21937681&pc=KDJ | 1416980-2025-02612 | 2025-04-02 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-04-30 04:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS. THE CAUSE OF PERITONITIS WAS NOT REPORTED. THE PATIENT WAS HOSPITALIZED FOR THE EVENT. TREATMENT WAS NOT REPORTED. THE PATIENT WAS HOSPITALIZED FOR FOUR DAYS AND DISCHARGED AFTER THE NEPHROLOGIST STATED THE PATIENT WAS "IN A BETTER STATE". THE DAY AFTER DISCHARGE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AND FURTHER REPORTED AS "SUDDEN DEATH". IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING BEFORE DEATH AND AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21937867&pc=NPT | 2025587-2025-03264 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-30 04:00 | NPT | EVOLUT FX VALVE | Material Invagination; Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC (LUNDERQUIST) GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: CONTINUATION OF D10: FIRST DCS PRODUCT ID: D-EVOLUTFX-34; PRODUCT LOT NUMBER: 0012423694; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE SECOND DCS PRODUCT ID: D-EVOLUTFX-34; PRODUCT LOT NUMBER: 0012449254; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE SECOND VALVE PRODUCT ID: EVOLUTFX-34; PRODUCT SERIAL NUMBER: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025 H10. RELATED REPORT NUMBERS MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. PRODUCT ID D-EVOLUTFX-34; PRODUCT LOT NUMBER: UNKNOWN; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE UPDATED DATA: B5. SECOND PARAGRAPH ADDED D10. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-34; PRODUCT LOT/SERIAL NUMBER NA; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE; EXPLANT DATE PRODUCT ID EVOLUTFX-34; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21937899&pc=NPT | 2025587-2025-03265 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-30 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Difficult to Advance; Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC (LUNDERQUIST) GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. Manufacturer Narrative: CONTINUATION OF D10. PRODUCT ID D-EVOLUTFX-34; PRODUCT LOT NUMBER: UNKNOWN; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE UPDATED DATA: B5. SECOND PARAGRAPH ADDED D4. D10. H4. H6. H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVOLUTFX-34; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE; EXPLANT DATE PRODUCT ID EVOLUTFX-34; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: CONTINUATION OF D10: FIRST VALVE PRODUCT ID: EVOLUTFX-34; PRODUCT SERIAL NUMBER: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTED SECOND DCS PRODUCT ID: D-EVOLUTFX-34; PRODUCT LOT NUMBER: 0012449254; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE SECOND VALVE PRODUCT ID: EVOLUTFX-34; PRODUCT SERIAL NUMBER: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025 H10. RELATED REPORT NUMBERS MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21937968&pc=NPT | 2025587-2025-03266 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-30 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC (LUNDERQUIST) GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC (LUNDERQUIST) GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: CONTINUATION OF D10: FIRST DCS PRODUCT ID: D-EVOLUTFX-34; PRODUCT LOT NUMBER: 0012423694; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE FIRST VALVE PRODUCT ID: EVOLUTFX-34; PRODUCT SERIAL NUMBER: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTED, SECOND DCS PRODUCT ID: D-EVOLUTFX-34; PRODUCT LOT NUMBER: 0012449254; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE H10. RELATED REPORT NUMBERS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6 .MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: SUPPLEMENTAL 002 WAS ERRONEOUSLY SUBMITTED WITH AN INCORRECT AWARE DATE. THE UPDATED INFORMATION WAS RECEIVED IN MAY, NOT APRIL. LET THIS CORRECTION REPORT REFLECT THE ACCURATE AWARE DATE OF MAY 14, 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. PRODUCT ID D-EVOLUTFX-34; PRODUCT LOT NUMBER: UNKNOWN; PRODUCT TYPE: 0195-HEART VALVES: NON-IMPLANTABLE DEVICE UPDATED DATA: B5. SECOND PARAGRAPH ADDED D10. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ID D-EVOLUTFX-34; PRODUCT LOT/SERIAL NUMBER NA; PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVOLUTFX-34; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21938003&pc=DSQ | 2916596-2025-02826 | 2025-04-19 04:00 | Death | THORATEC CORPORATION | 2025-04-30 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS FOLLOWED BY MULTI-SYSTEM ORGAN FAILURE (MSOF). THE PATIENT SUFFERED FROM SEPSIS FROM A NON-DEVICE RELATED SOURCE OR THE DRIVELINE, WHICH WAS NOT FURTHER DEFINED. SEPSIS LED TO PROGRESSIVE MSOF AND WAS RELATED TO THE OUTCOME. NO DEVICE ISSUES WERE MENTIONED AND THE DEVICE OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21938452&pc=OZD | 1220648-2025-27978 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-04-30 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA 5.5 PUMP PLACED FOR SUPPORT DURING A HIGH-RISK PROCEDURE. THE PUMP WAS SUPPORTING FOR SIX DAYS WHEN ACCIDENTALLY THE PUMP WAS PULLED OUT OF POSITION. THE PUMP WAS NOT RE-DELIVERED ACROSS THE VALVE TO THE LEFT VENTRICLE SO THE PUMP WAS EXPLANTED AND REPLACED. FURTHER DETAILS WERE RECEIVED STATING THAT THE PATIENT EXPIRED DUE TO A CAUSE UNRELATED TO THE PUMP.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: B1-ADVERSE EVENT B2-SELECTED DEATH AND ADDED DATE OF DEATH. B5-ADDITIONAL DETAILS PROVIDED STATES THAT THE PATIENT EXPIRED WHILE ON SUPPORT. H6-CLINICAL CODE 4580, IMPACT CODE 1802. H10-MEDICAL SAFETY OFFICER (MSO) REVIEW. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND ALTHOUGH THE CAUSE OF DEATH IS PROVIDED AND NOTED AS MULTIPLE ORGAN FAILURE (MOF), IT CANNOT BE SAID THAT THE LOSS/LACK OF HEMODYNAMIC SUPPORT HAD NOTHING TO DO WITH THE FINAL OUTCOME. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDED THE IMPELLA AS AN ASSOCIATED FACTOR TO THE OUTCOME. H1-UPDATED TO DEATH. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. CLINICAL STATES THAT THE PUMP WAS PULLED OUT INTO VENTRICLE WHEN THE PATIENT WAS CHANGING POSITION FOR AN X-RAY. PUMP WAS EXPLANTED AND REPLACED WITH ANOTHER PUMP. AS PER THE CLINICAL INFORMATION PROVIDED, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USE RELATED DUE TO IMPROPER TECHNIQUE. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27978-1. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27978. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21939467&pc=LZO | 3010536692-2025-00210 |  | Death | MICROPORT ORTHOPEDICS INC. | 2025-04-30 04:00 | LZO | PROFEMUR® TL CLASSIC STEM | Insufficient Information | Pain |  |  | 1 | Event Description: ALLEGEDLY, A LITERATURE DOCUMENT WAS RECEIVED DETAILING A STUDY CONDUCTED IN GREECE (TOTTAS (2) 2020 - MINIMA SHORT STEM VERSUS STANDARD PROFEMUR TL STEM IN PRIMARY TOTAL HIP REPLACEMENT A COMPARATIVE STUDY). THE DOCUMENT MENTIONS SIX CASES OF PAIN.Event Description: ALLEGEDLY, A LITERATURE DOCUMENT WAS RECEIVED DETAILING A STUDY CONDUCTED IN GREECE (TOTTAS (2) 2020 - MINIMA SHORT STEM VERSUS STANDARD PRO-FEMUR TL STEM IN PRIMARY TOTAL HIP REPLACEMENT A COMPARATIVE STUDY). THE DOCUMENT MENTIONS SIX CASES OF PAIN. Manufacturer Narrative: THIS EVENT WILL BE UPDATED ONCE THE INVESTIGATION IS COMPLETE. TRENDS WILL BE EVALUATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21939653&pc=NPS | 2135147-2025-02283 | 2024-04-05 04:00 | Death | ABBOTT MEDICAL | 2025-04-30 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Unspecified Infection; Respiratory Failure; Heart Failure/Congestive Heart Failure | P230007 | RWD2400493 | 11 | Event Description: IT WAS REPORTED THROUGH STS/ACC TVT REGISTRY - TRICLIP DATA THAT TRICLIP DEVICES MAY BE RELATED TO 11 ADVERSE EVENTS WHICH ARE CONSIDERED DEATH INCLUDING: HEART FAILURE, INFECTION, OTHER CARDIOVASCULAR REASON, OTHER NON-CARDIOVASCULAR REASON, PULMONARY, UNKNOWN DEATH. THE RELATIONSHIP OF THE DEATH TO THE TRICLIP DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. STS/ACC TVT REGISTRY - TRICLIP DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - RWD2400493. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 05 APRIL 2024 ¿ 16 SEPTEMBER 2024. PATIENTS¿ MEAN AGE IS 79 YEARS, RANGING FROM 64 TO 92 YEARS. 73% OF THE PATIENTS WERE MALE, 27% OF THE PATIENTS WERE FEMALE. AS OF 30 SEPT 2024, 477 PATIENTS WERE TREATED WITH THE TRICLIP SYSTEM.Event Description: Unique Complaint ID Number,Event Date,Date Entered,Manufacturer Aware Date,Brand Name,Generic Name,Model Number,Lot Number,Catalog Number,Serial Number,UDI Number,PMA / 510K Number,Date Returned to Manufacturer,Type of Reportable Event,Event Description,Manufacturer Narrative,Medical History,Patient Age,Patient Gender,Patient Weight,Date Implanted,Date Explanted,Health Effect Clinical Code,Health Effect Impact Code,Device Problem Code,Device Component Code,Investigation Type Code,Investigation Findings Code,Investigation Conclusion Code,Remedial Action Type,Latest Line Item Version
42134738,8/20/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 04/05/2024 in a 79 year old Male. On 08/20/2024, patient death was reported due to a non-cardiovascular reason. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to a non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,79,Male,70,4/5/2024,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
42559909,5/23/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 05/06/2024 in a 89 year old Male. On 05/23/2024, patient death was reported due to pulmonary issues. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to pulmonary issues was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,89,Male,82.3,5/6/2024,NI,2484,1802,2993,4755,4114;4119,3221,4315,NA,0
42597747,6/20/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 06/11/2024 in a 88 year old Male. On 06/20/2024, patient death was reported due to pulmonary issues. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to pulmonary issues was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,88,Male,79.8,6/11/2024,NI,2484,1802,2993,4755,4114;4119,3221,4315,NA,0
43071513,7/24/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 07/15/2024 in a 75 year old Male. On 07/24/2024, patient death was reported due to a non-cardiovascular reason. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to a non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,75,Male,70.8,7/15/2024,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
42546089,8/9/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 08/06/2024 in a 92 year old Male. On 08/09/2024, patient death was reported due to pulmonary issues. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to pulmonary issues was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,92,Male,66.9,8/6/2024,NI,2484,1802,2993,4755,4114;4119,3221,4315,NA,0
42546129,10/4/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 08/07/2024 in a 64 year old Female. On 10/04/2024, patient death was reported due to a non-cardiovascular reason. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to a non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,64,Female,92.08,8/7/2024,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
42933830,10/12/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 08/12/2024 in a 88 year old Female. On 10/12/2024, patient death was reported due to Infection. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to Infection was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,88,Female,57.3,8/12/2024,NI,1930,1802,2993,4755,4114;4119,3221,4315,NA,0
42630815,9/23/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 08/14/2024 in a 76 year old Male. On 09/23/2024, patient death was reported due to a cardiovascular reason. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to a cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,76,Male,93,8/14/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
42677160,8/19/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 08/14/2024 in a 87 year old Female. On 08/19/2024, patient death was reported due to pulmonary issues. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to pulmonary issues was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,87,Female,37.4,8/14/2024,NI,2484,1802,2993,4755,4114;4119,3221,4315,NA,0
42767018,11/6/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 09/05/2024 in a 67 year old Male. On 11/06/2024, patient death was reported due to Heart failure. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,67,Male,76.1,9/5/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
42834933,10/24/2024,4/30/2025,1/31/2025,TriClip G4 Delivery System (TCDS),"Tricuspid valve repair device, percutaneously delivered",Unknown,Unknown,Unknown,NI,NI,P230007,NA,D,"It was reported through STS/ACC TVT Registry data that a TriClip was implanted on 09/16/2024 in a 64 year old Male. On 10/24/2024, patient death was reported due to an unknown issue. The relationship of the Death to the TriClip could not be determined based on the limited data received from the registry.","An event of death due to an unknown issue was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,64,Male,70,9/16/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br> Manufacturer Narrative: THIS REPORT IS FOR QUARTER 1 TIMEFRAME BETWEEN JANUARY 1, 2025 AND MARCH 31 2025 11 ADVERSE EVENTS THAT COULD BE RELATED TO THE TRICLIP DEVICE WERE REPORTED INCLUDING HEART FAILURE, INFECTION, OTHER CARDIOVASCULAR REASON, OTHER NON-CARDIOVASCULAR REASON, PULMONARY, UNKNOWN DEATH. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED AND THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. A RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENT. ALL REPORTED DEVICE AND PATIENT COMPLICATIONS ARE INCLUDED IN THE DEVICE RISK ASSESSMENT; THEREFORE, THE COMPLAINT IS A FORESEEABLE EVENT. THE CURRENT OCCURRENCE RATES CONTINUE TO REMAIN WITHIN THE DOCUMENTED RATES IN THE DEVICE RISK ASSESSMENT. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT ISSUE THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2. THE DATE OF DEATH USED IS THE EARLIEST DEATH DATE. B3 - DATE OF EVENT USED IS THE EARLIEST EVENT DATE. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21940232&pc=BZD | 2518422-2025-105026 | 2025-04-07 04:00 | Death | RESPIRONICS, INC. | 2025-04-30 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTO SYSTEM ONE DEVICE HAS PASSED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION THAT A PATIENT WITH A REMSTAR AUTO SYSTEM ONE DEVICE HAS PASSED. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE SERVICE REPORT NOTED THAT THE DEVICE WAS FUNCTIONAL. A VISUAL INSPECTION WAS PERFORMED ON THE SOUND ABATEMENT FOAM, AND IT WAS STILL INTACT. THERE WAS NO EVIDENCE OF FOAM PARTICLES FOUND IN THE ASSEMBLY. THE SERVICE REPORT NOTED ERROR CODES IN THE DEVICE HISTORY BUT NONE THAT AFFECTED THE FUNCTIONALITY OF THE DEVICE. THE DEVICE WAS SCRAPPED BY THE THIRD-PARTY SERVICE CENTER. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21925171&pc=NPT | 2025587-2025-03180 | 2024-06-13 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Non specific EKG/ECG Changes; Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, SHORTLY AFTER THE VALVE PLACEMENT, ATRIO-VENTRICULAR (AV) BLOCK WAS OBSERVED. NO TREATMENT WAS REPORTED. OF NOTE, THE PATIENT HAD PRE-EXISTING RIGHT BUNDLE BRANCH BLOCK (RBBB). 1 YEAR AND 1 MONTH FOLLOWING IMPLANT OF THE VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, IT WAS UNKNOWN IF THE DEATH WAS RELATED TO THE VALVE OR VALVE IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21925519&pc=OZD | 1220648-2025-27928 | 2025-04-04 04:00 | Death | ABIOMED, INC. | 2025-04-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT EXPERIENCED LEG ISCHEMIA. THERE WAS NO FLOW PAST THE IMPELLA SHEATH. A FEMOROFEMORAL BYPASS WAS PERFORMED. THE PATIENT'S PULSE RECOVERED. THE PHYSICIAN CONSIDERED ESCALATING THE PATIENT TO AN IMPELLA 5.5 OR AXILLARY IMPELLA PLACEMENT.Event Description: IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED SEVEN MINUTES AFTER EXPLANT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS IT IS STILL SUPPORTING THE PATIENT. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LIMB ISCHEMIA: "LIMB ISCHEMIA CAN OCCUR DURING IMPELLA SUPPORT. WHEN MAKING A DETERMINATION AS TO THE CAUSE OF THE LIMB ISCHEMIA, THE FOLLOWING CLINICAL INFORMATION IS NECESSARY: PATIENTS PRE-MORBID CONDITION AND PERI-PROCEDURAL CONDITION. ANY CONCOMITANT MEDICATION. VASCULAR SIZE OR VARIATIONS THEREOF. DETAILED DESCRIPTION OF THE SYMPTOMS EXPERIENCED BY THE PATIENT. PHYSICAL EXAMINATION FINDINGS, INCLUDING PULSE ASSESSMENT AND VISUAL EXAMINATION OF THE AFFECTED LIMB. DIAGNOSTIC IMAGING RESULTS, SUCH AS DOPPLER ULTRASOUND, ANGIOGRAPHY, OR MAGNETIC RESONANCE ANGIOGRAPHY. LABORATORY TEST RESULTS, INCLUDING BLOOD TESTS FOR MARKERS OF INFLAMMATION OR CLOTTING DISORDERS. ANY RELEVANT INFORMATION ABOUT RISK FACTORS FOR PERIPHERAL ARTERIAL DISEASE, SUCH AS SMOKING, DIABETES, OR HYPERTENSION. PATIENT'S RESPONSE TO PREVIOUS TREATMENTS OR INTERVENTIONS FOR VASCULAR CONDITIONS. WITH A RETURNED IMPELLA, THE MAX OD COULD BE ASSESSED TO ENSURE IT IS WITHIN SPECIFICATION. THE PRODUCT COULD ALSO BE EVALUATED FOR ANY BURRS OR SHARP EDGES. ADDITIONALLY, THE CLINICAL INFORMATION ABOVE WOULD NEED TO BE CONSIDERED IN THE CONTEXT OF THE RETURNED PRODUCT. TO DETERMINE THE TRUE ROOT CAUSE OF LIMB ISCHEMIA, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED." B2 REVISED TO REFLECT THE PATIENT'S OUTCOME OF DEATH. B5 REVISED TO REFLECT THE PATIENT OUTCOME OF DEATH. D6B EXPLANT DATE ADDED. H1 REVISED TO DEATH. H6 HEALTH EFFECT - IMPACT CODE 1802 ADDED TO REFLECT PATIENT OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21925668&pc=OZD | 1220648-2025-27929 | 2025-04-09 04:00 | Death | ABIOMED, INC. | 2025-04-29 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem |  | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT DURING USE ON A PATIENT THE AUTOMATED IMPELLA CONTROLLER FAILURE "VERIFICATION FAILED, EXCHANGE THE CONSOLE" OCCURRED. BECAUSE THE HOSPITAL DID NOT HAVE A BACKUP CONSOLE, THE AIC WAS SWITCHED OFF AND COULD NOT SWITCHED ON AGAIN. THE PATIENT WAS IN A VERY CRITICAL CONDITION WITH CARDIAC ARREST AND EXPIRED. NO PATIENT INFORMATION WAS PROVIDED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - CONTROLLER FAILURE (RED ALARM) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. A2 PATIENT AGE UPDATED. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27929. Manufacturer Narrative: THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21925995&pc=DSQ | 2916596-2025-02834 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-04-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE REVIEW ARTICLE, "PULSATILITY OF PULMONARY ARTERY WAVEFORMS BEFORE AND AFTER LEFT VENTRICULAR ASSIST DEVICE IS ASSOCIATED WITH RIGHT VENTRICULAR FAILURE¿ THAT LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENT HAD AN INCIDENCE OF RIGHT VENTRICULAR FAILURE (RVF) OF ROUGHLY 35% CONTINUED TO PLAGUE CONTEMPORARY DEVICES AND IS ASSOCIATED WITH SIGNIFICANT LONG-TERM MORBIDITY AND MORTALITY. NUMEROUS SCORING SYSTEMS HAVE BEEN INVESTIGATED IN THE ATTEMPT TO PREDICT POSTIMPLANT RV FAILURE, BUT UNFORTUNATELY, DESPITE INCREASING SOPHISTICATION OF THESE SCORING SYSTEMS, OVERALL DISCRIMINATION OF PATIENTS AT RISK FOR RV FAILURE, PARTICULARLY LATE RV FAILURE, WAS MODEST. ALL PATIENTS WERE EVALUATED AT A SINGLE INSTITUTION UNDERGOING LVAD IMPLANTATION BETWEEN 2016 AND 2023 WHO HAD A PA PRESSURE SENSOR IN SITU AT THE TIME OF IMPLANTATION AND HAD AT LEAST TWO BUT UP TO TEN READINGS NOT REJECTED BY CARDIOMEMS FOR POOR SIGNAL, BEFORE AND AFTER LVAD IMPLANTATION (15/19, 79% HAD 10 READINGS BOTH BEFORE AND AFTER IMPLANT). ON AVERAGE, PATIENTS WERE 55 PLUS OR MINUS 13 YEARS OLD, PREDOMINANTLY WHITE (89%) MALES (68%) IMPLANTED WITH A HEARTMATE 3 (84%) AS DESTINATION THERAPY (68%). WITHIN THE COHORT, SEVEN PATIENTS (37%) EXPERIENCED RV FAILURE. PREIMPLANT PPI WAS ASSOCIATED WITH THE DEVELOPMENT OF RV FAILURE (AREA UNDER THE CURVE (AUC) = 0.79, P = 0.012) WITH OPTIMAL STRATIFICATION AT 0.25. SIMILARLY, POSTIMPLANT PPI WAS ASSOCIATED WITH THE DEVELOPMENT OF RV FAILURE (AUC = 0.71, P = 0.05) WITH OPTIMAL STRATIFICATION AT 0.35. IT WAS BELIEVED THAT THE HIGH NEGATIVE PREDICTIVE VALUE OF THE QUANTITY COMPLEMENTS CURRENT SCORING SYSTEMS, WHICH HAVE A HIGH POSITIVE PREDICTIVE VALUE FOR RVF ON LVAD THERAPY, SUCH AS MARKERS OF LIVER DYSFUNCTION. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B: DATE OF DEATH ESTIMATED AS EVENT DATE, WHICH HAS BEEN ENTERED AS (B)(6) 2023 AS PATIENTS WERE IMPLANTED BETWEEN 2016 AND 2023. AUTHOR INFORMATION: ROSENBAUM, A. N., BEHFAR, A., ROSSMAN, T. L., & LUNDGREN, S. W. (2025). PULSATILITY OF PULMONARY ARTERY WAVEFORMS BEFORE AND AFTER LEFT VENTRICULAR ASSIST DEVICE IS ASSOCIATED WITH RIGHT VENTRICULAR FAILURE. ASAIO JOURNAL, 71(3), E38-E40. DOI:HTTP://DX.DOI.ORG/10.1097/MAT.0000000000002272. DEPARTMENT OF CARDIOVASCULAR MEDICINE, DIVISION OF CIRCULATORY FAILURE, MAYO CLINIC, ROCHESTER, MN, UNITED STATES, WILLIAM J. VON LIEBIG CENTER FOR TRANSPLANTATION AND CLINICAL REGENERATION, MAYO CLINIC, ROCHESTER, MN, UNITED STATES, VAN CLEVE CARDIAC REGENERATIVE MEDICINE PROGRAM, CENTER FOR REGENERATIVE MEDICINE, MAYO CLINIC, ROCHESTER, MN, UNITED STATES ROSENBAUM.ANDREW@MAYO.EDU. DIVISION OF ENGINEERING, MAYO CLINIC, ROCHESTER, MN, UNITED STATES. DIVISION OF CARDIOVASCULAR MEDICINE, DEPARTMENT OF INTERNAL MEDICINE, UNIVERSITY OF NEBRASKA MEDICAL CENTER, OMAHA, NE, UNITED STATES. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE REPORTED INFORMATION WAS OBTAINED THROUGH A LITERATURE REVIEW. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AND OTHER SPECIFIC CASE/PATIENT INFORMATION ARE NOT AVAILABLE. THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926088&pc=MIH | 2247858-2025-00103 | 2025-04-03 04:00 | Death | BOLTON MEDICAL, INC. | 2025-04-29 04:00 | MIH | CUSTOM-MADE TREO ABDOMINAL STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding; Perforation of Vessels | P190015 |  | 1 | Event Description: "FENESTRATED TREO AUI WITH FENESTRATED LEG FOR RIGHT COMMON ILIAC ARTERY (CIA) - LEFT CIA WAS OCCLUDED. BEFORE BEGINNING THE FENESTRATED TREO PROCEDURE, DUE TO A STENOSIS IN THE RIGHT CIA, A BALLOON ANGIOPLASTY WAS PERFORMED. THE FENESTRATED TREO DELIVERY SYSTEM WAS THEN INSERTED, PLACED ACCURATELY AND DESHEATHED SUCCESSFULLY. CANNULATION OF THE VISCERAL VESSELS, FROM TOP ACCESS, WAS JUST ABOUT TO START. THE ANAESTHETIST INFORMED MR (B)(6) THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING. IT WAS ESTABLISHED THAT THE RCIA HAD BEEN DISSECTED AND MR (B)(6) DECIDED TO OPEN AND STOP THE BLEEDING. UNFORTUNATELY, THIS WAS NOT POSSIBLE AND THE PATIENT, VERY SADLY, DIED." PATIENT OUTCOME: "DEATH".Event Description: "FENESTRATED TREO AUI WITH FENESTRATED LEG FOR RIGHT COMMON ILIAC ARTERY (CIA) - LEFT CIA WAS OCCLUDED. BEFORE BEGINNING THE FENESTRATED TREO PROCEDURE, DUE TO A STENOSIS IN THE RIGHT CIA, A BALLOON ANGIOPLASTY WAS PERFORMED. THE FENESTRATED TREO DELIVERY SYSTEM WAS THEN INSERTED, PLACED ACCURATELY AND DESHEATHED SUCCESSFULLY. CANNULATION OF THE VISCERAL VESSELS, FROM TOP ACCESS, WAS JUST ABOUT TO START. THE ANAESTHETIST INFORMED MR (B)(6) THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING. IT WAS ESTABLISHED THAT THE RCIA HAD BEEN DISSECTED AND MR (B)(6) DECIDED TO OPEN AND STOP THE BLEEDING. UNFORTUNATELY, THIS WAS NOT POSSIBLE AND THE PATIENT, VERY SADLY, DIED." PATIENT OUTCOME: "DEATH." Manufacturer Narrative: BOLTON MEDICAL IS VOLUNTARILY REPORTING AN EVENT RELATED TO A TREO CUSTOM-MADE DEVICE. THE CUSTOM-MADE TREO DEVICES ARE NOT MARKETED IN THE US; HOWEVER, THEY ARE SIMILAR TO THE TREO ABDOMINAL STENT GRAFT DELIVERY SYSTEM APPROVED FOR SALE IN THE US (P190015). THE EVENT OCCURRED IN UNITED KINGDOM. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: BOLTON MEDICAL IS VOLUNTARILY REPORTING AN EVENT RELATED TO A TREO CUSTOM-MADE DEVICE. THE CUSTOM-MADE TREO DEVICES ARE NOT MARKETED IN THE US, HOWEVER THEY ARE SIMILAR TO THE TREO ABDOMINAL STENT GRAFT DELIVERY SYSTEM APPROVED FOR SALE IN THE US (P190015). THE EVENT OCCURRED IN UNITED KINGDOM. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926421&pc=DSP | MW5169609 | 2025-01-22 05:00 | Death | ARROW INTERNATIONAL, LLC (SUBSIDIARY OF TELEFLEX, INCORPORATED) | 2025-04-29 04:00 | DSP | IABP FIBEROPTIC AIB | Entrapment of Device; Material Rupture | Perforation of Vessels; Device Embedded In Tissue or Plaque |  |  | 1 | Event Description: IABP ALARMED FOR HELIUM LOSE AND SPOKE WITH TELEFLEX REGARDING SPECKS OF BLOOD ALONG THE TUBING. TELEFLEX ADVISED BALLOON RUPTURE AS EVIDENCED BY BLOOD SPECKS IN THE HELIUM TUBING. UNSUCCESSFUL ATTEMPTS TO REMOVE IABP CATHETER X 2 AND PT TAKEN EMERGENTLY TO THE OPERATING ROOM FOR REMOVAL AND REPAIR OF FEMORAL ARTERY. SURGERY SUCCESSFUL BUT PT DIED THE FOLLOWING DAY AND UNABLE TO EXCLUDE NEED FOR SURGICAL REMOVAL OF CATHETER AS A CONTRIBUTING FACTOR FOR PREVIOUSLY UNSTABLE PT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926480&pc=OZD | 1220648-2025-27937 | 2025-04-04 04:00 | Death | ABIOMED, INC. | 2025-04-29 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate; Device Sensing Problem | Arrhythmia; Ventricular Fibrillation; Ectopic Heartbeat | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS BROUGHT TO THE CATHETERIZATION LAB TODAY FOR A PLANNED PERCUTANEOUS CORONARY INTERVENTION OF THE SAPHENOUS VEIN GRAFT AND SELECTIVE SAPHENOUS VEIN GRAFT TO THE RIGHT CORONARY ARTERY. THE LEFT SUBCLAVIAN WAS SHOCKWAVED TO MAKE THE LEFT RADIAL APPROACH FEASIBLE. SAPHENOUS VEIN GRAFT TO OBTUSE MARGINAL WAS COMPLETED. DURING INTERVENTION, THE PATIENT HAD VENTRICULAR FIBRILLATION ARREST REQUIRING INTUBATION AND ADVANCED CARDIAC LIFE SUPPORT. THE IMPELLA CP WAS PLACED AS A BAILOUT. THE COMPLAINANT REPORTED THAT IMPELLA CP HAD ONGOING SUCTIONS WHILE IN USE. THE IMPELLA WAS NOT ABLE TO GO ABOVE P1 WITHOUT SUCTION. THE IMPELLA POSITION AND PRELOAD STATUS WERE REPORTEDLY GOOD. THE IMPELLA FLOWS APPEARED TO BE NORMAL. THE PHYSICIAN VERIFIED THE POSITION AND SUCTION ALARMS WERE TURNED OFF. THE PATIENT WAS ALSO SUPPORTED WITH EXTRACORPOREAL MEMBRANE OXYGENATION. IT WAS FURTHER REPORTED THAT ON DAY FIVE OF IMPELLA CP SUPPORT, THE PATIENT WAS TAKEN TO THE LAB FOR REPOSITIONING UNDER FLUOROSCOPY. THE IMPELLA ONLY MOVED FORWARD AT 2.7 CENTIMETERS THE PHYSICIAN COULD NOT MOVE IT ANY MORE FORWARD AND OPTED TO KEEP PLACEMENT AS IS. THE NEXT DAY, THE IMPELLA CP WAS STILL IN POOR POSITION AND WAS UNABLE TO ESTABLISH FULL FLOWS. THE PATIENT HAD RUNS OF TORSADE AND VENTRICULAR FIBRILLATION. THE IMPELLA WAS REMOVED IN AN ATTEMPT TO RELIEVE ECTOPY. AFTER THE IMPELLA WAS REMOVED, THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, LOW PUMP FLOW, VF/VT/AF/AT, AND OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. THE PRODUCT WAS NOT RETURNED FOR INVESTIGATION. SINCE THE CLINICAL REPORTS THAT ATTEMPTS TO REPOSITION THE PUMP DIDN'T RESOLVE THE ISSUE WAS CONFIRMED BY DATA LOGS, BUT WEANING EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) FLOWS ALIGNED WITH PUMP FLOW IMPROVEMENT ON DATA LOGS, THE ROOT CAUSE OF THE LOW PUMP FLOW WAS DETERMINED TO MOST LIKELY BE PATIENT CONDITION. ADDITIONALLY, THERE WERE NO ABNORMALITIES NOTED IN DATA LOGS AROUND THE TIME OF THE REPORTED COMPLICATION, SO THE ROOT CAUSE OF THE POSITIONING ISSUES COULD NOT BE DETERMINED. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. SINCE DATA LOG REVIEW SHOWED THE MOTOR CURRENT TO QUITE LOW DURING BOTH PHASES OF THE CARDIAC CYCLE EVEN THOUGH THE PLACEMENT SIGNAL (PS) WAS PULSATILE AND WITHIN SPECIFICATION, THE ROOT CAUSE OF THE OPTICAL SIGNAL ISSUE WAS DETERMINED TO MOST LIKELY BE A SUCTION CONDITION. THE SUCTION CONDITION WAS MOST LIKELY DUE TO THE PATIENT'S CONDITION, PER THE INVESTIGATION INTO LOW PUMP FLOW. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926559&pc=DSQ | 2916596-2025-02611 | 2025-04-04 04:00 | Death | THORATEC CORPORATION | 2025-04-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Liver Damage/Dysfunction; Hypovolemia; Hypervolemia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT POST-IMPLANT THE PATIENT HAD WORSENING LACTIC ACIDOSIS, LIVER DYSFUNCTION, FLUID OVERLOAD, AND HEIGHTENING PRESSOR REQUIREMENTS. THE FAMILY CHOSE TO MOVE TO COMFORT CARE, AND THE PATIENT PASSED PEACEFULLY AWAY ON (B)(6) 2025. THE PUMP WAS FUNCTIONING AS INTENDED.Event Description: THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT NOT PROVIDED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: H6: HEALTH EFFECT CLINICAL CODE, CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING HEPATIC DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926669&pc=BZD | 2518422-2025-104927 | 2025-04-04 04:00 | Death | RESPIRONICS, INC. | 2025-04-29 04:00 | BZD | BIPAP AUTO BIFLEX SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE SPOUSE OF A PATIENT REPORTED THE PATIENT PASSED AWAY. THE SPOUSE OF THE PATIENT REPORTED NO LONGER BEING IN POSSESSION OF THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION FOR THE DEVICE CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926767&pc=LYJ | 1644487-2025-10055 | 2025-04-04 04:00 | Death | HOUSTON | 2025-04-29 04:00 | LYJ | SYMMETRY GEN MODEL 8103 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUICIDE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: ADDITIONAL INFORMATION RECEIVED NOTING THAT THE DEATH IS NOT RELATED TO THE VNS DEVICE. THE DEATH FORM WAS FILLED OUT AND THE CAUSE OF DEATH WAS LISTED AS SUICIDE. THE VNS WAS NOT EXPLANTED AFTER THE PATIENT'S DEATH. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21926959&pc=OZD | 1220648-2025-27940 | 2025-04-03 04:00 | Death | ABIOMED, INC. | 2025-04-29 04:00 | OZD | IMPELLA CP | Fluid/Blood Leak | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTEDLY ARRIVED ONSITE WITH THE PATIENT ALREADY IN THE INTENSIVE CARE UNIT (ICU). INTRAOPERATIVELY, A MITRAL VALVE REPAIR WAS ADDED IN ADDITION TO CORONARY ARTERY BYPASS GRAFT TIMES FOUR. THE PATIENT DECOMPENSATED COMING OFF BYPASS AND AN INTRAAORTIC BALLOON PUMP (IABP) WAS INSERTED. IABP DID NOT PROVIDE ANY HEMODYNAMIC SUPPORT, SO THE PHYSICIAN WAS ASKED TO ASSIST IN INSERTING AN IMPELLA CP IN PLACE OF THE IABP. THE PATIENT WAS IN THE ICU ON FIVE TIMES INOTROPES/PRESSORS, SKIN MOTTLED AND WAS ACTIVELY CODING. DURING PREPARATION, THE AUTOMATED IMPELLA CONTROLLER (AIC) ALARMED THAT THE PURGE SYSTEM WAS OPEN WHEN IT WAS NOT, AFTER THE WHITE POWER CORD WAS PLUGGED IN, THE AIC KEPT SHOWING DIRECTIONS TO CONNECT THE GREY CABLE ARROW TO ARROW. A CONTROLLER FAILURE ALARM APPEARED. THE AIC WAS REPLACED AND THE SAME ISSUE OCCURRED. THE PURGE WAS DE-AIRED AND THE AIC AND PUMP WERE OPERATIONAL. THE CLINICAL REPRESENTATIVE DID NOT KNOW IF THIS WAS AN AIC ISSUE OR A PUMP ISSUE. THE PATIENT OUTCOME AT EXPLANT NOTED THAT THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE PUMP AND PURGE CASSETTE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: CLEANING, ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿ Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA AND PURGE LEAK ¿ PUMP HAS BEEN COMPLETED. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. THERE WAS NO PURGE LEAK-PUMP ISSUE FOUND DURING THE CASE SINCE SINCE PURGE LEAK ISSUE WAS RESOLVED AFTER SUCCESSFUL PRIMING WITH SECOND PURGE CASEATE. THE DEVICE FUNCTIONED/OPERATED TO SPECIFICATION DURING THE FIELD RUN AND RETURNED PRODUCT CONFIRMS NO ISSUE ON PURGE LEAK-PUMP.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927007&pc=LYJ | 1644487-2025-10056 | 2025-03-30 04:00 | Death | HOUSTON | 2025-04-29 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Pneumonia | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO PNEUMONIA COMPLICATIONS. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927292&pc=LWS | 2649622-2025-11560 | 2025-04-09 04:00 | Death | MPRI | 2025-04-29 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A LEAD INTEGRITY ALERT (LIA) ON THE RIGHT VENTRICULAR (RV) LEAD DUE TO HIGH-RATE NON-SUSTAINED EPISODES AND SHORT VENTRICULAR INTERVALS. IT WAS ALSO NOTED THERE WAS INTERMITTENT UNDERSENSING DURING TREATED FINE VENTRICULAR FIBRILLATION (VF) EPISODES WHICH DELAYED THERAPY. THE PATIENT PASSED AWAY DUE TO VENTRICULAR STORM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. ANALYSIS OF THE DEVICE MEMORY INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927366&pc=FKX | 0008030665-2025-00973 | 2025-04-18 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-04-29 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT EXPIRED ON (B)(6) 2025. THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT THIS PATIENT EXPIRED. THERE WAS NO FURTHER INFORMATION AVAILABLE AS THE PATIENT¿S FAMILY DID NOT WISH TO DISCLOSE ANY INFORMATION TO THE CLINIC. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. PER THE INFORMATION DOCUMENTED IN THE FILE, THIS EVENT IS UNRELATED TO PERITONEAL DIALYSIS (PD) THERAPY OR ANY FRESENIUS DEVICE(S) OR PRODUCT(S).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927450&pc=OZD | 1220648-2025-27944 | 2025-04-09 04:00 | Death | ABIOMED, INC | 2025-04-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE A PATIENT WAS IMPLANTED WITH AN IMPELLA CP, THE PATIENT EXHIBITED BLEEDING AT THE ACCESS SITE. THE VASCULAR SURGEON PLACED A TOBACCO POUCH SUTURE AND TREATED THE GROIN WITH A TABOTAMP. THREE UNITS OF BLOOD PRODUCTS WERE GIVEN TO THE PATIENT. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE ATTENDING PHYSICIAN STATED HE DOES NOT VIEW THE IMPELLA THERAPY AS A CAUSE OR IN CONNECTION WITH THE PATIENT'S DEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927924&pc=LWS | 2649622-2025-11571 | 2025-04-10 04:00 | Death | MPRI | 2025-04-29 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation; Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) DUE TO HIGH RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). THE RV LEAD ALSO EXHIBITED UNDERSENSING POSSIBLY DELAYING DETECTION AND THERAPY FOR A TREATED VENTRICULAR FIBRILLATION (VF) EPISODE. IT WAS NOTED THE RV LEAD ALSO EXHIBITED OVERSENSING. IN ADDITION, THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING AS WELL AS POSSIBLE SENSING OF VENTRICULAR ACTIVITY. THE RV LEAD AND RA LEAD REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD WAS REPROGRAMMED. TWO DAYS AFTER THE REPROGRAMMING THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: DDMB1D1 ICD IMPLANTED: (B)(6) 2019, 407652 LEAD IMPLANTED: (B)(6) 2011 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927955&pc=NPT | 2025587-2025-03189 | 2024-12-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-29 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Fever; Septic Shock; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A PERMANENT PACEMAKER WAS IMPLANTED DUE TO A COMPLETE HEART BLOCK (CHB). FOLLOWING THE IMPLANT OF THE PACEMAKER, THE PATIENT UNDERWENT REHABILITATION. IT WAS THEN REPORTED THAT APPROXIMATELY 19 DAYS FOLLOWING THE IMPLANT OF THE PERMANENT PACEMAKER, THE PATIENT PRESENTED WITH A FEVER AND MRSA BACTEREMIA. SUBSEQUENTLY, THE PACEMAKER WAS REMOVED, BUT THE INFECTION PERSISTED, AND THE PATIENT DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS SEPTIC SHOCK. Manufacturer Narrative: E1 - CORRECTED STREET ADDRESS FROM (B)(6) TO (B)(6). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21927993&pc=LYJ | 1644487-2025-00449 | 2023-03-30 04:00 | Death | LIVANOVA USA, INC. | 2025-04-29 04:00 | LYJ | PULSE GEN MODEL UNKNOWN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT A VNS PATIENT PASSED AWAY 3 YEARS AGO DUE TO A BLOOD CLOT. THE PATIENT COULD NOT BE IDENTIFIED WITH THE INFORMATION AVAILABLE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21928703&pc=OZD | 1220648-2025-27952 | 2025-04-09 04:00 | Death | ABIOMED, INC. | 2025-04-29 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT IMPELLA CP WAS PLACED EMERGENTLY ON A 44-YEAR-OLD MALE PATIENT. HEMOLYSIS WAS PRESENT ON LABS. BEDSIDE ECHO SHOWED DEPTH AT 5.5 CM. IMPELLA OUTLET ON AORTIC VALVE. PHYSICIAN NOTIFIED OF HEMOLYSIS AND REPOSITIONED AT BEDSIDE. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS: CLINICAL STATES THAT THE PUMP WAS PLACED WITH OUTLET ON THE AORTIC VALVE CAUSING LABS TO HEMOLYZE. PUMP WAS REPOSITIONED AND HEMOLYSIS RESOLVED. PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOGS SHOW "IMPELLA IN VENTRICLE" ALARMS AND PLACEMENT SIGNAL WAS SPREAD AND NO OTHER ABNORMALITIES WERE OBSERVED. THEREFORE, THE ROOT CAUSE OF THE HEMOLYSIS ISSUE WAS IMPROPER POSITIONING OF THE DEVICE DUE TO IMPROPER PLACEMENT TECHNIQUE. POSITIONING ISSUE: CLINICAL STATES THAT THE PUMP WAS IMPLANTED AND WAS FOUND TO FACING THE AORTIC VALVE WHICH WAS THEN REPOSITIONED. DATA LOGS SHOW INITIAL "IMPELLA IN VENTRICLE" ALARMS AT 15:22 ON (B)(6) 2025 WITH VENTRICULAR WAVEFORM WHICH RESOLVE SHORTLY AFTER. "IMPELLA IN VENTRICLE" ALARMS RETURN AGAIN AT ABOUT 17:00 AND ARE RESOLVED ONCE AGAIN BY 18:00. PUMP WAS NOT RETURNED. THE ROOT CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE THE PUMP WAS FOUND TO BE DEEP UPON ARRIVAL TO THE ICU AND UNKNOWN WHAT CAUSED IT TO GO DEEP. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21928817&pc=OZD | 1220648-2025-27953 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-04-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH A HISTORY OF VENTRICULAR FIBRILLATION (VF) AND WAS ON IMPELLA CP SUPPORT. BLEEDING FROM THE MEDIASTINUM OCCURRED POSSIBLY DUE TO CARDIOPULMONARY RESUSCITATION, AND THORACOTOMY WAS PERFORMED TO STOP BLEEDING. HEMODYNAMICS REMAINED UNSTABLE THEREAFTER AND VF RECURRED. AS A RESULT, THE PATIENT'S LIFE WAS DIFFICULT TO SAVE AND HE DIED. THE PATIENT'S CAUSE OF DEATH WAS FATAL ARRHYTHMIA. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21928831&pc=DSQ | 2916596-2025-02423 | 2025-03-31 04:00 | Death | THORATEC CORPORATION | 2025-04-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Hypoxia; Tachycardia; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT BECAME HYPOXIC WITH LACTIC UP TO 4.9. THEY WERE ABLE TO BE RE-INTUBATED FOR AIRWAY PROTECTION. THE PATIENT BECAME EXTENSIVELY HYPOTENSIVE PRIOR TO INTUBATION AND WAS STARTED ON LEVOPHED, VASOPRESSIN, AND EPINEPHRINE DROPS. THESE DROPS WERE MAXED OUT AND THE PATIENT REMAINED HYPOTENSIVE. THE PATIENT WAS IN VENTRICULAR TACHYCARDIA WITH HEART RATE IN 160S. AMIODARONE BOLUS WAS GIVEN. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) HAD CONTINUOUS LOW FLOWS WITH FLOWS BELOW 1.0. DESPITE ALL MEASURES, THE PATIENT'S HEMODYNAMICS DID NOT RESPOND. THEY WERE GIVEN NO HEMODYNAMIC IMPROVEMENT, AND THEIR HEART RATE WAS IN THE 50S. FLOWS WERE LESS THAN 0.5 CONSISTENTLY. ALL MEASURES WERE STOPPED, AND THE PATIENT WAS PRONOUNCED DECEASED AT 17:32. THE LVAD WAS TURNED OFF AND DISCONNECTED. Manufacturer Narrative: SECTION B1 CORRECTION MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE PROVIDED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING ARRHYTHMIA, RESPIRATORY FAILURE, AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 1 ALSO PROVIDES AN EXPLANATION OF EACH OF THE PUMP PARAMETERS, INCLUDING PUMP FLOW. THE IFU ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21928906&pc=NPT | 2025587-2025-03203 | 2024-11-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-29 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Patient Device Interaction Problem | Rupture; Cardiac Tamponade | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE MECHANISMS OF CARDIAC TAMPONADE DURING TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). OF THE 21 PATIENTS INCLUDED IN THE STUDY POPULATION, 15 UNDERWENT TAVR WITH A MEDTRONIC VALVE BRAND (COREVALVE, EVOLUT R, EVOLUT PRO, OR EVOLUT FX). VARIOUS TYPES OF GUIDEWIRES WERE USED DURING THE TAVR PROCEDURES, INCLUDING MEDTRONIC CONFIDA GUIDEWIRES. AMONG THE MEDTRONIC VALVE PATIENTS, SEVEN IN-HOSPITAL DEATHS OCCURRED FOLLOWING DIAGNOSIS AND INTERVENTIONS TAKEN TO ADDRESS CARDIAC TAMPONADE, YET NONE OF THE DEATHS WERE RECOUNTED IN DETAIL. FOR ONE OF THESE PATIENTS, THE TAMPONADE RESULTED FROM AN ANNULAR RUPTURE. OTHER ADVERSE EVENTS THAT OCCURRED IN PATIENTS TREATED WITH A MEDTRONIC VALVE: CARDIAC TAMPONADE NECESSITATING EITHER PERICARDIOCENTESIS OR BOTH PERICARDIOCENTESIS AND CARDIAC SURGERY. PACEMAKER-INDUCED RIGHT VENTRICLE PERFORATION (6 CASES), STIFF GUIDEWIRE-INDUCED LEFTVENTRICLE PERFORATION (5 CASES), SOFT GUIDEWIRE-INDUCED LEFT VENTRICLE PERFORATION (3 CASES), AND ANNULAR RUPTURE (1 CASE, AS NOTED ABOVE) WERE LISTED AS THE MECHANISMS OF TAMPONADE FOR THE MEDTRONIC PATIENTS. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH MAY HAVE BEEN IN USE WITH THE FOLLOWING: CONFIDA BRECKER CURVE GUIDEWIRE, PRODUCT ID: GWBC30, SERIAL/LOT: UNKNOWN, USE BY DATE: UNKNOWN, UDI: UNKNOWN. CITATION: NAOUM I, EITAN A, SLIMAN H, ET AL. CARDIAC TAMPONADE COMPLICATING TRANSCATHETER AORTIC VALVE REPLACEMENT: INSIGHTS FROM A SINGLE-CENTER REGISTRY. CJC OPEN. 2024;7(2):153-160. PUBLISHED 2024 NOV 14. DOI:10.1016/J.CJCO.2024.11.005 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. EARLIEST APPROVED MEDTRONIC PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21928910&pc=NIK | 3004209178-2025-07402 | 2013-06-19 04:00 | Death | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 2025-04-29 04:00 | NIK | VIVA¿ QUAD XT CRT-D | Device Sensing Problem | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. A PATIENT REPRESENTATIVE INDICATED THAT THE PATIENT WAS SHOCKED SIX TIMES BEFORE THE PATIENT DIED. ¿AN EVENT/SHOCK OCCURRED AFTER DIALYSIS, THE PATIENT REPORTEDLY WENT UP THE HILL TWO MILES AND STARTED TO GET SHOCKS, THE DIALYSIS WAS WEARING THE PATIENT DOWN. IT WAS THE NEXT MORNING THAT THE PATIENT COULD NOT WALK. ADDITIONALLY, IT WAS STATED THAT AT AN UNSPECIFIED TIME THE PATIENT WAS WORKING ON THE CAR, THERE WAS A FIRE AND WHEN THE PATIENT¿S SPOUSE WAS APPROACHING FROM THE HOME, THE PATIENT HAD RECEIVED A SHOCK. THE PATIENT¿S SON DROVE THE PATIENT UP TO (B)(6) WHEN THEY HAD ARRIVED AT THE DESTINATION THE PATIENT WAS UNRESPONSIVE AND PASSED AWAY.¿ ADDITIONAL INFORMATION WAS REQUESTED AND YIELDED NO INFORMATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929320&pc=MVK | 3008642652-2025-04365 | 2025-04-01 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-29 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 22:49:15 ON (B)(6) 2025. AT 09:19:20 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 09:19:56, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 09:20:21, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, TRANSITIONING TO AN IDIOVENTRICULAR RHYTHM @ 30 BPM. AT 09:24:00, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 09:25:05, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 09:34:53 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929334&pc=MVK | 3008642652-2025-04370 | 2025-01-31 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-29 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 21:38:22 ON (B)(6) 2025. AT 11:37:54 ON (B)(6) 2025 AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. AT 11:39:04, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 11:39:34 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929386&pc=MVK | 3008642652-2025-04369 | 2025-04-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-29 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SEVEN INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 20:21:17 ON (B)(6) 2025. AT 16:17:54 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION ARTIFACT. AT 16:19:11, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS AF @ 70 BPM WITH PVCS. AT 16:23:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 16:24:51, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY. AT 16:25:16, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY. AT 16:26:17, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 16:26:44, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 16:27:23, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 16:23:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 16:28:17, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 16:58:28 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929512&pc=NPT | 2015691-2025-03376 | 2025-04-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-29 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING PROCEDURE, THE VALVE WAS IMPLANTED IN THE INTENDED POSITION (90/10). HOWEVER, THIRTY MINUTES POST-PROCEDURE, THE PATIENT BLOOD PRESSURE COLLAPSED, ACCOMPANIED BY TACHYCARDIA AND PERICARDIAL EFFUSION. THIS WAS ATTRIBUTED TO AN ANNULAR RUPTURE, LEADING TO CARDIAC TAMPONADE, AND THE PATIENT WENT FOR SURGERY. UNFORTUNATELY, THE PATIENT EXPIRED TWO DAYS LATER. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS THAT IT WAS A (B)(6)-YEARS OLD FRAGILE PATIENT WITH SEVERE ANNULAR CALCIFICATION AND BULKY CALCIFICATION ON LEFT CORONARY CUSP. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS INDICATE AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS THAT IT WAS A (B)(6)-YEARS OLD FRAGILE PATIENT WITH SEVERE ANNULAR CALCIFICATION AND BULKY CALCIFICATION ON LEFT CORONARY CUSP. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929514&pc=DSQ | 2916596-2025-02873 | 2025-02-26 05:00 | Death | THORATEC CORPORATION | 2025-04-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Renal Failure; Tachycardia; Cardiogenic Shock; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2025 WITH REFRACTORY RIGHT VENTRICULAR FAILURE AND CARDIOGENIC SHOCK. THE PATIENT HAD PREVIOUSLY HAD BEEN ON HOME MILRINONE IN 2022 AND WAS DYING WITH RECURRENT VENTRICULAR TACHYCARDIA, REFRACTORY RIGHT VENTRICLE FAILURE, RENAL FAILURE, AND UPPER AND LOWER EXTREMITIES WITH DRY GANGRENE. PULMONARY ARTERY CATHETER WAS REMOVED. ON THE MORNING OF (B)(6) 2025, THE PATIENT EXPERIENCED CARDIAC ARREST AND VENTRICULAR TACHYCARDIA THAT REQUIRED SHOCK ALONG WITH LOW FLOW ALARMS. PRESSOR REQUIREMENTS INCREASED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: ADDITIONAL ANNEX E CODES: GENERALIZED DISORDERS E23: NECROSIS (E2327). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: B3 - DATE OF EVENT - CORRECTION. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED LOW FLOW HAZARD ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE SUBMITTED FOR REVIEW; HOWEVER, THE ACCOUNT REPORTED THAT THE ALARMS OCCURRED WHILE THE PATIENT WAS UNDER CARDIAC ARREST WHICH COULD HAVE CONTRIBUTED TO THE LOWER FLOWS. IT WAS REPORTED THAT THE PATIENT¿S OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND AN AUTOPSY WAS NOT PERFORMED. THE PUMP WILL NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING HEART FAILURE, RENAL DYSFUNCTION, ARRHYTHMIA, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, UNDER "RIGHT HEART FAILURE", DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. SECTION 6 ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929591&pc=DSQ | 3007042319-2025-00850 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO HEART DISEASE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO HEART DISEASE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929601&pc=DSQ | 3007042319-2025-00855 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929612&pc=DSQ | 3007042319-2025-00870 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS AN UNSPECIFICED NEUROLOGICAL DYSFUNCTION ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929615&pc=DSQ | 3007042319-2025-00881 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929617&pc=DSQ | 3007042319-2025-00869 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED CAUSE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929622&pc=DSQ | 3007042319-2025-00887 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED CIRCULATORY ISSUE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS AN UNSPECIFICED CIRCULATORY ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929643&pc=DSQ | 3007042319-2025-00896 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929650&pc=DSQ | 3007042319-2025-00910 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929655&pc=DSQ | 3007042319-2025-00909 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO A SUDDEN UNEXPLAINED UNSPECIFIED CIRCULATORY ISSUE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS AN UNSPECIFICED CIRCULATORY ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929665&pc=DSQ | 3007042319-2025-00921 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Unspecified Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A SUPERFICIAL BACTERIAL PERCUTANEOUS DRIVELINE INFECTION REQUIRING INTRAVENOUS (IV) MEDICATION AND A BLOOD STREAM INFECTION REQUIRING INTRAVENOUS (IV) MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE, INCLUDING CARE OF THE DRIVELINE EXIT SITE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929671&pc=DSQ | 3007042319-2025-00925 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - OUTFLOW GRAFT | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A VENTRICULAR ASSIST DEVICE (VAD) OUTFLOW GRAFT OBSTRUCTION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VAD AND OUTFLOW GRAFT REMAIN IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP OF UNKNOWN SERIAL NUMBER AND ONE (1) OUTFLOW GRAFT WITH UNKNOWN LOT NUMBER WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBERS WERE UNKNOWN. LOG FILE ANALYSIS COULD NOT BE PERFORMED SINCE LOG FILES COVERING THE REPORTED EVENT DATE WERE NOT AVAILABLE FOR ANALYSIS. THE REPORTED OUTFLOW GRAFT OBSTRUCTION EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING AN OUTFLOW GRAFT OBSTRUCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929702&pc=DSQ | 3007042319-2025-00961 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Pneumonia; Renal Failure; Sepsis; Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR PNEUMONIA, ACUTE STAGE 1 RENAL DYSFUNCTION, HEPATIC DYSFUNCTION, AND SEPSIS REQUIRING INTRAVENOUS (IV) MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING PNEUMONIA, ACUTE STAGE 1 RENAL DYSFUNCTION, HEPATIC DYSFUNCTION AND SEPSIS. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929703&pc=DSQ | 3007042319-2025-00960 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A SUSPECTED VENTRICULAR ASSIST DEVICE (VAD) THROMBUS AND ABNORMAL PUMP PARAMETERS. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO RIGHT HEART FAILURE. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. AS A RESULT, THE REPORTED ABNORMAL PUMP PARAMETERS COULD NOT BE CONFIRMED. THE REPORTED THROMBUS EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE ABNORMAL PUMP PARAMETERS EVENT INCLUDING BUT NOT LIMITED TO THROMBUS FORMATION/INGESTION, THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929712&pc=DSQ | 3007042319-2025-00963 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Respiratory Failure; Gastrointestinal Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR LOWER GASTROINTESTINAL BLEEDING REQUIRING AN UNSPECIFIED NON-SURGICAL INTERVEN TION AND RESPIRATORY FAILURE REQUIRING INTUBATION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING LOWER GASTROINTESTINAL BLEEDING. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929715&pc=DSQ | 3007042319-2025-00971 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Multiple Organ Failure; Unspecified Nervous System Problem; Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SUSTAINED VENTRICULAR ARRHYTHMIA REQUIRING AN UNSPECIFIED INTERVENTION, DELIRIUM WITHOUT CENTRAL NERVOUS SYSTEM (CNS) INJURY DIAGNOSED CLINICALLY, AND HEPATIC DYSFUNCTION. THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929725&pc=DSQ | 3007042319-2025-00956 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SUSTAINED VENTRICULAR ARRHYTHMIA REQUIRING AN UNSPECIFIED INTERVENTION. THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF AND SUPPORT WAS WITHDRAWN. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO CARDIAC ARRHYTHMIA. THE VAD WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE C ONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929726&pc=DSQ | 3007042319-2025-00958 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A BLOOD STREAM INFECTION REQUIRING INTRAVENOUS (IV) MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED AFTER WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929730&pc=DSQ | 3007042319-2025-00984 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO A SUDDEN UNEXPLAINED UNSPECIFIED CIRCULATORY ISSUE. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS AN UNSPECIFICED CIRCULATORY ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929732&pc=DSQ | 3007042319-2025-00964 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE CARDIOMYOPATHY. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO END STAGE CARDIOMYOPATHY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929734&pc=DSQ | 3007042319-2025-00981 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE DIAGNOSED VIA COMPUTERIZED TOMOGRAPHY (CT). THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING A CEREBRAL HEMORRHAGE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929738&pc=DSQ | 3007042319-2025-00979 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Stopped | Cardiac Arrest; Fall; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOLLOWING A FALL AT HOME. WHILE ADMITTED FOR MONITORING, IT WAS REPORTED THAT THE NURSING STAFF HAD ATTEMPTED TO CHANGE OUT POWER SOURCES FROM BATTERY TO CONTROLLER AC ADAPTER BUT HAD DIFFICULTY WITH DISCONNECTIONS. THE VENTRICULAR ASSIST DEVICE (VAD) BEGAN TO EXHIBIT VAD STOP ALARMS DESPITE HAVING BATTERIES STILL CONNECTED. THE PATIENT WAS STARTED ON DOBUTAMINE AND TRANSFERRED TO THE CRITICAL CARE UNIT. THE PATIENT BECAME UNRESPONSIVE. CARDIOPULMONARY RESUSCITATION WAS INITIATED BUT LATER STOPPED. THE PATIENT SUBSEQUENTLY EXPIRED. IT WAS NOTED THAT THE PATIENT HAD A KNOWN THROMBUS AND THAT THE VAD WAS INCLUDED IN AN UNSPECIFIED SUBGROUP POPULATION FOR DELAY OR FAILURE TO RESTART. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. THE REPORTED VAD STOP AND THROMBUS EVENTS COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING THROMBUS AND VAD STOPPED ALARMS; HOWEVER, THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A POSSIBLE ROOT CAUSE OF THE REPORTED VAD STOP EVENT CAN BE ATTRIBUTED BUT NOT LIMITED TO A DISCONNECTION OF BOTH POWER SOURCES FROM THE CONTROLLER, AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES, A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER, A CONTROLLER FAILURE, OR A PUMP FAILURE. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929740&pc=DSQ | 3007042319-2025-00990 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Renal Failure; Respiratory Failure; Heart Failure/Congestive Heart Failure; Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED DUE TO HEPATIC DYSFUNCTION, ACUTE STAGE 3 RENAL DYSFUNCTION, RESPIRATORY FAILURE, AND LATE-STAGE RIGHT HEART FAILURE REQUIRING INTRAVENOUS (IV) MEDICATION OR INOTROPIC SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF AND SUPPORT WAS WITHDRAWN. THE PATIENT SUBSEQUENTLY EXPIRED. THE VAD WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER THE VAD WAS TURNED OFF AND SUPPORT WAS WITHDRAWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929764&pc=DSQ | 3007042319-2025-00986 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Pneumonia; Urinary Tract Infection; Respiratory Failure; Unspecified Tissue Injury | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED TRAUMA/ACCIDENT, PNEUMONIA, A URINARY TRACT INFECTION REQUIRING INTRAVENOUS (IV) MEDICATION, AND RESPIRATORY FAILURE. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929835&pc=DSQ | 3007042319-2025-00983 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Convulsion/Seizure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A SEIZURE DIAGNOSED VIA COMPUTERIZED TOMOGRAPHY (CT). THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING A SEIZURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929844&pc=DSQ | 3007042319-2025-01010 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED CAUSE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929901&pc=DSQ | 3007042319-2025-01018 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929918&pc=DSQ | 3007042319-2025-01041 |  | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929933&pc=DSQ | 3007042319-2025-01083 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Stroke/CVA; Respiratory Failure; Gastrointestinal Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UPPER GASTROINTESTINAL BLEED, RESPIRATORY FAILURE REQUIRING REINTUBATION, A STROKE DIAGNOSED BY COMPUTERIZED TOMOGRAPHY (CT), AND AN UNSPECIFIED MAJOR BACTERIAL INFECTION REQUIRING AN UNSPECIFIED INTERVENTION. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO THE INFECTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: A VENTRICULAR ASSIST DEVICE (VAD) WITH AN UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS HOSPITALIZED FOR AN UPPER GASTROINTESTINAL BLEED, RESPIRATORY FAILURE REQUIRING REINTUBATION, A STROKE DIAGNOSED BY COMPUTERIZED TOMOGRAPHY (CT), AND AN UNSPECIFIED MAJOR BACTERIAL INFECTION REQUIRING AN UNSPECIFIED INTERVENTION AND SUBSEQUENTLY EXPIRED DUE TO THE INFECTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, BLEEDING, RESPIRATORY DYSFUNCTION, STROKE, INFECTION, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929947&pc=QAU | 1037905-2025-00349 | 2025-03-27 04:00 | Death | WILSON-COOK MEDICAL INC | 2025-04-29 04:00 | QAU | HEMOSPRAY ENDOSCOPIC HEMOSTAT | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Fever; Perforation | K200972 |  | 1 | Event Description: COOK BECAME AWARE OF A CLINICAL LITERATURE ARTICLE CASE STUDY INVOLVING COOK HEMOSPRAY ENDOSCOPIC HEMOSTAT. "INTRODUCTION: MALIGNANT TUMORS OF THE UPPER GASTROINTESTINAL (GI) TRACT ARE A FREQUENT CAUSE OF GI BLEEDING [1]. TREATMENT OF THESE PATIENTS IS CHALLENGING [2]. COMMON HEMOSTATIC METHODS ARE OFTEN INADEQUATE. APPLICATION OF LOCAL HEMOSTATIC AGENTS HAS BEEN REPORTED IN ACHIEVING HEMOSTASIS IN SELECT CASES (LIKE DIFFUSE HEMORRHAGE), MAINLY AS BRIDGING THERAPY BEFORE RADIOTHERAPY OR SURGICAL TREATMENT [3]. VISCERAL PERFORATION HAS BEEN REPORTED AS A VERY RARE COMPLICATION OF TC-325 AGENT USE [4]. CASE PRESENTATION: WE REPORT THE CASE OF AN 87-YEAR-OLD MALE WITH METASTATIC DIFFUSE STOMACH ADENOCARCINOMA UNDER CHEMOTHERAPY. THE PATIENT HAD BEEN PREVIOUSLY HOSPITALIZED DUE TO RECURRENT UPPER GI BLEEDING AND/OR ANEMIA ATTRIBUTED TO THE TUMOR THAT WAS NOT AMENABLE TO TRADITIONAL ENDOSCOPIC TREATMENT. ENDOSCOPIC HEMOSTASIS WAS PREVIOUSLY ACHIEVED WITH LOCAL APPLICATION OF TC-325 POWDER. AT PRESENT, THE PATIENT PRESENTED TO OUR HOSPITAL EMERGENCY DEPARTMENT WITH HEMATEMESIS, HEMATOCHEZIA AND HEMODYNAMIC INSTABILITY. THE PATIENT UNDERWENT UPPER GI ENDOSCOPY, AND WAS DIAGNOSED WITH ACTIVE GI BLEEDING FROM THE STOMACH MASS. ENDOSCOPIC HEMOSTASIS WAS ACHIEVED WITH THE USE OF TC ¿ 325 POWDER. FOLLOWING ENDOSCOPY, PATIENT COMPLAINED OF SEVERE DIFFUSE ABDOMINAL PAIN, WITH FEVER AND LEUKOCYTOSIS. ABDOMINAL COMPUTED TOMOGRAPHY (CT) WAS PERFORMED, WHICH SHOWED PRESENCE OF FREE AIR IN THE PERITONEAL CAVITY AND ABDOMINAL FLUID, INDICATING VISCERAL PERFORATION, WITHOUT ACTIVE EXTRAVASATION OF INTRAVENOUS CONTRAST INTO THE BOWEL LUMEN. DESPITE ACHIEVING HEMOSTASIS, THE PATIENT DECEASED DURING HOSPITALIZATION. CONCLUSIONS: USE OF GAS ¿ PROPELLED HEMOSTATIC POWDERS CAN CAUSE VISCERAL PERFORATION IN PATIENTS WITH ADVANCED MALIGNANT TUMORS OF THE UPPER GI TRACT. CAUTION NEEDS TO BE EXERCISED IN CASES OF DIFFUSE/LONG STANDING DISEASE." A CASE OF PERFORATION FOLLOWING THE APPLICATION OF HEMOSTATIC POWDER TC-325 FOR THE TREATMENT OF STOMACH CANCER BLEEDING. K PAPANTONIOU , T KAFENTZI , C SOTIROPOULOS , K THOMOPOULOS , C KONSTANTAKIS ENDOSCOPY 2025; 57(S 02): S413. DOI: 10.1055/S-0045-1806054.Event Description: COOK ENDOSCOPY WAS NOTIFIED OF THIS EVENT VIA A TWO SEPARATE CLINICAL LITERATURE ARTICLES. THE ARTICLES WERE PUBLISHED IN 2025. A CASE OF PERFORATION FOLLOWING THE APPLICATION OF HEMOSTATIC POWDER TC-325 FOR THE TREATMENT OF STOMACH CANCER BLEEDING ¿INTRODUCTION MALIGNANT TUMORS OF THE UPPER GASTROINTESTINAL (GI) TRACT ARE A FREQUENT CAUSE OF GI BLEEDING [1]. TREATMENT OF THESE PATIENTS IS CHALLENGING [2]. COMMON HEMOSTATIC METHODS ARE OFTEN INADEQUATE. APPLICATION OF LOCAL HEMOSTATIC AGENTS HAS BEEN REPORTED IN ACHIEVING HEMOSTASIS IN SELECT CASES (LIKE DIFFUSE HEMORRHAGE), MAINLY AS BRIDGING THERAPY BEFORE RADIOTHERAPY OR SURGICAL TREATMENT [3]. VISCERAL PERFORATION HAS BEEN REPORTED AS A VERY RARE COMPLICATION OF TC-325 AGENT USE [4]. CASE PRESENTATION WE REPORT THE CASE OF AN 87-YEAR-OLD MALE WITH METASTATIC DIFFUSE STOMACH ADENOCARCINOMA UNDER CHEMOTHERAPY. THE PATIENT HAD BEEN PREVIOUSLY HOSPITALIZED DUE TO RECURRENT UPPER GI BLEEDING AND/OR ANEMIA ATTRIBUTED TO THE TUMOR THAT WAS NOT AMENABLE TO TRADITIONAL ENDOSCOPIC TREATMENT. ENDOSCOPIC HEMOSTASIS WAS PREVIOUSLY ACHIEVED WITH LOCAL APPLICATION OF TC ¿ 325 POWDER. AT PRESENT, THE PATIENT PRESENTED TO OUR HOSPITAL EMERGENCY DEPARTMENT WITH HEMATEMESIS, HEMATOCHEZIA AND HEMODYNAMIC INSTABILITY. THE PATIENT UNDERWENT UPPER GI ENDOSCOPY, AND WAS DIAGNOSED WITH ACTIVE GI BLEEDING FROM THE STOMACH MASS. ENDOSCOPIC HEMOSTASIS WAS ACHIEVED WITH THE USE OF TC ¿ 325 POWDER. FOLLOWING ENDOSCOPY, PATIENT COMPLAINED OF SEVERE DIFFUSE ABDOMINAL PAIN, WITH FEVER AND LEUKOCYTOSIS. ABDOMINAL COMPUTED TOMOGRAPHY (CT) WAS PERFORMED, WHICH SHOWED PRESENCE OF FREE AIR IN THE PERITONEAL CAVITY AND ABDOMINAL FLUID, INDICATING VISCERAL PERFORATION, WITHOUT ACTIVE EXTRAVASATION OF INTRAVENOUS CONTRAST INTO THE BOWEL LUMEN. DESPITE ACHIEVING HEMOSTASIS, THE PATIENT DECEASED DURING HOSPITALIZATION. CONCLUSIONS USE OF GAS ¿ PROPELLED HEMOSTATIC POWDERS CAN CAUSE VISCERAL PERFORATION IN PATIENTS WITH ADVANCED MALIGNANT TUMORS OF THE UPPER GI TRACT. CAUTION NEEDS TO BE EXERCISED IN CASES OF DIFFUSE/LONG STANDING DISEASE.¿ AN ADDITIONAL ARTICLE WAS DISCOVERED REGARDING THIS SAME CASE STUDY ON 09 MAY 2025. PAPANTONIOU K, KAFENTZI T, SOTIROPOULOS C, ET AL. (MAY 06, 2025) VISCERAL PERFORATION FOLLOWING HEMOSTATIC POWDER TC-325 APPLICATION IN A PATIENT WITH GASTRIC CANCER: A CASE REPORT. CUREUS 17(5): E83611. DOI 10.7759/CUREUS.83611. ¿ABSTRACT: MALIGNANT TUMORS OF THE UPPER GASTROINTESTINAL (GI) TRACT ARE A COMMON CAUSE OF GI BLEEDING, OFTEN PRESENTING THERAPEUTIC CHALLENGES DUE TO DIFFUSE HEMORRHAGE AND POOR VISUALIZATION. LOCAL HEMOSTATIC AGENTS, SUCH AS TC-325 (HEMOSPRAY; WINSTON-SALEM, NC: COOK MEDICAL), HAVE EMERGED AS EFFECTIVE BRIDGING THERAPIES PRIOR TO DEFINITIVE TREATMENT. HOWEVER, VISCERAL PERFORATION, THOUGH RARE, IS A SERIOUS COMPLICATION ASSOCIATED WITH TC-325 USE. WE PRESENT A CASE OF AN 87-YEAR-OLD MALE WITH METASTATIC GASTRIC ADENOCARCINOMA WHO EXPERIENCED VISCERAL PERFORATION FOLLOWING TC-325 APPLICATION FOR RECURRENT UPPER GI BLEEDING. DESPITE ACHIEVING INITIAL HEMOSTASIS, THE PATIENT SUCCUMBED TO COMPLICATIONS OF PERFORATION. THIS STUDY HIGHLIGHTS THE NEED FOR CAUTION WHEN USING TC-325 IN PATIENTS WITH ADVANCED MALIGNANCIES AND FRIABLE TISSUE. INTRODUCTION NON-VARICEAL UPPER GASTROINTESTINAL BLEEDING (NVUGIB) IS A SIGNIFICANT CAUSE OF MORBIDITY AND MORTALITY WORLDWIDE. WHILE MALIGNANCIES OF THE UPPER GI TRACT ARE A LESS COMMON ETIOLOGY, THEY POSE UNIQUE CHALLENGES DUE TO DIFFUSE BLEEDING SURFACES, FRIABLE TISSUE, AND POOR ENDOSCOPIC VISUALIZATION [1]. TOPICAL HEMOSTATIC AGENTS, SUCH AS TC-325 (HEMOSPRAY; WINSTON-SALEM, NC: COOK MEDICAL), HAVE EMERGED AS EFFECTIVE ALTERNATIVES TO STANDARD ENDOSCOPIC THERAPIES, PARTICULARLY IN CASES OF MALIGNANT NVUGIB [2]. ALTHOUGH TC 325 IS GENERALLY CONSIDERED SAFE, RARE BUT SERIOUS COMPLICATIONS, INCLUDING VISCERAL PERFORATION, HAVE BEEN REPORTED [3,4]. HERE, WE PRESENT A CASE OF VISCERAL PERFORATION FOLLOWING TC-325 APPLICATION IN AN 87-YEAR-OLD MALE WITH METASTATIC GASTRIC ADENOCARCINOMA, UNDERSCORING THE NEED FOR CAUTION IN PATIENTS WITH ADVANCED MALIGNANCIES. THIS ARTICLE WAS PREVIOUSLY PRESENTED AS AN E-POSTER AT THE 44TH PANHELLENIC CONGRESS OF GASTROENTEROLOGY ON NOVEMBER 28, 2024, AND AT THE ANNUAL EUROPEAN SOCIETY OF GASTROINTESTINAL ENDOSCOPY (ESGE) DAYS CONGRESS ON APRIL 3, 2025. CASE PRESENTATION: AN 87-YEAR-OLD MALE WITH A HISTORY OF ARTERIAL HYPERTENSION, DYSLIPIDEMIA, AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE PRESENTED TO OUR EMERGENCY DEPARTMENT WITH HEMATEMESIS, HEMATOCHEZIA, AND HEMODYNAMIC INSTABILITY. LABORATORY TESTS REVEALED SEVERE ANEMIA WITH HEMOGLOBIN 4.9 G/DL (REFERENCE RANGE: 11.8-17 G/DL). THE PATIENT HAD BEEN PREVIOUSLY DIAGNOSED WITH METASTATIC GASTRIC ADENOCARCINOMA (LINITIS PLASTICA), WITH SECONDARY LESIONS IN THE LIVER AND LUNGS. HE HAD PREVIOUSLY UNDERGONE MULTIPLE HOSPITALIZATIONS, BOTH FOR RECURRENT UPPER GASTROINTESTINAL (GI) BLEEDING AND TRANSFUSION-DEPENDENT ANEMIA ATTRIBUTED TO THE TUMOR AND NOT AMENABLE TO TRADITIONAL ENDOSCOPIC TREATMENT. DURING A PRIOR HOSPITALIZATION, ACTIVE BLEEDING FROM THE GASTRIC MASS WAS SUCCESSFULLY CONTROLLED WITH THE APPLICATION OF TC-325 HEMOSTATIC POWDER (HEMOSPRAY) WITHOUT COMPLICATIONS. AT THE CURRENT PRESENTATION, COMPUTED TOMOGRAPHY (CT) ANGIOGRAPHY CONFIRMED THE GASTRIC TUMOR AS THE SOURCE OF BLEEDING, WITHOUT EVIDENCE OF ACTIVE CONTRAST EXTRAVASATION (FIGURE 1). GIVEN THE PATIENT¿S HEMODYNAMIC STATUS AND THE NEED FOR URGENT INTERVENTION, ENDOSCOPIC HEMOSTASIS WITH TC-325 WAS CHOSEN DUE TO ITS PREVIOUS SUCCESS AND THE LIMITATIONS OF ALTERNATIVE THERAPIES. THE KNOWN MALIGNANCY, WHICH WAS PRIMARILY LOCATED IN THE GASTRIC BODY, WAS IDENTIFIED DURING ENDOSCOPY. ACTIVE BLEEDING WAS OBSERVED FROM THE TUMOR SURFACE (FIGURE 2). THE PROCEDURE WAS CONDUCTED WITH UTMOST CARE TO MINIMIZE TRAUMA TO THE FRIABLE TISSUE: CO2 INFLATION WAS USED TO REDUCE THE RISK OF PERFORATION, AND NO ATTEMPT WAS MADE TO PASS THE PYLORUS, AVOIDING ADDITIONAL STRAIN ON THE TUMOR. GIVEN THE DIFFUSE NATURE OF THE BLEEDING AND THE SUBSTANTIAL TUMOR AREA, A 10FR CATHETER WAS SELECTED TO DISPENSE THE TC-325 POWDER, ACHIEVING HEMOSTASIS. IN THE IMMEDIATE POST-OPERATIVE PERIOD, THE PATIENT COMPLAINED OF EPIGASTRIC PAIN, WHICH WAS INITIALLY ATTRIBUTED TO THE VISCERAL DISTENTION ASSOCIATED WITH THE GAS PROPELLANT OF THE HEMOSTATIC DEVICE. SUBSEQUENTLY, THE PATIENT DEVELOPED SEVERE DIFFUSE ABDOMINAL PAIN, FEVER, AND LEUKOCYTOSIS. AN URGENT ABDOMINAL CT REVEALED FREE AIR AND FLUID IN THE PERITONEAL CAVITY, INDICATING VISCERAL PERFORATION (FIGURE 3). NO EXTRAVASATION OF INTRAVENOUS CONTRAST INTO THE BOWEL LUMEN WAS FOUND. DUE TO THE ADVANCED STAGE OF THE DISEASE AND THE HIGH RISK OF SURGICAL COMPLICATIONS, THE PATIENT WAS NOT DEEMED A CANDIDATE FOR SURGICAL INTERVENTION. DESPITE ACHIEVING INITIAL HEMOSTASIS, THE PATIENT SUCCUMBED TO COMPLICATIONS TWO DAYS LATER. DISCUSSION: ENDOSCOPY REMAINS THE CORNERSTONE OF TREATMENT FOR PATIENTS PRESENTING WITH NVUGIB, WITH RESUSCITATION USING IV FLUIDS AND BLOOD TRANSFUSIONS BEING CRITICAL TO STABILIZE PATIENTS PRIOR TO THE PROCEDURE [1]. CURRENT ESGE GUIDELINES RECOMMEND ENDOSCOPY WITHIN 24 HOURS OF ADMISSION, AS URGENT ENDOSCOPY (<12 HOURS) HAS NOT DEMONSTRATED SUPERIOR OUTCOMES COMPARED TO DELAYED INTERVENTION [5]. THE USE OF HEMOSPRAY (TC-325) IN SIMILAR CASES HAS BEEN ASSOCIATED WITH SEVERAL BENEFICIAL EFFECTS, INCLUDING SIGNIFICANT REDUCTION IN TRANSFUSION REQUIREMENTS, IMPROVED QUALITY OF LIFE, AND LONG-TERM ONCOLOGICAL OUTCOMES AND BRIDGING TO MORE DEFINITIVE TREATMENTS, SUCH AS RADIOTHERAPY OR SURGERY [6]. RECENT STUDIES HAVE HIGHLIGHTED THE EFFICACY OF TC-325 IN MANAGING MALIGNANCY-RELATED GI BLEEDING. A SYSTEMATIC REVIEW AND META-ANALYSIS BY CHAHAL ET AL. DEMONSTRATED THAT TC-325 SIGNIFICANTLY IMPROVES HEMOSTASIS RATES AND REDUCES TRANSFUSION NEEDS IN PATIENTS WITH UPPER GI BLEEDING [7]. SIMILARLY, HUSSEIN ET AL. REPORTED FAVORABLE OUTCOMES IN AN INTERNATIONAL REGISTRY OF PATIENTS TREATED WITH TC-325 FOR MALIGNANCY RELATED BLEEDING, WITH A HIGH RATE OF INITIAL HEMOSTASIS AND LOW RE-BLEEDING RATES [6]. KARNA ET AL. AND SAEED ET AL. FURTHER SUPPORTED THESE FINDINGS, SHOWING THAT TC-325 IS SUPERIOR TO STANDARD ENDOSCOPIC THERAPIES IN ACHIEVING HEMOSTASIS IN MALIGNANCY-RELATED BLEEDING [8,9]. HOWEVER, AS HIGHLIGHTED BY PITTAYANON ET AL., ENDOSCOPIC TREATMENT FOR PATIENTS WITH GI BLEEDING DUE TO MALIGNANT TUMORS MAY SOMETIMES BE FUTILE, PARTICULARLY IN THE SHORT-TO-MEDIUM TERM [10]. WANG ET AL., IN AN EDITORIAL ACCOMPANYING THE PITTAYANON STUDY, NOTED THAT WHILE SOME PATIENTS WITH TUMOR-RELATED BLEEDING MAY BE CANDIDATES FOR THERAPEUTIC ENDOSCOPIC OR SURGICAL RESECTION, MANY HAVE ADVANCED OR INCURABLE CANCERS. IN THESE PATIENTS, TC-325 APPEARS TO PROVIDE IMPROVED SHORT- TO MEDIUM-TERM HEMOSTASIS COMPARED TO STANDARD ENDOSCOPIC THERAPY. FOR PATIENTS WITH END-STAGE MALIGNANCIES, EVEN A FEW WEEKS OF PALLIATIVE HEMOSTASIS CAN PROVIDE VALUABLE TIME TO CONSIDER PALLIATIVE CARE OPTIONS AND MANAGE THEIR AFFAIRS BEFORE END-OF-LIFE [11]. DESPITE ITS BENEFITS, THE USE OF TC-325 IS NOT WITHOUT RISKS. IN THE PRESENT CASE, THE PATIENT EXPERIENCED VISCERAL PERFORATION, A RARE BUT SERIOUS COMPLICATION. ADVANCED CANCEROUS INVASION OF THE GASTRIC WALL LIKELY CONTRIBUTED TO THE FRIABILITY OF THE TISSUE, INCREASING THE RISK OF PERFORATION. CLINICIANS MUST EXERCISE CAUTION WHEN APPLYING TC-325 IN PATIENTS WITH ADVANCED MALIGNANCIES, PARTICULARLY THOSE WITH DIFFUSE OR LONG-STANDING DISEASE. THE FORCEFUL RELEASE OF THE HEMOSTATIC POWDER THROUGH THE ENDOSCOPE CAN CAUSE TISSUE RUPTURE AND SUBSEQUENT PNEUMOPERITONEUM [12]. THE USE OF A 10FR CATHETER FOR TC-325 APPLICATION IN THIS CASE MAY HAVE CONTRIBUTED TO THE RISK OF PERFORATION, GIVEN THE FRIABLE NATURE OF THE TUMOR AND THE ADVANCED STAGE OF THE DISEASE. WHILE THERE IS NO DIRECT EVIDENCE COMPARING THE PERFORATION RISK OF 10FR VS. 7FR CATHETERS IN MALIGNANCY-RELATED BLEEDING, THE HIGHER FORCE OF POWDER DELIVERY WITH A 10FR CATHETER COULD THEORETICALLY INCREASE MECHANICAL STRESS ON WEAKENED TISSUE. IN HIGH-RISK CASES, SUCH AS ADVANCED MALIGNANCIES, THE USE OF A 7FR CATHETER MAY BE PREFERABLE TO MINIMIZE TRAUMA, DESPITE POTENTIALLY REQUIRING MORE APPLICATIONS TO ACHIEVE HEMOSTASIS. FUTURE STUDIES COMPARING CATHETER SIZES IN THIS CONTEXT ARE NEEDED TO PROVIDE CLEARER GUIDANCE ON OPTIMAL TECHNIQUE. CONCLUSIONS: TC-325 HEMOSTATIC POWDER IS A VALUABLE TOOL FOR MANAGING NVUGIB IN PATIENTS WITH MALIGNANT UPPER GI TUMORS, OFFERING EFFECTIVE HEMOSTASIS IN CHALLENGING CASES. IT SIGNIFICANTLY REDUCES TRANSFUSION REQUIREMENTS, IMPROVES QUALITY OF LIFE, AND SERVES AS A BRIDGE TO MORE DEFINITIVE TREATMENTS SUCH AS RADIOTHERAPY OR SURGERY. HOWEVER, VISCERAL PERFORATION, THOUGH RARE, REMAINS A SERIOUS COMPLICATION, PARTICULARLY IN PATIENTS WITH ADVANCED DISEASE AND FRIABLE TISSUE. CLINICIANS MUST REMAIN VIGILANT FOR SIGNS OF PERFORATION, AS EARLY RECOGNITION AND INTERVENTION ARE CRITICAL TO OPTIMIZING OUTCOMES. THIS STUDY UNDERSCORES THE NEED FOR CAUTION WHEN USING TC-325 IN HIGH-RISK POPULATIONS AND HIGHLIGHTS THE IMPORTANCE OF BALANCING THE BENEFITS OF HEMOSTASIS WITH THE RISKS OF COMPLICATIONS IN PATIENTS WITH ADVANCED MALIGNANCIES.¿ Manufacturer Narrative: ADDITIONAL ARTICLE WAS PUBLISHED BY THE SAME AUTHOR ABOUT THE SAME PATIENT AND EVENT. SEE ATTACHED AND B5. VISCERAL PERFORATION FOLLOWING HEMOSTATIC POWDER TC-325 APPLICATION IN A PATIENT WITH GASTRIC CANCER: A CASE REPORT. K PAPANTONIOU, T KAFENTZI, C SOTIROPOULOS, K THOMOPOULOS, C KONSTANTAKIS. DOI: 10.7759/CUREUS.83611. 10 FR CATHETER WAS USED IN THE PROCEDURE. Manufacturer Narrative: INVESTIGATION EVALUATION: A PRODUCT EVALUATION WAS NOT PERFORMED IN RESPONSE TO THIS REPORT BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT PROVIDED TO COOK FOR EVALUATION. THE REPORT COULD NOT BE CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE CONDUCTED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. INVESTIGATION CONCLUSION: WE COULD NOT CONDUCT A COMPLETE INVESTIGATION BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT RETURNED FOR EVALUATION. THERE IS NO EVIDENCE TO SUGGEST THE REPORT IS DUE TO A DEVICE FAILURE. THE IFU PRECAUTIONS "ENSURE GASTROINTESTINAL LUMEN IS NOT DISTENDED BECAUSE HEMOSPRAY ADDS VOLUME IN EXCESS OF INSUFFLATION VOLUMES DURING PROCEDURE. CLOSELY MONITOR BOWEL DISTENSION AND BALANCE INSUFFLATION AND HEMOSPRAY VOLUMES AS NECESSARY." ADDITIONALLY, THE IFU PRECAUTIONS "LIKE OTHER MODALITIES, HEMOSPRAY MAY NOT BE EFFECTIVE FOR ALL TYPES OF BLEEDS. GASTROINTESTINAL BLEEDING MAY EXACERBATE EXISTING COMORBIDITIES, INCREASING THE POTENTIAL FOR ADVERSE EVENTS INCLUDING PATIENT MORTALITY." PRIOR TO DISTRIBUTION, ALL HEMOSPRAY ENDOSCOPIC HEMOSTAT ARE SUBJECTED TO A VISUAL INSPECTION TO ENSURE DEVICE INTEGRITY. CORRECTIVE ACTION: BASED ON THE QUALITY ENGINEERING RISK ASSESSMENT NO CORRECTIVE ACTION IS WARRANTED AT THIS TIME. A REVIEW OF THE COMPLAINT HISTORY WAS CONDUCTED. BASED ON THIS REVIEW, THE LIKELIHOOD OF THIS TYPE OF REPORT IS RARE. QUALITY ASSURANCE WILL CONTINUE TO MONITOR FOR COMPLAINT TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929948&pc=DSQ | 3007042319-2025-01072 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CARDIAC ARRHYTHMIA. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE C ONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929950&pc=DSQ | 2916596-2025-02730 | 2023-02-03 05:00 | Death | THORATEC CORPORATION | 2025-04-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿OUTCOMES FOR CHILDREN WITH CONGENITAL HEART DISEASE UNDERGOING VENTRICULAR ASSIST DEVICE IMPLANTATION¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH STROKE AND OTHER NEUROLOGICAL DYSFUNCTION, MAJOR BLEEDING, MAJOR INFECTION, RENAL AND HEPATIC DYSFUNCTION, ARRHYTHMIA, HEMOLYSIS, RESPIRATORY FAILURE, DEVICE MALFUNCTION, AND DEATH. THIS RETROSPECTIVE STUDY AIMED TO EVALUATE DIFFERENCES IN POST-VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION MORBIDITY AND MORTALITY FOR CHILDREN WITH CONGENITAL HEART DISEASE (CHD) COMPARED TO THOSE WITH NON-CHD, AS WELL AS DIFFERENCE IN MORBIDITY AND MORTALITY AMONG CHILDREN WITH BIVENTRICULAR CHD (BI-CHD) VS. UNIVENTRICULAR CHD (UNI-CHD). PEDIATRIC PATIENTS YOUNGER THAN 19 YEARS OLD FROM MULTIPLE CENTERS ENROLLED IN THE ADVANCED CARDIAC THERAPIES IMPROVING OUTCOMES NETWORK (ACTION) AND IMPLANTED WITH A VAD FROM 03APR2018 AND 02FEB2023 WERE INCLUDED WITH A MEDIAN FOLLOW-UP DURATION OF 2.6 MONTHS. THE MEDIAN AGE WAS 4.4 YEARS, THE MAJORITY (55.7%) WERE MALE, AND INTERMACS 1 (31.4%) AND 2 (53%). OUT OF 966 PATIENTS, 352 HAD CHD AND 614 HAD NON-CHD. OF THE 352 CHD PATIENTS, 230 HAD UNI-CHD AND 122 HAD BI-CHD, AND A TOTAL OF 33 PATIENTS WERE IMPLANTED WITH HM3. OF THE 614 NON-CHD PATIENTS, THE MAJORITY HAD DILATED CARDIOMYOPATHY OR MYOCARDITIS (499; 81.3%), AND A TOTAL OF 162 PATIENTS WERE IMPLANTED WITH HM3. OUTCOMES WERE NOT BROKEN DOWN BY VAD TYPE. RESULTS FOUND THAT CHD PATIENTS WERE MORE LIKELY TO EXPERIENCE STROKE (0.62 VS. 0.37 EVENTS PER PATIENT YEAR (EPPY)), OTHER NEUROLOGICAL DYSFUNCTION (0.27 VS. 0.12 EPPY), MAJOR BLEEDING (1.4 VS. 0.64 EPPY), MAJOR INFECTION (2.02 VS. 1.15 EPPY), RENAL DYSFUNCTION (0.36 VS 0.15 EPPY), HEPATIC DYSFUNCTION (0.23 VS 0.04 EPPY), RESPIRATORY FAILURE (0.71 VS. 0.24 EPPY), AND DEVICE MALFUNCTION (0.46 VS. 0.24 EPPY) (P < 0.001 FOR ALL). RESULTS FOR STROKE BETWEEN CHD AND NON-CHD PATIENTS NOTED HEMORRHAGIC (0.11 VS. 0.04 EPPY), ISCHEMIC (0.44 VS. 0.27 EPPY), AND ISCHEMIC WITH HEMORRHAGIC CONVERSION (0.04 VS. 0.02 EPPY). OTHER NEUROLOGICAL DYSFUNCTION EVENTS INCLUDED EXTRA-AXIAL HEMORRHAGIC DYSFUNCTION (0.13 VS. 0.06 EPPY), ACUTE ENCEPHALOPATHY (0.05 VS. 0.01 EPPY), HYPOXIC-ISCHEMIC ENCEPHALOPATHY (0.05 VS. 0.02 EPPY), AND TRANSIENT ISCHEMIC ATTACK (0.04 VS. 0.02 EPPY). RESULTS FOR INFECTION NOTED SEPSIS (0.83 VS. 0.38 EPPY), MEDIASTINITIS (0.16 VS. 0.03 EPPY), AND LOCALIZED NON-DEVICE RELATED INFECTION (0.84 VS. 0.43 EPPY). SUSTAINED SUPRAVENTRICULAR TACHYCARDIA (SVT) AND SUSTAINED VT REQUIRING DRUG TREATMENT OR CARDIOVERSION DID NOT DIFFER SIGNIFICANTLY BETWEEN CHD AND NON-CHD GROUPS. HEMOLYSIS OCCURRED AT A RATE OF 0.50 VS. 0.31 EPPY (P = 0.006). CHD PATIENTS ALSO HAD WORSE ADJUSTED SURVIVAL (91 VS. 51, P < 0.001). RESULTS ALSO FOUND THAT BI-CHD PATIENTS WERE MORE LIKELY TO EXPERIENCE MAJOR BLEEDING (1.95 VS. 1.19 EPPY) (P = 0.002); HOWEVER, RATES OF STROKE, MAJOR INFECTION, RENAL AND HEPATIC DYSFUNCTION, DEVICE MALFUNCTION, AND RESPIRATORY FAILURE WERE NOT SIGNIFICANTLY DIFFERENT BETWEEN BOTH GROUPS (P > 0.05 FOR ALL). THERE WERE NO DIFFERENCES IN ADJUSTED SURVIVAL IN UNI-CHD VS. BI-CHD PATIENTS (P > 0.05). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT THE TIME OF THE EVENT. SECTION B3: DATE OF EVENT HAS BEEN LISTED AS 02FEB2023 AS PATIENTS WERE IMPLANTED BETWEEN 03APR2018 AND 02FEB2023. DATE OF DEATH WAS NOT PROVIDED. AUTHOR INFORMATION: AMDANI, S. ET AL. (2025). OUTCOMES FOR CHILDREN WITH CONGENITAL HEART DISEASE UNDERGOING VENTRICULAR ASSIST DEVICE IMPLANTATION. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY, 85(8), 804¿814. HTTPS://DOI.ORG/10.1016/J.JACC.2024.10.083. CLEVELAND CLINIC CHILDREN¿S, CLEVELAND, OHIO, USA; CINCINNATI CHILDREN¿S HOSPITAL, CINCINNATI, OHIO, USA; ANN & ROBERT H. LURIE CHILDREN¿S HOSPITAL OF CHICAGO, NORTHWESTERN UNIVERSITY, CHICAGO, ILLINOIS, USA; LUCILE PACKARD CHILDREN¿S HOSPITAL, STANFORD UNIVERSITY, PALO ALTO, CALIFORNIA, USA; CINCINNATI CHILDREN¿S HOSPITAL MEDICAL CENTER, UNIVERSITY OF CINCINNATI COLLEGE OF MEDICINE, CINCINNATI, OHIO, USA; AND THE UNIVERSITY OF PENNSYLVANIA PERELMAN SCHOOL OF MEDICINE, CHILDREN¿S HOSPITAL OF PHILADELPHIA, PHILADELPHIA, PENNSYLVANIA, USA. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929956&pc=DSQ | 3007042319-2025-01079 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED CAUSE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21929978&pc=DSQ | 3007042319-2025-01073 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-04-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO A SUDDEN UNEXPLAINED UNSPECIFIED CIRCULATORY ISSUE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS AN UNSPECIFICED CIRCULATORY ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q4 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21930224&pc=NPT | 2025587-2025-03218 | 2024-01-04 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-29 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Dysphagia/ Odynophagia; Hemoptysis; Pleural Effusion; Pneumonia; Respiratory Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED WITH HEMOPTYSIS. SUBSEQUENTLY, THE PATIENT'S SYMPTOMS IMPROVED AND THE PATIENT WAS DISCHARGED FROM THE HOSPITAL. APPROXIMATELY 8 MONTHS AND 2 WEEKS LATER, THE PATIENT WAS HOSPITALIZED AGAIN DUE TO ASPIRATION PNEUMONIA WITH DYSPHAGIA. PLEURAL EFFUSION ACCUMULATION AND RESPIRATORY FAILURE WERE NOTED. THE PATIENT PASSED AWAY 9 DAYS AFTER THE START OF HOSPITALIZATION. THE CAUSE OF DEATH WAS PLEURAL EFFUSION AND RESPIRATORY FAILURE. PER THE PHYSICIAN, THE PLURAL EFFUSION AND RESPIRATORY FAILURE WERE DUE TO A PANCREATIC HEAD TUMOR. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21930498&pc=NVY | 2649622-2025-11589 | 2025-04-09 04:00 | Death | MPRI | 2025-04-29 04:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Tachycardia; Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. ON THE DATE OF DEATH, THE PATIENT EXPERIENCED MULTIPLE EPISODES OF VENTRICULAR FIBRILLATION (VF) AND NON-SUSTAINED VENTRICULAR TACHYCARDIA. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED POSSIBLE UNDERSENSING DURING THESE EPISODES WHICH MAY HAVE DELAYED DETECTION AND THERAPY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21930568&pc=NVN | 2017865-2025-40007 | 2025-04-06 04:00 | Death | ABBOTT | 2025-04-29 04:00 | NVN | TENDRIL STS | Capturing Problem | Arrhythmia; Respiratory Failure | P960013 |  | 1 | Event Description: FOLLOWING IMPLANT OF THE PACEMAKER, THE RIGHT ATRIAL (RA) LEAD, AND THE RIGHT VENTRICULAR (RV) LEAD, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS RESUSCITATED AND SUBSEQUENTLY HOSPITALIZED AND PUT ON LIFE SUPPORT BEFORE PASSING AWAY. DIAGNOSTIC IMAGING WAS PERFORMED AND NO EVIDENCE OF LEAD DISLODGEMENT OR OTHER MALFUNCTION WAS OBSERVED. INCREASED THRESHOLD WAS OBSERVED; IN THE PHYSICIAN'S OPINION THE INCREASED THRESHOLD WAS DUE TO PATIENT ANATOMY. ADDITIONAL INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE. Manufacturer Narrative: THE REPORTED EVENTS WERE UNACCEPTABLE THRESHOLD AND THE PATIENT WAS DECEASED. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN TWO PIECES. THE REPORTED EVENT OF UNACCEPTABLE THRESHOLD WAS NOT CONFIRMED. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL AND X-RAY EXAMINATIONS OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21930570&pc=DXY | 2017865-2025-40001 | 2025-04-06 04:00 | Death | ABBOTT | 2025-04-29 04:00 | DXY | ENDURITY PACEMAKER, DR, MRI RESTRICTED | Capturing Problem | Arrhythmia; Respiratory Failure | P880086 |  | 1 | Event Description: FOLLOWING IMPLANT OF THE PACEMAKER, THE RIGHT ATRIAL (RA) LEAD, AND THE RIGHT VENTRICULAR (RV) LEAD, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS RESUSCITATED AND SUBSEQUENTLY HOSPITALIZED AND PUT ON LIFE SUPPORT BEFORE PASSING AWAY. DIAGNOSTIC IMAGING WAS PERFORMED AND NO EVIDENCE OF LEAD DISLODGEMENT OR OTHER MALFUNCTION WAS OBSERVED. INCREASED THRESHOLD WAS OBSERVED; IN THE PHYSICIAN'S OPINION THE INCREASED THRESHOLD WAS DUE TO PATIENT ANATOMY. ADDITIONAL INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE.Event Description: IT WAS REPORTED THAT THE DEVICE WAS EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21930575&pc=NVN | 2017865-2025-40008 | 2025-04-06 04:00 | Death | ABBOTT | 2025-04-29 04:00 | NVN | ULTIPACE¿ LEAD, 65CM MRI | Capturing Problem | Arrhythmia; Respiratory Failure | P960013 |  | 1 | Event Description: FOLLOWING IMPLANT OF THE PACEMAKER, THE RIGHT ATRIAL (RA) LEAD, AND THE RIGHT VENTRICULAR (RV) LEAD, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS RESUSCITATED AND SUBSEQUENTLY HOSPITALIZED AND PUT ON LIFE SUPPORT BEFORE PASSING AWAY. DIAGNOSTIC IMAGING WAS PERFORMED AND NO EVIDENCE OF LEAD DISLODGEMENT OR OTHER MALFUNCTION WAS OBSERVED. INCREASED THRESHOLD WAS OBSERVED; IN THE PHYSICIAN'S OPINION THE INCREASED THRESHOLD WAS DUE TO PATIENT ANATOMY. ADDITIONAL INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE. Manufacturer Narrative: THE REPORTED EVENTS WERE UNACCEPTABLE THRESHOLD AND THE PATIENT WAS DECEASED. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN TWO PIECES. THE REPORTED EVENT OF HIGH PACING THRESHOLD WAS NOT CONFIRMED. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES HOWEVER AN INTERNAL SHORT WAS NOTED BETWEEN CONDUCTOR PATHS AT THE DISTAL PORTION DUE TO PROCEDURAL DAMAGE. VISUAL AND X-RAY EXAMINATIONS OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21930992&pc=NPT | 2025587-2025-03223 | 2024-03-10 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-29 04:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 9 YEARS AND 3 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE THE PATIENT DIED. THE DEATH TYPE WAS REPORTED AS CARDIOVASCULAR. NO FURTHER INFORMATION WAS AVAILABLE. THE PHYSICIAN INDICATED THE EVENT WAS NOT RELATED TO THE VALVE. AN AUTOPSY WAS NOT PERFORMED.Event Description: ADDITIONAL INFORMATION RECEIVED WHICH INDICATED THE DEATH OCCURRED APPROXIMATELY 9 YEARS AND 11 MONTHS FOLLOWING THE VALVE IMPLANT. THE OFFICIAL CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: UPDATED DATA: B2, B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21931644&pc=NKE | 2124215-2025-27327 | 2025-03-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-29 04:00 | NKE | VALITUDE X4 CRT-P | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P030005 |  | 1 | Event Description: IT WAS REPORTED BY AN AUTOPSY TECHNICIAN THAT THIS PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN, AND THE PATIENT'S PASSING WAS NOT WITNESSED. THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM HAD BEEN IMPLANTED THE SAME DAY AND WAS ACTIVE AT THE TIME OF THE PATIENT'S DEATH. THE COMPANY REPRESENTATIVE ASSOCIATED WITH THE IMPLANT HAD NO FURTHER INFORMATION. TECHNICAL SERVICES REVIEW OF THE SAVED DISK DATA OF THE DEVICE DEMONSTRATED THAT TWO EPISODES OF PACEMAKER-MEDIATED TACHYCARDIA (PMT) WERE RECORDED ON THE DAY OF DEATH/IMPLANT. THE FIRST PMT EPISODE APPEARED TO BE SINUS TACHYCARDIA AT THE MAXIMUM TRACKING RATE, WHILE THE SECOND EPISODE APPEARED TO BE TRUE PMT. A SIGNIFICANT CHANGE IN THE ATRIAL ELECTROGRAM (EGM) WAS OBSERVED, AS WELL AS POSSIBLE LOSS OF CAPTURE WITH BIVENTRICULAR PACING, AS EVIDENCED BY A LARGE DEFLECTION ON THE RIGHT VENTRICULAR EGM AFTER EVERY DELIVERED PACING COMPLEX. THE PMT TERMINATION ALGORITHM STOPPED WHAT WAS HAPPENING IN THE ATRIUM; HOWEVER, IT WAS DIFFICULT TO DETERMINE IF THERE WAS ATRIAL OR VENTRICULAR CAPTURE. BASED SOLELY ON THE AVAILABLE PMT EPISODES, THE DEVICE WAS OPERATING AS PROGRAMMED AND EXPECTED. THE DATA APPEARED TO SHOW A MARKED CHANGE IN CARDIAC FUNCTION POTENTIALLY CAUSING ATRIAL AND VENTRICULAR LOSS OF CAPTURE. THE DEVICE SYSTEM WAS EXPLANTED POST-MORTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21931818&pc=BZD | 2518422-2025-104982 | 2025-03-18 04:00 | Death | RESPIRONICS, INC. | 2025-04-29 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTO HAS PASSED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21931827&pc=GEI | 3006524618-2025-00201 | 2025-04-01 04:00 | Death | ARTHROCARE CORPORATION | 2025-04-29 04:00 | GEI | UNKNOWN COBLATION ENT DEV | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Pain |  |  | 1 | Event Description: IT WAS REPORTED THAT IN (B)(6) 2025, A 6-YEAR-OLD PATIENT UNDERWENT A ROUTINE TONSILLECTOMY IN WHICH AN UNKNOWN SMITH AND NEPHEW DEVICE WAS INVOLVED. THE SURGERY RESULTED IN SEVERE INTRAOPERATIVE BLEEDING THAT LED TO THE PATIENT BEING KEPT IN THE HOSPITAL FOR A COUPLE OF DAYS BEFORE BEING DISCHARGED HOME. SIX DAYS AFTER BEING DISCHARGED, THE PATIENT COMPLAINED OF PAIN IN HIS THROAT AND SUDDENLY STARTED VOMITING BLOOD, COLLAPSED, AND BECAME UNCONSCIOUS. THE PATIENT WAS RUSHED TO (B)(6) HOSPITAL, IN WHERE WAS DECLARED BRAIN DEAD, AND SUBSEQUENTLY DISCONNECTED FROM LIFE SUPPORT, PASSING AWAY. THE CAUSE AND FURTHER DETAILS ARE UNKNOWN AT THIS TIME. Manufacturer Narrative: H11: INTERNAL COMPLAINT REFERENCE: (B)(4). H3, H6: THE REPORTED DEVICE WAS NOT RETURNED TO THE DESIGNATED COMPLAINT UNIT FOR INDEPENDENT EVALUATION; THEREFORE, NO DEVICE DEFICIENCY COULD BE IDENTIFIED. INSUFFICIENT PRODUCT IDENTIFICATION INFORMATION WAS PROVIDED; AS A RESULT, DEVICE HISTORY, COMPLAINT HISTORY, RISK MANAGEMENT, AND INSTRUCTIONS FOR USE REVIEWS COULD NOT BE CONDUCTED. BASED ON THE LIMITED INFORMATION AVAILABLE, THERE WERE NO INDICATIONS TO SUGGEST THAT THE PRODUCT DID NOT MEET MANUFACTURING SPECIFICATION UPON RELEASE FOR DISTRIBUTION. A CLINICAL REVIEW STATES THAT INTRA-OPERATIVE RISKS EXIST FOR ANY SURGICAL PROCEDURE AND INCLUDE HEMORRHAGE. SHORT-TERM RISKS IN THE POST-TONSILLECTOMY PERIOD ARE WELL-DOCUMENTED AND ALTHOUGH RARE, INCLUDE BUT ARE NOT LIMITED TO FATAL COMPLICATIONS SUCH AS DELAYED HEMORRHAGE (PRIMARY AND/OR SECONDARY), INVOLVING INITIAL SCAB FORMATION AND WHEN THE SCAB DISLODGES (USUALLY BETWEEN 5 TO 10 DAYS). DEHYDRATION (OR LACK OF FLUIDS BY MOUTH POST-OPERATIVELY) IS KNOWN TO INCREASE THE RISK OF POST-TONSILLECTOMY PAIN; NEVERTHELESS, BASED ON THE LIMITED INFORMATION PROVIDED A DEFINITIVE CLINICAL ROOT CAUSE OF THE REPORTED FATAL EVENT CANNOT BE DETERMINED; THEREFORE, THE NEED FOR FURTHER ACTIONS IS NOT INDICATED AT THIS TIME. Manufacturer Narrative: H10: INTERNAL COMPLAINT REFERENCE: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21916774&pc=FRN | 3010293992-2025-00031 | 2025-04-19 04:00 | Death | EITAN MEDICAL LTD. | 2025-04-28 04:00 | FRN | SAPPHIRE M.T INFUSION PUMP - GERMAN | Use of Incorrect Control/Treatment Settings; Excess Flow or Over-Infusion | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K192860 |  | 1 | Event Description: THE COMPLAINT WAS REPORTED BY A CUSTOMER FROM GERMANY. DELIVERY ISSUE.Event Description: THIS COMPLAINT WAS REPORTED BY A CUSTOMER FROM GERMANY. A PROGRAMMING ERROR WHICH RESULTED IN PATIENT DEATH WAS REPORTED. Manufacturer Narrative: INVESTIGATION TYPE: EITAN MEDICAL INVESTIGATED THE PUMP'S EVENT LOG. INVESTIGATION FINDINGS: THE EVENT WAS CONFIRMED TO BE SOLELY A RESULT OF A PROGRAMMING ERROR. THE PUMP FUNCTIONED AS DESIGNED AND DELIVERED THE INFUSION IN ACCORDANCE WITH THE PROGRAMED PARAMETERS. CONCLUSION: THE EVENT WAS CONFIRMED TO BE A RESULT OF A PROGRAMMING ERROR. THE SAPPHIRE PUMP DESIGN MINIMIZES PROGRAMMING ERRORS BY REQUIRING USERS TO REVIEW AND APPROVE PARAMETERS BEFORE STARTING OR REPEATING INFUSIONS, WITH CLEAR ON-SCREEN DISPLAYS THROUGHOUT. THIS REPORTED EVENT OCCURRED OUTSIDE OF THE US ON A DEVICE THAT IS SIMILAR TO THE ONE MARKETED IN THE US AND FOR WHICH THERE IS NO DATA IN GUDID. DEVICE'S PREMARKET SUBMISSION NUMBER: K192860. Manufacturer Narrative: THIS REPORTED EVENT OCCURRED OUTSIDE OF THE US ON A DEVICE THAT IS SIMILAR TO THE ONE MARKETED IN THE US AND FOR WHICH THERE IS NO DATA IN GUDID. DEVICE'S PREMARKET SUBMISSION NUMBER: K161667.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21916819&pc=PNJ | 9612164-2025-02081 | 2024-12-20 05:00 | Death | MEDTRONIC IRELAND | 2025-04-28 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Renal Failure; Thrombocytopenia; Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT NINE DAYS POST IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DECEASED DUE TO THEIR TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE AND UNDERLYING MULTIPLE COMORBIDITIES INCLUDING TYPE TWO DIABETES MELLITUS, RENAL FAILURE, PERIPHERAL ARTERIAL DISEASE, THROMBOCYTOPENIA. IT WAS REPORTED THAT THE PATIENT'S DEATH WAS NOT RELATED TO DEVICE SYSTEM. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21916902&pc=PNJ | 9612164-2025-02084 | 2025-04-04 04:00 | Death | MEDTRONIC IRELAND | 2025-04-28 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED PERICARDIAL EFFUSION POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). IT WAS NOTED THAT THERE MAY HAVE BEEN A CARDIAC PERFORATION. THE LEADLESS IPG REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENTEvent Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. IT WAS NOTED THAT THE DEATH OCCURRED DUE TO A GRAM-POSITIVE COCCI (GPC) INFECTION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENTS DEATH WAS NOT RELATED TO THE LEADLESS IPG. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: [SERIAL (B)(6)). D9: <(><<)>NO> MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917023&pc=OZD | 1220648-2025-27870 | 2025-04-03 04:00 | Death | ABIOMED, INC. | 2025-04-28 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED AN ACCESS SITE HEMATOMA AND PUMP MALPOSITION. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. A PATIENT WITH ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE ON ADMISSION HAD A HEPARIN DRIP ON. ONE DAY PRIOR TO ADMISSION, THE PATIENT UNDERWENT A STRESS TEST WHICH REVEALED ANTERIOR ISCHEMIA. OVERNIGHT THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA ARREST WITH 10-MINUTE DOWN TIME. THE PATIENT WAS INTUBATED AND STARTED ON LEVOPHED AND VASOPRESSIN. THE FOLLOWING DAY, LEFT HEART CATHETERIZATION REVEALED NO NEED FOR INTERVENTION BUT EJECTION FRACTION WAS LESS THAN 10% ON HIGH DOSE PRESSORS. A DECISION WAS MADE TO PLACE A PULMONARY ARTERY CATHETER AND IMPELLA CP DEVICE FOR REST IN THE UNIT. THE IMPELLA CP WAS SUCCESSFULLY IMPLANTED AND FOLLOWING REMOVAL OF THE ARTERIAL LINE, A HEMATOMA WAS PRESENT AT ACCESS SITE. SYSTEMIC HEPARIN WAS DISCONTINUED, NEW FORWARD TENSION SUTURES WERE PLACED, AND FEM STOP WAS APPLIED TO ADDRESS THE HEMATOMA. FOLLOWING, IMPELLA POSITION IN AORTA ALARM TRIGGERED. THE TEAM DID NOT WANT TO GIVE VOLUME OR TURN IMPELLA PUMP FLOW DOWN WITH SUCTION ALARMS. AN ECHOCARDIOGRAM WAS COMPLETED AND THE IMPELLA WAS MEASURING 5.7 CENTIMETERS BELOW THE AORTIC VALVE. THE IMPELLA CP PUMP WAS REPOSITIONED AT BEDSIDE WITH ECHOCARDIOGRAM GUIDANCE AND THEREAFTER MEASURED 3.6 CENTIMETERS BELOW THE AORTIC VALVE. THE PATIENT WAS NOTED TO BE STILL HYPOTENSIVE WITH TACHYCARDIA. ELECTROCARDIOGRAM WAS OBTAINED FOR POSSIBLE CARDIOVERSION AND DECREASE IN MILRINONE WAS NOTED. APPROXIMATELY TWO DAYS LATER, THE PATIENT WOULD EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, ¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917055&pc=OZD | 1220648-2025-27871 | 2025-04-06 04:00 | Death | ABIOMED, INC. | 2025-04-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED LIMB ISCHEMIA DISTAL TO IMPELLA INSERTION SITE. NO INFORMATION ABOUT INTERVENTION WAS NOTED. HOWEVER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917217&pc=BZD | 2518422-2025-104845 | 2025-04-24 04:00 | Death | RESPIRONICS, INC. | 2025-04-28 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917336&pc=OZD | 1220648-2025-27875 | 2025-04-07 04:00 | Death | ABIOMED, INC. | 2025-04-28 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Difficult to Insert; Defective Device | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING THE INITIAL IMPLANT ATTEMPT OF AN IMPELLA CP PUMP, THE INTRODUCER KINKED AND SUBSEQUENTLY PREVENTED ADVANCING THE PUMP INTO POSITION. A COOK SHEATH WAS PLACED WHICH ALLOWED FOR SUCCESSFUL IMPLANT. FOLLOWING IMPLANT, IT WAS NOTED THAT THE SHEATH WAS OCCLUDING BLOOD FLOW TO THE LEG. A DISTAL PERFUSION CATHETER WAS PLACED TO TREAT THE CONDITION. AFTER ONE DAY OF SUPPORT, THE PUMP GENERATED SUCTION EVENTS FOLLOWING REPOSITIONING. ESCALATION TO AN IMPELLA 5.5 PUMP WAS DISCUSSED, BUT DUE TO THE PATIENT'S AGE AND LACK OF DURABLE OPTIONS, THE DECISION WAS MADE TO MOVE THE PATIENT TO COMFORT CARE AND THE PATIENT LATER PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DIFFICULTY INSERTING/REMOVING PUMP THROUGH INTRODUCER, INTRODUCER DAMAGE ¿ MECHANICAL, LIMB ISCHEMIA, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO LACK OF DATA LOGS OR PRODUCT RETURNED, THE ROOT CAUSE OF THE LOW PUMP FLOW CANNOT BE DETERMINED. THE ROOT CAUSE OF THE INTRODUCER DAMAGE - MECHANICAL IS MOST LIKELY DUE TO A SHEATH KINK DUE TO THE PATIENTS TORTUOUS VESSELS. THE ROOT CAUSE OF THE DIFFICULTY INSERTING/REMOVING PUMP THROUGH INTRODUCER IS MOST LIKELY DUE TO THE SHEATH KINK. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917383&pc=BZD | 2518422-2025-104848 | 2025-04-24 04:00 | Death | RESPIRONICS, INC. | 2025-04-28 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917434&pc=OZD | 1220648-2025-27876 | 2025-04-07 04:00 | Death | ABIOMED, INC. | 2025-04-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND DEVELOPED A HEMATOMA AT THE ACCESS SITE. ONE UNIT OF BLOOD PRODUCTS WERE GIVEN. SUPPORT WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE ANTICOAGULATION MANAGEMENT BECAUSE THE ACTIVATED CLOTTING TIME (ACT) VALUES (ACT 228) WERE NOTED BE HIGHER THE THERAPEUTIC RANGE AT THE TIME OF THE BLEEDING AND THE HEPARIN WAS KEPT ON HOLD TO MANGE THE BLEED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917585&pc=DTZ | 8010762-2025-0000193 | 2025-04-16 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-28 04:00 | DTZ | TUBING SET | Adverse Event Without Identified Device or Use Problem | Hemolysis; Septic Shock | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE INVOLVED CARDIOHELP DEVICE WAS NOISY. THUS, THE CARDIOHELP DEVICE WAS REPLACED. ON (B)(6) 2025 THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025 THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. THE AFFECTED CARDIOHELP DEVICE WILL BE INVESTIGATED IN COMPLAINT # (B)(4). HLS SET COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID #(B)(4).Event Description: (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025, THE INVOLVED CARDIOHELP DEVICE WAS NOISY. THUS, THE CARDIOHELP DEVICE WAS REPLACED. ON (B)(6) 2025 THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025 THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. DUE TO DEVICE EXCHANGES AND AS THE INVOLVED PATIENT PASSED AWAY A REPORT IS REQUIRED. THE AFFECTED HLS SET WAS INVESTIGATED IN THE GETINGE LABORATORY ON (B)(6) 2025 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES, AND NO ABNORMALITIES WERE DETECTED DURING TESTING. THUS, NO EXACT ROOT CAUSE COULD BE DETERMINED." THE INVOLVED CARDIOHELP DEVICE IS INVESTIGATED WITHIN COMPLAINT # (B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER # 8010762-2025-0000194. ON 2025-05-27, THE INFORMATION WAS RECEIVED THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. A MEDICAL REVIEW WAS PERFORMED ON 2025-06-20 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES A 46-YEAR-OLD MALE PATIENT (184 CM, 180 KG) UNDERGOING VVECMO THERAPY SUPPORTED CH-I AND HLS-SET. ON (B)(6) 2025, A MECHANICAL NOISE WAS NOTED DURING ONGOING THERAPY. AT THE TIME OF INITIAL OBSERVATION, NO DETERIORATION IN DEVICE PERFORMANCE WAS DETECTED. ECMO SUPPORT WAS TECHNICALLY STABLE, WITH CONSISTENT FLOW PARAMETERS (5.0 - 5.3 L/MIN), PUMP ROTATIONAL SPEEDS (3,400 - 4,100 RPM), AND DELTA P (35 - 38 MMHG). VENOUS DRAINAGE PRESSURES WERE WITHIN ACCEPTABLE NEGATIVE RANGES (-63 TO -78 MMHG). DUE TO PERSISTING ACOUSTIC ANOMALIES, THE CH-I WAS ELECTIVELY EXCHANGED ON (B)(6). NO CLINICAL OR TECHNICAL CHANGE RESULTED FROM THE INTERVENTION. SUBSEQUENTLY, A SUDDEN AND PRONOUNCED INCREASE IN HEMOLYSIS PARAMETERS WAS RECORDED, INCLUDING AN ELEVATION IN PLASMA-FREE HEMOGLOBIN (FHB) FROM 2 MG/DL TO 114 MG/DL, LDH FROM 400 TO 1,300 U/L, AND TOTAL BILIRUBIN FROM 1 TO 5 MG/DL. DESPITE THESE BIOCHEMICAL FINDINGS, NO ABNORMAL TRENDS IN ECMO FLOW OR PRESSURES WERE DETECTED. THAT SAID, THE LABORATORY RESULTS PROVIDED BY THE CUSTOMER ARE SUGGESTIVE OF HEMOLYSIS AND APPEAR TO CONFIRM THE ASSESSMENT OF THE ATTENDING CLINICIAN PRE-EXCHANGE. TECHNICAL EVALUATIONS OF BOTH THE CARDIOHELP CONSOLE AND THE RETURNED HLS-SET DID NOT IDENTIFY ANY MALFUNCTION OR PERFORMANCE DEVIATION. FUNCTIONAL, LEAKAGE, AND CALIBRATION CHECKS WERE COMPLETED ACCORDING TO FACTORY SPECIFICATIONS, AND NO DAMAGE, OCCLUSION, OR THROMBOSIS WAS EVIDENT. VISUAL INSPECTION OF THE HLS-SET REVEALED RESIDUAL BLOOD IN THE PUMP MODULE BUT NO IDENTIFIABLE MECHANICAL DEFECT. ALL PRESSURE AND FLOW PARAMETERS REMAINED WITHIN SPECIFICATION DURING LABORATORY TESTING. FOLLOWING THE ONSET OF HEMOLYSIS, CLINICAL MANAGEMENT INCLUDED THE REPLACEMENT OF THE HLS-SET. POST-EXCHANGE, A RAPID DECLINE IN HEMOLYSIS MARKERS WAS OBSERVED: FHB DECREASED FROM 114 TO 3 MG/DL, BILIRUBIN FROM 5 TO 3 MG/DL, AND HAPTOGLOBIN INCREASED FROM 0.1 TO 0.5 G/L. THIS IMMEDIATE BIOCHEMICAL NORMALIZATION APPEARS TO SUPPORT A TEMPORAL AND POTENTIALLY CAUSAL ASSOCIATION BETWEEN THE HLS-SET AND THE HEMOLYTIC REPORT. ALTHOUGH NO VISIBLE OR MECHANICAL CAUSE WAS IDENTIFIED DURING TECHNICAL ANALYSIS, THE CLINICAL FINDINGS SUGGEST THE POSSIBILITY OF BIOLOGICALLY MEDIATED OXYGENATOR IMPAIRMENT. THIS MAY INCLUDE MICRO THROMBOTIC DEPOSITION, FIBRIN ACCUMULATION, OR PROTEIN ADHESION ON THE MEMBRANE SURFACE, NONE OF WHICH ARE CONSISTENTLY DETECTABLE VIA ROUTINE POST-USE INSPECTION. SUCH MECHANISMS HAVE BEEN DESCRIBED IN PUBLISHED LITERATURE AS POTENTIAL CAUSES OF HEMOLYSIS DURING VV-ECMO, EVEN IN THE ABSENCE OF ELEVATED PRESSURE GRADIENTS OR DEVICE MALFUNCTION. THE PATIENT ULTIMATELY EXPIRED DUE TO SEPTIC SHOCK. CLINICAL INDICATORS, INCLUDING RISING PTT VALUES (46.4 TO 68.3 S), POINT TOWARD SYSTEMIC INFLAMMATORY PROGRESSION (4, 5). NO SIGNS OF OVERT DIC OR BLEEDING WERE DOCUMENTED PRIOR TO EXPIRATION. BASED ON THE CLINICAL TIMELINE, THE HEMOLYTIC EVENT PRECEDED THE WORSENING SEPSIS, AND WHILE IT MAY HAVE CONTRIBUTED TO CLINICAL DETERIORATION, IT WAS (REPORTEDLY) NOT THE DIRECT CAUSE OF DEATH. THEREFORE, A CAUSAL RELATIONSHIP BETWEEN THE PRODUCT USE AND THE PATIENT OUTCOME CANNOT BE ESTABLISHED." THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-13. ACCORDING TO THE FINAL TEST RESULTS, THE MODULE WITH LOT # 3000446977 AND UDI # (B)(4) PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED PRODUCT AND FAILURE. BASED ON THE RESULTS, THE REPORTED EVENT "INCREASE IN THE HEMOLYSIS PARAMETERS" WAS NOT RELATED TO A PRODUCT MALFUNCTION. THIS COMPLAINT WAS FOUND AS A SINGLE EVENT IN THE DATABASE OF CUSTOMER COMPLAINTS FOR THE INDICATED TIMEFRAME. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: THE AFFECTED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION AND WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-06-12 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES AND NO ABNORMALITIES WERE DETECTED DURING TESTING." THE INVOLVED CARDIOHELP DEVICE IS INVESTIGATED WITHIN COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000194. A MEDICAL REVIEW WAS REQUESTED BY GETINGE MEDICAL AFFAIRS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917669&pc=DTQ | 8010762-2025-0000194 | 2025-04-15 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-28 04:00 | DTQ | CARDIOHELP | Adverse Event Without Identified Device or Use Problem; Noise, Audible | Hemolysis; Septic Shock | K133598 |  | 1 | Event Description: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE CARDIOHELP DEVICE WAS NOISY. THUS, THE CARDIOHELP DEVICE WAS REPLACED. ON (B)(6) 2025 THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025 THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. THE AFFECTED HLS DEVICE WILL BE INVESTIGATED IN COMPLAINT # (B)(4). DUE TO DEVICE EXCHANGES AND AS THE INVOLVED PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID: (B)(4).Event Description: COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID#: (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE CARDIOHELP MAKES A MECHANICAL NOISE. FOR SAFETY REASONS THE CARDIOHELP WAS EXCHANGED. ON (B)(6) 2025, THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025, THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. DUE TO DEVICE EXCHANGES AND AS THE INVOLVED PATIENT PASSED AWAY A REPORT IS REQUIRED. NEW INFORMATION WAS RECEIVED ON 2025-05-27, THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT#: (B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER#: 8010762-2025-0000193. THE AFFECTED HLS SET WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-06-12 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES AND NO ABNORMALITIES WERE DETECTED DURING TESTING. THUS, NO EXACT ROOT CAUSE COULD BE DETERMINED." A MEDICAL REVIEW WAS PERFORMED ON 2025-06-20 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES A 46-YEAR-OLD MALE PATIENT (184 CM, 180 KG) UNDERGOING VVECMO THERAPY SUPPORTED CH-I AND HLS-SET. ON (B)(6) 2025, A MECHANICAL NOISE WAS NOTED DURING ONGOING THERAPY. AT THE TIME OF INITIAL OBSERVATION, NO DETERIORATION IN DEVICE PERFORMANCE WAS DETECTED. ECMO SUPPORT WAS TECHNICALLY STABLE, WITH CONSISTENT FLOW PARAMETERS (5.0 - 5.3 L/MIN), PUMP ROTATIONAL SPEEDS (3,400 - 4,100 RPM), AND DELTA P (35 - 38 MMHG). VENOUS DRAINAGE PRESSURES WERE WITHIN ACCEPTABLE NEGATIVE RANGES (-63 TO -78 MMHG). DUE TO PERSISTING ACOUSTIC ANOMALIES, THE CH-I WAS ELECTIVELY EXCHANGED ON (B)(6). NO CLINICAL OR TECHNICAL CHANGE RESULTED FROM THE INTERVENTION. SUBSEQUENTLY, A SUDDEN AND PRONOUNCED INCREASE IN HEMOLYSIS PARAMETERS WAS RECORDED, INCLUDING AN ELEVATION IN PLASMA-FREE HEMOGLOBIN (FHB) FROM 2 MG/DL TO 114 MG/DL, LDH FROM 400 TO 1,300 U/L, AND TOTAL BILIRUBIN FROM 1 TO 5 MG/DL. DESPITE THESE BIOCHEMICAL FINDINGS, NO ABNORMAL TRENDS IN ECMO FLOW OR PRESSURES WERE DETECTED. THAT SAID, THE LABORATORY RESULTS PROVIDED BY THE CUSTOMER ARE SUGGESTIVE OF HEMOLYSIS AND APPEAR TO CONFIRM THE ASSESSMENT OF THE ATTENDING CLINICIAN PRE-EXCHANGE. TECHNICAL EVALUATIONS OF BOTH THE CARDIOHELP CONSOLE AND THE RETURNED HLS-SET DID NOT IDENTIFY ANY MALFUNCTION OR PERFORMANCE DEVIATION. FUNCTIONAL, LEAKAGE, AND CALIBRATION CHECKS WERE COMPLETED ACCORDING TO FACTORY SPECIFICATIONS, AND NO DAMAGE, OCCLUSION, OR THROMBOSIS WAS EVIDENT. VISUAL INSPECTION OF THE HLS-SET REVEALED RESIDUAL BLOOD IN THE PUMP MODULE BUT NO IDENTIFIABLE MECHANICAL DEFECT. ALL PRESSURE AND FLOW PARAMETERS REMAINED WITHIN SPECIFICATION DURING LABORATORY TESTING. FOLLOWING THE ONSET OF HEMOLYSIS, CLINICAL MANAGEMENT INCLUDED THE REPLACEMENT OF THE HLS-SET. POST-EXCHANGE, A RAPID DECLINE IN HEMOLYSIS MARKERS WAS OBSERVED: FHB DECREASED FROM 114 TO 3 MG/DL, BILIRUBIN FROM 5 TO 3 MG/DL, AND HAPTOGLOBIN INCREASED FROM 0.1 TO 0.5 G/L. THIS IMMEDIATE BIOCHEMICAL NORMALIZATION APPEARS TO SUPPORT A TEMPORAL AND POTENTIALLY CAUSAL ASSOCIATION BETWEEN THE HLS-SET AND THE HEMOLYTIC REPORT. ALTHOUGH NO VISIBLE OR MECHANICAL CAUSE WAS IDENTIFIED DURING TECHNICAL ANALYSIS, THE CLINICAL FINDINGS SUGGEST THE POSSIBILITY OF BIOLOGICALLY MEDIATED OXYGENATOR IMPAIRMENT. THIS MAY INCLUDE MICRO THROMBOTIC DEPOSITION, FIBRIN ACCUMULATION, OR PROTEIN ADHESION ON THE MEMBRANE SURFACE, NONE OF WHICH ARE CONSISTENTLY DETECTABLE VIA ROUTINE POST-USE INSPECTION. SUCH MECHANISMS HAVE BEEN DESCRIBED IN PUBLISHED LITERATURE AS POTENTIAL CAUSES OF HEMOLYSIS DURING VV-ECMO, EVEN IN THE ABSENCE OF ELEVATED PRESSURE GRADIENTS OR DEVICE MALFUNCTION. THE PATIENT ULTIMATELY EXPIRED DUE TO SEPTIC SHOCK. CLINICAL INDICATORS, INCLUDING RISING PTT VALUES (46.4 TO 68.3 S), POINT TOWARD SYSTEMIC INFLAMMATORY PROGRESSION (4, 5). NO SIGNS OF OVERT DIC OR BLEEDING WERE DOCUMENTED PRIOR TO EXPIRATION. BASED ON THE CLINICAL TIMELINE, THE HEMOLYTIC EVENT PRECEDED THE WORSENING SEPSIS, AND WHILE IT MAY HAVE CONTRIBUTED TO CLINICAL DETERIORATION, IT WAS (REPORTEDLY) NOT THE DIRECT CAUSE OF DEATH. THEREFORE, A CAUSAL RELATIONSHIP BETWEEN THE PRODUCT USE AND THE PATIENT OUTCOME CANNOT BE ESTABLISHED." IN THE CARDIOHELP INSTRUCTIONS FOR USE (IFU) CHAPTER "USING THE EMERGENCY DRIVE WITH THE DISPOSABLE HLS RETAINER" IS STATED THAT THE EMERGENCY DRIVE CAN BE USED TO MANUALLY CONTROL THE BLOOD FLOW IN CASE OF A FAILED CARDIOHELP. IN THE INSTRUCTION FOR USE (IFU) IS STATED THAT IT IS NECESSARY TO DECREASE THE FLOW TO ZERO BEFORE DISCONNECTING AND CONNECTING THE DISPOSABLE TO THE DEVICE. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2025-06-24 FOR THE PERIOD OF 2014-09-22 TO 2025-04-22. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2014-09-22. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE CARDIOHELP. BASED ON THE RESULTS THE REPORTED FAILURE "MECHANICAL NOISE ON CARDIOHELP" COULD NOT BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: NEW INFORMATION WAS RECEIVED ON 2025-05-27, THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT # (B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000193. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THE SERIAL NUMBER OF THE CARDIOHELP WAS NOT PROVIDED. THUS, NO UDI NUMBER CAN BE PROVIDED.THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT# (B)(4). THE INVOLVED CARDIOHELP WILL INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21917980&pc=NGV | 2124215-2025-26635 | 2025-02-17 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-28 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA CLINICAL STUDY: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 24 MM WATCHMAN FLX CLOSURE DEVICE WAS IMPLANTED WITH DEPLOYED DEVICE DIAMETER OF 20.6 MM. THE PATIENT WAS DISCHARGED THE SAME DAY ON APIXABAN (10 MG/DAY) AND ASPIRIN (81 MG/DAY). 629 DAYS POST INDEX PROCEDURE, THE PATIENT DIED. AT THE TIME OF DEATH, THE PATIENT WAS ON ASPIRIN AND APIXABAN. THE CAUSE OF DEATH REMAINS UNKNOWN, AND NO FURTHER DETAILS WERE MADE AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21918017&pc=NPT | 2015691-2025-03306 | 2025-03-06 05:00 | Death | EDWARDS LIFESCIENCES | 2025-04-28 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM FRANCE, THE PATIENT UNDERWENT INTERVENTION FOR A 29 MM NON-EDWARDS VALVE BY TRANSFEMORAL APPROACH. AFTER IMPLANTATION, PARAVALVULAR LEAK WAS OBSERVED THAT PERSISTED AFTER POSTDILATION. IT WAS DECIDED TO IMPLANT A 26 MM SAPIEN 3 ULTRA VALVE BY TRANSFEMORAL APPROACH. DURING DEPLOYMENT OF THE EDWARDS VALVE, THE PATIENT HAD AN ANNULAR RUPTURE. A SURGICAL DRAIN WAS USED TO COLLECT AND REINJECT BLOOD. AT THE SAME TIME, IT WAS DECIDED TO IMPLANT A NEW VALVE AS A VIVIV WITH ANOTHER 26 MM SAPIEN 3 ULTRA TO TRY TO STOP THE BLOOD LEAK BUT THIS DID NOT WORK AND THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (LARGE AMOUNT OF AORTIC VALVE CALCIFICATION) AND OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21918430&pc=DTZ | 3013876692-2025-00030 | 2025-04-16 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-28 04:00 | DTZ | TUBING SET | Adverse Event Without Identified Device or Use Problem | Hemolysis; Septic Shock | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE INVOLVED CARDIOHELP DEVICE WAS NOISY. THUS, THE CARDIOHELP DEVICE WAS REPLACED. ON (B)(6) 2025- THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025 THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. THE AFFECTED CARDIOHELP DEVICE WILL BE INVESTIGATED IN COMPLAINT # (B)(4). HLS SET COMPLAINT ID # (B)(4).Event Description: COMPLAINT: (B)(4).Event Description: COMPLAINT ID #(B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE INVOLVED CARDIOHELP DEVICE WAS NOISY. THUS, THE CARDIOHELP DEVICE WAS REPLACED. ON (B)(6) 2025 THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025 THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. DUE TO DEVICE EXCHANGES AND AS THE INVOLVED PATIENT PASSED AWAY A REPORT IS REQUIRED. THE AFFECTED HLS SET WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-06-12 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES, AND NO ABNORMALITIES WERE DETECTED DURING TESTING. THUS, NO EXACT ROOT CAUSE COULD BE DETERMINED." THE INVOLVED CARDIOHELP DEVICE IS INVESTIGATED WITHIN COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000194. ON 2025-05-27, THE INFORMATION WAS RECEIVED THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. A MEDICAL REVIEW WAS PERFORMED ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES A 46-YEAR-OLD MALE PATIENT (184 CM, 180 KG) UNDERGOING VVECMO THERAPY SUPPORTED CH-I AND HLS-SET. ON APRIL 15TH, 2025, A MECHANICAL NOISE WAS NOTED DURING ONGOING THERAPY. AT THE TIME OF INITIAL OBSERVATION, NO DETERIORATION IN DEVICE PERFORMANCE WAS DETECTED. ECMO SUPPORT WAS TECHNICALLY STABLE, WITH CONSISTENT FLOW PARAMETERS (5.0 - 5.3 L/MIN), PUMP ROTATIONAL SPEEDS (3,400 - 4,100 RPM), AND DELTA P (35 - 38 MMHG). VENOUS DRAINAGE PRESSURES WERE WITHIN ACCEPTABLE NEGATIVE RANGES (-63 TO -78 MMHG). DUE TO PERSISTING ACOUSTIC ANOMALIES, THE CH-I WAS ELECTIVELY EXCHANGED ON APRIL 16TH. NO CLINICAL OR TECHNICAL CHANGE RESULTED FROM THE INTERVENTION. SUBSEQUENTLY, A SUDDEN AND PRONOUNCED INCREASE IN HEMOLYSIS PARAMETERS WAS RECORDED, INCLUDING AN ELEVATION IN PLASMA-FREE HEMOGLOBIN (FHB) FROM 2 MG/DL TO 114 MG/DL, LDH FROM 400 TO 1,300 U/L, AND TOTAL BILIRUBIN FROM 1 TO 5 MG/DL. DESPITE THESE BIOCHEMICAL FINDINGS, NO ABNORMAL TRENDS IN ECMO FLOW OR PRESSURES WERE DETECTED. THAT SAID, THE LABORATORY RESULTS PROVIDED BY THE CUSTOMER ARE SUGGESTIVE OF HEMOLYSIS AND APPEAR TO CONFIRM THE ASSESSMENT OF THE ATTENDING CLINICIAN PRE-EXCHANGE. TECHNICAL EVALUATIONS OF BOTH THE CARDIOHELP CONSOLE AND THE RETURNED HLS-SET DID NOT IDENTIFY ANY MALFUNCTION OR PERFORMANCE DEVIATION. FUNCTIONAL, LEAKAGE, AND CALIBRATION CHECKS WERE COMPLETED ACCORDING TO FACTORY SPECIFICATIONS, AND NO DAMAGE, OCCLUSION, OR THROMBOSIS WAS EVIDENT. VISUAL INSPECTION OF THE HLS-SET REVEALED RESIDUAL BLOOD IN THE PUMP MODULE BUT NO IDENTIFIABLE MECHANICAL DEFECT. ALL PRESSURE AND FLOW PARAMETERS REMAINED WITHIN SPECIFICATION DURING LABORATORY TESTING. FOLLOWING THE ONSET OF HEMOLYSIS, CLINICAL MANAGEMENT INCLUDED THE REPLACEMENT OF THE HLS-SET. POST-EXCHANGE, A RAPID DECLINE IN HEMOLYSIS MARKERS WAS OBSERVED: FHB DECREASED FROM 114 TO 3 MG/DL, BILIRUBIN FROM 5 TO 3 MG/DL, AND HAPTOGLOBIN INCREASED FROM 0.1 TO 0.5 G/L. THIS IMMEDIATE BIOCHEMICAL NORMALIZATION APPEARS TO SUPPORT A TEMPORAL AND POTENTIALLY CAUSAL ASSOCIATION BETWEEN THE HLS-SET AND THE HEMOLYTIC REPORT. ALTHOUGH NO VISIBLE OR MECHANICAL CAUSE WAS IDENTIFIED DURING TECHNICAL ANALYSIS, THE CLINICAL FINDINGS SUGGEST THE POSSIBILITY OF BIOLOGICALLY MEDIATED OXYGENATOR IMPAIRMENT. THIS MAY INCLUDE MICRO THROMBOTIC DEPOSITION, FIBRIN ACCUMULATION, OR PROTEIN ADHESION ON THE MEMBRANE SURFACE, NONE OF WHICH ARE CONSISTENTLY DETECTABLE VIA ROUTINE POST-USE INSPECTION. SUCH MECHANISMS HAVE BEEN DESCRIBED IN PUBLISHED LITERATURE AS POTENTIAL CAUSES OF HEMOLYSIS DURING VV-ECMO, EVEN IN THE ABSENCE OF ELEVATED PRESSURE GRADIENTS OR DEVICE MALFUNCTION. THE PATIENT ULTIMATELY EXPIRED DUE TO SEPTIC SHOCK. CLINICAL INDICATORS, INCLUDING RISING PTT VALUES (46.4 TO 68.3 S), POINT TOWARD SYSTEMIC INFLAMMATORY PROGRESSION (4, 5). NO SIGNS OF OVERT DIC OR BLEEDING WERE DOCUMENTED PRIOR TO EXPIRATION. BASED ON THE CLINICAL TIMELINE, THE HEMOLYTIC EVENT PRECEDED THE WORSENING SEPSIS, AND WHILE IT MAY HAVE CONTRIBUTED TO CLINICAL DETERIORATION, IT WAS (REPORTEDLY) NOT THE DIRECT CAUSE OF DEATH. THEREFORE, A CAUSAL RELATIONSHIP BETWEEN THE PRODUCT USE AND THE PATIENT OUTCOME CANNOT BE ESTABLISHED." THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-13. ACCORDING TO THE FINAL TEST RESULTS, THE MODULE WITH LOT# 3000446977 AND UDI# (B)(4) PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED PRODUCT AND FAILURE. BASED ON THE RESULTS, THE REPORTED EVENT "INCREASE IN THE HEMOLYSIS PARAMETERS" WAS NOT RELATED TO A PRODUCT MALFUNCTION. THIS COMPLAINT WAS FOUND AS A SINGLE EVENT IN THE DATABASE OF CUSTOMER COMPLAINTS FOR THE INDICATED TIMEFRAME. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073. Manufacturer Narrative: THE AFFECTED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION AND WAS INVESTIGATED IN THE GETINGE LABORATORY ON (B)(6) 2025 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES AND NO ABNORMALITIES WERE DETECTED DURING TESTING." THE INVOLVED CARDIOHELP DEVICE IS INVESTIGATED WITHIN COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000194. A MEDICAL REVIEW WAS REQUESTED BY GETINGE MEDICAL AFFAIRS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21918517&pc=DTQ | 3013876692-2025-00031 | 2025-04-15 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-04-28 04:00 | DTQ | CARDIOHELP | Adverse Event Without Identified Device or Use Problem; Noise, Audible | Hemolysis; Septic Shock | K133598 |  | 1 | Event Description: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE CARDIOHELP DEVICE WAS NOISY. THUS, THE CARDIOHELP DEVICE WAS REPLACED. ON (B)(6) 2025 THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025 THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. THE AFFECTED HLS DEVICE WILL BE INVESTIGATED IN COMPLAINT #(B)(4). DUE TO DEVICE EXCHANGES AND AS THE INVOLVED PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID: (B)(4).Event Description: COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID#: (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN GERMANY. IT WAS REPORTED THAT DURING PATIENT TREATMENT ON (B)(6) 2025 THE CARDIOHELP MAKES A MECHANICAL NOISE. FOR SAFETY REASONS THE CARDIOHELP WAS EXCHANGED. ON (B)(6) 2025, THERE WAS AN INCREASE IN THE HEMOLYSIS PARAMETERS. AFTER THIS NOTICE, THE HLS SET WAS REPLACED. ON (B)(6) 2025, THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY DUE TO A SEPTIC SHOCK. DUE TO DEVICE EXCHANGES AND AS THE INVOLVED PATIENT PASSED AWAY A REPORT IS REQUIRED. NEW INFORMATION WAS RECEIVED ON 2025-05-27, THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT#: (B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER#: 8010762-2025-0000193. THE AFFECTED HLS SET WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-06-12 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES AND NO ABNORMALITIES WERE DETECTED DURING TESTING. THUS, NO EXACT ROOT CAUSE COULD BE DETERMINED." A MEDICAL REVIEW WAS PERFORMED ON 2025-06-20 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES A 46-YEAR-OLD MALE PATIENT (184 CM, 180 KG) UNDERGOING VVECMO THERAPY SUPPORTED CH-I AND HLS-SET. ON (B)(6) 2025, A MECHANICAL NOISE WAS NOTED DURING ONGOING THERAPY. AT THE TIME OF INITIAL OBSERVATION, NO DETERIORATION IN DEVICE PERFORMANCE WAS DETECTED. ECMO SUPPORT WAS TECHNICALLY STABLE, WITH CONSISTENT FLOW PARAMETERS (5.0 - 5.3 L/MIN), PUMP ROTATIONAL SPEEDS (3,400 - 4,100 RPM), AND DELTA P (35 - 38 MMHG). VENOUS DRAINAGE PRESSURES WERE WITHIN ACCEPTABLE NEGATIVE RANGES (-63 TO -78 MMHG). DUE TO PERSISTING ACOUSTIC ANOMALIES, THE CH-I WAS ELECTIVELY EXCHANGED ON (B)(6). NO CLINICAL OR TECHNICAL CHANGE RESULTED FROM THE INTERVENTION. SUBSEQUENTLY, A SUDDEN AND PRONOUNCED INCREASE IN HEMOLYSIS PARAMETERS WAS RECORDED, INCLUDING AN ELEVATION IN PLASMA-FREE HEMOGLOBIN (FHB) FROM 2 MG/DL TO 114 MG/DL, LDH FROM 400 TO 1,300 U/L, AND TOTAL BILIRUBIN FROM 1 TO 5 MG/DL. DESPITE THESE BIOCHEMICAL FINDINGS, NO ABNORMAL TRENDS IN ECMO FLOW OR PRESSURES WERE DETECTED. THAT SAID, THE LABORATORY RESULTS PROVIDED BY THE CUSTOMER ARE SUGGESTIVE OF HEMOLYSIS AND APPEAR TO CONFIRM THE ASSESSMENT OF THE ATTENDING CLINICIAN PRE-EXCHANGE. TECHNICAL EVALUATIONS OF BOTH THE CARDIOHELP CONSOLE AND THE RETURNED HLS-SET DID NOT IDENTIFY ANY MALFUNCTION OR PERFORMANCE DEVIATION. FUNCTIONAL, LEAKAGE, AND CALIBRATION CHECKS WERE COMPLETED ACCORDING TO FACTORY SPECIFICATIONS, AND NO DAMAGE, OCCLUSION, OR THROMBOSIS WAS EVIDENT. VISUAL INSPECTION OF THE HLS-SET REVEALED RESIDUAL BLOOD IN THE PUMP MODULE BUT NO IDENTIFIABLE MECHANICAL DEFECT. ALL PRESSURE AND FLOW PARAMETERS REMAINED WITHIN SPECIFICATION DURING LABORATORY TESTING. FOLLOWING THE ONSET OF HEMOLYSIS, CLINICAL MANAGEMENT INCLUDED THE REPLACEMENT OF THE HLS-SET. POST-EXCHANGE, A RAPID DECLINE IN HEMOLYSIS MARKERS WAS OBSERVED: FHB DECREASED FROM 114 TO 3 MG/DL, BILIRUBIN FROM 5 TO 3 MG/DL, AND HAPTOGLOBIN INCREASED FROM 0.1 TO 0.5 G/L. THIS IMMEDIATE BIOCHEMICAL NORMALIZATION APPEARS TO SUPPORT A TEMPORAL AND POTENTIALLY CAUSAL ASSOCIATION BETWEEN THE HLS-SET AND THE HEMOLYTIC REPORT. ALTHOUGH NO VISIBLE OR MECHANICAL CAUSE WAS IDENTIFIED DURING TECHNICAL ANALYSIS, THE CLINICAL FINDINGS SUGGEST THE POSSIBILITY OF BIOLOGICALLY MEDIATED OXYGENATOR IMPAIRMENT. THIS MAY INCLUDE MICRO THROMBOTIC DEPOSITION, FIBRIN ACCUMULATION, OR PROTEIN ADHESION ON THE MEMBRANE SURFACE, NONE OF WHICH ARE CONSISTENTLY DETECTABLE VIA ROUTINE POST-USE INSPECTION. SUCH MECHANISMS HAVE BEEN DESCRIBED IN PUBLISHED LITERATURE AS POTENTIAL CAUSES OF HEMOLYSIS DURING VV-ECMO, EVEN IN THE ABSENCE OF ELEVATED PRESSURE GRADIENTS OR DEVICE MALFUNCTION. THE PATIENT ULTIMATELY EXPIRED DUE TO SEPTIC SHOCK. CLINICAL INDICATORS, INCLUDING RISING PTT VALUES (46.4 TO 68.3 S), POINT TOWARD SYSTEMIC INFLAMMATORY PROGRESSION (4, 5). NO SIGNS OF OVERT DIC OR BLEEDING WERE DOCUMENTED PRIOR TO EXPIRATION. BASED ON THE CLINICAL TIMELINE, THE HEMOLYTIC EVENT PRECEDED THE WORSENING SEPSIS, AND WHILE IT MAY HAVE CONTRIBUTED TO CLINICAL DETERIORATION, IT WAS (REPORTEDLY) NOT THE DIRECT CAUSE OF DEATH. THEREFORE, A CAUSAL RELATIONSHIP BETWEEN THE PRODUCT USE AND THE PATIENT OUTCOME CANNOT BE ESTABLISHED." IN THE CARDIOHELP INSTRUCTIONS FOR USE (IFU) CHAPTER "USING THE EMERGENCY DRIVE WITH THE DISPOSABLE HLS RETAINER" IS STATED THAT THE EMERGENCY DRIVE CAN BE USED TO MANUALLY CONTROL THE BLOOD FLOW IN CASE OF A FAILED CARDIOHELP. IN THE INSTRUCTION FOR USE (IFU) IS STATED THAT IT IS NECESSARY TO DECREASE THE FLOW TO ZERO BEFORE DISCONNECTING AND CONNECTING THE DISPOSABLE TO THE DEVICE. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2025-06-24 FOR THE PERIOD OF 2014-09-22 TO 2025-04-22. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2014-09-22. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE CARDIOHELP. BASED ON THE RESULTS THE REPORTED FAILURE "MECHANICAL NOISE ON CARDIOHELP" COULD NOT BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: NEW INFORMATION WAS RECEIVED ON 2025-05-27, THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000193. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THE SERIAL NUMBER OF THE CARDIOHELP WAS NOT PROVIDED. THUS, NO UDI NUMBER CAN BE PROVIDED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT# (B)(4). THE INVOLVED CARDIOHELP WILL INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21918652&pc=OZD | 1220648-2025-27895 | 2025-04-05 04:00 | Death | ABIOMED, INC. | 2025-04-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD BLEEDING AT THE ACCESS SITE. ANGLE MATCHING WAS PERFORMED, A FEMSTOP WAS PLACED, THREE UNITS OF PACKED RED BLOOD CELLS WERE GIVEN AND HEPARIN WAS WITHHELD. ADDITIONALLY, THE PATIENT HAD LIMB ISCHEMIA. STAFF WAS UNABLE TO DOPPLER PULSES IN THE BILATERAL LOWER EXTREMITIES AND BOTH LEGS WERE COLD AND MOTTLED. SUPPORT CONTINUED. THE PATIENT EXPIRED ON SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21918721&pc=DSQ | 2916596-2025-02670 | 2025-04-11 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN URGENT LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION DUE TO ACUTE IMPALLA 5.5 THROMBOSIS IN INTERMACS 1 AS A LAST TREATMENT OPTION. THE PATIENT PASSED AWAY. THE PUMP WAS WORKING AS EXPECTED AND NO PUMP RELATED ISSUES WERE REPORTED. THE DEVICE WOULD NOT RETURN. THE CAUSE OF DEATH WAS LONG LASTING MALPERFUSION WHICH WAS DUE TO THE IMPELLA 5.5 THROMBOSIS. THE PUMP WAS WORKING AS EXPECTED.Event Description: THE CAUSE OF DEATH WAS MULTI-SYSTEM ORGAN FAILURE CAUSED BY LONG LASTING MALPERFUSION WHICH WAS DUE TO THE IMPELLA 5.5 THROMBOSIS. THE PUMP WAS WORKING AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR ANALYSIS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS MULTIPLE ORGAN FAILURES/DYSFUNCTIONS AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919033&pc=NGV | 2124215-2025-26320 | 2025-04-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-28 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | Abdominal Pain; Pain | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. A 27 MM WATCHMAN FLX PRO LAA CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2025. THE POSITION OF THE CLOSURE DEVICE WAS GREAT WITH LESS THAN 1/3 SHOULDER AT THE 135-DEGREE VIEW. A TUG TEST WAS PERFORMED, AS WELL AS THE CLOSURE DEVICE WAS DOING AN AUTO TUG WITH THE HEART. THE COMPRESSIONS WERE 16%-20% AND THERE WAS ZERO FLOW AROUND THE DEVICE. POSITION, ANCHOR, SIZE AND SEAL (PASS) CRITERIA WERE MET, AND THE DEVICE WAS RELEASED. AFTER THE PROCEDURE THE PATIENT COMPLAINT OF SERIOUS BACK PAIN. THE PHYSICIAN CHECKED THE PATIENT AND EVERYTHING SEEMED OKAY. OVER THE WEEKEND, THE CLOSURE DEVICE WAS NOTED TO HAVE EMBOLIZED INTO THE PATIENT'S DESCENDING AORTA. THE DEVICE WAS PERCUTANEOUSLY REMOVED WITH A SNARE. THE PATIENT REMAINS INTUBATED IN THE HOSPITAL WITH COMPLICATIONS, BUT THE PATIENT IS EXPECTED TO FULLY RECOVER.Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2025. THE POSITION OF THE CLOSURE DEVICE WAS GREAT WITH LESS THAN 1/3 SHOULDER AT THE 135-DEGREE VIEW. A TUG TEST WAS PERFORMED, AS WELL AS THE CLOSURE DEVICE WAS DOING AN AUTO TUG WITH THE HEART. THE COMPRESSIONS WERE 16%-20% AND THERE WAS ZERO FLOW AROUND THE DEVICE. POSITION, ANCHOR, SIZE AND SEAL (PASS) CRITERIA WERE MET, AND THE DEVICE WAS RELEASED. AFTER THE PROCEDURE THE PATIENT COMPLAINT OF SERIOUS BACK PAIN. THE PHYSICIAN CHECKED THE PATIENT AND EVERYTHING SEEMED OKAY. OVER THE WEEKEND, THE CLOSURE DEVICE WAS NOTED TO HAVE EMBOLIZED INTO THE PATIENT'S DESCENDING AORTA. THE DEVICE WAS PERCUTANEOUSLY REMOVED WITH A SNARE. THE PATIENT REMAINS INTUBATED IN THE HOSPITAL WITH COMPLICATIONS, BUT THE PATIENT IS EXPECTED TO FULLY RECOVER. IT WAS FURTHER REPORTED THAT THE PATIENT WAS RECOVERING AND WAS ABOUT TO BE MOVED TO THE INTENSIVE CARE UNIT (ICU). THE PATIENT STARTED TO COMPLAIN OF ABDOMINAL PAIN, HAD AN EMERGENT SURGERY FOR ISCHEMIC BOWEL AND THE PATIENT PASSED AWAY AND NEVER RECOVERED. Manufacturer Narrative: B3: DATE OF EVENT IS AN APPROXIMATE DATE AS ACTUAL DATE OF EMBOLIZATION IS NOT KNOWN. D6B: EXPLANT DATE IS AN APPROXIMATE DATE AS THE ACTUAL DATE OF EXPLANT IS NOT KNOWN. Manufacturer Narrative: B2: DATE OF DEATH IS AN APPROXIMATE DATE AS THE ACTUAL DEATH DATE IS NOT KNOWN. B3: DATE OF EVENT IS AN APPROXIMATE DATE AS ACTUAL DATE OF EMBOLIZATION IS NOT KNOWN. D6B: EXPLANT DATE IS AN APPROXIMATE DATE AS THE ACTUAL DATE OF EXPLANT IS NOT KNOWN. B5: UPDATED. H6: IMPACT CODE ADDED. H6: PATIENT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919208&pc=LWS | 2649622-2025-11438 | 2025-04-05 04:00 | Death | MPRI | 2025-04-28 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Failure to Capture; Electrical /Electronic Property Problem; Over-Sensing; Under-Sensing | Arrhythmia | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD AND RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF CAPTURE. IT WAS ALSO REPORTED THAT THE RV LEAD EXHIBITED AN INTERMITTENT PATTERN OF VENTRICULAR SHORT-LONG SENSING NOTED ON SOME STORED EPISODES WHICH CAN BE EXPLAINED AS T-WAVE OVERSENSING (TWOS) OR R-WAVE DOUBLE COUNTING. ADDITIONALLY, IT WAS NOTED THAT THE RV LEAD APPEARED TO BE UNDERSENSING TORSADES FOR AN UNKNOWN DURATION WHERE THERAPIES WERE NOT DELIVERED. BOTH LEADS REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS FOUND DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED T-WAVE OVERSENSING. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS ELEVATED. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS RISING. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919366&pc=DSQ | 2916596-2025-02392 | 2024-10-31 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE AND THAT THEY WERE A HIGH-RISK SURGICAL CANDIDATE. THE PATIENT HAD COMPLICATIONS POST OPERATION LEADING TO THE MULTISYSTEM ORGAN FAILURE. DEATH WAS NOT CONSIDERED TO VENTRICULAR ASSIST DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND OPERATION MANUAL CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919604&pc=NPT | 2135147-2025-02218 | 2018-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-04-28 04:00 | NPT | PORTICO TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "PROPENSITY-MATCHED ANALYSIS OF EARLY AND LONG-TERM CLINICAL OUTCOMES WITH SELF-EXPANDABLE PROSTHESES IN TAVR: PORTICO VS. COREVALVE EVOLUT R", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY THAT DIRECTLY COMPARES THE SELF-EXPANDABLE (SE) VALVE DEVICES PORTICO AND COREVALVE EVOLUT R IN TERMS OF CLINICAL OUTCOMES AND EFFICACY AT A 30-DAY AND 3-YEAR FOLLOW-UP. DEVICES INCLUDED IN THE STUDY WERE PORTICO AND COREVALVE. THE ARTICLE CONCLUDED THAT IN TERMS OF EARLY CLINICAL OUTCOMES, NO STATISTICALLY SIGNIFICANT DIFFERENCES WERE OBSERVED BETWEEN THE TWO GROUPS OF SELF-EXPANDABLE VALVE PROSTHESES, EXCEPT FOR A SIGNIFICANTLY HIGHER RATE OF MINOR BLEEDING IN THE EVOLUT GROUP AT 30 DAYS. NOTABLY, THIS TREND OF INCREASED MINOR BLEEDING IN THE EVOLUT GROUP PERSISTED OVER THE 3-YEAR FOLLOW-UP PERIOD, ALTHOUGH THE DIFFERENCE DID NOT REACH STATISTICAL SIGNIFICANCE. BOTH GROUPS DEMONSTRATED LOW RATES OF ALL-CAUSE MORTALITY AND CLINICAL COMPLICATIONS AT LONG-TERM FOLLOW-UP. THE CHOICE OF VALVE SHOULD BE CUSTOMIZED TO THE INDIVIDUAL CHARACTERISTICS OF EACH PATIENT. [THE PRIMARY AND CORRESPONDING AUTHOR WAS UWE PRIMESSNIG, DEUTSCHES HERZZENTRUM DER CHARITÉ, DEPARTMENT OF CARDIOLOGY, ANGIOLOGY AND INTENSIVE CARE MEDICINE, CAMPUS VIRCHOW-KLINIKUM, 13353 BERLIN, GERMANY, WITH CORRESPONDING EMAIL: UWE.PRIMESSNIG@DHZC-CHARITE.DE] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2018 TO DECEMBER 2021. A TOTAL OF 396 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 312 (78.8%) RECEIVED AN ABBOTT DEVICE. OUT OF THE TOTAL COHORT, 79 PATIENTS WERE ASSIGNED TO EACH GROUP AFTER 1:1 PROPENSITY SCORE MATCHING BASED ON BASELINE PARAMETERS. IN THE PORTICO GROUP, THE AVERAGE AGE WAS 80.7 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE COREVALVE GROUP, THE AVERAGE AGE WAS 81.0 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED OBESITY, ARTERIAL HYPERTENSION, DIABETES MELLITUS, PRIOR CORONARY HEART DISEASE, ATRIAL FIBRILLATION, RENAL INSUFFICIENCY, HEART BLOCK, AORTIC VALVE STENOSIS. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: PROPENSITY-MATCHED ANALYSIS OF EARLY AND LONG-TERM CLINICAL OUTCOMES WITH SELF-EXPANDABLE PROSTHESES IN TAVR: PORTICO VS. COREVALVE EVOLUT R. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF PORTICO VALVES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING OBESITY, ARTERIAL HYPERTENSION, DIABETES MELLITUS, PRIOR CORONARY HEART DISEASE, ATRIAL FIBRILLATION, RENAL INSUFFICIENCY, HEART BLOCK, AORTIC VALVE STENOSIS. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), UNEXPECTED MEDICAL INTERVENTION (POST-DILATATION, RESUSCITATION), HOSPITALIZATION, RENAL FAILURE, HEART BLOCK, CARDIAC DECOMPENSATION (HEART FAILURE), BLEEDING, ATRIAL FIBRILLATION, PARAVALVULAR LEAK; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: PROPENSITY-MATCHED ANALYSIS OF EARLY AND LONG-TERM CLINICAL OUTCOMES WITH SELF-EXPANDABLE PROSTHESES IN TAVR: PORTICO VS. COREVALVE EVOLUT R.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919668&pc=NGV | 2135147-2025-02221 | 2025-03-17 04:00 | Death | ABBOTT MEDICAL | 2025-04-28 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Angina; Arrhythmia; Cardiac Arrest; Perforation; Pneumonia | P200049 |  | 1 | Event Description: IT WAS REPORTED ON 14 MARCH 2025, A 20MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT. THE OCCLUDER WAS IMPLANTED. ON (B)(6) 2025 EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED FOR THE PATIENT'S CHEST PAIN. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL. THE PATIENT PRESENTED WITH PNEUMONIA AND WAS COUGHING AND VOMITING. THE PATIENT WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT PASSED AWAY ON THE SAME DAY. AN AUTOPSY SHOWED A SMALL PUNCTURE IN THE PULMONARY ARTERY. THE OFFICIAL CAUSE OF DEATH WAS THE CARDIAC ARREST. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF CARDIAC ARREST AND DEATH WAS REPORTED. A RETURNED DEVICE INSPECTION, TO RULE OUT ANY DEVICE-RELATED CAUSES, COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE OFFICIAL CAUSE OF DEATH WAS THE CARDIAC ARREST. THE CAUSE OF CARDIAC ARREST COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES WHERE CPR/RESUSCITATION WAS PERFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. FROM MEDICAL REVIEW: THE PATIENT UNDERWENT CPR WHICH MAY HAVE CAUSED THE DEVICE TO BE PUSHED INTO THE PULMONARY ARTERY AND CAUSE A PUNCTURE IN THE PULMONARY ARTERY. HOWEVER, IT CANNOT BE DETERMINED FOR CERTAIN IF THE PULMONARY ARTERY INJURY MAY OCCURRED SPONTANEOUSLY PRIOR TO THE CPR. NOT DEVICE RELATED. UNCERTAIN IF PROCEDURE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919741&pc=NPT | 2025587-2025-03052 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 22 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE INFECTIONS. THE AVERAGE TIME TO EVENT WAS TWO HUNDRED NINETEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 56 TO 90 AND OVER. THERE WERE 14 MALES AND 8 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 22 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919785&pc=NPT | 2025587-2025-03053 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 4 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE INFECTIONS. THE AVERAGE TIME TO EVENT WAS TWENTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 79 TO 85. THERE WERE 4 MALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919804&pc=NPT | 2025587-2025-03054 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS INFECTION. THE TIME TO EVENT WAS ONE HUNDRED FIFTY-SEVEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 90 OR OVER. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919862&pc=NPT | 2025587-2025-03055 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 10 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE RENAL. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED SIXTY DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 51 TO 90 AND OVER. THERE WERE 4 MALES AND 6 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 10 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919864&pc=NPT | 2025587-2025-03056 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 3 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE RENAL. THE AVERAGE TIME TO EVENT WAS TWENTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 69 TO 79. THERE WAS 1 MALE AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 3 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919882&pc=NPT | 2025587-2025-03057 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 14 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE NEUROLOGIC. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED THIRTY-FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM (B)(6). THERE WERE 6 MALES AND 8 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 14 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919886&pc=NPT | 2025587-2025-03058 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS NEUROLOGIC. THE TIME TO EVENT WAS TWELVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 86. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919898&pc=NPT | 2025587-2025-03059 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 47 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED SEVENTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM (B)(6). THERE WERE 27 MALES AND 20 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 47 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919901&pc=NPT | 2025587-2025-03060 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 4 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS THIRTY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 77 TO 81. THERE WERE 4 MALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919902&pc=FZP | 9612501-2025-01160 | 2025-04-10 04:00 | Death | DAVIS & GECK CARIBE LTD | 2025-04-28 04:00 | FZP | ENDO CLIP III | Mechanics Altered | Laceration(s); Peritonitis | K100242 |  | 1 | Event Description: IT WAS REPORTED THAT ON A POST VETERINARY LAPAROSCOPIC GALLBLADDER SURGERY; WHILE CLOSING THE BILE DUCT, THE CLIP APPLIER DID NOT LOAD THE CLIPS PROPERLY. THE PRODUCT WAS EXPLANTED. IT WAS ALSO NOTED THAT THE HOSPITALIZATION AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. ALSO, THE INCISION EXTENDED DUE TO THE PRODUCT PROBLEM. TO RESOLVE THE ISSUE, EMERGENCY REINTERVENTION WAS DONE. IT WAS REPORTED THAT THE PATIENT EXPIRED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G3, H3, H6 H3 EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS AVAILABLE FOR EVALUATION. THE EVALUATION FOUND NO POTENTIALLY CONTRIBUTING FACTORS, AND THE SAMPLE MET ALL RELATED SPECIFICATIONS. IT WAS REPORTED THAT THE CLIPS DID NOT LOAD PROPERLY. THE REPORTED ISSUE COULD NOT BE CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE IDENTIFIED BECAUSE NO RELATED PROBLEM WAS DETECTED WITH THE DEVICE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919904&pc=NPT | 2025587-2025-03061 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 5 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS EIGHTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 79 TO 89. THERE WERE 4 MALES AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 5 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919905&pc=DQY | 2135147-2025-02223 | 2025-04-03 04:00 | Death | ABBOTT MEDICAL | 2025-04-28 04:00 | DQY | AMPLATZER TORQVUE DELIVERY SYSTEM | Use of Device Problem | Cardiac Arrest; Hemorrhage/Bleeding; Cardiac Tamponade; Pericardial Effusion; Unspecified Tissue Injury | K072313 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, AN UNKNOWN SIZE OCCLUDER WAS CHOSEN FOR A POSTINFARCT VENTRICULAR SEPTAL DEFECT (PIVSD) CLOSURE PROCEDURE USING A 10F AMPLATZER TORQVUE DELIVERY SYSTEM. DURING PROCEDURE, THE ACTIVATED CLOTTING LEVELS WERE >250S AND HEPARIN WAS ADMINISTERED. WHEN THE DELIVERY SHEATH WAS BEING ADVANCED THROUGH THE DEFECT, A LARGE PERICARDIAL EFFUSION WAS SEEN AND IMMEDIATE LOSS OF OUTPUT. AN EMERGENCY PERICARDIAL DRAIN WAS INSERTED THROUGH A SUB XIPHOID APPROACH (FIRST PASS) AND CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED. A LARGE VOLUME OF BLOOD WAS ASPIRATED BUT THERE WAS NO IMPROVEMENT IN THE HEMODYNAMICS AND NO REDUCTION IN THE PERICARDIAL EFFUSION WAS NOTED. THERE WAS CONTINUED AND ONGOING BLEEDING AND NO WAY OF TREATING THAT. THEY MENTIONED IT WAS DUE TO MYOCARDIAL RUPTURE AND NO MECHANICAL PROBLEM THAT THEY COULD FIX, THE WHOLE TEAM AGREED THAT IT WAS APPROPRIATE TO STOP CPR. THE PATIENT WAS REPORTED PASSED AWAY. THE OCCLUDER WAS NOT OPENED AND USED. THE PHYSICIAN THOUGHT THE CAUSE OF DEATH WAS VENTRICULAR RUPTURE DURING A PERCUTANEOUS ATTEMPT TO CLOSE A POST INFARCT VENTRICULAR SEPTAL RUPTURE (VSD).Event Description: N/A Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF USE OF DEVICE PROBLEM WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION, THE REPORTED USE OF DEVICE RELATED TISSUE DAMAGED COULD NOT BE DETERMINED. THE CAUSE OF DEATH WAS VENTRICULAR RUPTURE DURING A PERCUTANEOUS ATTEMPT TO CLOSE A POST INFARCT VENTRICULAR SEPTAL RUPTURE (VSD). THE PATIENT EFFECTS OF PERICARDIAL EFFUSION AND CARDIAC TAMPONADE COULD NOT BE CONCLUSIVELY DETERMINED. BLEEDING APPEARS TO RELATE TO PERICARDIAL EFFUSION/TAMPONADE, UNABLE TO TREAT, AND SUBSEQUENT CARDIOGENIC SHOCK (CARDIAC ARREST). THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: NA. THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919911&pc=NPT | 2135147-2025-02222 | 2025-04-03 04:00 | Death | ABBOTT MEDICAL | 2025-04-28 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Difficult to Fold, Unfold or Collapse; Migration or Expulsion of Device; Improper or Incorrect Procedure or Method | Low Blood Pressure/ Hypotension | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS CHOSEN FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE PATIENT HAD A CALCIUM NODULE IN THE ANNULUS/LVOT (LEFT VENTRICULAR OUTFLOW TRACT) AND THEY DECIDED TO PERFORM PRE-IMPLANTATION BALLOON AORTIC VALVULOPLASTY (PRE-BAV). THE RECOMMENDATION WAS A 22MM BALLOON BUT, THE PHYSICIAN WANTED TO USE A 20MM BALLOON. DURING PROCEDURE, THE PHYSICIAN DECIDED NOT DO A PRE-BAV BALLOON AT ALL AND CONTINUED TO IMPLANT THE VALVE IN CONTROLLED PACING OF 130 BPM USING THE WIRE AS A PACEMAKER. AFTER MAKING 80% MARK, THEY EVALUATED THE VALVE IN CUSP OVERLAP AND DECIDED THEY WERE AT ROUGHLY 3/4MM OF DEPTH. THEN THEY SWITCHED OVER TO 3-CUSP VIEW WITH THE PARALLAX REMOVED, THE VALVE WAS AT 4MM DEPTH ON THE LEFT CORONARY CUSP (LCC). THERE WAS A INCOMPLETE STENT OPENING ON THE NON-CORONARY CUSP (NCC) SIDE AND THE INDENTATION ON THE LEFT SIDE, ABBOTT REPRESENTATIVE RECOMMENDED TO PERFORM A COMPLETE RECAPTURE OF THE VALVE BUT, THE PHYSICIAN DECIDED TO RELEASE THE VALVE AND PLANNED TO PERFORM A POST BALLOON AORTIC VALVULOPLASTY. THE VALVE WAS THEN FULLY DEPLOYED WITH ALL TABS RELEASING AND VISUALIZED. THE FINAL IMPLANT DEPTH WAS 3MM ON THE NCC AND 4MM ON THE LCC. THE DELIVERY SYSTEM WAS THEN PULLED BACK AND RECAPTURED/ REMOVED FROM THE PATIENT. AFTER ABOUT 5 MINUTES (MIN) OF DEPLOYMENT, THE PHYSICIAN AND STAFF MENTIONED THAT THE VALVE HAD MIGRATED UP ABOVE THE ANNULUS. AFTER ABOUT ANOTHER 5 MIN, THE BOTTOM OF THE STENT FRAME OF THE VALVE WAS HALFWAY BETWEEN THE ANNULUS AND THE SINO-TUBULAR JUNCTION (STJ). THE PATIENT¿S PRESSURE STARTED TO DROP RAPIDLY AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. THEY DECIDED TO SNARE THE VALVE AND AN IMPELLA WAS PLACED INTO THE LEFT VENTRICLE (LV) TO TRY AND GAIN HEMODYNAMIC STABILITY. BUT THE PRESSURES REMAINED LOW. THE PHYSICIAN INSERTED 1 SNARE THROUGH THE GROIN AND GRASPED THE TAP OF THE VALVE BUT, UNABLE TO PULL THE VALVE UP ANY FURTHER SO THEY INSERTED A 2ND SNARE IN THE GROIN. THE PHYSICIAN WAS ABLE TO GRAB 2 DIFFERENT TABS ON TO OF THE STENT FRAME AND PULL THE VALVE UP. AFTER THE VALVE CAME UP, IT LANDED AT THE TOP OF THE ARCH NEAR THE GREAT VESSELS AND THEY ELECTED TO PULL THE VALVE TO THE DESCENDING AORTA TO KEEP FROM BLOCKING THE CAROTID ARTERIES. A NEW 29MM NAVITOR VALVE AND LARGE FLEXNAV DELIVERY SYSTEM WAS PREPARED. DURING PREPPING THE VALVE, THE PHYSICIAN ADVANCED A 22MM NON-ABBOTT BALLOON AND PREPPED THE ANNULUS PROPERLY. THE NEW VALVE/DELIVERY SYSTEM WAS THEN ADVANCED TO THE ANNULUS. THE IMPELLA WAS REMOVED PRIOR TO DEPLOYMENT OF THE VALVE. THE VALVE WAS THEN PLACED AT 3MM ON THE NCC (CHECKED IN CUSP OVERLAP) AND 3MM DEPTH ON THE LCC (CHECKED IN COPLANAR WITH PARALLAX REMOVED. THE VALVE WAS FULLY DEPLOYED AND THE DEPTH REMAINED THE SAME ON BOTH SIDES. THE IMPELLA WAS THEN READVANCED THROUGH THE NAVITOR VALVE TO TRY AND REGAIN HEMODYNAMIC STABILITY. THE PATIENT WAS THEN TRANSPORTED TO CARDIAC CARE UNIT (CCU) WITH THE IMPELLA IN PLACE. THE PATIENT WAS REPORTED PASSED AWAY PASSED AWAY A FEW HOURS LATER.Event Description: N/A. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT A DURING PARTIAL DEPLOYMENT OF 29 MM NAVITOR VALVE, INCOMPLETE STENT EXPANSION WAS NOTED. REPORTED FINAL IMPLANT DEPTH MEASUREMENTS WAS 4 MM LCC AND 3 MM NCC. UPON FULL RELEASE, THE VALVE WAS POSITIONED SUPRAMAMMILLARY, AND MIGRATION WAS OBSERVED. AT THIS TIME THE PATIENT BECAME HEMODYNAMIC INSTABILITY. TWO SNARES WERE USED TO RETRIEVE THE MIGRATED VALVE INTO THE DESCENDING AORTA TO KEEP FROM BLOCKING THE CAROTID ARTERIES. A SECOND 29 MM NAVITOR VALVE WAS THEN IMPLANTED AT 3 MM DEPTH AND REMAINED STABLE ACROSS THE ANNULUS. THE IMPELLA THAT WAS PLACED INTO THE LEFT VENTRICLE TO GAIN HEMODYNAMIC STABILITY, HAD BEEN WITHDRAWN PRIOR TO THE SECOND VALVE DEPLOYMENT, WAS REINTRODUCED INTO THE LEFT VENTRICLE AFTER IMPLANTATION OF THE SECOND VALVE. THE PATIENT WAS THEN TRANSFERRED TO CARDIAC CARE UNIT (CCU). REPORTEDLY, THE PATIENT PASSED AWAY A FEW HOURS LATER. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WILL NOT BE RETURNED FOR ANALYSIS. PROCEDURAL IMAGING WAS REVIEWED: INITIAL IMPLANT WAS DEEP WITH INCOMPLETE STENT OPENING. NO CLEAR EVIDENCE OF CORONARY OBSTRUCTION WAS NOTED. THE SECOND IMPLANT WAS ALSO DEEP, HOWEVER, NO SEVERE PVL WAS NOTED. BASED ON THE INFORMATION RECEIVED AND MEDICAL REVIEW PERFORMED AT ABBOTT, THE EXACT CAUSE OF DEATH COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, THE INITIAL VALVE MIGRATION LIKELY CONTRIBUTED TO THE HEMODYNAMIC INSTABILITY AND SUBSEQUENT CLINICAL DETERIORATION. THE INCOMPLETE STENT OPENING MAY HAVE CONTRIBUTED TO THE VALVE MIGRATION BUT THIS CANNOT BE CONFIRMED. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE THE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES, AS CPR WAS PERFORMED FOR DROP IN BLOOD PRESSURE AND THE VALVE FOR SNARED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. PLEASE NOTE THAT PER THE INSTRUCTIONS FOR USE, "ONCE THE VALVE IS FULLY DEPLOYED, REPOSITIONING AND RETRIEVAL OF THE VALVE IS NOT POSSIBLE. ATTEMPTED RETRIEVAL (E.G., USE OF A GUIDEWIRE, SNARE, OR FORCEPS) MAY CAUSE AORTIC ROOT, CORONARY ARTERY, AND/OR MYOCARDIAL DAMAGE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919923&pc=OZD | 1220648-2025-27902 | 2025-04-10 04:00 | Death | ABIOMED, INC. | 2025-04-28 04:00 | OZD | IMPELLA CP-JP | Adverse Event Without Identified Device or Use Problem | Hypovolemic Shock; Multiple Organ Failure | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AN IMPELLA PATIENT'S ARTERY WAS DAMAGED DURING CENTRAL VENOUS INSERTION. A BLOOD TRANSFUSION WAS PERFORMED, BUT THE BLOOD LOSS WAS SO GREAT THAT CIRCULATION COULD NOT BE MAINTAINED. THE PATIENT SUBSEQUENTLY PASSED AWAY FROM MULTIPLE ORGAN FAILURE DUE TO HEMORRHAGIC SHOCK. DUE TO ANTICOAGULATION IN THE PATIENT'S SYSTEM FROM IMPELLA SUPPORT, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919928&pc=NPT | 2025587-2025-03062 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 66 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FORTY-SIX DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 62 TO 90 AND OVER. THERE WERE 33 MALES AND 33 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 66 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919932&pc=NPT | 2025587-2025-03063 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 10 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 61 TO 88. THERE WERE 5 MALES AND 5 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 10 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919935&pc=NPT | 2025587-2025-03064 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 1 | Event Description: THE TYPES OF DEATH BEING REPORTED IS DEATH DUE TO OTHER CARDIAC REASON. THE TIME TO EVENT WAS ONE HUNDRED NINETY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 87. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919937&pc=NPT | 2025587-2025-03065 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 4 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FIFTY-SIX DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 72 TO 84. THERE WERE 3 MALES AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919954&pc=NPT | 2025587-2025-03066 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 4 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE LISTED AS UNKNOWN. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED SIXTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 81 TO 84. THERE WERE 2 MALES AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919956&pc=NPT | 2025587-2025-03068 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS LISTED AS UNKNOWN. THE TIME TO EVENT WAS THIRTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 77. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919962&pc=MCX | 2124215-2025-26898 | 2025-04-04 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-28 04:00 | MCX | ROTAPRO? | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiogenic Shock | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH TRIPLE VESSEL DISEASE (TVD). THE TARGET LESION WAS LOCATED IN THE PROXIMAL TO MIDDLE LEFT CIRCUMFLEX ARTERY (LCX). A 1.75 MM ROTAPRO WAS SELECTED FOR PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY (PTCA) WITH ROTABLATION. DURING THE PROCEDURE, THE PATIENT DEVELOPED CARDIOGENIC SHOCK, AND THE PATIENT CONDITION PROGRESSIVELY WORSENED. DESPITE MULTIPLE ATTEMPTS AT CARDIOPULMONARY RESUSCITATION (CPR) AND SHOCKS FROM AN AUTOMATED EXTERNAL DEFIBRILLATOR (AED), THE PATIENT'S CONDITION FAILED TO IMPROVE. THE PATIENT WAS SUBSEQUENTLY PLACED ON A VENTILATOR, AND THE PROCEDURE WAS DISCONTINUED DUE TO THE PATIENT'S CONDITION WORSENED. UNFORTUNATELY, THE PATIENT PASSED AWAY. THE OFFICIAL CAUSE OF DEATH IS CARDIAC ARREST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919966&pc=NPT | 2025587-2025-03069 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS LISTED AS UNKNOWN. THE TIME TO EVENT WAS THREE HUNDRED TWENTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 81. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919971&pc=NPT | 2025587-2025-03070 | 2025-04-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS DCS | Failure to Advance; Patient Device Interaction Problem | Extravasation; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER BIPROSTHETIC VALVE, INTO A PATIENT WITH CALCIFIED ARTERIES THAT WERE NOT COMPATIBLE WITH A 14F DELIVERY SYSTEM, THE PROCEDURE WAS PERFORMED THROUGH THE RIGHT FEMORAL ARTERY INSTEAD OF AN ALTERNATIVE ACCESS. DURING THE INTRODUCTION OF THE 14F MEDTRONIC SHEATH IN THE RIGHT FEMORAL ARTERY, THERE WAS DIFFICULTY. ANGIOPLASTY WAS PERFORMED IN THE COMMON ILIAC ARTERY AND THE RIGHT EXTERNAL ILIAC ARTERY WITH A 7X40 BALLOON. THE VALVE WAS LOADED AND A PRE-DILATION WAS PERFORMED WITH A NON-MEDTRONIC 18X40 BALLOON. THE SHEATH WAS REMOVED. THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO PROGRESS THROUGH THE RIGHT FEMORAL ARTERY THEREFORE A NEW ANGIOPLASTY WAS PERFORMED WITH A 7X120 BALLOON IN THE COMMON FEMORAL ARTERY AND THE EXTERNAL ILIAC ARTERY. AFTER IMAGE CONTROL, BLOOD EXTRAVASATION IN THE EXTERNAL ILIAC AND FEMORAL ARTERIES, LEADING TO HYPOTENSION. A 7X58 COVERED STENT WAS PLACED IN THE EXTERNAL ILIAC ARTERY. IMAGE CONTROL WAS PERFORMED AND EXTRAVASATION WAS NOTED IN THE COMMON FEMORAL ARTERY, NECESSITATING OPEN BYPASS SURGERY WITH GRAFT. THE PATIENT RECEIVED 5 RED BLOOD CELL CONCENTRATES, OROTRACHEAL INTUBATION, AND NOREPINEPHRINE IN A CONTINUOUS PUMP. IT WAS DETERMINED TO NOT IMPLANT THE VALVE DUE TO THE VASCULAR SEVERITY. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT, WHERE THEY EXPERIENCED TWO CARDIORESPIRATORY ARRESTS AND SUBSEQUENTLY DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DCS AND SHEATH DID CONTRIBUTE TO THE VASCULAR INJURY. THE PATIENT HAD ARTERIES WITH DIA METERS SMALLER THAN RECOMMENDED, DUE TO CALCIFICATION; HOWEVER, THE DECISION WAS MADE TO PERFORM THE PROCEDURE VIA THE RIGHT FEMORAL ROUTE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVPROPLUS-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: L-EVPROP23-29 (LOT: 0012084976); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: GWBC30 (LOT: 92106993); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE: N/A; EXPLANT DATE: N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919981&pc=NPT | 2025587-2025-03071 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 7 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE TRAUMA. THE AVERAGE TIME TO EVENT WAS TWO HUNDRED AND NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM (B)(6) AND OVER. THERE WERE 2 MALES AND 5 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038. QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 7 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21919994&pc=NPT | 2025587-2025-03072 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Myocardial Infarction | P130021 | E2014038 | 14 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE MYOCARDIAL INFARCTIONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED EIGHTY-EIGHT DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 66 TO 90 AND OVER. THERE WERE 9 MALES AND 5 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 14 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920005&pc=NPT | 2025587-2025-03073 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Myocardial Infarction | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS MYOCARDIAL INFARCTION. THE TIME TO EVENT WAS FOURTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 80. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920036&pc=NPT | 2025587-2025-03074 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 5 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE CARDIOVASCULAR PROCEDURE RELATED. THE AVERAGE TIME TO EVENT WAS THIRTY-FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 67 TO 88. THERE WAS 1 MALE AND 4 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 5 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920043&pc=NPT | 2025587-2025-03075 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE CARDIOVASCULAR PROCEDURE RELATED. THE AVERAGE TIME TO EVENT WAS THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 77 TO 86. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920045&pc=NPT | 2025587-2025-03076 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS CARDIOVASCULAR PROCEDURE RELATED. THE TIME TO EVENT WAS ZERO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 83. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920056&pc=NPT | 2025587-2025-03077 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Hemorrhage/Bleeding | P130021 | E2014038 | 8 | Event Description: THE TYPES OF DEATHS BEING REPORTED WERE HEMORRHAGIC AND CARDIOVASCULAR HEMORRHAGES. THE AVERAGE TIME TO EVENT WAS SEVENTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 70 TO 87. THERE WERE 4 MALES AND 4 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 8 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920067&pc=NPT | 2025587-2025-03078 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Hemorrhage/Bleeding | P130021 | E2014038 | 3 | Event Description: THE TYPES OF DEATHS BEING REPORTED WERE HEMORRHAGIC AND CARDIOVASCULAR HEMORRHAGES. THE AVERAGE TIME TO EVENT WAS TWELVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 76 TO 82. THERE WERE 2 MALES AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 3 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920092&pc=NPT | 2025587-2025-03079 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Hemorrhage/Bleeding | P130021 | E2014038 | 1 | Event Description: THE TYPES OF DEATH BEING REPORTED WAS HEMORRHAGIC. THE TIME TO EVENT WAS ONE HUNDRED FORTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 88. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920140&pc=DSQ | 2916596-2025-02626 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SPONTANEOUS INTRACRANIAL BLEED (ICB). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PATIENT SUFFERED FROM A SPONTANEOUS ICB DURING HAVING A COLD. ICB WAS RELATED TO THE OUTCOME AT LEAST. THE DEVICE WAS RUNNING AS EXPECTED. NO ISSUES PRIOR TO OUTCOME WERE MENTIONED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920141&pc=NPT | 2025587-2025-03081 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 28 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE HEART FAILURE. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FIFTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 61 TO 90 AND OVER. THERE WERE 16 MALES AND 12 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 28 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920145&pc=NPT | 2025587-2025-03082 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS HEART FAILURE. THE TIME TO EVENT WAS FIFTY DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 81. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920154&pc=NPT | 2025587-2025-03083 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS HEART FAILURE. THE TIME TO EVENT WAS FIFTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS (B)(6) OR OVER. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920162&pc=NPT | 2025587-2025-03086 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 3 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE HEART FAILURE. THE AVERAGE TIME TO EVENT WAS SEVENTY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 54 TO 85. THERE WAS 1 MALE AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 3 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920163&pc=NPT | 2025587-2025-03084 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS HEART FAILURE. THE TIME TO EVENT WAS TWO HUNDRED NINETY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 72. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY ((B)(6) REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920175&pc=NGV | 2124215-2025-25775 | 2025-04-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pain | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A PATIENT DEATH. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. PRE-IMAGING THAT WAS PERFORMED NOTED AN EXISTING PERICARDIAL EFFUSION. A WATCHMAN ACCESS SYSTEM (WAS) AND A 24 MM WATCHMAN FLX PRO LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN GAINED GROIN ACCESS AND PERFORMED THE TRANSEPTAL ACCESS WITH NO COMPLICATIONS. THE PHYSICIAN POSITIONED THE WAS IN THE LAA, AND THE EFFUSION WAS CHECKED AND REMAINED UNCHANGED. AN APPENDAGE PICTURE WITH CONTRAST WAS PERFORMED AND A FLASH OF CONTRAST WENT OUTSIDE THE HEART SHADOW. THE EFFUSION WAS CHECKED WITH INTRACARDIAC ECHOCARDIOGRAPHY (ICE) AND WITH TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THE PERICARDIAL EFFUSION CONTINUED TO REMAIN UNCHANGED. THE PATIENTS' VITALS REMAINED STABLE. THE PHYSICIAN WAITED 15 MINUTES, AND EVERYTHING WAS STILL STABLE AND UNCHANGED. THE PATIENTS CT WAS REVIEWED, AND IT DIDN'T INDICATE ANY PATIENT ANOMALY. THEY DECIDED TO CONTINUE THE PROCEDURE ASSUMING IT WAS CONTRAST WASHOUT IN THE VEINS. THE PROCEDURE WAS THEN SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. ONE HOUR POST PROCEDURE THE PATIENT COMPLAINT OF PAIN BETWEEN THEIR SHOULDER BLADES AND THEN THEY PROCEEDED TO GO INTO CARDIAC ARREST. THE PATIENT DID NOT RECOVER FROM THIS EVENT, AND THE PATIENT DIED.Event Description: IT WAS REPORTED THAT THERE WAS A PATIENT DEATH. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. PRE-IMAGING THAT WAS PERFORMED NOTED AN EXISTING PERICARDIAL EFFUSION. A WATCHMAN ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN GAINED GROIN ACCESS AND PERFORMED THE TRANSEPTAL ACCESS WITH NO COMPLICATIONS. THE PHYSICIAN POSITIONED THE WAS IN THE LAA, AND THE EFFUSION WAS CHECKED AND REMAINED UNCHANGED. AN APPENDAGE PICTURE WITH CONTRAST WAS PERFORMED AND A FLASH OF CONTRAST WENT OUTSIDE THE HEART SHADOW. THE EFFUSION WAS CHECKED WITH INTRACARDIAC ECHOCARDIOGRAPHY (ICE) AND WITH TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THE PERICARDIAL EFFUSION CONTINUED TO REMAIN UNCHANGED. THE PATIENTS VITALS REMAINED STABLE. THE PHYSICIAN WAITED 15 MINUTES, AND EVERYTHING WAS STILL STABLE AND UNCHANGED. THE PATIENTS CT WAS REVIEWED, AND IT DIDN'T INDICATE ANY PATIENT ANOMALY. THEY DECIDED TO CONTINUE THE PROCEDURE ASSUMING IT WAS CONTRAST WASHOUT IN THE VEINS. THE PROCEDURE WAS THEN SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. ONE HOUR POST PROCEDURE THE PATIENT COMPLAINT OF PAIN BETWEEN THEIR SHOULDER BLADES AND THEN THEY PROCEEDED TO GO INTO CARDIAC ARREST. THE PATIENT DID NOT RECOVER FROM THIS EVENT, AND THE PATIENT DIED. IT WAS FURTHER REPORTED THAT THE PATIENT WAS ON EIGHT LITERS OF OXYGEN CONSISTENTLY. Manufacturer Narrative: INFORMATION ADDED TO B5. PATIENT CODE ADDED TO H6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920190&pc=MGB | 2024168-2025-04380 | 2018-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-04-28 04:00 | MGB | PERCLOSE PROGLIDE¿ | Insufficient Information | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THIS EVENT IS FROM A RETROSPECTIVE STUDY COMPARING SURGICAL CUTDOWN WITH SINGLE-PERCLOSE PROGLIDE CLOSURE USING PERCUTANEOUS CANNULATION IN MINIMALLY INVASIVE CARDIAC SURGERY OF ATRIOVENTRICULAR VALVES. COMPLICATIONS NOTED FOR THE PROGLIDE WERE BLEEDING, VASCULAR OCCLUSION DUE TO THROMBOSIS, LYMPH FISTULA, STROKE/TRANSIENT ISCHEMIC ATTACK, AND IN-HOSPITAL MORTALITY. INTERVENTION FOR BLEEDING CONSISTED OF THE USE OF ANOTHER VASCULAR CLOSURE DEVICE, STITCHING, MANUAL COMPRESSION OR SURGICAL CUT-DOWN. IN THE STUDY GROUP, THE SUCCESS RATE OF SINGLE PROGLIDE STRATEGY WAS SIGNIFICANTLY IMPACTED BY OPERATOR EXPERIENCE. INTRAOPERATIVE BLEEDING, ACUTE VASCULAR COMPLICATIONS AND FEMORAL OBSTRUCTION; HOWEVER, WERE MORE COMMON WITH PROGLIDE USE. INCIDENTS OF ACCESS-STIE COMPLICATIONS WERE SIGNIFICANTLY HIGHER IN THE CUT-DOWN GROUP. THE REDUCTION OF VASCULAR COMPLICATIONS FROM PROGLIDE USE WAS CLEARLY ACHIEVED. THIS STUDY SUGGESTS THE USE OF PROGLIDE IS ASSOCIATED WITH REDUCED RATES OF LATE ACCESS SITE-RELATED WOUND HEALING DISORDERS. PROGLIDE USE WAS NOTED TO BE A SAFE AND FEASIBLE ALTERNATIVE TO FEMORAL CUT-DOWN WITH LOWER INCIDENCE OF ACCESS-SITE COMPLICATIONS FOR MINIMALLY INVASIVE CARDIAC SURGERY. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF THE PRODUCTION RECORD AND COMPLAINT HISTORY OF THE REPORTED LOT COULD NOT BE CONDUCTED BECAUSE THE PART AND LOT NUMBER WERE NOT PROVIDED. THE PATIENT EFFECT OF DEATH IS LISTED IN THE ELECTRONIC PROGLIDE INSTRUCTIONS FOR USE (EIFU), AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF THE DEVICE. BASED ON THE CASE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE LIKELY RELATED TO CASE CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B3, B6: DATES ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND DEVICE ISSUES REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: "FEMORAL VESSEL CANNULATION STRATEGIES IN MINIMALLY INVASIVE CARDIAC SURGERY: SINGLE PERCLOSE PROGLIDE VERSUS OPEN CUT¿DOWN".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920204&pc=NPT | 2025587-2025-03089 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Stroke/CVA | P130021 | E2014038 | 18 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE STROKE. THE AVERAGE TIME TO EVENT WAS EIGHTY DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 75 TO *(B)(6) AND OVER. THERE WERE 7 MALES AND 11 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 18 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920267&pc=NPT | 2025587-2025-03090 | 2025-02-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Stroke/CVA | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE STROKE. THE AVERAGE TIME TO EVENT WAS SIXTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 84 TO 86. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q1 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920419&pc=NPT | 2025587-2025-03106 | 2024-11-28 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Disseminated Intravascular Coagulation (DIC); Hepatitis; Renal Failure; Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S CARDIAC FUNCTION GRADUALLY IMPROVED. DESPITE IMPROVEMENT OF CARDIAC FUNCTION, INCREASED TOTAL BILIRUBIN, WORSENED LIVER FAILURE, SEPSIS, AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC) WERE OBSERVED. OF NOTE, THE PATIENT HAD SEVERAL COMORBIDITIES INCLUDING HYPERTENSIVE CARDIOVASCULAR DISEASE, NON-B NON-C HEPATITIS, CIRRHOSIS, AND CHRONIC KIDNEY DISEASE. APPROXIMATELY 19 AFTER THE PROCEDURE, THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. THE CAUSE OF DEATH WAS NOTED TO BE SEPTICEMIA AND DIC.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) COULD BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH. PER THE PHYSICIAN, THE PATIENT DIED DUE TO WORSENING COMORBIDITIES. Manufacturer Narrative: B3: DATE IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920646&pc=DSQ | 2916596-2025-02627 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY, THE CAUSE WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920719&pc=DSQ | 3007042319-2025-00839 | 2025-04-20 04:00 | Death | HEARTWARE, INC. | 2025-04-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Obstruction of Flow; Pumping Problem | Hematuria; Thrombosis/Thrombus; Lactate Dehydrogenase Increased | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH AN OUTFLOW GRAFT (OFG) AND VENTRICULAR ASSIST DEVICE (VAD) EXPERIENCED HIGH WATT ALARMS AT HOME, LEADING TO HOSPITAL ADMISSION. THE PATIENT PRESENTED WITH HEMATURIA, AND A COMPUTED TOMOGRAPHY (CT) SCAN INDICATED A POSSIBLE THROMBUS INSIDE THE OFG. THE PATIENT HAD SEVERAL SUBTHERAPEUTIC INTERNATIONAL NORMALISED RATIO LAB LEVELS OVER THE LAST SIX MONTHS, WITH A RECORDED LEVEL OF 1.6 ON (B)(6) 2025. ELEVATED LACTATE DEHYDROGENASE (LDH) LEVELS WERE ALSO NOTED. THE PATIENT WAS ON AN UNKNOWN DOSE OF WARFARIN, AND COMPLIANCE WITH ANTICOAGULATION THERAPY WAS UNKNOWN. MEDICAL TREATMENT INCLUDED THE ADMINISTRATION OF ANTICOAGULANTS. THE OFG AND VAD REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS NOT A CANDIDATE FOR PUMP EXCHANGE AND THE THROMBUS WAS CONFIRMED. THE VAD WAS TURNED OFF DUE TO PATIENT DEATH. NO ADDITIONAL INFORMATION WILL BE PROVIDED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE ASSOCIATED OUTFLOW GRAFT (LOT 2002672082) WERE NOT RETURNED FOR EVALUATION AS THEY REMAIN IMPLANTED. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED A SLIGHT INCREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS STARTING ON (B)(6) 2025, FOLLOWED BY A SHARP INCREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS STARTING ON (B)(6) 2025, LEADING TO PARAMETERS ABOVE NORMAL OPERATING RANGE. ADDITIONALLY, 109 HIGH WATT ALARMS WERE LOGGED SINCE (B)(6) 2025. AS A RESULT, THE REPORTED HIGH WATT EVENT WAS CONFIRMED. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS ADMINISTERED ANTICOAGULANTS TO TREAT THROMBUS; HOWEVER, THE PATIENT SUBSEQUENTLY EXPIRED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE AVAILABLE INFORMATION AND HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF THE HIGH POWER EVENT CAN BE ATTRIBUTED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF THROMBUS EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ OUTFLOW GRAFT D4: LOT #: 2002672082. H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - OUTFLOW GRAFT D4: MODEL#: 1125 / CATALOG#: 1125 / EXPIRATION DATE: 28-FEB-2021 / SERIAL OR LOT#: (B)(6) / D9: NO / H3: NO / H4: MFG DATE: 28-FEB-2019 / H5: YES / H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. ADDITIONAL PRODUCT: LOT# 2002672082 H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920804&pc=NPT | 2135147-2025-02236 | 2025-04-03 04:00 | Death | ABBOTT MEDICAL | 2025-04-28 04:00 | NPT | NAVITOR VISION | Perivalvular Leak | Low Blood Pressure/ Hypotension; Cardiac Tamponade | P190023 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025, A NAVITOR VALVE WAS IMPLANTED IN PATIENT. THE IMPLANT DEPTH OF THE VALVE WAS 4.94MM AT THE LCC AND 10MM AT THE NCC. THE SIZING OF THE ANNULAR DIMENSIONS OF THE NATIVE VALVE IS: ANNULUS DIAMETER 28.6MM, AREA 28.3MM, PERIMETER 90MM. THERE WAS SOME PERIVALVULAR LEAKAGE NOTED POST PROCEDURE AND POST DILATION WAS PERFORMED WITH A 28MM NON-ABBOTT DEVICE. AFTER THE DILATION, THE MEAN GRADIENT INCREASED. THEREFORE, THE PHYSICIAN DECIDED TO PERFORM AN ADDITIONAL DILATION WITH A 30MM NON-ABBOTT DEVICE WITH GOOD RESULTS. HOWEVER, AFTER A SHORT WHILE, THE PATIENT PRESSURE DROPPED AND DEVELOPED TAMPONADE. THE PATIENT WAS GIVEN A PERICARD SET AND TRANSFERRED TO THE ICU.
ON (B)(6) 2025, THE PATIENT EXPIRED. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION Manufacturer Narrative: IT WAS REPORTED THAT THERE WAS SOME PERIVALVULAR LEAKAGE NOTED POST PROCEDURE AND POST DILATION WAS PERFORMED; AFTER THE DILATION, THE MEAN GRADIENT INCREASED. THEREFORE, AN ADDITIONAL DILATION WAS PERFORMED WITH GOOD RESULTS. HOWEVER, AFTER A SHORT WHILE, THE PATIENT PRESSURE DROPPED AND DEVELOPED TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED AND PATIENT WAS TRANSFERRED TO THE ICU. THE PATIENT EXPIRED ON THE FOLLOWING DAY OF PROCEDURE. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. NO IMAGING WAS PROVIDED FOR EVALUATION. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED AND MEDICAL REVIEW, THE LIKELY CAUSE OF THE PATIENT DEATH APPEARS TO BE ANNULAR RUPTURE SECONDARY TO POST-DILATION. HOWEVER, THE EXACT CAUSE COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED MEDICAL INTERVENTION IS A CASE-SPECIFIC CIRCUMSTANCE AS POST DILATION WAS PERFORMED FOR PERIVALVULAR LEAKAGE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21920827&pc=DSQ | 2916596-2025-02899 | 2024-12-29 05:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pulmonary Emphysema; Hemorrhage/Bleeding; Pneumothorax; Respiratory Failure; Swelling/ Edema; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: THE PATIENT HAD WELL KNOWN ADVANCED HEART FAILURE WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND WAS DO NOT RESUSCITATE. THE PATIENT HAD TRACHEOSTOMY TUBE REPLACED/REPOSITIONED TODAY AND HAD SUBSEQUENT RESPIRATORY ISSUES AND BLEEDING. ON ARRIVAL, THE PATIENT WAS OBTUNDED, HYPERCAPNIC WITH BLOOD COMING FROM THE TRACHEA. TRACHEOSTOMY TUBE WAS PLACED BY GENERAL SURGEON IN OCTOBER AND THEY CALLED SURGERY TO COME HELP EVALUATE BUT THEIR SATURATION OF PERIPHERAL OXYGEN WAS 99% AND THEY COULD BE BAGGED, BUT THERE WAS SOME DIFFICULTY. BRONCHOSCOPE SHOWED TRACH TUBE IN CORRECT PLACE, AND THERE WAS LOTS OF FROTHY BLOOD COMING UP FROM LUNGS BUT NO SOURCE OF BLEEDING ABOVE. ON ARRIVAL THE PATIENT ALSO HAD IMMENSE SWELLING TO CHEST/HEAD/NECK W CREPITANCE. CHEST X-RAY SHOWED LARGE AMOUNT OF SUBCUTANEOUS EMPHYSEMA AND GOOD SIZED RIGHT PNEUMOTHORAX. PATIENT HAD SOMEWHAT AN ABRUPT CHANGE IN THEIR BLOOD PRESSURE AND FLOW ALARMS ON THEIR LVAD AND HAD CARDIOPULMONARY ARREST. AFTER DISCUSSION WITH THE FAMILY THEY WERE AGREEABLE TO ATTEMPTING TO RELIEVE THIS BY PLACING A CHEST TUBE ALTHOUGH THEY WERE A DO NOT RESUSCITATE. THEY DID HAVE IMPROVEMENT OF LVAD FLOW AND BLOOD PRESSURE AFTER THE CHEST TUBE PLACEMENT BUT THEY SUBSEQUENTLY CONTINUED TO DECLINE. HEART FAILURE AND CARDIOVASCULAR INTENSIVE CARE UNIT WAS CONSULTED. FAMILY EVENTUALLY ELECTED TO HALT INTERVENTIONS AND PROCEED WITH COMFORT CARE GIVEN THEIR SITUATION AND DID NOT WANT ANY FURTHER TREATMENT OR INTERVENTIONS. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. NO DEVICES WOULD RETURN. Manufacturer Narrative: OTHER ANNEX E CODES: METABOLISM AND NUTRITION E12: ELECTRICAL IMBALANCE E1202: BLOOD BICARBONATE INCREASED E120207. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT HAD A TRACHEOSTOMY (TRACH) TUBE REPLACED/REPOSITIONED AND HAD SUBSEQUENT RESPIRATORY ISSUES AND BLEEDING. THE PATIENT WAS NOTED TO HAVE HAD WELL KNOWN ADVANCED HEART FAILURE AND WAS MADE DO NOT RESUSCITATE (DNR). ON ARRIVAL, THE PATIENT WAS OBTUNDED AND HYPERCAPNIC WITH BLOOD COMING FROM THE TRACHEA. A TRACH TUBE HAD BEEN PLACED BY THE GENERAL SURGEON A FEW MONTHS PRIOR. SURGERY WAS CALLED TO COME HELP EVALUATE BUT THE PATIENT'S SATURATION OF PERIPHERAL OXYGEN (SPO2) WAS 99% AND THEY WERE ABLE TO BE BAGGED BUT WITH SOME DIFFICULTY. BRONCHOSCOPY SHOWED THE TRACH TUBE IN THE CORRECT PLACE, AND THERE WAS A LOT OF FROTHY BLOOD COMING UP FROM THE LUNGS BUT NO SOURCE OF BLEEDING ABOVE. ON ARRIVAL THE PATIENT ALSO HAD IMMENSE SWELLING TO THE CHEST/HEAD/NECK WITH CREPITUS. A CHEST X-RAY SHOWED A LARGE AMOUNT OF SUBCUTANEOUS EMPHYSEMA AND A SIZEABLE RIGHT PNEUMOTHORAX. THE PATIENT HAD SOMEWHAT AN ABRUPT CHANGE IN THEIR BLOOD PRESSURE WITH FLOW ALARMS ON THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND HAD CARDIOPULMONARY ARREST. AFTER DISCUSSION WITH THE FAMILY, THEY WERE AGREEABLE TO ATTEMPTING RELIEF BY PLACING A CHEST TUBE, ALTHOUGH THE PATIENT WAS DNR. THEY DID HAVE IMPROVEMENT OF LVAD FLOW AND BLOOD PRESSURE AFTER THE CHEST TUBE PLACEMENT, BUT THEY SUBSEQUENTLY CONTINUED TO DECLINE. HEART FAILURE AND CARDIOVASCULAR INTENSIVE CARE UNIT (CVICU) WAS CONSULTED. THE FAMILY EVENTUALLY ELECTED TO HALT INTERVENTIONS AND PROCEED WITH COMFORT CARE GIVEN THE SITUATION AND DID NOT WANT ANY FURTHER TREATMENT OR INTERVENTIONS. THE PATIENT ULTIMATELY EXPIRED ON (B)(6) 2024. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921104&pc=NIK | 2017865-2025-51364 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | NIK | GALLANT VR | Adverse Event Without Identified Device or Use Problem | Ischemic Heart Disease | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY DUE TO HEART DISEASE, CORONARY ARTERY DISEASE, AND ISCHEMIC CARDIOMYOPATHY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV) LEAD WERE EXPLANTED. IT IS UNKNOWN IF THE ICD OR RV LEAD CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. NO FURTHER INFORMATION WAS RECEIVED. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921106&pc=DSQ | 21921106 | 2024-09-14 04:00 | Death | MEDTRONIC | 2025-04-28 04:00 | DSQ | HEARTMATE 3 | Obstruction of Flow; Inappropriate Waveform; Infusion or Flow Problem; Protective Measures Problem | Hematuria; Thrombosis/Thrombus; Melena |  |  | 1 | Event Description: PATIENT ARRIVED AT THE HOSPITAL ON (B)(6) 2024 WITH LOW FLOW ALARMS AND NOTED DECREASES IN POWER. FLOWS 1-0.7, POWER 2.3. PT TAKES COUMADIN AS ORDERED. LDH BASELINE 150-180. PFHGB BASELINE 10-20. HEPARIN STARTED. INTRAVENOUS FLUIDS GIVEN. PT PULSATILE. THERE IS A SIGNIFICANT CHANGE TO RIGHT HEARTCATH NUMBERS FROM (B)(6) 2023 TO (B)(6) 2024 INDICATING POSSIBLE THROMBUS. ON (B)(6) URINE CHERRY RED. ON (B)(6) 2024 LARGE MELENA. HEPARIN STOPPED. POWER INCREASED GREATER THAN 6. SPEED ADJUSTED DOWN TO 3200 WITH POWER NOW IN 5S AND FLOWS IN 4S. NOT AN EXPLANT RE-IMPLANT CANDIDATE. NO HEMOLYSIS NOTED BY DISCHARGE WITH LDH 328. LOW DOES COUMADIN RESTARTED AT DISCHARGE. PT ADMITTED AGAIN ON (B)(6) 2024 DUE TO CHANGE IN SOUND IN VAD ON (B)(6), AND ONE EPISODE TEA COLORED URINE ON (B)(6). LDH 1078 ON ADMISSION AND PFHGB OF 78. VAD FLOWS 4.3 AND POWER 5.9. VAD SPEED SLOWED TO 3000 WITH IMPROVEMENT IN LDH BY DISCHARGE TO 668. PATIENT ADMITTED (B)(6) 2024 WITH LOW FLOW ALARMS AND ABNORMAL PULSATITY WAVEFORMS. LDH 759; T-BILI 2.1; ALT 127; AST 343. INCREASE OF LDH, T
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921112&pc=NVY | 2017865-2025-51365 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Ischemic Heart Disease | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY DUE TO HEART DISEASE, CORONARY ARTERY DISEASE, AND ISCHEMIC CARDIOMYOPATHY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV) LEAD WERE EXPLANTED. IT IS UNKNOWN IF THE ICD OR RV LEAD CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. NO FURTHER INFORMATION WAS RECEIVED. Manufacturer Narrative: A PARTIAL LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF ANY CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921296&pc=DSQ | 3007042319-2025-00841 | 2025-04-24 04:00 | Death | HEARTWARE, INC. | 2025-04-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) EXPERIENCED MULTISYSTEM ORGAN FAILURE, SEPSIS, BACTEREMIA, AND END STAGE HEART FAILURE, WHICH LED TO THE WITHDRAWAL OF CARE AND SUBSEQUENTLY THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPERIENCED MULTISYSTEM ORGAN FAILURE, SEPSIS, BACTEREMIA, AND END STAGE HEART FAILURE, WHICH LED TO THE WITHDRAWAL OF CARE AND SUBSEQUENTLY THE PATIENT EXPIRED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH, INFECTION, WORSENING HEART FAILURE ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: 419688 LEAD IMPLANT DATE: (B)(6) 2014, 0293 LEAD IMPLANT DATE: (B)(6) 2013, D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921356&pc=NPT | 2025587-2025-03172 | 2025-01-30 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 15 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPSIS. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921432&pc=OJX | 2649622-2025-11461 | 2025-04-12 04:00 | Death | MPRI | 2025-04-28 04:00 | OJX | ATTAIN PERFORMA MRI SURESCAN | Capturing Problem; Positioning Problem | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion; Heart Block; No Clinical Signs, Symptoms or Conditions | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEFT VENTRICULAR (LV) LEAD EXHIBITED INCREASED THRESHOLDS. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED WITH A CAUSE OF DEATH AS PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST. IT WAS NOTED THAT THERE WAS POSSIBLE LEAD MALPOSITION WITH A MODERATE AMOUNT OF PERICARDIAL FLUID. HOWEVER, A REPEAT COMPUTERIZED TOMOGRAPHY (CT) SHOWED ONLY A SMALL PERICARDIAL EFFUSION. IT WAS HYPOTHESIZED THAT THERE WAS A LARGE PERICARDIAL EFFUSION WITH TAMPONADE AND THEN A CARDIAC ARREST. DURING CARDIOPULMONARY RESUSCITATION, IT RUPTURED INTO THE MEDIASTINUM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921723&pc=DSQ | 2916596-2025-02324 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CARDIAC ARREST AT HOME, WAS ENROLLED IN HOME HOSPICE, AND HAD A DO NOT RESUSCITATE (DNR)/DO NOT INTUBATE (DNI) ON FILE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21921840&pc=NPT | 2025587-2025-03175 | 2024-09-02 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN.Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, SUDDEN DEATH OCCURRED, AND BOTH THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. IT WAS UNKNOWN WHETHER THE PATIENT'S UNDERLYING MEDICAL HISTORY WAS A FACTOR IN THE EVENT. Manufacturer Narrative: UPDATED: B5, D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922008&pc=NPT | 2015691-2025-03337 | 2025-04-02 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-28 04:00 | NPT | COMMANDER DS WITH SAPIEN 3 | Material Fragmentation; Difficult to Remove; Material Rupture; Detachment of Device or Device Component | Perforation of Vessels | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION. DURING THE PROCEDURE, A BALLOON RUPTURE CAUSED AN AVULSED ILIAC, RESULTING IN FATAL UNCONTROLLED BLEEDING. ADDITIONALLY, AN ANNULAR RUPTURE WAS SUCCESSFULLY SEALED WITH COILS. AS PER FOLLOW-UP WITH THE FCS, THE BALLOON RUPTURE WAS NOTICED DURING INFLATION. IN RESPONSE, THE BALLOON WAS PULLED INTO THE SHEATH, CAUSING IT TO COME APART. A SURGICAL CUT-DOWN AND HEMOSTATS WERE USED TO REMOVE THE BALLOON MATERIAL FROM THE CFA. NO DIFFICULTIES WERE ENCOUNTERED DURING DELIVERY SYSTEM INSERTION THROUGH THE SHEATH, BUT WITHDRAWAL DIFFICULTIES OCCURRED AT THE TIP OF THE SHEATH WITH THE MIDDLE MARKER IN THE TIP OF THE SHEATH. ULTIMATELY, THE DELIVERY SYSTEM AND ESHEATH WERE SURGICALLY REMOVED. THE PERCEIVED ROOT CAUSE OF THE BALLOON TEAR WAS ATTRIBUTED TO STJ CALCIUM. THE RIGHT ILIAC WAS REPAIRED WITH TWO OVERLAPPING STENTS AND THE ILIAC INJURY WAS RESOLVED FOR THE MOMENT. THE AVULSED ILIAC WAS BELIEVED TO HAVE RESULTED FROM THE WITHDRAWAL DIFFICULTY DURING SHEATH REMOVAL DUE TO THE RUPTURED BALLOON REMAINING OUTSIDE THE TIP. WITHDRAWAL DIFFICULTIES CONTRIBUTED TO THE COMPLICATION, REQUIRING ADDITIONAL BLOOD TRANSFUSIONS. THE REPORTED ANNULAR RUPTURE OCCURRED DURING POST DILATION WITH A NON-EDWARDS TRUE BALLOON WHICH WAS INSERTED IN THE (NON TAVR ACCESS SITE) LEFT GROIN. THE REPORTER DID NOT ALLEGE ANY DEFICIENCIES IN THE EW DEVICES RELATED TO THE ANNULAR RUPTURE. THE POST DILATION WAS REQUIRED AS THE VALVE WAS NOT FULLY EXPANDED DUE TO THE BALLOON BURST. THIS POST DILATION/ANNULAR RUPTURE WAS OCCURRING AT THE SAME TIME AS THE RIGHT ILIAC INJURY REPAIR AND BLEEDING. THE PATIENT WAS PLACED ON ECMO FOR HEMODYNAMIC SUPPORT AS THE PATIENT WAS STILL LOSING VOLUME. DESPITE ALL OF THIS, A SUCCESSFUL COILING OF THE RUPTURE FROM THE LEFT ACCESS WAS PERFORMED. THE TEAM THEN IMAGED THE RIGHT GROIN WHERE IT WAS NOTICED THE TWO OVERLAPPING STENTS HAD SEPARATED. THE ROOT CAUSE OF THIS SEPARATION IS UNKNOWN. BLOOD TRANSFUSIONS WERE ADMINISTERED THROUGHOUT THE PROCEDURE, INCLUDING BLOOD FROM A CELL SAVER. THE PROVIDED DEVICE-RELATED PHOTOS AND 3MENSIO IMAGING WERE REVIEWED. THE 3MENSIO IMAGES INDICATED CALCIFICATION IN THE LANDING ZONE. THE DS BALLOON APPEARED TO HAVE RUPTURED BOTH RADIALLY AND LONGITUDINALLY, WITH THE DISTAL TIP AND BALLOON MATERIAL SEPARATING FROM THE DELIVERY SYSTEM. ADDITIONALLY, THE BALLOON MATERIAL WAS INVERTED OVER THE NOSE TIP, AND THE BALLOON SPRING WAS ELONGATED. Manufacturer Narrative: INVESTIGATION IS ONGOING. THIS IS ONE OF TWO REPORTS BEING SUBMITTED FOR THIS CASE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE RELATED REPORT NUMBER. UPDATED B4, G3,, G6, H2, H10, AND H11. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAS BEEN UPDATED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE PROVIDED IMAGERY WAS EVALUATED AND REVEALED THE FOLLOWING: THE BALLOON APPEARS TO HAVE BURST RADIALLY AND LONGITUDINALLY WITH DISTAL-TIP AND BALLOON MATERIAL SEPARATED FROM THE DELIVERY SYSTEM, THE BALLOON MATERIAL WAS INVERTED OVER THE NOSE TIP, AND THE BALLOON SPRING WAS ELONGATED. 3MENSIO SHOWS CALCIFICATION IN THE LANDING ZONE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, BALLOON BURST EVENTS DURING VALVE DEPLOYMENT HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE OF THESE EVENTS HAS HISTORICALLY BEEN DUE TO CALCIFICATION IN THE LANDING ZONE OR OVER-INFLATION OF THE BALLOON. THE COMPLAINT FOR BALLOON BURST WAS CONFIRMED BASED ON THE IMAGERY REVIEW. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS 3MENSIO SHOWS CALCIFICATION IN THE LANDING ZONE. THE PRESENCE OF CALCIFICATION CAN CREATE A CHALLENGING ANATOMY FOR BALLOON INFLATION. WHILE THE BALLOONS ARE SUFFICIENTLY DESIGNED AND TESTED TO ENSURE THE BURST PRESSURE IS AT OR ABOVE THE RATED BURST PRESSURE, CALCIFIED NODULES CAN COMPROMISE THE STRUCTURE OF THE BALLOON WALL VIA FOLLOWING MECHANISMS SUCH AS PUNCTURE, LOCAL OVERSTRETCHING, OPEN CELL IMPINGEMENT, OR STRESS CONCENTRATION. THE COMPLAINTS FOR DIFFICULTY OR INABILITY TO WITHDRAW SYSTEM THROUGH SHEATH, DISTAL TIP SEPARATION, AND BALLOON SEPARATION WERE CONFIRMED BASED ON THE IMAGERY REVIEW. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (WITHDRAWAL OF BURST BALLOON, DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENTS. AS THE BALLOON BURST, THE ALTERED BALLOON PROFILE COULD HAVE MADE IT MORE SUSCEPTIBLE TO CATCH ON THE DISTAL END OF SHEATH TIP, WHICH WOULD HAVE THEN LED TO THE EXPERIENCED RETRIEVAL DIFFICULTY. IN SUCH A CONDITION, APPLYING ADDITIONAL PULL FORCE OR MANIPULATING THE DEVICE TO OVERCOME THE WITHDRAWAL DIFFICULTY CAN LEAD TO THE REPORTED SEPARATION. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922313&pc=NPT | 2025587-2025-03178 | 2025-04-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-28 04:00 | NPT | EVOLUT FX DCS | Device-Device Incompatibility; Material Twisted/Bent; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED WHICH WAS STANDARD FOR THE IMPLANTING PHYSICIAN. IT WAS NOTED THAT MINIMAL CALCIFICATION WAS OBSERVED ON THE ARCH. DURING ADVANCEMENT OF THE DELIVERY CATHETER SYSTEM (DCS) INTO THE AORTIC ARCH FROM THE DESCENDING AORTA, THE DCS WAS OBSERVED TO BE "STUCK" NEAR THE LEFT SUBCLAVIAN ARTERY OSTIUM. SUBSEQUENTLY, THE "STABILITY LAYER" OF THE DCS APPEARED TO BE BENT, AND THE DCS, ALONG WITH PART OF THE DCS CAPSULE, PUSHED INTO INTO THE LEFT SUBCLAVIAN ARTERY OSTIA. IT WAS REPORTED THAT THIS CAUSED AN IMMEDIATE LEFT SUBCLAVIAN ARTERY DISSECTION. THE PATIENT SUBSEQUENTLY "LOST PRESSURE" AND CARDIAC TAMPONADE WAS OBSERVED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR APPROXIMATELY 5-10 MINUTES; HOWEVER, THE PATIENT DIED. PER THE PHYSICIAN, THE IMPLANT PROCEDURE AND LEFT SUBCLAVIAN DISSECTION CAUSED THE DEATH. PER THE PHYSICIAN, THE PATIENT'S TORTUOUS ANATOMY ALSO CONTRIBUTED TO THE EVENT. IT WAS NOTED THAT THE NON-MEDTRONIC GUIDEWIRE WAS SUSPECTED TO BE "KINKED", OR THERE MAY HAVE NOT BEEN ENOUGH PRESSURE/TENSION ON THE GUIDEWIRE DURING DCS ADVANCEMENT WHICH MAY HAVE CONTRIBUTED TO THE BEND IN THE DCS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE NON-MEDTRONIC GUIDEWIRE CONTRIBUTED TO THE DISSECTION.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED WHICH WAS STANDARD FOR THE IMPLANTING PHYSICIAN. IT WAS NOTED THAT MINIMAL CALCIFICATION WAS OBSERVED ON THE ARCH. DURING ADVANCEMENT OF THE DELIVERY CATHETER SYSTEM (DCS) INTO THE AORTIC ARCH FROM THE DESCENDING AORTA, THE DCS WAS OBSERVED TO BE "STUCK" NEAR THE LEFT SUBCLAVIAN ARTERY OSTIUM. SUBSEQUENTLY, THE SHAFT OF THE DCS APPEARED TO BE BENT, AND THE DCS, ALONG WITH PART OF THE DCS CAPSULE, PUSHED INTO THE LEFT SUBCLAVIAN ARTERY OSTIA. IT WAS REPORTED THAT THIS CAUSED AN IMMEDIATE LEFT SUBCLAVIAN ARTERY DISSECTION. THE PATIENT SUBSEQUENTLY "LOST PRESSURE" AND CARDIAC TAMPONADE WAS OBSERVED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR APPROXIMATELY 5-10 MINUTES; HOWEVER, THE PATIENT DIED. PER THE PHYSICIAN, THE IMPLANT PROCEDURE AND LEFT SUBCLAVIAN DISSECTION CAUSED THE DEATH. PER THE PHYSICIAN, THE PATIENT'S TORTUOUS ANATOMY ALSO CONTRIBUTED TO THE EVENT. IT WAS NOTED THAT THE NON-MEDTRONIC GUIDEWIRE (SAFARI) WAS SUSPECTED TO BE "KINKED", OR THERE MAY HAVE NOT BEEN ENOUGH PRESSURE/TENSION ON THE GUIDEWIRE DURING DCS ADVANCEMENT WHICH MAY HAVE CONTRIBUTED TO THE BEND IN THE DCS. ADDITIONAL INFORMATION WAS RECEIVED THAT THE NON-MEDTRONIC GUIDEWIRE CONTRIBUTED TO THE DISSECTION. Manufacturer Narrative: IMAGE REVIEW: MEDIA IMAGES OF THE EVENT WERE PROVIDED. THERE WAS NO COMPUTED TOMOGRAPHY (CT) OR EXECUTIVE SUMMARY PROVIDED FOR ANATOMICAL CONSIDERATIONS. IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. EVIDENCE SHOWS POSSIBLE AORTA INSUFFICIENCY AND POSSIBLE AORTIC DISSECTION POST BAV. IT WAS REPORTED THAT DURING ADVANCEMENT OF THE DELIVERY CATHETER SYSTEM (DCS) INTO THE AORTIC ARCH FROM THE DESCENDING AORTA, THE DCS WAS OBSERVED TO BE "STUCK" NEAR THE LEFT SUBCLAVIAN ARTERY OSTIUM. SUBSEQUENTLY, THE SHAFT OF THE DCS APPEARED TO BE BENT, AND THE DCS, ALONG WITH PART OF THE DCS CAPSULE, PUSHED INTO THE LEFT SUBCLAVIAN ARTERY OSTIA. IT WAS REPORTED THAT THIS CAUSED AN IMMEDIATE LEFT SUBCLAVIAN ARTERY DISSECTION. EVIDENCE IS CONSISTENT WITH DESCRIPTION OF A DISSECTION. AS STATED IN THE INSTRUCTIONS FOR USE (IFU). POTENTIAL RISK ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT FX+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION), AND DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5. D4. FULL UDI WAS ADDED. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {EVFXPLUS-29}; PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922359&pc=NPT | 2015691-2025-03340 | 2025-04-02 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-28 04:00 | NPT | US S3 UR VALVE 26MM | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE WITH A 26MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION. DURING THE PROCEDURE, A BALLOON RUPTURE CAUSED AN AVULSED ILIAC, RESULTING IN FATAL UNCONTROLLED BLEEDING. ADDITIONALLY, AN ANNULAR RUPTURE WAS SUCCESSFULLY SEALED WITH COILS. AS PER FOLLOW-UP WITH THE FCS, THE BALLOON RUPTURE WAS NOTICED DURING INFLATION. IN RESPONSE, THE BALLOON WAS PULLED INTO THE SHEATH, CAUSING IT TO COME APART. A SURGICAL CUT-DOWN AND HEMOSTATS WERE USED TO REMOVE THE BALLOON MATERIAL FROM THE CFA. NO DIFFICULTIES WERE ENCOUNTERED DURING DELIVERY SYSTEM INSERTION THROUGH THE SHEATH, BUT WITHDRAWAL DIFFICULTIES OCCURRED AT THE TIP OF THE SHEATH WITH THE MIDDLE MARKER IN THE TIP OF THE SHEATH. ULTIMATELY, THE DELIVERY SYSTEM AND ESHEATH WERE SURGICALLY REMOVED. THE PERCEIVED ROOT CAUSE OF THE BALLOON TEAR WAS ATTRIBUTED TO STJ CALCIUM. THE RIGHT ILIAC WAS REPAIRED WITH TWO OVERLAPPING STENTS AND THE ILIAC INJURY WAS RESOLVED FOR THE MOMENT. THE AVULSED ILIAC WAS BELIEVED TO HAVE RESULTED FROM THE WITHDRAWAL DIFFICULTY DURING SHEATH REMOVAL DUE TO THE RUPTURED BALLOON REMAINING OUTSIDE THE TIP. WITHDRAWAL DIFFICULTIES CONTRIBUTED TO THE COMPLICATION, REQUIRING ADDITIONAL BLOOD TRANSFUSIONS. THE REPORTED ANNULAR RUPTURE OCCURRED DURING POST DILATION WITH A NON-EDWARDS BALLOON WHICH WAS INSERTED IN THE (NON TAVR ACCESS SITE) LEFT GROIN. THE REPORTER DID NOT ALLEGE ANY DEFICIENCIES IN THE EW DEVICES RELATED TO THE ANNULAR RUPTURE. THE POST DILATION WAS REQUIRED AS THE VALVE WAS NOT FULLY EXPANDED DUE TO THE BALLOON BURST. THIS POST DILATION/ANNULAR RUPTURE WAS OCCURRING AT THE SAME TIME AS THE RIGHT ILIAC INJURY REPAIR AND BLEEDING. THE PATIENT WAS PLACED ON ECMO FOR HEMODYNAMIC SUPPORT AS THE PATIENT WAS STILL LOSING VOLUME. DESPITE ALL OF THIS, A SUCCESSFUL COILING OF THE RUPTURE FROM THE LEFT ACCESS WAS PERFORMED. THE TEAM THEN IMAGED THE RIGHT GROIN WHERE IT WAS NOTICED THE TWO OVERLAPPING STENTS HAD SEPARATED. THE ROOT CAUSE OF THIS SEPARATION IS UNKNOWN. BLOOD TRANSFUSIONS WERE ADMINISTERED THROUGHOUT THE PROCEDURE, INCLUDING BLOOD FROM A CELL SAVER. THE PROVIDED DEVICE-RELATED PHOTOS AND 3MENSIO IMAGING WERE REVIEWED. THE 3MENSIO IMAGES INDICATED CALCIFICATION IN THE LANDING ZONE. THE DS BALLOON APPEARED TO HAVE RUPTURED BOTH RADIALLY AND LONGITUDINALLY, WITH THE DISTAL TIP AND BALLOON MATERIAL SEPARATING FROM THE DELIVERY SYSTEM. ADDITIONALLY, THE BALLOON MATERIAL WAS INVERTED OVER THE NOSE TIP, AND THE BALLOON SPRING WAS ELONGATED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PROCEDURAL FACTORS (POST DILATION WITH A NON-EDWARDS BALLOON WHICH WAS INSERTED IN THE (NON TAVR ACCESS SITE) LEFT GROIN) OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922462&pc=DSQ | 2916596-2025-02426 | 2010-01-01 05:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED IN THE RESEARCH ARTICLE TITLED, ¿LESS INVASIVE IMPLANTATION OF THIRD GENERATION LEFT VENTRICULAR ASSIST DEVICES: PRELIMINARY ANALYSIS¿, THE STUDY EVALUATED THE SHORT-TERM BENEFITS OF A LESS INVASIVE SURGICAL APPROACH (LIS) COMPARED TO TRADITIONAL FULL STERNOTOMY (FS) FOR IMPLANTING THIRD GENERATION LEFT VENTRICULAR ASSIST DEVICES (LVADS). A RETROSPECTIVE ANALYSIS WAS PERFORMED ON 115 CENTRIFUGAL PUMPS RECIPIENTS (HEART WARE HVAD AND HEARTMATE 3) BETWEEN 2010 AND 2023. OF THE 115 LVAD IMPLANTS ANALYZED, 79% (N=91) HAD PERFORMED USING FS, WHILE 21% (N=24) USED LIS. THE LIS COHORT WAS YOUNGER (MEDIAN AGE 57 VS. 61 YEARS, P=0.05) AND MORE FREQUENTLY RECEIVED THE HVAD DEVICE (83% VS. 17%, P=0.001). THEY PRESENTED WITH WORSE HEMODYNAMIC PARAMETERS AND A HIGHER FREQUENCY OF MODERATE TO SEVERE TRICUSPID REGURGITATION (83.3% VS. 45.6%, P=0.001). PERIOPERATIVELY, THE LIS GROUP HAD A LOWER USE AND SHORTER TIME OF EXTRACORPOREAL CIRCULATION (62% VS. 100%, P=0.001; MEDIAN TIME 60 VS. 92.5 MINUTES, P=0.007), LOWER INCIDENCE OF RIGHT HEART FAILURE (21.7% VS. 30.03%, P=0.044), AND FASTER RESPIRATORY WEANING EXTUBATION AT 48 HOURS, (70.8% VS. 52.7%, P=0.029). ADDITIONALLY, NO REVISION FOR BLEEDING WAS REQUIRED (0% VS. 15.5%, P=0.014), AND IN-HOSPITAL MORTALITY WAS LOWER (4.2% VS. 12.4%, P=0.003). AT SIX MONTHS, THE LIS GROUP EXPERIENCED FEWER DRIVELINE INFECTIONS (10.8% VS. 18.2%, P=0.05), CEREBRAL EVENTS (4.6% VS. 6.9%, P=0.044), AND PUMP THROMBOSIS (0% VS. 1.4%, P=0.039). THE LESS INVASIVE SURGICAL APPROACH FOR LVAD IMPLANTATION IS A FEASIBLE AND SAFE ALTERNATIVE TO FULL STERNOTOMY IN SELECTED PATIENTS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2010 AS THE PATIENTS WERE IMPLANTED BETWEEN 2010 AND 2023. DATE OF DEATH WAS NOT PROVIDED. AUTHOR INFORMATION: LOFORTE, A, ET AL. ¿LESS INVASIVE IMPLANTATION OF THIRD GENERATION LEFT VENTRICULAR ASSIST DEVICES: PRELIMINARY ANALYSIS.¿ JOURNAL OF CARDIOVASCULAR MEDICINE, VOL. 25, NO. SUPPLEMENT 1, 1 NOV. 2024, PP. E14¿E14, HTTPS://DOI.ORG/10.2459/01.JCM.0001096284.27626.65. ACCESSED 3 APR. 2025. UNIVERSITÀ DI TORINO, DIPARTIMENTO DI SCIENZE CHIRURGICHE, ASOU CITTÀ DELLA SALUTE E DELLA SCIENZA DI TORINO, SDCU DI CARDIOCHIRURGIA TORINO; SCDU CARDIOCHIRURGIA, OSPEDALE MOLINETTE, CITTÀ DELLA SALUTE E DELLA SCIENZA DI TORINO; CIRUGÍA CARDIOVASCULAR Y TORÁCICA, FUNDACIÓN FAVALORO, BUENOS AIRES, ARGENTINA, BUENOS AIRES; UNIVERSITÀ DI TORINO, DIPARTIMENTO DI SCIENZE MEDICHE, ASOU CITTÀ DELLA SALUTE E DELLA SCIENZA DI TORINO, SC DI CARDIOLOGIA TORINO; UNIVERSITÀ DI TORINO, DIPARTIMENTO DI ANESTESIA E RIANIMAZIONE, ASOU CITTÀ DELLA SALUTE E DELLA SCIENZA DI TORINO, U CARDIORIANIMAZIONE TORINO NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922800&pc=LYJ | 1644487-2025-10053 | 2025-04-03 04:00 | Death | HOUSTON | 2025-04-28 04:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922926&pc=NVN | 2017865-2025-51384 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE FOR ANALYSIS. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922928&pc=NVN | 2017865-2025-51385 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21922929&pc=LWP | 2017865-2025-51383 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. AS RECEIVED THE DEVICE WAS IN NORMAL RANGE OF OPERATION AND NO ANOMALIES WERE NOTED. ELECTRICAL, LONGEVITY, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21923229&pc=DSQ | 2916596-2025-02653 | 2025-03-27 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT'S AUTOMATIC IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (AICD) SHOCKED THEM WHILE IN HOSPICE CARE. THE PATIENT WAS THEN TAKEN TO THE HOSPITAL WHERE THE PATIENT'S FAMILY DECIDED TO REMOVE THEM FROM CARE AND END THEIR LEFT VENTRICULAR ASIST DEVICE THERAPY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21923244&pc=DSQ | 2916596-2025-02502 | 2025-03-27 04:00 | Death | THORATEC CORPORATION | 2025-04-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Fever; Sepsis; Chills; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS ON (B)(6) 2025. THE SOURCE OF INFECTION WAS UNKNOWN, ALTHOUGH THE PATIENT HAD A KNOWN DRIVELINE INFECTION. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. ACCORDING TO THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR, THE PATIENT CAME IN SEPTIC AND WAS TREATED WITH INTRAVENOUS (IV) ANTIBIOTICS. THE PATIENT STARTED HAVING PERSISTENT LOW FLOW ALARMS. THE POINT OF CARE ULTRASONOGRAPHY (POCUS) SHOWED A SMALL LEFT VENTRICLE CAVITY ON LEFT SIDE. THE PATIENT WAS ON LEVOPHED. THE CLINICAL SYMPTOMS IDENTIFIED WERE RIGORS, AND METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) BACTEREMIA. THE PATIENT WAS ALREADY DO NOT RESUSCITATE (DNR)/DO NOT INTUBATE (DNI) PRIOR TO ADMISSION. THE PATIENT'S FAMILY DID NOT ESCALATE CARE AND FOCUSED ON COMFORT CARE.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT'S DRIVELINE INFECTION STARTED ON (B)(6) 2021 (MFR #: 2916596-2025-04106). Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORT OF LOW FLOW ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE SUBMITTED FOR REVIEW. NO AUTOPSY WAS PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING SEPSIS AND DEATH. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, SEVERAL SECTIONS OF THE HEARTMATE 3 LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21923328&pc=PNJ | 2017865-2025-51402 | 2025-04-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Failure to Capture | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PRESENTED FOR IMPLANT PROCEDURE. DURING SURGERY, THE LEADLESS PACEMAKER (LP) EXHIBITED LOSS OF CAPTURE. WHEN THE LP WAS REPOSITIONED, THE BLOOD PRESSURE DROPPED AND EFFUSION WAS NOTED ON ECHOCARDIOGRAM WITH LACK OF CARDIAC MOTION. PERICARDIOCENTESIS WAS PERFORMED AND THE PHYSICIAN ATTEMPTED TO CLOSE THE TEAR IN THE RIGHT VENTRICULAR (RV) WALL. THE PATIENT WAS PLACED ON A CARDIAC BYPASS AND CARDIAC MASSAGE WAS PERFORMED. ATTEMPTS AT REVIVING THE PATIENT WERE UNSUCCESSFUL. THE PATIENT DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21923330&pc=PNJ | 2017865-2025-51403 | 2025-04-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-28 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PRESENTED FOR IMPLANT PROCEDURE. DURING SURGERY, THE LEADLESS PACEMAKER (LP) EXHIBITED LOSS OF CAPTURE. WHEN THE LP WAS REPOSITIONED, THE BLOOD PRESSURE DROPPED AND EFFUSION WAS NOTED ON ECHOCARDIOGRAM WITH LACK OF CARDIAC MOTION. PERICARDIOCENTESIS WAS PERFORMED AND THE PHYSICIAN ATTEMPTED TO CLOSE THE TEAR IN THE RIGHT VENTRICULAR (RV) WALL. THE PATIENT WAS PLACED ON A CARDIAC BYPASS AND CARDIAC MASSAGE WAS PERFORMED. ATTEMPTS AT REVIVING THE PATIENT WERE UNSUCCESSFUL. THE PATIENT DECEASED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21915897&pc=OZP | 2032227-2025-162970 | 2025-01-16 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-27 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS PASSED AWAY AT HOME ON (B)(6) 2025. THE CUSTOMER BLOOD GLUCOSE WAS UNKNOWN AT THE TIME OF EVENT. THE CAUSE OF THE DEATH WAS DUE TO DIABETIC KETOACIDOSIS. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-342, MMT-7040A, MMT-431AG. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER WORN INSULIN PUMP AT THE TIME OF PASSING. CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48HRS OF REPORTED EVENT AND CUSTOMER WAS NOT USING SENSORS. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-342. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. NO PRODUCT RETURN IS REQUIRED FOR MMT-431AG. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21907554&pc=NVZ | 3008973940-2025-02916 | 2025-03-23 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-25 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Septic Shock | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT EXPERIENCED SEPTIC SHOCK, PSEUDOMONAS BACTEREMIA, NECROTIZING FASCIITIS AND PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21907555&pc=DTB | 2649622-2025-11221 | 2025-03-23 04:00 | Death | MPRI | 2025-04-25 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Septic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT EXPERIENCED SEPTIC SHOCK, PSEUDOMONAS BACTEREMIA, NECROTIZING FASCIITIS AND PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST A PROXIMAL SEGMENT OF THE LEAD 6.6 CM LONG WAS RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21907557&pc=DTB | 2649622-2025-11222 | 2025-03-23 04:00 | Death | MPRI | 2025-04-25 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Septic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT EXPERIENCED SEPTIC SHOCK, PSEUDOMONAS BACTEREMIA, NECROTIZING FASCIITIS AND PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED A PROXIMAL SEGMENT OF THE LEAD 6.8 CM LONG WAS RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21908382&pc=OZD | 1220648-2025-27836 | 2025-04-05 04:00 | Death | ABIOMED, INC. | 2025-04-25 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS PLACED ON IMPELLA CP FOR HEMODYNAMIC SUPPORT. IT WAS NOTED THAT THE IMPELLA CP WAS BACKLOADED AND STARTED ON AUTO. HOWEVER, FLOWS WERE ONLY 2.2. THERE WAS CONCERNED THAT THE RIGHT VENTRICLE FUNCTION/VOLUME STATUS WAS CAUSING LOW FLOWS. CALCIUM AND VOLUME WERE GIVEN, FLOWS INCREASED TO 2.8. INTERMITTENT SUCTION ALARMS WERE REPORTED. NEUROLOGY WAS CONSULTED DUE TO POSSIBLE INTRACRANIAL HEMORRHAGE. HEPARIN WAS PAUSED. THERE WAS ALSO POOR FLOW DISTAL TO IMPELLA. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THERE WERE NO DATA LOGS OR PRODUCT RETURNED. BASED ON THE CLINICAL DESCRIPTION, THE ROOT CAUSE OF THE LOW PUMP FLOW IS MOST LIKELY DUE TO THE PATIENT'S CONDITION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21908593&pc=OZD | 1220648-2025-27838 | 2025-04-05 04:00 | Death | ABIOMED, INC. | 2025-04-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT ON IMPELLA CP SUPPORT DEVELOPED ISCHEMIA IN THEIR LOWER EXTREMITY AND HEMOLYSIS. THE PATIENT'S FAMILY OPTED TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE PATIENT'S PASSING WITH IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE IS NOT AVAILABLE FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21908921&pc=OZD | 1220648-2025-27840 | 2025-04-01 04:00 | Death | ABIOMED, INC. | 2025-04-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PUMP PLACED TO SUPPORT A PATIENT WITH CARDIAC HISTORY. THE IMPELLA CP PUMP WAS PLACED FEMORALLY DESPITE KNOWLEDGE THAT THE LEG WAS EDEMATOUS, NECROTIC AND INFECTED. THE TEAM ATTEMPTED TO PLACE DISTAL BILATERAL INFUSION TO THE LEG WITH ANTEGRADE SHEATHS. CARE ESCALATION TO THE IMPELLA 5.5 PUMP AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECPELLA) WERE DISCUSSED, ALONG WITH LIMB AMPUTATION POSSIBILITY, BUT RATHER THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21908936&pc=NPT | 2135147-2025-02151 | 2025-03-31 04:00 | Death | ABBOTT MEDICAL | 2025-04-25 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Heart Block | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VISOSN VALVE WAS CHOSEN FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. DURING PROCEDURE, THE VALVE WAS PREPPED AND LOADED AS PER INSTRUCTIONS FOR USE (IFU). AFTER PRE-DILATATION, THE PATIENT WENT TO COMPLETE HEART BLOCK (CHB) AND A TEMPORARY PACING WIRE INSERTED. THE VALVE WAS DEPLOYED USING NAVITOR CUSP OVERLAP TECHNIQUE. AT 80% VALVE DEPLOYMENT, OPTIMAL DEPTH, AN IMMEDIATE PRESSURE DROP WAS NOTED AND FLUID WAS GIVEN. THE VALVE WAS IMPLANTED AT A DEPTH OF 5MM AT NON-CORONARY CUSP (NCC) AND 5MM AT LEFT CORONARY CUSP (LCC). AFTER THE PROCEDURE, NO FLOW IN RIGHT CORONARY ARTERY (RCA) WAS NOTED, PATIENT WENT INTO CARDIAC ARREST AND CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED. THERE WERE MULTIPLE ATTEMPTS TRIED TO OPEN VESSEL BUT UNSUCCESSFUL. APPROXIMATELY AFTER 45 MINUTES, PATIENT WAS REPORTED PASSED AWAY. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS POSSIBLY RCA OCCLUSION WITH NATIVE LEAFLETS.Event Description: N/A Manufacturer Narrative: NA. THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF HYPOTENSION, HEART BLOCK, OCCLUSION, CARDIAC ARREST, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE FOR THE REPORTED HYPOTENSION, OCCLUSION, AND CARDIAC ARREST COULD NOT BE DETERMINED. THE CAUSE OF THE REPORTED DEATH WAS DUE TO THE PATIENTS ANATOMY. THE CAUSE OF THE REPORTED HEART BLOCK APPEARS TO BE RELATED TO PROCEDURAL CONDITIONS. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21909351&pc=CBK | 2518422-2025-104789 | 2023-01-01 05:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | CBK | TRILOGY DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION VIA PHONE CALL THAT A PATIENT (HUSBAND) PASSED AWAY IN 2023. THE CALLER STATES THAT THE "HUSBAND DID NOT PASS BY USE OF THE DEVICE. THE DURABLE MEDICAL EQUIPMENT (DME) COMPANY PICKED UP ALL OF THE EQUIPMENT THEY RECEIVED AN ENVELOPE WITH PAPERWORK, TRILOGY DEVICE QUESTIONS FROM TRACKED DEVICE. THE REPORTER DOES NOT HAVE DEVICE INFORMATION (IE., SN, LOT#, OR ANYTHING) ON THE LETTER. THEY NEED TO REPORT ANY DEVICE IN THEIR POSSESSION. THE ALLEGED HARM DEATH HAS BEEN ASSESSED BY THE CLINICAL EXPERT AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO DEVICE INFORMATION IS AVAILABLE. THE PRODUCT BRAND NAME, 510K NUMBER, SERIAL NUMBER, MODEL NUMBER, CATALOG ITEM IDENTIFIER, PRODUCT UDI, AND MANUFACTURE DATE WILL BE UPDATED IF AND WHEN THE INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21909521&pc=NPT | 2025587-2025-03021 | 2025-03-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-25 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Patient Device Interaction Problem | Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT DIED DUE TO A TYPE A AORTIC DISSECTION. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH.Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT DIED DUE TO A TYPE A AORTIC DISSECTION. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH OCCURRED ONE DAY FOLLOWING THE VALVE IMPLANT WHICH WAS IMPLANTED VALVE-IN-VALVE INTO A FAILED NON-MEDTRONIC (EDWARDS SAPIEN S3) TRANSCATHETER VALVE. PER THE PHYSICIAN, THE VALVE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DISSECTION OCCURRED DURING THE POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV). PER THE PHYSICIAN, THE DCS DID NOT CAUSE OR CONTRIBUTE TO THE DISSECTION.Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT DIED DUE TO A TYPE A AORTIC DISSECTION. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH OCCURRED ONE DAY FOLLOWING THE VALVE IMPLANT WHICH WAS IMPLANTED VALVE-IN-VALVE INTO A FAILED NON-MEDTRONIC (EDWARDS SAPIEN S3) TRANSCATHETER VALVE. PER THE PHYSICIAN, THE VALVE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DISSECTION OCCURRED DURING THE POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV). PER THE PHYSICIAN, THE DCS DID NOT CAUSE OR CONTRIBUTE TO THE DISSECTION. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE POST-IMPLANT BAV WAS PERFORMED DUE TO A MODERATE PARAVALVULAR LEAK. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {{D-EVOULTFX2329}}; PRODUCT LOT/SERIAL NUMBER {{UNKNOWN}}; PRODUCT TYPE: {{0195 HEART VALVE}}; IMPLANT DATE: {{NA}}; EXPLANT DATE: {{NA}}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED FOURTH PARAGRAPH. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21909737&pc=OZD | 1220648-2025-27848 | 2025-04-02 04:00 | Death | ABIOMED, INC. | 2025-04-25 04:00 | OZD | IMPELLA CP | Increase in Pressure; Infusion or Flow Problem |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND FOLLOWING THE IMPLANT THE IMPELLA CP PUMP EXPERIENCED LOW PURGE FLOW COINCIDING WITH A HIGH MOTOR CURRENT. THE DEVICE WAS REPLACED. HOWEVER, THE PATIENT EXPIRED THE FOLLOWING DAY. THE PATIENT WAS AT GYM SWIMMING WHEN THE PATIENT WENT INTO VENTRICULAR FIBRILLATION ARREST AND CODED. THE PATIENT ASPIRATED A LOT OF POOL WATER. THE PATIENT WAS BROUGHT INTO THE HOSPITAL AND WAS WORKED ON THROUGHOUT THE NIGHT. THE LEFT MAIN ARTERY WAS STENTED AND THE DECISION MADE TO PLACE IMPELLA CP DEVICE. THE PATIENT WAS SENT TO THE UNIT IN CRITICAL CONDITION ON IMPELLA MCS. SHORTLY AFTER THE IMPLANT OF THE IMPELLA CP, THERE WAS LOW PURGE FLOW COINCIDING WITH A HIGH MOTOR CURRENT. THE PUMP FLOW WAS ALSO NOTED TO BE 4.2 LITERS PER MINUTE FOR THE MAXIMUM, MINIMUM, AND AVERAGE VALUES. THE STAFF WAS ADVISED THAT THE PUMP COULD STOP BASED ON THE NOTED VALUES. THE DECISION WAS MADE TO REPLACE THE PUMP AFTER JUST UNDER THREE HOURS OF SUPPORT. THE PATIENT WAS CANNULATED FOR VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) AND WAS SENT BACK TO THE UNIT ON SUPPORT REMAINING IN CRITICAL CONDITION. CARE WAS CONTINUED. VV ECMO WAS SWITCHED TO VENOARTERIAL ECMO. THE PATIENT WAS VERY CRITICAL, WITH VERY LOW PRESSURES. THE PHYSICIAN WAS AWARE AND WAS USING THE IMPELLA CP DEVICE AS A VENT. THE FOLLOWING DAY, AS THE PATIENT CONTINUED TO DECOMPENSATE AND THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED PUMP FLOW WAS COMPROMISED. THE PATIENT MAY NOT HAVE RECEIVED APPROPRIATE HEMODYNAMIC SUPPORT. THE PUMP WAS REPLACED AND THE PATIENT EXPIRED APPROXIMATELY A DAY AND A HALF LATER. THE IMPACT TO OUTCOME IS UNKNOWN, AND THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE INITIAL (PUMP 1) IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21909788&pc=NKM | 2135147-2025-02159 | 2025-03-24 04:00 | Death | ABBOTT MEDICAL | 2025-04-25 04:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT PRESENTED WITH GRADE 4 FUNCTIONAL MITRAL REGURGITATION (MR), A RESTRICTED POSTERIOR 2 LEAFLET, AND HEART FAILURE FOR A MITRACLIP PROCEDURE. A MITRACLIP WAS IMPLANTED SUCCESSFULLY. THE FINAL MR WAS GRADE 1-2. THE EVENING AFTER THE PROCEDURE, A CARDIAC TAMPONADE DEVELOPED AND WAS DRAINED. THE PATIENT WAS STABILIZED. ON (B)(6) 2025, A BALLOON PUMP WAS PLACED. ON (B)(6) 2025, THE PATIENT'S STATUS DECLINED AND WAS PRONOUNCED DEAD. THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT DUE TO THE MITRACLIP DEVICES, BUT THE PATIENT'S COMORBIDITY (HEART FAILURE). THE CAUSE FOR THE CARDIAC TAMPONADE IS INCONCLUSIVE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE PROVIDED IMAGING WAS REVIEWED BY AN ABBOTT SENIOR STAFF CLINICAL R&D ENGINEER. BASED ON AVAILABLE INFORMATION, A CAUSE FOR THE CARDIAC TAMPONADE, CARDIAC ARREST, AND PERICARDIAL EFFUSION COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED DEATH IS RELATED TO PATIENT CONDITION (HEART FAILURE COMORBIDITY), PER CASE DETAILS. THE REPORTED PATIENT EFFECTS OF CARDIAC TAMPONADE, DEATH, CARDIAC ARREST, AND PERICARDIAL EFFUSION, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. CODES ADDED: H6 - HEALTH EFFECT - CLINICAL CODE: 3271 PERICARDIAL EFFUSION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21909869&pc=CBK | 2518422-2025-104798 | 2025-03-17 04:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | CBK | TRILOGY EVO, USA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION "THAT WHILE BEING TRANSPORTED ON THE WAY TO THE PATIENT'S PRIMARY CARE PHYSICIAN ON (B)(6) 2025 AT APPROXIMATELY 14:00, THE PATIENT (FEMALE, AGE 9) WAS NOTED TO HAVE GONE UNRESPONSIVE. UPON ARRIVAL AT AN INSTITUTIONAL FACILITY, THE PATIENT UNDERWENT CARDIOPULMONARY RESUSCITATION (CPR) AND SUBSEQUENTLY EXPIRED. A REQUEST WAS LODGED BY THE REPORTER TO PROVIDE A SUMMARY OF THE EVENT LOGS FOR FURTHER INVESTIGATION TO DETERMINE WHAT EXACTLY HAPPENED WITH THE VENTILATOR DURING THE TIME OF THE EVENT. NO DIRECT ALLEGATION OF MALFUNCTION OR FAILURE TO PERFORM TO MANUFACTURER DECLARED SPECIFICATIONS HAS BEEN LODGED AT THIS TIME. NOR HAS ANY ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED AT THIS TIME". CURRENTLY WITH THE INFORMATION AVAILABLE, CAUSE AND/OR CONTRIBUTION OF THE DEVICE TO THE PATIENT EXPIRATION CANNOT BE CONFIRMED NOR REFUTED PENDING ADDITIONAL GOOD FAITH EFFORTS AND DEVICE EVALUATION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". AFTER BEING INFORMED BY THE MANUFACTURER THAT THE SUMMARY OF THE VENTILATOR LOGS CANNOT BE PROVIDED WITHOUT FURTHER DEVICE EVALUATION, THE REPORTER AGREED TO SEND IN THE DEVICE TO THE MANUFACTURER FOR EVALUATION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21909983&pc=OZD | 1220648-2025-27851 | 2025-03-27 04:00 | Death | ABIOMED, INC. | 2025-04-25 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED A HEMATOMA AND POSITIONING ISSUES. THE PATIENT WOULD SUBSEQUENTLY EXPIRE AFTER 11 DAYS OF TOTAL MCS. THE PATIENT PRESENTED WITH A ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION ONE WEEK PRIOR AND AN INTRA-AORTIC BALLOON PUMP (IABP) WAS PLACED. THE PATIENT WAS ALSO ON PRESSERS BUT ACCORDING TO THE TEAM HAD MIXED CARDIOGENIC, SEPTIC SHOCK. THE SEPSIS IMPROVED. HOWEVER, THE PATIENT REMAINED ON DOBUTAMINE, LEVOPHED AND THE IABP. THE DECISION WAS MADE TO PLACE AN IMPELLA 5.5 DEVICE AS A BRIDGE. THE IMPELLA 5.5 WAS PLACED, AND THE PATIENT WAS SENT TO THE UNIT ON IMPELLA MCS. APPROXIMATELY SIX DAYS AFTER START OF THE SUPPORT, A HEMATOMA DEVELOPED AND WAS MANAGED WITH MANUAL PRESSURE AND HEPARIN WAS WITHHELD. THE FOLLOWING DAY POSITIONING DIFFICULTY WAS ENCOUNTERED. THERE WERE IMPELLA POSITION UNKNOWN ALARMS WITH REDUCED PULSATILITY. THE POSITION WAS VERIFIED 6.3 CENTIMETERS BUT THE LEFT VENTRICLE WAS DISTENDED. THE PATIENT HAD SEVERE MITRAL VALVE REGURGITATION. THE PATIENT WAS REINTUBATED YESTERDAY. PULMONARY EDEMA WAS WORSENING. GOALS OF CARE LEANED TOWARDS PALLIATIVE CARE INSTEAD OF LEFT VENTRICULAR ASSIST DEVICE AS THE PATIENT WAS NO LONGER A CANDIDATE FOR ANY ADVANCED SUPPORT. APPROXIMATELY FOUR DAYS LATER, THE PATIENT EXPIRED. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND NOTED THE PATIENT WAS IN CRITICAL CONDITION AND ALREADY IN SEPTIC SHOCK MIXED WITH CARDIOGENIC SHOCK. THERE WAS MINIMAL BLEEDING (SMALL HEMATOMA WITH MANUAL PRESSURE). HOWEVER, THE MAIN ISSUE WAS THE POSITIONING. THE PHYSICIAN WAS UNABLE TO GET THE PUMP IN THE CORRECT POSITION. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR. HOWEVER, TO NOTE THE DEMISE OUTCOME IS VERY UNLIKELY RELATED TO THE IMPELLA DEVICE USED. THE PATIENT WAS NOT A CANDIDATE FOR ANY ADVANCED SUPPORT FOR SURVIVAL. THE INVESTIGATION FOR THE ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910031&pc=PNJ | 2017865-2025-51220 | 2025-04-10 04:00 | Death | ABBOTT | 2025-04-25 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Capturing Problem | Arrhythmia; Cardiac Tamponade; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT HIGH CAPTURE THRESHOLD WAS OBSERVED DURING THE FIRST TWO FIXATION ATTEMPTS OF THE RIGHT VENTRICULAR DEVICE. THE DEVICE WAS FIXATED AT A THIRD LOCATION AND THE ELECTRICAL MEASUREMENTS WERE WITHIN THE ACCEPTABLE RANGES. FOLLOWING THE PROCEDURE, AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION THAT RESULTED IN TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED IN ATTEMPT TO RESOLVE THE EVENT, BUT THE PATIENT PASSED AWAY. THE DEVICE WAS DEACTIVATED AND LEFT IN SITU.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THE PATIENT EXPERIENCE CARDIAC ARREST AS A RESULT OF THE TAMPONADE. INTUBATION AND RESUSCITATION WERE PERFORMED, BUT THE PATIENT PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910036&pc=PNJ | 2017865-2025-51221 | 2025-04-10 04:00 | Death | ABBOTT | 2025-04-25 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Tamponade; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT HIGH CAPTURE THRESHOLD WAS OBSERVED DURING THE FIRST TWO FIXATION ATTEMPTS OF THE RIGHT VENTRICULAR DEVICE. THE DEVICE WAS FIXATED AT A THIRD LOCATION AND THE ELECTRICAL MEASUREMENTS WERE WITHIN THE ACCEPTABLE RANGES. FOLLOWING THE PROCEDURE, AN ECHOCARDIOGRAM REVEALED A PERICARDIAL EFFUSION THAT RESULTED IN TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED IN ATTEMPT TO RESOLVE THE EVENT, BUT THE PATIENT PASSED AWAY. THE DEVICE WAS DEACTIVATED AND LEFT IN SITU.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THE PATIENT EXPERIENCE CARDIAC ARREST AS A RESULT OF THE TAMPONADE. INTUBATION AND RESUSCITATION WERE PERFORMED, BUT THE PATIENT PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910064&pc=PSZ | 3006948883-2025-00097 | 2025-03-26 04:00 | Death | BD RAPIS DIAGNOSTICS (SUZHOU) CO. LTD. | 2025-04-25 04:00 | PSZ | VERITOR PRODUCT | False Negative Result | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED WHEN USING A VERITOR PRODUCT APPROXIMATELY 15 YEARS AGO, ONE (1) PATIENT DEATH WAS ASSOCIATED WITH A FALSE NEGATIVE FLU RESULT AFTER A H1N1 OUTBREAK. IN ADDITION, THERE WAS A HIGH RATE OF FALSE NEGATIVE FLU RESULTS OBTAINED. UPON PCR TESTING, A POSITIVE FLU RESULT WAS OBTAINED. THE CUSTOMER IS UNABLE TO PROVIDE A PRODUCT CATALOG NUMBER, A PRODUCT LOT NUMBER, AN EVENT OCCURRENCE DATE, A DATE OF DEATH, OR MEDICAL DETAILS SURROUNDING THE PATIENT'S ADVERSE EVENT. THE CUSTOMER'S PRODUCT ALLEGATION TIMELINE IS NOT ALIGNED WITH WHEN THE BD VERITOR PRODUCTS WERE RELEASED ON THE MARKET AND NO DATA COULD BE PROVIDED BY THE CUSTOMER TO CORROBORATE THE ALLEGATIONS. Manufacturer Narrative: B2. DATE OF DEATH IS UNKNOWN. THE DATE RECEIVED BY MANUFACTURER HAS BEEN USED FOR THIS FIELD. B3. DATE OF EVENT IS UNKNOWN. THE DATE RECEIVED BY MANUFACTURER HAS BEEN USED FOR THIS FIELD. D4. MEDICAL DEVICE EXPIRATION DATE: UNKNOWN, D4. UNIQUE IDENTIFIER (UDI) NUMBER: UNKNOWN, H4. DEVICE MANUFACTURE DATE: UNKNOWN. INVESTIGATION SUMMARY: THIS STATEMENT SUMMARIZES THE INVESTIGATION RESULTS REGARDING A COMPLAINT THAT ALLEGES FALSE NEGATIVE WHEN USING UNKNOWN PRODUCT (MATERIAL NUMBER UNKNOWN), BATCH NUMBER UNKNOWN. THE CUSTOMER EXPRESSED CONCERN OVER HIGH FALSE NEGATIVES WITH THE VERITOR FLU A AND B TEST KIT, ALLEGING THIS OCCURRED IN 2009. THEY INDICATED THAT IT COULD HAVE BEEN AS HIGH AS 85 PERCENT WHICH EVEN RESULTED IN A PATIENT DEATH. THE CUSTOMER WAS UNABLE TO CONFIRM A SKU NOR LOT NUMBER FOR THE PRODUCT USED IN 2009. THEY WERE ALSO UNABLE TO CONFIRM ANY DEMOGRAPHIC INFORMATION ABOUT THE PATIENT. BD QUALITY PERFORMS A SYSTEMATIC APPROACH TO INVESTIGATE ALL COMPLAINTS. THIS APPROACH INVOLVES REVIEW OF MANUFACTURING BATCH HISTORY RECORDS, VISUAL INSPECTION OF RETENTION SAMPLES, AND VISUAL INSPECTION OF CUSTOMER RETURNED SAMPLES, IF APPLICABLE. BATCH HISTORY RECORD (BHR) REVIEW AND RETAIN SAMPLE TESTING COULD NOT BE PERFORMED AS THERE WAS NO SKU NOR LOT NUMBER PROVIDED. ADDITIONALLY, NO PHOTOS OR SAMPLES WERE RECEIVED; THEREFORE, RETURN SAMPLE ANALYSIS COULD NOT BE PERFORMED. COMPLAINT HISTORY REVIEW WAS PERFORMED ON THE VERITOR FLU A AND B PRODUCT FAMILY PER AVAILABLE CUSTOMER INTAKE INFORMATION, AND THE COMPLAINTS FOR THIS PRODUCT FAMILY ARE UNDER CONTROL AND NO ADVERSE TRENDS WERE IDENTIFIED. THERE WERE ALSO NO FALSE NEGATIVE TRENDS NOTED. AT THE TIME OF THE ALLEGATION OCCURRING IN 2009, THE VERITOR FLU A AND B PRODUCT HAD NOT YET BEEN RELEASED. THE REPORTED ISSUE WAS UNABLE TO BE CONFIRMED. PER THE IFU, A NEGATIVE TEST IS PRESUMPTIVE AND IT IS RECOMMENDED THAT THESE RESULTS BE CONFIRMED BY VIRAL CULTURE OR AN FDA CLEARED INFLUENZA A AND B MOLECULAR ASSAY. NEGATIVE TEST RESULTS DO NOT PRECLUDE INFLUENZA VIRAL INFECTION AND SHOULD NOT BE USED AS THE SOLE BASIS FOR TREATMENT OR OTHER PATIENT MANAGEMENT DECISIONS. THERE WERE NO CORRECTIVE ACTIONS TAKEN AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910347&pc=DSQ | 2916596-2025-02547 | 2025-03-31 04:00 | Death | THORATEC CORPORATION | 2025-04-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED, "NON-VAD RELATED DEATH" PER COORDINATOR. IT WAS UNKNOWN IF THE DEVICE OPERATED AS EXPECTED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE SITE DID NOT BELIEVE THIS OUTCOME WAS RELATED TO THE DEVICE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910497&pc=DSQ | 2916596-2025-02373 | 2025-03-14 04:00 | Death | THORATEC CORPORATION | 2025-04-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS PREVIOUSLY ADMITTED ON (B)(6) 2025 DUE TO CONFUSION, AND IT WAS NOTED WHEN THEY ARRIVED THAT THEIR PUMP WAS OFF DUE TO HIS BATTERIES AND EMERGENCY BACKUP BATTERY BEING COMPLETELY DISCHARGED. IT WAS DETERMINED THAT HIS PUMP HAD BEEN OFF FOR APPROXIMATELY 8 HOURS. THE PATIENT HAD NO RECOLLECTION OF HEARING ANY ALARMS, ALTHOUGH IT WAS NOTED THAT THE PATIENT WAS OFTEN CONFUSED AND HAD BEEN HARD OF HEARING AND LIVED ALONE. THE PATIENT WAS DISCHARGED ON (B)(6) 2025. THE PATIENT WAS LATER FOUND DOWN IN THEIR HOME, AND IT WAS DETERMINED THAT THEY PASSED AWAY ON (B)(6) 2025 DUE TO A CARDIAC ARREST, ALTHOUGH THE CUSTOMER NOTING THAT THEY MOST LIKELY HAD BEEN DEAD FOR SEVERAL DAYS PRIOR TO BEING FOUND. WHETHER IT THE DEATH WAS DEVICE OR THERAPY RELATED WAS UNKNOWN, ALTHOUGH THE CUSTOMER STATED IT WAS POSSIBLY DUE TO A PUMP STOP AND GIVEN THE PATIENT'S HISTORY, THEY THOUGHT THAT BATTERY DEPLETION WAS THE LIKELY CAUSE. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: PREVIOUS ADMISSION ON (B)(6) 2025 IS CAPTURED UNDER MFR. REPORT # 2916596-2025-02881. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910592&pc=NPT | 2025587-2025-03028 | 2024-12-21 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-25 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR, 3 MONTHS, AND 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: UPDATED: B7, D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910624&pc=NKM | 2135147-2025-02182 |  | Death | ABBOTT MEDICAL | 2025-04-25 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Heart Failure/Congestive Heart Failure | P100009 |  | 1 | Event Description: THE ARTICLE "NOVEL RIGHT VENTRICULAR FUNCTION PARAMETERS CAN IDENTIFY SHORT-TERM NON-RESPONDERS TO TRANSCATHETER EDGE-TO-EDGE REPAIR FOR MITRAL REGURGITATION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE PROGNOSTIC VALUE OF ADVANCED RIGHT VENTRICULAR (RV) FUNCTION PARAMETERS IN PREDICTING MACE FOLLOWING PMVR USING THE MITRACLIP SYSTEM. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN PATIENTS UNDERGOING PMVR USING MITRACLIP, TAPSE NORMALIZED TO PULMONARY ARTERY PRESSURE OR RV AREA WERE STRONG AND INDEPENDENT PREDICTORS OF OUTCOME AT ONE-YEAR FOLLOW-UP. THIS SUGGESTS THE POTENTIAL ROLE OF THESE PARAMETERS COMBINING RV LONGITUDINAL FUNCTION WITH RV AFTERLOAD AND REMODELLING IN PRE-PROCEDURAL RISK STRATIFICATION. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARTIN PENICKA AT CARDIOVASCULAR CENTER AALST INSTITUTION WITH CORRESPONDING EMAIL MARTIN.PENICKA@OLVZ-AALST.BE]. THE TIME FRAME OF THE STUDY WAS NOT PROVIDED. A TOTAL OF 60 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 60 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DYSLIPIDEMIA, DIABETES, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, STROKE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DYSLIPIDEMIA, DIABETES, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, STROKE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910693&pc=DSQ | 2916596-2025-02444 | 2024-12-26 05:00 | Death | THORATEC CORPORATION | 2025-04-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO A PULMONARY EMBOLISM. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING THROMBOEMBOLISM AND DEATH. SECTION 6 ENTITLED ¿PATIENT CARE AND MANAGEMENT¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 ALSO PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR VALUES. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910831&pc=MOZ | MW5169518 |  | Death | ANIKA THERAPEUTICS, INC. | 2025-04-25 04:00 | MOZ | ORTHOVISC | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT IS DECEASED PER CRM (CIRCUMFERENTIAL RESECTION MARGIN).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21910866&pc=LYJ | 1644487-2025-10045 | 2024-08-27 04:00 | Death | HOUSTON | 2025-04-25 04:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024, NO FURTHER DETAILS WERE PROVIDED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21911143&pc=OZD | 1220648-2025-27859 | 2025-04-03 04:00 | Death | ABIOMED, INC. | 2025-04-25 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A THROMBUS WAS DISCOVERED IN THE PATIENT'S LEFT VENTRICLE ON THE LOWER PART OF THE IMPELLA DURING SUPPORT. IT WAS NOTED THAT THE PATIENT HAD BEEN OFF HEPARIN INTERMITTENTLY DUE TO SCHEDULED PROCEDURES. NO ADDITIONAL EFFECTS WERE NOTED RESULTING FROM THE THROMBUS AND THE DECISION WAS MADE TO CONTINUE SUPPORT. A FOLLOW UP TRANSESOPHAGEAL ECHOCARDIOGRAM WAS PERFORMED AND THE THROMBUS COULD NO LONGER BE FOUND. ALTHOUGH NOT VERIFIABLE, IT WAS BELIEVED THAT THE EARLIER OBSERVATION MAY HAVE BEEN AN ARTIFACT AND NOT A THROMBUS. ROUGHLY FIVE DAYS LATER, THE PATIENT SUFFERED A CEREBRAL HEMORRHAGE AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21911498&pc=MHX | 9610816-2025-000358 | 2025-04-01 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-04-25 04:00 | MHX | INTELLIVUE MX450 PATIENT MONITOR | No Audible Alarm | Insufficient Information | K130849 |  | 1 | Event Description: IT WAS REPORTED THE MONITOR DID NOT GENERATE A RED CRITICAL ALARM AND THE PATIENT PASSED AWAY; HOWEVER, IT WAS ALSO NOTED THE ECG ARRHYTHMIA ALARMS MAY HAVE BEEN TURNED OFF. ADDITIONAL INFORMATION RECEIVED INDICATED THERE WAS A DELAY IN CARE, BUT IT WAS NOT KNOWN HOW LONG THE DELAY IN CARE WAS FOR. IT WAS INDICATED THE LACK OF ALARMS MAY HAVE CONTRIBUTED TO THE DEATH, BUT NURSES ALSO INDICATED THE PATIENT DEATH MAY HAVE OCCURRED INDEPENDENT OF THE ALARM ISSUE. THE ALARMS WERE DEACTIVATED BUT THE NURSES STATED THE OPTION TO DEACTIVATE THE ALARMS SHOULD NOT BE READILY AVAILABLE. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE ANY MALFUNCTION OF THE DEVICE, AND IT CANNOT BE DETERMINED WHETHER THE LACK OF ALARM CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) EVALUATED THE DEVICE AND WAS UNABLE TO CONFIRM THE ALARM ISSUE. THE FSE PULLED FILES IN AN ATTEMPT TO REVIEW LOGS; HOWEVER, THE FILES PROVIDED WERE INSUFFICIENT; THEREFORE, THE TECHNICAL INVESTIGATION COULD NOT BE PERFORMED, AND IT COULD NOT BE CONFIRMED IF THE ALARMS WERE FUNCTIONING AT THE TIME OF THE ISSUE OR NOT.¿ THE ENGINEER WAS UNABLE TO PROVIDE THEIR ANALYSIS FINDINGS. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: E1: REPORTING INSTITUTION PHONE #: (B)(6). E1: REPORTER PHONE #: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21911591&pc=MNS | 2518422-2025-104809 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21911595&pc=MNS | 2518422-2025-104810 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912080&pc=MNS | 2518422-2025-104811 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912082&pc=MNS | 2518422-2025-104813 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Degraded; Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE WAS RECEIVED BY A THIRD-PARTY SERVICE CENTER, BUT THE EVALUATION HAS NOT YET BEGUN. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S THIRD-PARTY EVALUATION RESULTS. IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE RENTAL DREAMSTATION ST30 DEVICE WAS RETURNED AND EVALUATED BY A THIRD-PARTY SERVICE CENTER. THE DEVICE FOUND THAT NO FAULTS WERE REPORTED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED FOR PREVENTATIVE MAINTENANCE, AND THE ACCESSORY FLIP MODULE DOOR WAS MISSING FROM THE DEVICE. THE DEVICE WAS CLEANED AND TESTED, AND ALL TESTS PASSED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S THIRD-PARTY EVALUATION RESULTS. IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE RENTAL DREAMSTATION ST30 DEVICE WAS RETURNED AND EVALUATED BY A THIRD-PARTY SERVICE CENTER. THE DEVICE FOUND THAT NO FAULTS WERE REPORTED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED FOR PREVENTATIVE MAINTENANCE, AND THE ACCESSORY FLIP MODULE DOOR WAS MISSING FROM THE DEVICE. THE DEVICE WAS CLEANED AND TESTED, AND ALL TESTS PASSED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912593&pc=DXY | 2182208-2025-02054 | 2024-06-03 04:00 | Death | MEDTRONIC, INC. | 2025-04-25 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE EFFECTS OF DIGITALIS THERAPY IN PATIENTS WITH CARDIAC RESYNCHRONIZATION THERAPY (CRT). THE AUTHORS DESCRIBED PATIENT DEATHS IN THOSE ON DIGITALIS THERAPY AND THOSE THAT WERE NOT. THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS CARDIAC NON-ARRHYTHMIC DEATH DUE TO PROGRESSIVE HEART FAILURE, NON-CARDIOVASCULAR DEATHS, CARDIAC ARRHYTHMIC DEATHS, AND OTHER CARD IOVASCULAR DEATHS. THERE WERE PATIENTS IN BOTH GROUPS WHO WERE RE-HOSPITALIZED FOR CARDIAC REASONS OR WHO RECEIVED INAPPROPRIATE SHOCK THERAPY. THE CAUSES OF THE INAPPROPRIATE SHOCKS WERE UNKNOWN. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/67 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL IMPACT OF DIGITALIS THERAPY IN A LARGE MULTICENTER COHORT OF CRT-RECIPIENTS. JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE. 2024. 11, 173. DOI: 10.3390/JCDD11060173 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912616&pc=NIK | 2182208-2025-02053 | 2024-06-03 04:00 | Death | MEDTRONIC, INC. | 2025-04-25 04:00 | NIK | MEDTRONIC CRT-D | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE EFFECTS OF DIGITALIS THERAPY IN PATIENTS WITH CARDIAC RESYNCHRONIZATION THERAPY (CRT). THE AUTHORS DESCRIBED PATIENT DEATHS IN THOSE ON DIGITALIS THERAPY AND THOSE THAT WERE NOT. THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS CARDIAC NON-ARRHYTHMIC DEATH DUE TO PROGRESSIVE HEART FAILURE, NON-CARDIOVASCULAR DEATHS, CARDIAC ARRHYTHMIC DEATHS, AND OTHER CARD IOVASCULAR DEATHS. THERE WERE PATIENTS IN BOTH GROUPS WHO WERE RE-HOSPITALIZED FOR CARDIAC REASONS OR WHO RECEIVED INAPPROPRIATE SHOCK THERAPY. THE CAUSES OF THE INAPPROPRIATE SHOCKS WERE UNKNOWN. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/67 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL IMPACT OF DIGITALIS THERAPY IN A LARGE MULTICENTER COHORT OF CRT-RECIPIENTS. JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE. 2024. 11, 173. DOI: 10.3390/JCDD11060173 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912628&pc=DQX | 2025587-2025-03039 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-25 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem | Perforation | K132623 |  | 1 | Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE VENTRICULAR PERFORATION WAS TREATED WITH PERICARDIOCENTESIS. PER THE PHYSICIAN, THE VALVE AND THE DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH. ADDITIONALLY, THE PATIENT HAD A SMALL LEFT VENTRICLE THAT CONTRIBUTED TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A VENTRICULAR PERFORATION OCCURRED. PER THE PHYSICIAN, THE PERFORATION WAS CAUSED BY ADJUSTMENT OF THE GUIDEWIRE POSITION. SUBSEQUENTLY, THE PATIENT PASSED AWAY. Manufacturer Narrative: IMAGE REVIEW: 7 FLUOROSCOPIC CINE LOOPS OF THE IMPLANT PROCEDURE WERE PROVIDED FOR REVIEW. PATIENT SUMMARY WAS ALSO PROVIDED FOR ANATOMICAL REVIEW. ACCORDING TO THE REPORT, A GUIDEWIRE GWBC30 ¿ THE MEDTRONIC CONFIDA GUIDEWIRE ¿ WAS USED. IMAGES SHOW A LAO OPEN-ARCH PROJECTION WITH THE CONFIDA FLUSH IN THE AORTA¿S OUTER CURVE, WITH ITS TIP SITTING DEEP IN THE LEFT VENTRICLE. THE LEFT VENTRICLE IS REPORTED TO BE SMALL WHICH IS SUPPORTED BY IMAGES FROM THE EXECUTIVE SUMMARY REPORT. ALTHOUGH NO VISIBLE ADJUSTMENT OF THE GUIDEWIRE WAS MADE IN THE PROVIDED CINE LOOPS, ITS DISTAL END IS NOT PROPERLY POSITIONED IN A CURVE BUT CURLED UP AGAINST THE LEFT VENTRICLE-WALL, THE APPEARING TO BE TANGLED UP IN THE MV-CORDAE OR PAPILLARY MUSCLES. THERE ARE IMAGES THAT DISPLAY THE INITIAL SITUATION TO BE A FAILED PREVIOUS TAV-IN-SAV IMPLANT ATTEMPT, WITH THE SAV SHOWING MERELY AS A RING AND THE FIRST EVOLUT VALVE BEING HELD IN PLACE BY A SNARE. THE CONFIDA GUIDEWIRE TRACKS ALONG THE GREATER CURVE OF THE AORTA, OUTSIDE OF THE FIRST EVOLUT BIOPROSTHESIS. IN THE FINAL IMPLANT IMAGE, THE FIRST EVOLUT CAN BE SEEN TO BE SEVERELY TILTED IN THE ASCENDING AORTA, POKING INTO THE VESSEL WALL WITH IT¿S OUTFLOW CROWN. IN THE EVOLUT SYSTEM¿S INSTRUCTION FOR USE (IFU), POTENTIAL ADVERSE EVENTS INCLUDE, BUT ARE NOT LIMITED TO, DEATH AND CARDIAC TAMPONADE. THE IMAGES AND INFORMATION PROVIDED IN THE REPORT DON¿T ALLOW FOR IDENTIFICATION OF THE EXACT ROOT CAUSE OF THE VENTRICULAR PERFORATION LEADING TO THE PATIENT DEATH. HOWEVER, THE DEEP INSERTION OF THE CONFIDA GUIDEWIRE IN COMBINATION WITH ITS SUB-OPTIMAL TIP-PLACEMENT AND FORWARD PUSH ON THE GUIDEWIRE EXERTED DURING THE DELIVERY SYSTEM ADVANCEMENT ARE LIKELY TO HAVE CONTRIBUTED TO THE VENTRICULAR PERFORATION. IT¿S NOT PERCEIVABLE BY THE CINE LOOPS ALONE, HOW WELL THE GUIDEWIRE WAS CONTROLLED DURING THE IMPLANT ATTEMPT. IT SHOULD ALSO BE NOTED THAT THE FINAL RESTING POSITION OF THE FIRST EVOLUT BIOPROSTHESIS AND ITS RESULTING INTERACTION WITH THE VESSEL WALL WERE NOT OPTIMAL. THERE WAS NO MENTION OF MEDICATION BEING ADMINISTERED. IT WAS REPORTED THAT DESPITE THE APPLICATION OF CARDIO SYNTHESIS, THE PATIENT PASSED AWAY AFTER THE PROCEDURE. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912884&pc=MVK | 3008642652-2025-04286 | 2025-02-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 23:14:45 ON 2/16/2025. AT 09:50:06 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT. AT 09:50:52, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM WITH PVC¿S AND NSVT. AT 09:51:27, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. MULTIPLE COUNTING, MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 120 BPM WITH PVC¿S, TALL T WAVES, AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 120 BPM WITH PVC¿S, TALL T WAVES, AND MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 09:52:49 ON 2/17/2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912894&pc=MVK | 3008642652-2025-04289 | 2025-03-16 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 07:44:05 ON (B)(6) 2025. AT 01:20:48 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 240 BPM. AT 01:21:24, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ONE SINUS BEAT DEGRADING TO VT @ 200 BPM. AT 01:21:46, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ONE SINUS BEAT DEGRADING TO VT @ 220 BPM. AT 01:22:15, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE FOR FOUR SECONDS DEGRADING TO VT @ 250 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 01:22:47, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. THE RHYTHM THEN DEGRADES TO VT @ 260 BPM/VF. AT 01:23:18, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE TRANSITIONING TO SINUS BRADYCARDIA @ 40 BPM DEGRADING TO VF/FINE VF WITH VARYING AMPLITUDES. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 07:00:36 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912902&pc=MVK | 3008642652-2025-04290 | 2024-12-18 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 13:57:25 ON (B)(6) 2024. AT 21:11:30, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 240 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VT. RESPONSE BUTTON USE AND CPR/MOTION ARTIFACT PREVENTED THE LIFEVEST FROM DELIVERING A TREATMENT. THE ELECTRODE BELT WAS DISCONNECTED AT 21:12:26 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912903&pc=MVK | 3008642652-2025-04292 | 2025-03-30 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 11:34:15 ON (B)(6) 2025. AT 14:55:56 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 160 BPM WITH CPR/MOTION ARTIFACT. AT 14:57:36, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 160 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM. AT 14:58:09, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM. THE RHYTHM THEN DEGRADED TO VT FROM 170-150 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 14:58:56 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912904&pc=MVK | 3008642652-2025-04291 | 2025-04-04 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-04-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 22:12:47 ON (B)(6) 2025. AT 08:47:01 ON 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 70 BPM DEGRADING TO VF. AT 08:47:36, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 08:48:03, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, TRANSITIONING TO AN IDIOVENTRICULAR RHYTHM @ 40 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 10:08:25 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21912920&pc=KRD | 2135147-2025-02210 | 2025-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-04-25 04:00 | KRD | AMPLATZER VASCULAR PLUG II | Patient Device Interaction Problem | Aortic Valve Insufficiency/ Regurgitation | K071699 |  | 1 | Event Description: N/A. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD (DHR) AND COMPLAINT HISTORY REVIEWS WERE NOT PERFORMED BECAUSE THIS COMPLAINT WAS BASED ON AN ARTICLE REVIEW AND NO LOT INFORMATION WAS PROVIDED. SOME OF THE COMPLICATIONS REPORTED WERE RESIDUAL SHUNT, AORTIC VALVE REGURGITATION AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE WAS RECEIVED FOR ANALYSIS. BASED ON AVAILABLE INFORMATION AND DUE TO THE LIMITED INFORMATION AVAILABLE FROM THE ARTICLE, THE CAUSE OF THE REPORTED RESIDUAL SHUNT AND AORTIC VALVE REGURGITATION COULD NOT BE CONCLUSIVELY DETERMINED. DEATH WAS RELATED TO COMMODITIES OF THE INFANT REQUIRING ON-GOING ECMO SUPPORT AND RESULTING IN BLEEDING. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. B2: DEATH DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER OCCLUSION OF THE NEOAORTA TO AUGMENT MECHANICAL CIRCULATORY SUPPORT AFTER STAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21913045&pc=DSQ | 2916596-2025-02340 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-04-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE HOSPITAL ON (B)(6) 2025 WITH SHORTNESS OF BREATH. THE PATIENT WAS NOTED TO ALSO BE IN VENTRICULAR FIBRILLATION (V-FIB) AND WAS HAVING LOW FLOWS THROUGHOUT THE WEEK. THERE WAS NO HISTORY OF V-FIB PRE-IMPLANT, AND THE PATIENT DID NOT HAVE AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD). IT WAS UNCLEAR WHETHER THE ARRHYTHMIA WAS DEVICE OR THERAPY RELATED. THE PATIENT WAS TREATED WITH FLUIDS AND MEDICATIONS. THE PATIENT SUFFERED A CARDIAC ARREST AND PASSED AWAY ON (B)(6) 2025. LOG FILES WERE SUBMITTED FOR REVIEW DID NOT CAPTURE ANY PUMP FAULTS OR PUMP SPEED ISSUES, AND THE PUMP DID NOT APPEAR TO HAVE CONTRIBUTED TO THE PATIENT'S PRESENTATION. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE LOW FLOW EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED RELEVANT EVENTS ON (B)(6) 2025 FROM 09:42:58 ¿ 12:37:51, PER THE TIMESTAMPS. FROM THE BEGINNING OF THE FILE UNTIL 12:25:43, INTERMITTENT LOW FLOW EVENTS WERE CAPTURED, SEVERAL OF WHICH WERE ASSOCIATED WITH LOW FLOW HAZARD ALARMS. OF NOTE, PULSATILITY INDEX (PI) VALUES APPEARED ELEVATED DURING SOME OF THE LOW FLOW EVENTS. THE REMAINDER OF THE FILE CAPTURED A SUSTAINED LOW FLOW ALARM LASTING APPROXIMATELY 6 MINUTES. DURING THIS TIME, SEVERAL INTENTIONAL PUMP STOP FLAGS WERE CAPTURED ALONG WITH LVAD_OFF FAULTS, INDICATING THAT THE INTENTIONAL PUMP STOP BUTTON HAD BEEN PRESSED SEVERAL TIMES ON THE SYSTEM MONITOR. THE END OF THE FILE CAPTURED NO EXTERNAL POWER ALARMS DUE TO BOTH POWER CABLES BEING DISCONNECTED, FOLLOWED BY THE DISCONNECTION OF THE DRIVELINE AT 12:35:11. THESE EVENTS APPEAR CONSISTENT WITH ATTEMPTS TO TURN OFF THE PUMP DUE TO THE PATIENT'S EXPIRATION. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED PRIOR TO THE DRIVELINE DISCONNECT. IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE HOSPITAL ON (B)(6) 2025 WITH SHORTNESS OF BREATH. THE PATIENT WAS NOTED TO HAVE VENTRICULAR FIBRILLATION (V-FIB) AND WAS HAVING LOW FLOWS THROUGHOUT THE WEEK. THERE WAS NO PREVIOUS HISTORY OF V-FIB, AND THE PATIENT DID NOT HAVE AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD). IT WAS UNCLEAR WHETHER THE ARRHYTHMIA WAS DEVICE OR THERAPY RELATED. THE PATIENT WAS TREATED WITH FLUIDS AND MEDICATIONS. THE PATIENT ULTIMATELY EXPIRED AFTER SUFFERING FROM CARDIAC ARREST. IT WAS UNKNOWN WHETHER THE PATIENT'S OUTCOME WAS DEVICE OR THERAPY RELATED. THE DEVICE REPORTEDLY OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT THE PUMP WAS NOT RETRIEVED AND WOULD NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW. THIS SECTION EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. THIS SECTION ALSO NOTES THAT, IN GENERAL, THE MAGNITUDE OF THE PULSATILITY INDEX VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E., THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21913098&pc=BZD | 2518422-2025-104768 | 2024-12-04 05:00 | Death | RESPIRONICS, INC. | 2025-04-25 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Bowel Perforation | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY DUE TO BOWEL PERFORATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21913184&pc=DSQ | 2916596-2025-02492 | 2023-10-05 04:00 | Death | THORATEC CORPORATION | 2025-04-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: SECTION E1: CUSTOMER SITE WAS (B)(6). SECTION G2: CORRECTED. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21913327&pc=LWP | 2124215-2025-26583 | 2025-03-26 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-25 04:00 | LWP | ACCOLADE? MRI DR | Use of Device Problem | No Clinical Signs, Symptoms or Conditions | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT HAD UNDERGONE CORONARY ARTERY BYPASS GRAFTING (CABG) ON (B)(6) 2025, DURING WHICH TIME THE PACEMAKER HAD BEEN PROGRAMMED TO ELECTROCAUTERY PROTECTION MODE, DOO AND LRL AT 80 PPM. DEVICE TRANSMISSIONS FROM (B)(6) SHOWED FLAT ATRIAL AND VENTRICULAR EGMS, AS EXPECTED DUE TO THE PATIENT PASSING. THERE WAS ALSO ANOTHER DEVICE TRANSMISSION THE EVENING OF (B)(6) AT WHICH TIME THE PATIENT'S WIFE STATED THE PATIENT WAS STILL ALIVE, AND THOSE EGMS WERE ALSO FLAT. TECHNICAL SERVICES DISCUSSED THAT THIS WAS BECAUSE THE PACEMAKER WAS STILL IN ELECTROCAUTERY PROTECTION MODE. THE DOO MODE PROVIDES PACING IN THE ATRIAL AND VENTRICULAR CHANNELS, BUT NO SENSING OCCURS. ADDITIONALLY, THE ELECTROCAUTERY PROTECTION MODE IS NOT MEANT TO BE USED FOR LONG-TERM PACING BECAUSE IT CAN BE PRO-ARRHYTHMIC. DUE TO THE PROGRAMMING, IT WAS NOT POSSIBLE TO DETERMINE IF THE PATIENT'S DEATH WAS RELATED TO AN ARRHYTHMIA. THE PACEMAKER APPEARED TO BE FUNCTIONING AS PROGRAMMED AND THERE WAS NO DEVICE ALLEGATION.. GOOD FAITH EFFORT FOR FOLLOW-UP YIELDED NO ADDITIONAL INFORMATION FROM THE LOCAL SALES REPRESENTATIVE OR THE HEALTH CARE PROFESSIONAL (HCP) AT THE PATIENT'S DEVICE CLINIC. THE DEVICE SYSTEM WAS LIKELY NOT EXPLANTED POST-MORTEM. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21913328&pc=NVN | 2124215-2025-26597 | 2025-04-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-25 04:00 | NVN | INGEVITY MRI | Difficult to Remove; Defective Device | Pulmonary Embolism; Unspecified Infection; Ischemia Stroke | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS LEAD WAS BEING EXPLANTED DUE TO PATIENT COMPLICATIONS, SUCH AS INFECTION. HOWEVER, THE LEAD BROKE DURING REMOVAL. IT WAS NOTED THAT THE LEAD EMBOLIZED TO THE PATIENT'S LUNGS. IT WAS ALSO NOTED THAT THE PATIENT EXPERIENCED A STROKE. THE LEAD WAS ABLE TO BE REMOVED AFTER MULTIPLE ATTEMPTS. A TEMPORARY WIRE WAS IMPLANTED WHILST THE INFECTION CLEARS. THE LEAD IS EXPECTED TO BE RETURNED FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION PROVIDED FROM THE FIELD INDICATED THE PATIENT LATER PASSED AWAY DUE TO THE COMPLICATIONS INVOLVING THE REMOVAL OF THE RV LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS EVENT WILL BE UPDATED SHOULD THE RV LEAD BE RETURNED BACK FROM THE FIELD FOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21913634&pc=NPT | 2015691-2025-03292 | 2025-03-18 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-25 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS' AFFILIATE IN JAPAN, THE PATIENT UNDERWENT A TRANSFEMORAL (TF) TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) WITH A 29 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE. DUE TO LOW CARDIAC FUNCTION, THE TAVR OPERATION WAS PERFORMED UNDER EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT. THE ECMO WAS REMOVED AND REPLACED WITH AN INTRAAORTIC BALLOON PUMPING (IABP) AFTER THE TAVR, AND THE PATIENT RETURNED TO THE ICU. ON POSTOPERATIVE DAY (POD) 7, SEIZURES WERE OBSERVED. ON POD 8, THE PATIENT EXPIRED DUE TO CEREBRAL INFARCTION (CI). Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED, BUT INVESTIGATION RESULTS SUGGEST THAT FACTORS/MECHANISMS MENTIONED ABOVE AND PROCEDURAL FACTORS (PLACEMENT OF ECMO AND IABP) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21914033&pc=NPT | 2015691-2025-03298 | 2025-03-31 04:00 | Death | EDWARDS LIFESCIENCES | 2025-04-25 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Degraded | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST, APPROXIMATELY 4 YEARS AND 8 MONTHS AFTER A TAVR PROCEDURE WITH A SAPIEN 3 ULTRA VALVE, IT WAS NOTED THE VALVE WAS FAILING. THE PATIENT PRESENTED WITH SEVERE DYSPNEA, ORTHOPNEA, LOWER EXTREMITY EDEMA, AND EARLY SATIETY. THE PATIENT WAS ADMITTED FOR ACUTE CONGESTIVE HEART FAILURE EXACERBATION. THE VALVE EXHIBITED SEVERE VALVE DEGENERATION WITH SEVERE TO CRITICAL PROSTHETIC VALVE STENOSIS, MODERATE TO SEVERE CENTRAL VALVULAR REGURGITATION, AND THICKENED LEAFLETS. THE PATIENT WAS BEING WORKED UP FOR A VALVE IN VALVE PROCEDURE. HOWEVER, THE PROCEDURE WAS NOT PERFORMED. THE PATIENT WAS DEEMED TOO HIGH RISK DUE TO COMORBIDITIES INCLUDING SEVERE PULMONARY HYPERTENSION AND ACUTE KIDNEY INJURY ON CHRONIC KIDNEY DISEASE STAGE IIIB AND PALLIATIVE CARE WAS RECOMMENDED. THE PATIENT'S CONDITION CONTINUED TO DETERIORATE AND UPON DISCUSSION, THE FAMILY DECIDED TO TRANSITION THE PATIENT TO COMFORT CARE. THE PATIENT PASSED AWAY APPROXIMATELY 3 WEEKS AFTER ADMISSION. THE CAUSE OF DEATH WAS REPORTED AS ACUTE CONGESTIVE HEART FAILURE SECONDARY TO SEVERE PROSTHETIC AORTIC VALVE STENOSIS. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS G6, H2 AND CONCLUSIONS IN THIS SECTION HAVE BEEN UPDATED. H6 CODES WERE CORRECTED: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS. THE EVENTS REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE COMPLAINT FOR DEGENERATION WAS CONFIRMED BASED ON PROVIDED MEDICAL RECORDS. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD REFERS TO CHANGES INTRINSIC TO THE VALVE, AND CAN INCLUDE FAILURE MODES SUCH AS WEAR, CALCIFICATION, LEAFLET TEAR, STENT CREEP, LEAFLET DISRUPTION, OR LEAFLET RETRACTION. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A VERY COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT RELATED (E.G. PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. IN THIS CASE, THE PATIENT HAD VAST COMORBIDITIES (E.G. HTN, CHF, DM TYPE 2, HYPERLIPIDEMIA, CHRONIC KIDNEY DISEASE, MORBID OBESITY), WHICH ARE WELL-KNOWN RISK FACTORS FOR THE EARLY VALVE DEGENERATION. PER RECORDS, THE PATIENT HAS MEDICAL HISTORY OF HISTORY OF METABOLIC SYNDROME IN THE SETTING OF MORBID OBESITY WITH HYPERTENSION, HYPERLIPIDEMIA, AND TYPE 2 DIABETES MELLITUS, CHRONIC KIDNEY DISEASE STAGE. AS SUCH, AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (PRE-EXISTING COMORBIDITIES) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. HOWEVER, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21914255&pc=LWP | 2017865-2025-51248 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-25 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS DILATED CARDIOMYOPATHY. Manufacturer Narrative: THE DEVICE RETURNED FOR ANALYSIS DUE TO A PATIENT DEATH. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21914265&pc=NVN | 2017865-2025-51249 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-25 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS DILATED CARDIOMYOPATHY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21914365&pc=QFG | 3013756811-2025-77442 | 2025-04-10 04:00 | Death | TANDEM DIABETES CARE | 2025-04-25 04:00 | QFG | TANDEM MOBI INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K233044 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. PER THE CONTACT, THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS AN INTENTIONAL INSULIN OVERDOSE. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THANKS! Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21914477&pc=QZI | 2124215-2025-26626 | 2025-03-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-25 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER TO TREAT ATRIAL FIBRILLATION (A FIB) THE PATIENT PASSED AWAY. THE PULMONARY VEIN ISOLATION PROCEDURE WAS COMPLETED WITH NO COMPLICATIONS OCCURRING DURING TREATMENT. HOWEVER, ONCE THE PATIENT WAS CLOSED THEY WENT INTO VENTRICULAR FIBRILLATION AND EVENTUALLY ASYSTOLE. CHEST COMPRESSIONS WERE STARTED AND "THE CHEST WAS NOTED TO FEEL HARD". ONCE THE PATIENT WAS STABILIZED A HEART CATH WAS PERFORMED AND AN OCCLUSION WAS FOUND IN THE RIGHT CORONARY ARTERY, THERE WERE NO ADDITIONAL FINDINGS. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT AND PASSED AWAY " A COUPLE OF DAYS LATER". THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21914506&pc=DSQ | 3007042319-2025-00835 | 2025-03-06 05:00 | Death | HEARTWARE, INC. | 2025-04-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Obstruction of Flow; Deformation Due to Compressive Stress; Pumping Problem | Stenosis; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPERIENCED LOW FLOWS AND WORSENING HEART FAILURE SYMPTOMS AND WAS HOSPITALIZED. A COMPUTED TOMOGRAPHY (CT) SCAN IDENTIFIED OUTFLOW GRAFT (OFG) COMPRESSION FROM AROUND THE VAD OUTFLOW GRAFT, WHICH HAD A NON-MEDTRONIC GRAFT PLACED AROUND THE MEDTRONIC OFG, AND STENOSIS. AN ECHOCARDIOGRAM ALSO SHOWED BIVENTRICULAR FAILURE. THE PATIENT WAS TREATED WITH MEDICATION FOR HEART FAILURE SYMPTOMS AND WAS PLACED ON HOSPICE CARE. THE PATIENT SUBSEQUENTLY PASSED AWAY DUE TO BIVENTRICULAR FAILURE. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM OUTFLOW GRAFT. D4: MODEL #: 1125/ CATALOG #: 1125 / EXPIRATION DATE: 31-OCT-2022 / SERIAL OR LOT#: (B)(6) / D9: NO / H3: NO / H4: MFG DATE: 01-OCT-2017 / H5: YES / H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ### A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. ADDITIONAL PRODUCTS: OUTFLOW GRAFT 17044099-0558 H3: NO. H5: YES. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE ASSOCIATED OUTFLOW GRAFT (LOT NUMBER 17044099-0558) WERE NOT RETURNED FOR EVALUATION AS BOTH DEVICES REMAIN IMPLANTED. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THE REPORTED OUTFLOW GRAFT COMPRESSION AND LOW FLOW EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY DUE TO BIVENTRICULAR FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, WORSENING HEART FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS, IT WAS NOTED THAT THIS PATIENT HAD A HISTORY OF HEART FAILURE. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTI-PLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21915033&pc=DSQ | 2916596-2025-02486 | 2025-04-06 04:00 | Death | THORATEC CORPORATION | 2025-04-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE AND CARDIAC ARREST. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS NOT REPORTED WHETHER THE OUTCOME WAS THOUGHT TO BE DEVICE OR THERAPY RELATED. NO PRODUCT WAS RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21915062&pc=PNJ | 2017865-2025-51286 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-04-25 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension; Unspecified Infection; Respiratory Failure | P150035 |  | 1 | Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21897128&pc=EOQ | 2955842-2025-16755 | 2025-02-12 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-04-24 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT A CLINICAL STUDY REGISTRY PATIENT UNDERWENT AN UNEVENTFUL 19-MINUTE ION-ASSISTED LUNG BIOPSY. THE PROCEDURE WAS COMPLETED AS PLANNED, AND THE PATIENT WAS DISCHARGED THE SAME DAY. THERE WERE NO REPORTED DEVICE ISSUES OR COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. THE PATIENT¿S SPOUSE INFORMED THE PHYSICIAN THAT THE PATIENT EXPIRED 18 DAYS POST-PROCEDURE. THE REGISTRY STUDY PHYSICIAN DEEMED THE DEATH TO BE NOT RELATED TO THE PROCEDURE, NOT RELATED TO THE ION SYSTEM OR ACCESSORIES, AND PROBABLY RELATED TO THE PATIENT¿S PRE-EXISTING CO-MORBIDITIES, INCLUDING AN 80 PACK-YEAR HISTORY OF SMOKING. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: THE REGISTRY STUDY PHYSICIAN REPORTED THAT THE EVENT WAS PROBABLY RELATED TO THE PATIENT¿S EXISTING CONDITION(S), NOT RELATED TO THE PROCEDURE, NOT RELATED TO ION SYSTEM/ACCESSORIES. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WAS RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY AS PART OF A CLINICAL STUDY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS DISCHARGED HOME THE SAME DAY. THE BIOPSY CONFIRMED A SQUAMOUS CELL CANCER. THE PATIENT DIED 18 DAYS AFTER THE PROCEDURE. THE DEATH WAS ATTRIBUTED TO THEIR MULTIPLE UNDERLYING MEDICAL CO-MORBIDITIES INCLUDING CHRONIC DEMYELINATING POLYNEUROPATHY, HEPATITIS B, IRON DEFICIENCY ANEMIA, COPD/EMPHYSEMA AND MODERATE PULMONARY FIBROSIS. BASED ON THE AVAILABLE DATA THE DEATH WAS DUE TO UNDERLYING MEDICAL CONDITIONS AND ALTHOUGH THERE IS NO CLEAR EVIDENCE OF ASSOCIATION WITH THE INDEX PROCEDURE, THE POSSIBILITY THAT THE EVENT WAS PROCEDURE RELATED CANNOT BE EXCLUDED. THE DEATH WAS IDENTIFIED AS PART OF A CLINICAL STUDY. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21897671&pc=DSI | 3008973940-2025-02874 | 2025-03-14 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-04-24 04:00 | DSI | REVEAL LINQ | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | K132649 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A LINQ DEVICE EXPERIENCED A TERMINAL EPISODE ON (B)(6) 2025, RESULTING IN DEATH. THE PATIENT'S PREVIOUS DEVICE REACHED THE RECOMMENDED REPLACEMENT TIME (RRT), AND IT APPEARED THAT WHEN THE DEVICE WAS REPLACED, THE PATIENT NEVER PERFORMED AN INITIAL MANUAL TRANSMISSION WITH THE REMOTE MONITOR, WHICH WAS REQUIRED TO PAIR WITH THE NEW DEVICE, SO NO TRANSMISSIONS WERE COMING THROUGH. THE ONLY TRANSMISSIONS ON THE NETWORK WERE FROM THE LINQ MOBILE MANAGER AND CARELINK EXPRESS. THE PATIENT RECORDED A SYMPTOM ON (B)(6) 2025, BUT IT WAS NOT SEEN UNTIL THE LINQ DEVICE WAS INTERROGATED BY CARELINK EXPRESS POST-MORTEM. THERE WERE NO ACTIONS TAKEN ON THE MONITOR. IT WAS NOTED THAT THE MONITOR HAD BEEN PLUGGED IN, SUCCESSFULLY UPDATED THE FIRMWARE, AND REGULARLY CHECKED IN, INDICATING THAT THE MONITOR WAS WORKING AS EXPECTED. THE REMOTE MONITOR WAS DISCARDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21898707&pc=OZD | 1220648-2025-27794 | 2025-03-25 04:00 | Death | ABIOMED, INC. | 2025-04-24 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED A STROKE AND SUBSEQUENTLY EXPIRED. THE PATIENT WAS AN ATHLETE WHO CODED WHILE TRAINING FOR A TRIATHLON AND WAS INITIALLY IMPLANTED WITH AN IMPELLA CP PUMP FOR MCS. THE IMPELLA CP SUPPORTED THE PATIENT FOR APPROXIMATELY SIX HOURS BEFORE BEING REMOVED AND ESCALATED TO AN IMPELLA 5.5 DEVICE AT THE TERTIARY CENTER FOR INCREASE IN SUPPORT. TWO DAYS AFTER START OF THE IMPELLA 5.5 SUPPORT, A BRAIN BLEED PROMPTED THE TEAM TO STOP SYSTEMIC ANTICOAGULATION. HOWEVER, TWO DAYS LATER, THE PATIENT EXPIRED. CARE WAS WITHDRAWN AND ORGANS HARVESTED AFTER A TOTAL OF FOUR DAYS ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21899073&pc=KGG | MW5169480 | 2025-04-02 04:00 | Death | CERENOVUS, INC. | 2025-04-24 04:00 | KGG | TRUFILL N-BCA | Leak/Splash; Patient Device Interaction Problem | Abdominal Pain; Infiltration into Tissue; Ischemia; Necrosis; Diminished Pulse Pressure |  |  | 1 | Event Description: ON (B)(6) 2025, PT ADMITTED TO IR FOR COILING AND GLUE PLACEMENT FOR ENDOLEAK. FOLLOWING THE PROCEDURE, THE PATIENT HAD DISCHARGE ORDERS BUT STARTED TO EXPERIENCE ACUTE ABDOMINAL PAIN AND LOST THE PULSE IN HIS LEFT FOOT. PT WAS ADMITTED TO CCU AND CT ANGIOGRAM PERFORMED AND SHOWED HIGH-DENSITY MATERIAL IN THE RIGHT HEPATIC, DISTAL SPLENIC ARTERIES, AND DISTAL SMA. PATIENT WAS TAKEN TO OR EMERGENTLY FOR AORTOGRAM, SMA STENT, LEFT POPLITEAL CUTDOWN AND TIBIAL THROMBECTOMY, LEFT LOWER EXTREMITY ANGIOGRAM, LEFT DORSALIS PEDIS CUT DOWN AND TIBIAL PEDAL THROMBECTOMY. ON (B)(6), PATIENT ALSO HAD AN EXPLORATORY LAP WITH RIGHT HEMICOLECTOMY - FINDINGS INCLUDE ISCHEMIC ILEUM WITH A WELL PERFUSED JEJUNUM AND COLON. ON (B)(6), HE RETURNED TO THE OPERATING ROOM FOR A SECOND LOOK, WHICH INDICATED THAT THE MAJORITY OF THE SMALL BOWEL IS NONVIABLE AND THE GALLBLADDER IS NECROTIC. THERE IS LESS THAN 40CM OF VIABLE BOWEL AND EVEN THESE AREAS HAVE SMALL PATCHES OF ISCHEMIA. IT WAS THE UNDERSTANDING OF NURSING THAT THE GLUE TRAVELED THROUGH THE VASCULAR SYSTEM AND CAUSED THE ISCHEMIC LEG AND NECROTIC BOWEL. WHEN PHYSICIAN SPOKE WITH THE FAMILY, HE MENTIONED THAT HE WAS UNSURE HOW THIS COULD HAVE HAPPENED. PALLIATIVE WAS CONSULTED AND THE FAMILY DECIDED TO TRANSITION TO COMFORT CARE. THE PATIENT PASSED AWAY (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21899087&pc=QBJ | MW5169481 |  | Death | ABBOTT DIABETES CARE INC. | 2025-04-24 04:00 | QBJ | INTEGRATED CONTINUOUS GLUCOSE MONITORING SYSTEM, FACTORY CALIBRATED | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT PASSED AWAY (B)(4); REPORTED BY HCP (HEALTH CARE PROVIDER) DID NOT CONSENT TO FOLLOW UP ALL AVAILABLE INFORMATION IS PROVIDED IN THIS REPORT NO FURTHER CLARIFICATION IS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21899169&pc=OZD | 1220648-2025-27797 | 2025-03-13 04:00 | Death | ABIOMED, INC. | 2025-04-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE DECISION WAS MADE FOR IMPELLA CP PLACEMENT ON A PATIENT IN CARDIOGENIC SHOCK. THE PATIENT WENT IN TO VENTRICULAR FIBRILLATION, AFTER 15 MINUTES OF RESUSCITATION THE DECISION WAS MADE TO END ALL THERAPIES AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE DEVICE HAS BEEN RECEIVED FOR EVALUATION. UPON COMPLETION OF THE INVESTIGATION A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21899358&pc=MNS | 2518422-2025-104714 | 2025-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-04-24 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900220&pc=NPT | 2025587-2025-02974 | 2025-02-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-24 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Dyspnea; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS, 5 MONTHS, AND 18 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO DIFFICULTY BREATHING. TWO DAYS LATER, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS ACUTE HEART FAILURE. Manufacturer Narrative: CORRECTED D.4 AND H.4 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900415&pc=DSQ | 2916596-2025-02500 | 2025-04-05 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2025. THEY WERE MADE COMFORT CARE ON (B)(6) 2025 AFTER VENTRICULAR FIBRILLATION (VF/VFIB) CONTINUOUSLY WITH MORE THAN 55 SHOCKS WITHOUT SUCCESS TO SUSTAIN SINUS RHYTHM. IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF ATRIAL FIBRILLATION PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT AND HAD AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) IMPLANTED PRIOR TO LVAD AS WELL. THE ARRYTHMIA AND DEATH WAS NOT THOUGHT TO HAVE BEEN DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING ARRHYTHMIA, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE IFU ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900432&pc=OZD | 1220648-2025-27813 | 2025-04-04 04:00 | Death | ABIOMED, INC. | 2025-04-24 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT ON IMPELLA 5.5 EXPIRED ON SUPPORT. THE CAUSE OF DEATH WAS CIRCULATORY FAILURE DUE TO INFECTION. THE RELATIONSHIP BETWEEN DEATH AND IMPELLA IS UNKNOWN. ALTHOUGH THE POSSIBILITY OF INFECTION DUE TO THE DEVICE WAS CONSIDERED, THE PATIENT'S UNDERLYING HEART FUNCTION WAS VERY POOR AND THE PATIENT WAS IN A STATE WHERE HE WAS SUSCEPTIBLE TO INFECTION. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS: "THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN." SECTION: WARNINGS AND CAUTIONS: WARNINGS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900482&pc=NPT | 2025587-2025-02975 | 2025-01-21 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-24 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS, 7 MONTHS, AND 11 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO HEART FAILURE. THE PATIENT DIED APPROXIMATELY 3 WEEKS LATER. THE CAUSE OF DEATH WAS HEART FAILURE. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900690&pc=GDW | 3005075853-2025-02967 | 2025-03-01 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-24 04:00 | GDW | PROXIMATE*VASC LIN STAP RELOAD | Failure to Fire; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Hemorrhage/Bleeding | K020779 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STALING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. Manufacturer Narrative: (B)(4). DATE SENT 4/24/2025. D4 BATCH#: UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: PLEASE PROVIDE THE ACCOUNT/FACILITY NAME WAS THERE ANY OTHER ISSUE NOTED WITH THE STAPLE LINE OTHER THAN BLEEDING/OOZING (MALFORMED STAPLES, STAPLE LINE MISSING STAPLES, ETC.)? HOW WAS THE BLEEDING CONTROLLED? HOW MUCH BLOOD WAS LOST (ML)? DID THE PATIENT REQUIRE A TRANSFUSION? COULD YOU PLEASE CLARIFY IF THERE WERE ANY PATIENT CONSEQUENCES? COULD YOU PLEASE CLARIFY IF THERE WERE ANY CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 4/30/2025. CORRECTED DATA: B2, H1. ADDITIONAL INFORMATION RECEIVED: HEALTH IMPACT: DEATH. DURING LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STAPLING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? DID THE PATIENT RECEIVE ANY PREOPERATIVE CHEMOTHERAPY OR RADIATION? WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? WERE THERE ANY ISSUES WITH DEVICE USE/FIRING? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? HOW WAS THE BLEEDING ADDRESSED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WHAT IS THE DATE OF DEATH? WOULD THE SURGEON LIKE TO MEET WITH ETHICON TO DISCUSS THE EVENT? WHAT WAS THE PROXIMAL WIDTH OF THE RIGHT HEPATIC VEIN? DID THE SURGEON INSPECT THE STAPLE LINE BY RELEASING THE DEVICE PRIOR TO DIVIDING THE VESSEL? DID THE SURGEON PLACE A VASCULAR CLAMP ON THE VESSEL PRIOR TO DIVIDING? DID THE SURGEON ENSURE PROXIMAL AND DISTAL CONTROL OF THE VESSEL PRIOR TO PLACING THE STAPLER? WHAT IS THE SURGEONS EXPERIENCE WITH THE DEVICE? WAS IT CONFIRMED THAT THE SURGEON INDEPENDENTLY CLOSED THE CLOSURE TRIGGER AND FIRED THE FIRING TRIGGER SEPARATELY? Manufacturer Narrative: (B)(4). DATE SENT: 5/12/2025. CORRECTED DATA: B1, B2, H1. UPON REVIEW IT WAS IDENTIFIED THAT THIS IS A DUPLICATE REPORT. ALL REPORTED INFORMATION FOR THIS EVENT WAS REPORTED UNDER 3005075853-2025-02969. MOVING FORWARD ALL ADDITIONAL INFORMATION WILL BE FILED UNDER 3005075853-2025-0263.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900692&pc=GDW | 3005075853-2025-02968 | 2025-03-01 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-24 04:00 | GDW | PROXIMATE*VASC LIN STAP RELOAD | Failure to Fire; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Hemorrhage/Bleeding | K020779 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STALING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. Manufacturer Narrative: (B)(4). DATE SENT: 4/30/2025. CORRECTED DATA: B2, H1. ADDITIONAL INFORMATION RECEIVED: HEALTH IMPACT: DEATH DURING LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STAPLING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? DID THE PATIENT RECEIVE ANY PREOPERATIVE CHEMOTHERAPY OR RADIATION? WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? WERE THERE ANY ISSUES WITH DEVICE USE/FIRING? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? HOW WAS THE BLEEDING ADDRESSED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WHAT IS THE DATE OF DEATH? WOULD THE SURGEON LIKE TO MEET WITH ETHICON TO DISCUSS THE EVENT? WHAT WAS THE PROXIMAL WIDTH OF THE RIGHT HEPATIC VEIN? DID THE SURGEON INSPECT THE STAPLE LINE BY RELEASING THE DEVICE PRIOR TO DIVIDING THE VESSEL? DID THE SURGEON PLACE A VASCULAR CLAMP ON THE VESSEL PRIOR TO DIVIDING? DID THE SURGEON ENSURE PROXIMAL AND DISTAL CONTROL OF THE VESSEL PRIOR TO PLACING THE STAPLER? WHAT IS THE SURGEONS EXPERIENCE WITH THE DEVICE? WAS IT CONFIRMED THAT THE SURGEON INDEPENDENTLY CLOSED THE CLOSURE TRIGGER AND FIRED THE FIRING TRIGGER SEPARATELY? Manufacturer Narrative: (B)(4). DATE SENT: 5/12/2025. CORRECTED DATA: B1, B2, H1. UPON REVIEW IT WAS IDENTIFIED THAT THIS IS A DUPLICATE REPORT. ALL REPORTED INFORMATION FOR THIS EVENT WAS REPORTED UNDER 3005075853-2025-02969. MOVING FORWARD ALL ADDITIONAL INFORMATION WILL BE FILED UNDER 3005075853-2025-0263. Manufacturer Narrative: (B)(4). DATE SENT: 4/24/2025. D4 BATCH#: UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: PLEASE PROVIDE THE ACCOUNT/FACILITY NAME WAS THERE ANY OTHER ISSUE NOTED WITH THE STAPLE LINE OTHER THAN BLEEDING/OOZING (MALFORMED STAPLES, STAPLE LINE MISSING STAPLES, ETC.)? HOW WAS THE BLEEDING CONTROLLED? HOW MUCH BLOOD WAS LOST (ML)? DID THE PATIENT REQUIRE A TRANSFUSION? COULD YOU PLEASE CLARIFY IF THERE WERE ANY PATIENT CONSEQUENCES? COULD YOU PLEASE CLARIFY IF THERE WERE ANY CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900697&pc=DYE | 9612164-2025-02050 | 2025-03-27 04:00 | Death | MEDTRONIC MEXICO | 2025-04-24 04:00 | DYE | AVALUS ULTRA BIOPROSTHESIS | Patient Device Interaction Problem; Central Regurgitation | Ventricular Fibrillation; Aortic Valve Insufficiency/ Regurgitation | P170006 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT OF THIS 25MM AORTIC BIOPROSTHETIC VALVE, IT WAS REPORTED THAT AFTER ATTACHING THE VALVE TO A COMPOSITE GRAFT FOR A BENTALL PROCEDURE, THE CENTER OF THE VALVE APPEARED MORE OPEN THAN USUAL. MYOCARDIAL FLUID WAS INJECTED INTO THE GRAFT BUT DID NOT ENTER THE CORONARY DUE TO AORTIC REGURGITATION, AND A CONSIDERABLE AMOUNT OF PRESSURE WAS REQUIRED. THE SURGERY WAS COMPLETED AND AN ULTRASOUND DISCOVERED THERE WAS NO REGURGITATION. IT WAS FURTHER REPORTED THAT HEART-LUNG DETACHMENT WAS ATTEMPTED, HOWEVER VENTRICULAR FIBRILLATION DID NOT SUBSIDE. THE PATIENT REMAINED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WITH IMPELLA. SUBSEQUENTLY, 13 DAYS LATER, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THE DEVICE WAS EXPLANTED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: PRODUCT ANALYSIS: UPON RECEIPT AT MEDTRONIC'S QUALITY LABORATORY, VISUAL EXAMINATION REVEALED THAT THE AVALUS VALVE WAS SUTURED INTO A FLEXIBLE GRAFT. PLEDGETS AND BLUE MONOFILAMENT SUTURES WERE FOUND ON THE DISTAL END OF THE VALVE-GRAFT SPECIMEN. ADDITIONAL PLEDGETS AND BLUE MONOFILAMENT SUTURES WERE FOUND ON THE OUTFLOW OF THE GRAFT SUTURED INTO THE NATIVE ASCENDING AORTA. THE GRAFT APPEARS TO HAVE BEEN CUT AND/OR TORN DURING THE EXPLANT EXPOSING PART OF THE VALVE. DUE TO ITS RECEIPT STATE, THE LEAFLETS WILL BE DESIGNATED AS L1, L2, L3 ARBITRARILY AS LEAFLET IDENTIFICATION THROUGH VISUAL ASSESSMENT COULD BE PERFORMED. ALL LEAFLETS WERE STIFF DUE TO HOST TISSUE GROWTH ON THE OUTFLOW. LEAFLET 3 SHOWED A CUT TO THE HOST OVERGROWTH WITH REMNANT PIECE RESTING ON TOP OF THE LEAFLET; THE EXPOSED LEAFLET APPEARED INTACT. ALL POSTS SHOWED VARYING DEGREES OF HOST TISSUE OVERGROWTH, WITH ONE POST SHOWING LESS COVERAGE THAN THE OTHER TWO. THE CONDITION OF THE COMMISSURES COULD NOT BE DETERMINED DUE THE HOST TISSUE OVERGROWTH ON ALL COMMISSURES. ALL LEAFLETS WERE IN THE CLOSED POSITION WITH A SMALL GAP BETWEEN LEAFLET 2 AND LEAFLET 3. TAN AND MAROON VEGETATIVE APPEARING HOST TISSUE COVERED THE OUTFLOW OF ALL LEAFLETS AND COMMISSURES EXTENDING TO THE OUTFLOW RAIL AND POSTS. TAN VEGETATIVE APPEARING HOST TISSUE APPEARED TO HAVE ADHERED TO THE INNER LUMEN OF THE GRAFT. TAN AND MAROON VEGETATIVE APPEARING HOST TISSUE WAS NOTED ON THE INFERIOR COAPTIVE AREAS AND MARGIN OF ATTACHMENTS OF ALL LEAFLETS WITH A LAYER EXTENDING TO THE BELLY OF LEAFLET 3. HOST TISSUE OVERGROWTH ON THE LEAFLETS APPEARED TO BE VEGETATION. REMNANTS OF HOST MYOCARDIAL TISSUE WERE OBSERVED AROUND THE GRAFT AS WELL AS PROXIMAL AND DISTAL TO THE VALVE-GRAFT. COAGULATED BLOOD WAS FOUND ON THE GRAFT SURROUNDING ONE OF THE CORONARY ARTERIES. THE INFLOW ASPECT SHOWED A RESTRICTED INFLOW ORIFICE AREA. D9: UPDATED. H3: DEVICE EVALUATED? UPDATED H6: UPDATED THE CODES MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900698&pc=GDW | 3005075853-2025-02969 | 2025-03-01 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-04-24 04:00 | GDW | PROXIMATE*VASC LIN STAP RELOAD | Failure to Fire; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Hemorrhage/Bleeding | K020779 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STALING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. Manufacturer Narrative: (B)(4). DATE SENT: 4/30/2025. CORRECTED DATA: B2, H1. ADDITIONAL INFORMATION RECEIVED: HEALTH IMPACT: DEATH. DURING LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STAPLING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? DID THE PATIENT RECEIVE ANY PREOPERATIVE CHEMOTHERAPY OR RADIATION? WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? WERE THERE ANY ISSUES WITH DEVICE USE/FIRING? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? HOW WAS THE BLEEDING ADDRESSED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WHAT IS THE DATE OF DEATH? WOULD THE SURGEON LIKE TO MEET WITH ETHICON TO DISCUSS THE EVENT? WHAT WAS THE PROXIMAL WIDTH OF THE RIGHT HEPATIC VEIN? DID THE SURGEON INSPECT THE STAPLE LINE BY RELEASING THE DEVICE PRIOR TO DIVIDING THE VESSEL? DID THE SURGEON PLACE A VASCULAR CLAMP ON THE VESSEL PRIOR TO DIVIDING? DID THE SURGEON ENSURE PROXIMAL AND DISTAL CONTROL OF THE VESSEL PRIOR TO PLACING THE STAPLER? WHAT IS THE SURGEONS EXPERIENCE WITH THE DEVICE? WAS IT CONFIRMED THAT THE SURGEON INDEPENDENTLY CLOSED THE CLOSURE TRIGGER AND FIRED THE FIRING TRIGGER SEPARATELY? Manufacturer Narrative: (B)(4). DATE SENT: 5/12/2025. CORRECTED DATA: B1, B2, H1. UPON REVIEW IT WAS IDENTIFIED THAT THIS IS A DUPLICATE REPORT. ALL REPORTED INFORMATION FOR THIS EVENT WAS REPORTED UNDER 3005075853-2025-02969. MOVING FORWARD ALL ADDITIONAL INFORMATION WILL BE FILED UNDER 3005075853-2025-0263. Manufacturer Narrative: (B)(4). DATE SENT 4/24/2025. D4 BATCH # UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: PLEASE PROVIDE THE ACCOUNT/FACILITY NAME. WAS THERE ANY OTHER ISSUE NOTED WITH THE STAPLE LINE OTHER THAN BLEEDING/OOZING (MALFORMED STAPLES, STAPLE LINE MISSING STAPLES, ETC.)? HOW WAS THE BLEEDING CONTROLLED? HOW MUCH BLOOD WAS LOST (ML)? DID THE PATIENT REQUIRE A TRANSFUSION? COULD YOU PLEASE CLARIFY IF THERE WERE ANY PATIENT CONSEQUENCES? COULD YOU PLEASE CLARIFY IF THERE WERE ANY CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21900972&pc=KRH | 2025587-2025-02977 | 2025-03-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-04-24 04:00 | KRH | CONTOUR 3D | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Pneumonia; Insufficient Information | K101212 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 4 YEARS AND 1 MONTH POST IMPLANT OF THIS 32MM TRICUSPID ANNULOPLASTY RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE ANNULOPLASTY RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT REPORTED THE CAUSE OF DEATH AS PNEUMONIA. IT WAS FURTHER REPORTED THAT THE DEVICE WAS NOT RELATED TO THE EVENT. NO OTHER ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CHANGED THE EVENT DESCRIPTION AND DEVICE RELATEDNESS OF THIS PREVIOUSLY REPORTED EVENT. THERE IS NO EVIDENCE TO SUGGEST THE DEVICE CAUSED OR CONTRIBUTED TO A DEATH. UPDATED DATA: B.5: EVENT DESCRIPTION H.6: CODING MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21901005&pc=OZD | 1220648-2025-27816 | 2025-03-31 04:00 | Death | ABIOMED, INC. | 2025-04-24 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT A PATIENT HAD A CORONARY ARTERY BYPASS GRAFT AND A VENTRICULAR TACHYCARDIA (VT) ABLATION THE DAY BEFORE. HE WAS INCREASINGLY HEMODYNAMICALLY UNSTABLE AND SHOWED REDUCED LEFT VENTRICLE (LV) FUNCTION. THE DECISION WAS MADE FOR IMPELLA 5.5. THE IMPELLA WAS IMPLANTED WITHOUT COMPLICATIONS. DURING SUPPORT IT WAS REPORTED THAT TWO DAYS AFTER THE IMPELLA IMPLANTATION, A LEFT VENTRICLE THROMBUS WAS SEEN ON COMPUTED TOMOGRAPHY (CT). THIS WAS NOT PRESENT/NOT VISIBLE/NOT KNOWN BEFORE IMPLANTATION. ADDITIONAL INFORMATION WAS RECEIVED THAT LATER INDICATED THE FOLLOWING: THE PHYSICIAN REPORTED ABOUT A NEW CT SCAN. THERE WAS NO LV THROMBUS AND THEY STATED THAT THE FIRST DIAGNOSIS WAS WRONG. THE PATIENT HAD PROLONGED VT EPISODES THE DAY BEFORE. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VT/VF/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21901230&pc=DSQ | 2916596-2025-02497 | 2023-09-22 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO HOSPITAL POLICY, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CUSTOMER. THE HEARTMATE 3 LVAS IFU, REV. B, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21901847&pc=DSQ | 2916596-2025-02524 | 2025-04-06 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RUPTURED THORACIC AORTIC ANEURYSM. THE DEATH WAS EXPECTED SINCE THE PATIENT DID NOT HAVE AN INTENTION OF TRANSPLANTING, THE CASE WAS ESSENTIALLY MANAGED AS DESTINATION THERAPY. THE DEVICE WAS NOT EXPLANTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC ANEURYSM WAS GENETIC IN ORIGIN. THERE WAS NO DEVICE INVOLVEMENT IN THE PATIENT'S DEATH AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION D1: BRAND NAME CORRECTED. SECTION D4: MODEL AND CATALOG NUMBERS CORRECTED, UDI NUMBER CORRECTED. SECTION E1: CUSTOMER SITE WAS (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21901899&pc=DSQ | 3007042319-2025-00828 | 2024-04-16 04:00 | Death | HEARTWARE, INC. | 2025-04-24 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Obstruction of Flow; Pumping Problem | Fatigue; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT PRESENTED WITH MANY LOW FLOW ALARMS AND BELOW NORMAL POWER AS WELL AS BEING DESCRIBED AS EXHAUSTED. IMAGING TESTS, INCLUDING AN ECHOCARDIOGRAM AND A COMPUTED TOMOGRAPHY SCAN, WERE CONDUCTED. THE PRESENCE OF A THROMBUS INSIDE THE OUTFLOW GRAFT WAS NOTED. THE PATIENT WAS IN THE CATHETER LAB FOR A PROCEDURE TO DILATE AND STENT THE OUTFLOW GRAFT, THE ATTEMPT TO RE-PERMEABILIZE THE GRAFT WAS UNSUCCESSFUL DESPITE THE USE OF SEVERAL STENTS. SUBSEQUENTLY, THE PATIENT DIED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION AND A CORRECTION TO H6. ADDITIONAL PRODUCTS: OUTFLOW GRAFT D4: EXPIRATION DATE: 28-FEB-2021 / LOT#: 2002675065 D6A: IMPLANTED DATE: (B)(6) 2021 H4: MFG DATE: 28-FEB-2019 H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. ADDITIONAL PRODUCTS: D4: SERIAL OR LOT#: (B)(6) H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: (B)(6) AND THE ASSOCIATED OUTFLOW GRAFT (LOT NO. 2002675065) WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE AVAILABLE AUTOLOGS REPORT REVEALED A DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS THROUGHOUT THE ANALYZED PERIOD, LEADING TO PARAMETERS BELOW NORMAL OPERATING RANGE, AND ONE HUNDRED (100) LOW FLOW ALARMS LOGGED SINCE (B)(6) 2025. RAW LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED LOW FLOW/POWER EVENT WAS CONFIRMED; HOWEVER, THE REPORTED OUTFLOW GRAFT OCCLUSION COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT AN ATTEMPT TO STENT THE OUTFLOW GRAFT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY EXPIRED. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF THROMBUS EVENTS. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW/POWER EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SP EED. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM OUTFLOW GRAFT D4: MODEL #: 1125 / CATALOG #: 1125 / EXPIRATION DATE: UNKNOWN / SERIAL OR LOT#: UNKNOWN D9: NO H3: N0 H4: MFG DATE: UNKNOWN H5: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21902012&pc=DSQ | 2916596-2025-02489 | 2025-01-11 05:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING VARIOUS TYPES OF ORGAN FAILURES/DYSFUNCTIONS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21902064&pc=DSQ | 2916596-2025-02716 | 2025-04-14 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PULMONARY FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT NO FURTHER DETAILS WOULD BE SHARED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21902297&pc=DSQ | 2916596-2025-02339 | 2024-11-07 05:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Coagulation Disorder; Hemorrhage/Bleeding; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. THEY HAD PROFOUND COAGULOPATHY WITH MASSIVE DIFFUSE BLEEDING LEADING TO MULTISYSTEM ORGAN FAILURE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS FUNCTIONING AS EXPECTED AND HAD NO ISSUES. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS BLEEDING AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21902305&pc=LGW | 3006705815-2025-02973 | 2025-04-06 04:00 | Death | ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC) | 2025-04-24 04:00 | LGW | OCTRODE TRIAL LEAD KIT, 60CM LENGTH | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING AN SCS TRIAL LEAD IMPLANT ON (B)(6) 2025, THE PATIENT WAS HOSPITALIZED THE WEEKEND OF (B)(6) 2025 DUE TO A BLOOD CLOT AND LATER PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: A PATIENT DEATH WAS REPORTED TO ABBOTT. EVENT DETAILS DESCRIBE HEALTH ISSUES UNRELATED TO THE IMPLANTED SYSTEM. ADDITIONALLY, NO SCS SYSTEM COMPLAINTS WERE REPORTED NOR WERE IMPLANTED PRODUCTS RETURNED FOR ANALYSIS. ALL EVENT INFORMATION PERTAINING TO THIS DEVICE HAS BEEN REVIEWED AND NO PRODUCT INVESTIGATION IS NECESSARY AT THIS TIME AS THE CIRCUMSTANCES OF THE EVENT HAVE NOT BEEN ATTRIBUTED TO THE IMPLANTED SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21902577&pc=DSQ | 2916596-2025-02665 | 2025-04-08 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO MULTIPLE SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE WOULD NOT RETURN. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED SECONDARY TO MULTIPLE SYSTEM ORGAN FAILURE. THE PATIENT OUTCOME WAS NOT DEVICE OR THERAPY RELATED. IT WAS FURTHER COMMUNICATED THAT THE PATIENT¿S DEVICE HAD OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING VARIOUS TYPES OF ORGAN FAILURES/DYSFUNCTIONS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21902962&pc=OZD | 1220648-2025-27826 | 2025-04-03 04:00 | Death | ABIOMED, INC | 2025-04-24 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Pumping Stopped; Malposition of Device | Atrial Fibrillation; Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN ETHNICITY) WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED A POSITIONING ISSUE, SUCTION EVENTS, ATRIAL FIBRILLATION WITH CARDIOVERSION BY IMPLANTABLE DEVICE AND A PUMP STOP WITH A RESTART. THE PATIENT WOULD EVENTUALLY EXPIRE. THE PATIENT PRESENTED WITH ACUTE ON CHRONIC HEART FAILURE EXACERBATION STATUS POST RIGHT HEART CATHETERIZATION. PATIENT WAS ON INOTROPIC SUPPORT AND AFTERLOAD REDUCTION DRIPS NOW WITH INTRA-AORTIC BALLOON PUMP (IABP) BEING EVALUATED FOR LEFT VENTRICULAR ASSIST DEVICE (LVAD). THERE WERE ELEVATED PULMONARY ARTERY PRESSURES DESPITE IABP, AND THE DECISION WAS MADE TO IMPLANT AN IMPELLA 5.5 DEVICE FOR IMPROVED LEFT VENTRICLE OFF-LOADING AND PULMONARY ARTERY PRESSURES OPTIMIZATION. APPROXIMATELY SEVEN DAYS AFTER START OF THE SUPPORT, THE IMPELLA 5.5 PUMP REQUIRED REPOSITIONING DUE TO PLACEMENT SIGNAL LOW ALARM AS WELL AS ONE SUCTION, SECURED AT 45CM. THE PATIENT HAD BEEN CARDIOVERTED DUE TO ATRIAL FIBRILLATION. THE FOLLOWING DAY THERE WAS ANOTHER SUCTION EVENT THE PATIENT ENCOUNTERED WHEN LEANING FORWARD. THE PERFORMANCE LEVEL WAS DECREASED FROM P-9 TO P-8 WITH RESOLUTION. TWO DAYS LATER, IT WAS NOTED THE PATIENT WAS UNFORTUNATELY NO LONGER A CANDIDATE FOR ADVANCED THERAPIES GIVEN PERSISTENT FEVER AND MULTI-ORGAN DYSFUNCTION. THE PATIENT REMAINED UNDER SUPPORTED DESPITE BEING ON PERFORMANCE LEVEL P-8 ON NEAR MAXIMAL PHARMACOLOGICAL SUPPORT. IMPELLA SUPPORT CONTINUED PER THE PLAN. HOWEVER, LATER THAT AFTERNOON, THE IMPELLA 5.5 DEVICE PUMP STOPPED ALARM TRIGGERED THREE TIMES, THE PUMP STOPPED AND RESTARTED AT PERFORMANCE LEVEL P-8. IT WAS FURTHER EXPLAINED THAT THE EVENT OCCURRED WHILE THE PATIENT WAS SITTING IN THE CHAIR BEING WHEELED OUTSIDE, WHICH WAS ABORTED; PATIENT TAKEN BACK TO THE ROOM THE PATIENT WAS ALERT AND TALKING VISITING WITH FAMILY. ALL DRIPS REMAIN WHERE THEY WERE ALL DAY. THE PHYSICIAN NOTED THERE IS NO BACK UP PLAN IF THE PUMP STOPS AGAIN AND FAILS TO RESTART AS THE PATIENT IS A DO NOT RESUSCITATE STATUS AND THERE IS NO ESCALATING CARE OTHER THAN MAINTAINING THE CURRENT CHEMICAL SUPPORT. THE FOLLOWING DAY IT WAS NOTED NO ADDITIONAL PUMP ALARMS; HOWEVER ADDITIONAL PRESSORS WERE ADDED TO HELP MAINTAIN BLOOD PRESSURE WHILE THE PATIENT MANAGED AFFAIRS PRIOR TO PLAN FOR COMFORT CARE LATER IN THE DAY, WHICH SUBSEQUENTLY OCCURRED, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, TEMPORARY PUMP STOP, AND VF/VT/AF/AT HAS BEEN COMPLETED. THERE WAS NO PROBLEM FOUND DURING THE CASE WITH PUMP RELATIVE TO PUMP STOP EVENT. THE DEVICE FUNCTIONED/OPERATED TO SPECIFICATION. THE ROOT CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL INFORMATION. THE ROOT CAUSE OF THE ARRHYTHMIA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPACT TO PERFORMANCE AND OUTCOME AS A RESULT OF THE IMPELLA RELATED EVENTS (E.G., PUMP STOP ALARMS AND STOPPING OF THE PUMP) IS UNKNOWN. THERE IS NOT ENOUGH TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR. THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF THE ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. SECTION: DIRECT AORTIC INSERTION, ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21903287&pc=LWS | 1028232-2025-01981 | 2025-04-07 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-04-24 04:00 | LWS | IFORIA 5 VR-T PROMRI DF-1 | Failure to Interrogate | No Clinical Signs, Symptoms or Conditions | P050023 |  | 1 | Event Description: IT WAS REPORTED THAT THE ICD WAS EXPLANTED AFTER THE PATIENT DIED. IT WAS NOT INTERROGATABLE.
PLEASE NOTE THIS ICD IS AFFECTED BY A FIELD SAFETY CORRECTIVE ACTION, BIO-LQC, INITIATED IN MARCH 2021. Manufacturer Narrative: THE ICD AND THE LEAD WERE RETURNED FOR ANALYSIS. PRIOR TO THE ANALYSIS OF THE DEVICES, THE QUALITY DOCUMENTS ACCOMPANYING THE MANUFACTURING PROCESS FOR THESE DEVICES WERE RE-INVESTIGATED. ALL PRODUCTION STEPS WERE PERFORMED ACCORDINGLY, AND IN PARTICULAR THE FINAL ACCEPTANCE TESTS PROVED THE DEVICES FUNCTIONS TO BE AS SPECIFIED. ANALYSIS OF THE ICD: UPON RECEIPT, THE ICD WAS VISUALLY INSPECTED. THE INSPECTION REVEALED A SPOT OF MOLTEN TITANIUM ON THE ICD HOUSING. THE ICD WAS SUBJECTED TO AN ELECTRICAL ANALYSIS, CONFIRMING THAT THE DEVICE COULD NOT BE INTERROGATED. THEREFORE, BASED ON THE OBSERVED SYMPTOMS, IT IS REASONABLE TO ASSUME THAT A SHOCK DELIVERY INTO AN EXTERNAL SHORT CIRCUIT LED TO A DAMAGE OF THE ELECTRICAL MODULE. DUE TO THE DAMAGE OF THE ELECTRICAL MODULE THE DEVICE COULD NOT BE INTERROGATED AND THE MEMORY CONTENT OF THE DEVICE WAS NOT RESTORABLE. ANALYSIS OF THE LEAD: UPON RECEIPT, THE LEAD UNDER COMPLAINT WAS SUBJECTED TO AN EXTENSIVE ANALYSIS. THE PERFORMANCE OF THE LEAD WAS SCRUTINIZED, INCLUDING A VISUAL, MECHANICAL AND ELECTRICAL INSPECTION. IN THE COURSE OF THE ANALYSIS, MULTIPLE SIGNS OF ABRASION WERE NOTED ALONG THE LEAD BODY. THE INSULATION WAS STILL INTACT IN THESE REGIONS. ADDITIONALLY EXTRACTION DAMAGES WERE FOUND SUCH AS MULTIPLE CUTS IN THE INSULATION AND A TORN UP LEAD BODY. THE LEAD BODY WAS COVERED IN CALCIFIED TISSUE INDICATING SEVERE INGROWTH OF THE LEAD BODY. FURTHERMORE THE SVC SHOCK COIL SHOWED A MARK OF ARCING OVER 42 CM DISTAL TO THE IS-1 CONNECTOR PIN. CONCLUSION: THE ANALYSIS REVEALED MARKS OF ARCING OVER ON THE ICD HOUSING AND ON THE SVC SHOCK COIL OF THE LEAD. BASED ON THE LOCATION OF THE MARK ON THE LEAD IT IS HIGHLY LIKELY THAT IT OCCURRED DURING THE EXTRACTION PROCEDURE. THE ICD WAS STILL ACTIVATED, MOST LIKELY RESULTING IN A SHOCK DELIVERED INTO AN EXTERNAL SHORT CIRCUIT WHEN THE SHOCK COIL CONTACTED THE HOUSING DURING THE EXPLANTATION. AS A RESULT, THE ELECTRICAL MODULE WAS DAMAGED AND THE ICD WAS NO LONGER INTERROGATABLE. THE ANALYSIS DID NOT REVEAL ANY DAMAGES OF THE DEVICES WHICH MIGHT HAVE CONTRIBUTED TO THE PATIENT¿S DEATH. FURTHERMORE THERE WAS NO SIGN OF ANY MATERIAL OR MANUFACTURING PROBLEM OF THE DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21903450&pc=DSQ | 2916596-2025-02522 | 2025-04-05 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Lethargy | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 SECONDARY TO CARDIAC ARREST. THE PATIENT ACUTELY DECOMPENSATED WITH LOW MEAN ARTERIAL PRESSURES (MAPS), LETHARGY, AND NO VENTRICULAR ASSIST DEVICE (VAD) ALARMS, AND WAS SENT BACK TO THE INTENSIVE CARE UNIT (ICU). THE FAMILY DECIDED TO PLACE THE PATIENT ON COMFORT MEASURES AND NOT INTUBATE THE PATIENT OR ESCALATE CARE. SYMPTOMS WERE UNKNOWN AT THE TIME OF DEATH. IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF VENTRICULAR TACHYCARDIA PRIOR TO IMPLANT. THE ARRHYTHMIA IN THIS EVENT WAS NOT THOUGHT TO BE DEVICE OR THERAPY RELATED. AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) WAS IMPLANTED PRIOR TO VAD. THE DEVICE OPERATED AS INTENDED Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21903519&pc=FRN | 1416980-2025-02472 | 2024-12-14 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-04-24 04:00 | FRN | NOVUM IQ | Filling Problem | Cardiac Arrest; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | K211122 |  | 1 | Event Description: IT WAS REPORTED THAT A NOVUM IQ LARGE VOLUME PUMP UNDER-INFUSED DURING PATIENT INFUSION. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: RELATED COMPLAINTS: DEVICE 2: 1416980-2025-02471 AND DEVICE 3: 1416980-2025-02469. ADDITIONAL INFORMATION: B1, B2, B3, B5, D10, F10, AND H1. B5: UPON FOLLOW UP INFORMATION, IT WAS FURTHER REPORTED THREE (3) PUMPS WERE IN USE ON ONE PATIENT. THE PATIENT REPORTEDLY WAS RECEIVING LEVOPHED, DOBUTAMINE, VASOPRESSIN AND IV FLUIDS ON THE NOVUM PUMPS. ON AN UNSPECIFIED DATE, THE PATIENT REPORTEDLY EXPERIENCED A ¿DECLINE IN CONDITION.¿ THIS PATIENT WAS TRANSFERRED TO A ¿HIGHER LEVEL OF CARE.¿ ON THE DATE OF THE EVENT, THE PATIENT REPORTEDLY ¿CODED¿ AND DIED. IT WAS NOT REPORTED WHICH PUMP WAS RUNNING WHICH MEDICATED SOLUTION OR WHICH PUMP UNDER-INFUSED RELATED TO THE PATIENT¿S REPORTED DEATH. ADDITIONALLY, THE CLINICAL ENGINEERING DEPARTMENT OF THE REPORTING FACILITY EVALUATED THE PUMPS ONSITE AND ¿NO ISSUES WERE IDENTIFIED WITH THE THREE PUMPS INVOLVED.¿ THIS COMPLIANT IS FOR DEVICE ONE (1). SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3, H6 AND H11. H11: THE DEVICE WAS RECEIVED FOR EVALUATION. VISUAL INSPECTION DID NOT IDENTIFY ANY ABNORMALITIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED CONDITION. THE SET UNDERWENT FUNCTIONAL TESTING BY RUNNING A SHORT, SIMULATED THERAPY OF 25ML AT 25ML/HOUR AND PERFORMED ACCORDING TO PRODUCT SPECIFICATIONS. PRE-SERVICING AND FLOW ACCURACY TESTING WAS PERFORMED AND NO ISSUES NOTED. THE EVENT HISTORY WAS REVIEWED, AND THE REPORTED INFUSION WAS CONFIRMED. THE REPORTED DRUG VASOPRESSIN WAS IDENTIFIED IN THE LOGS AS A PRIMARY INFUSION AT THE RATE OF 12ML/HOUR WITH VOLUME TO BE INFUSED OF 60 ML ON THE EVENT DATE. ONE DOWNSTREAM OCCLUSION ALARM WAS FOUND ON THE EVENT DATE. NO AIR IN LINE ALARMS OCCURRED. NO MOTOR STALL ALERTS WERE IDENTIFIED. STANDBY ALARMS WERE NOT IDENTIFIED ON OR AROUND THIS INFUSION. THE INFUSION DID NOT RUN TO COMPLETION USER STOPPED THE PUMP AND SHUTDOWN THE DEVICE. A SERVICE HISTORY REVIEW WAS PERFORMED AND REVEALED THAT THE DEVICE HAS NO PREVIOUS SERVICE EVENTS; THEREFORE, SERVICING DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. THE REPORTED CONDITION WAS NOT VERIFIED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21903831&pc=DSQ | 2916596-2025-02327 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-04-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypoxia; Septic Shock; Cardiogenic Shock; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO ACUTE HYPOXIC RESPIRATORY FAILURE, MIXED WITH CARDIOGENIC SHOCK AND POTENTIAL COMPONENT OF SEPTIC SHOCK. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, SEPSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, THE IFU ALSO LISTS INFECTION AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21903976&pc=BZD | 2518422-2025-104758 | 2025-04-04 04:00 | Death | RESPIRONICS, INC. | 2025-04-24 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21904125&pc=DRA | 2124215-2025-26173 | 2025-03-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-24 04:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation of Vessels |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH THE PATIENT PASSED AWAY. AFTER THE PROCEDURE THE PATIENT EXPERIENCED BLOOD LOSS FOLLOWING THE PROCEDURE. THEY UNDERWENT A CT SCAN WHICH IDENTIFIED A PUNCTURE "HIGH UP IN THE ILIAC ARTERY" AND ANOTHER IN THE SUPERIOR EPIGASTRIC ARTERY. COILING WAS ATTEMPTED TO COMBAT THE BLEEDING BUT ULTIMATELY THE PATIENT PASSED AWAY. THE PUNCTURES WERE FOUND ON THE LEFT SIDE OF THE BODY, THE FARADRIVE SHEATH WAS INSERTED IN THE RIGHT GROIN. THE SHEATH WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS.Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH THE PATIENT PASSED AWAY. AFTER THE PROCEDURE THE PATIENT EXPERIENCED BLOOD LOSS FOLLOWING THE PROCEDURE. THEY UNDERWENT A CT SCAN WHICH IDENTIFIED A PUNCTURE "HIGH UP IN THE ILIAC ARTERY" AND ANOTHER IN THE SUPERIOR EPIGASTRIC ARTERY. COILING WAS ATTEMPTED TO COMBAT THE BLEEDING BUT ULTIMATELY THE PATIENT PASSED AWAY. THE PUNCTURES WERE FOUND ON THE LEFT SIDE OF THE BODY, THE FARADRIVE SHEATH WAS INSERTED IN THE RIGHT GROIN. THE SHEATH WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. CORRECTION TO THE INITIAL MDR IN BLOCK H6: PATIENT CODES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21904325&pc=OPR | 2032493-2025-90108 | 2025-04-07 04:00 | Death | MICROVENTION, INC. | 2025-04-24 04:00 | OPR | WEB SL | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P170032 |  | 1 | Event Description: AS REPORTED: A WEB 7X3 WAS SUCCESSFULLY IMPLANTED INTO A PREVIOUSLY HIGH GRADE RUPTURED (X2) ANEURYSM. CONTRAST STASIS WAS OBSERVED IN THE ANEURYSM AND THE WEB SHOWED GOOD COMPRESSION. NO ENCROACHMENT ON BRANCH VESSELS AND POST ANGIO IMAGES LOOKED GOOD. THE PROCEDURE CONCLUDED AND THE PATIENT LEFT THE ANGIO SUITE. AT SOME POINT LATER THAT NIGHT OR EARLY MORNING THE HEMATOMA HAD EXPANDED AND THE RUPTURED ANEURYSM REBLED. NO FURTHER SURGICAL INTERVENTIONS WERE PERFORMED.Event Description: AS REPORTED TO MICROVENTION: "NOTHING ELSE WAS ABLE TO BE DONE POST REBLEED, AS THE HEMATOMA EXPANDED TOO QUICKLY AND IT WAS A HIGH GRADE. THE PATIENT UNFORTUNATELY PASSED." THE PROCEDURE OCCURRED EARLY MORNING. THE PATIENT PASSED AWAY ONE OR TWO DAYS AFTER THE PROCEDURE. THE OP REPORT HAS BEEN REQUESTED, BUT NOT RECEIVED. Manufacturer Narrative: AS THE LOT NUMBER WAS NOT PROVIDED, A SEARCH FOR PRODUCTION-RELATED NCRS COULD NOT BE PERFORMED. THE DEVICE WAS IMPLANTED AND THEREFORE NOT AVAILABLE FOR RETURN AND INVESTIGATION BY THE MANUFACTURER. THE INVESTIGATION IS ONGOING AND UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL MDR WILL BE SUBMITTED. Manufacturer Narrative: THE PHYSICAL DEVICE WAS NOT AVAILABLE FOR EVALUATION TO INVESTIGATE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. SUPPLEMENTAL IMAGING WAS ALSO UNAVAILABLE FOR REVIEW; WITHOUT IMAGING, THE INVESTIGATION CANNOT VERIFY THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EXAMINE THE CAUSE OF THE REPORTED EVENT. THIS INFORMATION MAY BE UPDATED IF ADDITIONAL INFORMATION IS PROVIDED AT A LATER DATE. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21904882&pc=PCU | 3005099803-2025-01661 | 2017-10-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-04-24 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING AXIOS STENTS AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS THROUGH THE ARTICLE "PREDICTIVE FACTORS FOR EARLY MORTALITY AFTER EUS-GUIDED GASTROENTEROSTOMY IN MALIGNANT GASTRIC OUTLET OBSTRUCTION" BY HYUK LEE ET AL PER THE ARTICLE, BETWEEN OCTOBER 2017 AND OCTOBER 2022, A STUDY OF PATIENTS SUFFERING FROM MALIGNANT GASTRIC OUTLET OBSTRUCTION (MGOO) WHO UNDERGONE ENDOSCOPIC ULTRASOUND-GUIDED GASTROENTEROSTOMY (EUS-GE) WITH AXIOS STENTS WAS MADE. ACCORDING TO THE STUDY, SEVEN PATIENTS DIED WITHIN THIRTY DAYS AND CLINICAL DETERIORATION DUE TO DIRECT OR INDIRECT EFFECTS OF THE PROCEDURE CANNOT BE EXCLUDED. Manufacturer Narrative: BLOCK B3: THE EXACT DATE OF THE EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY DATE OF OCTOBER 1, 2017, IS USED FOR THE ESTIMATED DATE OF EVENT BETWEEN OCTOBER 2017 AND OCTOBER 2022. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WERE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: HYUK LEE ET AL. "PREDICTIVE FACTORS FOR EARLY MORTALITY AFTER EUS-GUIDED GASTROENTEROSTOMY IN MALIGNANT GASTRIC OUTLET OBSTRUCTION" ON BEHALF OF ENDOSC INT OPEN 2025; 13: A24749802., HTTPS://DOI.ORG/10.1055/A-2474-9802 BLOCK H6: IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF 7 MORTALITIES WITHIN 30 DAYS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21905539&pc=QEW | 3005168196-2025-00135 | 2023-09-04 04:00 | Death | PENUMBRA, INC. | 2025-04-24 04:00 | QEW | INDIGO SYSTEM SEPARATOR 12 | Adverse Event Without Identified Device or Use Problem | Hemoptysis; Great Vessel Perforation |  |  | 1 | Event Description: IT WAS REPORTED THAT A POTENTIAL ADVERSE EVENT INVOLVING A PENUMBRA PRODUCT OCCURRED DURING A PROCEDURE. BASED ON THE LIMITED INFORMATION AVAILABLE AT THIS TIME, THE PRODUCT MAY HAVE BEEN A PENUMBRA SEPARATOR; AND THE ADVERSE EVENT MAY HAVE RESULTED IN PATIENT HEMOPTYSIS AND DEATH.Event Description: IT WAS REPORTED THAT A POTENTIAL ADVERSE EVENT INVOLVING A PENUMBRA PRODUCT OCCURRED DURING A PROCEDURE. BASED ON THE LIMITED INFORMATION AVAILABLE AT THIS TIME, THE PRODUCT MAY HAVE BEEN A PENUMBRA SEPARATOR; AND THE ADVERSE EVENT MAY HAVE RESULTED IN PATIENT HEMOPTYSIS AND DEATH. ON 05JUN25, PENUMBRA RECEIVED A VOLUNTARY MEDWATCH REPORT SUBMITTED TO FDA BY A FORMER PENUMBRA EMPLOYEE. BASED ON THE REPORT, ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT THE PATIENT SUFFERED A FATAL PULMONARY ARTERY PERFORATION DURING THE PROCEDURE INVOLVING AN INDIGO SYSTEM CAT12 ASPIRATION CATHETER AND AN INDIGO SYSTEM SEP12 SEPARATOR. Manufacturer Narrative: A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED AS THE DEVICE WAS NOT RETURNED. THE PRODUCT LOT NUMBER WAS NOT PROVIDED; THEREFORE, THE MANUFACTURING RECORDS COULD NOT BE REVIEWED. FROM THE LIMITED INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTIONS ARE BEING UPDATED BASED ON ADDITIONAL INFORMATION RECEIVED ON 05-JUN-2025 FROM A VOLUNTARY MEDWATCH REPORT: 1. SECTION A. A1) PATIENT IDENTIFIER AND A2) AGE OF TIME OF EVENT. 2. SECTION B. B5) DESCRIBE EVENT OR PROBLEM. 3. SECTION D. D1) BRAND NAME. 4. SECTION E. INITIAL REPORTER. 5. SECTION H. H6) HEALTH EFFECT CLINICAL CODE 1. THE INFORMATION IN THE VOLUNTARY MEDWATCH REPORT SUBMITTED TO FDA WAS PROVIDED BY A FORMER PENUMBRA EMPLOYEE. PRIOR TO RECEIVING THE VOLUNTARY MEDWATCH REPORT FROM FDA, PENUMBRA ATTEMPTED TO CONFIRM THE EVENT AND ASSOCIATED DETAILS WITH THE HOSPITAL. THE HOSPITAL HAS NOT RESPONDED TO PENUMBRA'S REQUEST FOR INFORMATION. THEREFORE, THE EVENT AND DETAILS OF THE EVENT HAVE NOT BEEN CONFIRMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21905613&pc=NPT | 2135147-2025-02149 | 2025-02-15 05:00 | Death | ABBOTT MEDICAL | 2025-04-24 04:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device | Dyspnea; Heart Block; Aortic Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 35MM NAVITOR VISION VALVE WAS CHOSEN FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE ANNULAR DIMENSIONS OF THE NATIVE VALVE WAS 87.4MM. DURING PROCEDURE, THE FINAL IMPLANT DEPTH AT 80% WAS 5/5MMMM AND THE PATIENT HAD HEART BLOCK AND THE VALVE WAS IMPLANTED AT A DEPTH OF 8MM. AFTER THE PROCEDURE, A DUAL-CHAMBER PERMANENT PACEMAKER WAS IMPLANTED IN THE PATIENT. THREE WEEKS AFTER, THE PATIENT HAD HEART FAILURE SYMPTOMS SUCH AS MODERATE SHORTNESS OF BREATH AND AORTIC INSUFFICIENCY (AI). ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO HEART FAILURE (HF) SYMPTOMS. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS THE VALVE MIGRATED DEEPER CAUSE TORE RISK AI AND HF BUT STATED NO IMAGING WAS DONE TO CHECK VALVE DEPTH.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS THE VALVE MIGRATED DEEPER CAUSING TORRENTIAL AI AND HF BUT STATED NO IMAGING WAS DONE TO CHECK VALVE DEPTH. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT THREE WEEKS AFTER NAVITOR IMPLANT THE PATIENT HAD HEART FAILURE SYMPTOMS SUCH AS MODERATE SHORTNESS OF BREATH AND AORTIC INSUFFICIENCY. THE PATIENT EXPIRED DUE TO HEART FAILURE SYMPTOMS. IT WAS REPORTED THAT THE CAUSE OF PATIENT DEATH WAS DUE TO THE VALVE MIGRATED DEEPER CAUSING TORRENTIAL AORTIC INSUFFICIENCY AND HEART FAILURE BUT NO IMAGING WAS DONE TO CHECK VALVE DEPTH. INFORMATION FROM FIELD INDICATED THAT DURING PROCEDURE THE PATIENT HAD HEART BLOCK AND THE VALVE WAS IMPLANTED AT A DEPTH OF 8 MM. A DUAL-CHAMBER PERMANENT PACEMAKER WAS IMPLANTED IN THE PATIENT AFTER THE PROCEDURE. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. BASED ON THE MEDICAL REVIEW PERFORMED AT ABBOTT, REVIEW OF PRE PROCEDURAL CT DEMONSTRATED COMPLETE ABSENCE OF CALCIFICATION OF LEAFLETS/ANNULUS. PRE OR POST PROCEDURAL IMAGING WAS NOT PROVIDED FOR FURTHER ANALYSIS TO DETERMINE CAUSE OF REPORTED EVENT. THE PATIENT HAD HISTORY OF RIGHT BUNDLE BRANCH BLOCK WHICH COULD HAVE BEEN EXACERBATED BY THE PROCEDURE LEADING TO THE HEART BLOCK BEING REPORTED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF PATIENT DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE PATIENT EFFECTS OF DYSPNEA AND AORTIC INSUFFICIENCY APPEAR TO BE CASCADING EFFECT OF DEVICE MIGRATION. IT IS POSSIBLE THAT DEVICE MIGRATION IS RELATED TO THE PROCEDURAL AND ANATOMICAL CONDITIONS (IMPLANT DEPTH DEEPER THAN 3 MM AND ABSENCE OF CALCIFICATION). HOWEVER, THIS CANNOT BE CONFIRMED. THE REPORTED SURGICAL INTERVENTION IS RESULT OF CASE-SPECIFIC CIRCUMSTANCE AS PERMANENT PACEMAKER WAS IMPLANTED FOR HEART BLOCK DURING PROCEDURE. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21905905&pc=NIK | 3004209178-2025-07157 | 2025-04-02 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-04-24 04:00 | NIK | AMPLIA MRI QUAD CRT-D SURESCAN | Device Alarm System; Failure to Interrogate; Pacing Problem | Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS UNABLE TO BE INTERROGATED, AND SURFACE ELECTROCARDIOGRAM AND MONITORS DID NOT SHOW ANY INDICATION OF PACING. IT WAS ALSO REPORTED THAT THE MOBILE PROGRAMMER APPLICATION WAS UNABLE TO IDENTIFY THE CRT-D. IT WAS NOTED THAT THE PATIENT WAS NOT ENROLLED IN THE REMOTE MONITORING NETWORK WEBSITE. WHEN A MAGNET WAS APPLIED, THERE WAS NO MAGNET RESPONSE TONE. THE PHYSICIAN BELIEVED THE CRT-D BATTERY WAS JUST DEPLETED. IT WAS NOTED THAT NO PRODUCT PERFORMANCE ISSUES LED TO THE BATTERY DEPLETION. NO FURTHER ACTIONS WERE TAKEN TO DATE. THE CRT-D REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE CLINICIAN WAS WAITING TO PERFORM A DEVICE CHANGE-OUT AS THE PATIENT WAS BEING ASSESS FOR LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. HOWEVER, IT WAS NOTED THAT PATIENT LATER PASSED AWAY DUE TO HEART FAILURE BEFORE THEY COULD RECEIVE A REPLACEMENT OR LVAD. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.