## Feb'25
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21390484&pc=KDI | 2521402-2025-00010 | 2025-01-20 05:00 | Death | B. BRAUN AVITUM AG | 2025-02-15 05:00 | KDI | DIALOG+® | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: AS REPORTED BY THE USER FACILITY: THE CUSTOMER REPORTED THAT A PATIENT HAD DIED DURING DIALYSIS. IN THE ADDITIONAL INFORMATION OBTAINED, THE CAUSE OF DEATH WAS NAMED AS CARDIOPULMONARY ARREST. SOME INFORMATION ABOUT THE PATIENT'S MEDICAL HISTORY DESCRIBED END-STAGE RENAL DISEASE, A HISTORY OF DIABETES, DEEP VEIN THROMBOSIS AND A PARTIAL FOOT AMPUTATION. A TECHNICAL SAFETY INSPECTION OF THE MACHINE WAS PERFORMED. THE TSI WAS PASSED, CONFIRMING THAT THE MACHINE WORKS AS INTENDED. THE INVESTIGATION OF THE RECEIVED MACHINE DATA RECORD SHOWS THAT AT; 1:01:30 (TREND TIME) ALL MACHINE SELF-TESTS WERE PASSED SUCCESSFULLY. 1:54:43 (TREND TIME) AND 1:59:58 (TREND TIME) AUTOMATIC BLOOD PRESSURE MEASUREMENTS (ABPM) WERE TAKEN, THE FIRST WITH THE RESULT 175/163 MMHG, PULSE 49 BPM AND THE SECOND WITH THE RESULT 139/121 MMHG, PULSE 68 BPM. AS THE DIASTOLIC VALUE OF THE FIRST MEASUREMENT WAS TOO HIGH, THE DEVICE TRIGGERED THE ALARM ABPM: DIA PRESS. IS TOO HIGH'. AFTER THESE TWO MEASUREMENTS, NO FURTHER ABPM MEASUREMENTS WERE PERFORMED. DURING THE PREPARATION, THE SELF-TESTS WERE PASSED AND THE THERAPY WAS STARTED. 1:59:58 (TREND TIME) MAIN FLOW WAS ACTIVATED. 2:26:24 (TREND TIME) CHANGE OF FINAL CONDUCTIVITY FROM 138 MMOL/L TO 150 MMOL/L. APPROXIMATELY 5 MINUTES LATER, THE ARTERIAL PRESSURE (PA) SUDDENLY DROPPED FROM APPROXIMATELY -170 MMHG TO -240 MMHG, WHEREUPON THE DEVICE TRIGGERED THE ARTERIAL PRESSURE - LOWER LIMIT' ALARM AND SWITCHED TO PATIENT-SAFE MODE. THE ALARM WAS CONFIRMED 15 SECONDS LATER. HOWEVER, THE PA SUDDENLY DROPPED FURTHER TO -430 MMHG, WHEREUPON THE DEVICE TRIGGERED THE ALARMS ARTERIAL PRESSURE - LOWER LIMIT' AND ARTERIAL PRESSURE - LOWER LIMIT (SUP)' AND SWITCHED TO PATIENT-SAFE MODE. ALL ALARMS WERE CONFIRMED. 2:46:30 (TREND TIME) FINAL CONDUCTIVITY WAS CHANGED TO 148,5 MMOL/L. 3:05:30 (TREND TIME) MINIMUM ULTRAFILTRATION (UF) WAS ACTIVATED. TWO MINUTES LATER, 3:07:42 (TREND TIME) FINAL CONDUCTIVITY WAS SET TO 147,0 MMOL/L. TEN SECONDS LATER, 3:07:52 (TREND TIME) THE VENOUS PRESSURE (PV) DECREASED FROM APPROXIMATELY 250 MMHG TO 160 MMHG, WHEREUPON THE DEVICE TRIGGERED THE ALARM "VENOUS PRESSURE - LOWER LIMIT - CHECK ACCESS" AND SWITCHED TO PATIENT-SAFE MODE. SIX SECONDS LATER, THE ALARM WAS ACKNOWLEDGED. THEN PV INCREASED TO 207 MMHG AND EXCEEDED THE UPPER LIMIT OF 193 MMHG. THEREFORE, AT 3:08:34 (TREND TIME) THE MACHINE TRIGGERED THE ALARM "VENOUS PRESSURE - UPPER LIMIT". SOME SECONDS LATER, 3:08:42 (TREND TIME) THIS ALARM WAS ACKNOWLEDGED. 3:10:04 (TREND TIME) THE PV DECREASED AGAIN TO CIRCA 140 MMHG AND THE DEVICE TRIGGERED THE ALARM "VENOUS PRESSURE - LOWER LIMIT - CHECK ACCESS". THREE SECONDS LATER THE ALARM WAS CONFIRMED. 3:10:39 (TREND TIME) VENOUS PRESSURE INCREASED TO 203 MMHG CAUSING THE ALARM "VENOUS PRESSURE - UPPER LIMIT". EIGHT SECONDS LATER, THE ALARM WAS CONFIRMED. 3:17:17 (TREND TIME), CIRCA 1:14 HOURS AFTER THERAPY START, THE ULTRAFILTRATION VOLUME WAS CHANGED TO 2500 ML AND MINIMUM UF WAS DEACTIVATED. 3:29:34 (TREND TIME), CIRCA 12 MINUTES LATER, THE FINAL CONDUCTIVITY WAS CHANGED TO 145,5 MMOL/L. 3:36:08 (TREND TIME), CIRCA 7 MINUTES LATER THE UF VOLUME WAS CHANGED TO 2300 ML. 3:44:08 (TREND TIME). APPROXIMATELY 1:41 HOURS AFTER THERAPY START, MINIMUM UF WAS AGAIN ACTIVATED. 3:45:48 (TREND TIME): THE RED DETECTOR (RDV) OF THE SAFETY AIR DETECTOR (SAD) SWITCHED FROM "RED" TO "NO MORE RED". THIS IS AN INDICATION THAT THE BLOOD LINE WAS REMOVED. THREE SECONDS LATER, THE VENOUS PRESSURE DECREASED FROM CIRCA 230 MMHG TO APPROXIMATELY 145 MMHG AND THEREFORE THE ALARM "VENOUS PRESSURE - LOWER LIMIT - CHECK ACCESS" WAS TRIGGERED. THIS ALARM WAS ACKNOWLEDGED THREE SECONDS LATER. 3:46:01 (TREND TIME), THE PA DECREASED TO -427 MMHG AND THE MACHINE TRIGGERED THE ALARMS "ARTERIAL PRESSURE - LOWER LIMIT" AND "ARTERIAL PRESSURE - LOWER LIMIT (SUP)". 3:47:59 (TREND TIME), THE MACHINE TRIGGERED THE ALARM "BLOOD PUMP IS STATIONARY (SUP)" BECAUSE THE BLOOD PUMP HAD NOT BEEN RUNNING FOR 2 MINUTES. THE ALARM WAS ACKNOWLEDGED, BUT THE VENOUS PRESSURE INCREASED TO 343 MMHG AND THE MACHINE TRIGGERED THE ALARMS "VENOUS PRESSURE - UPPER LIMIT" AND "VENOUS PRESSURE - UPPER LIMIT (SUP)". THE BLOOD PUMP REMAINED STOPPED FOR THE REST OF THIS MACHINE DATA RECORD. 3:55:29 (TREND TIME), THAT IS APPROXIMATELY 1:52 HOURS AFTER THERAPY START, THE DEVICE WAS SWITCHED FROM THERAPY TO END OF THERAPY PHASE. THE ANALYSIS OF THE MACHINE DATA RECORD AND ALL RELEVANT PARAMETERS DOCUMENTED THEREIN REVEALED NO ANOMALIES OR INDICATIONS OF A MALFUNCTION OF THE DEVICE. THERE IS NO PRODUCT DEVIATION. IN GENERAL, ALL SAFETY RELEVANT FUNCTIONS ARE TESTED DURING THE SELF-TESTS IN PREPARATION. WITHOUT HAVING PASSED THE SELF-TESTS A THERAPY START IS NOT POSSIBLE. IF DURING THERAPY A SAFETY RELEVANT MALFUNCTION OCCURS, THE MACHINE TRIGGERS AN ALARM AND SWITCHES TO PATIENT-SAFE MODE. IN CONCLUSION, BASED ON THE ANALYSIS AND EVALUATION OF THE REPORTED AND PROVIDED INFORMATION THERE WAS NOT MALFUNCTION OF THE DEVICE AND THE DEVICE HAS NOT LED TO THE PATIENT'S DEATH. THIS IS IN LINE WITH THE CUSTOMER'S STATEMENT. THIS CASE IS CLASSIFIED AS NOT REPORTABLE TO THE COMPETENT AUTHORITY BY THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21382410&pc=MSD | 3006260740-2025-00652 | 2021-11-04 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-02-14 05:00 | MSD | EQUISTREAM ALPHACURVE DL CATHETER 14.5FR, 28CM | Fluid/Blood Leak; Material Split, Cut or Torn | Exsanguination | K090101 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE THE PATIENT'S CAREGIVER WAS CHANGING THE DRESSING AROUND THE DIALYSIS CATHETER, THE TUBING WAS INADVERTENTLY CUT BENEATH ONE OF THE PORTS. IT WAS FURTHER REPORTED THAT THE PATIENT BEGAN BLEEDING OUT THROUGH THE CUT CATHETER END AND COLLAPSED. REPORTEDLY, THE CAREGIVER CALLED 911 AND THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM. THE BLEEDING COULD NOT BE STOPPED UNTIL THE TUBE WAS CLAMPED IN THE EMERGENCY ROOM. DESPITE ALL EFFORTS INCLUDING TRANSFUSION, THE PATIENT COULD NOT BE RESUSCITATED. THE PATIENT EXPIRED DUE TO EXSANGUINATION.Event Description: IT WAS REPORTED THAT WHILE THE PATIENT'S CAREGIVER WAS CHANGING THE DRESSING AROUND THE DIALYSIS CATHETER, THE TUBING WAS INADVERTENTLY CUT BENEATH ONE OF THE PORTS. IT WAS FURTHER REPORTED THAT THE PATIENT BEGAN BLEEDING OUT THROUGH THE CUT CATHETER END AND COLLAPSED. REPORTEDLY, THE CAREGIVER CALLED 911 AND THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM. THE BLEEDING COULD NOT BE STOPPED UNTIL THE TUBE WAS CLAMPED IN THE EMERGENCY ROOM. DESPITE ALL EFFORTS INCLUDING TRANSFUSION, THE PATIENT COULD NOT BE RESUSCITATED. THE PATIENT EXPIRED DUE TO EXSANGUINATION.Event Description: IT WAS REPORTED THAT WHILE THE PATIENT'S CAREGIVER WAS CHANGING THE DRESSING AROUND THE DIALYSIS CATHETER, THE TUBING WAS INADVERTENTLY CUT BENEATH ONE OF THE PORTS. IT WAS FURTHER REPORTED THAT THE PATIENT BEGAN BLEEDING OUT THROUGH THE CUT CATHETER END AND COLLAPSED. REPORTEDLY, THE CAREGIVER CALLED 911 AND THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM. THE BLEEDING COULD NOT BE STOPPED UNTIL THE TUBE WAS CLAMPED IN THE EMERGENCY ROOM. DESPITE ALL EFFORTS INCLUDING TRANSFUSION, THE PATIENT COULD NOT BE RESUSCITATED. THE PATIENT EXPIRED DUE TO EXSANGUINATION. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H11: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED CUT, FLUID LEAK, AND PATIENT DEATH, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FROM THE REPORT, THE EVENT IS STATED TO BE DUE TO INADVERTENT CUTTING OF THE CATHETER. THEREFORE, IT APPEARS THE MOST LIKELY ROOT CAUSE FOR THE REPORTED MATERIAL CUT, LEAKAGE, AND PATIENT DEATH IS SHARP INSTRUMENT DAMAGE DUE TO USER ERROR. IT IS UNKNOWN WHETHER PATIENT OR PRODUCT ISSUES CONTRIBUTED TO THE REPORTED EVENT. AT THIS TIME, A RELATIONSHIP BETWEEN THE CATHETER AND THE REPORTED MATERIAL CUT AND DEATH CANNOT BE CONFIRMED. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (UNIQUE IDENTIFIER (UDI) #), D5, E3, G3, H4, H6 (DEVICE, COMPONENT, METHOD). SECTION A THROUGH F: THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: B2 (DATE OF DEATH). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21382462&pc=MIH | 3002808486-2025-00027 | 2019-04-03 04:00 | Death | WILLIAM COOK EUROPE | 2025-02-14 05:00 | MIH | ZENITH ALPHA ENDOVASCULAR GRAFT |  | Hemorrhage/Bleeding; Ruptured Aneurysm |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO WCE-1713: ZENITH ALPHA THORACIC POST MARKET RETROSPECTIVE STUDY SYNOPSIS: PROCEDURE IMAGING REVEALED A TYPE IV ENDOLEAK WHICH INCREASED 7 DAYS POST-PROCEDURE, LEADING TO AORTIC RUPTURE AND DEATH. ON (B)(6) 2019, THE PATIENT WAS ROUTINELY TREATED FOR AORTIC DISSECTION TYPE A WITH PLACEMENT OF 4 ZTA PROXIMAL COMPONENTS AND 2 NON-COOK GRAFTS. PROCEDURE IMAGING DATA INDICATES NO EVIDENCE OF DEVICE ISSUE AND PATENT THORACIC AND ABDOMINAL FALSE LUMENS, WITH SOURCE OF FLOW NOTED AS UNKNOWN. THE POST-OP SCAN ON (B)(6) 2019 REVEALED A TYPE IV ENDOLEAK IN THE THORACIC PORTION OF THE STENT RE-FEEDING THE FALSE CHANNEL WITH A PERIAOTIC HAEMATOMA (ALREADY PRESENT PRE-OPERATIVELY ON THE SCAN OF (B)(6) 2019). ON (B)(6) 2019, THE PATIENT EXPERIENCED BOWEL ISCHEMIA, TREATED WITH BOWEL RESECTION. THIS EVENT IS NOTED AS RELATED TO THE TREATED AORTIC DISEASE AND THE STUDY PROCEDURE, NOTING THE RELATION TO THE PROCEDURE IS JUSTIFIED BY THE FACT THAT THIS ISCHAEMIA OCCURRED AFTER THE PROCEDURE. THE CT SCAN SHOWS HAEMOPERITONEUM DUE TO RELEASE OF THE RIGHT ILIAC SUTURE ASSOCIATED WITH AN ISCHAEMIC ASPECT OF THE SIGMOID. STUDY DOCUMENTATION INDICATES TYPE IV ENDOLEAK, WAS STABLE ON (B)(6) 2019. ON (B)(6) 2019, THE PATIENT EXPERIENCED AORTIC RUPTURE AT THE STENTED SEGMENT, TREATED WITH BALLOON ANGIOPLASTY OF ALL ZTA DEVICES. THIS EVENT WAS NOTED AS RELATED TO THE STUDY DEVICE AND THE TREATED AORTIC DISEASE, AND IT LED TO THE PATIENT¿S DEATH. THE EMERGENCY SCAN SHOWED AN INCREASE IN THE ENDOLEAK ALREADY PRESENT ON (B)(6) 2019 ASSOCIATED WITH A NON-OCCLUSIVE AORTIC THROMBOSIS EXTENDING TO THE FEMORAL AXES AND THE SUPERIOR MESENTERIC ARTERY. DECISION TO PERFORM EMERGENCY SURGERY FOR RUPTURED PERIPROSTHETIC HAEMATOMA (ANGIOPLASTY OF THE ZTA THORACIC ENDOPROSTHESIS). PATIENT OUTCOME: THE PATIENT DIED ON (B)(6) 2019, 7 DAYS POST-PROCEDURE. THE CAUSE OF DEATH IS NOTED AS AORTIC RUPTURE. THE DEATH WAS CONSIDERED RELATED TO PREEXISTING CONDITION PRIOR TO PROCEDURE, PRIMARY DISEASE PROGRESSION, AND PRODUCT COMPLICATION.Event Description: NO ADDITIONAL INFORMATION REGARDING THE PATIENT AND/OR EVENT HAS BEEN RECEIVED SINCE THE PREVIOUS MEDWATCH REPORT WAS SENT. Manufacturer Narrative: MANUFACTURER REF# (B)(4). BLANK FIELDS ON THIS FORM INDICATE THE INFORMATION IS UNKNOWN OR UNAVAILABLE. D3) DENMARK, E1) FRANCE, G1) DENMARK AND G4) SIMILAR TO DEVICE MARKETED UNDER PMA/510(K): P140016. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: MANUFACTURER¿S REF# (B)(4). SUMMARY OF INVESTIGATIONAL FINDINGS: DURING A POST MARKET CLINICAL FOLLOW-UP (PMCF) STUDY COOK BECAME AWARE OF THE EVENT OF A TYPE IV ENDOLEAK WHICH INCREASED 7 DAYS POST-PROCEDURE AND LED TO AORTIC RUPTURE AND DEATH. ON (B)(6) 2019, THE PATIENT AS A CLASSIFICATION OF 4 (A PATIENT WITH SEVERE SYSTEMIC DISEASE THAT IS A CONSTANT THREAT TO LIFE) WAS ROUTINELY TREATED FOR AORTIC DISSECTION TYPE A WITH PLACEMENT OF 4 ZTA PROXIMAL COMPONENTS (FROM ZONE 1 TO ZONE 5) AND 2 NON-COOK GRAFTS (TO ZONE 9 (INFRARENAL ABDOMINAL AORTA)). CAROTID-CAROTID BYPASS AND AORTO-BILIAC BYPASS WERE ALSO PERFORMED. THE LOCATION OF THE DISSECTION WAS FROM ZONE 2 TO ZONE 10. PRE-PROCEDURE IMAGING DATA HAS NOT YET BEEN ENTERED. PROCEDURE IMAGING DATA INDICATES NO EVIDENCE OF DEVICE ISSUES AND PATENT THORACIC AND ABDOMINAL FALSE LUMENS, WITH SOURCE OF FLOW NOTED AS ENDOLEAK. THE POST-OP SCAN ON (B)(6) 2019 REVEALED A TYPE IV ENDOLEAK (CURRENT COMPLAINT) IN THE THORACIC PORTION OF THE STENT RE-FEEDING THE FALSE CHANNEL WITH A PERIAORTIC HEMATOMA (ALREADY PRESENT PRE-OPERATIVELY ON THE SCAN OF (B)(6) 2019). THE ENDOLEAK WAS STABLE ON (B)(6) 2019. THE SAME DAY THE PATIENT EXPERIENCED BOWEL ISCHEMIA (RELATED COMPLAINT), TREATED WITH BOWEL RESECTION. THIS EVENT IS NOTED AS RELATED TO THE TREATED AORTIC DISEASE AND THE STUDY PROCEDURE, NOTING ¿THE RELATION TO THE PROCEDURE IS JUSTIFIED BY THE FACT THAT THIS ISCHEMIA OCCURRED AFTER THE PROCEDURE. THE EVENT WAS NOTED AS NOT TO OCCUR DUE TO DEVICES DEFICIENCY. THE CT (COMPUTER TOMOGRAPHY) SCAN SHOWED HEMOPERITONEUM DUE TO RELEASE OF THE RIGHT ILIAC SUTURE ASSOCIATED WITH AN ISCHEMIC ASPECT OF THE SIGMOID. ON (B)(6) 2019, THE THORACIC DRAIN (2 L) SUDDENLY EXTERIORIZED. THE EMERGENCY SCAN SHOWED AN INCREASE IN THE ENDOLEAK ALREADY PRESENT ON (B)(6) 2019 ASSOCIATED WITH A NON-OCCLUSIVE AORTIC THROMBOSIS EXTENDING TO THE FEMORAL AXES AND THE SUPERIOR MESENTERIC ARTERY. THE PATIENT EXPERIENCED AORTIC RUPTURE (RUPTURE OF PERIAORTIC HEMATOMA) AT THE STENTED SEGMENT, TREATED WITH BALLOON ANGIOPLASTY OF ALL ZTA DEVICES. THIS EVENT WAS NOTED AS RELATED TO PREEXISTING CONDITION PRIOR TO PROCEDURE, PRIMARY DISEASE PROGRESSION AND DEVICE COMPLICATION, AND IT LED TO THE PATIENT¿S DEATH. REVIEW OF THE DEVICE HISTORY RECORDS GAVE NO INDICATION OF THE DEVICES BEING PRODUCED OUT OF SPECIFICATION. CLINICAL ASSESSMENT WAS PERFORMED BY THE MEDICAL ADVISOR DUE TO THE PATIENT´S DEATH. THIS STUDY INVOLVES RETROSPECTIVE DATA COLLECTION. NO IMAGING WAS PROVIDED FOR THE INVESTIGATION, AND BASED ON THE PROVIDED INFORMATION AND CLINICAL ASSESSMENT IT WAS NOT POSSIBLE TO DETERMINE THE CAUSE OF THE ENDOLEAK. HOWEVER, TYPE IV ENDOLEAKS ARE A SPECIFIC TYPE OF ENDOLEAK THAT ARISES FROM THE NATURAL POROSITY OF SOME GRAFT MATERIALS, LEADING TO A TRANSIENT BLOOD LEAK. ACCORDING TO THE INSTRUCTIONS FOR USE, ENDOLEAK IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH ZTA STENT GRAFT OR THE IMPLANTATION PROCEDURE. REPORTED PREOPERATIVE PERIAORTIC HEMATOMA OFTEN ASSOCIATED WITH ACUTE AORTIC DISSECTIONS AND IT'S A COMPLICATION THAT CAN LEAD TO SIGNIFICANT CLINICAL ISSUES AND MORTALITY. SINCE ZTA IS NOT INDICATED FOR TREATMENT OF DISSECTION AND SINCE SAFETY AND EFFECTIVENESS OF THE ZTA HAVE NOT BEEN EVALUATED IN THE PATIENTS WITH DISSECTIONS, IT CAN NOT BE RULED OUT THAT THE USE OF THE COMPLAINT DEVICES OUTSIDE OF THE INTENDED USE COULD BE A CONTRIBUTING FACTOR TO THE REPORTED RUPTURE. COOK MEDICAL WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21382474&pc=MSD | 3006260740-2025-00653 | 2022-02-25 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-02-14 05:00 | MSD | GLIDEPATH | Fluid/Blood Leak | Cardiac Arrest; Exsanguination | K051748 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A SCHEDULED HEMODIALYSIS TREATMENT THROUGH THE HEMODIALYSIS CENTRAL VENOUS CATHETER, THE NURSE IDENTIFIED A POOL OF BLOOD OF APPROXIMATELY 100-300ML UNDER THE PATIENT'S CHAIR DURING THE FIRST SEVERAL MINUTES OF TREATMENT INITIATION. IT WAS FURTHER REPORTED THAT THE BLOOD PUMP WAS STOPPED, AND IT WAS DISCOVERED THAT THE BLOOD LOSS WAS ORIGINATING FROM THE VENOUS BLOODLINE/VENOUS LUMEN CONNECTION SITE. REPORTEDLY, THE PATIENT'S CAROTID PULSE WAS PALPABLE, HEMODIALYSIS TREATMENT WAS TERMINATED, AND BLOOD WAS RETURNED USING ARTERIAL LUMEN OF THE CATHETER. THE PATIENT BECAME UNRESPONSIVE, THE EMERGENCY RESPONSE SERVICES WAS CALLED, AND CPR WAS INITIATED. UPON ARRIVAL, THE EMERGENCY RESPONSE SERVICES CONTINUED RESUSCITATIVE EFFORTS WHILE TRANSPORTING THE PATIENT TO THE HOSPITAL. THE PATIENT REPORTEDLY EXPIRED SHORTLY AFTER.Event Description: IT WAS REPORTED THAT DURING A SCHEDULED HEMODIALYSIS TREATMENT THROUGH THE HEMODIALYSIS CENTRAL VENOUS CATHETER, THE NURSE IDENTIFIED A POOL OF BLOOD OF APPROXIMATELY 100-300ML UNDER THE PATIENT'S CHAIR DURING THE FIRST SEVERAL MINUTES OF TREATMENT INITIATION. IT WAS FURTHER REPORTED THAT THE BLOOD PUMP WAS STOPPED, AND IT WAS DISCOVERED THAT THE BLOOD LOSS WAS ORIGINATING FROM THE VENOUS BLOODLINE/VENOUS LUMEN CONNECTION SITE. REPORTEDLY, THE PATIENT'S CAROTID PULSE WAS PALPABLE, HEMODIALYSIS TREATMENT WAS TERMINATED, AND BLOOD WAS RETURNED USING ARTERIAL LUMEN OF THE CATHETER. THE PATIENT BECAME UNRESPONSIVE, THE EMERGENCY RESPONSE SERVICES WAS CALLED, AND CPR WAS INITIATED. UPON ARRIVAL, THE EMERGENCY RESPONSE SERVICES CONTINUED RESUSCITATIVE EFFORTS WHILE TRANSPORTING THE PATIENT TO THE HOSPITAL. THE PATIENT REPORTEDLY EXPIRED SHORTLY AFTER. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FLUID LEAK AND PATIENT DEATH, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED FLUID LEAK AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21382968&pc=BTT | 9611451-2025-00137 | 2025-01-16 05:00 | Death | FISHER & PAYKEL HEALTHCARE LTD | 2025-02-14 05:00 | BTT | FISHER & PAYKEL HEALTHCARE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Low Oxygen Saturation |  |  | 1 | Event Description: ON 16 JANUARY 2025, A DISTRIBUTOR IN JAPAN REPORTED ON BEHALF OF A HEALTHCARE FACILITY THAT A PT101 AIRVO 2 HUMIDIFIER (AIRVO 2) GENERATED AN ALARM DURING PATIENT USE. THE HEALTHCARE FACILITY ALSO REPORTED THAT THE PATIENT'S SPO2 DECREASED AND A DEVICE WAS REPLACED. IT WAS FURTHER REPORTED THAT THE PATIENT LATER DECEASED AFTER AN UNSPECIFIED PERIOD OF TIME. FISHER & PAYKEL HEALTHCARE (F&P) HAS REQUESTED FOR THE AIRVO 2 DEVICE TO BE RETURNED FOR EVALUATION AND IS SEEKING FURTHER INFORMATION FROM THE HEALTHCARE FACILITY. F&P WILL PROVIDE A FOLLOW UP REPORT UPON COMPLETION OF OUR INVESTIGATION.Event Description: ON 16 JANUARY 2025, A DISTRIBUTOR IN JAPAN REPORTED ON BEHALF OF A HEALTHCARE FACILITY THAT A PT101 AIRVO 2 HUMIDIFIER (AIRVO 2) GENERATED AN ALARM DURING PATIENT USE. FURTHER INFORMATION PROVIDED BY THE HEALTHCARE FACILITY ON 18 FEBRUARY 2025 STATED THAT THE PATIENT STARTED THERAPY ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT WAS TRANSFERRED FROM THE ICU TO THE HOSPITAL WARD. IT WAS REPORTED THAT ON (B)(6) 2025, THE AIRVO 2 GENERATED AN ALARM AT 12:31AM AND THE PATIENT'S SPO2 DECREASED TO 94%. IT WAS REPORTED THAT THE SUBJECT DEVICE WAS THEN REPLACED. IT WAS FURTHER REPORTED BY THE HEALTHCARE FACILITY THAT ON (B)(6) 2025 THE PATIENT PASSED FROM SEPTIC SHOCK WITH HEMORRHAGIC SHOCK AS THE MEDICAL CAUSE OF DEATH. Manufacturer Narrative: (B)(4) [CORRECTED DATA: B5,D9,G3,G6,H1,H2,H3,H6] [D4,G4: PT101JP IS NOT SOLD IN THE USA BUT IT IS SIMILAR TO A PRODUCT WHICH IS SOLD IN THE USA. THE 510(K) FOR THE SIMILAR PRODUCT IS K131895.] PRODUCT BACKGROUND: THE AIRVO 2 DEVICE IS AN ACTIVE HUMIDIFIER WITH INTEGRATED FLOW GENERATOR THAT DELIVERS HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES TO SPONTANEOUSLY BREATHING PATIENTS THROUGH A VARIETY OF PATIENT INTERFACES. IT'S INTENDED USE IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES. THE AIRVO 2 DEVICE SHOULD NOT BE USED FOR LIFE SUPPORT PURPOSES, AND APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES. METHOD: THE SUBJECT DEVICE WAS RECEIVED AT FISHER & PAYKEL HEALTHCARE (F&P) IN NEW ZEALAND WHERE IT WAS INSPECTED BY A TRAINED F&P INVESTIGATION ENGINEER. F&P'S INVESTIGATION IS BASED ON THE INFORMATION PROVIDED BY THE HEALTHCARE FACILITY, F&P'S EVALUATION OF THE SUBJECT DEVICE, PREVIOUS INVESTIGATIONS OF SIMILAR COMPLAINTS, AND F&P'S KNOWLEDGE OF THE PRODUCT. RESULTS: VISUAL INSPECTION OF THE SUBJECT DEVICE REVEALED NO SIGNS OF EXTERNAL DAMAGE. THE HEALTHCARE FACILITY STATED THAT THE PATIENT STARTED THERAPY ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT WAS TRANSFERRED FROM THE ICU TO THE HOSPITAL WARD. IT WAS REPORTED THAT ON (B)(6) 2025, THE AIRVO 2 GENERATED AN ALARM AT 12:31AM AND THE PATIENT'S SPO2 DECREASED TO 94%. IT WAS REPORTED THAT THE SUBJECT DEVICE WAS THEN REPLACED. IT WAS FURTHER REPORTED BY THE HEALTHCARE FACILITY THAT ON (B)(6) 2025 THE PATIENT PASSED FROM SEPTIC SHOCK WITH HEMORRHAGIC SHOCK AS THE MEDICAL CAUSE OF DEATH. CONCLUSION: ON THE BASIS OF THE REPORTED INFORMATION, THERE IS NO CONCLUSIVE EVIDENCE OF A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH, AND THE SUBJECT DEVICE. DURING THE MANUFACTURING PROCESS, QUALITY CONTROL MEASURES ENSURE EACH MANUFACTURED AIRVO 2 MEETS SPECIFICATION. THESE QUALITY CONTROL MEASURES ARE PERFORMED AT THE END OF THE FINAL ASSEMBLY PROCESS ON 100% OF THE MANUFACTURED UNITS. ANY UNIT THAT FAILS ANY OF THESE TESTS WILL BE REJECTED. THE SUBJECT DEVICE THEREFORE WOULD HAVE MET THE REQUIRED SPECIFICATIONS AT THE TIME OF RELEASE. THE USER INSTRUCTIONS WHICH ACCOMPANY THE AIRVO 2 PROVIDE GUIDANCE ON TROUBLESHOOTING ALARMS GENERATED BY THE DEVICE. THE USER INSTRUCTIONS ALSO STATE THE FOLLOWING: "ALARM CONDITIONS THAT AFFECT OXYGEN DELIVERY REQUIRE AN IMMEDIATE RESPONSE TO ASSESS THE PATIENT'S SATURATION LEVELS" "THE UNIT IS NOT INTENDED FOR LIFE SUPPORT" "USE CONTINUOUS OXYGEN MONITORING ON PATIENTS WHO WOULD DESATURATE SIGNIFICANTLY IN THE EVENT OF DISRUPTION TO THEIR OXYGEN SUPPLY." "APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES" "NEVER OPERATE THE UNIT IF IT IS NOT WORKING PROPERLY" "NEVER OPERATE THE UNIT IF IT HAS BEEN DROPPED OR DAMAGED". Manufacturer Narrative: (B)(6). [D4,G4: PT101JP IS NOT SOLD IN THE USA BUT IT IS SIMILAR TO A PRODUCT WHICH IS SOLD IN THE USA. THE 510(K) FOR THE SIMILAR PRODUCT IS K131895.]. FISHER & PAYKEL HEALTHCARE (F&P) HAS REQUESTED FOR THE PT101 AIRVO 2 HUMIDIFIER (AIRVO 2) TO BE RETURNED FOR EVALUATION AND IS SEEKING FURTHER INFORMATION FROM THE HEALTHCARE FACILITY. F&P WILL PROVIDE A FOLLOW UP REPORT UPON COMPLETION OF OUR INVESTIGATION. PRODUCT BACKGROUND: THE AIRVO 2 DEVICE IS AN ACTIVE HUMIDIFIER WITH INTEGRATED FLOW GENERATOR THAT DELIVERS HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES TO SPONTANEOUSLY BREATHING PATIENTS THROUGH A VARIETY OF PATIENT INTERFACES. IT'S INTENDED USE IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES. THE AIRVO 2 DEVICE SHOULD NOT BE USED FOR LIFE SUPPORT PURPOSES, AND APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21383080&pc=QSK | 9612515-2025-00015 | 2023-10-01 04:00 | Death | VASCUTEK LTD | 2025-02-14 05:00 | QSK | THORAFLEX HYBRID | Adverse Event Without Identified Device or Use Problem; Insufficient Information |  | P210006 |  | 1 | Event Description: IMPLANT DATE : (B)(6) 2023. EVENT WAS REPORTED FROM EXTEND POST MARKET STUDY (B)(6)) AS ARTERIAL CLOTS.Event Description: IMPLANT DATE: (B)(6) 2023. EVENT WAS REPORTED FROM EXTEND POST MARKET STUDY (B)(4) AS ARTERIAL CLOTS. ADDITIONAL INFORMATION WAS RECEIVED ON 10TH MAR 2025: THE AE (ADVERSE EVENT) CHANGED FROM ARTERIAL CLOTS TO DEEP VEIN THROMBOSIS. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP # 1 FOR MANUFACTURING REPORT 9612515-2025-00015 TO PROVIDE EVENT CLOSURE INFORMATION FOR TAG0141. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4871 - CORONARY OBSTRUCTION/OCCLUSION - EVENT OF ARTERIAL CLOTS REPORTED FROM EXTEND POST MARKET STUDY. IMPACT CODE: 4648 - INSUFFICIENT INFORMATION - PATIENT OUTCOME ONGOING / UNRESOLVED STILL TREATING . DEVICE PROBLEM CODE : 3190 - INSUFFICIENT INFORMATION - NO DEVICE DEFICIENCY HAS BEEN COMMUNICATED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW THORAFLEX HYBRID SALES (B)(6) 2021 -(B)(6) 2024 VS OCCLUSION/ THROMOBOSIS > ARTERY EVENTS JAN 21 - JAN 25 GAVE AN (B)(4). ACTIONS WILL BE TAKEN ON THIS TREND. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW WILL BE PERFORMED. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION-INVESTIGATION IS ONGOING AND RESULTS ARE NOT YET AVAILABLE. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE -A CONCLUSION HAS YET TO BE ESTABLISHED AS THE INVESTIGATION IS INCOMPLETE. Manufacturer Narrative: CLINICAL CODE: 4871 - CORONARY OBSTRUCTION/OCCLUSION - EVENT OF ARTERIAL CLOTS REPORTED FROM EXTEND POST MARKET STUDY. IMPACT CODE: 1802 - DEATH: THE SITE CONFIRMED ON (B)(6) 2025, THE PATIENT DIED DUE TO SEPTIC SHOCK. DEVICE PROBLEM CODE: 2993 - ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM - THE SITE REPORTED DATA INDICATES NO DEVICE DEFECTS BEFORE USE AND NO ISSUES DURING DEVICE IMPLANT. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW THORAFLEX HYBRID SALES (B)(6) 2021 - (B)(6) 2024 VS OCCLUSION/ THROMBOSIS > ARTERY EVENTS (B)(6) 2021 - (B)(6) 2025 GAVE AN OCCURRENCE RATE OF (B)(4). ACTIONS WILL BE TAKEN ON THIS TREND. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION WAS RECEIVE FROM THE SITE CONFIRMED THE BELOW: · THE AE CHANGED FROM ARTERIAL CLOTS TO DEEP VEIN THROMBOSIS. · THE SITE REPORTED DATA INDICATES NO EVENTS RELATED TO PATIENTS ARTERY DAMAGE DURING IMPLANT. · THERE WAS NO CALCIFICATION IN THE ARTERY. · THE SITE REPORTED DATA INDICATES NO DEVICE DEFECTS BEFORE USE. · THE PATIENT MEDICAL HISTORY CONFIRMED ATRIAL FIBRILLATION · THE PATIENT DIED DUE TO SEPTIC SHOCK. · THE PATIENT HAD AN ANEURYSM OF POPLITEAL ARTERY PRIOR TO IMPLANT OF THORAFLEX HYBRID. ADDITIONALLY PAGE 41 ALSO CONFIRMED POST OPERATIVE "RIGHT SUPERFICIAL FEMORAL ARTERY OCCLUSION" ONSET 04OCT2023, REQUIRING ANTICOAGULANT THERAPY (REMAINED UNRESOLVED AT TIME OF DEATH). · INDICATES DISTAL SEAL WAS OBTAINED AT THE END OF PROCEDURE HOWEVER, THE ATTACHED MEDICAL NOTES INDICATED CT FINDINGS (FROM (B)(6)2023) AS "...UNCLEAR IF THIS IS TYPE 2 RETROGRADE FROM THE LEFT SUBCLAVIAN ARTERY OR A TYPE 1 FROM AN UNSEALED END OF GRAFT, HOWEVER, MAXIMAL AXIAL DIMENSION OF THE ANEURYSM ITSELF WAS UNCHANGED AT 74, COMPARED TO 70MM PREVIOUSLY." CONFIRMING SITE WERE AWARE. · CONFIRMS CAUSE OF DEATH WAS SEPTIC SHOCK, ALSO FURTHER CONFIRMS SEQUENCE OF EVENTS AND PRELIMINARY COD (CAUSE OF DEATH) AS SEPTIC SHOCK. 3331 - ANALYSIS OF PRODUCTION RECORDS: COMPREHENSIVE BATCH REVIEW WAS PERFORMED, NO ISSUES IDENTIFIED DURING MANUFACTURING PROCESS. DEVICE WAS MANUFACTURED TO SPECIFICATION. NO CAUSAL LINK BETWEEN THE DEVICE AND EVENT WAS IDENTIFIED. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. 4112 - ANALYSIS OF INFORMATION PROVIDED BY USER/THIRD PARTY: SCAN REVIEW WAS PERFORMED (B)(6) 2025, THIS REVIEW CONCLUDED: · A THORAFLEX HYBRID DEVICE WAS IMPLANTED TO TREAT AN ANEURYSM OF THE DISTAL ARCH AND PROXIMAL DESCENDING AORTA. THE DEVICE WAS IMPLANTED WITH A SUTURE COLLAR ATTACHMENT SITE IN ZONE 0/1. THE STENTED SECTION OF THE THORAFLEX HYBRID DEVICE IS NOT OBTAINING A DISTAL SEAL AND IS LOCATED WITHIN THE ANEURYSM SAC. A SINGLE BRANCH HAS BEEN USED TO RECONSTRUCT THE BCT AND LCCA, WITH THE LSA BEING OVER STENTED BUT REMAINING PATENT DUE TO RETROGRADE PERFUSION. · THE THORACOABDOMINAL AORTA IS ECTATIC, WITH A NOTABLE REGION OF THROMBUS AT THE LEVEL OF THE SMA. THERE IS AN INDWELLING EVAR DEVICE, WHICH IS FULLY PATENT. DISTALLY, THE RIGHT SUPERFICIAL FEMORAL AND POPLITEAL ARTERIES ARE OCCLUDED, AS IS THE LEFT POPLITEAL ARTERY. IT IS UNCLEAR IF THIS CONDITION WAS PRESENT PRE IMPLANT OF THE THORAFLEX HYBRID DEVICE. · ASSESSMENT OF THE IMAGING CONFIRMS THE THORAFLEX HYBRID DEVICE IS PERFORMING AS INTENDED, ALBEIT A DISTAL SEAL IS NOT ACHIEVED HOWEVER AS THIS CASE IS PART OF THE EXTEND STUDY, FUTURE DISTAL EXTENSION IS ANTICIPATED. THIS EVENT IS THEREFORE LIKELY TO BE PROCEDURE RELATED. INVESTIGATION FINDINGS: 213 - NO DEVICE PROBLEM FOUND: COMPREHENSIVE BATCH REVIEW WAS PERFORMED, NO ISSUES IDENTIFIED DURING MANUFACTURING PROCESS. DEVICE WAS MANUFACTURED TO SPECIFICATION. NO CAUSAL LINK BETWEEN THE DEVICE AND EVENT WAS IDENTIFIED. THE SITE REPORTED DATA INDICATES NO DEVICE DEFECTS BEFORE USE AND NO ISSUES DURING DEVICE IMPLANT. INVESTIGATION CONCLUSION: 4311 - ADVERSE EVENT RELATED TO PROCEDURE: THE ROOT CAUSE OF THIS INVESTIGATION WAS DETERMINED TO BE PROCEDURE RELATED. THIS IS DUE TO THE A SCAN REVIEW PERFORMED BY SME (SUBJECT MATTER EXPERT) WHICH CONFIRMED THAT THIS EVENT IS LIKELY TO BE PROCEDURE RELATED. 4323 - DEVICE PROBLEM EXCLUDED: COMPREHENSIVE BATCH REVIEW WAS PERFORMED, NO ISSUES IDENTIFIED DURING MANUFACTURING PROCESS. DEVICE WAS MANUFACTURED TO SPECIFICATION. NO CAUSAL LINK BETWEEN THE DEVICE AND EVENT WAS IDENTIFIED. THE SITE REPORTED DATA INDICATES NO DEVICE DEFECTS BEFORE USE AND NO ISSUES DURING DEVICE IMPLANT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21383090&pc=QSK | 9612515-2025-00016 | 2025-01-22 05:00 | Death | VASCUTEK LTD. | 2025-02-14 05:00 | QSK | THORAFLEX HYBRID | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Stenosis; Thrombosis/Thrombus | P2100006 |  | 1 | Event Description: THROMBOSIS OF THE PROSTHESIS, IN THE APPARENT ABSENCE OF ABNORMAL CONDITIONS (NO COAGULOPATHIES, NO RECENT INFECTIOUS CONDITIONS, NO EVIDENCE OF SURGICAL FITTING ALTERATIONS, NO KINKING OF THE PROSTHESIS. ON (B)(6) 2025 MASSIVE THROMBOSIS OF ENTIRE PROSTHESIS AND SUB-OCCLUSIVE THROMBUS OF THE ASCENDING AORTA WITHIN THE SURGICAL MONTAGE AND OF THE DESCENDING THORACIC AORTA WITH STENOSIS EXTENSION TO THE LEFT COMMON CAROTID ARTERY. THROMBOASPIRATION FAILURE WITH PENUMBRA ST STENT SYSTEM (12F). PLACEMENT OF ADVANTA V12 CAROTID STENT TO IMPROVE CEREBRAL PERFUSION. PATIENT IN CARDIAC ARREST AT THE END OF THE PROCEDURE , PLACED ON ECMO. DIED A FEW HOURS LATER.Event Description: THROMBOSIS OF THE PROSTHESIS , IN THE APPARENT ABSENCE OF ABNORMAL CONDITIONS( NO COAGULOPATHIES, NO RECENT INFECTIOUS CONDITIONS , NO EVIDENCE OF SURGICAL FITTING ALTERATIONS, NO KINKING OF THE PROSTHESIS. ON (B)(6) 2025 MASSIVE THROMBOSIS OF ENTIRE PROSTHESIS AND SUB-OCCLUSIVE THROMBUS OF THE ASCENDING AORTA WITHIN THE SURGICAL MONTAGE AND OF THE DESCENDING THORACIC AORTA WITH STENOSIS EXTENSION TO THE LEFT COMMON CAROTID ARTERY. THROMBOASPIRATION FAILURE WITH PENUMBRA ST STENT SYSTEM (12F). PLACEMENT OF ADVANTA V12 CAROTID STENT TO IMPROVE CEREBRAL PERFUSION . PATIENT IN CARDIAC ARREST AT THE END OF THE PROCEDURE , PLACED ON ECMO . DIED A FEW HOURS LATER. THIS REPORT IS BEING SUBMITTED AS A FOLLOW UP #1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00016 TO PROVIDE EVENT CLOSURE INFORMATION FOR TAG0145. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4440 -THROMBUS -THROMBUS OF ENTIRE PROSTHESIS + SUB OCCLUSIVE THROMBUS OF THE ASCENDING AORTA WITHIN THE SURGICAL MONTAGE AND THE ASCENDING AORTA WITH STENOSIS EXTENSION TO THE LEFT COMMON CAROTID ARTERY.. IMPACT CODE: 4624 - SURGICAL INTERVENTION - THROMBOASPIRATION FAILURE WITH PENUMBRA ST STENT SYSTEM(12F). PLACEMENT OF ADVANTA V12 CAROTID STENT TO IMPROVE CEREBRAL PERFUSION . PATIENT IN CARDIAC ARREST AT END OF PROCEDURE, PLACED ON ECMO (PERCEIVED TO BE EXTRACORPOREAL MEMBRANE OXYGENATION). 1802 - DEATH - PATIENT PASSED AWAY 6 YEARS POST SURGERY. DEVICE PROBLEM CODE: 3190 - INSUFFICIENT INFORMATION - THROMBUS OF THE PROSTHESIS IN THE APPARENT ABSENCE OF ABNORMAL CONDITIONS (NO COAGULOPATHIES, NO RECENT INFECTIOUS CONDITIONS) NO KINKING OF THE PROSTHESIS. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW THORAFLEX HYBRID SALES (B)(6) VS OCCLUSION/ THROMOBOSIS > ARTERY EVENTS JAN 21 - JAN 25 GAVE AN OCCURRENCE RATE OF (B)(4). ACTIONS WILL BE TAKEN ON THIS TREND. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW WILL BE PERFORMED. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION-INVESTIGATION IS ONGOING AND RESULTS ARE NOT YET AVAILABLE. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE -A CONCLUSION HAS YET TO BE ESTABLISHED AS THE INVESTIGATION IS INCOMPLETE. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4871 - CORONARY OBSTRUCTION / OCCLUSION - THROMBUS REPORTED THROUGHOUT THE ENTIRE PROSTHESIS. 2263 - STENOSIS OF THE VESSEL - NARROWING OF THE LEFT COMMON CAROTID ARTERY BRANCH. 4440 -THROMBUS -THROMBUS OF ENTIRE PROSTHESIS + SUB OCCLUSIVE THROMBUS OF THE ASCENDING AORTA WITHIN THE SURGICAL MONTAGE AND THE ASCENDING AORTA WITH STENOSIS TO THE LEFT COMMON CAROTID ARTERY BRANCH. IMPACT CODE: 4624 - SURGICAL INTERVENTION - THROMBOASPIRATION FAILURE WITH PENUMBRA ST STENT SYSTEM(12F). PLACEMENT OF ADVANTA V12 CAROTID STENT TO IMPROVE CEREBRAL PERFUSION . PATIENT IN CARDIAC ARREST AT END OF PROCEDURE , PLACED ON ECMO ( PERCEIVED TO BE EXTRACORPOREAL MEMBRANE OXYGENATION). 1802 - DEATH - PATIENT PASSED AWAY 6 YEARS POST-SURGERY. DEVICE PROBLEM CODE : 2993 -ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM - FULL BATCH REVIEW WAS PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT WHICH CONFIRMED THE DEVICE WAS MANUFACTURED TO SPECIFICATION. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW THORAFLEX HYBRID SALES JAN 21 - DEC 24 VS HYBRID OCCLUSION / THROMBOSIS <EMBOLI / STENOSIS JAN 21 - JAN 25 GAVE AN OCCURRENCE RATE OF 0.029%. NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION WAS RECEIVED - THE PATIENT HAD A TYPE A DISSECTION WITH AN ENTRY TEAR EXTENDING FROM THE ARCH TO THE ISTHMUS. THE PATIENT, WHO WAS ON A REGIMEN OF 75 MG OF ASPIRIN FOR THROMBUS MANAGEMENT, HAD A HISTORY OF RENAL FAILURE DUE TO COMPRESSION OF THE TRUE CANAL DURING A DISSECTION IN 2019. THEIR LAST RECORDED CHRONIC KIDNEY DISEASE (CKD) LEVEL WAS 20 ML/MIN. NO THROMBECTOMY WAS PERFORMED AS THE LUMEN OF THE PROSTHESIS HAD NEVER BEEN RE-PERMEABLE. THERE WERE NO COMPLICATIONS OR KINKING DETECTED IN THE AORTIC SURGICAL MONTAGE. THE VESSEL DIAMETERS WERE MEASURED AT 29 MM PROXIMALLY AND 26 MM DISTALLY. THE PATIENT HAD NO KNOWN ALLERGIES TO THE GRAFT COMPONENTS, AND THE GRAFT WAS NOT CUT BY CAUTERY. THE ARTERY WAS NOT CALCIFIED, AND NO PARTICULATES WERE OBSERVED ON THE GRAFT SURFACE DURING IMPLANTATION. THE GRAFT WAS HANDLED MINIMALLY, AND THE PATIENT HAD NO PRE-EXISTING VASCULAR DEVICES. THE DESCENDING AORTA WAS NOT PARTICULARLY TORTUOUS AROUND THE LANDING ZONE OR THE DISTAL SEALING RING. THE IMPLANT WAS INTENTIONALLY OVERSIZED TO REMODEL THE TRUE LUMEN, WHICH WAS SUCCESSFUL AS THE DIAMETER OF THE TRUE LUMEN AT THE END OF THE PROCEDURE WAS 26 MM. THE MOST RECENT CT SCAN SHOWED THROMBOSIS OF THE FALSE LUMEN. 3331 - ANALYSIS OF PRODUCTION RECORDS: ALL RETAINED MANUFACTURING RECORDS WERE REVIEWED AND SHOWED NO ISSUES WITH THE MANUFACTURE OF THE DEVICE. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED . INVESTIGATION FINDINGS: 213 - NO DEVICE PROBLEM WAS FOUND -FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT CONFIRMED THE BATCH WAS MANUFACTURED TO SPECIFICATION. ROOT CAUSE ANALYSIS WAS PERFORMED , AS INFECTION COULD POTENTIALLY CAUSE THROMBUS BUILD UP , PACKAGING INTEGRITY AND DELIVERY WAS REVIEWED , PRODUCT WAS ACCEPTED BY THE CUSTOMER WITH NO ISSUES REPORTED . STERILISATION OF THE DEVICE WAS REVIEWED AND ALL PROCESS TEST RESULTS WERE WITHIN SPECIFICATION, A REVIEW OF MANUFACTURING CONDITIONS - BIOLOGICAL INDICATORS ENDOTOXIN RESULTS WERE WITHIN SPECIFICATION. INAPPROPRIATE GEOMETRY CAN PROMOTE THE GENERATION OF THROMBUS INSIDE THE DEVICE - ( GEOMETRY OF THE DEVICE ) RETAINED DEVICE HISTORY RECORDS FOR THE DEVICE WERE REVIEWED WHICH SHOWED THE DEVICE WAS MANUFACTURES TO SPECIFICATION WITH NO ISSUES NOTED . (GEOMETRY OF THE IMPLANT SITE ) SCANS WERE REVIEWED AND ALTHOUGH THE SITE REPORTED "THE DESCENDING AORTA WAS NOT PARTICULARLY TORTUOUS AROUND THE LANDING ZONE OR THE DISTAL SEALING RING". - THE FIRST IMAGE ON THE SCAN REVIEW SHOWS TORTUOUS ARCH ANATOMY. INCORRECT LENGTH OF DEVICE CAN POTENTIALLY CAUSE THROMBUS , IT WAS CONFIRMED THE DEVICE WAS THE CORRECT LENGTH FOR THE PROCEDURE. INCORRECT OVERSIZE - THE OVERSIZE COULD NOT BE DETERMINED AS NO PRE-OP SCANS WERE PROVIDED FOR REVIEW. ALL PHYSICAL TESTING WAS REVIEWED FOR BASE FABRIC WITH TEST RESULTS WITHIN SPECIFICATION. DEVICE HISTORY RECORD FOR THE JOINED DEVICE WAS REVIEWED WHICH SHOWED NO ISSUES WITH THE SEWING WHICH COULD POTENTIALLY CAUSE TURBULENT BLOOD FLOW RESULTING IN THROMBUS FORMATION. POST IMPLANT IMAGING DEMONSTRATED A COMPRESSION OF THE DEVICE AT THE PROXIMAL ANASTOMOSIS SITE AND AT THE LEVEL OF THE LCCA ( LEFT COMMON CAROTID ARTERY) BRANCH. THE DISTAL STENT RING WAS ALSO INVERTED, WHICH MAY ALL HAVE CONTRIBUTED TO REDUCED FLOW THROUGH THE DEVICE, RESULTING IN THROMBUS FORMATION. DEVICE RELATIONSHIP - UNDETERMINED INVESTIGATION CONCLUSION: 4315 - CAUSE NOT ESTABLISHED - NO CAUSAL LINK BETWEEN THE EVENT AND THE DEVICE COULD BE ESTABLISHED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21383379&pc=PNJ | 9612164-2025-00788 | 2025-01-23 05:00 | Death | MEDTRONIC IRELAND | 2025-02-14 05:00 | PNJ | MICRA AV | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Myocardial Infarction; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENTS¿ BLOOD PRESSURE DROPPED AND EXPERIENCED AN ACUTE INFERIOR WALL MYOCARDIAL INFARCTION, PERICARDIAL TAMPONADE FROM CARDIAC PERFORATION AND PERICARDIAL EFFUSION. RESUSCITATION ATTEMPTS WERE MADE WITH MEDICATION AND A PERICARDIOCENTESIS WAS PERFORMED WITH CARDIOPULMONARY RESUSCITATION AND INTUBATION. RESCUE WAS INEFFECTIVE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID: [MC1AVR1-DELSYS] (SERIAL: [(B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21383775&pc=NIP | 2183870-2025-00080 | 2025-02-02 05:00 | Death | COVIDIEN | 2025-02-14 05:00 | NIP | EVERFLEX SELF-EXPANDING PERIPHERAL STENT WITH ENTRUST DELIVERY SYSTEM LONG | Activation, Positioning or Separation Problem; Material Deformation | No Clinical Signs, Symptoms or Conditions | P110023 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE DEPLOYMENT OF AN EVERFLEX ENTRUST SELF-EXPANDED STENT IN THE SUPERFICIAL FEMORAL ARTERY AND POPLITEAL ARTERY, SEVERAL ISSUES OCCURRED. THE STENT EXPERIENCED DEFORMATION DURING DEPLOYMENT AND WAS UNABLE TO BE DEPLOYED AS INTENDED. THE THUMBWHEEL BECAME STUCK AFTER ONE TURN, AND DURING MANUAL DEPLOYMENT USING THE PIN AND PUSH TECHNIQUE, THE STENT BECAME ELONGATED IN VIVO. THE ELONGATED STENT WAS ULTIMATELY LEFT IN THE PATIENT. THE PROCEDURE WAS COMPLETED, AND NO PATIENT SYMPTOMS OR COMPLICATIONS WERE ASSOCIATED WITH THIS EVENT. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SUPERFICIAL FEMORAL ARTERY IS THE ASSOCIATED VESSEL, RIGHT DISTAL. TYPE OF LESION WAS LONG CALCIFIED WITH NO DEGREE OF TORTUOSITY, SEVERE DEGREE OF CALCIFICATION AND 70-80% LESION STENOSIS, ARTERY DIAMETER WAS 6 MM AND LESION LENGTH WAS 80 MM. NO ABNORMALITIES REPORTED IN RELATION TO ANATOMY. A BALKIN 6F SHEATH AND A SUPER STIFF 0.035 GUIDEWIRE WAS USED. NO EMBOLIC PROTECTION WAS USED. NO DAMAGE WAS NOTED TO PACKAGING (I.E. SHELF CARTON, HOOP/TRAY) AND NO ISSUES NOTED WHEN REMOVING THE DEVICE FROM THE HOOP/TRAY. THE IFU WAS FOLLOWED AND THE DEVICE WAS PREPPED WITH NO ISSUES IDENTIFIED. THE LESION WAS PRE-DILATED. NO RESISTANCE WAS ENCOUNTERED DURING DELIVERY TO LESION. NO EXCESSIVE FORCE WAS USED AND THE DEVICE DID PASS THROUGH A PREVIOUSLY DEPLOYED STENT. THE THUMBSCREW/LOCK-PIN WAS CHECKED FOR SECUREMENT PRIOR TO PROCEDURE AND THE LOCK-PIN WAS REMOVED BEFORE ATTEMPTED DEPLOYMENT. WHEN TRYING TO DEPLOY MANUALLY USING THE PIN AND PUSH TECHNIQUE THE STENT GOT ELONGATED ("LIKE A SUPERA"). THE PROCEDURE WAS STARTED WITH A COMPETITIVE STENT THAT ALSO DID NOT DEPLOY CORRECTLY. THE ENTRUST WAS ADDED TO SECURE THE FIRST FAILED STENT. PATIENT SYMPTOMS WERE NOT IMMEDIATE, IN THE EVENING OF THE PROCEDURE, PATIENT EXPERIENCED SEVERE PAIN. ON SUNDAY (B)(6)2025, PATIENT ARRIVED WITH SEVERE INFECTION AND UNDERWENT A MAJOR AMPUTATION ATK ON THURSDAY (B)(6)2025. PATIENT PASSED AWAY ON (B)(6)2025. THE HOSPITAL HAS NO CLAIMS/ALLEGATION TOWARD THE DEVICES USED AND THE REPORTED DEATH. THE PATIENT WAS VERY COMPLEX TO BEGIN WITH MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE ANALYSIS ONE STILL IMAGE WAS RETURNED FOR ANALYSIS. IMAGE SHOWS THE RIGHT FEMORAL TO POPLITEAL ARTERY WITH ONE LONG STENT IMPLANTED WITHIN A SECOND SHORT STENT. THE LONG STENT APPEARS TO BE DEFORMED AT THE DISTAL END. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS ONLY THE CATHETER PORTION OF THE DEVICE WAS RETURNED, THE DEVICE WAS IDENTIFIED BY THE STRAIN RELIEF, THE PULL WIRE WAS STILL ATTACHED TO THE REST OF THE DEVICE, A SERIES OF KINKS WERE NOTED ALONG THE LENGTH OF THE SILVER OUTER SHEATH THE BACK HUB / LEUR WAS ALSO RETURNED SEPARATE TO THE DEVICE, DUE TO THE CONDITION OF THE RETURNED DEVICE, NO FURTHER TESTING COULD BE PERFORMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21384012&pc=BZD | 2518422-2025-102051 | 2024-11-04 05:00 | Death | RESPIRONICS, INC. | 2025-02-14 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A PHILIPS CPAP DEVICE. THE PATIENT'S WIFE CONTACTED US, INDICATING THAT SHE HAS A REPLACEMENT DEVICE AND HAS ALREADY RETURNED THE AFFECTED UNIT. THE RESPIRATORY THERAPIST INFORMED THE PATIENT THAT THE REPLACEMENT DEVICE DID NOT NEED TO BE RETURNED. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. ADDITIONALLY, NO DEVICE INFORMATION WAS PROVIDED REGARDING THE CPAP UNIT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21384080&pc=LYJ | 1644487-2025-00165 | 2025-01-06 05:00 | Death | LIVANOVA USA, INC. | 2025-02-14 05:00 | LYJ | PULSE GEN MODEL UNKNOWN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: DURING A REVIEW OF SCIENTIFIC LITERATURE TITLED: LONG-TERM SEIZURE REDUCTION ASSOCIATED WITH VAGAL NERVE STIMULATION IN DRAVET SYNDROME; IT WAS REPORTED THAT A PATIENTS IMPLANTED WITH VNS EXPERIENCED RESPIRATORY COMPLICATIONS DURING AN ADMISSION FOR STATUS EPILEPTICUS- WHICH ULTIMATELY LED TO DEATH. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LONG-TERM SEIZURE REDUCTION ASSOCIATED WITH VAGAL NERVE STIMULATION IN DRAVET SYNDROME. SUNANJAY BAJAJ1, ALINA IVANIUK2, TOBIAS BRUENGER1, ÉMILE MOURA COELHO DA SILVA1, EMILY HUTH1, LUDOVICA MONTANUCCI1, COSTIN LEU1, GARY TAYLOR1, MOUSUMI SINHA1, RAHIL A. TAI1, MANISH N. SHAH3,4, MICHAEL W. WATKINS1, JEREMY E. LANKFORD1, INDIRA M. KOMMURU1, SANDIPAN PATI1, PRAKASH KOTAGAL5, ANDREAS ALEXOPOULOS5, SAMDEN D. LHATOO1, ELIA PESTANA KNIGHT5, GRETCHEN VON ALLMEN6, DENNIS LAL. LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21384603&pc=NIK | 2017865-2025-14129 | 2024-01-06 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-14 05:00 | NIK | GALLANT HF | Inappropriate/Inadequate Shock/Stimulation | Discomfort | P030054 |  | 1 | Event Description: RELATED MFR REPORT NUMBER: 2017865-2024-03204. IT WAS REPORTED THAT A PATIENT WAS IMPLANTED THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) ON (B)(6) 2024. IT WAS REPORTED THAT THE ICD AND RIGHT VENTRICULAR (RV) LEAD WERE DEFECTIVE AND INCORRECTLY IMPLANTED LEADING TO INAPPROPRIATELY SHOCKED THE PATIENT RESULTING IN PERMANENT PAIN, MENTAL ANGUISH, AND DISFIGUREMENT UP UNTIL HIS DEATH ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21384616&pc=FTL | 1213643-2025-090081 | 2005-11-16 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. | 2025-02-14 05:00 | FTL | KUGEL PATCH | Defective Device; Patient Device Interaction Problem | Abscess; Adhesion(s); Bacterial Infection; Cellulitis; Fistula; Seroma; Hernia | K963141 |  | 1 | Event Description: PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: (B)(6) 2003 - PATIENT WITH A COMPLICATED HISTORY OF MULTIPLE HERNIA REPAIRS WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF KUGEL PATCH (DEVICE #1). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS DISSECTED, AND ADHESIONS WERE TAKEN DOWN. A KUGEL PATCH (DEVICE #1) WAS PLACED AND SUTURED.¿ (B)(6) 2003 - PATIENT WAS DIAGNOSED WITH SEROMA. (B)(6) 2004 - PATIENT HAD CELLULITIS AROUND THE WOUND THAT HAD BEEN TREATED WITH ANTIBIOTICS WAS COMPLETELY GONE AWAY. (B)(6) 2004 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT REPAIR WITH IMPLANT OF COMPOSIX KUGEL (DEVICE #2). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS OPENED AND MULTIPLE FASCIAL DEFECTS IDENTIFIED. MULTIPLE BOWEL ADHESIONS FROM THE ANTERIOR ABDOMINAL WALL TO OLD GORE-TEX MESH AND OLD PACKS WAS FOUND. ONCE ALL THE ADHESIONS WERE CLEARED, A COMPOSIX KUGEL (DEVICE #2) WAS PLACED AND SUTURED.¿ (B)(6) 2005 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF SEPRAMESH IP (DEVICE #3). PER OPERATIVE NOTES, ¿THE BOWEL ADHESIONS TO THE OLD KUGEL MESHES (DEVICE #1 & #2) PLACED SUPERIORLY WERE TAKEN DOWN. THE HERNIA SAC WAS EXCISED AND A SEPRAMESH IP (DEVICE #3) WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH, VENTRAL HERNIA THEREBY UNDERWENT REMOVAL OF ALL ABDOMINAL WALL MESHES (DEVICE #1, #2, & #3) AND IMPLANT OF BIOLOGICAL MESH. PER OPERATIVE NOTES, ¿THERE WAS AN ABSCESS AND FISTULA NOTED. BOWEL INTERLOOP ADHESIONS FROM PREVIOUS ABDOMINAL WALL MESH WERE TAKEN DOWN. TH MESH ALONG WITH THE PROLENE SUTURES, OLD METAL TACKS WERE ISOLATED, AND HERNIAS REDUCED. ALL THE ABDOMINAL WALL MESHES (DEVICE #1, #2 & #3) WERE RESECTED AND A BIOLOGICAL MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH WITH COLONIC FISTULA THEREBY UNDERWENT REMOVAL OF BIOLOGICAL MESH AND IMPLANT OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿THE BIOLOGICAL MESH AND COLONIC FISTULA WERE REMOVED. SMALL BOWEL RESECTION AND ANASTOMOSIS WAS DONE. A SYNTHETIC MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH PNEUMONIA. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT REMOVAL OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿A LARGE AMOUNT OF SUCCUS COMING FROM THE INTRAABDOMINAL REGION. THERE WERE INTRA-ABDOMINAL LOOP ABSCESSES. THE END ILEOSTOMY BECAME A MUCOUS FISTULA WAS SUTURED. THE SYNTHETIC MESH WAS REMOVED AND DEBRIDED ABDOMINAL WALL, THE MUSCLE, SKIN, AND FASCIA THAT WAS NECROTIC. A TEMPORARY VAC DRESSING PLACED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN THEREBY UNDERWENT ABDOMEN WASHOUT, VAC PLACEMENT AND ABDOMINAL WALL DEBRIDEMENT. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN AND ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT OPEN REPAIR OF ABDOMINAL WOUND, ABDOMEN WASHOUT AND ENTEROCUTANEOUS FISTULA. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, PLUS RETRACTION OF ILEOSTOMY THEREBY UNDERWENT ABDOMINAL WASHOUT, ILEOSTOMY REVISION, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT ABDOMINAL WASHOUT, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTERO ATMOSPHERIC FISTULA THEREBY UNDERWENT WOUND VAC CHANGE. (B)(6) 2018 - PATIENT PASSED AWAY. PER DEATH SUMMARY, ¿76 Y/O FEMALE WITH INFECTED ABDOMINAL MESH WAS TAKEN TO THE OPERATING ROOM FOR EXPLORATORY LAPAROTOMY MESH REMOVAL COMPLEX ABDOMINAL WALL REPAIR, BOWEL RESECTIONS ON (B)(6) 2017 WAS TAKEN BACK TO THE OPERATING ROOM ON (B)(6) 2017 FOR INTRA-ABDOMINAL SEPSIS ON ENTEROCUTANEOUS FISTULA. PATIENT DEVELOPED ANOTHER FISTULA SUBSEQUENT TO THIS AND HAD MULTIPLE TRIPS TO THE OPERATING ROOM FOR ABDOMINAL WALL WASHOUTS AND ATTEMPTS AT ENTEROCUTANEOUS FISTULA REPAIR. PATIENT DID DEVELOP A PNEUMONIA AND WAS ON MULTIPLE ANTIBIOTICS FOR SEVERAL WEEKS WHILE IN THE HOSPITAL. PATIENT WAS UNABLE TO BE WEANED FROM THE VENTILATOR AND THE FAMILY ULTIMATELY MADE A DECISION TO TERMINALLY EXTUBATE THE PATIENT AND MAKE HER COMFORT MEASURES ONLY AFTER SEVERAL WEEKS OF BEING ON THE VENTILATOR. PATIENT WENT ALMOST 48 HOURS AFTER EXTUBATION AND DIED AT 12:36 PM ON (B)(6) 2018.¿ ATTORNEY ALLEGED THAT THE PATIENT HAD ADHESIONS, BOWEL OBSTRUCTION, DEATH, FISTULAE, INFECTION, PAIN AND HERNIA RECURRENCE. IT WAS ALSO ALLEGED THAT THE PATIENT HAD BULGE, WOUND VAC, TRANSVERSE COLOSTOMY, ABDOMINAL WALL DEBRIDEMENT, COLON REPAIR, SEPARATION OF ABDOMINAL WALL COMPONENTS, END-ILEOSTOMY, ENDOSOTOMY. Manufacturer Narrative: BASED ON THE INFORMATION RECEIVED, NO CONCLUSIONS CAN BE MADE. PER MEDICAL RECORDS REVIEW, ABOUT 2 YEARS POST IMPLANT OF SINGLE USE STERILE KUGEL PATCH, PATIENT WAS DIAGNOSED WITH FISTULA, ABSCESS, INFECTION, SEROMA, HERNIA RECURRENCE, CELLULITIS AND ADHESIONS THEREBY UNDERWENT REPAIR WITH REMOVAL OF MESH. PATIENT WAS PASSED AWAY ON (B)(6) 2018. NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. THE INSTRUCTIONS-FOR-USE SUPPLIED WITH THE DEVICE LISTS FISTULA, SEROMA, HERNIA RECURRENCE AND ADHESIONS AS POSSIBLE COMPLICATIONS. IN REGARDS TO INFECTION THE WARNING SECTION STATES, "IF AN INFECTION DEVELOPS, TREAT THE INFECTION AGGRESSIVELY. CONSIDERATION SHOULD BE GIVEN REGARDING THE NEED TO REMOVE THE MESH. AN UNRESOLVED INFECTION MAY REQUIRE REMOVAL OF THE DEVICE". NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. THIS EMDR REPRESENTS THE KUGEL PATCH (DEVICE #1). AN ADDITIONAL SUPPLEMENTAL EMDR WAS SUBMITTED TO REPRESENTS THE MESH ¿ COMPOSIX KUGEL (DEVICE #2) AND AN EMDR WAS SUBMITTED TO REPRESENT THE SEPRAMESH IP (DEVICE #3). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21384622&pc=FTL | 1213643-2025-090082 | 2017-12-14 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. | 2025-02-14 05:00 | FTL | SEPRAMESH IP | Defective Device; Patient Device Interaction Problem | Abscess; Adhesion(s); Bacterial Infection; Fistula | K063739 |  | 1 | Event Description: PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: (B)(6) 2003 - PATIENT WITH A COMPLICATED HISTORY OF MULTIPLE HERNIA REPAIRS WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF KUGEL PATCH (DEVICE #1). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS DISSECTED, AND ADHESIONS WERE TAKEN DOWN. A KUGEL PATCH (DEVICE #1) WAS PLACED AND SUTURED.¿ (B)(6) 2003 - PATIENT WAS DIAGNOSED WITH SEROMA. (B)(6) 2004 - PATIENT HAD CELLULITIS AROUND THE WOUND THAT HAD BEEN TREATED WITH ANTIBIOTICS WAS COMPLETELY GONE AWAY. (B)(6) 2004 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT REPAIR WITH IMPLANT OF COMPOSIX KUGEL (DEVICE #2). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS OPENED AND MULTIPLE FASCIAL DEFECTS IDENTIFIED. MULTIPLE BOWEL ADHESIONS FROM THE ANTERIOR ABDOMINAL WALL TO OLD GORE-TEX MESH AND OLD PACKS WAS FOUND. ONCE ALL THE ADHESIONS WERE CLEARED, A COMPOSIX KUGEL (DEVICE #2) WAS PLACED AND SUTURED.¿ (B)(6) 2005 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF SEPRAMESH IP (DEVICE #3). PER OPERATIVE NOTES, ¿THE BOWEL ADHESIONS TO THE OLD KUGEL MESHES (DEVICE #1 & #2) PLACED SUPERIORLY WERE TAKEN DOWN. THE HERNIA SAC WAS EXCISED AND A SEPRAMESH IP (DEVICE #3) WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH, VENTRAL HERNIA THEREBY UNDERWENT REMOVAL OF ALL ABDOMINAL WALL MESHES (DEVICE #1, #2, & #3) AND IMPLANT OF BIOLOGICAL MESH. PER OPERATIVE NOTES, ¿THERE WAS AN ABSCESS AND FISTULA NOTED. BOWEL INTERLOOP ADHESIONS FROM PREVIOUS ABDOMINAL WALL MESH WERE TAKEN DOWN. TH MESH ALONG WITH THE PROLENE SUTURES, OLD METAL TACKS WERE ISOLATED, AND HERNIAS REDUCED. ALL THE ABDOMINAL WALL MESHES (DEVICE #1, #2 & #3) WERE RESECTED AND A BIOLOGICAL MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH WITH COLONIC FISTULA THEREBY UNDERWENT REMOVAL OF BIOLOGICAL MESH AND IMPLANT OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿THE BIOLOGICAL MESH AND COLONIC FISTULA WERE REMOVED. SMALL BOWEL RESECTION AND ANASTOMOSIS WAS DONE. A SYNTHETIC MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH PNEUMONIA. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT REMOVAL OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿A LARGE AMOUNT OF SUCCUS COMING FROM THE INTRAABDOMINAL REGION. THERE WERE INTRA-ABDOMINAL LOOP ABSCESSES. THE END ILEOSTOMY BECAME A MUCOUS FISTULA WAS SUTURED. THE SYNTHETIC MESH WAS REMOVED AND DEBRIDED ABDOMINAL WALL, THE MUSCLE, SKIN, AND FASCIA THAT WAS NECROTIC. A TEMPORARY VAC DRESSING PLACED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN THEREBY UNDERWENT ABDOMEN WASHOUT, VAC PLACEMENT AND ABDOMINAL WALL DEBRIDEMENT. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN AND ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT OPEN REPAIR OF ABDOMINAL WOUND, ABDOMEN WASHOUT AND ENTEROCUTANEOUS FISTULA. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, PLUS RETRACTION OF ILEOSTOMY THEREBY UNDERWENT ABDOMINAL WASHOUT, ILEOSTOMY REVISION, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT ABDOMINAL WASHOUT, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTERO ATMOSPHERIC FISTULA THEREBY UNDERWENT WOUND VAC CHANGE. (B)(6) 2018 - PATIENT PASSED AWAY. PER DEATH SUMMARY, ¿76 Y/O FEMALE WITH INFECTED ABDOMINAL MESH WAS TAKEN TO THE OPERATING ROOM FOR EXPLORATORY LAPAROTOMY MESH REMOVAL COMPLEX ABDOMINAL WALL REPAIR, BOWEL RESECTIONS ON (B)(6) 2017 WAS TAKEN BACK TO THE OPERATING ROOM ON (B)(6) 2017 FOR INTRA-ABDOMINAL SEPSIS ON ENTEROCUTANEOUS FISTULA. PATIENT DEVELOPED ANOTHER FISTULA SUBSEQUENT TO THIS AND HAD MULTIPLE TRIPS TO THE OPERATING ROOM FOR ABDOMINAL WALL WASHOUTS AND ATTEMPTS AT ENTEROCUTANEOUS FISTULA REPAIR. PATIENT DID DEVELOP A PNEUMONIA AND WAS ON MULTIPLE ANTIBIOTICS FOR SEVERAL WEEKS WHILE IN THE HOSPITAL. PATIENT WAS UNABLE TO BE WEANED FROM THE VENTILATOR AND THE FAMILY ULTIMATELY MADE A DECISION TO TERMINALLY EXTUBATE THE PATIENT AND MAKE HER COMFORT MEASURES ONLY AFTER SEVERAL WEEKS OF BEING ON THE VENTILATOR. PATIENT WENT ALMOST 48 HOURS AFTER EXTUBATION AND DIED AT 12:36 PM ON (B)(6) 2018.¿ ATTORNEY ALLEGED THAT THE PATIENT HAD ADHESIONS, BOWEL OBSTRUCTION, DEATH, FISTULAE, INFECTION, PAIN AND HERNIA RECURRENCE. IT WAS ALSO ALLEGED THAT THE PATIENT HAD BULGE, WOUND VAC, TRANSVERSE COLOSTOMY, ABDOMINAL WALL DEBRIDEMENT, COLON REPAIR, SEPARATION OF ABDOMINAL WALL COMPONENTS, END-ILEOSTOMY, END OSOTOMY. Manufacturer Narrative: BASED ON THE INFORMATION RECEIVED, NO CONCLUSIONS CAN BE MADE. PER MEDICAL RECORDS REVIEW, ABOUT 12 YEARS POST IMPLANT OF SINGLE USE STERILE SEPRAMESH IP, PATIENT WAS DIAGNOSED WITH FISTULA, ABSCESS, INFECTION AND ADHESIONS THEREBY UNDERWENT REPAIR WITH REMOVAL OF MESH. PATIENT WAS PASSED AWAY ON (B)(6) 2018. NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. THE INSTRUCTIONS-FOR-USE SUPPLIED WITH THE DEVICE LISTS FISTULA AND ADHESIONS AS POSSIBLE COMPLICATIONS. IN REGARDS TO INFECTION THE WARNING SECTION STATES, "IF AN INFECTION DEVELOPS, TREAT THE INFECTION AGGRESSIVELY. CONSIDERATION SHOULD BE GIVEN REGARDING THE NEED TO REMOVE THE PROSTHESIS. AN UNRESOLVED INFECTION MAY REQUIRE REMOVAL OF THE PROSTHESIS". NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. THIS EMDR REPRESENTS THE SEPRAMESH IP (DEVICE #3). AN ADDITIONAL SUPPLEMENTAL EMDR WAS SUBMITTED TO REPRESENTS THE MESH ¿ COMPOSIX KUGEL (DEVICE #2) AND AN EMDR WAS SUBMITTED TO REPRESENT THE KUGEL PATCH (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21385396&pc=DQO | 2015691-2025-01054 |  | Death | EDWARDS LIFESCIENCES PR | 2025-02-14 05:00 | DQO | SWAN-GANZ CCOMBO V | Difficult to Remove; Use of Device Problem; Physical Resistance/Sticking | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Laceration(s) | K160084 |  | 1 | Event Description: EDWARDS LIFESCIENCES WAS NOTIFIED VIA A DIRECT VOLUNTARY INCIDENT REPORT SUBMISSION TO THE FDA THAT A CUSTOMER EXPERIENCED RESISTANCE WITH REMOVAL OF A SWAN GANZ CATHETER. PATIENT WAS REPOSITIONED WHICH ALLOWED THE CATHETER TO BE REMOVED SMOOTHLY. AFTER REMOVAL, PATIENT'S BLOOD PRESSURE SUDDENLY DROPPED AND CHEST TUBE NOTED TO HAVE LARGELY INCREASED SANGUINEOUS OUTPUT. PATIENT WAS AWAKE THROUGHOUT THE EVENT. PLANS WERE MADE TO GO TO THE OPERATING ROOM. THERE WAS APPROXIMATELY A 3 X 2 CM DEFECT IN THE ANTERIOR WALL OF THE SVC, JUST SUPERIOR TO THE CAVOATRIAL JUNCTION. THE INJURY ENCOMPASSED ABOUT 40% OF THE CIRCUMFERENCE OF THE SUPERIOR VENA CAVA (SVC). THE SVC HAD BEEN CLAMPED FOR A FEW MINUTES AND THERE WAS EVIDENCE OF VENOUS CONGESTION IN THE INNOMINATE VEIN. PATCH REPAIR OF LARGE SUPERIOR VENA CAVA LACERATION WAS DONE. PATIENT PROVED TO HAVE TOLERATED THE PROCEDURE SATISFACTORILY. PATIENT WAS THEN TRANSFERRED TO INTENSIVE CARE UNIT. IT WAS SUSPECTED THAT THE SWAN HAD BEEN LOOPED BY THE SUTURE AT TIME OF OPERATION EITHER IN THE SVC OR RIGHT ATRIUM AS THIS WOULD LEAVE A SMALL DEFECT IN THE WALL. PATIENT EXPERIENCED HEMORRHAGE/BLEEDING, LOW BLOOD PRESSURE/ HYPOTENSION, LACERATIONS, AND DEATH. PATIENT WAS UNDERGOING REMOVAL OF MITRAL VALVE NODULE, A REDO AORTIC VALVE REPLACEMENT, AND TRANSESOPHAGEAL ECHOCARDIOGRAM CENTRAL LINE INSERTION WHEN ISSUE OCCURRED. THE VOLUNTARY INCIDENT REPORT DID NOT INCLUDE CUSTOMER CONTACT INFORMATION, PATIENT INFORMATION OR PRODUCT INFORMATION RELATED TO THE REPORTED EVENT. Manufacturer Narrative: EDWARDS LIFESCIENCES WAS UNABLE TO PERFORM DUE DILIGENCE EFFORTS TO RETRIEVE THE SUSPECT DEVICE FOR ANALYSIS AS THERE WAS NO CUSTOMER CONTACT INFORMATION PROVIDED IN THE VOLUNTARY INCIDENT REPORT. THE LOT NUMBER WAS NOT PROVIDED THUS A DEVICE HISTORY RECORD WAS NOT REVIEWED. NO PRODUCT WAS RETURNED FOR EVALUATION; THEREFORE, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED, AND CORRECTIVE OR PREVENTIVE ACTION IS NOT INDICATED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS A PART OF THE MONTHLY REVIEW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21385465&pc=DSI | 3010157426-2025-00032 | 2025-01-13 05:00 | Death | SPACELABS HEALTHCARE | 2025-02-14 05:00 | DSI | SPACELABS ULTRAVIEW SL COMMAND MODULE | Capturing Problem | Arrhythmia | K103142 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT A CRITICALLY ILL PATIENT THAT OVERDOSED ON CALCIUM CHANNEL BLOCKERS AND A TRANSCUTANEOUS PACER WAS ADMITTED TO A BEDSIDE MONITOR. THE PACER DETECTION OF THE MODULE WAS TURNED ON, HOWEVER, THERE WAS STILL SIGNIFICANT ARTIFACT SEEN ON THE MODULE. A TRANSVENOUS WIRE WAS PLACED AND THE CUSTOMER STATED THAT THE INTERFERENCE IMPACTED THE CLINICAL STAFF'S ABILITY TO CONFIRM PACER FUNCTIONALITY. Manufacturer Narrative: CLINICAL ACCESS RETROSPECTIVE DATA WAS REVIEWED AND CONFIRMED THAT THERE WAS SIGNIFICANT ARTIFACT ON THE ECG SIGNAL THAT DISPLAYED AS MULTIPLE PACER SPIKES. THE FREQUENCY OF THE PACER SPIKES INDICATES THAT THE INTERFERENCE WAS DUE TO AN UNKNOWN EXTERNAL DEVICE. Manufacturer Narrative: THE CUSTOMER BIOMED TESTED THE DEVICE AFTER THE EVENT AND COULD NOT FIND ANY ISSUES WITH THE PACER DETECTION OF THE COMMAND MODULE. A DATABASE INCLUDING THE PATIENT WAVEFORM WAS REQUESTED TO EVALUATE THE ARTIFACT PRESENT AT THE TIME OF THE REPORTED EVENT. AT THIS TIME THE DATABASE HAS NOT BEEN RECEIVED. A FOLLOW UP REPORT WILL BE SUBMITTED IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21385471&pc=OZD | 1220648-2025-26148 | 2025-01-20 05:00 | Death | ABIOMED, INC. | 2025-02-14 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED LIMB ISCHEMIA AND ACCESS SITE BLEEDING REQUIRING A SURGICAL CUTDOWN AND MASSIVE BLOOD TRANSFUSION. THE PATIENT WOULD LATER EXPIRE, BUT PRIOR THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 WITH CONTINUED COMPLICATION. THE PATIENT ARRIVED AT THE HOSPITAL VIA LIFE FLIGHT AND WAS INTUBATED. PATIENT WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION TO THE CATHETERIZATION LAB FOR INTERVENTION CODING MULTIPLE TIMES IN ROUTE AND ON ARRIVAL. THE IMPELLA CP WAS PLACED VIA THE RIGHT FEMORAL ARTERY. LEFT VENTRICLE ACCESS WAS CHALLENGING DUE TO MODERATE TO SEVERE AORTIC STENOSIS. AFTER MULTIPLE ATTEMPTS, THE PHYSICIAN WAS FINALLY ABLE TO CROSS THE AORTIC VALVE INTO THE LEFT VENTRICLE. THE IMPELLA CP WAS INSERTED WITHOUT AN ISSUE. FLOWS 3.7L, MEAN ARTERIAL PRESSURE WAS 66 ON EPINEPHRINE AND LEVOPHED. MULTIPLE BLOOD PRODUCTS WERE INFUSED AT THIS TIME DUE TO PATIENT HEMOGLOBIN OF 4.9 AT THE START OF THE CASE. TWO STENTS WERE IMPLANTED TO THE LEFT ANTERIOR DESCENDING ARTERY. POST PERCUTANEOUS CORONARY INTERVENTION, THE 14F X 13 CM PEEL-AWAY SHEATH INTRODUCER WAS REMOVED AND REPOSITIONING SHEATH PLACED. DIGITAL SUBTRACTION ANGIOGRAPHY WAS CHECKED TO ASSESS FOR RIGHT LOWER EXTREMITY FLOW; OCCLUSION NOTED. FEM-FEM ANTEGRADE SHEATHS WERE PLACED AND PERFUSION RESTORED. A LARGE HEMATOMA WAS NOTED AND PROFOUND BLEEDING AROUND THE REPOSITIONING SHEATH OCCURRED. DECISION MADE TO EXPLANT THE IMPELLA CP IN CATHETERIZATION LAB WITH VASCULAR SURGERY AT LAB TABLE FOR POSSIBLE DRY CLOSURE FOR RIGHT COMMON FEMORAL ARTERY (RCFA). THE DECISION WAS MADE TO TAKE THE PATIENT TO THE OPERATING ROOM FOR A SURGICAL CUT DOWN TO RCFA. THE PATIENT'S HEMOGLOBIN WAS 4.9 ON ARRIVAL AND OTHER BLEEDING SOURCES WERE QUESTIONED. NO FURTHER SCANS WERE COMPLETED PRIOR TO ARRIVING IN THE CATH LAB. THE PATIENT RECEIVED MULTIPLE UNITS OF PACK RED BLOOD CELLS, FRESH FROZEN PLASMA AND CRYOPRECIPITATE WERE GIVEN. IT WAS NOTED NO ISSUES OR RESISTANCE WERE NOTED DURING INSERTION NOR DURING THE CASE; PATIENT'S HEMODYNAMICS IMPROVED WITH THE IMPELLA CP. HOWEVER, THE PATIENT WAS IN CRITICAL CONDITION. THE FOLLOWING DAY, THE PATIENT WAS PLACED ON AN IMPELLA 5.5. DURING THE PROCEDURE THE LEFT VENTRICLE ACCESS WAS OBTAINED AND THE 0.018 GUIDEWIRE PLACED. THE IMPELLA 5.5 WAS PLACED IN LEFT VENTRICLE AND INITIALLY WAS CAUGHT UP ON THE MITRAL VALVE. IT WAS MANEUVERED UNDER ECHO GUIDANCE AND ULTIMATELY WAS 4.63 CM IN THE CENTRAL LEFT VENTRICLE. THE DEVICE WAS NOTED TO BE WORKING WELL. ADDITIONAL LINES WERE PLACED AT THE END OF THE CASE WHICH CAUSED SEVERAL EPISODES OF VENTRICULAR TACHYCARDIA WHICH REQUIRED SHOCK. ONCE THE SWAN WAS PLACED, THE PATIENT WENT SYSTOLIC. A TRANSVENOUS PACING WIRE WAS PLACED AND ULTIMATELY RHYTHM WAS RESTORED AS WELL AS RETURN OF SPONTANEOUS CIRCULATION. IMPELLA 5.5 FLOWS INCREASED BACK TO 4.5 LPM. THE PATIENT HAD SIGNIFICANT METABOLIC ACIDOSIS LEAVING THE OPERATING ROOM WHICH CONTINUED TO BE CORRECTED. THREE DAYS LATER THE PATIENT WOULD EXPIRE. MEDICAL SAFETY REVIEW OF THE EVENT DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. WHETHER THE PATIENT RECOVERED OR NOT WITH ESCALATION IS INCONSEQUENTIAL BECAUSE THE PATIENT WAS VERY SICK AND EXPERIENCED "PROFOUND" BLEEDING AS A RESULT OF THE IMPELLA CP REQUIRING MULTIPLE BLOOD PRODUCTS, AND THE DEMISED OUTCOME CANNOT BE EXCLUDED FROM IMPELLA CP EVEN LONG TERM. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THIS EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-26084 FOR THE IMPELLA 5.5 DEVICE REPORT. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS: "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY" (INCLUDING VENTRICULAR PERFORATION). Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21385677&pc=MVK | 3008642652-2025-01659 | 2025-01-21 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-14 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 06:48:52 ON (B)(6) 2025. AT 17:44:03, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. AT 17:45:20, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. AT 17:46:34, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. AT 17:47:42, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE DEVICE WAS SHUT DOWN AT 17:52:39 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21385682&pc=NPT | 2015691-2025-01056 | 2024-11-16 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-14 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Obstruction/Occlusion | P140031 |  | 1 | Event Description: AS REPORTED THROUGH REVIEW OF RECEIVED MEDICAL RECORDS, REGARDING A 20MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION. PER MEDICAL RECORD REVIEW, THE PATIENT EXPERIENCED CHEST PAIN AND HYPOTENSION S/P TAVR. THE PATIENT WAS TAKEN BACK TO THE CATH LAB YIELDING NO SIGNIFICANT CAD. HYPOTENSION AND VASOPRESSOR SUPPORT CONTINUED. AN ECHO REVEALED MECHANICAL OBSTRUCTION OF THE LM CORONARY ARTERY. THE PATIENT UNDERWENT CABG. THE PATIENT HAD A PROGRESSIVELY DYSFUNCTIONAL COURSE COMPLICATED BY SEIZURE, LEFT HEMIPLEGIA AND DIFFUSE ENCEPHALOPATHY. A CT BRAIN SHOWED A LARGE RIGHT PARIETAL, FRONTAL, AND TEMPORAL ISCHEMIC STROKE AND LEFT PARIETAL, OCCIPITAL, INFERIOR FRONTAL STROKES AND PONTINE STROKE WITHOUT MASS EFFECT OR HEMORRHAGIC TRANSFORMATION. THE PATIENT WAS NOT A CANDIDATE FOR NEURO ENDOVASCULAR INTERVENTION WITH THE LARGE STROKE BURDEN. A CAROTID ULTRASOUND REVEALED A RIGHT DISTAL COMMON CAROTID ARTERY COMPLETE OCCLUSION/SUB-OCCLUSIVE WITH NO EVIDENCE OF SIGNIFICANT LEFT CAROTID STENOSIS. UNFORTUNATELY, THE PROGNOSIS WAS POOR. AFTER MEETING WITH THE PATIENTS FAMILY, PALLIATIVE CARE WAS DECIDED AND THE PATIENT PASSED AWAY 10 DAYS POST VALVE IMPLANTATION. IT IS TO BE NOTED A POST TAVR ECHO REVEALED A WELL-FUNCTIONING BIOPROSTHETIC AO VALVE. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO CLINICAL INVESTIGATION FINDINGS. SECTIONS: G3, G6, H2, H6 INVESTIGATION CONCLUSIONS, H11 HAVE BEEN UPDATED. REGARDING THE CORONARY OCCLUSION, PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE ROOT CAUSE FOR THE CORONARY OBSTRUCTION WAS MECHANICAL OBSTRUCTION OF THE LM CORONARY ARTERY, MOST LIKELY THE FRAME OF THE EWS VALVE, WHICH REQUIRED SURGERY. THE PATIENT HAD NO SIGNIFICANT CAD. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. REGARDING THE STROKE, PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE ROOT CAUSE FOR THE ISCHEMIC STROKE WAS THE TAVR ITSELF AND THE ADDITIONAL SURGERY REQUIRED TO CORRECT THE CORONARY OBSTRUCTION. SURGERY INCREASED THE RISK FOR EMBOLUS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: REGARDING THE CORONARY OCCLUSION, PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (CAD) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. REGARDING THE STROKE, PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (GENDER, AGE) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21386265&pc=DSQ | 2916596-2025-00703 | 2024-01-14 05:00 | Death | THORATEC CORPORATION | 2025-02-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RIGHT VENTRICULAR FAILURE (RVF) AND SUBSEQUENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21387049&pc=DSQ | 2916596-2025-00820 | 2024-06-08 04:00 | Death | THORATEC CORPORATION | 2025-02-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pneumonia; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO ASPIRATION PNEUMONIA.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS ON DRUG THERAPY, INTUBATION, AND VENTILATOR SUPPORT. THE PATIENT PASSED AWAY IN THE HOSPITAL. THE ASPIRATION PNEUMONIA WAS DUE TO A DECLINE IN ACTIVITIES OF DAILY LIVING (ADL) AND WAS NOT CONSIDERED TO BE DEVICE RELATED.Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2024 DUE TO ASPIRATION PNEUMONIA. THE PATIENT WAS PUT ON DRUG THERAPY, INTUBATION, AND VENTILATOR SUPPORT. THE PATIENT DEVELOPED RESPIRATORY FAILURE ON (B)(6) 2024 DUE TO THE ASPIRATION PNEUMONIA AND WAS INTUBATED FOR 16 DAYS. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT CORRECTED. SECTION D4: PRIMARY UDI CORRECTED. SECTION D6B: DATE OF EXPLANT CORRECTED. SECTION D9: CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), WAS RETURNED FOR EVALUATION AFTER THE PATIENT PASSED AWAY DUE TO ASPIRATION PNEUMONIA. THE ASPIRATION PNEUMONIA WAS DUE TO A DECLINE IN ACTIVITIES OF DAILY LIVING AND WAS NOT CONSIDERED TO BE DEVICE RELATED. THE PATIENT DEVELOPED RESPIRATORY FAILURE DUE TO THE PNEUMONIA AND WAS INTUBATED PRIOR TO PASSING AWAY IN THE HOSPITAL. EVALUATION OF THE RETURNED PUMP DID NOT REVEAL ANY FINDINGS THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 6¿ FROM THE PUMP HOUSING. THE MODULAR CABLE AND THE REMAINDER OF THE PUMP CABLE WERE NOT RETURNED. THE APICAL CUFF WAS NOT RETURNED. APPROXIMATELY 1¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE BEND RELIEF ENGAGED TO THE GRAFT HARDWARE. UPON DISASSEMBLY OF THE RETURNED PUMP, VISUAL INSPECTION OF THE BLOOD CONTACTING SURFACES DID NOT REVEAL ANY ADHERED DEPOSITIONS OR THROMBUS FORMATIONS. THE RETRIEVED LVAD (LEFT VENTRICULAR ASSIST DEVICE) LOG FILES FROM THE RETURNED DEVICE SHOWED THE PUMP OPERATING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS INFECTION AND RESPIRATORY FAILURE AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATIONS. SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4 - PATIENT WEIGHT NOT PROVIDED. B5 NARRATIVE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21387512&pc=NPT | 2025587-2025-01250 | 2025-01-23 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-14 05:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Calcified; Gradient Increase; Patient Device Interaction Problem | Cardiogenic Shock; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WERE HIGH GRADIENTS AND CALCIFIED LEAFLETS ASSOCIATED WITH THE 31MM COREVALVE TRANSCATHETER AORTIC VALVE APPROXIMATELY 9 YEARS AND 6 MONTHS FOLLOWING THE PROCEDURE. INTERVENTION WAS REQUIRED.Event Description: IT WAS REPORTED THAT THERE WERE HIGH GRADIENTS AND CALCIFIED LEAFLETS ASSOCIATED WITH THE 31MM COREVALVE TRANSCATHETER AORTIC VALVE APPROXIMATELY 9 YEARS AND 6 MONTHS FOLLOWING THE PROCEDURE. INTERVENTION WAS REQUIRED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT DIED WHILE BEING EVALUATED FOR TREATMENT/INTERVENTION. WHEN THE DEATH OCCURRED AND THE CAUSE OF DEATH WERE NOT PROVIDED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT WHILE THE PATIENT WAS PENDING VALVE-IN-VALVE TREATMENT, THE PATIENT SUBSEQUENTLY DIED DUE TO CARDIOGENIC SHOCK.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN; THE VALVE DID NOT CAUSE OR CONTRIBUTE TO THE CARDIOGENIC SHOCK AND SUBSEQUENT DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21387550&pc=NPT | 2025587-2025-01253 | 2023-12-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-14 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 11 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO ACUTE HEART FAILURE. DURING HOSPITALIZATION, THE PATIENT PASSED AWAY. PER THE PHYSICIAN, IT IS UNKNOWN WHETHER THE VALVE OR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: CORRECTION: D4 AND H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B3: DATE IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388124&pc=DSQ | 2916596-2025-00818 | 2025-01-27 05:00 | Death | THORATEC CORPORATION | 2025-02-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO CEREBROVASCULAR ACCIDENT (CVA). NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388288&pc=MHY | 2182207-2025-00460 | 2024-12-24 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-02-14 05:00 | MHY | IMPLANTABLE NEUROSTIMULATOR | Malposition of Device; Adverse Event Without Identified Device or Use Problem; Insufficient Information | Stroke/CVA; Fatigue; Headache; Hematoma; Hemorrhage/Bleeding; Intracranial Hemorrhage; Unspecified Infection; Paresis; Dizziness; Dysphasia; Ambulation Difficulties; Cognitive Changes; Confusion/ Disorientation; Hydrocephalus; Insufficient Information; No Clinical Signs, Symptoms or Conditions | P960009 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING "COMPARATIVE ANALYSIS OF HEMORRHAGIC COMPLICATIONS IN DBS: MICROELECTRODE RECORDING VS. MACROSTIMULATION - DOES THE TECHNIQUE MATTER?" BACKGROUND: HEMORRHAGE IS A DANGEROUS COMPLICATION OF DEEP BRAIN STIMULATION (DBS) SURGERY, BUT A COMPREHENSIVE EXPLANATION OF THE A SSOCIATED RISK FACTORS REMAINS INCONCLUSIVE, PARTICULARLY APPLICATION OF MICROELECTRODE RECORDING (MER) COMPARED WITH MACROSTIMULATION (NON-MER)-BASED DBS PROCEDURES. WE CONDUCTED A COMPARATIVE ANALYSIS OF MER AND MACROSTIMULATION TECHNIQUES, EVALUATING THE IMPACT OF BRAIN PENETRATIONS BY MICROELECTRODE GUIDING CANNULAS ON THE OCCURRENCE OF INTRACRANIAL HEMORRHAGIC EVENTS. METHODS: THIS RETROSPECTIVE STUDY INCLUDED ALL DBS PROCEDURES (MER-BASED AND NON-MEREBASED) PERFORMED AT A SINGLE MEDICAL CENTER FROM NOVEMBER 2008 TO JUNE 2023. PATIENTS UNDERWENT COMPREHENSIVE PREOPERATIVE AND POSTOPERATIVE EVALUATIONS, INCLUDING POSTOPERATIVE COMPUTED TOMOGRAPHY TO ASSESS LEAD PLACEMENT AND INTRACRANIAL BLEEDING, CATEGORIZED AS SYMPTOMATIC OR ASYMPTOMATIC. STATISTICAL ANALYSES WERE CONDUCTED TO ASCERTAIN THE CORRELATION BETWEEN DIFFERENT SURGICAL TECHNIQUES AND THE RISK OF HEMORRHAGIC COMPLICATIONS. RESULTS: IN A COHORT OF 618 PATIENTS, 1096 ELECTRODES WERE IMPLANTED DURING 672 CONSECUTIVE DBS PROCEDURES. THERE WERE 21 ASYMPTOMATIC AND 6 SYMPTOMATIC INTRACRANIAL HEMORRHAGES (ICHS), INCLUDING 1 DEATH. THE BLEEDING RATE WAS 4.37% PER PATIENT AND 2.92% PER ELECTRODE. PATIENTS WHO DEVELOPED HEMATOMAS WERE MORE PREVALENT IN THE MER GROUP (25 CASES) COMPARED WITH THE NON-MER GROUP (2 CASES). THE MER GROUP HAD A HIGHER MEAN NUMBER OF BRAIN PENETRATIONS BY GUIDING CANNULAS. CONCLUSIONS: DBS PROCEDURES ARE GENERALLY SAFE, WITH ONLY 0.97% RESULTING IN LONG-TERM OR PERMANENT DEFICITS. ALL SYMPTOMATIC ICH OC CURRED IN MER-BASED PROCEDURES, WHEREAS MACROSTIMULATION WAS ASSOCIATED WITH ONLY 2 ASYMPTOMATIC ICH CASES. THE HIGHER NUMBER OF BRAIN PENETRATIONS IN MER CORRELATES WITH HIGHER BLEEDING RATES. REPORTED EVENTS: OF THE 27 PATIENTS WITH INTRACRANIAL HEMORRHAGE (ICH), 22 PATIENTS HAD PARKINSON'S DISEASE, IN WHICH, 5 PATIENTS EXPERIENCED PERMANENT DEFICITS AFTER THE ICH AND 1 DEATH. ALL PATIENTS DIAGNOSED WITH ICH WERE CLOSELY MONITORED AND MANAGED IN THE INTENSIVE CARE UNIT. ALL PATIENTS WITH SYMPTOMATIC ICH, WHO SURVIVED THE OPERATION, REQUIRED COMPREHENSIVE REHABILITATION THROUGHOUT THEIR EXTENDED HOSPITALIZATION IN THE NEUROSURGERY DEPARTMENT, FOLLOWED BY A NEUROLOGIC REHABILITATION FACILITY AT THE INSTITUTION. WITHIN THE STUDY GROUP, 5 PATIENTS SHOWED HEMORRHAGE ASSOCIATED WITH BOTH ELECTRODES, ACCOUNTING FOR 32 ELECTRODES DIRECTLY LINKED TO INTRACRANIAL BLEED AND AN OVERALL BLEEDING RATE OF 2.92% PER IMPLANTED ELECTRODE. - PATIENT 12: 66/M, TYPE OF ICH AND LOCATION: INTRAVENTRICULAR BLEEDING INTO THE THIRD, FOURTH, AND LATERAL VENTRICLES, INTRAPARENCHYMAL BLEEDING IN THE LEFT FRONTAL LOBE, BLEEDING TO CEREBELLOMEDULLARY CISTERN, RESULTING IN DEATH. Manufacturer Narrative: B3. PLEASE NOTE THAT THIS DATE IS BASED OFF OF THE DATE THAT THE ARTICLE WAS ACCEPTED FOR PUBLICATION AS THE EVENT DATES WERE NOT PROVIDED IN THE PUBLISHED LITERATURE. B5. IT WAS NOT POSSIBLE TO ASCERTAIN SPECIFIC DEVICE INFORMATION FROM THE ARTICLE OR TO MATCH THE EVENTS REPORTED WITH PREVIOUSLY REPORTED EVENTS. CORRESPONDENCE WILL BE SENT TO THE AUTHOR OF THE ARTICLE INQUIRING ABOUT INDIVIDUAL PATIENT INFORMATION AND ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS. D10. SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: NEU_UNKNOWN_LEAD, SERIAL/LOT #: UNKNOWN, CITATION: SOBSTYL, M., KARAMON, K. STAPINSKA-SYNIEC, A. COMPARATIVE ANALYSIS OF HEMORRHAGIC COMPLICATIONS IN DBS: MICROELECTRODE RECORDING VS. MACROSTIMULATION - DOES THE TECHNIQUE MATTER? WORLD NEUROSURG. 123634 (2024) DOI: 10.1016/J.WNEU.2024.123634. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388353&pc=OOB | 2124215-2025-08549 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIOGENIC SHOCK, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388364&pc=OOB | 2124215-2025-08518 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: SUDDEN CARDIAC DEATH. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388365&pc=LOX | 2124215-2025-08553 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIOGENIC SHOCK, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388381&pc=OOB | 2124215-2025-08522 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388382&pc=NIQ | 2124215-2025-08526 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388384&pc=NWX | 2124215-2025-08534 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIOGENIC SHOCK, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388385&pc=LOX | 2124215-2025-08524 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388386&pc=OOB | 2124215-2025-08535 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE; CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388572&pc=NPT | 2015691-2025-01075 | 2024-11-09 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-14 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 |  | 1 | Event Description: AS REPORTED BY PATIENT SUPPORT, THE CARE ADVOCATE INDICATED THE PATIENT WAS FOUND UNCONSCIOUS ON POSTOPERATIVE DAY (POD) 3 POST TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 26MM SAPIEN 3 ULTRA VALVE IN THE NATIVE AORTIC POSITION. THE PATIENT WAS TAKEN TO THE EMERGENCY ROOM AND A SUBSEQUENT COMPUTED TOMOGRAPHY (CT) SCAN REVEALED THE PATIENT HAD AN ISCHEMIC STROKE. THE PATIENT PASSED AWAY ON POD 4. ACCORDING TO THE DEATH CERTIFICATE, THE CAUSE OF DEATH IS ISCHEMIC STROKE AND PULMONARY HYPERTENSION. Manufacturer Narrative: THV/TVT REGISTRY PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS ARE INCONCLUSIVE DUE TO THE LIMITED INFORMATION PROVIDED. HOWEVER, PATIENT FACTORS MIGHT HAVE CONTRIBUTED TO THE EVENT, INCLUDING THE PATIENT'S ADVANCED AGE OF 77 YEARS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388976&pc=NGV | 2124215-2025-06077 | 2025-01-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | NGV | WATCHMAN FLX PRO | Device Dislodged or Dislocated | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A DEVICE EMBOLIZATION AND A PATIENT DEATH. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 35MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. IT WAS NOTED THERE WAS A GROIN COMPLICATION POST PROCEDURE WHICH REQUIRED ADDITIONAL INTERVENTION, BUT FURTHER DETAIL BEYOND THAT WAS NOT PROVIDED. SIX (6) DAYS POST PROCEDURE THE CLOSURE DEVICE EMBOLIZED TO THE MITRAL VALVE OF THE PATIENT AND THE PATIENT WENT INTO CARDIAC ARREST. THE PHYSICIAN SNARED THE CLOSURE DEVICE FROM THE MITRAL VALVE BUT THEN IT MIGRATED TO THE DESCENDING AORTA. ANOTHER DEVICE WAS THEN USED TO RETRIEVE THE CLOSURE DEVICE TO A LARGER SHEATH WHERE IT WAS THEN REMOVED FROM THE PATIENT. DURING THE RETRIEVAL OF THE CLOSURE DEVICE THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS TO TREAT THE EFFUSION AND DRAIN BLOOD FROM THE PATIENT. THE PATIENT WAS LEFT ON BLOOD PRESSERS AND WAS TRANSFERRED TO THE ICU. THE PATIENT DID NOT RECOVERED, AND AS A RESULT THE PATIENT DIED FROM COMPLICATIONS OF THIS EVENT.Event Description: IT WAS REPORTED THAT THERE WAS A DEVICE EMBOLIZATION AND A PATIENT DEATH. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 35MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. IT WAS NOTED THERE WAS A GROIN COMPLICATION POST PROCEDURE WHICH REQUIRED ADDITIONAL INTERVENTION, BUT FURTHER DETAIL BEYOND THAT WAS NOT PROVIDED. SIX (6) DAYS POST PROCEDURE THE CLOSURE DEVICE EMBOLIZED TO THE MITRAL VALVE OF THE PATIENT AND THE PATIENT WENT INTO CARDIAC ARREST. THE PHYSICIAN SNARED THE CLOSURE DEVICE FROM THE MITRAL VALVE BUT THEN IT MIGRATED TO THE DESCENDING AORTA. ANOTHER DEVICE WAS THEN USED TO RETRIEVE THE CLOSURE DEVICE TO A LARGER SHEATH WHERE IT WAS THEN REMOVED FROM THE PATIENT. DURING THE RETRIEVAL OF THE CLOSURE DEVICE THE PATIENT EXPERIENCED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS TO TREAT THE EFFUSION AND DRAIN BLOOD FROM THE PATIENT. THE PATIENT WAS LEFT ON BLOOD PRESSERS AND WAS TRANSFERRED TO THE ICU. THE PATIENT DID NOT RECOVER, AND AS A RESULT THE PATIENT DIED FROM COMPLICATIONS OF THIS EVENT. Manufacturer Narrative: DEVICE EVAL BY MFR.: THE RETURNED DEVICE CONSISTED OF A WATCHMAN FLX PRO IMPLANT ONLY WITH THE FABRIC SATURATED IN BLOOD. INSPECTION FOUND FRACTURED STRUTS AND TORN FABRIC. PRODUCT ANALYSIS COULD NOT CONFIRM THE REPORTED EVENT AS CLINICAL CIRCUMSTANCES COULD NOT BE REPLICATED. MEDICAL MEDIA WAS PROVIDED TO AID IN THE INVESTIGATION AND WAS REVIEWED BY MEMBERS OF THE BSC TRAINING TEAM, MEDICAL SAFETY, AND CLINICAL SPECIALISTS. THE MEDIA REVIEW CONCLUDED: THERE WAS ONE SINGLE IMAGE PROVIDED THAT SHOWED THE RETRIEVED DEVICE ON THE EXAMINATION TABLE. THE DEVICE WAS DAMAGED DUE TO THE EXTRACTION MANEUVER. A PROCEDURAL TRANSESOPHAGEAL ECHO (TEE) IMAGING DATED (B)(6) 2025, WAS REVIEWED AND CONCLUDED: THE RELEASED DEVICE WAS NOT SHOWN BUT THE POSITION OF THE DEPLOYED DEVICE SEEMED ADEQUATE AT LOW ANGLE VIEWS BUT INCREASED TO ABOUT 1/3 AT THE 135-DEGREE VIEW. THE COMPRESSION WAS MEASURED AT ABOUT 20%. IN LOWER ANGLES VIEW, THE DEVICE SEEMED TO SIT UNDER A LIMBUS OVERHANG WHICH COULD RENDER THE TUG TEST LESS RELIABLE, BUT THE TUG TEST ITSELF WAS NOT SHOWN. OVERALL, THERE WAS NO CLEAR INDICATION OF WHAT CONTRIBUTED TO THE DEVICE EMBOLIZATION DETECTED SIX (6) DAYS AFTER IMPLANT. A TRANSTHORACIC ECHO (TTE) DATED (B)(6) 2025, WAS REVIEWED AND CONCLUDED: AN EMBOLIZED DEVICE WAS ENTANGLED IN THE MITRAL VALVE CHORDAE WHICH SEEMED TO PARTIALLY OBSTRUCT THE LEFT VENTRICLE OUTFLOW TRACT (LVOT) DURING SYSTOLE. NO PERICARDIAL EFFUSION WAS SHOWN IN THIS IMAGING WHICH WAS LIKELY PERFORMED BEFORE DEVICE EXTRACTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388977&pc=LOX | 2124215-2025-08532 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIOGENIC SHOCK, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21388978&pc=NIQ | 2124215-2025-08546 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-14 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21389150&pc=PRL | 2017233-2025-05836 | 2025-01-16 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-14 05:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Failure to Advance | Low Blood Pressure/ Hypotension; Paralysis; Paraplegia | P160021 |  | 1 | Event Description: ON (B)(6) 2025, PATIENT UNDERWENT A THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) PROCEDURE TO TREAT AN ANEURYSM UTILIZING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS AND REDUCED PROFILE GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS DEVICES (VBX DEVICES). DURING THE PROCEDURE, THERE WAS DIFFICULTY ADVANCING A 6 MM X 79 MM VBX DEVICE INTO THE LEFT RENAL. IT WAS REPORTED, THE ANGLE AT THE OSTIUM OF THE RENAL ARTERY WAS ACUTE AND ARTERY AFTER THE BIFURCATION WAS SHORT, THEREFORE IT WAS DIFFICULT TO GET ENOUGH WIRE PURCHASE INTO THE LANDING ZONE. THERE WERE MULTIPLE ATTEMPTS MADE TO RE-ACCESS THE ARTERY. IT WAS REPORTED, THE PHYSICIAN CHOSE TO USE A SHORTER STENT (BRAND UNKNOWN), AND IT TRACKED AND DEPLOYED INTO THE RENAL ARTERY. THE PHYSICIAN SUCCESSFULLY TREATED THE SUPERIOR MESENTERIC ARTERY (SMA) AND THE CELIAC ARTERY WITH NO ISSUES. IT WAS REPORTED THERE WAS NO VBX DEVICE IMPLANTED INTO THE CELIAC ARTERY DUE TO OCCLUSION SHOWN ON PREOPERATIVE SCAN. IT WAS REPORTED, THE PHYSICIAN USED A 46 MM CODA BALLOON TO CREATE A BACKBOARD AND A 1 CM TIP AMPLATZ WIRE, IN WHICH PHYSICIAN WAS ABLE TO BRIDGE A VBX DEVICE IN THE RIGHT RENAL ARTERY TO THE TAMBE DEVICE. PATIENT TOLERATED THE PROCEDURE. AN ANGIOGRAPHY WAS PERFORMED, AND THE PHYSICIAN WAS SATISFIED WITH THE RESULTS. THE PATIENT THEN PRESENTED BACK LATER THAT AFTERNOON AND BLEEDING IN THE LEFT KIDNEY WAS OBSERVED. THE PHYSICIAN SUSPECTS THE CAUSE OF BLEEDING IN THE LEFT KIDNEY WAS A RESULT FROM TRAUMA FROM THE NON-GORE WIRE. PATIENT BECAME HYPOTENSIVE, WITH SOME LIMB ISCHEMIA AS WELL. FEMORAL ENDARTERECTOMY WAS PERFORMED. AFTER BECOMING HYPOTENSIVE, PATIENT ALSO EXPERIENCED SOME PARAPLEGIA AND SHOCK LIVER WAS ALSO REPORTED. NO OTHER RECOVERY PROCEDURES WERE PERFORMED AND PATIENT PASSED AWAY THE FOLLOWING DAY ON (B)(6) 2025. IT WAS ALSO REPORTED CAUSE OF ISCHEMIA DID NOT INVOLVE A DEVICE, AS ALL DEVICES WERE PATENT. A CLOT WAS JUST OBSERVED AT SOME POINT DURING THE PROCEDURE. PHYSICIAN BELIEVES THE PARAPLEGIA IS DUE TO HYPOTENSION, AS A RESULT FROM HEMORRHAGIC EVENT OF THE LEFT KIDNEY. PATIENT DEATH WAS DUE TO SHOCK LIVER. NO FURTHER PATIENT INFORMATION OR IMAGES ARE AVAILABLE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. A2: PATIENT AGE/DOB WAS REQUESTED BUT NOT MADE AVAILABLE. A3A/A3B: PATIENT SEX/GENDER REQUESTED BUT NOT MADE AVAILABLE. A4: PATIENT WEIGHT WAS REQUESTED BUT NOT MADE AVAILABLE. H6: CODE B20: DEVICE REMAINS IMPLANTED, AND NO IMAGES WERE PROVIDED; THEREFORE, DIRECT PRODUCT ANALYSIS WAS NOT POSSIBLE. H6: CODE C21 RESULTS PENDING COMPLETION OF INVESTIGATION (DEVICE HISTORY RECORD REVIEW PENDING). W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. H6: CODE C19 A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. INVESTIGATION CONCLUSIONS: D15 AND D12 REMOVED AND ADDED D1001. ENGINEERING EVALUATION: THE PRIMARY REPORTED COMPLAINT COULD NOT BE INDEPENDENTLY CONFIRMED BECAUSE NO PRODUCT NOR IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. THEREFORE, AN INDEPENDENT ASSESSMENT OF PRODUCT PERFORMANCE COULD NOT BE CONDUCTED. AS REPORTED, THE DEVICE WAS UNABLE TO BE ADVANCED; THE ANGLE AT THE OSTIUM OF THE RENAL ARTERY WAS ACUTE AND THE ARTERY AFTER THE BIFURCATION WAS SHORT, THEREFORE IT WAS DIFFICULT TO GET ENOUGH WIRE PURCHASE INTO THE LANDING ZONE. THEREFORE, THE CAUSE FOR THE PRIMARY REPORTED COMPLAINT CAN BE ATTRIBUTED TO THE PATIENT CONDITION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21389630&pc=OZD | 1220648-2025-26155 | 2025-01-25 05:00 | Death | ABIOMED, INC. | 2025-02-14 05:00 | OZD | IMPELLA 5.5 | Pumping Stopped | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA 5.5 PUMP PLACED TO SUPPORT A PATIENT IN FOR SURGERY. THE PUMP WAS SUPPORTING FOR THREE DAYS WHEN THE PUMP STOPPED AND COULD NOT BE RESTARTED DESPITE ALL TROUBLESHOOTING MEASURES. THE MEDICAL DOCTOR AT BEDSIDE DISCONTINUED IMPELLA, SURGICAL CLOSURE AND STAPLED SHUT. THE FAMILY DECIDED TO WITHDRAW CARE AS THE PATIENT STATUS CONTINUED TO DECLINE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: DIRECT AORTIC INSERTION ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE PUMP STOP HAS BEEN COMPLETED. THE PUMP WAS INSPECTED AND BIOMATERIAL WAS FOUND IN THE OUTLET CAGE AND AROUND THE IMPELLER. THE BIOMATERIAL DID NOT CLEAR UPON SOAKING. THE DATA LOGS FROM THE CASE SHOWED THE IMPELLA STOPPED - MOTOR CURRENT HIGH ALARMS. THERE WAS A MOTOR CURRENT SPIKE WITH SPEED DEVIATION AND PLACEMENT SIGNAL SHIFT CHARACTERISTIC OF BIOMATERIAL INGESTION. THE ROOT CAUSE OF THE PUMP STOP WAS DETERMINED TO BE DUE TO BIOMATERIAL INGESTION. B.7 INFORMATION WAS ADDED THAT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26155. H.6 ADDED CODE TO HEALTH EFFECT - IMPACT CODE AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26155.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21389643&pc=LWR | 3004478276-2025-00120 | 2024-11-30 05:00 | Death | CORCYM CANADA CORP. | 2025-02-14 05:00 | LWR | PERCEVAL SUTURELESS AORTIC HEART VALVE | Insufficient Information | Insufficient Information | P150011 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED OF THE FOLLOWING EVENT THROUGH PERCEVAL JAPAN CLINICAL STUDY. REPORTEDLY, A PERCEVAL VALVE, PVS25 WAS IMPLANTED IN THE PATIENT ON (B)(6) 2019. BASED ON INFORMATION RECEIVED, PATIENT DIED DUE TO UNKNOWN REASON ON (B)(6) 2024. IT IS UNKNOWN IF THE EVENT IS RELATED TO THE IMPLANTED DEVICE AND TO THE INITIAL IMPLANT PROCEDURE. NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE MANUFACTURING AND MATERIAL RECORDS FOR THE DEVICE INVOLVED IN THIS EVENT, AS THEY PERTAIN TO THE REPORTED ISSUE, WERE RETRIEVED AND REVIEWED BY THE MANUFACTURER¿S QUALITY ENGINEERING. THE RESULTS CONFIRMED THAT THIS PROSTHESIS SATISFIED ALL REQUIRED MATERIAL, VISUAL, AND PERFORMANCE STANDARDS AT THE TIME OF MANUFACTURE AND RELEASE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEFINITIVE ROOT CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED AT THIS TIME. HOWEVER, FROM THE DOCUMENT REVIEW PERFORMED, NO MANUFACTURING DEFICIENCIES WERE NOTED. SHOULD FURTHER INFORMATION BE RECEIVED IN THE FUTURE, A FOLLOW UP REPORT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21389844&pc=PYX | 1220648-2025-26156 | 2025-01-23 05:00 | Death | ABIOMED, INC. | 2025-02-14 05:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate; Pumping Stopped; Failure to Sense | Hemorrhage/Bleeding | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS INITIALLY IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MEDICAL CIRCULATORY SUPPORT (MCS) AND UNEXPECTEDLY THE PUMP FLOW STARTED DECREASING WITH INACCURATE PLACEMENT SIGNAL NUMBERS AND THEN PUMP FAILED. THE RP FLEX WAS EMERGENTLY REPLACED WITH BLEEDING AT THE ACCESS SITE. THE PATIENT WOULD EXPIRE THE NEXT DAY. HOWEVER, PRIOR THE PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR BIPELLA SUPPORT AND WOULD EXPERIENCE LIMB ISCHEMIA. THE PATIENT CAME PRESENTED TO THE EMERGENCY ROOM AS A ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND WAS EMERGENTLY TAKEN TO THE CATHETERIZATION LAB. THREE STENTS WERE IMPLANTED IN THE RIGHT CORONARY ARTERY WITH PATIENT CODING. THE DECISION WAS MADE TO IMPLANT AN IMPELLA P FLEX AFTER PUTTING IN A SWAN AND SEEING SIGNS OF RIGHT VENTRICULAR FAILURE. THE PATIENT WAS TRANSFERRED TO THE UNIT ON IMPELLA MCS. APPROXIMATELY 10 HOURS LATER THE PATIENT NEEDED LEFT VENTRICULAR UNLOADING WITH A LEFT SIDED IMPELLA CP DEVICE. THE PATIENT SENT BACK TO THE UNIT NOW ON BIPELLA MCS. THE PLAN WAS TO UNLOAD AND REST LEFT VENTRICLE, TREND DOWN DRIPS AS ABLE, AND KEEP UP WITH VOLUME. PATIENT TO PLACED ON CONTINUOUS RENAL REPLACEMENT THERAPY IF URINE OUTPUT REMAINED LOW. THE PATIENT WAS NOTED TO BE IN VERY CRITICAL CONDITION WITH NO LONG-TERM PLAN, AS POLY SUBSTANCE USE EXCLUDES FROM SOME LONG-TERM THERAPIES. CONSTANT DISCUSSIONS WERE GOING ON WITH FAMILY AND PHYSICIANS TO DETERMINE NEXT STEPS IF PATIENT STARTS IMPROVING AFTER BIPELLA SUPPORT. IT WAS ALSO NOTED THE PATIENT HAD MULTIPLE BLEEDS, SOME OF UNKNOWN ORIGIN. CHEST AND ABDOMEN DISTENDED IN MULTIPLE AREAS, SHOWING SIGNS OF BLEEDING FROM INTERNAL JUGULAR RP FLEX ACCESS SITE AND OTHER MAJOR BLEEDING IN MULTIPLE AREAS OF UNKNOWN ORIGIN. ALBUMIN, AND SEVERAL UNITS OF BLOOD WERE GIVEN (4 UNITS OF PACKED RED BLOOD CELLS). EVENTUALLY A MASSIVE TRANSFUSION PROTOCOL WAS STARTED. LATER IN THE DAY NOW APPROXIMATELY ONE DAY AFTER START OF IMPELLA RP FLEX SUPPORT, RP FLEX FLOWS AND CATHETER WERE RUNNING AS EXPECTED, THEN SUDDENLY STARTED DROPPING RAPIDLY WITH INACCURATE PLACEMENT SIGNAL NUMBERS DISPLAYING. THE FLOW DROPPED FROM 3.0-3.5, THEN DROPPED DOWN TO 2.5 AND KEPT DROPPING DOWN OVER A FEW MINUTES EVENTUALLY DOWN TO 0.0. THERE WERE NO ALARMS, AND MOTOR CURRENT WAS SHOWING THE MOTOR WAS RUNNING. THE PERFORMANCE LEVEL WAS STILL P-9, AND WAS BROUGHT DOWN. THE PERFORMANCE LEVEL WAS DROPPED TO P-0, THEN INCREASED BACK UP TO P-4 AS AN ATTEMPT TO CORRECT THE ISSUE BUT THIS WAS UNSUCCESSFUL; THERE WAS NO EFFECT. AFTER THE PUMP STOPPED WORKING, FAILED, PATIENT WAS BROUGHT DOWN TO THE CATH LAB TO REPLACE WITH ANOTHER IMPELLA RP FLEX, ATTEMPTING TO USE THE SAME SITE FIRST, AND IF NOT SUCCESSFUL, MOVE TO A NEW SITE IN THE RIGHT FEMORAL VEIN. THE PHYSICIAN COMMENTED: WHEN INFORMED OF THE SITUATION, QUICKLY SAID TO REPLACE THE DEVICE. FURTHER, THE PHYSICIAN MENTIONED THE INTENSIVE CARE UNIT DID NOT START SYSTEMIC HEPARIN WHICH HE WAS UNAWARE OF; HOWEVER, THE ICU SAID DUE TO THE LAB VALUES, IT WAS FELT ADDITIONAL HEPARIN WAS NOT NEEDED FOR THE PATIENT AT SUCH TIME. WHEN PATIENT CAME DOWN FOR REPLACEMENT OF THE FAILED RP FLEX CATHETER, THE DISTAL RIGHT EXTREMITY WAS NOTED TO BE SLIGHTLY MOTTLED BUT HAD A PULSE WHEN CHECKED PREVIOUSLY WHILE ON THE UNIT. AT THE END OF THE RP FLEX REPLACEMENT, THE LEG LOOKED WORSE AND MORE MOTTLED. AN ANGIOGRAM WAS PERFORMED SHOWING BLOCKAGE OF FLOW PASSED THE IMPELLA REPO SHEATH INSERTION SITE. AN EXTERNAL BYPASS FOR DISTAL PERFUSION WAS DISCUSSED; HOWEVER, BEFORE GETTING ACCESS FOR IT, THE PLAN WAS CHANGED TO REMOVE THE CURRENT IMPELLA CP AND PLACE A NEW ONE ON THE LEFT SIDE. AFTER REMOVAL ON THE RIGHT, AND PLACEMENT OF A NEW IMPELLA CP ON THE LEFT, BOTH EXTREMITIES HAD ADEQUATE DISTAL FLOW WITH PULSES CONFIRMED WITH DOPPLER AND THE COLOR OF THE RIGHT LOWER EXTREMITY IMPROVED. DURING THE SWAP OF THE RP FLEX A LARGE HEMATOMA DEVELOPED FROM SWAPPING THE RP FLEX CATHETERS AT THE SAME SITE REQUIRING TWO UNITS OF BLOOD. THE PATIENT NOW BACK ON BIPELLA SUPPORT WITH A NEWLY PLACED IMPELLA CP. HOWEVER, FLOWS LOWER THAN EXPECTED DESPITE PLENTY OF VOLUME GOING IN AND FULL RP FLEX FLOW AT PERFORMANCE LEVEL P-9. DISCUSSION WAS MADE REGARDING POSSIBLE MAJOR BLEEDS AND CHEST/ABDOMEN VERY DISTENDED IN MULTIPLE AREAS. FAMILY WAS INFORMED OF CRITICAL STATUS OF PATIENT AND URGENCY OF END-OF-LIFE DISCUSSIONS IF SITUATION DID NOT DRAMATICALLY IMPROVE VERY QUICKLY. THE PATIENT WOULD EXPIRE THIS SAME DAY. MEDICAL SAFETY REVIEW OF THE EVENT DETERMINED THREE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE RP FLEX THAT FAILED AS AN ASSOCIATED FACTOR, LEADING TO THE DEMISED OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW, PLACEMENT SIGNAL ISSUE, AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE LOW PUMP FLOW FAILURE MODE WAS REMOVED AS REASON FOR REPORTED LOW PUMP FLOW IS ENCAPSULATED UNDER THE PLACEMENT SIGNAL ISSUE. THE RETURNED PRODUCT WAS INSPECTED AND THERE WAS NO BIOMATERIAL IN THE INLET CAGE AND ON THE SENSOR. THE DATA LOGS WERE REVIEWED AND DP SENSOR DRIFT WAS OBSERVED WITH A RAPID DRIFT UP OF PLACEMENT SIGNAL SIMULTANEOUS WITH A RAPID DRIFT DOWN OF PUMP FLOWS. PLACEMENT SIGNAL AND CVP WERE BOTH NOISY WITH CVP NEGATIVE FOR MOST OF THE SUPPORT DURATION BEFORE TRENDING UP DURING PS AND FLOW DRIFT. THE PUMP WAS TESTED IN THE FLOW LOOP AND THE SENSOR DRIFT REPRODUCED WITH #1051 PLACEMENT SIGNAL NOT RELIABLE - DP SENSOR ALARM. THE ROOT CAUSE OF THE PLACEMENT SIGNAL ISSUE WAS DETERMINED TO BE DUE TO SENSOR DRIFT AS ISSUE REPRODUCED ON RETURNED PRODUCT. THE ROOT CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS NOT DETERMINED AS LIMITED CLINICAL INFORMATION RELATED TO FAILURE WAS PROVIDED B1 WAS REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26156 TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26156 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H6 MEDICAL DEVICE PROBLEM CODE 1503 WAS INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26156. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF THREE DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THIS EVENT AND THIS REPORT IS FOR THE INITIALLY IMPLANTED IMPELLA RP FLEX DEVICE THAT FAILED. REFER TO THE FOLLOWING RELATED INITIAL MEDICAL DEVICE REPORTS: IMPELLA RP FLEX SERIAL NUMBER (B)(6) WITH DATE OF EVENT 23-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 24-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). IMPELLA RP FLEX SERIAL NUMBER (B)(6) WITH DATE OF EVENT 24-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21371278&pc=MIH | 2017233-2025-05825 | 2025-02-01 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-13 05:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Perforation of Vessels | P040043 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2024, THE PATIENT UNDERWENT AN URGENT ENDOVASCULAR TREATMENT OF A RUPTURE OF GRAFT ANASTOMOSIS ANEURYSM AFTER A DESCENDING AORTA GRAFT REPLACEMENT USING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. ON (B)(6), 2025, A RUPTURE OF AN AORTA WAS OBSERVED. THE RUPTURE WAS OBSERVED WHERE MORE DISTAL OF THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. REPORTEDLY, AN INFECTION WAS ALSO CONFIRMED, BUT IT IS UNKNOWN WHEN THE INFECTION OCCURRED AND THE ROUTE OF THE INFECTION. AN EMERGENCY REINTERVENTION (TEVAR) WAS PLANNED ON (B)(6) 2025. HOWEVER, THE PATIENT DIED BEFORE THE REINTERVENTION. IT IS REPORTED THAT THE INFECTION MIGHT HAVE CAUSED THE RUPTURE OF AORTA. Manufacturer Narrative: B3: THE ACTUAL EVENT DATE IS UNKNOWN, THEREFORE 01-FEB-2025 IS USED AS THE EVENT DATE. H6: CODE D12: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: INFECTION, AORTIC RUPTURE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21371667&pc=OZP | 2032227-2025-125753 | 2024-12-04 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-13 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Cytomegalovirus (CMV) Infection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS RELATED TO THE CUSTOMER UNDERGOING STERNOTOMY SURGERY. NO CONTRIBUTING FACTORS WERE IDENTIFIED FROM THE PREEXISTING MEDICAL CONDITIONS. THE INSULIN PUMP WAS IN USE AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1884, UNOMEDICAL, MMT-7841ZN, MMT-7040A. TROUBLESHOOTING WAS PERFORMED. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A.Event Description: UPDATED SUMMARY: THE CAUSE OF DEATH WAS RELATED TO THE CUSTOMER SPLEEN RUPTURED. CUSTOMER DID NOT SURVIVE SURGERY TO REMOVE RUPTURED SPLEEN .UNDERGOING STERNOTOMY SURGERY. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INFORMATION HAS BEEN CORRECTED WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. THE INFORMATION HAS BEEN PROVIDED IN SECTION B5 AND H6 (HEALTH EFFECT - CLINICAL CODE & HEALTH EFFECT - IMPACT CODE) WITH THIS REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21373442&pc=DTB | 3008973940-2025-01059 | 2025-01-21 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-02-13 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE WHEN THE RIGHT VENTRICULAR (RV) LEAD WAS PLACED THE PATIENT HEART RATE DECREASED E EVENTUALLY LEADING TO CARDIAC ARREST. MEDICATION WAS ADMINISTERED, PERCUTANEOUS PACING AND TEMPORARY PACING WERE PERFORMED HOWEVER THE HEART RATE DID NOT RECOVER AND THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21373517&pc=NPT | 2015691-2025-00977 | 2024-11-21 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-13 05:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Fluid/Blood Leak; Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: IT WAS REPORTED BY PATIENT SUPPORT (PCS), THAT APPROXIMATELY 2 WEEKS AND 3 DAYS POST TAVR WITH A 29MM SAPIEN 3 VALVE THE PATIENT DIED OF A STROKE. THE PATIENT DID NOT RECOVER FROM THE IMPLANTING SURGERY AND HAD BEEN DISCHARGED HOME TO HOSPICE WHERE HE PASSED AWAY. PER MEDICAL RECORDS, THE ECHO SHOWED THE PATIENT HAD MODERATE REGURGITATION OF THE VALVE POST OPERATIVE DAY 1. THE PATIENT HAD CONGESTIVE HEART FAILURE AND WAS A FORMER SMOKER. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING AND CLINICAL EVALUATION FINDINGS. SECTIONS G6, H2, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND CONCLUSIONS IN THIS SECTION HAVE BEEN UPDATED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE COMPLAINT FOR REGURGITATION WAS CONFIRMED BASED ON THE PROVIDED MEDICAL RECORDS. A REVIEW OF THE DHR, LOT HISTORY AND COMPLAINT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF THE IFU REVEALED NO DEFICIENCIES. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. PER THE INSTRUCTIONS FOR USE (IFU), VALVE REGURGITATION, INCLUDING PARAVALVULAR LEAK (PVL) AND TRANSVALVULAR LEAK (CENTRAL), ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (THV) PROCEDURE. CENTRAL REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO SEVERAL ISSUES INCLUDING PATIENT-RELATED FACTORS, STRUCTURAL VALVE DETERIORATION (E.G. CALCIFICATION, LEAFLET THICKENING), AND/OR NONSTRUCTURAL DYSFUNCTION (FIBROSIS, PANNUS FORMATION). PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE CARDIAC TISSUE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. DUE TO LIMITED INFORMATION PROVIDED; A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME.PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST (HYPOTHYROIDISM, HYPERCHOLESTEROL AND SMOKING) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21373866&pc=BZD | 2518422-2025-102006 | 2024-12-13 05:00 | Death | RESPIRONICS, INC. | 2025-02-13 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21374079&pc=HAW | 1723170-2025-00653 | 2024-12-24 05:00 | Death | MEDTRONIC NAVIGATION, INC | 2025-02-13 05:00 | HAW | MEDTRONIC NAVIGATION | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: SOBSTYL, M., KARAMON, K., STAPINSKA-SYNIEC, A., COMPARATIVE ANALYSIS OF HEMORRHAGIC COMPLICATIONS IN DEEP BRAIN STIMULATION: MICROELECTRODE RECORDING VERSUS MACROSTIMULATIONE-DOES THE TECHNIQUE MATTER? WORLD NEUROSURGERY, 2025. 194(123634) HTTPS://DOI.ORG/10.1016/ J.WNEU.2024.123634 - BACKGROUND: HEMORRHAGE IS A DANGEROUS COMPLICATION OF DEEP BRAIN STIMULATION (DBS) SURGERY, BUT A COMPREHENSIVE EXPLANATION OF THE ASSOCIATED RISK FACTORS REMAINS INCONCLUSIVE, PARTICULARLY APPLICATION OF MICRO ELECTRODE RECORDING (MER) COMPARED WITH MACRO STIMULATION (NON-MER)-BASED DBS PROCEDURES. WE CONDUCTED A COMPARATIVE ANALYSIS OF MER AND MACRO STIMULATION TECHNIQUES, EVALUATING THE IMPACT OF BRAIN PENETRATIONS BY MICROELECTRODE GUIDING CANNULAS ON THE OCCURRENCE OF INTRACRANIAL HEMORRHAGIC EVENTS. - METHODS: THIS RETROSPECTIVE STUDY INCLUDED ALL DBS PROCEDURES (MER-BASED AND NON-MEREBASED) PERFORMED AT A SINGLE MEDICAL CENTER FROM NOVEMBER 2008 TO JUNE 2023. PATIENTS UNDERWENT COMPREHENSIVE PREOPERATIVE AND POSTOPERATIVE EVALUATIONS, INCLUDING POSTOPERATIVE COMPUTED TOMOGRAPHY TO ASSESS LEAD PLACEMENT AND INTRACRANIAL BLEEDING, CATEGORIZED AS SYMPTOMATIC OR ASYMPTOMATIC. STATISTICAL ANALYSES WERE CONDUCTED TO ASCERTAIN THE CORRELATION BETWEEN DIFFERENT SURGICAL TECHNIQUES AND THE RISK OF HEMORRHAGIC COMPLICATIONS. - RESULTS: IN A COHORT OF 618 PATIENTS, 1096 ELECTRODES WERE IMPLANTED DURING 672 CONSECUTIVE DBS PROCEDURES. THERE WERE 21 ASYMPTOMATIC AND 6 SYMPTOMATIC INTRACRANIAL HEMORRHAGES (ICHS), INCLUDING 1 DEATH. THE BLEEDING RATE WAS 4.37% PER PATIENT AND 2.92% PER ELECTRODE. PATIENTS WHO DEVELOPED HEMATOMAS WERE MORE PREVALENT IN THE MER GROUP (25 CASES) COMPARED WITH THE NON-MER GROUP (2 CASES). THE MER GROUP HAD A HIGHER MEAN NUMBER OF BRAIN PENETRATIONS BY GUIDING CANNULAS. - CONCLUSIONS: DBS PROCEDURES ARE GENERALLY SAFE, WITH ONLY 0.97% RESULTING IN LONG-TERM OR PERMANENT DEFICITS. ALL SYMPTOMATIC ICH OCCURRED IN MER-BASED PROCEDURES, WHEREAS MACROSTIMULATION WAS ASSOCIATED WITH ONLY 2 ASYMPTOMATIC ICH CASES. THE HIGHER NUMBER OF BRAIN PENETRATIONS IN MER CORRELATES WITH HIGHER BLEEDING RATES. REPORTABLE EVENTS: 1 PATIENT DIED; INTRAVENTRICULAR BLEEDING INTO THE THIRD, FOURTH, AND LATERAL VENTRICLES, INTRAPARENCHYMAL BLEEDING IN THE LEFT FRONTAL LOBE, BLEEDING TO CEREBELLOMEDULLARY CISTERN. Manufacturer Narrative: B3) EVENT DATE IS THE ACCEPTED DATE OF THE LITERATURE ARTICLE. D4) DEVICE LOT NUMBER, OR SERIAL NUMBER, UNAVAILABLE. G4) 510(K) IS DEPENDENT UPON THE DEVICE MODEL NUMBER AND THEREFORE, UNAVAILABLE. H3,H6) NO PARTS HAVE BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H4) DEVICE MANUFACTURING DATE IS DEPENDENT ON LOT NUMBER/SERIAL NUMBER, THEREFORE, UNAVAILABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21374421&pc=DXY | 2182208-2025-00724 | 2015-01-24 05:00 | Death | MEDTRONIC, INC. | 2025-02-13 05:00 | DXY | MEDTRONIC ICD | Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s) |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. A REPRESENTATIVE FOR THE PATIENT INDICATED THAT THE PATIENT'S HEART WAS "SO BAD AND WEAK" AND THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) MALFUNCTIONED AND KEPT SHOCKING THE PATIENT. THE PATIENT'S SPOUSE TRIED TO GET THE ICD TO STOP SHOCKING THE PATIENT BUT WAS UNABLE TO DO SO AND THE PATIENT PASSED AWAY. NO FURTHER DETAILS ARE AVAILABLE. Manufacturer Narrative: B2, B3: THE DATE OF DEATH AND EVENT DATE ARE ESTIMATES BASED ON INCOMING INFORMATION THAT NOTED THE EVENT OCCURRED "ABOUT TEN YEARS AGO". MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21374448&pc=NPT | 2015691-2025-00981 | 2025-01-21 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-13 05:00 | NPT | COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning | Cardiac Arrest; Tachycardia | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS AFFILIATES IN FRANCE, THIS CASE INVOLVED THE IMPLANTATION OF A 26 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION VIA A SUBCLAVIAN APPROACH. THE PATIENT HAD EXPERIENCED A CARDIAC ARREST THREE DAYS PRIOR TO THE PROCEDURE, WITH AN EJECTION FRACTION OF 40%, AND WAS UNSTABLE BEFORE THE PROCEDURE. DURING THE PROCEDURE, WITH THE SAFARI GUIDEWIRE CORRECTLY POSITIONED IN THE LEFT VENTRICLE, THE PATIENT EXPERIENCED HEMODYNAMIC FAILURE JUST BEFORE CROSSING THE NATIVE VALVE, LEADING TO CARDIAC ARREST WITH VENTRICULAR TACHYCARDIA. DUE TO SEVERE CALCIFICATION OF THE NATIVE VALVE, IT WAS IMPOSSIBLE TO IMPLANT THE VALVE, WHICH BECAME DEFORMED DUE TO THE MANEUVERS ATTEMPTED TO CROSS IT. THE DECISION WAS MADE TO RELEASE THE VALVE IN THE ASCENDING AORTA AND REMOVE THE COMMANDER DELIVERY SYSTEM. A PRE-DILATION WITH A 20 MM BALLOON WAS PERFORMED BEFORE ATTEMPTING TO USE A SECOND VALVE. UNFORTUNATELY, THE PATIENT REMAINED IN CARDIAC ARREST WITH ELECTROMECHANICAL DISSOCIATION AND PASSED AWAY BEFORE THE SECOND VALVE COULD BE INSERTED. Manufacturer Narrative: CORRECTION TO H6; TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSION. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE REPORTED EVENT OF CROSSING DIFFICULTY WAS UNABLE TO BE CONFIRMED DUE TO THE UNAVAILABILITY OF IMAGERY OR MEDICAL RECORDS. IN THIS CASE, AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS REPORTED INFORMATION INDICATED A "SEVERE CALCIFICATION OF THE NATIVE VALVE". PATIENT FACTORS SUCH AS CALCIFICATION MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERE WITH THE PASSAGE OF THE DELIVERY SYSTEM AND CAUSE DIFFICULTY DURING TRACKING/CROSSING. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS'S DEVICE-RELATED BLEED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PATIENT FACTORS (THE PATIENT WENT INTO CARDIAC ARREST THREE DAYS BEFORE THE PROCEDURE) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21374581&pc=DSQ | 2916596-2025-00657 | 2016-11-01 04:00 | Death | THORATEC CORPORATION | 2025-02-13 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿THE RESULTS OF MINIMALLY INVASIVE IMPLANTATION OF THE HEARTMATE 3 LVAD LEFT VENTRICULAR ASSIST DEVICE¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH INCREASED INTRAOPERATIVE USE OF FRESH FROZEN PLASMA (FFP) (P = .02) AND PLATELET CONCENTRATES (PCS), POSTOPERATIVE BLEEDING, POSTOPERATIVE WOUND INFECTIONS, OTHER INFECTIOUS COMPLICATIONS, POSTOPERATIVE RIGHT VENTRICULAR FAILURE, AND DEATH. THIS RETROSPECTIVE STUDY EVALUATED 47 PATIENTS WHO UNDERWENT THE HEARTMATE 3 LVAD IMPLANTATION BY STERNOTOMY AND 26 PATIENTS WHO HAD HEARTMATE 3 IMPLANTED VIA MINIMALLY INVASIVE METHOD. THE OBSERVATION LASTED FROM NOV2016 TO MAY2020. SIGNIFICANTLY HIGHER INTRAOPERATIVE USE OF FFP (P = .02) AND PCS (P = .01) WAS OBSERVED AMONG PATIENTS TREATED WITH THE MINIMALLY INVASIVE METHOD. THERE WAS NO SIGNIFICANT DIFFERENCE IN THE NUMBER OF POSTOPERATIVE BLEEDINGS REQUIRING SURGICAL INTERVENTION BETWEEN THE STUDY GROUPS. HOWEVER, THERE WAS A TENDENCY FOR MORE FREQUENT POSTOPERATIVE BLEEDINGS THAT REQUIRED TRANSFUSIONS BUT NOT SURGERY AMONG PATIENTS TREATED WITH MINIMALLY INVASIVE ACCESS. 6% OF PATIENTS WHO UNDERWENT STERNOTOMY EXPERIENCED POSTOPERATIVE WOUND INFECTIONS COMPARED TO 0% IN THE MINIMALLY INVASIVE APPROACH GROUP. OTHER INFECTIOUS COMPLICATIONS DID NOT DIFFER SIGNIFICANTLY BETWEEN THE GROUPS. 10 PATIENTS (21.28%) WHO UNDERWENT STERNOTOMY EXPERIENCED POSTOPERATIVE RIGHT VENTRICULAR FAILURE AND 4 PATIENTS (15.38%) WHO UNDERWENT THE MINIMALLY INVASIVE APPROACH EXPERIENCED POSTOPERATIVE RIGHT VENTRICULAR FAILURE. THERE WAS NO SIGNIFICANT DIFFERENCE IN SURVIVAL BETWEEN THE GROUPS. THE MEAN SURVIVAL TIME AMONG PATIENTS UNDERGOING STERNOTOMY WAS 430.0 DAYS (+/- 291.77 DAYS). PATIENTS UNDERGOING MINIMALLY INVASIVE SURGERY HAD A MEAN SURVIVAL TIME OF 633.15 DAYS (+/- 426.84). THERE WAS NO SIGNIFICANT DIFFERENCE IN ONE 1 YEAR SURVIVAL BETWEEN GROUPS (84% VS 84%, P = .24). Manufacturer Narrative: SECTION A: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN NOVEMBER 2016 TO MAY 2020. SILESIAN CENTER FOR HEART DISEASES, ZABRZE, POLAND; AND DEPARTMENT OF CARDIAC, VASCULAR, AND ENDOVASCULAR SURGERY AND TRANSPLANTOLOGY, MEDICAL UNIVERSITY OF SILESIA, KATOWICE, POLAND. ANTONCZYK, R., BIELKA, A., KALINOWSKI, M., SMIGIEL-IRAS, M., & KUCZAJ, A. (2024). THE RESULTS OF MINIMALLY INVASIVE IMPLANTATION OF THE HEARTMATE 3 LVAD LEFT VENTRICULAR ASSIST DEVICE. TRANSPLANTATION PROCEEDINGS, 56(4), 854¿859. HTTPS://DOI.ORG/10.1016/J.TRANSPROCEED.2024.03.024. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 88197). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21374640&pc=MKJ | 1220908-2025-00479 | 2025-01-23 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-13 05:00 | MKJ | X SERIES | Incorrect Interpretation of Signal | Cardiac Arrest; Tachycardia | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO CARDIOVERT A FEMALE PATIENT (AGE UNKNOWN), THE DEVICE GAVE A "NO SHOCK ADVISED" PROMPT FOR A RHYTHM CLINICIANS BELIEVED WAS SHOCKABLE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION WHERE THE CLINICAL FILE ANALYSIS WAS REVIEWED FOR EACH OUTCOME. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. REVIEW OF THE CLINICAL DATA LOGS SHOWED ANALYSIS EVENTS 1 AND 3, THE ALGORITHM DETERMINED THE PRESENTING WAVEFORMS AS NON-SHOCKABLE FORMS OF VENTRICULAR TACHYCARDIA, MOSTLY DUE TO LOW RATE AND WIDE COMPLEXES. ANALYSIS EVENT 2, 21 AND 23 WERE RAPID SHOCK EVENTS THAT WERE DETERMINED TO BE NON-SHOCKABLE. ALSO PRESENTING VENTRICULAR TACHYCARDIA WITH LOW RATE AND WIDE COMPLEXES. ANALYSIS EVENTS 4 AND 19 WERE RAPID SHOCK EVENTS WHICH DID RESULT IN SHOCKABLE DECISIONS. HOWEVER, THE SHOCK CONVERSION ESTIMATOR (SCE) FEATURE WAS ENABLED, WHICH OVERRODE THE DECISION TO SHOCK IN EACH ANALYSIS EVENT. THE SCE FUNCTION LOOKS AT THE WAVEFORM SPECTRALLY AND DETERMINES ITS PROBABILITY TO BE CONVERTED TO AN ORGANIZED RHYTHM. THE SCE WILL OVERRIDE THE SHOCK ADVISEMENT AND RESULT IN A NO SHOCK ADVISED EVEN IF THE MAJORITY OR ALL SEGMENTS OF THE ANALYSIS POINT TO SHOCKABLE. IN BOTH ANALYSES, THE ALGORITHM ALTERED ITS SHOCK DECISION BASED ON SCE AND REVERSED ITS SHOCK DECISION. BASED ON OUR REVIEW OF THE DATA WE HAVE CONCLUDED THAT THE ANALYSIS PROGRAM WORKED AS DESIGNED WITHIN ITS LIMITATIONS AND THERE WAS NO INDICATION OF A MALFUNCTION WITH THE DEVICE. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375049&pc=KDI | 3003464075-2025-00041 | 2025-01-19 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-02-13 05:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K170469 |  | 1 | Event Description: A REPORT WAS RECEIVED ON 21 JAN 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 58-YEAR-OLD FEMALE PATIENT APPROVED FOR PERFORMING SOLO HEMODIALYSIS, WITH A MEDICAL HISTORY INCLUDING DIABETES, CARDIAC COMORBIDITIES AND END STAGE RENAL DISEASE WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 28 JAN 2025 FROM THE HOME THERAPY NURSE (HTN) WHO STATED THAT THE PATIENT WAS FOUND EXPIRED WITH RINSEBACK COMPLETED AND STILL CONNECTED TO THE DEVICE. PER THE HTN, THE CAUSE OF DEATH IS LISTED AS 'ESRD IN DIABETES' AND WAS UNRELATED TO NXSTAGE PRODUCT OR THERAPY. Manufacturer Narrative: THERE WAS NO DEVICE MALFUNCTION. THE CYCLER WAS RECEIVED FOR EVALUATION AND SUCCESSFULLY PASSED TESTING. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. AVAILABLE LOG FILES WERE RETRIEVED AND ANALYZED WHICH SHOWED DEVICE PERFORMANCE WAS WITHOUT DEFICIENCY AND WAS UNREMARKABLE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375073&pc=NPT | 2135147-2025-00705 | 2025-01-17 05:00 | Death | ABBOTT MEDICAL | 2025-02-13 05:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE ANATOMICAL SIZING OF THE PATIENT IS: DIAMETER OF VALSALVA L:31.5MM, R:30.1MM, N:32.0MM, HEIGHT OF VALSALVA L:17.5MM, R:20.0MM, EFFECTIVE ORIFICE AREA 425.3 MM2, PERIMETER OF ANNULUS 74.7MM, ASCENDING AORTA DIAMETER 29.9MM, AND LENGTH OF MEMBRANOUS SEPTUM 2.0MM. THERE WAS NO CALCIFICATION CONFIRMED FROM THE ANNULUS TO THE SUB VALVULAR TO LVOT. THE VALVE WAS ABLE TO BE SUCCESSFULLY IMPLANTED AT A DEPTH OF N3MM, L5MM. ON (B)(6) 2025, THE PATIENT WAS DISCHARGED WITH BAYASPIRIN. ON AN (B)(6) 2025, THE PATIENT SUDDENLY DETERIORATED AND PASSED AWAY DUE TO CARDIAC ARREST. CT WAS PERFORMED AND SHOWED NO ABNORMAL FINDINGS AROUND THE HEART. THE PHYSICIAN STATED THAT THE CAUSE OF DEATH IS UNKNOWN, BUT THERE IS A POSSIBILITY THAT IT WAS DUE TO HYPOTHERMIA.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE ANATOMICAL SIZING OF THE PATIENT IS: DIAMETER OF VALSALVA L:31.5MM, R:30.1MM, N:32.0MM, HEIGHT OF VALSALVA L:17.5MM, R:20.0MM, EFFECTIVE ORIFICE AREA 425.3 MM2, PERIMETER OF ANNULUS 74.7MM, ASCENDING AORTA DIAMETER 29.9MM, AND LENGTH OF MEMBRANOUS SEPTUM 2.0MM. THERE WAS NO CALCIFICATION CONFIRMED FROM THE ANNULUS TO THE SUB VALVULAR TO LVOT. THE VALVE WAS ABLE TO BE SUCCESSFULLY IMPLANTED AT A DEPTH OF N3MM,L5MM. ON (B)(6) 2025, THE PATIENT WAS DISCHARGED WITH BAYASPIRIN. ON AN (B)(6) 2025, THE PATIENT SUDDENLY DETERIORATED AND PASSED AWAY DUE TO CARDIAC ARREST. CT WAS PERFORMED AND SHOWED NO ABNORMAL FINDINGS AROUND THE HEART. THE PHYSICIAN STATED THAT THE CAUSE OF DEATH IS UNKNOWN, BUT THERE IS A POSSIBILITY THAT IT WAS DUE TO HYPOTHERMIA. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF DEATH AND CARDIAC ARREST WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS.
THE EVENT WAS REVIEWED BY AN ABBOTT DIRECTOR OF MEDICAL AFFAIRS. BASED ON ALL AVAILABLE INFORMATION, CAUSES FOR THE REPORTED DEATH AND CARDIAC ARREST COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375107&pc=HFA | MW5166109 | 2023-06-05 04:00 | Death | KARL STORZ SE & CO., KG. | 2025-02-13 05:00 | HFA | KARL STORZ FETOSCOPY SET |  | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Abdominal Cramps |  |  | 1 | Event Description: A 32 Y.O. G3P0020 FEMALE AT 23W6D GESTATION (EDD(ESTIMATED DATE OF DELIVERY) IS (B)(6) 2023) WHO WAS REFERRED FOR CONCERNS OF TWIN TO TWIN TRANSFUSION SYNDROME (TTTS). HER US(OBSTRETIC ULTRASOUND) FINDINGS SHOWED TWIN A (DONOR) HAD A VISIBLE BLADDER BUT HAD PREDOMINANTLY ABSENT END DIASTOLIC FLOW IN THE UMBILICAL ARTERY (ALTHOUGH SOME CYCLES DEMONSTRATE PRESERVED END DIASTOLIC FLOW), WITH OCCASIONAL REVERSED END DIASTOLIC FLOW. THE DV(DUCTUS VENOSUS) AND MCA (MIDDLE CEREBRAL ARTERY) WERE OTHERWISE NORMAL. ALL DOPPLERS IN TWIN B(RECIPIENT) WAS ALSO NORMAL AND BLADDER WAS VISIBLE AS WELL. THEREFORE, THESE FINDINGS WERE CONSISTENT WITH TTTS STAGE III. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON (B)(6) 2023 WHEN SHE WAS 24 WEEKS PREGNANT. SHE HAD A POSTERIOR PLACENTA WITH VASCULAR ANASTOMOSIS BETWEEN RECIPIENT AND DONOR CORD INSERTIONS. A TOTAL OF 22 VASCULAR COMMUNICATIONS WERE IDENTIFIED INCLUDING: 4 SMALL D-R, 5 SMALL R-D, 6 MEDIUM D-R, 1 MEDIUM R-D, 3 LARGE R-D, AND 3 EXTRA-LARGE D-R CONNECTION. SHE ALSO UNDERWENT AN AMNIOREDUCTION. HER CERVIX WAS SHORT SO THERE WAS ALSO A CERCLAGE PLACED ON THE SAME OPERATIVE DAY. POST-OPERATIVELY THERE WERE NO COMPLICATIONS. ON POD#5, HER US SHOWED TWO VIABLE FETUSES WITH NORMALIZING AMNIOTIC FLUID AND ALSO NO CRITICALLY ABNORMAL WAVEFORMS WERE IDENTIFIED. HOWEVER, ON REPEAT US ON POD#10(POST-OPERATIVE DAY #5), SHE WAS FOUND TO HAVE A DONOR DEMISE. WHEN SHE WAS 29W1D, SHE PRESENTED WITH VAGINAL BLEEDING AND CONTRACTIONS. FORMER RECIPIENT WAS DELIVERED AND AS OF NOW IS DOING WELL AND IS 1.5 YEARS OF AGE. REFERENCE REPORT: MW5166110.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375117&pc=HFA | MW5166110 | 2023-06-05 04:00 | Death | KARL STORZ SE & CO., KG. | 2025-02-13 05:00 | HFA | KARL STORZ FETOSCOPY SET |  | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Abdominal Cramps |  |  | 1 | Event Description: A 32 Y.O. G3P0020 FEMALE AT 23W6D GESTATION (EDD (ESTIMATED DATE OF DELIVERY) IS (B)(6) 2023) WHO WAS REFERRED FOR CONCERNS OF TWIN TO TWIN TRANSFUSION SYNDROME (TTTS). HER US(OBSTRETIC ULTRASOUND) FINDINGS SHOWED TWIN A (DONOR) HAD A VISIBLE BLADDER BUT HAD PREDOMINANTLY ABSENT END DIASTOLIC FLOW IN THE UMBILICAL ARTERY (ALTHOUGH SOME CYCLES DEMONSTRATE PRESERVED END DIASTOLIC FLOW), WITH OCCASIONAL REVERSED END DIASTOLIC FLOW. THE DV(DUCTUS VENOSUS) AND MCA(MIDDLE CEREBRAL ARTERY) WERE OTHERWISE NORMAL. ALL DOPPLERS IN TWIN B(RECIPIENT) WAS ALSO NORMAL AND BLADDER WAS VISIBLE AS WELL. THEREFORE, THESE FINDINGS WERE CONSISTENT WITH TTTS STAGE III. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON (B)(6) 2023 WHEN SHE WAS 24 WEEKS PREGNANT. SHE HAD A POSTERIOR PLACENTA WITH VASCULAR ANASTOMOSIS BETWEEN RECIPIENT AND DONOR CORD INSERTIONS. A TOTAL OF 22 VASCULAR COMMUNICATIONS WERE IDENTIFIED INCLUDING: 4 SMALL D-R, 5 SMALL R-D, 6 MEDIUM D-R, 1 MEDIUM R-D, 3 LARGE R-D, AND 3 EXTRA-LARGE D-R CONNECTION. SHE ALSO UNDERWENT AN AMNIOREDUCTION. HER CERVIX WAS SHORT SO THERE WAS ALSO A CERCLAGE PLACED ON THE SAME OPERATIVE DAY. POST-OPERATIVELY THERE WERE NO COMPLICATIONS. ON POD#5, HER US SHOWED TWO VIABLE FETUSES WITH NORMALIZING AMNIOTIC FLUID AND ALSO NO CRITICALLY ABNORMAL WAVEFORMS WERE IDENTIFIED. HOWEVER, ON REPEAT US ON POD#10(POST-OPERATIVE DAY #5), SHE WAS FOUND TO HAVE A DONOR DEMISE. WHEN SHE WAS 29W1D, SHE PRESENTED WITH VAGINAL BLEEDING AND CONTRACTIONS. FORMER RECIPIENT WAS DELIVERED AND AS OF NOW IS DOING WELL AND IS 1.5 YEARS OF AGE. REFERENCE REPORT: MW5166109.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375284&pc=CAW | 2515872-2025-00005 | 2022-04-21 04:00 | Death | DEVILBISS HEALTHCARE | 2025-02-13 05:00 | CAW | DEVILBISS HEALTHCARE | Infusion or Flow Problem | Insufficient Information | K071397 |  | 1 | Event Description: DEVILBISS HEALTHCARE WAS RECENTLY NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY HEALTH CANADA, WHICH STATED "OXYGEN COMPRESSOR (USED AT PATIENT'S HOME) HAD LOW FLOW DUE TO BACTERIAL OVERGROWTH. THE USER APPARENTLY CONTACTED ONTARIO MEDICAL SUPPLY AS THE WARNING LIGHT WAS FLASHING "YELLOW". AT THE TIME OF DEATH ((B)(6) 2022) THERE HAD BEEN NO RESPONSE, APPARENTLY, FROM ONTARIO MEDICAL SUPPLY. ALSO, NO MAINTENANCE RECORDS WERE AVAILABLE FROM ONTARIO MEDICAL SUPPLY DURING THE INVESTIGATION (OMS REPRESENTATIVE STATED THEY HAD BEEN DESTROYED IN A HACKING INCIDENT)." THE DEVICE IS EQUIPPED WITH A YELLOW INDICATOR LIGHT WITH INTERMITTENT AUDIBLE BEEPING THAT ACTIVATES WHEN THE OXYGEN CONCENTRATION BEGINS TO FALL BELOW SPECIFICATION. THIS LIGHT IS AN INDICATOR THAT THE EQUIPMENT MAY BE IN NEED OF ROUTINE MAINTENANCE. IT IS UNCLEAR HOW BACTERIAL OVERGROWTH WAS DETERMINED OR HOW SUCH A CONDITION WOULD CONTRIBUTE TO THE STATED LOW FLOW. THE DEVICE CAN BE USED WITH A HUMIDIFIER ACCESSORY TO PROVIDE MOISTURE TO THE OXYGEN. IT IS POSSIBLE THAT THE "BACTERIAL OVERGROWTH" WAS IN THE HUMIDIFIER ACCESSORY AS THAT IS VISIBLE TO USER. THE DEVICE IS NOT A LIFE-SUPPORTING OR LIFE-SUSTAINING DEVICE. THE DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER WAS NOTIFIED OF THE INCIDENT WITH NO RESPONSE. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. DEVILBISS HEALTHCARE WILL FILE AN UPDATE IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375366&pc=OZD | 1220648-2025-26127 | 2024-10-01 04:00 | Death | ABIOMED, INC. | 2025-02-13 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT MULTIPLE COMPLICATIONS WERE ENCOUNTERED DURING IMPELLA CP SUPPORT. IT WAS DOCUMENTED THAT THE PATIENT ENDURED VENTRICULAR ARRHYTHMIA WITH A PROBABLE RELATIONSHIP TO THE IMPELLA DEVICE. THE PATIENT WAS BOLUSED WITH 50MG OF AMIODARONE AND CARDIOVERSION WAS PERFORMED WITH NOTED IMPROVEMENT OF THEIR HEARTRATE. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT EXPERIENCED BLEEDING WITH A PROBABLE RELATIONSHIP TO THE IMPELLA DEVICE. THE PATIENT WAS GIVEN 2 UNITS OF PACKED RED BLOOD CELLS TREAT THE DROPPING HEMOGLOBIN LEVELS. DUE TO THE PATIENT'S CONDITION, THE FAMILY DECIDED TO MOVE THE PATIENT TO COMFORT CARE WHERE THEY EXPIRED. IT IS UNKNOWN IF THE CONDITIONS CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION FOR THE VASCULAR DAMAGE AND ARRHYTHMIA HAS BEEN COMPLETED. THE IMPELLA WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING COULD NOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED, THE ROOT CAUSE OF THE ARRHYTHMIA WAS NOT DETERMINED B.2 ADDED REQUIRED INTERVENTION AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26127. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE (IFUS) WERE REPORTED ON THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26127 INCORRECTLY. NO IFUS ARE NEEDED TO BE REPORTED FOR THIS REPORT IN ACCORDANCE WITH UPDATED PROCEDURES. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375493&pc=JCW | 2124215-2025-06615 | 2025-01-15 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-13 05:00 | JCW | AMS 700 LGX TENACIO PUMP | Insufficient Information | Insufficient Information | N970012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS INFLATABLE PENILE PROSTHESIS (IPP) WAS DISCHARGED TO HOME AND EXPIRED IN THEIR SLEEP THAT EVENING. THERE WERE NO REPORTED PROCEDURAL COMPLICATIONS OR DEVICE ISSUES. THE CAUSE OF THE DEATH IS UNKNOWN.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS INFLATABLE PENILE PROSTHESIS (IPP) WAS DISCHARGED TO HOME AND EXPIRED IN THEIR SLEEP THAT EVENING. THERE WERE NO REPORTED PROCEDURAL COMPLICATIONS OR DEVICE ISSUES. THE CAUSE OF THE DEATH IS UNKNOWN. DESPITE EFFORTS, NO FURTHER INFORMATION COULD BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375525&pc=HAW | 1723170-2025-00661 | 2024-11-22 05:00 | Death | MEDTRONIC NAVIGATION, INC. | 2025-02-13 05:00 | HAW | STEALTHSTATION¿ S8 SYSTEM | Incorrect, Inadequate or Imprecise Result or Readings | No Clinical Signs, Symptoms or Conditions | K162309 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A NAVIGATION SYSTEM BEING USED DURING A SACROILIAC AND THORACOLUMBAR PROCEDURE. IT WAS REPORTED THAT THERE WAS AN ALLEGED NAVIGATION INACCURACY THAT MAY BE LINKED TO A BLINKING LIGHT-EMITTING DIODE (LED) ON THE IMAGING SYSTEM DURING ACQUISITION. THERE WAS LESS THAN ONE HOUR DELAY TO THE CASE. ADDITIONAL INFORMATION WAS RECEIVED. IT WAS REPORTED THAT THE SITE REPORTED A MINOR INACCURACY. THERE WAS NO REPORTED IMPACT ON PATIENT OUTCOME.Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE HEALTH CARE PROFESSIONAL NOTICED A LACK OF PRECISION IN THEIR NAVIGATION AND THOUGHT IT WOULD BE RELATED TO THE BLINKING LEDS DURING THE SCAN ACQUISITION. BUT THEY STATED ALSO THAT THE REFERENCE FRAME COULD HAVE MOVED BY MISTAKE. Manufacturer Narrative: H2-3) THE SOFTWARE ANALYSIS CONCLUDED THAT THERE WAS INSUFFICIENT INFORMATION TO DETERMINE IF A SOFTWARE ANOMALY CAUSED THE REPORTED EVENT. THE SOFTWARE LOGS WERE NOT HELPFUL IN DIAGNOSING AUTO-REGISTRATION ACCURACY ISSUES. FRAME MOVEMENT AFTER REGISTRATION WAS A COMMON CAUSE OF ACCURACY ISSUES BUT COULDN'T BE DETECTED IN LOGFILES OR ARCHIVES. CODES: B01, C19, D15 H2) PLEASE SEE SECTION D9 FOR WHEN THE SOFTWARE LOGS WERE AVAILABLE FOR EVALUATION. PLEASE SEE SECTION D3-4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H2, CORRECTION: D3-4 UPDATED. H1 IN THE INITIAL REGULATORY REPORT INCORRECTLY INDICATED A PATIENT DEATH, BUT THERE WAS NO PATIENT THAT ACTUALLY DIED. AS PREVIOUSLY REPORTED, THERE WAS NO REPORTED IMPACT ON PATIENT OUTCOME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: 9735740, VERSION: 2.1.0, UBD: UNKNOWN, UDI#: UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H2: ADDITIONAL INFORMATION WAS SUBMITTED IN 3006544299-2024-01068 THAT WAS ADDED TO B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H3,H6) NO PARTS HAVE BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. FOR THE IMAGING SYSTEM, THE COMPLAINT IS REFERENCED IN REG REPORT: 3006544299-2024-01068. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375528&pc=HFA | MW5166112 | 2024-01-14 05:00 | Death | KARL STORZ SE & CO. / KARL STORZ ENDOSCOPY-AMERICA, INC. | 2025-02-13 05:00 | HFA | KARL STORZ MINI STRAIGHT-FORWARD TELESCOPE 0°,2 MM |  | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Premature Labor; Prolapse; Bowel Perforation |  |  | 1 | Event Description: (B)(6) IS A 31 YO G5P2002 AT 18W5D (EDD AT (B)(6) 2024) WITH MONOCHORIONIC DIAMNIOTIC TWINS WHO WAS REFERRED FROM OSH FOR TWIN TO TWIN TRANSFUSION SYNDROME. SHE HAD TTTS STAGE III WITH FINDINGS OF TWIN A THE RECIPIENT HAVING REVERSED A-WAVE DV DIAGNOSED AT 17W3 BUT STILL HAD SIGNIFICANT MEMBRANE SEPARATION. TWIN B WAS NOTED TO HAVE INTERMITTENT ABSENT END-DIASTOLIC FLOW (EDF) OF THE UMBILICAL ARTERY (UA). THE CHORIOAMNIOTIC SEPARATION IMPROVED BY 18W5D. SHE UNDERWENT LASER PHOTOCOAGULATION AND AMNIOREDUCTION ON (B)(6) 2023. SHE HAD AN ANTERIOR PLACENTA AND WAS FOUND TO HAVE 17 VASCULAR ANASTOMOSES INCLUDING: 6 SMALL HAIRS,2 SMALL R-D, 2 MEDIUM R-D, 7 MEDIUM D-R. PRE-OPERATIVE FETAL HEART RATES WAS 142 BPM FOR RECIPIENT (TWIN A) AND 164 BPM FOR DONOR (TWINB). PRE-OPERATIVE MITRAL VALVE PROLAPSE (MVP) WAS 14.2CM. POST-OP MVP WAS 14.3CM FOR THE RECIPIENT. POST-OPERATIVE FETAL HEART RATES 142 FOR THE RECIPIENT (TWIN A) AND 148 BPM FOR DONOR (TWINB). PRE-OPERATIVE MIDDLE CEREBRAL ARTERY (MCA) FOR RECIPIENT TWIN A WAS 19.9CM/S AND FOR DONOR TWIN B WAS 14.0 CM/S AND POST-OPERATIVE MCA FOR RECIPIENT TWIN A WAS 19.5CM/S AND DONOR TWIN B WAS 14.0 CM/S. POST-OPERATIVELY, WHEN SHE WAS 22W4D SHE WAS FOUND TO HAVE CERVICAL SHORTENING AND UNDERWENT A RESCUE MCDONALD CERVICAL CERCLAGE. AT 24W ((B)(6) 2024) SHE SPONTANEOUSLY LABORED AND DELIVERED BOTH TWINS. TWIN B FORMER DONOR PASSED ON DOL#7 SECONDARY TO GRADE IV IVH BILATERALLY AND BOWEL PERFORATION. TWIN A FORMER RECIPIENT AND AT THE 1 YEAR FOLLOW UP IS DOING WELL. REFERENCE REPORT: MW5166113.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375544&pc=HFA | MW5166113 | 2024-01-14 05:00 | Death | KARL STORZ SE & CO. / KARL STORZ ENDOSCOPY-AMERICA, INC. | 2025-02-13 05:00 | HFA | KARL STORZ OPERATING SHEATH, 9 FR. |  | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Premature Labor; Prolapse; Bowel Perforation |  |  | 1 | Event Description: (B)(6) IS A 31 YO G5P2002 AT 18W5D (EDD AT (B)(6) 2024) WITH MONOCHORIONIC DIAMNIOTIC TWINS WHO WAS REFERRED FROM OSH FOR TWIN-TO-TWIN TRANSFUSION SYNDROME. SHE HAD TTTS STAGE III WITH FINDINGS OF TWIN A THE RECIPIENT HAVING REVERSED A-WAVE DV DIAGNOSED AT 17W3 BUT STILL HAD SIGNIFICANT MEMBRANE SEPARATION. TWIN B WAS NOTED TO HAVE INTERMITTENT ABSENT END-DIASTOLIC FLOW (EDF) OF THE UMBILICAL ARTERY (UA). THE CHORIOAMNIOTIC SEPARATION IMPROVED BY 18W5D. SHE UNDERWENT LASER PHOTOCOAGULATION AND AMNIOREDUCTION ON (B)(6) 2023. SHE HAD AN ANTERIOR PLACENTA AND WAS FOUND TO HAVE 17 VASCULAR ANASTOMOSES INCLUDING: 6 SMALL HAIRS,2 SMALL R-D, 2 MEDIUM R-D, 7 MEDIUM D-R. PRE-OPERATIVE FETAL HEART RATES WAS 142 BPM FOR RECIPIENT (TWIN A) AND 164 BPM FOR DONOR (TWINB). PRE-OPERATIVE MITRAL VALVE PROLAPSE (MVP) WAS 14.2CM. POST-OP MVP WAS 14.3CM FOR THE RECIPIENT. POST-OPERATIVE FETAL HEART RATES 142 FOR THE RECIPIENT (TWIN A) AND 148 BPM FOR DONOR (TWINB). PRE-OPERATIVE MIDDLE CEREBRAL ARTERY (MCA) FOR RECIPIENT TWIN A WAS 19.9CM/S AND FOR DONOR TWIN B WAS 14.0 CM/S AND POST-OPERATIVE MCA FOR RECIPIENT TWIN A WAS 19.5CM/S AND DONOR TWIN B WAS 14.0 CM/S. POST-OPERATIVELY, WHEN SHE WAS 22W4D SHE WAS FOUND TO HAVE CERVICAL SHORTENING AND UNDERWENT A RESCUE MCDONALD CERVICAL CERCLAGE. AT 24W ((B)(6) 2024) SHE SPONTANEOUSLY LABORED AND DELIVERED BOTH TWINS. TWIN B FORMER DONOR PASSED ON DOL#7 SECONDARY TO GRADE IV IVH BILATERALLY AND BOWEL PERFORATION. TWIN A FORMER RECIPIENT AND AT THE 1 YEAR FOLLOW UP IS DOING WELL. REFERENCE REPORT: MW5166112.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375568&pc=LWS | 2649622-2025-04269 | 2025-01-25 05:00 | Death | MPRI | 2025-02-13 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Defibrillation/Stimulation Problem; Under-Sensing | Myocardial Infarction; Cardiogenic Shock | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING NOTED ON STORED OR CURRENT ELECTROGRAMS (EGM) RESULTING IN DELAYED DELIVERY OF THERAPY. ADDITIONALLY, THE LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH-RATE, NON-SUSTAINED EPISODESAND SHORT VENTRICULAR INTERVALS. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK AND ELEVATION MYOCARDIAL INFARCTION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375928&pc=HFA | MW5166114 | 2023-03-03 05:00 | Death | KARL STORZ SE & CO KG. | 2025-02-13 05:00 | HFA | KARL STORZ TELESCOPE |  | Failure to Anastomose |  |  | 1 | Event Description: A 29 Y/O G2P1001 AT 17W0D WHO WAS DIAGNOSED WITH STAGE III TTTS ("POLY-OLY" AND ABSENT BLADDER IN TWIN B(DONOR) WITH ABNORMAL DOPPLERS OF INTERMITTENT ABSENT OF THE END DIASTOLIC FLOW OF THE UA. MVP WAS 8.5CM AND 0.5CM RESPECTIVELY FOR THE RECIPIENT AND DONOR. ALL OTHER DOPPLERS WERE NORMAL. THE PATIENT AND HER FAMILY UNDERWENT EXTENSIVE COUNSELING, AND THEY ELECTED TO PROCEED WITH FETOSCOPIC-DIRECTED LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSES. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON (B)(6) 2023. THE CASE WAS UNEVENTFUL. SHE HAD A POSTERIOR PLACENTA AND HAD A TOTAL OF 14 VASCULAR CONNECTIONS IDENTIFIED: 4 SMALL D-R, 2 MEDIUM D-R, 6 MEDIUM R-D, 1 LARGE D-R, AND 1 EXTRALARGE VV CONNECTION. SHE UNDERWENT AN AMNIOREDUCTION AS WELL. PRE-OPERATIVE FETAL HEART RATES WERE 159 BPM FOR DONOR (TWIN B) AND 134 BPM FOR RECIPIENT (TWINA). PRE-OPERATIVE MVP WAS 8.5CM, POST-OP MVP WAS 7.2CM FOR THE RECIPIENT. POST-OPERATIVE FETAL HEART RATES WERE 153 BPM FOR THE DONOR AND 144 BPM FOR THE RECIPIENT. UNFORTUNATELY, ON POD #1 ((B)(6) 2023) SHE WAS FOUND TO HAVE A DEMISE OF TWIN A. SHE CONTINUED WITH PREGNANCY AND EVENTUALLY DELIVERED TWIN B AT 37W. AS OF NOW, THE BABY IS DOING WELL AT 1.5 YEARS. REFERENCE REPORT: MW5166115.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21375940&pc=HFA | MW5166115 | 2023-03-03 05:00 | Death | KARL STORZ SE & CO., KG. | 2025-02-13 05:00 | HFA | KARL STORZ SHEATH |  | Failure to Anastomose |  |  | 1 | Event Description: A 29 Y/O G2P1001 AT 17W0D WHO WAS DIAGNOSED WITH STAGE III TTTS ("POLY-OLY" AND ABSENT BLADDER IN TWIN B(DONOR) WITH ABNORMAL DOPPLERS OF INTERMITTENT ABSENT OF THE END DIASTOLIC FLOW OF THE UA. MVP WAS 8.5CM AND 0.5CM RESPECTIVELY FOR THE RECIPIENT AND DONOR. ALL OTHER DOPPLERS WERE NORMAL. THE PATIENT AND HER FAMILY UNDERWENT EXTENSIVE COUNSELING, AND THEY ELECTED TO PROCEED WITH FETOSCOPIC-DIRECTED LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSES. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON (B)(6) 2023. THE CASE WAS UNEVENTFUL. SHE HAD A POSTERIOR PLACENTA AND HAD A TOTAL OF 14 VASCULAR CONNECTIONS IDENTIFIED: 4 SMALL D-R, 2 MEDIUM D-R, 6 MEDIUM R-D, 1 LARGE D-R, AND 1 EXTRALARGE VV CONNECTION. SHE UNDERWENT AN AMNIOREDUCTION AS WELL. PRE-OPERATIVE FETAL HEART RATES WERE 159 BPM FOR DONOR (TWIN B) AND 134 BPM FOR RECIPIENT (TWINA). PRE-OPERATIVE MVP WAS 8.5CM, POST-OP MVP WAS 7.2CM FOR THE RECIPIENT. POST-OPERATIVE FETAL HEART RATES WERE 153 BPM FOR THE DONOR AND 144 BPM FOR THE RECIPIENT. UNFORTUNATELY, ON POD #1 ((B)(6) 2023) SHE WAS FOUND TO HAVE A DEMISE OF TWIN A. SHE CONTINUED WITH PREGNANCY AND EVENTUALLY DELIVERED TWIN B AT 37W. AS OF NOW, THE BABY IS DOING WELL AT 1.5 YEARS. REFERENCE REPORT: MW5166114.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376059&pc=OZD | 1220648-2025-26130 | 2025-01-14 05:00 | Death | ABIOMED, INC. | 2025-02-13 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT A PATIENT DEVELOPED HEMOLYSIS DURING IMPELLA CP SUPPORT. IT WAS INDICATED THAT THERE WAS A DEFINITE RELATIONSHIP BETWEEN THE CONDITION AND THE IMPELLA DEVICE USED. THE PATIENT'S PLASMA FREE HEMOGLOBIN LEVELS WERE ELEVATED AT THE TIME OF THE DIAGNOSIS. THE PATIENT WAS TRANSFUSED WITH ONE UNIT OF PACKED RED BLOOD CELLS TO TREAT THE CONDITION. DUE TO THE PATIENT'S OVERALL CONDITION, THE PATIENT'S FAMILY MADE THE DECISION TO MAKE THE PATIENT DO NOT RESUSCITATE AND THE PUMP WAS REMOVED. THE PATIENT SUBSEQUENTLY PASSED AWAY. IT CAN NOT BE DETERMINED IF THE HEMOLYSIS CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376098&pc=DSQ | 2916596-2025-01080 | 2014-11-01 04:00 | Death | THORATEC CORPORATION | 2025-02-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Ischemia; Pneumonia; Renal Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "SURGICAL INTERVENTIONS FOR ACCUMULATION OF BIO-DEBRIS CAUSING OUTFLOW GRAFT OBSTRUCTION AND THROMBOSIS FOLLOWING HEARTMATE 3 (HM3)" RETROSPECTIVELY REVIEWED 298 HM3 PATIENTS AND IDENTIFIED EIGHT, 2.7%, THAT WERE IMPLANTED BETWEEN NOVEMBER 2014 AND DECEMBER 2022 WHICH REQUIRED CORRECTIVE SURGERY. PATIENT 8 WAS IMPLANTED AT AGE 60.3 YEARS AS DESTINATION THERAPY AND PRESENTED WITH SYMPTOMATIC LOW FLOW ALARMS 1.1 YEARS FOLLOWING IMPLANT. COMPUTERIZED TOMOGRAPHY ANGIOGRAPHY (CTA) IDENTIFIED 100% OCCLUSION RESULTING IN CARDIOGENIC SHOCK AND NECESSITATING A PUMP EXCHANGE. PRE-INTERVENTION HM3 PARAMETERS WERE REPORTED AS; SPEED 5100 ROTATIONS PRE-MINUTE (RPM), FLOW N/A, PI N/A AND POWER N/A. POST INTERVENTION PARAMETERS WERE REPORTED AS: SPEED 5200 RPM, FLOW N/A, PI N/A AND POWER N/A. DURATION OF FOLLOW-UP WAS 1.1 YEARS POST HM3. POST INTERVENTION 9 DAYS WAS COMPLICATED BY ACUTE LIMB ISCHEMIA, PNEUMONIA AND ACUTE RENAL FAILURE REQUIRING DIALYSIS. THEIR LAST KNOWN STATUS WAS REPORTED AS DECEASED. Manufacturer Narrative: B2: DATE OF DEATH WAS NOT PROVIDED. SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS: NOV2014 AS PATIENTS WERE IMPLANTED BETWEEN NOV2014 AND DEC2022. VINOGRADSKY, A., HYNDS, M., KAKU, Y., LEB, J., YUZEFPOLSKAYA, M., COLOMBO, P., SAYER, G., URIEL, N., NAKA, Y., & TAKEDA, K. (2024B). SURGICAL INTERVENTIONS FOR ACCUMULATION OF BIODEBRIS CAUSING OUTFLOW GRAFT OBSTRUCTION AND THROMBOSIS FOLLOWING HEARTMATE 3. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S287. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1233 DIVISION OF CARDIAC, THORACIC & VASCULAR SURGERY, DEPARTMENT OF SURGERY, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY . COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY . DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY . THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL INFORMATION. THE INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MANUFACTURER REPORT #2916596-2025-01663.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376122&pc=LYJ | 1644487-2025-00154 | 2025-01-07 05:00 | Death | LIVANOVA USA, INC. | 2025-02-13 05:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY RECENTLY. THE CAUSE OF DEATH IS UNKNOWN. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: IT WAS LATER REPORTED THAT THE PATIENTS CAUSE OF DEATH WAS ACUTE CARDIOPULMONARY ARREST DUE TO OR AS A CONSEQUENCE OF: EPILEPSY. NO AUTOPSY WAS PREFORMED AND THE DEVICE WAS EXPLANTED BUT WAS RETURNED TO THE FAMILY. DEVICE IS NOT AVAILABLE FOR RETURN. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE. Manufacturer Narrative: B5, DESCRIBE EVENT OR PROBLEM, CORRECTED DATA: INITIAL REPORT INADVERTENTLY DID NOT INCLUDE CAUSE OF DEATH AND VNS EXPLANATION. CAUSE OF DEATH AND DEVICE EXPLANATIONS WAS REPORTED PRIOR TO INITIAL MDR SUBMISSION. F10, HEALTH EFFECT - CLINICAL CODE. CORRECTED DATA: INITIAL REPORT INADVERTENTLY DID NOT INCLUDE THE CAUSE OF DEATH. DEATH DUE TO CARDIAC IS ADDED. F10, HEALTH EFFECT. CORRECTED DATA: INITIAL REPORT INADVERTENTLY DID NOT INCLUDE DEVICE EXPLANATION. FULL DEVICE EXPLANATION WILL BE ADDED TO THE REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376648&pc=NPT | 2015691-2025-01005 | 2025-01-22 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-13 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Hematoma; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES AFFILIATE IN GERMANY, REGARDING AN IMPLANT OF A 29MM SAPIEN 3 VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. THE VALVE WAS DEPLOYED IN A 90:10 POSITION DESPITE A NON CAPTURE DURING RAPID PACING WITH GOOD RESULT. HOWEVER, AFTER 1-2 MIN, THE BLOOD PRESSURE DROPPED AND PERICARDIAL EFFUSION WAS IDENTIFIED. A PERICARDIOCENTESIS WAS PERFORMED AND ABOUT 1L OF BLOOD WAS DRAINED. THE PATIENT WAS HEMODYNAMICALLY STABLE OVER THE TIME WITH CATECHOLAMINE. AFTER 30 MINUTES, IT WAS DECIDED TO CONVERT INTO A SURGERY. DURING THE SURGERY, A HEMATOMA ON THE LEFT LATERAL WALL WAS OBSERVED MOST PROBABLY CAUSED BY A WIRE PERFORATION. THE PATIENT PASSED AWAY THAT EVENING. Manufacturer Narrative: E1 PHONE NUMBER (B)(6). PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (DEVICE MANIPULATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376675&pc=NIK | 3004209178-2025-02739 | 2025-02-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-13 05:00 | NIK | COMPIA MRI QUAD CRT-D SURESCAN | Failure to Deliver Shock/Stimulation; Defibrillation/Stimulation Problem; Device Sensing Problem; Data Problem | Cardiac Arrest; Tachycardia; Ventricular Fibrillation | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) INTERROGATION REPORT DISPLAYED QUESTION MARKS INSTEAD OF A THRESHOLD VALUE WHEN THE DEVICE FEATURE THAT AUTOMATICALLY MONITORS PACING THRESHOLDS AT PERIODIC INTERVALS IS PROGRAMMED OFF. IT WAS FURTHER REPORTED THAT WHILE SLEEP IN BED IN THE EARLY MORNING, THE PATIENT WENT INTO A POLYMORPHIC VENTRICULAR TACHYCARDIA (VT) AS STATED BY THE PHYSICIAN. IT WAS NOTED THAT THE PATIENT DEVELOPED AN UNSTABLE VENTRICULAR TACHYCARDIA (VT) THAT DETERIORATED INTO VENTRICULAR FIBRILLATION (VF). IT WAS ALSO NOTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) FEATURE THAT HELPS PREVENT THE INAPPROPRIATE DETECTION OF ATRIAL FIBRILLATION (AF) WITH RAPID VENTRICULAR RESPONSE AS A VENTRICULAR TACHYARRHYTHMIA BY EVALUATING THE STABILITY OF VENTRICULAR INTERVALS, INITIALLY DELAYED THERAPY EVEN AFTER SEVERAL ROUNDS OF ANTI-TACHYCARDIA PACING (ATP). THE THERAPY NEVER LED TO A DEFINITIVE SHOCK AND THE PATIENT ULTIMATELY PASSED AWAY AS A RESULT OF CARDIAC ARREST. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. THE DEVICE MEMORY INDICATED PACING CAPTURE MANAGEMENT OF THE LEFT VENTRICLE WAS MISSING/HAD INVALID DATA. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376797&pc=NPT | 2025587-2025-01226 | 2025-02-05 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-13 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Cardiac Arrest; Cardiac Tamponade | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, UPON THE FIRST TWO DEPLOYMENT ATTEMPTS, THE VALVE DISLODGED AND WAS RECAPTURED AND REPOSITIONED. ON THE THIRD DEPLOYMENT ATTEMPT, THE VALVE WAS IMPLANTED WITH GOOD HEMODYNAMIC RESULTS. SHORTLY AFTER THE PROCEDURE ENDED, A CARDIOCIRCULATORY ARREST WAS OBSERVED. RESUSCITATION MANEUVERS, INCLUDING A PERICARDIOCENTESIS, WERE PERFORMED, BUT IT WAS NOT POSSIBLE TO STABILIZE THE PATIENT'S CONDITION, RESULTING IN THE PATIENT'S DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC DISSECTION AT THE SINO-TUBULAR JUNCTION RESULTING CARDIAC TAMPONADE. PER THE PHYSICIAN, THE VALVE CAN NOT BE EXCLUDED AS CONTRIBUTING TO THE PATIENT DEATH. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM (DCS) CAN BE EXCLUDED AS CONTRIBUTING TO THE PATIENT DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT HE PERICARDIOCENTESIS WAS PERFORMED FOR THE AORTIC DISSECTION AT THE SINO-TUBULAR JUNCTION AND RESULTING CARDIAC TAMPONADE. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012558051); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE UPDATED DATA: B5., H6., H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21376830&pc=DSQ | 2916596-2025-00816 | 2025-01-25 05:00 | Death | THORATEC CORPORATION | 2025-02-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure; Hypervolemia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE AND WITHDRAWAL OF SUPPORT. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR FLUID VOLUME OVERLOAD, LEADING TO RESPIRATORY DISTRESS AND INTUBATION. MULTIPLE ATTEMPTS WERE MADE TO EXTUBATE THE PATIENT; HOWEVER, THE PATIENT DID NOT TOLERATE ANY TRIALS AND EXPERIENCED SIGNIFICANT RESPIRATORY DISTRESS EACH TIME. THE MEDICAL TEAM DISCUSSED TRACHEOSTOMY AS AN OPTION FOR LONG-TERM AIRWAY SUPPORT WITH FAMILY. AFTER THOROUGH DISCUSSION, THE FAMILY DECLINED KNOWING THE PATIENT WOULD NOT APPROVE AND ELECTED TO WITHDRAW FURTHER INTERVENTIONS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT NOTED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING RESPIRATORY FAILURE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21377092&pc=QBJ | 3004753838-2025-037151 | 2025-01-16 05:00 | Death | DEXCOM, INC. | 2025-02-13 05:00 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | No Audible Alarm | Hypoglycemia | DEN170088 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THE REPORTER STATED THE PATIENT EXPERIENCED SEVERE HYPOGLYCEMIA DUE TO A FAILED ALERT, AND THAT THE PATIENT PASSED AWAY. THE REPORTER WAS INITIALLY CONTACTED BY DEXCOM ON (B)(6) 2025 AND REQUESTED TO BE CALLED BACK ON THE (B)(6) 2025. WHEN THE REPORTER WAS CALLED ON THE PROVIDED PHONE NUMBER ON THE (B)(6) 2025, THE REPORTER´S OFFICE WAS REACHED, AND THE DEXCOM TECHNICAL SUPPORT REPRESENTATIVE WAS INFORMED THAT THE REPORTER WAS ON SICK LEAVE. THE REPORTER WAS SENT AN EMAIL ON THE (B)(6) 2025 TO OBTAIN FURTHER EVENT DETAILS. NO REPLY TO THE EMAIL WAS RECEIVED, BUT IT WAS CONFIRMED ON (B)(6) 2025, THAT THE REPORTER WOULD BE BACK IN THE OFFICE ON THE (B)(6) 2025 AND COULD BE REACHED AT THAT MOMENT. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. PROBLEM COULD NOT BE CONFIRMED AND PROBABLE CAUSE CANNOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21377230&pc=CBK | 2518422-2025-102029 | 2025-02-11 05:00 | Death | RESPIRONICS, INC. | 2025-02-13 05:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM (B)(6) OF THE (B)(6) IN (B)(6) THAT A 66-YEAR-OLD, MALE PATIENT ON VENTILATION THERAPY USING A TRILOGY EVO VENTILATOR WAS FOUND DECEASED ON (B)(6) 2025, IN (B)(6). THE PATIENT WAS REPORTEDLY FOUND DECEASED WHILE STILL WEARING THE [FULL FACE] VENTILATOR MASK. HOWEVER, THE VENTILATOR WAS UNPLUGGED, AND THE INTERNAL BATTERY AND THE DETACHABLE BATTERY WERE BOTH FULLY DISCHARGED (DEAD). IT IS UNKNOWN AT THIS TIME IF THE DEVICE ALARMED AND WHAT THE DEVICE SETTINGS WERE. THE CAUSE OF DEATH IS ALSO UNKNOWN AT THIS TIME. THERE HAS BEEN NO ALLEGATION OF A DEVICE MALFUNCTION AT THIS TIME. DEVICE EVENT AND ALARM LOGS ARE BEING REQUESTED. RETURN OF THE DEVICE AND CIRCUIT HAVE BEEN REQUESTED FOR EVALUATION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION; THEREFORE, THE FUNCTIONALITY OF THE DEVICE IS NOT ABLE TO BE CONFIRMED AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION HAS BEEN COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21377244&pc=GDW | 3005075853-2025-01199 | 2025-01-29 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-13 05:00 | GDW | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Difficult to Open or Close; Activation Failure; Activation Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K141952 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON WAS UNABLE TO DISENGAGE THE STAPLER. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED. Manufacturer Narrative: (B)(4) DATE SENT: 02/13/25 D4: BATCH #UNK D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. H4: THE DEVICE MANUFACTURE DATE IS CURRENTLY NOT AVAILABLE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WAS THE DEVICE LOCKED ON TISSUE WITH THE JAWS CLOSED? IF YES, HOW WAS THE DEVICE REMOVED? WHAT TROUBLESHOOTING STEPS WERE ATTEMPTED TO OPEN THE DEVICE (E.G. SQUEEZING THE CLOSING TRIGGER WHILE SIMULTANEOUSLY DEPRESSING THE ANVIL RELEASE BUTTON, KNIFE REVERSE SWITCH, MANUAL OVERRIDE)? DID THE JAWS EVENTUALLY OPEN? IS THE CURRENT PATIENT STATUS KNOWN? WAS THERE ANY PATIENT CONSEQUENCE OR CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NO LOT OR BATCH NUMBER WAS PROVIDED THEREFORE A DEVICE HISTORY COULD NOT BE DONE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 2/25/2025. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: 1. WAS THE DEVICE LOCKED ON TISSUE WITH THE JAWS CLOSED? IF YES, HOW WAS THE DEVICE REMOVED? YES, IT COULD NOT BE REMOVED. 2. WHAT TROUBLESHOOTING STEPS WERE ATTEMPTED TO OPEN THE DEVICE (E.G. SQUEEZING THE CLOSING TRIGGER WHILE SIMULTANEOUSLY DEPRESSING THE ANVIL RELEASE BUTTON, KNIFE REVERSE SWITCH, MANUAL OVERRIDE)? ACCORDING TO THE SURGEON, THE VENDOR WAS CONTACTED, HE WAS SPEAKING TO THE VENDOR DURING THE CASE. HE ATTEMPTED TO KNIFE REVERSE SWITCH (DIDN¿T WORK) FOLLOWED BY THE MANUAL OVERRIDE AND INSTRUCTIONS RELAYED BY THE VENDOR OVER THE PHONE. 3. DID THE JAWS EVENTUALLY OPEN? NO (WAS REMOVED WHEN PATIENT EXPIRED). 4. IS THE CURRENT PATIENT STATUS KNOWN? YES, DECEASED. 5. WAS THERE ANY PATIENT CONSEQUENCE OR CHANGE IN THE POST-OPERATIVE CARE OF THE PATIENT AS A RESULT OF THE EVENT? PATIENT EXPIRED. Manufacturer Narrative: (B)(4). DATE SENT: 2/28/2025. UPON REVIEW OF THE ADDITIONAL INFORMATION PROVIDED, IT WAS CONCLUDED THAT THIS EVENT IS A DUPLICATE OF THE EVENT REPORTED UNDER MDR#3005075853-2025-01040.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21378389&pc=NPT | 2025587-2025-01228 | 2025-02-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-13 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | High Blood Pressure/ Hypertension; Renal Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT FAMILY MEMBER THAT THREE MONTHS AFTER UNDERGOING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS ALSO NOTED THAT THE PATIENT EXPERIENCED HYPERTENSION AND KIDNEY DECLINE FOLLOWING THE TAVI. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21378590&pc=NPT | 2025587-2025-01229 | 2025-01-21 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-13 05:00 | NPT | EVOLUT FX DCS | Perivalvular Leak; Patient Device Interaction Problem | Cardiac Arrest; Hemorrhage/Bleeding; Vascular Dissection; Heart Block; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH MILD ANEMIA, THROMBOCYTOPENIA, A HISTORY OF PREVIOUS ENDOVASCULAR ANEURYSM REPAIR, AND CALCIFICATION AT THE COMMON ILIAC ARTERY (CIA), ACCESS WAS OBTAINED VIA THE RIGHT LOWER LIMB. IT WAS NOTED THAT THE DIAMETER OF THE ACCESS SITE BLOOD VESSELS WERE NARROW, AND AN ENDOSKELETON STENT GRAFT WAS PREVIOUSLY PLACED IN THE ABDOMEN FURTHER COMPLICATING ACCESS. DESPITE THIS, AN 18 FR INTRODUCER SHEATH WAS INSERTED INTO THE PATIENT IN AN ATTEMPT TO ADDRESS THE AREAS OF CALCIFICATION AT THE CIA AND STENOSIS AT THE GRAFT BIFURCATION. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED, DURING WHICH COMPLETE HEART BLOCK (CHB) WAS NOTED. THE PATIENT ENTERED CARDIAC ARREST FOLLOWING THE BAV; HOWEVER, THE CARDIAC BEAT RESUMED WITH TEMPORARY PACING. ALTHOUGH THE ANNULUS PERIMETER WAS SIZED FOR A 26 MM VALVE, THE DECISION WAS MADE TO USE A 23 MM VALVE. THIS DECISION WAS MADE AS THE LEFT INTERNAL THORACIC ARTERY-LEFT ANTERIOR DESCENDING CORONARY ARTERY WAS NEARLY OCCLUDED. ADDITIONALLY, THERE WAS A THREE-VESSEL BYPASS WITH THE SAFENA-RIGHT CORONARY ARTERY AND LEFT CIRCUMFLEX ARTERY. PER THE PHYSICIAN, THIS WAS DONE TO LIMIT THE RISK OF OCCLUSION. SUBSEQUENTLY, THE VALVE WAS IMPLANTED IN THE PATIENT. FOLLOWING VALVE IMPLANT THE MEAN PRESSURE GRADIENT WAS 5 MM HG. MODERATE PARAVALVULAR LEAK (PVL) WAS OBSERVED AND COMPARED TO THE TRIVIA L PRE-OPERATIVE PVL. A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED. DUE TO CALCIFICATION AT THE SINOTUBULAR JUNCTION (STJ), THERE WAS CONSIDERATION FOR INCREASING THE BALLOON SIZE FOR THE POST-IMPLANT BAV. THE PHYSICIAN ELECTED NOT TO INCREASE THE BALLOON BEYOND THE SIZE OF THE STJ SHORT AXIS GIVEN THE PATIENT'S THROMBOCYTOPENIA, THIN BLOOD VESSELS IN THE LEG, THE RISK OF RUPTURE, AND LENGTHY BLOOD FLOW OCCLUSION TIME. DIGITAL SUBTRACTION ANGIOGRAPHY (DSA) REVEALED A PROXIMAL COMMON ILIAC ARTERY DISSECTION WHICH WAS TREATED USING A 7 MM ENDOVASCULAR THERAPY BARE METAL STENT. THE PUNCTURE SITE WAS CHECKED WITH BOTH DSA AND A ECHOCARDIOGRAM, BUT NO BLEEDING WAS OBSERVED. THE INCISION WAS CLOSED AND COMPRESSION HEMOSTASIS WAS PERFORMED. THE PATIENT WAS DISCHARGED TO THE INTENSIVE CARE UNIT (ICU) WITH THE TEMPORARY PACEMAKER STILL IN PLACE DUE TO THE CHB. THE ANESTHESIOLOGIST NOTED THAT THE PATIENTS HEMOGLOBIN VALUE HAD DROPPED AND THAT ANEMIA WAS MILDLY PROGRESSING. AFTER WAKING FROM ANESTHESIA, THE PATIENT BEGAN TO COMPLAIN OF LEFT SIDED PAIN. THE HEMOGLOBIN LEVEL SUDDENLY DECREASED FROM 8.9 AT THE TIME OF ICU ADMISSION TO 6.6, AND AN EMERGENCY COMPUTED TOMOGRAPHY SCAN REVEALED BLEEDING AT THE PUNCTURE SITE WHICH HAD SPREAD TO THE RETROPERITONEUM. CARDIAC ARREST OCCURRED AND VASCULAR REPAIR SURGERY WAS PERFORMED WHILE ALSO PERFORMING CARDIAC MASSAGE. HEMOSTASIS WAS ACHIEVED, HOWEVER, THE HEMOGLOBIN DID NOT INCREASE AND LACTATE ROSE TO 11. ACIDOSIS PROGRESSED RAPIDLY, LEADING TO ORGAN DAMAGE AND CARDIAC FUNCTION IMPAIRMENT. THE PATIENT PASSED AWAY AT 8:27 AM THE MORNING AFTER THE PROCEDURE.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH MILD ANEMIA, THROMBOCYTOPENIA, A HISTORY OF PREVIOUS ENDOVASCULAR ANEURYSM REPAIR, AND CALCIFICATION AT THE COMMON ILIAC ARTERY (CIA), ACCESS WAS OBTAINED VIA THE RIGHT LOWER LIMB. IT WAS NOTED THAT THE DIAMETER OF THE ACCESS SITE BLOOD VESSELS WERE NARROW, AND AN ENDOSKELETON STENT GRAFT WAS PREVIOUSLY PLACED IN THE ABDOMEN FURTHER COMPLICATING ACCESS. DESPITE THIS, AN 18 FR INTRODUCER SHEATH WAS INSERTED INTO THE PATIENT IN AN ATTEMPT TO ADDRESS THE AREAS OF CALCIFICATION AT THE CIA AND STENOSIS AT THE GRAFT BIFURCATION. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED, DURING WHICH COMPLETE HEART BLOCK (CHB) WAS NOTED. THE PATIENT ENTERED CARDIAC ARREST FOLLOWING THE BAV; HOWEVER, THE CARDIAC BEAT RESUMED WITH TEMPORARY PACING. ALTHOUGH THE ANNULUS PERIMETER WAS SIZED FOR A 26 MM VALVE, THE DECISION WAS MADE TO USE A 23 MM VALVE. THIS DECISION WAS MADE AS THE LEFT INTERNAL THORACIC ARTERY-LEFT ANTERIOR DESCENDING CORONARY ARTERY WAS NEARLY OCCLUDED. ADDITIONALLY, THERE WAS A THREE-VESSEL BYPASS WITH THE SAFENA-RIGHT CORONARY ARTERY AND LEFT CIRCUMFLEX ARTERY. PER THE PHYSICIAN, THIS WAS DONE TO LIMIT THE RISK OF OCCLUSION. SUBSEQUENTLY, THE VALVE WAS IMPLANTED IN THE PATIENT. FOLLOWING VALVE IMPLANT THE MEAN PRESSURE GRADIENT WAS 5 MM HG. MODERATE PARAVALVULAR LEAK (PVL) WAS OBSERVED AND COMPARED TO THE TRIVIA L PRE-OPERATIVE PVL. A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED. DUE TO CALCIFICATION AT THE SINOTUBULAR JUNCTION (STJ), THERE WAS CONSIDERATION FOR INCREASING THE BALLOON SIZE FOR THE POST-IMPLANT BAV. THE PHYSICIAN ELECTED NOT TO INCREASE THE BALLOON BEYOND THE SIZE OF THE STJ SHORT AXIS GIVEN THE PATIENT'S THROMBOCYTOPENIA, THIN BLOOD VESSELS IN THE LEG, THE RISK OF RUPTURE, AND LENGTHY BLOOD FLOW OCCLUSION TIME. DIGITAL SUBTRACTION ANGIOGRAPHY (DSA) REVEALED A PROXIMAL COMMON ILIAC ARTERY DISSECTION WHICH WAS TREATED USING A 7 MM ENDOVASCULAR THERAPY BARE METAL STENT. THE PUNCTURE SITE WAS CHECKED WITH BOTH DSA AND A ECHOCARDIOGRAM, BUT NO BLEEDING WAS OBSERVED. THE INCISION WAS CLOSED AND COMPRESSION HEMOSTASIS WAS PERFORMED. THE PATIENT WAS DISCHARGED TO THE INTENSIVE CARE UNIT (ICU) WITH THE TEMPORARY PACEMAKER STILL IN PLACE DUE TO THE CHB. THE ANESTHESIOLOGIST NOTED THAT THE PATIENTS HEMOGLOBIN VALUE HAD DROPPED AND THAT ANEMIA WAS MILDLY PROGRESSING. AFTER WAKING FROM ANESTHESIA, THE PATIENT BEGAN TO COMPLAIN OF LEFT SIDED PAIN. THE HEMOGLOBIN LEVEL SUDDENLY DECREASED FROM 8.9 AT THE TIME OF ICU ADMISSION TO 6.6, AND AN EMERGENCY COMPUTED TOMOGRAPHY SCAN REVEALED BLEEDING AT THE PUNCTURE SITE WHICH HAD SPREAD TO THE RETROPERITONEUM. CARDIAC ARREST OCCURRED AND VASCULAR REPAIR SURGERY WAS PERFORMED WHILE ALSO PERFORMING CARDIAC MASSAGE. HEMOSTASIS WAS ACHIEVED, HOWEVER, THE HEMOGLOBIN DID NOT INCREASE AND LACTATE ROSE TO 11. ACIDOSIS PROGRESSED RAPIDLY, LEADING TO ORGAN DAMAGE AND CARDIAC FUNCTION IMPAIRMENT. THE PATIENT PASSED AWAY AT 8:27 AM THE MORNING AFTER THE PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 3.6MMX5.3MM. THE PATIENT WENT INTO CARDIAC ARREST AFTER THE PRE- IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV). PER THE PHYSICIAN, DUE TO THROMBOCYTOPENIA, A PUNCTURE CUT-DOWN ACCESS WAS AVOIDED RESULTING IN INADEQUATE HEMOSTASIS AND BLEEDING TO THE RETROPERITONEAL SIDE. ACIDOSIS PROGRESSED DUE TO NECROSIS OF THE INTESTINAL TRACT DUE TO THE PLACEMENT OF A STENT NEAR THE FEMORAL HEAD MID DURING THE TREATMENT OF THE LOWER LIMB AT MULTIPLE CENTERS AND THE PRESENCE OF A THICK SHEATH IN THE LONG-TERM ACCESS BLOOD VESSELS, RESULTING IN DEATH DUE TO CARDIAC ARREST COUPLED WITH HEMORRHAGIC SHOCK. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012199763); PRODUCT TYPE: 0195-HEART VALVES; SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX VALVE; PRODUCT ID EVOLUTFX-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: INFORMATION WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. 2. OUTCOMES ATTRIBUTED TO ADVERSE EVENT UPDATED ADDITIONAL CODES. ANNEX F CODES ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21378805&pc=DSQ | 2916596-2025-00872 | 2020-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿A MACHINE LEARNING-DERIVED RISK SCORE IMPROVES PREDICTION OF OUTCOMES AFTER LVAD IMPLANTATION: AN ANALYSIS OF THE INTERMACS DATABASE?¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE RELATED TO DEATH, HEART TRANSPLANTATION AND LEFT VENTRICULAR ASSIST DEVICE (LVAD) EXPLANT. MACHINE LEARNING ASSESSMENT OF RISK AND EARLY MORTALITY IN HEART FAILURE (MARKER-HF) SCORE WAS CALCULATED USING VARIABLES COLLECTED ON THE DAY OF SURGERY, BEFORE LVAD IMPLANTATION. PATIENTS WITH MISSING DATA ON DISCHARGE STATUS (N=1467) OR LENGTH OF STAY LOS (N =1083) AND THOSE WITH CLEARLY ERRONEOUS DATA (N=907), WHO UNDERWENT HEART TRANSPLANTATION OR LVAD EXPLANTATION DURING THE INDEX ADMISSION (N=268), WITH MISSING INTERAGENCY REGISTRY OF MECHANICALLY ASSISTED CIRCULATORY SUPPORT (INTERMACS) PROFILE (N=86), OR UNKNOWN DATE OF DEATH (N =1) WERE EXCLUDED FROM THE STUDY. A TOTAL OF 25,365 PATIENTS WERE INCLUDED IN THE FINAL ANALYSIS, PATIENTS AGED 18-80 YEARS WHO UNDERWENT LVAD IMPLANTATIONS BETWEEN 2006 AND 2020. THE STUDY OBJECTIVES WERE TO ASSESS THE ABILITY OF MARKER-HF TO PREDICT OUTCOMES AFTER LVAD IMPLANTATION, COMPARE ITS PREDICTIVE ACCURACY WITH INTERMACS PROFILES, AND DETERMINE WHETHER COMBINING THE SCORES WOULD HELP TO PREDICT OUTCOMES. THE PRIMARY END POINT WAS IN-HOSPITAL MORTALITY. SECONDARY END POINTS WERE LOS AFTER LVAD IMPLANTATION AND 1-YEAR POST DISCHARGE AND 1-YEAR ALL-CAUSE MORTALITY ACCORDING TO THE MARKER-HF SCORE. THE 1-YEAR ALL-CAUSE MORTALITY ALSO INCLUDED PATIENTS WHO DIED WITHIN THE INDEX HOSPITALIZATION, WHEREAS THE 1-YEAR POST DISCHARGE MORTALITY INCLUDED ONLY THOSE PATIENTS WHO SURVIVED TO DISCHARGE. PATIENTS WITH AN HM3 DEVICE EXPERIENCED A HIGHER IN HOSPITAL MORTALITY (6.5% VS 5.2%, P < .001), BUT MARKEDLY LOWER 1-YEAR POST DISCHARGE MORTALITY (5.8% VS 13.4%, P < .001) THAN THOSE WITH OTHER DEVICE TYPES, RESULTING IN A LOWER OVERALL 1-YEAR ALL-CAUSE MORTALITY (11.9% VS 17.9%, P < .001). THE 1-YEAR POST DISCHARGE MORTALITY RATES INCREASED AS MARKER-HF SCORES INCREASED, RANGING FROM 8.2% IN THE HIGHEST RISK GROUP TO 4.1% IN THE LOWEST. AS INTERMACS PROFILES DECLINED, THE 1-YEAR MORTALITY RATES IMPROVED FROM 8.0% TO 3.6%. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AND OTHER SPECIFIC CASE/PATIENT INFORMATION ARE NOT AVAILABLE AND WERE NOT REQUESTED. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION WERE NOT PROVIDED AND WERE UNABLE TO BE DETERMINED THROUGH THIS EVALUATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2020 SINCE THE PATIENTS WERE IMPLANTED BETWEEN 2006 AND 2020. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. JIN JOO PARK ET AL., JOURNAL OF CARDIAC FAILURE 00(2024)1:10. DOI:10.1016/J.CARDFAIL.2024.09.013 CARDIOLOGY DEPARTMENT, UNIVERSITY OF CALIFORNIA, SAN DIEGO, LAJOLLA, CALIFORNIA, USA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21379195&pc=NGV | 2135147-2025-00718 | 2025-01-24 05:00 | Death | ABBOTT MEDICAL | 2025-02-13 05:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device; Patient-Device Incompatibility | Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypoxia; Shock; Cardiac Tamponade; Vascular Dissection; No Clinical Signs, Symptoms or Conditions | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 28MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT. THE PATIENT'S CONDITION PRIOR TO PROCEDURE WAS REPORTED AS HAVING RECURRENT GASTROINTESTINAL BLEEDING REQUIRING TRANSFUSION UNDER NOAK IN RECURRENT ATRIAL FIBRILLATION. FLUID WAS GIVEN PRIOR TO TRANSSEPTAL PUNCTURE. DURING PROCEDURE, TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) AND ANGIOGRAPHY WERE UTILIZED. TENSION TEST WAS PERFORMED AND THEN THE AMULET WAS RELEASED. THE AMULET WAS "NOT COMPRESSED ENOUGH" AT TIME OF RELEASE. IMMEDIATELY AFTER RELEASE, TEE AND ANGIOGRAPHY SHOWED THAT DEVICE EMBOLIZATION OCCURRED AND THE AMULET HAD EMBOLIZED FROM THE LEFT ATRIAL APPENDAGE TO THE MITRAL VALVE. THE EMBOLIZED AMULET DID NOT CAUSE PARTIAL OR COMPLETE OBSTRUCTION. CLOSE CRITERIA HAD NOT BEEN ACHIEVED PRIOR TO RELEASE; THE AMULET WAS "NOT DEEP ENOUGH UNDER CX" AND THERE WAS "NOT ENOUGH DISTANCE BETWEEN DISC AND CORPUS." THE AMULET WAS PERCUTANEOUSLY RETRIEVED USING A SNARE VIA THE 14F AMPLATZER AMULET SHEATH. THE PATIENT WAS REPORTED TO HAVE EXPERIENCED HEMODYNAMIC INSTABILITY DURING PROCEDURE. NO DEVICE WAS ULTIMATELY IMPLANTED. ON (B)(6) 2025, THE PATIENT WAS REPORTED TO HAVE PASSED.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED AND THE PROCEDURE DESCRIPTION WAS UPDATED: IT WAS REPORTED THAT ON (B)(6) 2025, A 28MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT. THE PATIENT'S CONDITION PRIOR TO PROCEDURE WAS REPORTED AS HAVING RECURRENT GASTROINTESTINAL BLEEDING REQUIRING TRANSFUSION UNDER ORAL ANTICOAGULANT DUE TO RECURRENT ATRIAL FIBRILLATION. IT WAS NOTED THAT FLUID WAS GIVEN PRIOR TO TRANSSEPTAL PUNCTURE. DURING PROCEDURE, TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) AND ANGIOGRAPHY WERE UTILIZED. CLOSE CRITERIA HAD NOT BEEN ACHIEVED PRIOR TO RELEASE; THE AMULET WAS "NOT DEEP ENOUGH UNDER CX" AND THERE WAS "NOT ENOUGH DISTANCE BETWEEN DISC AND CORPUS." TENSION TEST WAS PERFORMED AND THEN THE AMULET WAS RELEASED. THE AMULET WAS "NOT COMPRESSED ENOUGH" AT TIME OF RELEASE. UNFORTUNATELY, THREE MINUTES AFTER RELEASE, TEE AND ANGIOGRAPHY SHOWED THAT DEVICE EMBOLIZATION OCCURRED AND THE AMULET HAD EMBOLIZED FROM THE LEFT ATRIAL APPENDAGE TO THE LEFT VENTRICLE. THE EMBOLIZED AMULET DID NOT CAUSE PARTIAL OR COMPLETE OBSTRUCTION. AFTER SEVERAL ATTEMPTS, THE AMULET DEVICE WAS PERCUTANEOUSLY RETRIEVED USING A SNARE VIA THE 14F AMPLATZER AMULET SHEATH. THE PATIENT WAS REPORTED TO HAVE EXPERIENCED HEMODYNAMIC INSTABILITY DURING PROCEDURE. PRIOR TO RETRIEVAL THE PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION (CPR) WITH RETURN OF SPONTANEOUS CIRCULATION (ROSC) AFTER 5-8 MINUTES. ECHOCARDIOGRAPHY SHOWED MODERATE TO SEVERE MITRAL INSUFFICIENCY WHEREAS PREVIOUSLY IT WAS ONLY MILD MITRAL INSUFFICIENCY. THE PRESENCE OF THE DEVICE FLOATING SUGGESTED A PARTIAL TEAR OF THE MITRAL VALVE LEAFLET. PERICARDIAL EFFUSION DEVELOPED AND A PERICARDIOCENTESIS WAS PERFORMED. INITIALLY TRAINING 900ML OF BLOODY FLUID. AFTER TWO HOURS, INCREASING PERICARDIAL EFFUSION WITH ON-GOING PERICARDIAL DRAINAGE AND NOREPINEPHRINE WAS REQUIRED. THE PATIENT REQUIRED CPR FOR 25 MINUTES. DUE TO RECURRENT PERICARDIAL EFFUSION/TAMPONADE WITH INCREASING HEMODYNAMIC INSTABILITY, THE PATIENT WAS TRANSFERRED TO THE CARDIAC SURGERY DEPARTMENT. NO DEVICE WAS ULTIMATELY IMPLANTED. THE PATIENT WAS DIAGNOSED WITH CARDIOGENIC SHOCK CAUSED BY PERICARDIAL TAMPONADE FOLLOWING IATROGENIC HEART INJURY. THERE WAS MODERATE MITRAL INSUFFICIENCY DUE TO CHORDAE TENDINEAE RUPTURE OF THE ANTERIOR LEAFLET. THE PATIENT WAS STATUS POST IATROGENIC ATRIAL SEPTAL DEFECT (ASD), LEFT HEART FAILURE NEW YORK HEART ASSOCIATION (NYHA) CLASSIFICATION TYPE THREE. ON (B)(6) 2025, THE PATIENT WAS REPORTED TO HAVE A SUDDEN DROP IN BLOOD PRESSURE AND DESATURATION AND CPR WAS INITIATED. THE PATIENT PASSED DUE TO POOR PROGNOSIS AND FAILURE TO ACHIEVE ROSC WHILE IN THE INTENSIVE CARE UNIT. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF MIGRATION OR EXPULSION OF DEVICE (DEVICE EMBOLIZATION) AND PATIENT-DEVICE INCOMPATIBILITY (IMPLANT-RELATED TISSUE DAMAGE) WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR INCIDENTS REPORTED FROM THIS LOT. THE PROVIDED PROCEDURAL IMAGING AND INFORMATION WERE REVIEWED BY MEDICAL AFFAIRS. BASED ON THE INFORMATION RECEIVED, THE CAUSE FOR THE REPORTED IMPLANT-RELATED TISSUE DAMAGE COULD NOT BE DETERMINED. THE REPORTED DEVICE EMBOLIZATION APPEARS TO BE RELATED TO PROCEDURAL CONDITIONS. THE PATIENT EFFECTS OF DEATH DUE TO POOR PROGNOSIS AND FAILURE TO ACHIEVE ROSC WHILE IN THE INTENSIVE CARE UNIT, PER DETAIL EVENTS. PER CASE DETAILS, THE PATIENT WAS DIAGNOSED WITH CARDIOGENIC SHOCK CAUSED BY PERICARDIAL TAMPONADE FOLLOWING IATROGENIC HEART INJURY. THE PATIENT EFFECTS OF CARDIAC ARREST, CARDIAC TAMPONADE, HYPOXIA, VASCULAR DISSECTION, AND HYPOTENSION COULD NOT DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. CODES REMOVED: H6 HEALTH EFFECT-CLINICAL CODE: 4582 NO CLINICAL SIGNS, SYMPTOMS OR CONDITIONS. CODES ADDED: H6 HEALTH EFFECT-CLINICAL CODE:1914,1762, 2072, 1918, 2226, 3160. H6 MEDICAL DEVICE PROBLEM CODE: 2682.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21379241&pc=DSQ | 2916596-2025-01046 | 2021-07-01 04:00 | Death | THORATEC CORPORATION | 2025-02-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "PERICARDIAL CLOSURE WITH EXPANDED POLYTETRAFLUOROETHYLENE PATCH IN LEFT VENTRICULAR ASSIST DEVICE SURGERY" THAT 166 PATIENTS UNDERWENT LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION, AND 116 PATIENTS UNDERWENT PERICARDIAL CLOSURE USING AN EXPANDED POLYTETRAFLUOROETHYLENE (EPTFE) PATCH TO INVESTIGATE OUTCOMES OF PERICARDIAL CLOSURE. OVERALL, 21 PATIENTS (13%) DEVELOPED MEDIASTINITIS. CANDIDA ALBICANS WAS DETECTED IN THREE CASES (14%), CORYNEBACTERIUM SPECIES IN TWO CASES (10%), AND ENTEROCOCCUS FAECIUM IN ONE CASE (5%). PSEUDOMONAS AERUGINOSA WAS DETECTED IN THREE CASES (14%). STAPHYLOCOCCUS AUREUS WAS DETECTED IN THREE CASES (14%), STAPHYLOCOCCUS EPIDERMIDIS IN TWO CASES (10%), AND STAPHYLOCOCCUS HAEMOLYTICUS IN ONE CASE (5%). STREPTOCOCCUS ORALIS WAS DETECTED IN ONE CASE (5%) AND SERRATIA MARCESCENS IN ONE CASE (5%). IN SEVEN CASES (33%) NO PATHOGEN WAS DETECTED. NO SIGNIFICANT DIFFERENCE IN OVERALL SURVIVAL BETWEEN BOTH GROUPS WAS FOUND IN THE KAPLAN¿MEIER ANALYSIS (P = 0.29). 61 PATIENTS (37%) DEVELOPED DRIVELINE INFECTIONS (DLI). DLI OCCURRED MORE OFTEN IN PATIENTS WITH AN EPTFE PATCH. OVERALL, 49 OF 166 PATIENTS (29%) REQUIRED RETURN TO THE OPERATING THEATER IN 30 DAYS. SPECIFICALLY, 2 PATIENTS (1%) UNDERWENT DELAYED STERNAL CLOSURE, 40 (24%) UNDERWENT REDO STERNOTOMY, AND 9 (5%) UNDERWENT SUBXIPHOID PERICARDIOTOMY FOR PERICARDIAL EFFUSION. IN TOTAL, 21 PATIENTS DEVELOPED MEDIASTINITIS, 12 OF WHOM (57%) REQUIRED RETURN TO THE THEATER IN 30 DAYS. OF THE 166 PATIENTS, 18 HAD A STROKE POST OP, 82 HAD BLEEDING POST-OP, 40 REQUIRED A TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE POST OP AND 14 PATIENTS HAD PUMP THROMBOSIS POST OP. IN THIS RETROSPECTIVE SINGLE-CENTER STUDY, PATIENTS WHO UNDERWENT LVAD IMPLANTATION AND HAD PERICARDIAL CLOSURE WITH AN EPTFE PATCH, DEVELOPED MEDIASTINITIS (OR PUMP POCKET INFECTION, WHICH IS CONSIDERED A SUBTYPE OF MEDIASTINITIS) MORE FREQUENTLY THAN PATIENTS WITHOUT. FROM THIS STUDY, IT IS CURRENTLY CAUTIONED AGAINST THE USE OF EPTFE PATCHES DURING LVAD IMPLANTATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DEPARTMENT OF CARDIAC SURGERY, HEIDELBERG UNIVERSITY HOSPITAL, HEIDELBERG, GERMANY. SCHNOERING L, KHATTAB MA, AKHYARI P, ET AL. PRESSURE-DIMENSION INDEX AND LEFT VENTRICULAR SPHERICITY INDEX FOLLOWING HEARTMATE II AND HEARTMATE 3 IMPLANTATION. ESC HEART FAILURE. 2024;11(5):3012-3022. HTTPS://DIALOG.PROQUEST.COM/PROFESSIONAL/DOCVIEW/3065006187?ACCOUNTID=152602. DOI: HTTP://DX.DOI.ORG/10.1002/EHF2.14839. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21379484&pc=DTB | 2649622-2025-04310 | 2024-12-12 05:00 | Death | MPRI | 2025-02-13 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis; Septic Shock | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BACTERIAL POCKET INFECTION REQUIRING ANTIBIOTICS. ADDITIONALLY, THE PATIENT UNDERWENT A IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) POCKET DEBRIDEMENT SURGERY. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED SEPSIS LEADING TO SEPTIC SHOCK. THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 6721S-50 LEAD AND 4968-60 LEAD; IMPLANTED: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21379489&pc=DTB | 2649622-2025-04311 | 2024-12-12 05:00 | Death | MPRI | 2025-02-13 05:00 | DTB | CAPSURE EPICARDIAL LEAD | Adverse Event Without Identified Device or Use Problem | Sepsis; Septic Shock | P950024 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BACTERIAL POCKET INFECTION REQUIRING ANTIBIOTICS. ADDITIONALLY, THE PATIENT UNDERWENT A IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) POCKET DEBRIDEMENT SURGERY. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED SEPSIS LEADING TO SEPTIC SHOCK. THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTION: H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD, IMPLANTED: (B)(6) 2008 AND 6721S-50 LEAD, IMPLANTED: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21379508&pc=LWS | 2182208-2025-00736 | 2024-12-12 05:00 | Death | MEDTRONIC, INC. | 2025-02-13 05:00 | LWS | UNIPOLAR, EPICARDIAL, OVAL PATCH LEAD, DF1 | Adverse Event Without Identified Device or Use Problem | Sepsis; Septic Shock | P900061 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BACTERIAL POCKET INFECTION REQUIRING ANTIBIOTICS. ADDITIONALLY, THE PATIENT UNDERWENT A IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) POCKET DEBRIDEMENT SURGERY. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED SEPSIS LEADING TO SEPTIC SHOCK. THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD, IMPLANTED: (B)(6) 2008 AND 4968-60 LEAD, IMPLANTED: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21381394&pc=OZD | 1220648-2025-26132 | 2025-01-19 05:00 | Death | ABIOMED, INC. | 2025-02-13 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Perforation | P140003 |  | 1 | Event Description: COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SMART ASSIST SUPPORT AND DURING SUPPORT, THE PATIENT'S LEFT VENTRICLE RUPTURED. THE CAUSE OF THE RUPTURE IS UNKNOWN AND THE RELATIONSHIP BETWEEN THE RUPTURE AND THE IMPELLA IS UNKNOWN. THE VENTRICLE RUPTURE CAUSED THE PATIENT TO EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS PER THE CLINICAL INFORMATION PROVIDED, AFTER TWO DAYS OF PUMP INSERTION, THERE WAS A VENTRICLE PERFORATION OBSERVED IE. LEFT VENTRICLE RUPTURE, BUT THE RELATION TO IMPELLA IS UNKNOWN. THEREFORE, THE ROOT CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26132 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 2152 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26132. NEW CODE WAS ADDED TO HEALTH EFFECT - CLINICAL CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21381407&pc=OZD | 1220648-2025-26136 | 2025-01-25 05:00 | Death | ABIOMED, INC. | 2025-02-13 05:00 | OZD | IMPELLA CP | Restricted Flow rate; Pumping Stopped | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. HOWEVER, FOLLOWING THE IMPLANT, IT WAS NOTED THAT SUPPORT WAS INADEQUATE ABOVE PERFORMANCE LEVEL P-6. THE PATIENT WAS TAKEN INTO THE INTENSIVE CARE UNIT WHERE THE POSITION WAS ADJUSTED, AND BOTH THE VOLUME AND HEART STATUS WERE EVALUATED. DURING TROUBLESHOOTING, THE PUMP UNEXPECTEDLY STOPPED AND WAS ABLE TO BE TURNED BACK ON. HOWEVER, FLOWS WOULD NOT INCREASE BEYOND .3-.5. DURING THIS TIME, THE PATIENT CODED MULTIPLE TIMES. THE PATIENT'S FAMILY ULTIMATELY MADE THE DECISION TO WITHDRAW CARE, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT, LOW PUMP FLOW, AND TEMPORARY PUMP STOP HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG SHOWS LOW PUMP FLOW WITH AN ACTIVE SUCTION ALARM WHILE RUNNING ON A HIGHER P-LEVEL SETTING. THE PUMP FLOW ISSUE WAS RESOLVED AFTER TURNING TO P2. A PUMP STOP WAS REPORTED WITH A MOTOR CURRENT HIGH ALARM, BUT IT WAS ABLE TO RESTART. HOWEVER, THE PUMP WAS UNABLE TO ACHIEVE PROPER FLOW ON P2 AFTER THE RESTART. ADDITIONALLY, THE PATIENT WAS CODED MULTIPLE TIMES WHILE ON SUPPORT. THE CAUSE OF THE TEMPORARY PUMP ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON DATA LOG REVIEW. THE CAUSE OF THE LOW PUMP ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON DATA LOG REVIEW. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26136 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21381586&pc=LRG | 1061932-2025-00018 | 2025-01-17 05:00 | Death | BECKMAN COULTER | 2025-02-13 05:00 | LRG | DXM 1040 MICROSCAN WALKAWAY | Incorrect, Inadequate or Imprecise Result or Readings | No Clinical Signs, Symptoms or Conditions | K911400 |  | 1 | Event Description: THE CUSTOMER REPORTED GETTING HIGH PROBABILITY MISIDENTIFICATION (HPID) OF CITROBACTER FARMERI ON PATIENT SAMPLES USING NEG URINE COMBO 90 (NUC90, P/N C48640, LOT # 2025-11-11) PANELS PROCESSED IN THEIR DXM 1040 MICROSCAN WALKAWAY INSTRUMENT S/N (B)(6). THE CUSTOMER STATED THE MISIDENTIFICATION IMPACTED ONE PATIENT SAMPLE, MULTIPLE CULTURES. THE CUSTOMER NOTIFIED THE PHYSICIAN OF THE ERRONEOUS RESULTS OBTAINED FROM URINE SAMPLE. THE CUSTOMER STATED THE PATIENT WAS PUT ON HOSPICE AND PASSED AWAY ON (B)(6) 2025. IT WAS FURTHER NOTED THAT THE PATIENT DID NOT HAVE CAPACITY AND WAS REFUSING TREATMENT. THE PATIENT¿S FAMILY OPTED FOR COMFORT FOCUSED MEASURES. THE CUSTOMER INDICATED THEY REACHED OUT TO THEIR INTERNAL TEAM, AND THEY STATED THE MISIDENTIFICATION CAUSED A DELAY IN APPROPRIATE TREATMENT FOR THE PATIENT. Manufacturer Narrative: THE CAUSE OF DEATH IS UNKNOWN. THE INFORMATION AVAILABLE SUGGESTS THE MICROSCAN SYSTEM FAILED TO IDENTIFY THE ORGANISM CORRECTLY BASED ON THE ATYPICAL BEHAVIOR OF THE ISOLATE ITSELF AND NEGATIVE LYS REACTION, WHICH IS A KEY BIOCHEMICAL REACTION USED BY THE SYSTEM TO IDENTIFY AN E. COLI. AS INSTRUMENT AND PANEL QUALITY CONTROL TESTING WERE ALL IN RANGE, A PRODUCT PERFORMANCE ISSUE OR MALFUNCTION IS NOT INDICATED. PANEL DATA DID NOT INDICATE AN OBVIOUS INSTRUMENT OR PANEL ISSUE AND CONFIRMED THE ATYPICAL BEHAVIOR OF THE SAMPLE. THEREFORE, THE PROBABLE CAUSE OF THE REPORTED ORGANISM MISIDENTIFICATION WAS AN ATYPICAL ISOLATE. SECTION A4 AND A5: INFORMATION NOT PROVIDED BY THE CUSTOMER. THE BECKMAN COULTER INTERNAL IDENTIFIER IS (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21381605&pc=NPT | 2015691-2025-01041 | 2025-01-24 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-13 05:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture; Cardiac Tamponade | P140031 |  | 1 | Event Description: AS REPORTED FROM OUR AFFILIATES IN AUSTRIA, THIS WAS A CASE WITH A 29MM SAPIEN 3 VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. BEFORE STARTING THE PROCEDURE, THE PATIENT WAS IN POOR CONDITION WITH NAUSEA AND VOMITING. DURING THE PROCEDURE, THE VALVE WAS DEPLOYED WITHOUT ISSUE, UNDERFILLING THE BALLOON BY -1 ML. THE PATIENT WAS STABLE UNTIL THE CLOSURE OF THE FEMORAL VESSEL. DURING THE CLOSURE OF THE FEMORAL VESSEL, THE PATIENT SUDDENLY HAD A CIRCULATORY COLLAPSE AND HAD TO BE RESUSCITATED. A CARDIAC TAMPONADE WAS DETECTED BY ECHO. UNFORTUNATELY, THERE WAS NO TIME FOR OPEN SURGERY AND THE PATIENT PASSED AWAY DURING THE RESUSCITATION. THE CAUSE OF DEATH WAS POSSIBLY DUE TO AN ANNULAR RUPTURE OR LVOT RUPTURE BUT REMAINED UNKNOWN. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE CAUSE OF THE RUPTURE CANNOT BE CONFIRMED; HOWEVER, MAY BE RELATED TO PATIENT FACTORS AND/OR THE MECHANISM DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21381843&pc=NPT | 2025587-2025-01240 | 2024-05-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-13 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Chest Pain; Fever; Hematoma; Hemoptysis; Pseudoaneurysm | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A PATIENT WHO UNDERWENT SUCCESSFUL TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) OF A MEDTRONIC 26-MM EVOLUT PRO BIOPROSTHETIC VALVE. A FOLLOW-UP COMPUTED TOMOGRAPHY AT 12 MONTHS AFTER TAVI HAD NO ABNORMAL FINDINGS. APPROXIMATELY TWO YEARS POST-TAVI, THE PATIENT PRESENTED WITH CHEST DISCOMFORT. A COMPUTED TOMOGRAPHY REVEALED TWO PSEUDOANEURYSMS IN THE ASCENDING AORTA AND A HEMATOMA IN THE MEDIASTINUM. A FINAL BLOOD CULTURE REPORT HAD NEGATIVE RESULTS. ECHOCARDIOGRAPHY DID NOT REVEAL ANY VEGETATIONS, AND THE MODIFIED DUKE CRITERIA WERE NEGATIVE FOR ENDOCARDITIS. AFTER ADMISSION, THE PATIENT HAD SEVERAL FEVER SPIKES AND MILD HEMOPTYSIS. AS THE PATIENT WAS DEEMED TOO FRAIL FOR SURGICAL INTERVENTION, A VASCULAR PLUG CLOSURE PROCEDURE WAS CHOSEN BY THE MEDICAL TEAM. IN THE PROCEDURE, TWO NON-MEDTRONIC VASCULAR PLUG DEVICES WERE IMPLANTED TO COVER THE ORIFICE OF BOTH PSEUDOANEURYSMS. FINAL AORTOGRAPHY SHOWED ALMOST COMPLETE DISAPPEARANCE OF BLOOD FLOW INTO THE PSEUDOANEURYSMS. A COMPUTED TOMOGRAPHY CONFIRMED THAT BOTH VASCULAR PLUG WERE FIRMLY IN PLACE WITHOUT MIGRATION. AFTER THE PROCEDURE THE PATIENT HAD IMPROVEMENT IN HIS CHEST DISCOMFORT, ALTHOUGH HE CONTINUED TO HAVE MILD HEMOPTYSIS. HOWEVER, ON POSTOPERATIVE DAY 4, THE PATIENT EXPERIENCED SUDDEN MASSIVE HEMOPTYSIS AND DIED. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: KATADA ET AL. PLUG CLOSURE FOR ASCENDING AORTIC PSEUDOANEURYSMS AFTER TAVI USING SELF-EXPANDABLE VALVE. ANN THORAC SURG SHORT REP. 2024 MAY 23;2(4):776-778. DOI: 10.1016/J.ATSSR.2024.04.026. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21362597&pc=MVK | 3008642652-2025-01524 | 2025-01-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-12 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 09:12:07 ON (B)(6) 2025. AT 03:05:05 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VF. CPR/ELECTRODE LEAD FALL OFF AND HEART RATE NOT IN DETECTION SEQUENCE LONG ENOUGH PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 03:05:15, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 03:09:22 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21362938&pc=GAM | 1219930-2025-00707 | 2024-12-05 05:00 | Death | COVIDIEN LP LLC NORTH HAVEN | 2025-02-12 05:00 | GAM | UNKNOWN VLOC PRODUCT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE ANALYSIS OF 435 PATIENTS WHO UNDERWENT ROBOTIC PANCREATICODUODENECTOMY (WHIPPLE) BETWEEN 2012 AND 2024 WAS COMPLETED. V-LOC SUTURE WAS USED FOR HEPATICOJEJUNOSTOMY, PANCREATICOJEJUNOSTOMY, AND DUODENOJEJUNOSTOMY REANASTOMOSIS. MORTALITIES WITHIN 90 DAYS WERE MENTIONED IN 37 CASES WITH REASONS UNSPECIFIED; 27 OF THESE MORTALITIES OCCURRED WITHIN 30 DAYS. AN INSTITUTIONAL ANALYSIS OF HOSPITAL READMISSION FOLLOWING A ROBOTIC PANCREATICODUODENECTOMY DOI.ORG/10.1007/S11701-024-02186-0. Manufacturer Narrative: D10. CONCOMITANT PRODUCTS: UNKNOWN-VLOC UNKNOWN VLOC PRODUCT (LOT UNKNOWN); UNKNOWN-VLOC UNKNOWN VLOC PRODUCT (LOT UNKNOWN) DENIS GRATSIANSKIY, DHARTI PATEL, ISWANTO SUCANDY, TARA M. PATTILACHAN, MARIA CHRISTODOULOU, ALEXANDER ROSEMURGY, SHARONA B. ROSS, "AN INSTITUTIONAL ANALYSIS OF HOSPITAL READMISSION FOLLOWING A ROBOTIC PANCREATICODUODENECTOMY", JOURNAL OF ROBOTIC SURGERY, (2025) 19:20, HTTPS://DOI.ORG/10.1007/S11701-024-02186-0. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21362975&pc=NVZ | 3008973940-2025-01024 |  | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-02-12 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. Manufacturer Narrative: B3: EVENT DATE IS NOT KNOWN. PLEASE SEE B5 FOR APPROXIMATE DATE RANGE, IF APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21363358&pc=DSY | 9612515-2025-00013 | 2022-04-21 04:00 | Death | VASCUTEK LTD. | 2025-02-12 05:00 | DSY | GELSOFT | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | K955230 |  | 1 | Event Description: EVENT DATE: 21 APR 2022. IMPLANT DATE: (B)(6) 2020. EVENT REPORTED FROM PANTHER POST MARKET STUDY ((B)(6)). GELSOFT PLUS STRAIGHT IMPLANTED IN A MALE PATIENT (FOR ABOVE KNEE FEMOROPOPLITEAL) WITH PRE-CONDITION REPORTED AS CORONARY ARTERY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, HYPERTENSION (CONTROLLED BY MEDICATION), HYPERLIPIDEMIA (CONTROLLED BY MEDICATION), TYPE I DIABETES (CONTROLLED BY MEDICATION), RENAL INSUFFICIENCY, PERIPHERAL ARTERIAL OCCLUSIVE DISEASE ("PAD" - INTERPRETED AS PERIPHERAL ARTERIAL DISEASE), ASA (PHYSICAL STATUS CLASSIFICATION SYSTEM), CLASS III, CURRENT SMOKER. INDICATION: OCCLUSIVE DISEASE, TYPE: DISABLING INTERMITTENT CLAUDICATION. THE STUDY SITE REPORTED THE EVENT OF DEATH. CAUSE IS UNKNOWN. THERE WAS NO AUTOPSY. DATE OF DEATH WAS REPORTED AS 2(B)(6)2022. NO DEVICE DEFICIENCY REPORTED BY THE SITE IN THE STUDY DATABASE.Event Description: EVENT DATE: (B)(6) 2022. IMPLANT DATE: (B)(6) 2020. EVENT REPORTED FROM PANTHER POST MARKET STUDY (B)(4). GELSOFT PLUS STRAIGHT IMPLANTED IN A MALE PATIENT (FOR ABOVE KNEE FEMOROPOPLITEAL) WITH PRE-CONDITION REPORTED AS CORONARY ARTERY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, HYPERTENSION (CONTROLLED BY MEDICATION), HYPERLIPIDEMIA (CONTROLLED BY MEDICATION), TYPE I DIABETES (CONTROLLED BY MEDICATION), RENAL INSUFFICIENCY, PERIPHERAL ARTERIAL OCCLUSIVE DISEASE ("PAD" - INTERPRETED AS PERIPHERAL ARTERIAL DISEASE), ASA (PHYSICAL STATUS CLASSIFICATION SYSTEM), CLASS III, CURRENT SMOKER. INDICATION: OCCLUSIVE DISEASE, TYPE: DISABLING INTERMITTENT CLAUDICATION. THE STUDY SITE REPORTED THE EVENT OF DEATH. CAUSE IS UNKNOWN. THERE WAS NO AUTOPSY. DATE OF DEATH WAS REPORTED AS (B)(6) 2022. NO DEVICE DEFICIENCY REPORTED BY THE SITE IN THE STUDY DATABASE. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP#1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00013 TO PROVIDE EVENT CLOSURE INFORMATION FOR TAG0152. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4580 - INSUFFICIENT INFORMATION - EVENT OF PATIENT DEATH REPORTED WITH NO CAUSE OF DEATH AND NO DEVICE DEFICIENCY. IMPACT CODE: 1802 - DEATH - REPORTED AS DEATH EVENT. DEVICE PROBLEM CODE: 3190 - INSUFFICIENT INFORMATION - NO CAUSE OF DEATH AND NO DEVICE FAILURE /DEFICIENCY HAS BEEN REPORTED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW SHOWED THIS IS THE ONLY EVENT OF GELSOFT PLUS / MEDICAL EVENT / DEATH ((B)(6)2021 - (B)(6) 2025) 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW WILL BE PERFORMED. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION-INVESTIGATION IS ONGOING AND RESULTS ARE NOT YET AVAILABLE. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE -A CONCLUSION HAS YET TO BE ESTABLISHED AS THE INVESTIGATION IS INCOMPLETE. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 4580 - INSUFFICIENT INFORMATION - EVENT OF PATIENT DEATH REPORTED WITH NO CAUSE OF DEATH AND NO DEVICE DEFICIENCY. IMPACT CODE: 1802 - DEATH - REPORTED AS DEATH EVENT NO AUTOPSY WAS PERFORMED. DEVICE PROBLEM CODE : 2993 -ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM- NO CAUSE OF DEATH AND NO DEVICE FAILURE /DEFICIENCY HAS BEEN REPORTED NO AUTOPSY WAS PERFORMED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW COULD NOT BE PERFORMED AS THIS IS THE ONLY EVENT OF GELSOFT /MEDICAL EVENT / DEATH (UNKNOWN CAUSE NO DEVICE DEFICIENCY REPORTED (B)(6) 2021 - (B)(6)2025) 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION WAS RECEIVED STATING THE PATIENT SUFFERED POST-OPERATIVE COMPLICATIONS OF ACUTE PULMONARY EDEMA (OAP) DUE TO OVERLOAD. GOOD PROGRESS AND WEANING OFF OXYGEN AFTER DAILY DIALYSIS FOR 3 DAYS + PROBABILISTIC ANTIBIOTICS DUE TO PROBABLE SUPERINFECTION. ON (B)(6)2021 PATIENT HAD A FOLLOW-UP AWAITING KIDNEY TRANSPLANT. GOOD PERMEABILITY ON THE RIGHT BUT STENOSES ON THE LEFT. HYPERIMMUNIZATION IN (B)(6) 2021, CLASS 1, TGI 73%. - CONTRAINDICATION MAINTAINED FOR PERIOPERATIVE BYPASS. THE SITE HAVE CONFIRMED THEY HAVE NO MORE INFORMATION. THEY WERE AWARE OF THE DEATH VIA A DATABASE IN FRANCE, WHERE DATE OF DEATH IS WRITTEN. AS IT DID NOT HAPPEN AT THE HOSPITAL, THEY ARE NOT ABLE TO RETRIEVE FURTHER INFORMATION, WHICH IS WHY THEY SELECTED "UNKNOWN". NO AUTOPSY WAS PERFORMED. ADDITIONAL INFORMATION ON THE POSSIBLE INFECTION AND IS SWABS WERE TAKEN WAS REQUESTED ON (B)(6)2025 SITE RESPONDED ON (B)(6) - THE SITE HAVE RESPONDED TO THE QUESTION ABOUT THE INFECTION: "ACUTE RESPIRATORY DISTRESS DUE TO FLUID OVERLOAD IN THE CONTEXT OF DELAYED DIALYSIS, WITH A POSSIBLE INFECTIOUS COMPONENT UNDER INVESTIGATION. HE ALSO RECEIVED PROBABILISTIC ANTIBIOTIC THERAPY FOR A PROBABLE PULMONARY SUPERINFECTION WITH TAZOCILLINE AND ROVAMYCINE. NO CLEAR INFECTION IDENTIFIED; THE TREATMENT IS PROBABILISTIC. A BACTERIOLOGICAL SAMPLE WAS TAKEN, BUT NO RESULTS ARE REFERENCED IN THE MEDICAL RECORD. THE COVID PCR TEST IS NEGATIVE. SCANS WERE RECEIVED AND REVIEWED FOR THIS EVENT , HOWEVER THE IMAGING PROVIDED ONLY COVERS THE CHEST AND THE GRAFT WAS USED IN THE LEG. ADDITIONAL SCANS WERE REQUESTED AND UPON RECEIPT THESE WERE FOUND TO BE ULTRASOUND SCANS SO NO REVIEW WAS POSSIBLE. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW WAS PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT WHICH CONFIRMED THE GRAFT WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 213-NO DEVICE PROBLEM FOUND - NO DEVICE DEFICIENCY OR CAUSE OF DEATH WAS REPORTED AND FULL BATCH REVIEW SHOWED GRAFT WAS MANUFACTURED TO SPECIFICATION AS SCANS WERE ULTRASOUND NO REVIEW OF THESE COULD BE PERFORMED. INVESTIGATION CONCLUSION: 67 - NO PROBLEM DETECTED, IT WAS NOT CONFIRMED THERE WAS A PROBLEM WITH THE DEVICE -NO CAUSAL LINK BETWEEN THE EVENT AND THE DEVICE COULD BE ESTABLISHED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21363804&pc=LWS | 2649622-2025-04083 | 2025-01-24 05:00 | Death | MPRI | 2025-02-12 05:00 | LWS | SPRINT | Under-Sensing; Capturing Problem | Cardiac Arrest; Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THERE WAS AN ELECTROCARDIOGRAM (ECG) WAVEFORM RECORD THAT APPEARED TO SHOW AN IMPLANT ABLE CARDIOVERTER DEFIBRILLATOR (ICD) PACING FAILURE. FOURTEEN MINUTES LATER, THE ECG APPEARED TO SHOW VENTRICULAR FIBRILLATION (VF); HOWEVER, THE ICD DID NOT DELIVER A SHOCK. THE PATIENT RETURNED TO SINUS RHYTHM, BUT THEY SUBSEQUENTLY DIED OF CARDIAC ARREST. WHEN THE ICD WAS CHECKED, THERE WERE NO EPISODES RECORDED. IT WAS INDICATED THERE HAD BEEN A RIGHT VENTRICULAR (RV) LEAD THRESHOLD INCREASE AND SUSPECTED UNDERSENSING. IT WAS NOTED THAT IT WAS UNKNOWN WHETHER THERE WAS A PROBLEM WITH THE MYOCARDIUM OR THE ICD. Manufacturer Narrative: CONTINUATION OF D10: 6940-52 LEAD IMPLANTED: (B)(6) 2004. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS RISING. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS UNSTABLE. ANALYSIS OF THE DEVICE MEMORY INDICATED DIMINISHED RIGHT VENTRICULAR SENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21363962&pc=HFA | MW5165994 | 2023-12-15 05:00 | Death | KARL STORZ SE & CO. KG. | 2025-02-12 05:00 | HFA | FETOSCOPY SET |  | Hematoma; Hemorrhage/Bleeding; Necrosis; Uterine Perforation; Bronchial Hemorrhage |  |  | 1 | Event Description: A 35 YO G2P1001 AT 19W4D WITH MO-DI (MONOCHORIONIC DIAMNIOTIC) TWINS COMPLICATED BY TWIN-TO-TWIN TRANSFUSION SYNDROME, TTTS STAGE III AND FIRST TRIMESTER BLEEDING WHICH LED TO A LARGE SUBCHORIONIC HEMATOMA. US FINDINGS FOR TWIN A, THE RECIPIENT HAD REVERSED AWAVE OF THE DV (DUCTUS VENOSUS) AND THETWIN B, THE DONOR HAD INTERMITTENT ABSENT EDF OF THE UA (END-DIASTOLIC FLOW OF THE UMBILICAL ARTERY.). THE GROWTH DISCORDANCE WAS 18%. SHE UNDERWENT FETOSCOPIC LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS FOR TTTS ON (B)(6) 2023. THE RISK OF HER SURGERY WAS DISCUSSED IN DETAIL BUT THAT INCLUDED FURTHER MEMBRANE SEPARATION THAT COULD LEAD TO PREMATURE PRETERM RUPTURE OF MEMBRANES AFTER THE SURGERY. SHE WAS ALSO AT RISK FOR PLACENTAL ABRUPTION. AT THE SAME TIME IF SHE DID NOT PERFORM SURGERY, SHE WAS AT RISK LOSING ONE OR BOTH TWINS. THE SURGERY ITSELF WAS UNEVENTFUL. ONE POD#1(POST-OPERATIVE DAY ONE) SHE WAS FOUND TO HAVE SIGNIFICANT AMOUNT OF CHORIOAMNIOTIC SEPARATION. SHE WAS NOT LEAKING BUT LATER THAT AFTERNOON SHE RUPTURED HER MEMBRANES. SHE REMAINED STABLE AND THE TWINS RECOVERED FROM SURGERY WITH CONCORDANT FLUID, BLADDER SEEN AND NORMAL DOPPLERS. SHE RESEALED AND ACCUMULATED FLUID IN BOTH TWINS. SHE HAD LATENCY ANTIBIOTICS. HOWEVER, AT 22W1D, THE DAYS PRIOR SHE HAD SEVERE PREECLAMPSIA AND UNDERWENT C-SECTION FOR DELIVERY OF HER TWINS ON (B)(6) 2023 AT 22W1D. TWIN B, FORMER DONOR WAS GROWTH RESTRICTED BORN AT 0.4KG. ON DOL #11 (DAY OF LIFE 11) HE DEVELOPED NECROTIZING ENTEROCOLITIS WITH PNEUMATOSIS REQUIRING RAPID ESCALATION OF MANAGEMENT OVER THE COURSE OF THE DAY. THIS INCLUDED ADMINISTRATION OF MULTIPLE BLOOD PRODUCTS, VASOPRESSOR SUPPORT, ANTIBIOTICS, VENTILATION TITRATION. HE DEVELOPED SEVERE PULMONARY HEMORRHAGE THAT EVENING RESULTING IN A CODE EVENT. THE NICU TEAM ALONG WITH PARENTS MADE THE JOINT DECISION TO COMPASSIONATELY EXTUBATE. HE PASSED ON DOL#12(DAY OF LIFE 12). REFERENCE REPORT: MW5165995.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21363970&pc=HFA | MW5165995 | 2023-12-15 05:00 | Death | KARL STORZ SE & CO. KG. | 2025-02-12 05:00 | HFA | FETOSCOPY SET |  | Cardiac Arrest; Hematoma; Necrosis; Uterine Perforation; Rupture; Bronchial Hemorrhage |  |  | 1 | Event Description: A 35 YO G2P1001 AT 19W4D WITH MO-DI (MONOCHORIONIC DIAMNIOTIC) TWINS COMPLICATED BY TWIN-TO-TWIN TRANSFUSION SYNDROME, TTTS STAGE III AND FIRST TRIMESTER BLEEDING WHICH LED TO A LARGE SUBCHORIONIC HEMATOMA. US FINDINGS FOR TWIN A, THE RECIPIENT HAD REVERSED AWAVE OF THE DV(DUCTUS VENOSUS) AND THE TWIN B, THE DONOR HAD INTERMITTENT ABSENT EDF OF THE UA (END-DIASTOLIC FLOW OF THE UMBILICAL ARTERY.). THE GROWTH DISCORDANCE WAS 18%. SHE UNDERWENT FETOSCOPIC LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS FOR TTTS ON (B)(6) 2023. THE RISK OF HER SURGERY WAS DISCUSSED IN DETAIL BUT THAT INCLUDED FURTHER MEMBRANE SEPARATION THAT COULD LEAD TO PREMATURE PRETERM RUPTURE OF MEMBRANES AFTER THE SURGERY. SHE WAS ALSO AT RISK FOR PLACENTAL ABRUPTION. AT THE SAME TIME IF SHE DID NOT PERFORM SURGERY, SHE WAS AT RISK LOSING ONE OR BOTH TWINS. THE SURGERY ITSELF WAS UNEVENTFUL. ONE POD#1(POST-OPERATIVE DAY ONE) SHE WAS FOUND TO HAVE SIGNIFICANT AMOUNT OF CHORIOAMNIOTIC SEPARATION. SHE WAS NOT LEAKING BUT LATER THAT AFTERNOON SHE RUPTURED HER MEMBRANES. SHE REMAINED STABLE AND THE TWINS RECOVERED FROM SURGERY WITH CONCORDANT FLUID, BLADDER SEEN AND NORMAL DOPPLERS. SHE RESEALED AND ACCUMULATED FLUID IN BOTH TWINS. SHE HAD LATENCY ANTIBIOTICS. HOWEVER, AT 22W1D, THE DAYS PRIOR SHE HAD SEVERE PREECLAMPSIA AND UNDERWENT C-SECTION FOR DELIVERY OF HER TWINS ON (B)(6) 2023 AT 22W1D. TWIN B, FORMER DONOR WAS GROWTH RESTRICTED BORN AT 0.4KG. ON DOL #11(DAY OF LIFE 11) HE DEVELOPED NECROTIZING ENTEROCOLITIS WITH PNEUMATOSIS REQUIRING RAPID ESCALATION OF MANAGEMENT OVER THE COURSE OF THE DAY. THIS INCLUDED ADMINISTRATION OF MULTIPLE BLOOD PRODUCTS, VASOPRESSOR SUPPORT, ANTIBIOTICS, VENTILATION TITRATION. HE DEVELOPED SEVERE PULMONARY HEMORRHAGE THAT EVENING RESULTING IN A CODE EVENT. THE NICU TEAM ALONG WITH PARENTS MADE THE JOINT DECISION TO COMPASSIONATELY EXTUBATE. HE PASSED ON DOL#12(DAY OF LIFE 12). REFERENCE REPORT: MW5165994.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21364232&pc=LZS | MW5165996 | 2024-08-15 04:00 | Death | UNK | 2025-02-12 05:00 | LZS | LASIK LASER | Use of Device Problem; Patient-Device Incompatibility; Therapeutic or Diagnostic Output Failure | Corneal Clouding/Hazing; Headache; Blurred Vision; Visual Impairment; Halo; Suicidal Ideation |  |  | 1 | Event Description: (B)(6) WENT FOR A ROUTINE LASIK PROCEDURE ON (B)(6) 2024. AFTER COMPLETION, (B)(6) DEVELOPED DOUBLE/BLURRY VISION, GHOSTING, HALOS, INTENSE MIGRAINE HEADACHES, AND SENSITIVITY TO LIGHT. ALTHOUGH THE PROCEDURE NEGATIVELY AFFECTED BOTH EYES, THE RIGHT ONE HAD SEVERE DEGRADATION. (B)(6) DESCRIBED IT AS LOOKING THROUGH FROSTED GLASS OR CELLOPHANE. AFTER COMMUNICATING WITH (B)(6), THEY WERE VERY DISMISSIVE AND (B)(6) WAS TOLD HE NEEDED TIME TO HEAL. THESE ADVERSE EVENTS PREVENTED (B)(6) FROM BEING ABLE TO FUNCTION IN HIS DAILY LIFE OR COMPLETE HIS DUTIES AS A (B)(6) POLICE OFFICER. UNABLE TO WORK, HE SOUGHT SHORT TERM DISABILITY. AS MONTHS PAST, (B)(6) SAW NO IMPROVEMENT AND EVENTUALLY FILED FOR LONG TERM DISABILITY WHICH WAS DENIED BY (B)(6) ON (B)(6) 2025. (B)(6) SOUGHT HELP FROM (B)(6) BUT NO SOLUTIONS CAME. THEY OFFERED TO DO AN ENHANCEMENT WHICH IS NOT FDA APPROVED AND THEN OFFERED HIM A REFUND WHICH WAS NOT ACCEPTABLE AS A FIX. HE POSTED A REVIEW ABOUT HIS NEGATIVE EXPERIENCE AND (B)(6) FIRED HIM AS A PATIENT. (B)(6) WENT TO SEE A CORNEAL SPECIALIST AT (B)(6). THE PHYSICIANS EVALUATION CONFIRMED (B)(6) WAS NOT A CANDIDATE FOR LASIK AND THE PROCEDURE SHOULD HAVE NEVER BEEN PERFORMED. (B)(6) CORNEAS WERE TOO THIN. THE (B)(6) DOCTOR ALSO REPORTED TOO MUCH TISSUE WAS ALTERED AND THE LASER CREATED HOA'S. (B)(6) WAS LOVED BY ALL THAT HE CAME IN CONTACT WITH HIM. AFTER DOING MONTHS OF RESEARCH, (B)(6) FOUND MANY PEOPLE SUFFERING FROM THE SAME NEGATIVE LIFE ALTERING EFFECTS FROM LASIK. ALTHOUGH, HE FOUND A COMMUNITY, (B)(6) SAW THAT NO ONE HAD ANY IMPROVEMENT OR RESOLUTION TO THESE DEBILITATING SIDE EFFECTS. ON (B)(6) 2025, (B)(6) TOOK HIS OWN LIFE DUE TO GROSS NEGLIGENCE AND BECAUSE OF A PRODUCT THAT SHOULD HAVE NEVER COME TO MARKET. THE ORIGINAL DATA THAT THE FDA APPROVED NEEDS A MAJOR UPDATE WARNING THE PUBLIC OF THE SEVERE CONSEQUENCES AND DAMAGES CAUSE BY LASERS AND THE LASIK PROCEDURE. (B)(6) TOUTS AN IMPRESSIVE SUCCESS RATE BY HIDING NEGATIVE REVIEWS AND NOT COMMUNICATING PROPER RISKS TO PATIENTS. IF A PROCEDURE PRODUCES A NEGATIVE OUTCOME, LASIK DOCTORS GASLIGHT THE PATIENT AND TELL THEM THEY'VE NEVER SEEN THIS BEFORE. HOWEVER, I'VE SPOKEN A MULTITUDE OF PATIENTS NOW SOME THAT HAVE GONE TO THE SAME DOCTORS AND HAD THE SAME NEGATIVE OUTCOMES. IT IS JUST ANOTHER BLACK STAIN ON THE MEDICAL ESTABLISHMENT. IT IS TIME FOR THE FDA TO DO THE RIGHT THING WHICH SHOULD ALSO EXCLUDE THE OPINIONS OF THE MONSTERS PERFORMING THIS PROCEDURE. WHEN THE ACTUAL DATA SHOWS 60% OF PATIENTS SUFFER FROM FLUCTUATING VISION IN DIM LIGHT, 50% WITH GRITTY EYES, 40% LIGHT SENSITIVITY, 20% WITH FLUCTUATING VISION IN BRIGHT AND DIM LIGHT AND 15% SUFFERING FROM DEGRADED NIGHT DRIVING BECAUSE OF GLARE AND HALOS. I ASK THE FDA TO STOP TURNING A BLIND EYE ON LASIK BECAUSE WITHOUT VISION...THERE IS NO FUTURE. (B)(6) ALSO COMMENTED ON THE 2022 FDA LASIK GUIDANCE DOCUMENT 3 MONTHS AFTER HIS SURGERY. HERE IS THE LINK: HTTPS://WWW.REGULATIONS.GOV/COMMENT/FDA-2022-D-1253-0786. Event Description: ADDITIONAL INFORMATION RECEIVED ON 2/12/2024 FOR REPORT MW5165996. AGAIN, IT'S TIME TO DO THE RIGHT THING. I'VE ATTACHED (B)(6) COMMENT ON THE 2022 FDA GUIDANCE LASIK DOCUMENT. THIS DOCUMENT COULD HAVE SAVED HIS LIFE BUT IT'S HANGING IN SPACE WHILE THE BUTCHERING OF HEALTHY EYES CONTINUES. (B)(6). I CALL ON YOU TO PROTECT PATIENTS AND NOT THE INDUSTRY AND THEIR PROFITS. THANK YOU, (B)(6).Event Description: ADDITIONAL INFORMATION RECEIVED ON 2/25/2025 FOR REPORT MW5165996 TO UPDATE EMAIL ADDRESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21364271&pc=DTB | 2649622-2025-04104 | 2025-01-24 05:00 | Death | MPRI | 2025-02-12 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pericardial Effusion | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF THE PACING LEAD, PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY DURING THE PROCEDURE. THE LEAD WAS REMOVED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED TO REGAIN PULSE. TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS DONE AND BLOOD WAS OBSERVED AROUND PERICARDIUM. POSSIBLE SMALL PERICARDIAL EFFUSION WAS SUSPECTED. PATIENT HAD DROPPING BLOOD PRESSURE IN THE INTENSIVE CARE UNIT (ICU) AND WAS EXPIRED POST-OPERATIVELY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21364637&pc=DSQ | 2916596-2025-00901 | 2023-05-01 04:00 | Death | THORATEC CORPORATION | 2025-02-12 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "EARLY POSTOPERATIVE CHANGES IN VON WILLEBRAND FACTOR ACTIVITY ARE ASSOCIATED WITH FUTURE BLEEDING AND STROKE IN HEARTMATE 3 PATIENTS" THAT HEARTMATE 3 MAY BE ASSOCIATED WITH GASTROINTESTINAL BLEEDING, HEMORRHAGIC STROKE, ISCHEMIC STROKE, AND DEATH. A PROSPECTIVE SINGLE-CENTER STUDY INCLUDING 74 HEARTMATE 3 DEVICE RECIPIENTS BETWEEN 2016 AND 2023 WAS UNDERTAKEN. A TOTAL OF 113 PATIENTS UNDERWENT HEARTMATE 3 LVAD IMPLANTATION BETWEEN SEPTEMBER 2016 AND FEBRUARY 2023. SEVENTY-FOUR (97.4%) PATIENTS SURVIVED FOR AT LEAST 60 DAYS AND WERE THEREFORE INCLUDED. CLINICAL FOLLOW-UP WAS PERFORMED UNTIL MAY 2023. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01MAY2023 AS CLINICAL FOLLOW-UP WAS PERFORMED UNTIL MAY 2023. JAHANGIRI, P., VEEN, K. M., VAN MOORT, I., BUNGE, J. H., CONSTANTINESCU, A., SJATSKIG, J., DE MAAT, M., KLUIN, J., LEEBEEK, F., & CALISKAN, K. (2024). EARLY POSTOPERATIVE CHANGES IN VON WILLEBRAND FACTOR ACTIVITY ARE ASSOCIATED WITH FUTURE BLEEDING AND STROKE IN HEARTMATE 3 PATIENTS. ASAIO JOURNAL. HTTPS://DOI.ORG/10.1097/MAT.0000000000002250. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21364829&pc=NIK | 2017865-2025-13943 | 2025-01-17 05:00 | Death | ABBOTT | 2025-02-12 05:00 | NIK | UNIFY ASSURA ICD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY. THE CAUSE OF THE PATIENT'S DEATH AND ANY RELATION TO THE ABBOTT DEVICE REMAINS UNKNOWN AT THIS TIME. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21365884&pc=NVN | 2182208-2025-00686 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-02-12 05:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEFT BUNDLE BRANCH AREA PACING. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE MORTALITY. THERE IS NO ALLEGATION OF LEAD-DEATH RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND LEAD-RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS AND TRANSIENT RIGHT BUNDLE BRANCH BLOCK. THERE WERE LEADS WHICH EXHIBITED INCREASES IN THRESHOLDS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS FEMALE/81 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: COMPARISON OF CLINICAL AND ECHOCARDIOGRAPHIC OUTCOMES BETWEEN LEFT BUNDLE BRANCH AREA PACING AND RIGHT VENTRICULAR PACING IN OLDER PATIENTS. PACING AND CLINICAL ELECTROPHYSIOLOGY. 2024; 47:1346¿1357. DOI: 10.1111/PACE.15056. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21366126&pc=FNL | 1824206-2025-00281 | 2025-01-08 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-02-12 05:00 | FNL | CENTRELLA BED | Failure to Sense | Fall; Cognitive Changes; Ptosis | NA |  | 1 | Event Description: IT WAS REPORTED BY THE CUSTOMER A PATIENT FELL AFTER EXITING A CENTRELLA BED WHEN THE BED ALARM FAILED TO TRIGGER, AND THE PATIENT SUBSEQUENTLY DIED. ON AN UNKNOWN DATE THE PATIENT WAS HOSPITALIZED FOR A LOWER GI BLEED. ON THE DAY OF THE EVENT (REPORTED AS ¿EARLIER IN THE DAY¿), THE PATIENT EXITED THE CENTRELLA BED AN UNKNOWN NUMBER OF TIMES, AND THE ALARM TRIGGERED EACH TIME. LATER THAT DAY, THE PATIENT EXITED THE BED FOR AN UNKNOWN CAUSE AND THE PATIENT FELL. THE BED ALARM WAS NOTED TO HAVE BEEN SET; HOWEVER, THIS TIME THE ALARM DID NOT TRIGGER WHEN THE PATIENT EXITED THE BED. THE PATIENT WAS ASSESSED FOR ANY VISIBLE INJURIES, AND NONE WERE NOTED. AFTER ¿APPROXIMATELY 2-3 HOURS AFTER THE FALL¿, THE PATIENT ¿SHOWED COGNITIVE DECLINE AND RIGHT SIDED DROOP¿. PER HOSPITAL PROTOCOL, VITAL SIGNS WERE TAKEN, A HEAD-TO-TOE ASSESSMENT WAS REDONE, AND A HEAD COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED. THE CT WAS FOUND TO BE NEGATIVE. THE PATIENT WAS DESCRIBED BY THE NURSING STAFF AS ¿FINE¿; HOWEVER, AFTER AN UNSPECIFIED AMOUNT OF TIME, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT REPORTED; NOR, WAS IT REPORTED IF AN AUTOPSY WAS PERFORMED. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21366135&pc=DSQ | 2916596-2025-00843 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-12 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Hemorrhage/Bleeding; Unspecified Infection; Cardiogenic Shock; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿TIME-VARYING ANALYSES OF SURVIVAL AND OUTCOMES IN PATIENTS WITH HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICES¿ IDENTIFYING THAT HM3 MAY BE RELATED TO NEUROLOGICAL EVENTS, GASTROINTESTINAL BLEEDING, NON-GASTROINTESTINAL BLEEDING, HEMOLYSIS, RIGHT HEART FAILURE, VENTRICULAR ARRHYTHMIA, INFECTION, PUMP EXCHANGE OR EXPLANT DUE MALFUNCTIONS, TRANSPLANT, AND DEATH. A TOTAL OF 431 HM3 PATIENTS IMPLANTED FROM OCTOBER 2015 TO JANUARY 2023 WERE EVALUATED. SURVIVAL WAS ESTIMATED TO 5 YEARS POST-IMPLANT. TIME-VARYING RISKS OF DEATH, NEUROLOGICAL EVENTS, GASTROINTESTINAL BLEEDING, DEVICE-RELATED INFECTIONS, AND OTHER ADVERSE EVENTS WERE CHARACTERIZED USING MULTIPHASE HAZARD MODELLING. SURVIVAL ON HEARTMATE 3 AT 1 AND 5¿YEARS WAS 88% AND 58%, RESPECTIVELY. CUMULATIVE NUMBER OF NEUROLOGICAL EVENTS/PATIENT AT 1 YEAR AND 5 YEARS WAS 0.13 AND 0.29, RESPECTIVELY, AND CUMULATIVE NUMBER OF GASTROINTESTINAL BLEEDING EVENTS/PATIENT AT 1 YEAR AND 5 YEARS WAS 0.32 AND 0.78, RESPECTIVELY. DEVICE-RELATED INFECTIONS DEVELOPED IN 136 PATIENTS. ONE- AND 5-YEAR FREEDOM FROM DEVICE-RELATED INFECTION WAS 77% AND 45%, RESPECTIVELY. IN TOTAL, 205 RHF EVENTS OCCURRED IN 128 PATIENTS. CUMULATIVE INCIDENCE PER PATIENT AT 1 MONTH, 1 YEAR AND 5 YEARS WAS 0.12, 0.37 AND 0.85, RESPECTIVELY. THERE WERE 157 SVA EVENTS IN 106 PATIENTS. CUMULATIVE INCIDENCE PER PATIENT AT 1 MONTH, 1 YEAR AND 5 YEARS IS 0.12, 0.24 AND 0.71, RESPECTIVELY. FOUR PATIENTS UNDERWENT HM3 TO HM3 EXCHANGE AND WERE CENSORED AT PUMP EXCHANGE WITH EACH DEVICE AND POSTOPERATIVE COURSE ANALYZED SEPARATELY. AT LAST FOLLOW-UP, 213 REMAINED ON HM3 SUPPORT, 67 HAD UNDERGONE HEART TRANSPLANTATION, 9 HAD UNDERGONE HM3 EXPLANT OR EXCHANGE, 41 HAD TRANSFERRED CARE TO ANOTHER INSTITUTION, AND 101 HAD DIED. SURVIVAL AT 1, 3 AND 5 YEARS WAS 88%, 72% AND 58%, RESPECTIVELY. IN TOTAL, 13 MAJOR DEVICE MALFUNCTIONS OCCURRED IN 13 PATIENTS. THIS INCLUDED SEVEN DEVICE THROMBOSES, WHICH REQUIRED HM3 EXCHANGE IN THREE AND ANTICOAGULATION ALONE IN ONE; THE OTHER THREE WERE ASYMPTOMATIC WITH DISCOVERY OF THROMBOSIS AT HM3 EXPLANT FOR HEART TRANSPLANT. THE REMAINING SIX DEVICE MALFUNCTIONS INCLUDED BATTERY FAILURE LEADING TO DEATH IN ONE, INFLOW MALPOSITION REQUIRING SURGICAL INTERVENTION IN FOUR, AND DRIVELINE DAMAGE IN ONE PATIENT. DEATH AND ALL CAUSES OF DEATH ARE DESCRIBED IN THE TABLE 3 OF SUPPLEMENTARY INFORMATION AND WILL BE ADDRESSED UNDER DEATH REGULATORY REPORT DECISION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2023 SINCE THE PATIENTS WERE IMPLANTED BETWEEN (B)(6) 2015 AND (B)(6) 2023. DEPARTMENT OF THORACIC AND CARDIOVASCULAR SURGERY, CLEVELAND CLINIC, CLEVELAND, OH, USA. AARON J. WEISS ET AL., PUBLISHED ONLINE ON 09 JANUARY 2025 EUROPEAN JOURNAL OF HEART FAILURE (2025). DOI:10.1002/EJHF.3577. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED "TIME-VARYING ANALYSES OF SURVIVAL AND OUTCOMES IN PATIENTS WITH HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICES" REPORTED A STUDY THAT AIMED TO CHARACTERIZE TEMPORAL VARIATIONS IN RISKS OF MORTALITY AND ADVERSE OUTCOMES IN PATIENTS WITH A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE STUDY INCLUDED A TOTAL OF 431 CONSECUTIVE HEARTMATE 3 DEVICES IMPLANTED IN 427 ADULTS AT CLEVELAND CLINIC MAIN CAMPUS FROM 15OCT2015 TO 01JAN2023. THE ARTICLE INDICATED THAT OF THESE PATIENTS, 101 HAD DIED. SURVIVAL AT 1, 3 AND 5 YEARS WAS 88%, 72% AND 58%, RESPECTIVELY. NON-CARDIOVASCULAR CAUSES OF DEATH INCLUDED NEUROLOGICAL DYSFUNCTION, MAJOR INFECTION, MULTI-SYSTEM ORGAN FAILURE, MAJOR BLEEDING, RESPIRATORY FAILURE, BOWEL ISCHEMIA, WITHDRAWAL OF SUPPORT, HEMATOLOGICAL, AND HEPATIC DYSFUNCTION. CARDIOVASCULAR CAUSES OF DEATH INCLUDED HEART FAILURE, CARDIAC ARRHYTHMIA, SUDDEN UNEXPLAINED DEATH, CARDIAC ARREST, AND CARDIOGENIC SHOCK. THE ARTICLE CONCLUDED THAT LONG-TERM SURVIVAL ON HEARTMATE 3 SUPPORT WAS FAVORABLE, ALIGNING WITH PREVIOUSLY PUBLISHED TRIAL AND NATIONAL REGISTRY DATA. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING NEUROLOGICAL EVENTS (NOT STROKE-RELATED), INFECTION (DRIVELINE, PUMP POCKET OR PSEUDO POCKET, AS WELL AS LOCALIZED), MULTIPLE ORGAN FAILURES/DYSFUNCTION (INCLUDING RESPIRATORY FAILURE AND HEPATIC DYSFUNCTION), BLEEDING, CARDIAC ARRHYTHMIA, AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. THIS SECTION ALSO LISTS INFECTION, ARRHYTHMIA, AND NEUROLOGICAL DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21366789&pc=OZP | 2032227-2024-238133 | 2024-08-15 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-12 05:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hyperglycemia; Hypoglycemia; Diabetic Ketoacidosis | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2024. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS HYPERGLYCEMIA, HYPOGLYCEMIA AND DIABETES KETOACIDOSIS. THE CUSTOMER¿S BLOOD GLUCOSE WAS 600 MG/DL. THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2024, PRIOR TO PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KPK, MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. MMT-1780KPK WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. THE INSULIN PUMP WILL BE RETURNED FOR PRODUCT ANALYSIS. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21367033&pc=OZD | 1220648-2025-26117 | 2025-01-23 05:00 | Death | ABIOMED, INC. | 2025-02-12 05:00 | OZD | IMPELLA CP | Failure to Advance | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A DIAGNOSED ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION PATIENT WAS INTUBATED IN THE EMERGENCY DEPARTMENT PRIOR TO TRANSPORTING TO THE CATHETERIZATION LAB FOR IMPELLA CP PLACEMENT. THE PATIENT CODED AFTER GAINING ACCESS AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. THE RETURN OF SPONTANEOUS CIRCULATION WAS GAINED BRIEFLY. THEN THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AND THEN CODED FOR 45 MINUTES WITH SUSTAINED PULSELESS ELECTRICAL ACTIVITY ARREST. THE IMPELLA CP WAS OPENED AND PRIMED BUT NEVER INSERTED. THE PATIENT EXPIRED AND THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED DELIVERY ISSUE AND VENTRICULAR TACHYCARDIA (VT) HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THE ROOT CAUSE OF DELIVERY ISSUE COULD NOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED, THE ROOT CAUSE OF THE VT WAS NOT DETERMINED. B.2 REVISED AS A SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26117. H.1 REVISED AS SELECTION WAS PREVIOUSLY INCORRECT ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26117.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21367278&pc=NIQ | 2024168-2025-01595 | 2019-08-12 04:00 | Death | ABBOTT VASCULAR | 2025-02-12 05:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED THAT BETWEEN (B)(6) 2017, AND (B)(6) 2019, 445 PATIENTS WERE RANDOMIZED AT 22 SITES IN CHINA FOR THE BIOHEART-II TRIAL AND DOCUMENTED IN THE ARTICLE TITLED, "A RANDOMIZED COMPARISON OF BIOHEART SIROLIMUS-ELUTING BIORESORBABLE SCAFFOLD AND EVEROLIMUS-ELUTING STENTS". THE TRIAL USED THE BIOHEART BIORESORBABLE SCAFFOLD (BRS), THE XIENCE STENT, THE DRAGONFLY OCT (OPTICAL COHERENCE TOMOGRAPHY) CATHETER, AND THE C7-XR OPTIS OCT SYSTEM. TWO PATIENTS FROM THE XIENCE ARM REQUIRED BAILOUT STENTING (DISSECTION). OUT OF THE 216 PATIENTS ON THE XIENCE ARM, CLINICAL RESULTS WERE CAPTURED AT 1 MONTH, 1 YEAR AND 3 YEARS. THERE WAS A LOW OCCURRENCE FOR ALL CAUSE DEATH, MYOCARDIAL INFARCTION, STENT THROMBOSIS, AND REVASCULARIZATION. LOW RATE OF RESTENOSIS WAS FOUND AT 12 MONTHS POST STENT. DEVICE/LESION SUCCESS WAS 100% ON THE XIENCE ARM. THERE WERE LOWER RATES OF STRUT MALPOSITION. PLEASE SEE THE ARTICLE FOR FULL DETAILS. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. B2 - DATE OF DEATH: ESTIMATED. B3 - DATE OF EVENT: ESTIMATED. D4 - PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: ESTIMATED. LITERATURE ATTACHMENT: ARTICLE TITLED; " A RANDOMIZED COMPARISON OF BIOHEART SIROLIMUS-ELUTING BIORESORBABLE SCAFFOLD AND EVEROLIMUS-ELUTING STENTS THE BIOHEART-II TRIAL".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21367542&pc=NPT | 2015691-2025-00955 | 2025-01-26 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-12 05:00 | NPT | EDWARDS LIFESCIENCES COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning | Vascular Dissection | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), THAT APPROXIMATELY 7 YEARS, 4 MONTHS AND 4 DAYS POST TAVR WITH A 23 MM SAPIEN 3 VALVE THE PATIENT DEVELOPED SEVERE STENOSIS. THE STENOSIS WAS TREATED WITH A VALVE IN VALVE USING A 20MM SAPIEN 3 ULTRA RESILIA VALVE. PER FOLLOW UP THE PATIENT UNDERWENT AN URGENT TAVR ON (B)(6) 2025, WHICH WAS COMPLICATED BY AN EXTENSIVE TYPE A DISSECTION. WHEN THE IMPLANTER WAS CROSSING THE EXISTING VALVE, WITH THE NEW VALVE, THE VALVE HAD SLIGHT TROUBLE CROSSING, GETTING HUNG UP ON THE PATIENTS EXISTING VALVE. THE ANGLE OF THE COMMANDER WIRE POSITION WAS CHANGED, AND THE VALVE SLID THROUGH. AFTER THE AORTOGRAM THERE WAS NO OBVIOUS DISSECTION. THEN THE PHYSICIAN ENGAGED THE RIGHT CORONARY. AT THE FINAL AORTOGRAM, A JET CONTRAST WAS NOTICED NEAR THE NON-SINUS. ON THE ECHO, THERE WAS NO PERICARDIAL EFFUSION, AND THE PATIENT'S BLOOD PRESSURE WAS STABLE. THE PATIENT WAS EXTUBATED AND SENT FOR A CT AFTER THE CASE. THE PATIENT DIED THAT EVENING. THE PROCEDURE WAS PLANNED FOR (B)(6) 2025, BUT THE PATIENT DEVELOPED NARROW COMPLEX VT IN THE 140'S ON (B)(6) 2025. TTE SHOWED THE AORTIC VALVE HAD TRACE REGURGITATION PRESENT, A PERCUTANEOUS STENTED BIO PROSTHESIS PRESENTS IN THE AORTIC POSITION, AND THE PEAK AND MEAN TRANSVALVULAR GRADIENTS ARE 105 MMHG AND 63 MMHG. THERE IS SEVERE PROSTHETIC VALVE STENOSIS. THE ACCELERATION TIME IS GREATER THAN 170MS AND DVI IS 0.19. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING AND CLINICAL EVALUATION FINDINGS. SECTIONS G6, H2, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND CONCLUSIONS IN THIS SECTION HAVE BEEN UPDATED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED TO EDWARDS FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE ENGINEERING SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT FOR DIFFICULTY TO CROSS NATIVE ANNULUS WITH SUBSEQUENT AORTIC INJURY AND DEATH WAS UNABLE TO BE CONFIRMED DUE TO NO RELEVANT DEVICE OR IMAGERY WAS PROVIDED FOR EVALUATION. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (SEVERE PROSTHETIC VALVE STENOSIS) LIKELY CONTRIBUTED TO THE DIFFICULTIES CROSSING THE ANNULUS. PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, SMALL VESSEL SIZE, PRE-EXISTING VASCULAR STENT) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. ADDITIONALLY, PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ASCENDING AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, IT IS POSSIBLE THAT AORTIC RUPTURE MAY HAVE OCCURRED DUE TO MULTIPLE ATTEMPTS TO CROSS THE VALVE AND/OR EXCESSIVE DEVICE MANIPULATION IN RESPONSE TO THE REPORTED CROSSING DIFFICULTY BROUGHT BY THE SEVERE PROSTHETIC VALVE STENOSIS. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THIS IS ONE OF 2 REPORTS BEING SUBMITTED FOR THIS CASE. THIS INVESTIGATION IS STILL ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21367633&pc=MKJ | 1220908-2025-00571 | 2025-01-22 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-12 05:00 | MKJ | R SERIES | Failure to Deliver Shock/Stimulation | Ventricular Fibrillation; Asystole | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 48-YEAR-OLD FEMALE PATIENT, THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE AND NEW SET OF ELECTRODE PADS TO CONTINUE TREATING THE PATIENT. HOWEVER, THEY WERE UNSUCCESSFUL AT DISCHARGING THE SELECTED ENERGY. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE AND DEFIB PADS USED DURING THE EVENT WERE RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE DEVICE PASSED DEFIB CYCLE, BENCH HANDLING, IMPEDANCE CALIBRATION AND ALL FUNCTIONAL TESTS. THE PADS PASSED CONTINUITY TESTING. A VISUAL INSPECTION DID NOT IDENTIFY ANY ABNORMALITIES OR INDICATIONS OF POOR PATIENT PREPARATION. SLIGHT DISCOLORATION IN AREAS OF THE ELECTRODE TINS WERE OBSERVED CONSISTENT WITH ARCING BETWEEN THE PATIENT AND ELECTRODE. A DHR REVIEW OF THE PADS LOT WAS PERFORMED AND DETERMINED THE PADS WERE ASSEMBLED TO SPECIFICATION. THE CLINICAL LOG WAS NOT AVAILABLE FOR REVIEW; HOWEVER, THE ACTIVITY LOG SHOWED A DIFFERENCE IN PATIENT IMPEDANCE AND DEFIB IMPEDANCE WHICH IS AN INDICATION OF POOR COUPLING. ELECTRODE LABELING STATES THE IMPORTANCE OF GOOD PLACEMENT ON THE PATIENT AND PROVIDES INSTRUCTION FOR PROPER ELECTRODE APPLICATION TECHNIQUE. ZOLL RECOMMENDS THAT PATIENTS ARE CLEANED AND HAIR IS CLIPPED PRIOR TO APPLYING ELECTRODE PADS TO ASSURE GOOD COUPLING OF ELECTRODE TO SKIN CONTACT. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21367671&pc=DYB | 2015691-2025-00957 | 2025-01-07 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-12 05:00 | DYB | EDWARDS ESHEATH PLUS INTRODUCER SHEATH | Material Puncture/Hole; Material Integrity Problem; Physical Resistance/Sticking | Insufficient Information | K200258 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 26 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION, THE VALVE WAS UNABLE TO PASS THROUGH THE ESHEATH+ AND A STRUT BENT WHILE TRYING TO PUSH THE VALVE THROUGH IT. A SECOND VALVE WAS ALSO USED IN THE SAME MANNER AND ALSO DAMAGED. THE DELIVERY SYSTEM WAS UNABLE TO ADVANCE THE FIRST TIME OUT OF THE PARTIAL EXPANDING PORTION AND THE SECOND TIME IT WAS STUCK JUST DISTAL TO THE PARTIAL EXPANDING PORTION. IT'S NOTED THERE WAS A TEAR AT THE PARTIALLY EXPANDING TO FULLY EXPANDING TRANSITION POINT OF THE FIRST ESHEATH+. THE SECOND ESHEATH+ WAS DAMAGED AFTER THE CASE IN THE PARTIALLY EXPANDING PORTION DUE TO TRYING TO GET THE VALVE OUT TO EXAMINE. THE PATIENT HAD RAPID DECOMPENSATION WITH NO EFFUSION NOTED IN THE CHEST. THE PHYSICIAN BELIEVES CAUSE WOULD POSSIBLY BE FROM RETROPERITONEAL BLEED FROM THE ILIAC ARTERY, HOWEVER NO ANGIO WAS TAKEN DURING CPR AND LITTLE TO NO FLOW MADE IT INTO THE ILLIACS TO EXAMINE IF/ WHERE THE PATIENT WAS BLEEDING FROM. ADDITIONAL INFORMATION RECEIVED NOTED IT'S STILL AN UNKNOWN CAUSE OF DEATH AND AN AUTOPSY IS BEING DONE, THOUGH THE RESULTS HAVE NOT YET BEEN MADE AVAILABLE. AFTER REVIEWING THE CASE WITH THE CT SURGEON, IT WAS NOTED IN THE ANESTHESIA CHART THAT THE PATIENT'S PRESSURE WAS MAPPING IN THE 30-40 FOR A TOTAL OF 8 MINUTES BEFORE ANY ACLS OR ACTION WAS TAKEN. IT'S UNKNOWN IF THE PATIENT HAD A BLEED SOMEWHERE OR IF THEY JUST DECOMPENSATED AND COULD NOT RECOVER. THE PERCEIVED ROOT CAUSE OF THE BENT STRUTS WAS THAT THE STRUTS CAUSE ON THE ILIAC STENTS. THE PERCEIVED ROOT CAUSE OF THE RESISTANCE WAS EITHER THE STENTS IN THE ILIAC ARTERY OR CALCIUM BURDEN. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR A COMPLETION TO THE ENGINEERING EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED OR CORRECTED: B4, G3, G6, H2, H3, H6, H11. THE PRODUCT WAS RETURNED; THEREFORE, A PRODUCT EVALUATION INVESTIGATION WAS COMPLETED. FUNCTIONAL TESTING WAS PERFORMED ALONG WITH A VISUAL EXAMINATION. THE ASSOCIATED COMMANDER DELIVERY SYSTEM WAS ADVANCED THROUGH THE SHEATH AND THE REMAINING SHEATH EXPANDED AND THE TIP OPENED, AS DESIGNED. AS THE ASSOCIATED DELIVERY SYSTEM WAS ABLE TO BE FULLY ADVANCED THROUGH THE SHEATH WITH NO ISSUES, DIMENSIONAL TESTING INCLUDING MEASURING THE DELIVERY SYSTEM FLEX TIP OUTER DIAMETER (OD) WAS NOT NECESSARY TO BE PERFORMED. THE RETURNED DEVICES WERE VISUALLY EXAMINED, AND THE FOLLOWING WAS OBSERVED: SHEATH SHAFT CURVATURE. EXPOSED THV THROUGH PUNCTURED STRAIN RELIEF APPROXIMATELY 1.5" FROM COMNUT. ADDITIONAL SHEATH PUNCTURE APPROXIMATELY 0.7" FROM DISTAL STRAIN RELIEF. LINER PARTIALLY EXPANDED APPROXIMATELY 4.5" FROM STRAIN RELIEF. REMAINING LINER UNEXPANDED AND DISTAL TIP UNOPENED. SHEATH SHAFT KINK APPROXIMATELY 1" FROM DISTAL STRAIN RELIEF. PROVIDED IMAGERY WAS REVIEWED, CONFIRMING THE COMPLAINT EVENT. THE 3MENSIO IMAGERY WAS PROVIDED AND THE FOLLOWING WAS OBSERVED: PATIENT'S ACCESS VESSELS HAD PRESENCE OF TORTUOSITY, CALCIFICATION/STENTS, AND UNDERSIZED VESSEL DIAMETERS. THE 14F ESHEATH+ IS INDICATED FOR USE WITH VESSEL DIAMETERS '5.5MM. AS THE RETURNED DEVICE MEETS SPECIFICATION WITH NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT HAS CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. WHILE THE IFU/TRAINING MANUALS SPECIFIC TO THIS COMPLAINT EVENT ARE NOT LISTED IN THE TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, REVIEW OF THE INSTRUCTIONS REVEALS SIMILAR CONTENT AS DOCUMENTED IN THE TECHNICAL SUMMARY. THEREFORE, THE REVIEW OF THE INSTRUCTIONS/MANUALS WITHIN THE TECHNICAL SUMMARY REMAIN EQUIVALENT TO THE INSTRUCTIONS/MANUALS FOR THIS COMPLAINT EVENT. THE CONTENT ADEQUATELY PROVIDES WARNINGS AND INSTRUCTIONS FOR USE REGARDING THE REPORTED COMPLAINT EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, IFU WARNINGS AND CAUTIONS, AND PHYSICIAN TRAINING. NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION.' THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR INABILITY TO ADVANCE THROUGH SHEATH WAS CONFIRMED BASED ON THE EVALUATION OF THE RETURNED DEVICE AND PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (UNDERSIZED VESSEL, CALCIFICATION/STENTS, TORTUOSITY) LIKELY CONTRIBUTED TO THE EVENT AS THESE PATIENT FACTORS WERE CONFIRMED VIA 3MENSIO IMAGERY. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. THE COMPLAINT FOR SHEATH PUNCTURED WAS CONFIRMED BASED ON THE EVALUATION OF THE RETURNED DEVICE AND PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (UNDERSIZED VESSEL, CALCIFICATION, TORTUOSITY) AND/OR PROCEDURAL FACTORS (DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT AS INDICATED BY THE CONDITION OF THE RETURNED DEVICE. IT IS POSSIBLE THAT ADDITIONAL DEVICE MANIPULATION TO OVERCOME THE RESISTANCE MAY BE APPLIED DURING THE PROCEDURE RESULTING IN DAMAGE TO THE SHEATH. DURING DELIVERY SYSTEM ADVANCEMENT THROUGH A CHALLENGING PATHWAY, IT IS POSSIBLE THAT THE DEVICES MAY NOT BE COAXIALLY ALIGNED. THIS CAN LEAD TO THE CRIMPED VALVE TO CATCH ONTO THE SHEATH HDPE AND/OR STRAIN RELIEF AND LEAD TO DAMAGE. EXCESSIVE DEVICE MANIPULATION APPLIED TO OVERCOME THE RESISTANCE CAN FURTHER DAMAGE THE SHEATH THROUGH CONTINUED INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. VESSEL CALCIFICATION AND/OR TORTUOSITY IF PRESENT CAN EXASPERATE THE INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. SHARP CALCIFIED NODULES CAN DIRECTLY WEAKEN THE SHEATH SHAFT MAKING IT MORE SUSCEPTIBLE TO DAMAGE AS THE DELIVERY SYSTEM WITH CRIMPED VALVE IS ADVANCED THROUGH. TORTUOSITY CAN SUBJECT THE SHEATH TO SUBOPTIMAL ANGLES THAT CAN LEAD TO SHAFT KINKS. EXCESSIVE MANIPULATION CAN LEAD TO SHEATH SHAFT DAMAGE (I.E. KINKS, SCRATCHES, TIP DAMAGE) IF COMPOUNDED WITH VESSEL CALCIFICATION AND/OR TORTUOSITY. THE PRODUCT RISK ASSESSMENT (PRA) IS CAPTURED IN THE TECHNICAL SUMMARY, WHICH APPLIES TO THIS COMPLAINT EVENT. THE RISKS OUTLINED IN THE PRA REMAIN THE SAME. THE PRA DOCUMENTS THE DIFFICULTY ADVANCING THE DELIVERY SYSTEM THROUGH THE SHEATH, RESULTING IN DIFFICULTY OR THE INABILITY TO INTRODUCE DELIVERY SYSTEM/LOADER AND ADVANCE THROUGH SHEATH, PROLONGING PROCEDURE, WHICH DID OCCUR DURING THIS EVENT. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. AS THE COMPLAINT DID NOT EXCEED THE PRA RE-ESCALATION THRESHOLD, NO FURTHER ACTION IS REQUIRED. A CAPA/SCAR IS NOT REQUIRED BECAUSE AN EDWARDS/SUPPLIER DEFECT WHICH COULD HAVE RESULTED IN THE COMPLAINT WAS NOT CONFIRMED. PRIOR CORRECTIVE ACTION, DETAILED IN THE TECHNICAL SUMMARY, CAPTURES PRIOR HIGH PUSH FORCE INVESTIGATION AND CORRECTIVE/PREVENTATIVE ACTION. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. NO FURTHER ESCALATION IS NEEDED AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21367882&pc=DXF | 2124215-2025-07813 | 2025-01-23 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-12 05:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Air Embolism; Arrhythmia; Bradycardia; Low Blood Pressure/ Hypotension; Hypoxia; Ventricular Fibrillation; Asystole | K150709 |  | 1 | Event Description: DURING THE PROCEDURE, A PATIENT DEATH OCCURRED. IT HAS BEEN REPORTED THAT A VERSACROSS CONNECT LAAC ACCESS SOLUTION KIT WAS SELECTED FOR USE FOR A WATCHMAN PROCEDURE. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS DONE IMMEDIATELY PRIOR TO PLACING THE PATIENT UNDER TWILIGHT CONSCIOUS SEDATION. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS PERFORMED WHILE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS INSERTED WITH VERSACROSS CONNECT TO PERFORM TRANSSEPTAL PUNCTURE (TSP). TSP WENT SMOOTHLY AND NO SUSPICION OF EVENT FROM TSP. THE INTRAARTERIAL SEPTUM (IAS) WAS FLOSSED WITH THE WAS AND VCC AND THE WAS AND ICE CATHETER WAS POSITIONED ACROSS THE IAS INTO THE LEFT ATRIUM (LA). DURING A FIVE (5) MINUTE PERIOD, THE WAS AND ICE REMAINED IN THE LA UNATTENDED AS THE PHYSICIAN NEEDED TO LEAVE THE ROOM BRIEFLY, AND WHEN THE PHYSICIAN RETURNED THE WAS WAS NOT ABLE TO BE VISUALIZED ON IMAGING. THE QRS INTERVAL BEGAN WIDENING, BRADYCARDIA WAS NOTED AND THEN ASYSTOLE. AS STAFF WAS PREPARING TO PERFORM CARDIOPULMONARY RESUSCITATION (CPR), VENTRICULAR FIBRILLATION WAS OBSERVED. THE STAFF AND PHYSICIAN CODED THE PATIENT FOR TWO (2) HOURS USING A TEMPORARY PACEMAKER, INTRA-AORTIC BALLOON PUMP (IABP), DIAGNOSTIC CARDIAC CATHETERIZATION (CC), AND DEFIBRILLATION. TEE WAS REINSERTED AND THERE WAS NO EVIDENCE OF PERICARDIAL EFFUSION. THE TEMPORARY PACEMAKER WAS FIRST UNABLE TO CAPTURE BETWEEN COMPRESSION CYCLES AND ULTIMATELY ONE LEAD WAS PLACED AND CAPTURED SUCCESSFULLY. THE DIAGNOSTIC CC REVEALED NO SIGN OF CORONARY DISEASE. ALTHOUGH IABP WAS SUCCESSFULLY PLACED, THE PATIENT MAINTAINED ONLY A WEAK BLOOD PRESSURE AND POOR OXYGEN SATURATION MAINTENANCE AT 70 OR BELOW. DUE TO THE LONG PERIOD OF TIME OF LOW OXYGEN SATURATION, THE PHYSICIAN DECIDED TO STOP RESUSCITATING THE PATIENT AND THE PATIENT DIED. IT WAS NOTED THAT THERE WAS NOT ANY HEAVY SNORING, NO GASPS OF AIR TAKEN BY THE PATIENT, AND NO CONTRAST INJECTIONS PERFORMED WITHIN THE LA. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IN THE PHYSICIAN'S OPINION, THE DEVICE/PROCEDURE DID NOT CONTRIBUTE TO THE PATIENT COMPLICATION.Event Description: DURING THE PROCEDURE, A PATIENT DEATH OCCURRED. IT HAS BEEN REPORTED THAT A VERSACROSS CONNECT LAAC ACCESS SOLUTION KIT WAS SELECTED FOR USE FOR A WATCHMAN PROCEDURE. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS DONE IMMEDIATELY PRIOR TO PLACING THE PATIENT UNDER TWILIGHT CONSCIOUS SEDATION. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS PERFORMED WHILE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS INSERTED WITH VERSACROSS CONNECT TO PERFORM TRANSSEPTAL PUNCTURE (TSP). TSP WENT SMOOTHLY AND NO SUSPICION OF EVENT FROM TSP. THE INTRAARTERIAL SEPTUM (IAS) WAS FLOSSED WITH THE WAS AND VCC AND THE WAS AND ICE CATHETER WAS POSITIONED ACROSS THE IAS INTO THE LEFT ATRIUM (LA). DURING A FIVE (5) MINUTE PERIOD, THE WAS AND ICE REMAINED IN THE LA UNATTENDED AS THE PHYSICIAN NEEDED TO LEAVE THE ROOM BRIEFLY, AND WHEN THE PHYSICIAN RETURNED THE WAS NOT ABLE TO BE VISUALIZED ON IMAGING. THE QRS INTERVAL BEGAN WIDENING, BRADYCARDIA WAS NOTED AND THEN ASYSTOLE. AS STAFF WAS PREPARING TO PERFORM CARDIOPULMONARY RESUSCITATION (CPR), VENTRICULAR FIBRILLATION WAS OBSERVED. THE STAFF AND PHYSICIAN CODED THE PATIENT FOR TWO (2) HOURS USING A TEMPORARY PACEMAKER, INTRA-AORTIC BALLOON PUMP (IABP), DIAGNOSTIC CARDIAC CATHETERIZATION (CC), AND DEFIBRILLATION. TEE WAS REINSERTED AND THERE WAS NO EVIDENCE OF PERICARDIAL EFFUSION. THE TEMPORARY PACEMAKER WAS FIRST UNABLE TO CAPTURE BETWEEN COMPRESSION CYCLES AND ULTIMATELY ONE LEAD WAS PLACED AND CAPTURED SUCCESSFULLY. THE DIAGNOSTIC CC REVEALED NO SIGN OF CORONARY DISEASE. ALTHOUGH IABP WAS SUCCESSFULLY PLACED, THE PATIENT MAINTAINED ONLY A WEAK BLOOD PRESSURE AND POOR OXYGEN SATURATION MAINTENANCE AT 70 OR BELOW. DUE TO THE LONG PERIOD OF TIME OF LOW OXYGEN SATURATION, THE PHYSICIAN DECIDED TO STOP RESUSCITATING THE PATIENT AND THE PATIENT DIED. IT WAS NOTED THAT THERE WAS NOT ANY HEAVY SNORING, NO GASPS OF AIR TAKEN BY THE PATIENT, AND NO CONTRAST INJECTIONS PERFORMED WITHIN THE LA. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IN THE PHYSICIAN'S OPINION, THE DEVICE/PROCEDURE DID NOT CONTRIBUTE TO THE PATIENT COMPLICATION. IT HAS BEEN FURTHER MENTIONED THAT THERE WAS NO ISSUE WITH TRANSSEPTAL OR THE VERSACROSS DEVICES. AUTOPSY SUSPECTS AIR EMBOLISM IS THE CAUSE OF DEATH. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION IS IN PROGRESS, BUT IT WAS NOT AVAILABLE AS OF YET. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: DUE TO ADDITIONAL INFORMATION RECEIVED, THIS REPORT IS BEING SUBMITTED FOR THE UPDATED FIELDS DESCRIBE EVENT OR PROBLEM (B5) IN SECTION B - ADVERSE EVENT/PRODUCT PROBLEM, AND PATIENT CODES (F10 AND H6), IN SECTIONS F - FOR USE BY UF/IMPORTER AND H - DEVICE MANUFACTURERS ONLY. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368008&pc=NPT | 2025587-2025-01211 | 2025-01-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-12 05:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE EVFXPLUS-29, THE PATIENT DEVELOPED AN EFFUSION. THE PROCEDURE WAS ABORTED TO STABILIZE THE PATIENT, WHO WAS THEN PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION AND TRANSFERRED TO THE INTENSIVE CARE UNIT. SUBSEQUENTLY, THE PATIENT WAS PUT ON COMFORT CARE AND LATER PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN, AND IT WAS UNCLEAR IF THE PROCEDURE OR THE MEDTRONIC DEVICE CONTRIBUTED TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE EXACT CAUSE OF THE EFFUSION WAS UNKNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368065&pc=DSQ | 2916596-2025-00698 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-12 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO SEVERE STROKE LEFT HEMISPHERE. PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE REPORTED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: G2 - REPORT SOURCE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368332&pc=DSQ | 2916596-2025-00900 | 2024-01-15 05:00 | Death | THORATEC CORPORATION | 2025-02-12 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS UNKNOWN.Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO HOSPITAL FROM SKILLED NURSING FACILITY WITH COVID-19 INFECTION. THE PATIENT WAS INTUBATED AND THEN HAD TRAUMATIC HEMOPTYSIS AFTER PULLING THEIR OWN NASOGASTRIC TUBE LEADING TO MASSIVE BLEEDING AND ASPIRATION. THEY THEN WERE MADE COMFORT CARE BY FAMILY DUE TO POOR REHAB POTENTIAL. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: IT WAS REPORTED THAT THE PATIENT EXPERIENCED MASSIVE BLEEDING AFTER PULLING THEIR NASOGASTRIC TUBE. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED BLEEDING COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS NOT RETURNED FOR EVALUATION. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING BLEEDING. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368441&pc=LGW | 1627487-2025-00641 | 2025-01-25 05:00 | Death | ABBOTT MEDICAL | 2025-02-12 05:00 | LGW | ORION RC SCS IPG | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Myocardial Infarction | P010032 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THEIR IMPLANT PROCEDURE ON (B)(6) 2025, THE PATIENT PASSED AWAY ON (B)(6) 2025 FROM A MYOCARDIAL INFARCTION AND CEREBROVASCULAR ACCIDENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368794&pc=FSA | 3009402404-2025-00008 | 2025-01-15 05:00 | Death | JOERNS HEALTHCARE | 2025-02-12 05:00 | FSA | HOYER LIFT | Appropriate Term/Code Not Available | Fall |  |  | 1 | Event Description: IT WAS REPORTED TO THE MANUFACTURER, BY THE END USER, PER THE END USER, THAT CUSTOMER STATES WHILE USING THE HPL700-S2 LIFT THERE WAS SLACK IN THE LOOPS. ALLEGEDLY, THIS LIFT DOES NOT HAVE SAFETY CLIPS. THE STAFF DID NOT ENSURE THE LOOPS WERE ALL THE WAY ON THE SIDE. PATIENT FELL OFF THE LIFT ON THEIR HEAD. THEY WERE HOSPITALIZED AND BROKE THEIR NECK. CUSTOMER REPORTS THAT THE PATIENT WAS IN THE HOSPITAL AND RELEASED. A DAY AFTER RELEASE, THE PATIENT PASSED AWAY. THE FOLLOWING INJURIES ALLEGEDLY OCCURRED: PATIENT BROKE THEIR NECK. REPORTER REPORTED THAT THE INJURED PARTY WAS TREATED AT A HOSPITAL/CLINIC AND RELEASED. THE PRODUCT HAS BEEN TAKEN OUT OF USE. THE RESIDENT'S WEIGHT IS 177.5 LBS. 2 OTHER PEOPLE, AIDS, WERE PRESENT DURING THE INCIDENT. WHEN ASKED ABOUT AN IDEAL OUTCOME, FACILITY WANTS HOYER TO BE ADVISED ABOUT THIS ISSUE WITH NO SAFETY CLIPS. FOLLOW UP NEEDS TO TAKE PLACE WITHIN 24 HOURS. JOERNS IS NOT AWARE OF THE SLING THE CUSTOMER WAS USING AND HAS REQUESTED THAT INFORMATION. NO RESPONSE HAS BEEN GIVEN BY THE CUSTOMER. COMPLAINT # (B)(4) AND RA# (B)(4) WERE ENTERED INTO OUR SYSTEM TO HAVE THE PRODUCTS RETURNED FOR INVESTIGATION. AS OF THIS WRITING, THE PRODUCTS HAVE NOT BEEN RETURNED. Manufacturer Narrative: THIS REPORT OR OTHER INFORMATION SUBMITTED BY JOERNS HEALTHCARE UNDER 21 CFR PART 803, AND RELEASE BY FDA OF THAT REPORT INFORMATION, DOES NOT REFLECT A CONCLUSION OR ADMISSION BY JOERNS HEALTHCARE , ITS EMPLOYEES, ITS CONTRACT SERVICE FIRMS, OR THEIR EMPLOYEES, FINISHED DEVICE SUPPLIERS, OR THEIR EMPLOYEES CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368834&pc=CBK | 2023050-2025-00005 | 2025-01-17 05:00 | Death | NEWPORT MEDICAL INSTRUMENT, INC. | 2025-02-12 05:00 | CBK | HT70 VENTILATOR | No Flow | Respiratory Failure | K111146 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE IN USE ON AN ELDERLY, TERMINALLY ILL PATIENT, THE HT70 VENTILATOR GENERATED AN ALARM AND VENTILATION STOPPED. WHILE BEING TRANSFERRED TO AN ALTERNATE VENTILATOR, THE PATIENT EXPERIENCED CARDIAC ARREST AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, BUT THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: SECTION G9 FDA 806 NOTIFICATION NUMBER: 2024500-05-13-2025-001-R UPDATE TO SECTION D4 UNIQUE IDENTIFIER (UDI)# SECTION D9 WAS UPDATED DUE TO PART RETURN SECTION H6 EVALUATION CODES WERE UPDATED DUE TO ANALYSIS OF RETURNED PARTS METHOD CODE (ANNEX B) ADDITIONAL CODES ADDED RESULT CODE (ANNEX C) ADD ADDITIONAL CODE CONCLUSION CODE (ANNEX D) REMOVE D01 AND ADD D02 H3 DEVICE EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. IT WAS REPORTED THAT WHILE IN USE ON A TERMINALLY ILL PATIENT, THE HT70 VENTILATOR GENERATED AN ALARM AND VENTILATION STOPPED. THE SERVICE PERSONNEL (SP) INSPECTED THE VENTILATOR AND CONFIRMED THE REPORTED ISSUE. SP FOUND A SYSTEM ERROR AND REPLACED THE SINGLE BOARD COMPUTER (SBC) AND CONTROL BOARD. ADDITIONALLY, SP FOUND THE SYSTEM LEAK VALUE WAS OUTSIDE THE ALLOWABLE RANGE, SO REPLACED THE PUMP AND MANIFOLD ASSEMBLY. SP ALSO FOUND THAT THE POSITIVE END-EXPIRATORY PRESSURE (PEEP) VALUE WAS OUTSIDE THE PERMISSIBLE RANGE, SO REPLACED THE ON/OFF VALVE. THE VENTILATOR PASSED ALL CALIBRATIONS AND TESTS PER MANUFACTURER SPECIFICATIONS AT THE TIME OF SERVICE. THE PUMP AND MANIFOLD ASSEMBLY AND ON/OFF VALVE WERE NOT RETURNED TO MEDTRONIC. ONE CONTROL BOARD WAS RETURNED TO MEDTRONIC FOR ANALYSIS. VISUAL INSPECTION OF THE RETURNED CONTROL BOARD, IN ADDITION TO THE IMAGE PROVIDED, SHOWED SIGNS OF DAMAGE TO CAPACITOR C92 ON THE CONTROL BOARD. IF A FAILURE OF CAPACITOR C92 OCCURS WHILE IN USE ON A PATIENT, THE VENTILATOR RESPONSE WOULD BE A LOSS OF POWER/VENTILATION FOR THE UNIT. THE SBC WAS RETURNED TO MEDTRONIC FOR ANALYSIS. THE COMPONENT WAS VISUALLY INSPECTED IN ADDITION TO THE IMAGE PROVIDED AND FUNCTIONALLY TESTED WITH NO MALFUNCTION OR PRODUCT DEFICIENCY OBSERVED. THE CAUSE OF THE EVENT WAS ISOLATED TO A FAULTY C92 CAPACITOR ON THE HT70 CONTROL BOARD. THE VENTILATOR IS IN THE SCOPE OF THE FIELD CORRECTIVE ACTION. ADDITIONAL SYSTEM LEAK AND PEEP ISSUE WAS ISOLATED TO POTENTIAL FAULT IN PUMP AND MANIFOLD ASSEMBLY AND ON/OFF VALVE. THE ADDITIONAL EVENTS ARE INCLUDED IN A DATA MONITORING PLAN. A MEDICAL SAFETY REVIEW (MSR) WAS PERFORMED AND DETERMINED THAT THE DEVICE COULD CONTRIBUTE TO THE REPORTED EVENT, AND THE RESULTS SHOW THAT REPORTED HARMS AND SEVERITIES ARE PROPERLY ASSESSED AND THE RESULTS ARE WITHIN THE RISK DOCUMENT AND THE CAUSAL RELATIONSHIP TO THE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H3: IMAGES OF THE SINGLE BOARD COMPUTOR (SBC) BOARD AND THE MAIN CONTROL BOARD WERE RECEIVED FOR REVIEW. THERE WAS THERMAL DAMAGE OBSERVED WHICH INCLUDED THE C92 CAPACITOR ON THE MAIN CONTROL BOARD. THERE IS AN EXISTING PREVIOUS INTERNAL INVESTIGATION REGARDING THIS ISSUE. THE ACTUAL SBC AND MAIN CONTROL BOARDS WERE RECEIVED BY MEDTRONIC AND ARE UNDERGOING ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21368959&pc=OZD | 1220648-2025-26121 | 2025-01-17 05:00 | Death | ABIOMED, INC. | 2025-02-12 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND POST IMPLANT, SAME DAY, OOZING WAS NOTED WHEN THE PATIENT LIFTED THE IMPELLA LEG. SYSTEMIC HEPARIN WAS NOT STARTED AS THE PHYSICIAN WAS TO START SYSTEMIC HEPARIN DRIP WHEN ACTIVATED CLOTTING TIME IS IN RANGE OF 160-180 AS IT WAS 260. IMPELLA SITE WAS MONITORED FOR CONTINUED OOZING. FIVE DAYS LATER, THE PATIENT WAS CANNULATED FOR VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) OVERNIGHT. THE IMPELLA CP WAS AT THIS TIME BEING USED FOR LEFT VENTRICLE UNLOADING. IN THE AFTERNOON, THE PATIENT WAS TRANSFERRED TO AN OUTSIDE HOSPITAL FOR A LEFT VENTRICULAR ASSIST DEVICE (LVAD) WORKUP. PATIENT ESCALATION TO IMPELLA 5.5 WAS NOTED. LATER, THAT EVENING, THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR THE IMPELLA 5.5 ESCALATION. THE PHYSICIAN UTILIZED A DOUBLE BARREL TECHNIQUE TO EXCHANGE THE IMPELLA CP FOR THE IMPELLA 5.5 WITH SUCCESSFUL FLOW SWAP. THE IMPELLA CP WAS EXPLANTED WITHOUT COMPLICATIONS, RE-ACCESS PORT UTILIZED FOR PERCLOSURE. HOWEVER, POST PERCLOSURE ON RIGHT FEMORAL ARTERY FOR IMPELLA CP EXPLANT, DOPPLER PULSES WERE EXTREMELY MONOPHASIC AND DIFFICULT TO FIND. AN ORDER WAS MADE FOR AN ARTERIAL ULTRASOUND STUDY FOR RIGHT LOWER EXTREMITY. THE PATIENT WAS TRANSFERRED TO THE UNIT ON SUPPORT. THE STUDY SHOWED FLOW. THE FOLLOWING DAY, IT WAS NOTED THE PATIENT WAS NOT BEING WORKED UP FOR A DURABLE LVAD OR TRANSPLANT, AS THE TEAM STATED THE PATIENT WOULD ONLY QUALIFY IF RENAL RECOVERY WAS PRESENT AND DECIDED TO ATTEMPT WEANING DEVICES FIRST TO MONITOR RECOVERY. THE FOLLOWING DAY AN ATTEMPT AT ECMO WEANING/DECANNULATION WHILE CONTINUING TO RIDE 5.5 SUPPORT WAS MADE. PATIENT ON CONTINUOUS RENAL REPLACEMENT THERAPY. THE NEXT DAY THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR DECANNULATION OF ECMO; ECMO WAS REMOVED AND IMPELLA 5.5 UP TO SUPPORT LEVEL P-8 SUPPORTING THE PATIENT WITHOUT INCIDENT. NO ADVANCED THERAPIES WERE OFFERED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION ARREST THE FOLLOWING DAY, CODED FOR APPROXIMATELY 20 MINUTES BEFORE THE TIME OF DEMISE WAS CALLED. ALTHOUGH THE PATIENT WAS VERY ILL NO ADVANCED THERAPIES AVAILABLE, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA CP DEVICE AS A POSSIBLE ASSOCIATED FACTOR. Manufacturer Narrative: THE INVESTIGATION FOR THE LIMB ISCHEMIA AND ACCESS SITE BLEED HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL DETAILS WERE SUFFICIENT TO DETERMINE THE ROOT CAUSE FOR THE ACCESS SITE BLEEDING. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY PATIENT MOVEMENT. SOME OOZING WAS NOTED AFTER THE PATIENT LIFTED LEG. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ B.7 INFORMATION WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26121. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26121 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4614 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26121. A NEW CODE WAS ADDED. H.10 PART OF THE INSTRUCTIONS FOR USE WER INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26121 AND WERE ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21369148&pc=DSQ | 2916596-2025-00702 | 2024-04-12 04:00 | Death | THORATEC CORPORATION | 2025-02-12 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21369910&pc=QBJ | 3004753838-2025-036386 | 2025-01-14 05:00 | Death | DEXCOM, INC. | 2025-02-12 05:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Insufficient Information | Insufficient Information | K213919 |  | 1 | Event Description: IT WAS REPORTED BY A HEALTHCARE PROVIDER VIA MAUDE (MW5164857) AND BY THE FOOD AND DRUG ADMINISTRATION THAT A PATIENT PASSED AWAY ON (B)(6) 2025. THE MAUDE REPORT STATES THAT THE PATIENT WAS SUPPOSED TO COME TO THE HEALTHCARE PROVIDER (HCP) OFFICE BUT CANCELLED AND DID NOT RESCHEDULE. THE DATE OF DEATH WAS NOT PROVIDED. THERE IS NO REPORTER INFORMATION TO GATHER ADDITIONAL INFORMATION AS THE HCP DID NOT WISH TO FOLLOW UP AND THERE IS NO PATIENT INFORMATION AVAILABLE. NO DATA OR PRODUCT WAS PROVIDED FOR INVESTIGATION. THE ALLEGATION WAS UNDETERMINED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21369994&pc=DSQ | 2916596-2025-00673 | 2023-09-30 04:00 | Death | THORATEC CORPORATION | 2025-02-12 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR (RV) FAILURE, MULTISYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND WAS DUE TO RV AND MULTI-SYSTEM ORGAN FAILURE. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21371020&pc=EOQ | 2955842-2025-01565 | 2025-01-22 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-02-12 05:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED A CARDIAC ARREST AND THE PROCEDURE WAS ABORTED. ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS ADMINISTERED, RESULTING IN THE RETURN OF SPONTANEOUS CIRCULATION (ROSC). HOWEVER, THE PATIENT CODED FIVE MORE TIMES IN THE PROCEDURE ROOM. EACH TIME, ACLS WAS PERFORMED UNTIL THE PATIENT WAS STABLE ENOUGH FOR TRANSFER TO THE INTENSIVE CARE UNIT. THE PATIENT CODED AGAIN IN THE ICU AND AT AN UNSPECIFIED TIME LATER PASSED AWAY. THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE (ETA) WAS PRESENT DURING THE ION PROCEDURE. THE PHYSICIAN INFORMED THE ETA THAT THE CARDIAC ARREST AND THE PATIENT OUTCOME WERE NOT RELATED TO THE ION SYSTEM AND WOULD LIKELY HAVE OCCURRED WITH ANOTHER MODALITY. THE PHYSICIAN ATTRIBUTED THE EVENT TO THE PATIENT'S EXTENSIVE MULTIPLE COMORBID CONDITIONS; THE PATIENT'S MEDICAL HISTORY WAS NOT PROVIDED. THERE WAS NO DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. Manufacturer Narrative: THE PHYSICIAN STATED THAT THE PATIENT OUTCOME WAS RELATED TO EXTENSIVE, MULTIPLE CO-MORBIDITIES. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND THERE WERE NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH EXTENSIVE MEDICAL COMORBIDITIES NOT OTHERWISE SPECIFIED UNDERWENT AN ION LUNG BIOPSY. DURING THE PROCEDURE THE PATIENT SUFFERED MULTIPLE EPISODES OF CARDIAC ARREST WITH EVENTUAL ROSC THEN TRANSFERRED TO THE ICU WHERE THE PATIENT ARRESTED AGAIN AND ULTIMATELY DIED. THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT OR ACCESSORIES. THE AVAILABLE DATA SUGGESTS THE ADVERSE EVENT WAS RELATED TO THE PROCEDURE IN A PATIENT WITH SIGNIFICANT UNDERLYING CO-MORBIDITIES. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. ---FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). ---OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. ---FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. ---KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. ---BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. FOR SECTION B2: DATE OF DEATH IS ESTIMATED AS SAME DAY AS PROCEDURE(B)(6) 2025 OR FOLLOWING DAY (B)(6) 2025. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21371160&pc=NAY | 2955842-2025-01566 | 2025-01-06 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-02-12 05:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LEFT ADRENALECTOMY PROCEDURE, BLEEDING OCCURRED AND THE PROCEDURE WAS CONVERTED TO OPEN. IT WAS REPORTED THAT THE PATIENT EXPIRED. IT IS UNKNOWN IF THE PATIENT EXPIRED INTRA-OPERATIVELY OR POST-OPERATIVELY. THE CUSTOMER REPORTED THAT THIS EVENT HAD NOTHING TO DO WITH THE DA VINCI PRODUCTS USED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, PERMANENT CAUTERY HOOK, SMALL GRASPING RETRACTOR, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS, LONG BIPOLAR GRASPER, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN ADRENALECTOMY. NEITHER THE DIAGNOSIS, NOR COMORBIDITIES, DETAILS OF THE CASE, POST OPERATIVE COURSE, OR CAUSE OF DEATH ARE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21373474&pc=QZI | MW5166062 | 2025-01-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-12 05:00 | QZI | FARAWAVE PULSED FIELD ABLATION (PFA) SYSTEM |  | Tachycardia; Electrolyte Imbalance; Cardiogenic Shock; Unspecified Respiratory Problem |  |  | 1 | Event Description: IT WAS REPORTED DURING AN AFIB CASE, THE PATIENT WENT INTO CARDIOGENIC SHOCK. CALLER REPORTED THAT THE PATIENT¿S POTASSIUM LEVELS WERE ¿WAY OUT OF RANGE, SUPER HIGH¿ AND THE PATIENT KEPT GOING INTO VT(VENTRICULAR TACHYCARDIA). CALLER STATED THEY CALLED A CODE AND DECIDED TO GIVE THE PATIENT A BALLOON PUMP. CALLER STATED THE PATIENT WAS SENT TO THE ICU AND AS OF LAST NIGHT WAS NOT AWAKE YET. THE PHYSICIAN DID NOT CONFIRM THE CAUSE OF INJURY. CALLER STATED THE PHYSICIAN THOUGHT THE PATIENT MAY BE TOO UNSTABLE FOR ANESTHESIA. THE SOUNDSTAR CATHETER (B)(6) AND THE DECA NAV CATHETER WAS IN THE BODY AT THE TIME OF THE INJURY. PFA PROCEDURE USING THE FARAPULSE SYSTEM. THE PATIENT FULLY RECOVERED FROM SAID PROCEDURE BUT LATER DIED DUE TO PRE-EXISTING CONDITIONS. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION BECAUSE THE PATIENT WAS PUT ON A VENTILATOR FOR WHAT WAS BELIEVED TO BE UNRELATED CONDITIONS. THE SOUNDSTAR CATHETER (B)(6) AND THE DECA NAV CATHETER WAS IN THE BODY AT THE TIME OF THE INJURY. IT WAS A PFA(PULSED FIELD ABLATION) PROCEDURE USING THE FARAPULSE SYSTEM. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS PATIENT CONDITION. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21352601&pc=NIQ | 3002806593-2025-00002 | 2025-01-14 05:00 | Death | BIOSENSORS EUROPE SA (BESA) | 2025-02-11 05:00 | NIQ | BIOFREEDOM DRUG COATED CORONARY STENT SYSTEM | Obstruction of Flow | Myocardial Infarction; Thrombosis/Thrombus | P190020 |  | 1 | Event Description: THIS IS A SERIOUS ADVERSE EVENT REPORTED IN A POST-MARKET CLINICAL INVESTIGATIONAL STUDY CONDUCTED BY BIOSENSORS IN ASIA RELATED TO LEADERS FREE IV, IN WHICH THE BIOFREEDOM DEVICE SERVED AS A CONTROL DEVICE. PATIENT: FEMALE, 44 YEARS OLD. CO-MORBIDITIES: HYPERTENSION, TYPE 2 DIABETES, CHRONIC KIDNEY DISEASE, HEART FAILURE, REDUCED EJECTION FRACTION, GASTRITIS AND ANEMIA. THE INDEX PCI WAS PERFORMED ON (B)(6) 2025. THE PATIENT WAS LOADED WITH ASPIRIN 100MG AND TICAGRELOR 180MG. HEPARIN 3,000 UNITS DURING THE PROCEDURE. MILD DISEASE OF DISTAL LMS, 80% STENOSIS OSTIAL LAD AND 90% STENOSIS PLAD (OSTIAL PROXIMAL LESION). PERCUTANEOUS CORONARY INTERVENTION (PCI) DONE AT LM-LAD VESSEL. THE TARGET LESION NOT PREVIOUSLY TREATED, THERE WAS NO TOTAL OCCLUSION AND NO BIFURCATION. PCI APPROACHED VIA RIGHT RADIAL ACCESS. BY VISUAL ESTIMATION, THE LESION HAD A LENGTH OF 60 MM, REFERENCE VESSEL DIAMETER OF 3MM AND 80% STENOSIS. PRE-PROCEDURE TIMI FLOW GRADE WAS 3 AND LESION IS CLASSIFIED AS TYPE B2 LESION (ACC/AHA CLASS). PRE-DILATED THE PLAD TO DLMS USING 2.5X15MM AT 18ATM (NC BALLOON), 3.0X15MM AT 18ATM (NC BALLOON) AND 3.5X15MM AT 18ATM (SCORING BALLOON). MEDTRONIC EBU 3.0 GUIDE CATHETER WAS USED WITH GOOD SUPPORT. SION BLUE GUIDEWIRE CROSSED INTO LAD. TWO BIOFREEDOM STENTS (BFR1-3036 & BFR1-3528) WAS IMPLANTED, IN AN OVERLAPPING MANNER. POST-DILATATION WAS DONE USING NC 3.5X15MM BALLOON AT 18ATM AND NC 4.0X15MM BALLOON AT 16 ATM. INTRAVASCULAR ULTRASOUND (IVUS) WAS NOT PERFORMED. NO DIFFICULTIES DURING DELIVERY, STENT WAS PROPERLY ASSESSED ANGIOGRAPHICALLY AND DETERMINED APPOSED TO THE WALL. POST-PROCEDURE % STENOSIS WAS 0% BY VISUAL ESTIMATION AND TIMI FLOW GRADE WAS III. PATIENT POST-PROCEDURE ADMITTED TO WARD. SHE DEVELOPED CHEST PAIN AT 04:00 A.M. ON (B)(6) 2025 AND HAD CARDIOGENIC SHOCK. PATIENT WAS RESUSCITATED, HOWEVER PATIENT SUCCUMBED AT 08:00 A.M. LACK OF EFFICACY/WORSENING OF TREATED CONDITION AND PATIENT DIED ON (B)(6) 2025. CAUSE OF DEATH FROM THE CLINICAL PRESENTATION AND ECG CHANGES (NEW ONSET RBBB AND ST ELEVATION AT AVR) LIKELY REPRESENT MYOCARDIAL INFARCTION INVOLVING LM STENT. Manufacturer Narrative: THE EVENT-RELATED STENTS WAS IMPLANTED IN THE PATIENT, THUS NOT RETURNED FOR BIOSENSORS INVESTIGATION. BASED ON THE EVENT DESCRIBED IN CONJUNCTION WITH ANGIOGRAPHY RECORD PROVIDED TO BIOSENSORS: IT APPEARS THAT THIS IS A CASE OF STENT THROMBOSIS, LEADING TO A MYOCARDIAL INFARCTION AND THE PATIENT'S DEMISE AFTER THE PLACEMENT OF BIOFREEDOM 3.00 X 36 MM AND BIOFREEDOM 3.50 X 28 MM STENTS. THERE IS A POSSIBILITY THAT STENT THROMBOSIS RESULTED FROM DISSECTION AT THE STENT EDGE, EITHER PROXIMALLY OR DISTALLY IN THE PROCEDURE. MOREOVER, THE ANTICOAGULANT ADMINISTRATION MAY HAVE BEEN INADEQUATE. THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED COMPLICATION IS RELATED TO THE CHARACTERISTICS OF THE STENT. BIOSENSORS ANALYSIS OF THE EVENT IS CONCLUDED THAT DEVICE HISTORY RECORD (DHR) AND MANUFACTURING PROCESS DID NOT SHOW ANY EVIDENCE OF PRODUCT DEFICIENCY. THE STENT THROMBOSIS AND MYOCARDIAL INFARCTION EVENTS ARE RECOGNIZED POTENTIAL HAZARDOUS SITUATION AND DOCUMENTED IN MANUFACTURER'S PRODUCT ANALYSIS AND INSTRUCTION FOR USE. THE RISKS ARE ASSESSED AND IT IS ACCEPTABLE AS BENEFITS OUTWEIGH RESIDUAL RISK. THE EVENT IS MORE LIKELY RELATED TO PATIENT AND/OR PROCEDURAL RELATED FACTORS RATHER THAN THE DEVICE'S FAULT. THERE IS NO EVIDENCE OF DEVICE OR MANUFACTURING ACTIVITIES DEFICIENCY IDENTIFIED. ACCORDINGLY, THERE IS NO CORRECTION OR CORRECTIVE ACTION TO BE IMPLEMENTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21352613&pc=NIQ | 3002806593-2025-00003 | 2025-01-14 05:00 | Death | BIOSENSORS EUROPE SA (BESA) | 2025-02-11 05:00 | NIQ | BIOFREEDOM DRUG COATED CORONARY STENT SYSTEM | Obstruction of Flow | Myocardial Infarction; Thrombosis/Thrombus | P190020 |  | 1 | Event Description: THIS IS A SERIOUS ADVERSE EVENT REPORTED IN A POST-MARKET CLINICAL INVESTIGATIONAL STUDY CONDUCTED BY BIOSENSORS IN ASIA RELATED TO LEADERS FREE IV, IN WHICH THE BIOFREEDOM DEVICE SERVED AS A CONTROL DEVICE. PATIENT: FEMALE, 44 YEARS OLD. CO-MORBIDITIES: HYPERTENSION, TYPE 2 DIABETES, CHRONIC KIDNEY DISEASE, HEART FAILURE, REDUCED EJECTION FRACTION, GASTRITIS AND ANEMIA. THE INDEX PCI WAS PERFORMED ON (B)(6) 2025. THE PATIENT WAS LOADED WITH ASPIRIN 100MG AND TICAGRELOR 180MG. HEPARIN 3,000 UNITS DURING THE PROCEDURE. MILD DISEASE OF DISTAL LMS, 80% STENOSIS OSTIAL LAD AND 90% STENOSIS PLAD (OSTIAL PROXIMAL LESION). PERCUTANEOUS CORONARY INTERVENTION (PCI) DONE AT LM-LAD VESSEL. THE TARGET LESION NOT PREVIOUSLY TREATED, THERE WAS NO TOTAL OCCLUSION AND NO BIFURCATION. PCI APPROACHED VIA RIGHT RADIAL ACCESS. BY VISUAL ESTIMATION, THE LESION HAD A LENGTH OF 60 MM, REFERENCE VESSEL DIAMETER OF 3MM AND 80% STENOSIS. PRE-PROCEDURE TIMI FLOW GRADE WAS 3 AND LESION IS CLASSIFIED AS TYPE B2 LESION (ACC/AHA CLASS). PRE-DILATED THE PLAD TO DLMS USING 2.5X15MM AT 18ATM (NC BALLOON), 3.0X15MM AT 18ATM (NC BALLOON) AND 3.5X15MM AT 18ATM (SCORING BALLOON). MEDTRONIC EBU 3.0 GUIDE CATHETER WAS USED WITH GOOD SUPPORT. SION BLUE GUIDEWIRE CROSSED INTO LAD. TWO BIOFREEDOM STENTS (BFR1-3036 & BFR1-3528) WAS IMPLANTED, IN AN OVERLAPPING MANNER. POST-DILATATION WAS DONE USING NC 3.5X15MM BALLOON AT 18ATM AND NC 4.0X15MM BALLOON AT 16 ATM. INTRAVASCULAR ULTRASOUND (IVUS) WAS NOT PERFORMED. NO DIFFICULTIES DURING DELIVERY, STENT WAS PROPERLY ASSESSED ANGIOGRAPHICALLY AND DETERMINED APPOSED TO THE WALL. POST-PROCEDURE % STENOSIS WAS 0% BY VISUAL ESTIMATION AND TIMI FLOW GRADE WAS III. PATIENT POST-PROCEDURE ADMITTED TO WARD. SHE DEVELOPED CHEST PAIN AT 04:00 A.M. ON (B)(6) 2025 AND HAD CARDIOGENIC SHOCK. PATIENT WAS RESUSCITATED, HOWEVER PATIENT SUCCUMBED AT 08:00 A.M. LACK OF EFFICACY/WORSENING OF TREATED CONDITION AND PATIENT DIED ON 14 JAN 2025. CAUSE OF DEATH FROM THE CLINICAL PRESENTATION AND ECG CHANGES (NEW ONSET RBBB AND ST ELEVATION AT AVR) LIKELY REPRESENT MYOCARDIAL INFARCTION INVOLVING LM STENT. Manufacturer Narrative: THE EVENT-RELATED STENTS WAS IMPLANTED IN THE PATIENT, THUS NOT RETURNED FOR BIOSENSORS INVESTIGATION. BASED ON THE EVENT DESCRIBED IN CONJUNCTION WITH ANGIOGRAPHY RECORD PROVIDED TO BIOSENSORS: IT APPEARS THAT THIS IS A CASE OF STENT THROMBOSIS, LEADING TO A MYOCARDIAL INFARCTION AND THE PATIENT'S DEMISE AFTER THE PLACEMENT OF BIOFREEDOM 3.00 X 36 MM AND BIOFREEDOM 3.50 X 28 MM STENTS. THERE IS A POSSIBILITY THAT STENT THROMBOSIS RESULTED FROM DISSECTION AT THE STENT EDGE, EITHER PROXIMALLY OR DISTALLY IN THE PROCEDURE. MOREOVER, THE ANTICOAGULANT ADMINISTRATION MAY HAVE BEEN INADEQUATE. THERE IS NO EVIDENCE TO SUGGEST THAT THE REPORTED COMPLICATION IS RELATED TO THE CHARACTERISTICS OF THE STENT. BIOSENSORS ANALYSIS OF THE EVENT IS CONCLUDED THAT DEVICE HISTORY RECORD (DHR) AND MANUFACTURING PROCESS DID NOT SHOW ANY EVIDENCE OF PRODUCT DEFICIENCY. THE STENT THROMBOSIS AND MYOCARDIAL INFARCTION EVENTS ARE RECOGNIZED POTENTIAL HAZARDOUS SITUATION AND DOCUMENTED IN MANUFACTURER'S PRODUCT ANALYSIS AND INSTRUCTION FOR USE. THE RISKS ARE ASSESSED AND IT IS ACCEPTABLE AS BENEFITS OUTWEIGH RESIDUAL RISK. THE EVENT IS MORE LIKELY RELATED TO PATIENT AND/OR PROCEDURAL RELATED FACTORS RATHER THAN THE DEVICE'S FAULT. THERE IS NO EVIDENCE OF DEVICE OR MANUFACTURING ACTIVITIES DEFICIENCY IDENTIFIED. ACCORDINGLY, THERE IS NO CORRECTION OR CORRECTIVE ACTION TO BE IMPLEMENTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21353491&pc=NVN | 1000165971-2025-00089 | 2024-11-22 05:00 | Death | MICROPORT CRM S.R.L. | 2025-02-11 05:00 | NVN | VEGA | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P130010 |  | 1 | Event Description: REPORTEDLY, THE PATIENT, HAD DUAL CHAMBER PM (PM + 2 LEADS) SYSTEM IMPLANTED (B)(6) 2024, PATIENT WAS ADMITTED AGAIN ON (B)(6) 2024. THE PATIENT DIED ON (B)(6) 2025, FROM SEPSIS\SYSTEMIC INFECTION. THE PM SYSTEM WAS NOT EXPLANTED IN ECHO REPORT IT IS DESCRIBED THAT SOME ¿THIN STRINGS¿ ARE VISIBLE BUT NO HARD EVIDENCE FOR ENDOCARDITIS. NO POST-MORTEM AUTOPSY PERFORMED NO RECENT INTERROGATION OF PM DONE AS OTHER HEALTH ISSUES WERE PRIORITIZED ALSO PM WAS NOT THE TROUBLEMAKER AND NO MALFUNCTION WAS SUSPECTED. Manufacturer Narrative: ALL THE DEVICES HAVE BEEN STERILIZED AND RELEASED ACCORDING TO ALL APPLICABLE PROCEDURES. BOREA DR 1500 (SN: (B)(6)) WAS SOLD AND LABELED AS STERILE ARE MANUFACTURED, STERILIZED AND RELEASED ACCORDING TO APPLICABLE STANDARD OPERATING PROCEDURES. MANUFACTURING DATE: 02 JULY 2024 USE BEFORE DATE: 02 JANUARY 2027 STERILIZATION LOT: S0690 - 3775 VEGA R52 (SN: (B)(6)) WAS SOLD AND LABELED AS STERILE ARE MANUFACTURED, STERILIZED AND RELEASED ACCORDING TO APPLICABLE STANDARD OPERATING PROCEDURES. MANUFACTURING DATE: 12 DECEMBER 2022 USE BEFORE DATE: 16 NOVEMBER 2025 STERILIZATION LOT: 25131-S0578 VEGA R58 (SN: (B)(6)) WAS SOLD AND LABELED AS STERILE ARE MANUFACTURED, STERILIZED AND RELEASED ACCORDING TO APPLICABLE STANDARD OPERATING PROCEDURES. MANUFACTURING DATE: 10 JUNE 2024 USE BEFORE DATE: 23 MAY 2027 STERILIZATION LOT: 30659-S0683 IT SHOULD BE NOTED THAT INFECTION IS A WELL-KNOWN COMPLICATION OF CARDIAC PACING/DEFIBRILLATION SYSTEMS, WHICH IS WELL DESCRIBED IN THE LITERATURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21353497&pc=NVN | 1000165971-2025-00090 | 2024-11-22 05:00 | Death | MICROPORT CRM S.R.L. | 2025-02-11 05:00 | NVN | VEGA | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P130010 |  | 1 | Event Description: REPORTEDLY, THE PATIENT, HAD DUAL CHAMBER PM (PM + 2 LEADS) SYSTEM IMPLANTED (B)(6) 2024, PATIENT WAS ADMITTED AGAIN ON (B)(6) 2024. THE PATIENT DIED ON (B)(6) 2025, FROM SEPSIS\SYSTEMIC INFECTION. THE PM SYSTEM WAS NOT EXPLANTED IN ECHO REPORT IT IS DESCRIBED THAT SOME ¿THIN STRINGS¿ ARE VISIBLE BUT NO HARD EVIDENCE FOR ENDOCARDITIS. NO POST-MORTEM AUTOPSY PERFORMED NO RECENT INTERROGATION OF PM DONE AS OTHER HEALTH ISSUES WERE PRIORITIZED ALSO PM WAS NOT THE TROUBLE MAKER AND NO MALFUNCTION WAS SUSPECTED. Manufacturer Narrative: ALL THE DEVICES HAVE BEEN STERILIZED AND RELEASED ACCORDING TO ALL APPLICABLE PROCEDURES. BOREA DR 1500 (SN:(B)(6) WAS SOLD AND LABELED AS STERILE ARE MANUFACTURED, STERILIZED AND RELEASED ACCORDING TO APPLICABLE STANDARD OPERATING PROCEDURES. MANUFACTURING DATE: 02 JULY 2024 USE BEFORE DATE: 02 JANUARY 2027 STERILIZATION LOT: S0690 - 3775 VEGA R52 (SN:(B)(6) WAS SOLD AND LABELED AS STERILE ARE MANUFACTURED, STERILIZED AND RELEASED ACCORDING TO APPLICABLE STANDARD OPERATING PROCEDURES. MANUFACTURING DATE: 12 DECEMBER 2022 USE BEFORE DATE: 16 NOVEMBER 2025 STERILIZATION LOT: 25131-S0578 VEGA R58 (SN: (B)(6) WAS SOLD AND LABELED AS STERILE ARE MANUFACTURED, STERILIZED AND RELEASED ACCORDING TO APPLICABLE STANDARD OPERATING PROCEDURES. MANUFACTURING DATE: 10 JUNE 2024 USE BEFORE DATE: 23 MAY 2027 STERILIZATION LOT: 30659-S0683 IT SHOULD BE NOTED THAT INFECTION IS A WELL-KNOWN COMPLICATION OF CARDIAC PACING/DEFIBRILLATION SYSTEMS, WHICH IS WELL DESCRIBED IN THE LITERATURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21353527&pc=LJT | 3006260740-2025-00517 | 2025-01-30 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-02-11 05:00 | LJT | POWERPORT CLEARVUE SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | Fracture | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K122899 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE; THE PORT CATHETER HAD FRACTURED. IT WAS FURTHER REPORTED THAT PATIENT WAS EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE; THE PORT CATHETER HAD FRACTURED. IT WAS FURTHER REPORTED THAT PATIENT WAS EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. HOWEVER, THE CURRENT STATUS OF THE PATIENT IS UNKNOWN. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE SMDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THE REPORTED CATHETER FRACTURE AND DEATH THEREFORE CANNOT BE CONFIRMED. NO CAUSAL RELATIONSHIP BETWEEN THE CATHETER FRACTURE OR THE PATIENT DEATH HAS BEEN ESTABLISHED. BASED ON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE OF THE CATHETER FRACTURE AND DEATH IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. G3, H6 (COMPONENT). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21353557&pc=BZD | 2518422-2025-014277 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-02-11 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation; Cardiovascular Insufficiency; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21354179&pc=OTD | 9612164-2025-00708 |  | Death | MEDTRONIC IRELAND | 2025-02-11 05:00 | OTD | ENDO ANCHOR SYSTEM - HELI-FX AAA | Adverse Event Without Identified Device or Use Problem | Rupture | K140036 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH AN ABDOMINAL AORTIC ANEURYSM UNDERWENT A PROCEDURE INVOLVING A HELI-FX ENDOANCHORING SYSTEM ON AN UNKNOWN DATE. THE PATIENT SUBSEQUENTLY DIED, AND THE CAUSE OF DEATH IS CURRENTLY UNKNOWN. IT WAS SUSPECTED THAT THE DEATH COULD BE LINKED TO ENDOANCHORS RUPTURING THE AORTA A FEW DAYS POST PROCEDURE. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED; IT WAS CONFIRMED THAT THE ENDOANCHORS WERE USED PROPHYLACTICALLY AND NOT TO TREAT AN ENDOLEAK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21354464&pc=NPT | MW5165894 | 2025-01-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-11 05:00 | NPT | EVOLUT FX |  | Low Blood Pressure/ Hypotension; Aortic Valve Insufficiency/ Regurgitation |  |  | 1 | Event Description: A TRANSFEMORAL TAVR PROCEDURE WAS ATTEMPTED TO BE PERFORMED WITH A NON-(B)(6) 29MM EVOLUT FX+ VALVE. THE VALVE ENFOLDED AND WAS RECAPTURED AND REMOVED. THE PATIENT BECAME HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY. A SECOND NON-(B)(6) 29MM EVOLUT FX+ VALVE WAS PREPPED AND WOULD NOT CROSS THE VALVE DURING CPR. CPR CONTINUED FOR APPROXIMATELY 20 MINUTES. AN (B)(6) 23MM SAPIEN 3 ULTRA VALVE WAS PREPPED AND IMPLANTED WITHOUT ISSUE AND NO LEAKS. THE IMPLANT WAS CONSIDERED A BAILOUT IMPLANT AS IT TOOK PLACE DURING AN ACTIVE CODE EVENT WITH ACTIVE RESUSCITATION EFFORTS FROM THE FAILED EVOLUT VALVE IMPLANT ATTEMPTS. THE PATIENT DID NOT RECOVER AND EXPIRED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5165895.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21354487&pc=NPT | MW5165895 | 2025-01-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-11 05:00 | NPT | EVOLUT FX |  | Cardiac Arrest; Low Blood Pressure/ Hypotension; Failure of Implant; Aortic Valve Insufficiency/ Regurgitation |  |  | 1 | Event Description: A TRANSFEMORAL TAVR PROCEDURE WAS ATTEMPTED TO BE PERFORMED WITH A NON-(B)(6) 29MM EVOLUT FX+ VALVE. THE VALVE ENFOLDED AND WAS RECAPTURED AND REMOVED. THE PATIENT BECAME HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY. A SECOND NON-(B)(6) 29MM EVOLUT FX+ VALVE WAS PREPPED AND WOULD NOT CROSS THE VALVE DURING CPR. CPR CONTINUED FOR APPROXIMATELY 20 MINUTES. AN (B)(6) 23MM SAPIEN 3 ULTRA VALVE WAS PREPPED AND IMPLANTED WITHOUT ISSUE AND NO LEAKS. THE IMPLANT WAS CONSIDERED A BAILOUT IMPLANT AS IT TOOK PLACE DURING AN ACTIVE CODE EVENT WITH ACTIVE RESUSCITATION EFFORTS FROM THE FAILED EVOLUT VALVE IMPLANT ATTEMPTS. THE PATIENT DID NOT RECOVER AND EXPIRED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5165894.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21354509&pc=CAW | 2518422-2025-011992 | 2023-10-03 04:00 | Death | RESPIRONICS, INC. | 2025-02-11 05:00 | CAW | SIMPLYGO SYSTEM-INTERNATIONAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K111885 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF SIMPLYGO SYSTEM-INTERNATIONAL ALLEGING DEATH. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CA BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21354858&pc=NPT | 2025587-2025-01148 | 2024-07-22 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-11 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem; Device Stenosis | Arrhythmia; Ischemia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 18 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED DUE TO A NON-OCCLUSIVE MESENTERIC ISCHEMIA AND ARRHYTHMIA. PER THE PHYSICIAN, THIS WAS CAUSED BY AORTIC VALVE STENOSIS. Manufacturer Narrative: UPDATED DATA: D4 H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21354948&pc=MNS | 2518422-2025-014361 | 2023-12-07 05:00 | Death | RESPIRONICS, INC. | 2025-02-11 05:00 | MNS | DREAMSTATION BIPAP AUTOSV | Degraded; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP THAT THE PATIENT ALLEGES THAT THE DEVICE IS HAVING A HUMIDIFIER ERROR DISPLAYED. A DEVICE WAS RETURNED TO A MANUFACTURER/THIRD-PARTY SERVICE CENTER. DURING THE EVALUATION OF THE DEVICE, THEY FOUND OUT THAT THE COMPLAINT WAS CONFIRMED, HEATER PLATE WITH ELECTRICAL DAMAGE, SCRATCHED UI BEZEL PLUS. DEVICE IS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP THAT THE PATIENT ALLEGES THAT THE DEVICE IS HAVING A HUMIDIFIER ERROR DISPLAYED. A DEVICE WAS RETURNED TO A MANUFACTURER / THIRD-PARTY SERVICE CENTER. DURING THE EVALUATION OF THE DEVICE, THEY FOUND OUT THAT THE COMPLAINT WAS CONFIRMED, HEATER PLATE WITH ELECTRICAL DAMAGE, SCRATCHED UI BEZEL PLUS. DEVICE IS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE CORRECT B5 SHOULD BE- THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION BIPAP AUTOSV DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. IN THE PREVIOUSLY SUBMITTED REPORT THE REPORTER COUNTRY INFORMATION WAS MISSING, IN THIS REPORT THE INFORMATION HAS BEEN UPDATED AND CORRECTED. APPROPRIATE CODE UPDATED/CORRECTED. IN THE INITIAL REPORT IT WAS WRONGLY REPORTED UNDER RECALL, CORRECTED IN THIS FOLLOW-UP REPORT. THIS REPORT IS NOT A PART OF RECALL. Manufacturer Narrative: AFTER FURTHER INVESTIGATION, IT HAS BEEN DETERMINED THAT THE REPLACEMENT DEVICE WAS RECEIVED BY THE REPORTER AFTER THE PATIENT PASSED AWAY. THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THIS REPORT AND NO INCREASED RISK TO THE INTENDED USER. BASED ON THE INFORMATION AVAILABLE, IT IS A NON-REPORTABLE COMPLAINT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21355096&pc=NPT | 2025587-2025-01152 | 2024-11-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-11 05:00 | NPT | EVOLUT FX VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Ischemia; Vomiting; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 12 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AFTER THE PATIENT WAS DISCHARGED HOME, THE PATIENT PRESENTED TO THE EMERGENCY ROOM AFTER EXPERIENCING VOMITING. A COMPUTED TOMOGRAPHY SCAN AND EXAMINATION REVEALED NON-OCCLUSIVE INTESTINAL ISCHEMIA. THE PATIENT WAS HOSPITALIZED AND TREATED FOR THE ISCHEMIA, HOWEVER, THE PATIENT'S CONDITION WORSENED. THE PATIENT PASSED AWAY IN THE HOSPITAL. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE NON-OCCLUSIVE INTESTINAL ISCHEMIA.Event Description: IT WAS REPORTED THAT APPROXIMATELY 12 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AFTER THE PATIENT WAS DISCHARGED HOME, THE PATIENT PRESENTED TO THE EMERGENCY ROOM AFTER EXPERIENCING VOMITING. A COMPUTED TOMOGRAPHY SCAN AND EXAMINATION REVEALED NON-OCCLUSIVE INTESTINAL ISCHEMIA (NOMI). THE PATIENT WAS HOSPITALIZED AND TREATED FOR THE ISCHEMIA, HOWEVER, THE PATIENT'S CONDITION WORSENED. THE PATIENT PASSED AWAY IN THE HOSPITAL. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE NON-OCCLUSIVE INTESTINALISCHEMIA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PATIENT WAS RECEIVING HEMODIALYSIS TREATMENT. PER THE PHYSICIAN, THIS CAUSED THE NOMI. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT RECEIVED SHOWS THERE IS NO INFORMATION TO REASONABLY SUGGEST THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. NON-OCCLUSIVE MESENTERIC ISCHEMIA (NOMI). NOMI IS MOST COMMONLY DUE TO PRIMARY MESENTERIC ARTERIAL VASOCONSTRICTION AND OFTEN INVOLVES ARTERIAL SPASM. IT IS OFTEN ASSOCIATED WITH HEMODIALYSIS PATIENTS, WHICH WAS NOTED IN THIS CASE. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. B5. A SECOND PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE EXPIRATION DATE; SERIAL NUMBER; UDI# E. FACILITY STREET ADDRESS WAS MISSING THE STREET NUMBER H4. DEVICE MFG DATE H6. ADDITIONAL CODES MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21355224&pc=DSQ | 2916596-2025-00577 | 2023-06-01 04:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿CYSTATIN C-BASED ESTIMATED GLOMERULAR FILTRATION RATE IS INDEPENDENTLY ASSOCIATED WITH LONG- TERM MORTALITY IN HEART MATE 3 PATIENTS¿ THAT THE HEARTMATE 3 MAY BE RELATED TO DEATH. 228 PATIENTS WERE ENROLLED IN THE STUDY FROM (AGE60 ± 13YEARS, 15% FEMALE, 21% BLACK, 43% ISCHEMIC, 58% INTERMACS PROFILE 1 OR 2). OVER A MEDIAN FOLLOW-UP OF 3 YEARS (INTERQUARTILE RANGE 1.5 TO 5), DEATH OCCURRED IN 41 (18%) PATIENTS. COMPARED TO SERUM CREATININE (SCR), CYSTATIN C (CYSC) RECLASSIFIED 61% OF THE PATIENTS TO A DIFFERENT ESTIMATED GLOMERULAR FILTRATION RATE (EGFR) CATEGORY, LEADING TO A HIGHER PREVALENCE OF EGFR <60ML/MIN/1.73 M2. EGFR WAS ASSOCIATED WITH MORTALITY ONLY WHEN ASSESSED BY CYSC (HR PER10ML/MIN/1.73M2 DECREASE IN SCR-BASED EGFR (EGFRSCR) 1.15, P=0.2 AND CYSC-BASED EGFR (EGFRCYSC) 1.44,P=0.01). THE STUDY INVESTIGATED THE ASSOCIATION OF PRE-IMPLANT EGFRSCR AND EGFRCYSC WITH LONG-TERM MORTALITY IN A COHORT OF PATIENTS IMPLANTED WITH HM3. THEY FOUND THAT LOWER PRE-IMPLANT EGFRCYSC, BUT NOT EGFRSCR, IS IN DEPENDENTLY ASSOCIATED WITH INCREASED LONG-TERM MORTALITY IN HM3 RECIPIENTS. THE FINDINGS SUGGEST THAT THIS ALTERNATIVE ESTIMATE OF KIDNEY FUNCTION IMPROVES RISK STRATIFICATION DURING ASSESSMENT OF HM3 CANDIDACY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D4: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN FEB 2016 TO JULY 2023. MONDELLINI, G., PINSINO, A., HYNDS, M., VINOGRADSKY, A., LADANYI, A., SAYER, G., KAKU, Y., URIEL, N., TAKEDA, K., HUSAIN, S., MOHAN, S., COLOMBO, P., & YUZEFPOLSKAYA, M. (2024). CYSTATIN C-BASED ESTIMATED GLOMERULAR FILTRATION RATE IS INDEPENDENTLY ASSOCIATED WITH LONG-TERM MORTALITY IN HEARTMATE3 PATIENTS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S135¿S136. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.275. COLUMBIA UNIVERSITY AND IRVING MEDICAL CENTER, NEW YORK, NY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21355320&pc=FKX | 0008030665-2025-00332 | 2025-01-30 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-02-11 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS THAT THIS PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THE PDRN CONTACTED FRESENIUS FOR ASSISTANCE WITH ACCESSING TREATMENT DATA FOLLOWING THIS PATIENT'S PASSING. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) IN THE INITIAL REPORTING. MULTIPLE ATTEMPTS WERE MADE TO ACQUIRE FURTHER DETAILS CONCERNING THIS PATIENT'S DEATH; HOWEVER, NO ADDITIONAL INFORMATION WAS OBTAINED.Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS THAT THIS PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THE PDRN CONTACTED FRESENIUS FOR ASSISTANCE WITH ACCESSING TREATMENT DATA FOLLOWING THIS PATIENT'S PASSING. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) IN THE INITIAL REPORTING. MULTIPLE ATTEMPTS WERE MADE TO ACQUIRE FURTHER DETAILS CONCERNING THIS PATIENT¿S DEATH; HOWEVER, NO ADDITIONAL INFORMATION WAS OBTAINED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE CYCLER WAS RETURNED FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE. THE HEATER TRAY WAS DISCOLORED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. THE VALVE ACTUATION AND SYSTEM AIR LEAK TESTS WERE PERFORMED AND PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THERE WERE NO MALFUNCTIONS FOUND DURING TESTING THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL REVIEW: MULTIPLE ATTEMPTS WERE MADE TO ACQUIRE FURTHER DETAILS CONCERNING THIS PATIENT¿S DEATH; HOWEVER, NO ADDITIONAL INFORMATION WAS OBTAINED. IN THE INTAKE, IT WAS REPORTED THAT AN OUTPATIENT CLINIC PDRN COULD NOT RETRIEVE FINAL TREATMENT DATA FROM THIS PATIENT¿S LIBERTY SELECT CYCLER FOLLOWING THEIR PASSING. THIS PROVIDED EVIDENCE THAT THE PATIENT¿S DEATH WAS NOT FOCUSED ON THE PERFORMANCE OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S), RATHER IT WAS FOR REQUIRED DATA RETRIEVAL. AS THE PDRN WOULD NOT PROVIDE PATIENT INFORMATION FROM THE OUTPATIENT CLINIC DUE TO THE TECHNICAL ISSUE WITH DATA RETRIEVAL, ADDITIONAL PATIENT-FOCUSED INFORMATION WAS NOT AVAILABLE. WITH LIMITED INFORMATION, THERE WAS NO INDICATION THIS EVENT HAD A TEMPORAL OR CAUSAL RELATIONSHIP TO PD THERAPY OR THAT IT INVOLVED THE PERFORMANCE OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). PRESENTLY THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A SERIOUS INJURY, PATIENT DEATH, OR OTHER ADVERSE EVENT(S) OCCURRED RELATED TO A FRESENIUS PRODUCT(S) OR DEVICE(S). ADDITIONALLY, THERE IS NO ALLEGATION A FRESENIUS PRODUCT(S) OR DEVICE(S) DEFICIENCY OR MALFUNCTION CAUSED OR CONTRIBUTED TO THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356046&pc=QEZ | 1319211-2025-00002 | 2025-01-14 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-11 05:00 | QEZ | ALPHAVAC | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SCIENTIFIC AND CLINICAL AFFAIRS MANAGER REPORTED A PATIENT EXPIRATION, OCCURING AFTER USE OF AN ALPHAVAC MULTIPURPOSE MECHANICAL ASPIRATION 25F / C180. THE FOLLOWING WAS REPORTED: "PATIENT SCHEDULED FOR ALPHAVAC PROCEDURE FOR RA THROMBUS AND PE. PRECASE PLAN INCLUDED USING F22180 FOR THE RA THROMBUS AND THEN EXCHANGING FOR F1885 IF CHOOSING TO PROCEED WITH THE PE PORTION. AT PROCEDURE START PATIENT WAS SEDATED AND INTUBATED BY ANESTHESIA WITH VITALS NOTED TO BE HR 120S, NIBP 120S/70S. O2 SATS 97%. ACCESS WAS RFV WITH 26FR. GORE SHEATH WITH NO APPRECIABLE ISSUES. F22180 DEVICE WAS INTRODUCED INTO THE RA OVER THE AMPLATZ SS WIRE AND ASPIRATION WAS ATTEMPTED WITH TEE GUIDANCE. SOME MATERIAL WAS ENGAGED AND REMOVED WITH THE DEVICE. SOME MATERIAL WAS THEN ENGAGED AND NOT ABLE TO BE REMOVED DUE TO THE SIZE AND DENSITY OF THE MATERIAL. WE ATTEMPTED TO REMOVE THE MATERIAL WITH NEGATIVE PRESSURE LOCK ON THE DEVICE THROUGH THE GORE SHEATH. THE MATERIAL WAS REMOVED FROM THE HEART BUT BECAME TRAPPED IN THE LOWER PART OF THE IVC. THE ALPHAVAC DEVICE WAS THEN REMOVED AND NOT RE-INSERTED FOR THE REMAINDER OF THE PROCEDURE. TEE SHOWED NO MATERIAL IN THE RA AND LIVE TEE WAS MAINTAINED TO ASSESS THIS FOLLOWING. THE GORE SHEATH WAS EXCHANGED FOR AN INARI SHEATH AND CANNULA AND ASPIRATION WAS ATTEMPTED AGAIN. THIS PROVED UNSUCCESSFUL AS WELL AND THE CLOTRIEVER XL DEVICE WAS THEN INTRODUCED INTO THE IVC TO CAPTURE AND REMOVE THE MATERIAL. THIS WAS SUCCESSFUL ON THE FIRST PASS WITH THE MATERIAL IN THE CAVA BEING REMOVED AND VENOGRAM NOW SHOWED A CLEAR IVC PER THE PHYSICIAN. CONTRAST INJECTED MOVED RATHER SLOWLY THROUGH THE IVC AND A SMALL AMOUNT OF AIR WAS NOTED IN THE IVC. THEY INSERTED THE INARI CANNULA AGAIN AND ASPIRATED THIS AIR. THEY CHOSE TO PROCEED WITH THE PE PART OF THE CASE USING A BALLOON CATHETER TO CROSS THE TV AND INTO THE PA. ARTERIOGRAM WAS PERFORMED ON THE LEFT PA AND SHOWED THROMBUS PRESENT. THEY CHOSE TO PROCEED WITH THE INARI DEVICE SINCE IT WAS ALREADY OPEN AND PERFORMED 2 ASPIRATION ATTEMPTS WITH NO MATERIAL REMOVED ON INSPECTION. THE PATIENT'S BLOOD PRESSURE (NIBP) WAS NOW NOTED TO BE LOWER AT 70S/50S. THEY PAUSED THE PROCEDURE AND ATTEMPTED TO REASSESS THE BLOOD PRESSURE BUT WERE UNABLE TO OBTAIN A NEW ONE. PATIENT'S HEART RATE HAD ALSO DECLINED TO AROUND 100 NOW. UNABLE TO OBTAIN A GOOD BLOOD PRESSURE AND WITH A DECLINING HR INTO THE 70S THEY BEGAN CPR. THEY WERE UNABLE TO RESUSCITATE THE PATIENT AND ULTIMATELY HE EXPIRED. POST CONFERENCE WITH PHYSICIAN, HE WAS UNABLE TO PINPOINT EXACTLY WHAT MAY HAVE CAUSED THE EVENT BUT DID NOT BELIEVE OUR DEVICE'S USE AT THE BEGINNING OF THE PROCEDURE HAD CONTRIBUTED TO THIS. PRODUCT WAS NOT COLLECTED AND RETURNED DUE TO NO DEVICE FAILURE BEING NOTED. " THE PROCEDURE WAS PERFORMED TO THE DIRECTIONS FOR USE (DFU) AND DR. KNAPP HAS EXTENSIVE INTERVENTIONAL RADIOLOGY EXPERIENCE. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT EMBOLIZATION AND SUBSEQUENT EXPIRATION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). THERE WERE NO REPORTS OF ALPHAVAC DEVICE MALFUNCTION DURING THE PROCEDURE; THEREFORE, THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE PHYSICIAN REPORTED THEY WERE UNABLE TO PINPOINT EXACTLY WHAT MAY HAVE CAUSED THE EVENT BUT DID NOT BELIEVE THE ALPHAVAC DEVICE'S USE AT THE BEGINNING OF THE PROCEDURE HAD CONTRIBUTED TO THE SAE. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: DISTAL EMBOLIZATION OF THROMBUS. PULMONARY EMBOLISM. CARDIAC ARREST. DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4). Manufacturer Narrative: THE DEVICE WAS DISPOSED; THEREFORE, IT WILL NOT BE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356493&pc=DSQ | 2916596-2025-00990 | 2021-11-01 04:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | THORATEC¿¿ HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: THE ARTICLE, GUIDELINE-DIRECTED MEDICAL THERAPY (GDMT) USE IN A CONTEMPORARY MULTICENTER COHORT OF HM3 LVAD RECIPIENTS, OUTLIES THE BP-LVAD STUDY, WHICH CONDUCTED A PROSPECTIVE, OBSERVATIONAL INVESTIGATION OF MEDICAL MANAGEMENT OF BLOOD PRESSURE (BP) AND ASSOCIATION WITH CLINICAL OUTCOMES IN HM3 LVAD RECIPIENTS AT 4 TERTIARY CENTERS. STUDIED WAS GDMT USE AT INDEX DISCHARGE, AND 3-MONTHS AFTER DISCHARGE. ENROLLED AT THE TIME OF ARTICLE PUBLICATION WERE 111 HM3 LVAD SUBJECTS (AGE 58±12 YEARS, 26% FEMALE, 55% WHITE, 35% BLACK, 73% INTERMACS PROFILES 1 OR 2) WHICH WERE IMPLANTED BETWEEN 11/2021 - 8/2023 AND DISCHARGED. 2 SUBJECTS DIED BEFORE 3-MONTHS. MEAN DOPPLER BP WAS 82.7 ± 9.2 MMHG AT DISCHARGE AND 83.9 ± 9.1 MMHG AT 3-MONTHS. AT DISCHARGE, THE MOST FREQUENTLY PRESCRIBED MEDICATIONS WERE ALDOSTERONE ANTAGONIST (40 % OF SUBJECTS), BETA-BLOCKER (34%), ANGIOTENSIN II RECEPTOR BLOCKER (20%), AND SODIUM-GLUCOSE COTRANSPORTER-2 INHIBITOR (SGLT2I; 18%) (FIGURE, LEFT). AT 3-MONTHS, THE MOST COMMON MEDICATIONS WERE ALDOSTERONE ANTAGONIST (47%), BETA-BLOCKER (42%), SGLT2I (27%), AND ANGIOTENSIN RECEPTOR-NEPRILYSIN INHIBITOR (ARNI; 21%). CONCLUSION: GDMT USE AT DISCHARGE WAS LOW WITH =40 % OF SUBJECTS RECEIVING EACH OF THE 4GDMT CLASSES. THE NUMBER OF SUBJECTS RECEIVING SGLT2I AND ARNI WAS HIGHER AT 3MONTHS AFTER DISCHARGE. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS: 01NOV2021 AS PATIENTS WERE IMPLANTED BETWEEN (B)(6) 2021 AND (B)(6) 2023. AUTHOR INFORMATION: ZHAO, M., CEVASCO, M., BILLARD, H., DEVORE, A., MILANO, C., SHEIKH, F., GENUARDI, M., BLUE, L., DOUGHERTY, K., MARBLE, J., KURCIK, K., HALLMAN, M., SHARKOSKI, T., & VIDULA, H. (2024). GUIDELINE-DIRECTED MEDICAL THERAPY USE IN A CONTEMPORARY MULTICENTER COHORT OF HM3 LVAD RECIPIENTS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S394. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1275. 1 UNIVERSITY OF PENNSYLVANIA, PHILADELPHIA, PA. 2 HOSPITAL OF THE UNIVERSITY OF PENNSYLVANIA, PHILADELPHIA, PA. 3 DUKE UNIV MED CTR, DURHAM, NC. 4 DUKE UNIVERSITY, DURHAM, NC. 5 DUKE UNIVERSITY MEDICAL CENTER, DURHAM, NC. 6 MEDSTAR HEART AND VASCULAR INSTITUTE, GEORGETOWN UNIVERSITY SCHOOL OF MEDICINE, WASHINGTON DC, DC. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE STUDY SUGGESTS THAT LONGER FOLLOW-UP IS NEEDED TO DETERMINE THE EFFECTS OF GDMT ON CLINICAL OUTCOMES AND THE OPTIMAL MEDICAL REGIMEN FOR HM3 LVAD RECIPIENTS. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356499&pc=DSQ | 2916596-2025-00743 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED.Event Description: IT WAS REPORTED THAT THE MULTIPLE ORGAN FAILURE WAS DUE TO BACTEREMIA, AND IT WAS LIKELY FROM A GASTROINTESTINAL SOURCE. IT WAS STATED THAT THE MULTIORGAN FAILURE WAS DUE TO DISEASE PROGRESSION, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: H6-3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED INFECTION COULD NOT BE DETERMINED. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B, AND THE PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, AND HEPATIC DYSFUNCTION), AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356547&pc=MNS | 2518422-2025-014431 | 2024-04-10 04:00 | Death | RESPIRONICS, INC. | 2025-02-11 05:00 | MNS | DREAMSTATION BIPAP AUTOSV | Degraded | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION BIPAP AUTOSV DEVICE WITH AN ALLEGATION STATING THAT THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO DEATH. THERE WAS NO SERIOUS PATIENT HARM OR INJUURY. THERE WAS NO MEDICAL INTERVENTION REPORTED. THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. Manufacturer Narrative: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION BIPAP AUTOSV DEVICE WITH AN ALLEGATION STATING THAT THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO DEATH. THERE WAS NO SERIOUS PATIENT HARM OR INJUURY. THERE WAS NO MEDICAL INTERVENTION REPORTED. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED AS PRODUCT PROBLEM (IN THE TYPE OF REPORTED COMPLAINT) WHICH SHOULD HAVE BEEN FILED AS DEATH ONLY, AND THERE WAS NO REPORT OF SERIOUS OR PERMANENT PATIENT HARM OR INJURY. IN THIS REPORT, BOX H: TYPE OF REPORTED COMPLAINT HAS BEEN UPDATED CORRECTLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356827&pc=MVK | 3015185344-2025-00036 | 2025-02-07 05:00 | Death | KESTRA MEDICAL TECHNOLOGIES, INC. | 2025-02-11 05:00 | MVK | ASSURE WEARABLE CARDIOVERTER DEFIBRILLATOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200037 |  | 1 | Event Description: PATIENT'S CAREGIVER CALLED IN AS THE PATIENT PASSED AWAY OVERNIGHT AT AROUND 2AM WHILE WEARING THE WCD SYSTEM. CAREGIVER DID NOT INDICATE CAUSE OF DEATH, BUT EMS WAS DISPATCHED AND REMOVED THE SYSTEM FROM THE PATIENT. MDR FILED DUE TO REPORTED PATIENT DEATH. WCD ROLE IN PATIENT DEATH IS PENDING ADDITIONAL MEDICAL INFORMATION AND PENDING EVALUATION OF TO-BE-RETURNED DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356871&pc=NPT | 2025587-2025-01170 | 2024-12-12 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-11 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Renal Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 20 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED AT A DIFFERENT HOSPITAL. THE CAUSE OF DEATH WAS NOTED TO BE CARDIAC ARREST DUE TO GENERAL WEAKNESS AND HIGH POTASSIUM CAUSED BY RENAL FAILURE. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21356965&pc=NPT | 2015691-2025-00938 | 2025-01-24 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-11 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning; Inflation Problem; Difficult to Remove | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY OUR FRENCH AFFILIATES, DURING IMPLANT OF A 20MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSAORTIC APPROACH, THE COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE WAS STUCK IN VALVE CUSP. THE BALLOON WAS INFLATED WITH 2 ML AND THE DISTAL PART OF THE VALVE WAS DEPLOYED. THE COMMANDER WITH VALVE WAS STUCK SO THE PATIENT WAS CONVERTED TO SURGERY. THE SURGERY WAS COMPLETED WITH GOOD RESULTS, BUT PATIENT PASSED AWAY THE FOLLOWING DAY.Event Description: AS REPORTED BY OUR FRENCH AFFILIATES, DURING IMPLANT OF A 20MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSAORTIC APPROACH, DURING THE PROCEDURE THE VALVE WAS CORRECTLY ALIGNED BETWEEN THE DISTAL AND PROXIMAL MARKERS ON THE INFLATION BALLOON, BUT THE COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE WAS STUCK IN THE VALVE CUSP AND IT WAS UNABLE TO CROSS THE NATIVE ANNULUS. THE BALLOON WAS INFLATED WITH 2 ML TO FACILITATE CROSSING THE NATIVE ANNULUS; HOWEVER, THE DISTAL PART OF THE VALVE WAS DEPLOYED AND THE COMMANDER WITH VALVE WAS STUCK. THE PATIENT WAS CONVERTED TO SURGERY AND IT WAS COMPLETED WITH GOOD RESULTS. HOWEVER, PATIENT PASSED AWAY THE FOLLOWING DAY.Event Description: AS REPORTED BY OUR FRENCH AFFILIATES, DURING IMPLANT OF A 20MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSAORTIC APPROACH, DURING THE PROCEDURE THE VALVE WAS CORRECTLY ALIGNED BETWEEN THE DISTAL AND PROXIMAL MARKERS ON THE INFLATION BALLOON, BUT THE COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE WAS STUCK IN VALVE CUSP AND IT WAS UNABLE TO CROSS THE NATIVE ANNULUS. THE BALLOON WAS INFLATED WITH 2 ML TO FACILITATE CROSSING THE NATIVE ANNULUS, HOWEVER, THE DISTAL PART OF THE VALVE WAS DEPLOYED IN THE ASCENDING AORTA CLOSE TO THE AORTIC ANNULUS AND THE COMMANDER WITH VALVE WAS STUCK. THE PATIENT WAS CONVERTED TO SURGERY AND IT WAS COMPLETED WITH GOOD RESULTS. HOWEVER, PATIENT PASSED AWAY THE FOLLOWING DAY CAUSED BY THE LENGTH OF THE PROCEDURE CONTAINING RESCUE AND SURGERY. Manufacturer Narrative: A FINAL MDR IS BEING SUBMITTED FOR A COMPLETION TO THE ENGINEERING EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED: UPDATE TO B4, B5, G3, G6, H2, H6, H11. THE PRODUCT WAS NOT RETURNED; THEREFORE, A NO PRODUCT RETURN INVESTIGATION WAS COMPLETED. AS A DEVICE WAS NOT RETURNED, VISUAL INSPECTION, FUNCTIONAL TESTING, AND DIMENSIONAL TESTING WERE UNABLE TO BE DONE. THE PROVIDED IMAGERY WAS EVALUATED, AND THE FOLLOWING WAS OBSERVED: PATIENT PRESENTED AN HORIZONTAL AORTA ANATOMY. AS NO DEVICE WAS RETURNED AND THERE IS NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT POTENTIALLY CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. THE COMMANDER DELIVERY SYSTEM WITH S3U IFU WAS REVIEWED, ALONG WITH THE DEVICE PROCEDURAL AND PATIENT SCREENING MANUALS. WARNINGS INCLUDE, 'THE PROCEDURE SHOULD BE CONDUCTED UNDER FLUOROSCOPIC GUIDANCE. SOME FLUOROSCOPICALLY GUIDED PROCEDURES ARE ASSOCIATED WITH A RISK OF RADIATION INJURY TO THE SKIN. THESE INJURIES MAY BE PAINFUL, DISFIGURING, AND LONG-LASTING'. POTENTIAL ADVERSE EVENTS NOTE, 'DEATH' AND 'REOPERATION'. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, IFU WARNINGS AND CAUTIONS, AND PHYSICIAN TRAINING. NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION.' THE COMPLAINTS FOR DIFFICULTY OR INABILITY TO CROSS NATIVE ANNULUS AND/OR BIOPROSTHESIS AND DIFFICULTY OR INABILITY TO WITHDRAW CATHETER FROM DEPLOYED VALVE WERE UNABLE TO BE CONFIRMED. AS THE DEVICE WAS NOT RETURNED, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL EXAMINATION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THEREFORE, THE PRESENCE OF A MANUFACTURING NON-CONFORMANCE WAS UNABLE TO BE DETERMINED. HOWEVER, REVIEW OF THE DHR DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED, 'IT WAS A 20MM SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IMPLANT IN AORTIC POSITION BY TRANSAORTIC APPROACH IN A PATIENT WITH COMPLEX ANATOMY AND AN AORTIC ANGLE AT 110 DEGREES. ['] BUT THE COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE WAS STUCK IN VALVE CUSP AND IT WAS UNABLE TO CROSS THE NATIVE ANNULUS. THE BALLOON WAS INFLATED WITH 2 ML TO FACILITATE CROSSING THE NATIVE ANNULUS; HOWEVER, THE DISTAL PART OF THE VALVE WAS DEPLOYED IN THE ASCENDING AORTA CLOSE TO THE AORTIC ANNULUS AND THE COMMANDER WITH VALVE WAS STUCK'. EDWARDS HAS PROVIDED GUIDELINES AND INSTRUCTIONS TO PHYSICIANS ON VISUALIZATION, ASSESSMENT, AND PRESERVATION OF THE VASCULATURE, ADEQUATE PREPARATION OF THE VASCULATURE, THE OPTIONAL USE OF PREDILATATION OF THE TARGET VALVE, AND PROPER USE OF FLEXION WHILE PERFORMING NAVIGATION THROUGH THE ANATOMY IN THE IFU AND TRAINING MATERIALS/REFRESHER TRAINING. IN ADDITION, EDWARDS PHYSICIAN TRAINING PROGRAM INCLUDES CASE OBSERVATION/REVIEW, PROCTOR QUALIFICATION AND REFRESHER TRAINING. PER EVALUATION OF PROVIDED IMAGERY, THE PATIENT PRESENTED AN HORIZONTAL AORTA ANATOMY. PER TRAINING MANUALS AND SCREENING MANUALS 'HORIZONTAL AORTA' IS IDENTIFIED AS ONE OF THE FACTORS THAT MAY CONTRIBUTE TO THE DIFFICULTIES CROSSING NATIVE VALVE AND ACHIEVE PROPER VALVE POSITIONING. IN THIS CASE, IT IS LIKELY THAT ANGULATED AORTA MAY HAVE PREVENT THE COMMANDER DELIVERY SYSTEM AND CRIMPED VALVE ADVANCE THROUGH THE AORTIC ROOT AND VALVE. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (HORIZONTAL AORTA) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. PER TRAINING MANUALS, ONE OF THE RECOMMENDED TECHNIQUES TO TRY TO RESOLVE THE CROSSING DIFFICULTIES IS ADDING A SMALL AMOUNT OF FLUID TO THE DELIVERY SYSTEM TO INFLATE THE DISTAL BALLOON. IN THIS CASE, THE BALLOON WAS INFLATED WITH 2ML OF FLUID, WHICH CAUSE A PARTIALLY DEPLOYMENT OF THE VALVE IN THE DESCENDING AORTA. SINCE THE VALVE WAS NOT TOTALLY DEPLOYED, IT IS LIKELY THAT PARTIALLY EXPANDED BALLOON MATERIAL BECOMING CAUGHT ON THE THV FRAME. AS SUCH, AVAILABLE INFORMATION SUGGESTS THE PROCEDURAL FACTORS (DEVICE MANIPULATION, BALLOON CAUGHT ON PARTIALLY DEPLOYED THV) MAY HAVE CAUSE OR CONTRIBUTED TO THE REPORTED EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED FROM A FOLLOW UP. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED: UPDATE TO B4, B5, D4, G3, G6, H2, H6, H11. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21357132&pc=DSQ | 3007042319-2025-00485 | 2025-01-29 05:00 | Death | HEARTWARE, INC. | 2025-02-11 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Brain Injury | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A MASSIVE HEAD BLEED AND WAS INTUBATED WITH NO SIGNIFICANT BRAIN FUNCTION AND WAS PRONOUNCED BRAIN DEAD. IT WAS NOTED THAT, MAY PROLONG THINGS WITH TRACH/PERCUTANEOUS ENDOSCOPIC GASTROSTOMY. RECOVERY IS UNLIKELY.Event Description: IT WAS FURTHER REPORTED THAT AFTER A WHILE WITHOUT BRAIN ACTIVITY THE PATIENT WAS PRONOUNCED DEAD AND PLACED ON ORGAN PROCUREMENT SERVICES. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED. UPDATED B5 AND ADDITIONAL CODES. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING A MASSIVE HEAD BLEED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, BLEEDING, NEUROLOGICAL DYSFUNCTION, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF NEUROLOGICAL DYSFUNCTION AND BLEEDING. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: 5076-52, 694765 LEADS IMPLANTED: (B)(6) 2011, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21357263&pc=OZD | 1220648-2025-26093 | 2025-01-22 05:00 | Death | ABIOMED, INC. | 2025-02-11 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED HEMOLYSIS DURING IMPELLA CP SUPPORT. FOUR UNITS OF RED BLOOD CELLS WERE GIVEN TO THE PATIENT TO TREAT THE CONDITION. SUPPORT CONTINUED WITH THE IMPLANTED PUMP FOLLOWING THE INTERVENTION. IT WAS LATER DOCUMENTED THAT THE PATIENT EXPIRED. IT IS UNKNOWN IF THE CONDITION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE DATA LOGS SHOWED PUMP RAN WITHOUT ISSUE DURING SUPPORT. THERE WERE POSITIONING AND SUCTION ALARMS TRIGGERED DURING THE CASE BUT RESOLVED QUICKLY. THE PRODUCT WAS NOT RETURNED FOR REVIEW. THE ROOT CAUSE OF HEMOLYSIS WAS MOST LIKELY PATIENT CONDITION RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21357464&pc=DSQ | 2916596-2025-00994 | 2021-11-01 04:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Sepsis; Cardiogenic Shock; Multiple Organ Failure; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿AORTIC REGURGITATION IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS: DOES AORTIC ROOT DILATATION CONTRIBUTE TO VALVE INCOMPETENCE?¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH DEATH, AORTIC REGURGITATION, MALIGNANCY, SEPTIC MULTIORGAN FAILURE, GASTROINTESTINAL BLEEDING, MYOCARDIAL INFARCTION, AORTIC DISSECTION, CARDIOGENIC SHOCK, STROKE, BLEEDING, AND HEART TRANSPLANTATION. FROM NOVEMBER 2015 TO NOVEMBER 2021, A TOTAL OF 119 PATIENTS UNDERWENT LVAD IMPLANTATION USING HM3. 31 PATIENTS WITH INSUFFICIENT FOLLOW UP (EARLY MORTALITY, INCOMPLETE ECHO DATA, AND LOST TO FOLLOW UP) WERE EXCLUDED FROM THE ANALYSIS. 68 PATIENTS WERE SELECTED FOR THE FINAL ANALYSIS, WITH A MEAN AGE OF 60 +- 10 YEARS AND THEY WERE PREDOMINANTLY MALE. DURING FOLLOW UP, 23 PATIENTS DEVELOPED AT LEAST MODERATE AORTIC REGURGITATION (AR). THE MAJORITY OF PATIENTS SHOWED NON-EXCENTRIC AR WITH A CENTRAL ORIGIN OF THE REGURGITANT JET (DATA NOT SHOWN). ACTUARIAL INCIDENCE WAS 8% AT 1 YEAR, 29% AT 3 YEARS, AND 41% AT 5 YEARS, MEAN FOLLOW-UP TIME UNTIL THE FIRST OCCURRENCE OF AT LEAST MODERATE AR WAS 30 MONTHS. DURING LVAD SUPPORT, 21 PATIENTS DIED. CAUSES OF DEATH WERE MALIGNANCY, SEPTIC MULTIORGAN FAILURE, MYOCARDIAL INFARCTION, GASTROINTESTINAL BLEEDING, AORTIC DISSECTION, OR CARDIOGENIC SHOCK. IN SEVEN INSTANCES, THE CAUSE OF DEATH WAS UNKNOWN. FURTHER ADVERSE EVENTS WERE STROKE IN 5 PATIENTS, AND NON-LVAD RELATED BLEEDING. OF THOSE, 1 STROKE AND 2 BLEEDING EVENTS OCCURRED IN PATIENTS WITH MODERATE OR SEVERE AR DURING FOLLOW UP. 14 PATIENTS UNDERWENT HEART TRANSPLANTATION DURING FOLLOW UP. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. SECTION E: REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING SEPSIS, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, THE IFU ALSO LISTS INFECTION AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5:SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DATE OF EVENT IS APPROXIMATED AS NOVEMBER 1, 2015 AS THE DATA WERE COLLECTED BETWEEN NOVEMBER 2015 AND NOVEMBER 2021. DEVICE WAS IMPLANTED AT THE TIME OF THE EVENT: DEPARTMENT OF CARDIAC SURGERY, SAARLAND UNIVERSITY MEDICAL CENTER, HOMBURG/SAAR, GERMANY. DEPARTMENT OF RADIOLOGY, UNIVERSITY HOSPITAL JENA, FRIEDRICH SCHILLER UNIVERSITY, JENA, GERMANY. DEPARTMENT OF CARDIOTHORACIC SURGERY, UNIVERSITY HOSPITAL JENA, FRIEDRICH SCHILLER UNIVERSITY, JENA, GERMANY. DEPARTMENT OF CARDIAC SURGERY, HERZ-UND KREISLAUFZENTRUM, ROTENBURG AN DER FULDA, GERMANY. FÄRBER G, SCHNEIDER U, GRÄGER S, ELAYAN Y, SCHWAN I, TKEBUCHAVA S, ET AL. AORTIC REGURGITATION IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS: DOES AORTIC ROOT DILATATION CONTRIBUTE TO VALVE INCOMPETENCE? ARTIF. ORGANS. 2025;49:292¿299. HTTPS://DOI.ORG/10.1111/AOR.14873. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21357541&pc=DSQ | 3007042319-2025-00489 | 2025-01-30 05:00 | Death | HEARTWARE, INC. | 2025-02-11 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Unspecified Heart Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) EXPERIENCED A FATAL EVENT. THE CAUSE OF DEATH WAS IDENTIFIED AS RIGHT HEART FAILURE, AND SUPPORT WAS TERMINATED WITH FAMILY MEMBERS PRESENT AT THE BEDSIDE. NO DEVICE ALLEGATIONS WERE MADE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RIGHT VENTRICULAR FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF WORSENING HEART FAILURE. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10. CONCOMITANT PROD. DDBB1D1 - DEFIB IMPLANTED ON (B)(6) 2014 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21357657&pc=FED | 3014683069-2025-00003 | 2024-12-02 05:00 | Death | CALYXO, INC. | 2025-02-11 05:00 | FED | CVAC ASPIRATION SYSTEM | Adverse Event Without Identified Device or Use Problem | Sepsis | K233472 |  | 1 | Event Description: ON (B)(6) 2024, A 66-YEAR-OLD MORBIDLY OBESE PATIENT WITH SIGNIFICANT CO-MORBIDITIES (BMI >50, DIABETES AND HYPERTENSION) UNDERWENT AN URETEROSCOPIC STONE REMOVAL PROCEDURE WITH THE CVAC ASPIRATION SYSTEM TO TREAT MULTIPLE STONES IN THE RIGHT KIDNEY THAT WERE PRESUMED TO BE INFECTIOUS. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE DEVICE. AT AN UNSPECIFIED TIME POST PROCEDURE, THE PATIENT BECAME SEPTIC AND WAS HOSPITALIZED IN THE ICU. IT WAS REPORTED THAT THE PATIENT WAS RESUSCITATED, INTUBATED, AND REQUIRED VASOPRESSOR SUPPORT. THE PATIENT WAS GIVEN INTRAVENOUS ANTIBIOTICS FOR A POSITIVE BLOOD CULTURE AND WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND CONTINUOUS VENO-VENOUS HEMODIALYSIS (CVVHD). ON (B)(6) 2024, THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT REPORTED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE LOT HISTORY REVIEW (LHR) FOR LOT NUMBER P02077 WAS CONDUCTED WHICH CONFIRMED THAT THERE WERE NO NON-CONFORMANCES DURING THE MANUFACTURING PROCESS THAT COULD BE RELATED TO THE REPORTED EVENT. THE REVIEW INDICATED THAT THE DEVICE MET ALL DESIGN AND MANUFACTURING SPECIFICATIONS WHEN RELEASED FOR DISTRIBUTION. IT WAS REPORTED THAT A PATIENT EXPIRED FOLLOWING THE SURE PROCEDURE WITH THE CVAC SYSTEM. NO MALFUNCTION OF THE CVAC SYSTEM WAS REPORTED. THE PATIENT HAD A COMPLEX MEDICAL HISTORY, INCLUDING MORBID OBESITY, DIABETES, AND HYPERTENSION. THE STONES TREATED WERE PRESUMED TO BE INFECTIOUS AND THE PATIENT DEVELOPED POST-OPERATIVE SEPSIS. THE TREATING SURGEON REPORTED THAT THE PATIENT OUTCOME WAS DUE TO "PATIENT SELECTION" MEANING THIS PATIENT WAS VERY SICK AND NOT A GOOD CANDIDATE FOR URS. ADDITIONALLY, THE TREATING SURGEON SHARED IN DISCUSSION WITH THE MANUFACTURER THAT THIS EVENT IS NOT DEVICE-RELATED. SEPSIS IS A COMPLICATION THAT CAN OCCUR FOLLOWING ANY KIDNEY STONE PROCEDURE. OCCURRENCE IS UP TO ~5% (BHOJANI, BJU INT. 2023 AUG;132(2):210-216) FOLLOWING URETEROSCOPY (URS) AND UP TO 4.5% (FALAHATKAR, ASIAN J UROL 2024; 11(2): 253-260) FOLLOWING PCNL. MORTALITY RATES FOLLOWING URS AND PCNL HAVE BEEN REPORTED TO RANGE FROM 0.02 - 0.07% (DE LA ROSETTE, J ENDOUROL 2014; 28(2): 131-139 AND BHOJANI, BJU INT. 2023 AUG;132(2):210-216) AND 0.03 TO 0.7% (DE LA ROSETTE, J ENDOUROL 2011; 25(1): 11-17, UNSAL, UROLOGY 2012; 79(1): 55-60 AND JOHNSTON, J ENDOUROL 2019; 33(9): 704-711), RESPECTIVELY. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS CURRENTLY IN PROGRESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21357957&pc=DSQ | 2916596-2025-01034 | 2023-06-16 04:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH ¿REDUCED LEFT VENTRICULAR ASSIST DEVICE (LVAD) UTILIZATION AND POST-TRANSPLANT SURVIVAL AFTER DONOR ALLOCATION POLICY CHANGE¿ THAT HEARTMATE 3 PATIENTS EXPERIENCED TRANSPLANTATION WHILE SEEKING TO PROVIDE AN ESTIMATE ON WAITLIST AND POST-TRANSPLANT SURVIVAL TRENDS IN PATIENTS SUPPORTED ON DURABLE LEFT VENTRICULAR ASSIST DEVICES IN RELATION TO THE UNITED NETWORK FOR ORGAN SHARING (UNOS) DONOR HEART ALLOCATION POLICY REVISION. PATIENTS IMPLANTED WITH HEARTMATE 3, HEARTMATE II, AND HEARTWARE AND LISTED FOR ISOLATED HEART TRANSPLANTATION BETWEEN 18OCT2013 TO 16JUN2023 WERE IDENTIFIED IN THE UNOS DATABASE AND STRATIFIED INTO PRE AND POST POLICY REVISION (18OCT2018) GROUPS. WAITLIST AND POST-TRANSPLANT SURVIVAL WAS ANALYZED AND COMPARED BETWEEN THESE TWO GROUPS. 7,664 PATIENTS WERE LISTED WITH LVAD (PRE: 4,144, POST: 3,520), AND 7,253 PATIENTS UNDERWENT ISOLATED HEART TRANSPLANT (PRE: 4,853, POST: 2,400). POLICY REVISIONS WERE ASSOCIATED WITH DECREASE IN PATIENTS SUPPORTED ON LVADS (42.3% VS. 19.8%, P<0.01). FOR THOSE LISTED WITH AN LVAD, FEWER PATIENTS EVENTUALLY UNDERWENT TRANSPLANTATION (74% VS. 63.3%, P<0.01). WAITLIST SURVIVAL WAS COMPARABLE AMONG PATIENTS SUPPORTED ON LVADS (P=0.7, FIGURE 1A). FOLLOWING THE POLICY CHANGE, PATIENTS SUPPORTED WITH AN LVAD AT THE TIME OF TRANSPLANT HAD A WORSE SURVIVAL COMPARED TO PRIOR TO THE POLICY CHANGE (P<0.01). Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT IS APPROXIMATED AS 01OCT2021 AS THE DATA WERE COLLECTED BETWEEN 18OCT2013 AND 16JUN2023. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. DIVISION OF CARDIAC SURGERY, (B)(6), UNITED STATES, 2DEPARTMENT OF SURGERY, (B)(6), UNITED STATES, 3DIVISION OF CARDIOLOGY, (B)(6), UNITED STATES. D PRITTING, C., WEBER, M., STOROZYNSKY, E., J EISEN, H., RAME J EDUARDO, ALVAREZ, R., RAJAGOPAL, K., W ENTWISTLE, J., & TCHANTCHALEISHVILI, V. (N.D.). REDUCED LEFT VENTRICULAR ASSIST DEVICE UTILIZATION AND POST-TRANSPLANT SURVIVAL AFTER DONOR ALLOCATION POLICY CHANGE. ASAIO. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED TO END DATE OF DATA COLLECTION SECTION B3: DATE OF DEATH WAS NOT REPORTED. SECTIONS E2 AND E3: CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21358466&pc=FTM | 1000306051-2025-00021 | 2022-01-30 05:00 | Death | LIFECELL | 2025-02-11 05:00 | FTM | STRATTICE RECONSTRUCTIVE TISSUE MATRIX, 20 X 25, FIRM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K070560 |  | 1 | Event Description: ON 27/JAN/2025, PMQA RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHO UNDERWENT A ¿VENTRAL¿ HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT: SP100381-101¿ ON (B)(6) 2016. THE PATIENT UNDERWENT AN ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH¿ ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT: SP100537-078¿ ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN ¿INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT: SP100595." THE PATIENT THEN UNDERWENT AN ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH¿ ON (B)(6) 2018 AND IMPLANTED WITH A NON-ABBVIE DEVICE, ¿XENMATRIX LOT: HUCR0748.¿ THE MESH WAS REVISED DUE TO ¿RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, ¿BARD MESH LOT#: HUCQ0919.¿ THE MESH WAS REVISED ON (B)(6) 2022 DUE TO ¿OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION¿ AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT: SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH¿ ON OR ABOUT ON (B)(6) 2016. THE PATIENT ALSO UNDERWENT A ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE)¿ ON OR ABOUT ON (B)(6) 2017. THE PATIENT WAS TREATED FOR ¿INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE)¿ BETWEEN ON (B)(6) 2018. THE PATIENT UNDERWENT ¿OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN¿ BETWEEN ON (B)(6) 2018. THE PATIENT WAS TREATED FOR ¿RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH¿ ON OR ABOUT ON (B)(6) 2018. THE PATIENT WAS SEEN FOR ¿ABDOMINAL PAIN¿ BETWEEN 2018 AND ON (B)(6) 2022. THE PATIENT WAS TREATED FOR ¿BOWEL OBSTRUCTION¿ BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ¿ABDOMINAL PAIN/ADHESIONS¿ BETWEEN ON (B)(6) 2016 AND ON (B)(6) 2022. THE PATIENT WAS SEEN FOR ¿BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS¿ BETWEEN ON (B)(6) 2016 AND ON (B)(6) 2022. THE PATIENT WAS SEEN FOR ¿HERNIA SYMPTOMS¿ FROM APPROXIMATELY 2018 TO ON (B)(6) 2022. THE PATIENT WAS TREATED ¿OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION¿ BETWEEN ON (B)(6) 2022. THE PATIENT RECEIVED ¿POST-OP REHAB¿ BETWEEN 2018-2019. PER DEATH CERTIFICATE, PATIENT DIED ON (B)(6) 2022. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD FOR STRATTICE LOT: SP200312 HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE EVENT IS A PHYSIOLOGICAL COMPLICATION AND ANALYSIS OF THE DEVICE GENERALLY DOES NOT ASSIST ABBVIE IN DETERMINING A PROBABLE CAUSE FOR THIS EVENT. DUE TO THE LEGAL PROCESS, IF ADDITIONAL INFORMATION IS MADE AVAILABLE DURING LEGAL PROCEEDINGS, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21358667&pc=DSQ | 2916596-2025-00767 | 2025-01-24 05:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, JP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS RUSHED TO THE HOSPITAL IN THE STATE OF CARDIAC ARREST. DEATH WAS CONFIRMED AT THE HOSPITAL.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS FOUND IN A STATE OF CARDIAC ARREST AND EMERGENCY SERVICES WAS CALLED AFTER THE "BATTERY WAS REPLACED". IT WAS NOTED THAT THE CAUSE OF DEATH WAS VENTRICULAR ASSIST DEVICE FAILURE, IT WAS SUSPECTED THAT THE SYSTEM CONTROLLER MALFUNCTIONED AS THE DEVICE DID NOT OPERATE EVEN AFTER CONNECTING A CHARGED BATTERY. IT WAS UNKNOWN IF THE DEVICE OPERATED AS EXPECTED AND THE PUMP WOULD NOT RETURN.Event Description: IT WAS REPORTED THAT IT WAS UNKNOWN IF THE PATIENT WAS CONNECTED TO BATTERY POWER WHILE IN TRANSIT TO THE HOSPITAL. IT WAS UNKNOWN WHAT ACTIONS WERE PERFORMED AT ARRIVAL TO THE HOSPITAL AND AFTER DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A DISTAL PORTION OF THE DRIVELINE WAS RETURNED MEASURING APPROXIMATELY 25¿ IN LENGTH. ELECTRICAL CONTINUITY TESTING OF THE RETURNED PORTION OF THE DRIVELINE DID NOT REVEAL ANY DISCONTINUITIES OR SHORTS. IT WAS REPORTED THAT THE PATIENT WAS RUSHED TO THE HOSPITAL IN A STATE OF CARDIAC ARREST, AND DEATH WAS CONFIRMED AT THE HOSPITAL. IT WAS COMMUNICATED THAT EMERGENCY SERVICES WERE CALLED AFTER THE BATTERY WAS REPLACED, AND IT IS SUSPECTED THAT THE SYSTEM CONTROLLER MALFUNCTIONED AS THE DEVICE DID NOT OPERATE AFTER CONNECTING THE CHARGED BATTERY. THIS EVENT IS FURTHER ADDRESSED UNDER THE SYSTEM CONTROLLER INVESTIGATION. IT WAS COMMUNICATED THAT THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21358668&pc=JDI | 1038671-2025-00834 | 2025-01-24 05:00 | Death | EXACTECH, INC. | 2025-02-11 05:00 | JDI | PROSTHESIS, HIP, SEMI-CONSTRAINED, METAL/POLYMER, CEMENTED | Adverse Event Without Identified Device or Use Problem | Pneumonia; Sepsis |  |  | 1 | Event Description: IT WAS REPORTED THAT AN INDIVIDUAL WAS INFORMED BY CHANNEL 8 NEWS ABOUT ISSUES WITH THE IMPLANT EQUIPMENT. THEIR MOTHER RECEIVED A HIP IMPLANT AT (B)(6) HOSPITAL IN 2017. SHE HAD 5 SURGERIES RELATED TO THAT IMPLANT IN THE SAME CALENDAR YEAR. AS A RESULT OF THAT IMPLANT, SHE CONTRACTED BONE SEPSIS AND SUBSEQUENTLY DIED IN 2020 OF SEPSIS AND PNEUMONIA. I DO NOT KNOW IF YOUR ADVERTISEMENT PERTAINS TO ME AND HER SITUATION OR NOT, BUT I WANTED TO INFORM YOU OF IT. NO FURTHER INFORMATION.Event Description: THIS IS #6 OF 6 REPORTS. ADDITIONAL EVENTS REPORTED UNDER: 1ST - 1038671-2025-00819, 2ND - 1038671-2025-00818, 3RD - 1038671-2025-00816, 4TH - 1038671-2025-00826, 5TH - 1038671-2025-00827. Manufacturer Narrative: H3: THE REPORTED EVENT WAS UNABLE TO BE CONFIRMED DUE TO LIMITED INFORMATION RECEIVED FROM THE CUSTOMER. NO DEVICE WAS RETURNED FOR EVALUATION; FURTHER, PHOTOGRAPHS AND/OR RADIOGRAPH IMAGES WERE NOT PROVIDED FOR REVIEW. OPERATIVE NOTES AND/OR MEDICAL RECORDS WERE NOT PROVIDED FOR REVIEW OF USAGE/ TECHNIQUE. A REVIEW OF MANUFACTURING DATA WAS UNABLE TO BE PERFORMED AS THE LOT INFORMATION OF THE PRODUCT INVOLVED IN THE EVENT WAS NOT AVAILABLE. A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AS THE NECESSARY INFORMATION TO ADEQUATELY INVESTIGATE THE REPORTED EVENT WAS NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: B5.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21359268&pc=CBK | 3016723884-2025-00003 | 2025-01-23 05:00 | Death | HAMILTON MEDCIAL AG | 2025-02-11 05:00 | CBK | HAMILTON-H900 | Contamination | Insufficient Information |  |  | 1 | Event Description: INFANT WAS DIAGNOSED WITH A BACTERIAL INFECTION PSEUDOMONAS AERUGINOSA. IT WAS BELIEVED THAT THE ORIGIN WAS FROM THE HAMILTON H-900 DUE TO THE RAIN OUT. THE H900 WAS LATER DEEMED NOT TO BE THE SOURCE OF THE BACTERIA AND WAS NOT LIABLE. NO CULTURES ON PRODUCT, TESTING WAS COMPLETED (CONTEXT OF TESTING IS BIOBURDEN TESTING).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360264&pc=OZD | 1220648-2025-26097 | 2025-01-22 05:00 | Death | ABIOMED, INC. | 2025-02-11 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE IMPELLA SITE STARTED TO OOZE. TWO UNITS OF PACKED RED BLOOD CELLS WERE TRANSFUSED. IT WAS DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEED COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360548&pc=NGV | 2135147-2025-00684 | 2025-01-19 05:00 | Death | ABBOTT MEDICAL | 2025-02-11 05:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Hemoptysis; Pulmonary Edema | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 31MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT USING A 14F AMPLATZER TORQVUE DELIVERY SYSTEM (LOT: 9115457). COMPUTED TOMOGRAPHY (CT) AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS USED TO DETERMINE THE SIZE. DURING PROCEDURE, AFTER MULTIPLE UNSUCCESSFUL ATTEMPTS TO IMPLANT THE AMULET DUE TO BEING TOO LARGE WITH SUBOPTIMAL POSITION, IT WAS REMOVED. THE 31MM AMULET WAS EXCHANGED FOR A NEW 28MM AMPLATZER AMULET (LOT: 9051988). THE SHEATH WAS NOT EXCHANGED PER PHYSICIAN PREFERENCE AND ASPIRATED. THE 28MM AMULET WAS UNSHEATHED TO BALL, BUT WAS NOT DEPLOYED AS A SMALL THROMBUS WAS OBSERVED AT THE TIP OF THE SHEATH. NO DEPLOYMENTS WERE ATTEMPTED WITH THIS DEVICE AND IT WAS WITHDRAWN FROM THE SHEATH. THE SHEATH WAS AGAIN ASPIRATED, BUT AGAIN NOT EXCHANGED DUE TO PHYSICIAN PREFERENCE. ANOTHER NEW 28MM AMPLATZER AMULET (LOT: 8875300) WAS SELECTED AND ATTEMPTED TO BE IMPLANTED, HOWEVER THIS WAS UNSUCCESSFUL DUE TO SUBOPTIMAL POSITION AND WAS REMOVED FROM THE PATIENT. A NEW TRANSSEPTAL PUNCTURE (TSP) WAS ELECTED TO OPTIMIZE THE APPROACH; THE NEW TSP PROVIDED A MORE FAVORABLE APPROACH. A THIRD NEW 28MM AMPLATZER AMULET (LOT: 8898897) WAS SELECTED FOR IMPLANT USING A NEW 14F AMPLATZER TORQVUE DELIVERY SYSTEM (LOT: 10256632). THE THIRD 28MM AMULET WAS PLACED AND PASSED TENSION TEST. TEE ASSESSMENT WAS COMPLETED. SECOND TENSION WAS PERFORMED TO ENSURE DEVICE STABILITY AND CLOSE CRITERIA WAS ACHIEVED. THE DEVICE WAS SUCCESSFULLY IMPLANTED. THE PATIENT'S CONDITION POST-IMPLANT WAS STABLE AND WELL. ON (B)(6) 2025, THE PATIENT PRESENTED EMERGENTLY TO THE HOSPITAL WITH PULMONARY EDEMA AND COUGHING WITH BLOODY SPUTUM. THE PHYSICIAN SUSPECTED DEVICE EMBOLIZATION. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF PULMONARY EDEMA, HEMOPTYSIS, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. THE PROVIDED IMAGING WAS REVIEWED BY MEDICAL AFFAIRS AND ABBOTT CLINICAL ENGINEERING MANAGER. BASED ON THE INFORMATION RECEIVED, A CAUSE FOR THE REPORTED PULMONARY EDEMA, HEMOPTYSIS, AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360608&pc=NGV | 2135147-2025-00687 | 2025-01-23 05:00 | Death | ABBOTT MEDICAL | 2025-02-11 05:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 20MM AMPLATZER AMULET WAS IMPLANTED WITH NO REPORTED COMPLICATIONS FOLLOWING A SINGLE DEPLOYMENT. ON (B)(6) 2025, THE PATIENT EXPERIENCED CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED WITHOUT SUCCESS AND THE PATIENT PASSED. THE CAUSE OF THE PATIENT PASSING WAS PRESUMED CARDIAC ARREST. TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED NO EFFUSION. THE AMULET REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF CARDIAC ARREST AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. THE PROVIDED IMAGING WAS REVIEWED BY MEDICAL AFFAIRS. BASED ON THE INFORMATION RECEIVED, A CAUSE FOR THE REPORTED CARDIAC ARREST AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360694&pc=MIH | 3008011247-2025-00014 | 2025-01-03 05:00 | Death | ENDOLOGIX SANTA ROSA | 2025-02-11 05:00 | MIH | OVATION IX | Complete Blockage | Ischemia; Obstruction/Occlusion | P120006 |  | 1 | Event Description: THE PATIENT WAS INITIALLY TREATED FOR AN ABDOMINAL AORTIC ANEURYSM (AAA) WITH THE IMPLANT OF AN ALTO ABDOMINAL STENT GRAFT SYSTEM ON (B)(6) 2023. ON (B)(6) 2025, A CLINICAL STUDY REPORTED THAT THE RIGHT LIMB OF THE PATIENT'S ENDOGRAFT WAS ACUTELY OCCLUDED, PROMPTING AN OPEN THROMBECTOMY OF THE RIGHT LOWER EXTREMITY (RLE), A FEMORAL-FEMORAL BYPASS, AND A MECHANICAL THROMBECTOMY DUE TO ACUTE THROMBUS. THE PATIENT, DIAGNOSED WITH LUNG CANCER AND BRAIN METASTASES, IS LIKELY HYPERCOAGULABLE. ACUTE THROMBUS WAS NOTED IN THE RIGHT COMMON FEMORAL ARTERY (CFA), DISTAL ABOVE-KNEE POPLITEAL ARTERY, AND TIBIAL VESSELS, WITH CHRONIC OCCLUSION OBSERVED IN THE EXTERNAL ILIAC ARTERY (EIA). DESPITE REVASCULARIZATION ON (B)(6) 2025, RE-OCCLUSION OCCURRED. DUE TO SEVERE PERIPHERAL VASCULAR DISEASE (PVD) IN SMALL VESSELS AND UNDERLYING MEDICAL COMORBIDITIES, A RIGHT ABOVE-KNEE AMPUTATION (AKA) WAS REQUIRED. THE PATIENT REMAINS HOSPITALIZED AND IN STABLE CONDITION AS OF (B)(6) 2025.Event Description: THE PATIENT WAS INITIALLY TREATED FOR AN ABDOMINAL AORTIC ANEURYSM (AAA) WITH THE IMPLANT OF AN ALTO ABDOMINAL STENT GRAFT SYSTEM ON (B)(6) 2023. ON (B)(6) 2025, A CLINICAL STUDY REPORTED THAT THE RIGHT LIMB OF THE PATIENT'S ENDOGRAFT WAS ACUTELY OCCLUDED, PROMPTING AN OPEN THROMBECTOMY OF THE RIGHT LOWER EXTREMITY (RLE), A FEMORAL-FEMORAL BYPASS, AND A MECHANICAL THROMBECTOMY DUE TO ACUTE THROMBUS. THE PATIENT, DIAGNOSED WITH LUNG CANCER AND BRAIN METASTASES, IS LIKELY HYPERCOAGULABLE. ACUTE THROMBUS WAS NOTED IN THE RIGHT COMMON FEMORAL ARTERY (CFA), DISTAL ABOVE-KNEE POPLITEAL ARTERY, AND TIBIAL VESSELS, WITH CHRONIC OCCLUSION OBSERVED IN THE EXTERNAL ILIAC ARTERY (EIA). DESPITE REVASCULARIZATION ON (B)(6) 2025, RE-OCCLUSION OCCURRED. DUE TO SEVERE PERIPHERAL VASCULAR DISEASE (PVD) IN SMALL VESSELS AND UNDERLYING MEDICAL COMORBIDITIES, A RIGHT ABOVE-KNEE AMPUTATION (AKA) WAS REQUIRED. THE PATIENT REMAINS HOSPITALIZED AND IN STABLE CONDITION AS OF (B)(6) 2025. ADDITIONAL INFORMATION: IT WAS REPORTED THE PATIENT DIED ON (B)(6) 2025. THE CAUSE OF DEATH WAS IDENTIFIED AS END-STAGE METASTATIC NON-SMALL CELL LUNG CANCER. Manufacturer Narrative: THE DEVICES INVOLVED IN THIS EVENT WILL NOT BE RETURNED FOR EVALUATION AND REMAIN IMPLANTED IN THE PATIENT. PATIENT MEDICAL RECORDS AND IMAGING STUDIES WILL BE REQUESTED FOR FURTHER EVALUATION BY A CLINICAL SPECIALIST. IF ADDITIONAL INFORMATION PERTINENT TO THE INCIDENT IS OBTAINED, A FOLLOW-UP REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. WHERE POSSIBLE, IT IS ENDOLOGIX PRACTICE TO MAKE AT LEAST THREE GOOD FAITH EFFORTS TO RETRIEVE A REPORTED ADVERSE EVENT/INCIDENT-RELATED DEVICE AS WELL AS MEDICAL RECORDS AND MEDICAL IMAGING. AN EVALUATION OF THE MANUFACTURING RECORD WAS COMPLETED. A REVIEW OF THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION SHOWS THE DEVICE TO HAVE BEEN PROPERLY MANUFACTURED AND RELEASED IN ACCORDANCE WITH THE DEVICE MASTER RECORD. THE REVIEW CONFIRMS THERE WERE NO MANUFACTURING OR PROCESSING NON-CONFORMITIES IDENTIFIED THAT WOULD CONTRIBUTE TO THE REPORTED ADVERSE EVENT/INCIDENT. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT WAS COMPLETED. AN EXAMINATION OF MEDICAL RECORDS AND/OR MEDICAL IMAGING RECEIVED BY ENDOLOGIX SHOWS THE RIGHT ILIAC LIMB OCCLUSION, ADDITIONAL ENDOVASCULAR PROCEDURE (THROMBECTOMY) ADDITIONAL SURGICAL PROCEDURE (FEMORAL TO FEMORAL BYPASS), AND RIGHT LIMB ABOVE THE KNEE AMPUTATION ARE CONFIRMED. THIS IS CONSISTENT WITH THE REPORTED ADVERSE EVENT/INCIDENT. THE CLINICAL EVALUATION ALSO SHOWS REASONABLE EVIDENCE TO SHOW THE PATIENT HAS EXPIRED. THE PATIENT'S DEATH WAS DISCOVERED DURING A REVIEW OF THE DEATH CERTIFICATE. THE REPORTED CAUSE OF DEATH WAS END-STAGE METASTATIC NON-SMALL CELL LUNG CANCER. THE THROMBUS CAUSING THE OCCLUSION WAS REPORTED TO BE LIKELY A RESULT OF THE PATIENT BEING IN A HYPERCOAGULABLE STATE DUE TO CANCER. THE INTRACRANIAL HEMORRHAGE WAS REPORTED TO BE RELATED TO THE HEPARIN ADMINISTERED DURING THE PROCEDURE. THE COMPLAINT IS MOST LIKELY ANATOMY-RELATED. THE PROCEDURE-RELATED HARM IDENTIFIED WAS AN INTRACRANIAL HEMORRHAGE. THE FINAL PATIENT STATUS WAS REPORTED AS DECEASED ON JANUARY 21, 2025. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS. CORRECTIONS: B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENTS - UPDATED B5: DESCRIBE EVENT OR PROBLEM - UPDATED G3: AWARENESS DATE - UPDATED H1: TYPE OF REPORTABLE EVENT - UPDATED H6: INVESTIGATION FINDING CODES - REMOVE CODE 3233 H6: INVESTIGATION CONCLUSION CODES - REMOVE CODE 11.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360913&pc=DSQ | 2916596-2025-00730 | 2025-01-22 05:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS AND CARDIOVASCULAR RESYNCHRONIZATION THERAPY WITH LEFT VENTRICULAR ASSIST DEVICES: A MOMENTUM 3 TRIAL ANALYSIS¿ IDENTIFYING THAT HM3 MAY BE RELATED TO VENTRICULAR ARRHYTHMIA AND DEATH. OF THE 1387 PATIENTS ANALYZED, 699 (50.4%) HAD AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) ALONE AND 535 (38.6%) HAD A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D), AND 153 (11.0%) HAD NO DEVICE. VENTRICULAR ARRHYTHMIA OCCURRED MORE FREQUENTLY IN THE ICD/CRT-D GROUP THAN THE NO DEVICE GROUP (HR 2.4, 95% CI 1.3-4.3, P=0.006), AND MORE FREQUENTLY IN THE CRT-D GROUP THAN THE ICD GROUP (HR 1.3, 95% CI 1.0-1.7, P=0.03). ADJUSTED COX PROPORTIONAL HAZARDS ANALYSIS SHOWED NO DIFFERENCE OVER 2-YEAR SURVIVAL (AHR:1.3, 95% CI: 0.8-2.1, P=0.36) BETWEEN PATIENTS WITH ICD/CRT-D AND THOSE WITHOUT A DEVICE. LIKEWISE, NO DIFFERENCES WERE SEEN WHEN THE ICD GROUP WAS COMPARED TO THE NO DEVICE GROUP (AHR 0.8, 95% CI 0.4-1.7, P=0.60), THE CRT-D GROUP WAS COMPARED TO THE NO DEVICE GROUP (AHR 1.4, 95% CI 0.8-2.4, P=0.25) OR WHEN THE CRT-D GROUP WAS COMPARED TO THE ICD GROUP (AHR 1.1, 95% CI 0.9-1.5, P=0.36). AN EXPLORATORY ANALYSIS SHOWED INCREASED RATES OF REHOSPITALIZATIONS FOR HF AMONG THE ICD GROUP AND THE COMBINED ICD/CRT-D GROUP COMPARED TO THE NO DEVICE GROUP, AS WELL AS INCREASED REHOSPITALIZATIONS FOR CARDIAC ARRHYTHMIAS (NOT LIMITED TO VA ALONE) AMONG THE ICD GROUP, CRT-D GROUP AND ICD/CRT-D GROUP COMPARED TO THE NO DEVICE GROUP. Manufacturer Narrative: SECTION B2: DATE OF DEATH ESTIMATED AS EVENT DATE SAYER G, ET AL. IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS AND CARDIOVASCULAR RESYNCHRONIZATION THERAPY WITH LEFT VENTRICULAR ASSIST DEVICES: A MOMENTUM 3 TRIAL ANALYSIS. J CARD FAIL. 2025 JAN 22:S1071-9164(25)00012-0. DOI: 10.1016/J.CARDFAIL.2024.12.011. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS SAME AS PUBLISHED DATE, 22JAN2025, SINCE THE DATE OF DATA COLLECTION WAS NOT PROVIDED. GABRIEL SAYER ET AL., PUBLISHED ONLINE ON 22 JANUARY 2025 JOURNAL OF CARDIAC FAILURE (2025), DOI: HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.12.011 COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360921&pc=DSQ | 2916596-2025-00716 | 2024-12-07 05:00 | Death | THORATEC CORPORATION | 2025-02-11 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO SEPTIC SHOCK. THE VENTRICULAR ASSIST DEVICE (VAD) WAS DECOMMISSIONED, AND THE VAD FUNCTIONED AS INTENDED. THE SOURCE OF INFECTION WAS THE DRIVELINE EXIT SITE. PSEUDOMONAS WAS IDENTIFIED THROUGH CLINICAL CULTURES. NO DEVICE WILL RETURN FOR EVALUATION. Manufacturer Narrative: E3 OCCUPATION CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21360970&pc=NPS | 2135147-2025-00690 | 2020-06-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-11 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "REAL-WORLD EXPERIENCE IN TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR: TRANSCATHETER TRICUSPID VALVE REPAIR IN SPAIN REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE STUDY, MULTI CENTER STUDY, TO EVALUATE ACUTE AND SHORT-TERM CARDIOVASCULAR OUTCOMES OF TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) WITH DEDICATED DEVICES IN A REAL-WORLD SETTING. DEVICES MENTIONED INCLUDE TRICLIP AND PASCAL. THE ARTICLE CONCLUDED THAT IN A REAL-WORLD CONTEMPORARY SETTING, TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR WITH DEDICATED DEVICES EMERGED AS EFFECTIVE THERAPEUTIC OPTION FOR PATIENTS WITH SEVERE TR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MANUEL BARREIRO-PEREZ AT UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN WITH CORRESPONDING EMAIL MANUELBARREIROPEREZ@GMAIL.COM. THE TIME FRAME OF THE STUDY WAS JUNE 2020 TO MAY 2023. . A TOTAL OF 283 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 223 (79%) RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED HEART FAILURE, PERIPHERAL EDEMA, TRICUSPID REGURGITATION, ARTERIAL HYPERTENSION, DIABETES, DYSLIPIDEMIA, CORONARY ARTERY DISEASE, PRVIOUS CABG, PREVIOUS HEART VALVE SURGERY, PREVIOUS STROKE, CHRONIC KIDNEY DISEASE, COPD, PERIPHERAL ARTERY DISEASE, PREVIOUS STROKE, PACEMAKER, HEART FAILURE MEDICATION, ANTICOAGULANT THERAPY. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED SINGLE LEAFLET DEVICE ATTACHMENT, ACUTE MYOCARDIAL INFARCTION, MAJOR BLEEDING, VASCULAR COMPLICATION, DEATH, HEART FAILURE HOSPITALIZATION, RECURRENT TRICUSPID REGURGITATION, UNEXPECTED MEDICAL INTERVENTION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEART FAILURE, PERIPHERAL EDEMA, TRICUSPID REGURGITATION, ARTERIAL HYPERTENSION, DIABETES, DYSLIPIDEMIA, CORONARY ARTERY DISEASE, PREVIOUS CABG, PREVIOUS HEART VALVE SURGERY, PREVIOUS STROKE, CHRONIC KIDNEY DISEASE, COPD, PERIPHERAL ARTERY DISEASE, PREVIOUS STROKE, PACEMAKER, HEART FAILURE MEDICATION, ANTICOAGULANT THERAPY. COMPLICATIONS REPORTED INCLUDED SINGLE LEAFLET DEVICE ATTACHMENT, ACUTE MYOCARDIAL INFARCTION, MAJOR BLEEDING, VASCULAR COMPLICATION, DEATH, HEART FAILURE HOSPITALIZATION, RECURRENT TRICUSPID REGURGITATION, UNEXPECTED MEDICAL INTERVENTION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTION REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE: ARTICLE TITLED "REAL-WORLD EXPERIENCE IN TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR_ TRANSCATHETER TRICUSPID VALVE REPAIR IN SPAIN REGISTRY".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21361017&pc=MIH | 2017233-2025-05819 | 2025-01-16 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-11 05:00 | MIH | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Insufficient Information | Low Blood Pressure/ Hypotension; Paralysis | P230023 |  | 1 | Event Description: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM UTILIZING GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS AND A REDUCED PROFILE GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE). DURING THE PROCEDURE, THERE WAS TROUBLE ADVANCING A REDUCED PROFILE GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) INTO THE LEFT RENAL AND MAINTAINING WIRE ACCESS. THERE WERE MULTIPLE ATTEMPTS MADE TO RE-ACCESS THE ARTERY. PHYSICIAN CHOSE TO USE A SHORTER STENT, AND IT TRACKED AND DEPLOYED INTO THE RENAL ARTERY. PHYSICIAN THEN ATTEMPTED TO PLACE THE BRIDGING STENT THAT WAS IN THE LEFT RENAL ARTERY, AND IT WOULDN¿T TRACK. PHYSICIAN THEN ELECTED TO MOVE TO THE SMA AND CELIAC ARTERIES AND WAS ABLE TO SUCCESSFULLY TREAT BOTH ARTERIES WITH NO ISSUES. PHYSICIAN WENT BACK TO THE RIGHT RENAL ¿ USED A 46MM CODA BALLOON TO CREATE A BACKBOARD AND A 1CM TIP AMPLATZ WIRE, IN WHICH PHYSICIAN WAS ABLE TO BRIDGE THE RENAL STENT TO THE TAMBE DEVICE. PATIENT TOLERATED THE PROCEDURE. AN ANGIO WAS PERFORMED, AND THE PHYSICIAN WAS SATISFIED WITH THE RESULTS. THE PATIENT THEN PRESENTED BACK LATER THAT AFTERNOON AND BLEEDING IN THE LEFT KIDNEY WAS OBSERVED. PATIENT BECAME HYPOTENSIVE, WITH SOME LIMB ISCHEMIA AS WELL. FEMORAL ENDARTERECTOMY WAS PERFORMED. AFTER BECOMING HYPOTENSIVE, PATIENT ALSO EXPERIENCED SOME PARAPLEGIA AND SHOCK LIVER WAS ALSO REPORTED. NO OTHER RECOVERY PROCEDURES WERE PERFORMED AND PATIENT PASSED AWAY THE FOLLOWING DAY ON (B)(6) 2025. IT WAS ALSO REPORTED CAUSE OF ISCHEMIA DID NOT INVOLVE A DEVICE, AS ALL DEVICES WERE PATENT. A CLOT WAS JUST OBSERVED AT SOME POINT DURING THE PROCEDURE. PHYSICIAN BELIEVES THE CAUSE OF BLEEDING IN THE LEFT KIDNEY WAS A RESULT FROM TRAUMA FROM THE NON-GORE WIRE. PHYSICIAN BELIEVES THE PARAPLEGIA IS DUE TO HYPOTENSION, AS A RESULT FROM HEMORRHAGIC EVENT OF THE LT KIDNEY. PATIENT DEATH WAS DUE TO SHOCK LIVER. THE CAUSE OF DIFFICULTY INSERTING STENT INTO LEFT RENAL ARTERY WAS DUE TO THE ANGLE AT THE OSTIUM OF THE RENAL WAS ACUTE, AND RENAL ARTERY AFTER BIFURCATION (OUR LANDING ZONE) WAS SHORT, WHICH IMPACTED WIRE PURCHASE. NO FURTHER PATIENT INFORMATION OR IMAGES ARE AVAILABLE. Manufacturer Narrative: H6: CODE C19-A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE MANUFACTURING SPECIFICATIONS. AS THE DEVICE WAS NOT ACCESSIBLE, THE PRODUCT HISTORY REVIEW (PHR) REVIEW WAS THE EXTEND OF THE INVESTIGATION. NO DEVICE PROBLEM WAS FOUND PER REVIEW OF THESE RECORDS. H6: CODE B20 ¿ THE DEVICE REMAINS IMPLANTED AND, THEREFORE, WAS NOT AVAILABLE FOR DIRECT ANALYSIS BY GORE. H6: CODE A26-USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE IN RELATION TO CAUSE OF DEATH. IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21361087&pc=FED | 3014683069-2025-00004 | 2024-05-28 04:00 | Death | CALYXO, INC. | 2025-02-11 05:00 | FED | CVAC ASPIRATION SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K233472 |  | 1 | Event Description: ON (B)(6) 2024, AN ASA CLASS III 83-YEAR-OLD NURSING HOME PATIENT WAS TREATED WITH THE SURE PROCEDURE FOR BILATERAL INFECTIOUS KIDNEY STONES. COMPLEX MEDICAL HISTORY INCLUDED MYOCARDIAL INFARCTION LEADING TO EMERGENT CARDIAC CATHETERIZATION DURING A FAILED PERCUTANEOUS NEPHROLITHOTOMY 8 MONTHS PRIOR AND RECURRENT HOSPITALIZATIONS FOR URINARY TRACT INFECTIONS AND PYELONEPHRITIS AS RECENTLY AS 1 MONTH PRIOR TO PROCEDURE. URINARY TRACT INFECTION WAS TREATED PRIOR TO THE PROCEDURE AND ANTIBIOTIC PROPHYLAXIS WAS ADMINISTERED. ABOUT 30 MINUTES INTO THE PROCEDURE, THE PHYSICIAN NOTICED A DECREASE IN FLOW, REMOVED THE DEVICE, AND NOTICED A TEAR IN THE OUTER JACKET OF THE CVAC CATHETER SHAFT. A FRAGMENT OF THE OUTER JACKET WAS LOCATED IN THE RENAL PELVIS AND REMOVED. THE PATIENT WAS STABLE, AND THE PROCEDURE CONTINUED WITH A NEW DEVICE. AT APPROXIMATELY 90 MINUTES INTO THE PROCEDURE, THE PHYSICIAN COMPLETED THE REMOVAL OF STONES FROM ONE KIDNEY AND WAS BEGINNING TO TREAT THE KIDNEY ON THE OTHER SIDE WHEN THE PATIENT WAS NOTED TO HAVE LOW BLOOD PRESSURE, AND THE PHYSICIAN SUSPECTED THE PATIENT EXPERIENCED A MYOCARDIAL INFARCTION. THE PROCEDURE WAS HALTED, VASOPRESSORS WERE ADMINISTERED, AND THE PATIENT WAS STABILIZED FOR TRANSFER TO THE CARDIAC ICU. AN ULTRASOUND REVEALED INTACT, UNDAMAGED KIDNEY; URINALYSIS AND PERITONEAL TAP WERE CLEAR. CENTRAL LINE PLACEMENT WAS COMPLICATED BY WHAT THE TREATING PHYSICIAN BELIEVES WAS A LACERATED CAROTID ARTERY AND SIGNIFICANT BLOOD LOSS AT THE BEDSIDE. THE PATIENT SUFFERED FROM MULTIORGAN FAILURE AND TWO HOURS AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. AUTOPSY WAS DECLINED; HOWEVER, THE OFFICIAL CAUSE OF DEATH PROVIDED BY THE TREATING PHYSICIAN WAS HEMORRHAGIC SHOCK AND SECONDARY CAUSE WAS COAGULOPATHY. THE TREATING PHYSICIAN STATES THAT THE PATIENT'S DEATH WAS UNRELATED TO THE CVAC DEVICE BUT MAY HAVE BEEN RELATED TO COMPROMISED VASCULAR ACCESS. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR INVESTIGATION. ANALYSIS OF THE RETURNED DEVICE REVEALED THAT A TEAR ~1CM LONG WAS EVIDENT AT THE DISTAL TIP, CONFIRMING THE REPORTED EVENT OF A TEAR IN THE DEVICE'S CATHETER SHAFT. IT IS MOST LIKELY THAT THE CATHETER TEAR WAS THE RESULT OF THE OUTER JACKET LUMEN RUBBING ON THE ACCESS SHEATH DURING THE PROCEDURE. THE OBSERVED TEAR IS UNRELATED TO THE REPORTED PATIENT ADVERSE EVENT. THE LOT HISTORY REVIEW (LHR) FOR LOT# P01256 WAS CONDUCTED, WHICH CONFIRMED THAT THERE WERE NO NON-CONFORMANCES DURING THE MANUFACTURING PROCESS THAT COULD BE RELATED TO THE REPORTED EVENT. THE REVIEW INDICATED THAT THE DEVICE MET ALL DESIGN AND MANUFACTURING SPECIFICATIONS WHEN RELEASED FOR DISTRIBUTION. KIDNEY STONE PROCEDURES ARE GENERALLY SAFE BUT NOT RISK-FREE. MORTALITY RATES FOLLOWING URETEROSCOPY (URS) AND PCNL HAVE BEEN REPORTED TO RANGE FROM 0.02 - 0.07% FOLLOWING URS AND 0.03 TO 0.7% FOLLOWING PCNL. MULTIPLE FACTORS INCLUDING MEDICAL HISTORY MUST BE CONSIDERED WHEN TREATING A PATIENT AS A RECENT META-ANALYSIS FOUND THAT COMORBIDITIES OF DIABETES, CARDIOVASCULAR DISEASE, AND RECURRENT UTIS INCREASED THE RISK OF MORTALITY FOLLOWING URS. Manufacturer Narrative: CALYXO HAS RECEIVED THE REPORTED DEVICE. THE REPORTED TEAR IN THE OUTER JACKET OF THE DEVICE WAS VISIBLY OBSERVED; HOWEVER, OUR INVESTIGATION IS STILL IN PROCESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21363611&pc=FRO | MW5165991 | 2024-07-26 04:00 | Death | INTEGRA LIFESCIENCES CORPORATION | 2025-02-11 05:00 | FRO | MEDIHONEY GEL |  | Bacterial Infection; Sepsis |  |  | 1 | Event Description: MY MOTHER, (B)(6) WENT INTO THE HOSPITAL WITH SEVERE SEPSIS AND STAPHYLOCOCCUS EPIDERMIS BACTERIA. PRIOR TO THIS SHE WAS BEING TREATED AT HOME FOR WOUND CARE. THE HOME HEALTH AID WAS USING MEDIHONEY GEL TO TREAT HER WOUNDS AT THE TIME SHE CONTRACTED THE INFECTION. IT WAS DETERMINED THAT THE INFECTION STARTED AT A WOUND ON HER LEG WHERE SHE WAS BEING TREATED WITH THE MEDIHONEY GEL. TODAY I RECEIVED A LETTER FROM CARDINAL HEALTH THAT WAS A NOTICE FOR URGENT MEDICAL RECALL. IT STATES THAT THERE WAS A MANUFACTURING ERROR WITH THE SEAL AND IT WAS POSSIBLE THAT IT COULD GET CONTAMINATED. IT STATES THE DATE AND LOT LUMBER OF THE MEDIHONEY IN QUESTION OF WHICH I HAVE IN MY POSSESSION AND UNOPENED BOX OF THE PRODUCT THAT MATCHES THE DATE AND LOT NUMBER OF THE PRODUCT IN QUESTION. THE DATES OF USE OF THIS PRODUCT CORRESPOND TO THE TIMES THAT MEDIHONEY WAS ADMINISTERED TO (B)(6) AND SUBSEQUENTLY ADMITTED TO THE HOSPITAL WITH THE INFECTION. THIS HAPPENED ON TWO OCCASIONS. ONCE HOME FROM THE HOSPITAL THE FIRST TIME SHE WAS AGAIN ADMINISTERED MEDIHONEY BY THE HOME HEALTH AID. SHE AGAIN WAS ADMITTED TO THE HOSPITAL WITH SEPSIS AND A STAPH INFECTION. AMONG HER OTHER CONDITIONS I SUSPECT THAT THIS LED TO HER PASSING ON (B)(6) 2024. TREATING A SMALL WOUND ON HER LEG. REFERENCE REPORT: #MW5165992.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21363681&pc=FRO | MW5165992 | 2024-07-26 04:00 | Death | INTEGRA LIFESCIENCES CORPORATION | 2025-02-11 05:00 | FRO | MEDIHONEY GEL | Nonstandard Device; Device Contaminated During Manufacture or Shipping; Patient Device Interaction Problem | Bacterial Infection; Sepsis |  |  | 1 | Event Description: MY MOTHER, (B)(6) WENT INTO THE HOSPITAL WITH SEVERE SEPSIS AND STAPHYLOCOCCUS EPIDERMIS BACTERIA. PRIOR TO THIS SHE WAS BEING TREATED AT HOME FOR WOUND CARE. THE HOME HEALTH AID WAS USING MEDIHONEY GEL TO TREAT HER WOUNDS AT THE TIME SHE CONTRACTED THE INFECTION. IT WAS DETERMINED THAT THE INFECTION STARTED AT A WOUND ON HER LEG WHERE SHE WAS BEING TREATED WITH THE MEDIHONEY GEL. TODAY I RECEIVED A LETTER FROM CARDINAL HEALTH THAT WAS A NOTICE FOR URGENT MEDICAL RECALL. IT STATES THAT THERE WAS A MANUFACTURING ERROR WITH THE SEAL AND IT WAS POSSIBLE THAT IT COULD GET CONTAMINATED. IT STATES THE DATE AND LOT LUMBER OF THE MEDIHONEY IN QUESTION OF WHICH I HAVE IN MY POSSESSION AND UNOPENED BOX OF THE PRODUCT THAT MATCHES THE DATE AND LOT NUMBER OF THE PRODUCT IN QUESTION. THE DATES OF USE OF THIS PRODUCT CORRESPOND TO THE TIMES THAT MEDIHONEY WAS ADMINISTERED TO (B)(6) AND SUBSEQUENTLY ADMITTED TO THE HOSPITAL WITH THE INFECTION. THIS HAPPENED ON TWO OCCASIONS. ONCE HOME FROM THE HOSPITAL THE FIRST TIME SHE WAS AGAIN ADMINISTERED MEDIHONEY BY THE HOME HEALTH AID. SHE AGAIN WAS ADMITTED TO THE HOSPITAL WITH SEPSIS AND A STAPH INFECTION. AMONG HER OTHER CONDITIONS I SUSPECT THAT THIS LED TO HER PASSING ON (B)(6) 2024. TREATING A SMALL WOUND ON HER LEG. REFERENCE REPORT: #MW5165991.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21341858&pc=CBK | 9611500-2025-00066 | 2025-01-11 05:00 | Death | DRÄGERWERK AG & CO. KGAA | 2025-02-10 05:00 | CBK | OXYLOG 3000PLUS | Failure to Run on Battery; Failure to Deliver; Complete Loss of Power | Hypoxia | K103625 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE HAD BEEN LEFT IN DEVICE CHECK SCREEN WHILE ON CHARGE, AT SOME POINT THE MAINS LEAD HAD SEPARATED FROM THE POWER SUPPLY TO NO LONGER MAKE CONTACT. ONCE PUT INTO USE THE BATTERIES WERE FLAT. FURTHERMORE, IT WAS REPORTED BY THE USER THAT THE DEVICE' "BATTERY WENT FLAT DURING RESUSCITATION IN EU, RESULTING IN A DELAY IN PROVIDING THERAPY AS MANUAL VENTILATION VIA A BAG AND MASK WAS REQUIRED. IT IS POSSIBLE THE INCIDENT HASTENED THE PATIENT'S DEMISE BUT ONLY BY MINUTES OR AT BEST HOURS IN THE OPINION OF THE MEDICS."Event Description: IT WAS REPORTED THAT THE DEVICE HAD BEEN LEFT IN DEVICE CHECK SCREEN WHILE ON CHARGE, AT SOME POINT THE MAINS LEAD HAD SEPARATED FROM THE POWER SUPPLY TO NO LONGER MAKE CONTACT. ONCE PUT INTO USE THE BATTERIES WERE FLAT. FURTHERMORE, IT WAS REPORTED BY THE USER THAT THE DEVICE' "BATTERY WENT FLAT DURING RESUSCITATION IN EU, RESULTING IN A DELAY IN PROVIDING THERAPY AS MANUAL VENTILATION VIA A BAG AND MASK WAS REQUIRED. IT IS POSSIBLE THE INCIDENT HASTENED THE PATIENT'S DEMISE BUT ONLY BY MINUTES OR AT BEST HOURS IN THE OPINION OF THE MEDICS." Manufacturer Narrative: THE INVESTIGATION HAS JUST STARTED; RESULTS WILL BE PROVIDED IN A FOLLOW-UP REPORT. Manufacturer Narrative: THE LOG FILE OF THE AFFECTED DEVICE WAS PROVIDED AND ANALYZED. IN ADDITION, THE POWER SUPPLY INCLUDING A POWER CORD WERE RETURNED FOR FURTHER INVESTIGATION. BASED ON THE LOG FILE ANALYSIS, A DEVICE MALFUNCTION COULD BE EXCLUDED AS ROOT CAUSE OF THE REPORTED EVENT (NEITHER DEVICE NOR INTERNAL BATTERY FAULT). ON THE (B)(6) 2025 THE LOG FILE SHOWS THAT THE DEVICE WAS TURNED ON AT 00:43 A.M. AND A DEVICE CHECK WAS SUCCESSFULLY PERFORMED. AT 07:46 A.M. A SECOND DEVICE CHECK WAS CARRIED OUT SUCCESSFULLY. AT 08:40:48 A.M., THE CONFIGURATION MODE WAS EXITED INDICATING THE START OF THE VENTILATION MODE. IMMEDIATELY AFTER SETTING THE VENTILATION MODE, AT 08:40:49 A.M., THE DEVICE ALARMED FOR ¿CHARGE INT. BATTERY¿ WITH MID PRIORITY (10 MINS OF BATTERY RUNTIME LEFT ACC. TO THE IFU) AND AGAIN AT 08:50:49 A.M.. AT 08:54:16 A.M., THE ALARM ¿INT. BATTERY DISCHARGED¿ WAS GENERATED BY THE DEVICE AND THE VENTILATION STOPPED DUE TO LOSS OF POWER. AT 08:56:12 A.M. THE DEVICE WAS MANUALLY SHUTDOWN BY USER. THE DEVICE IS EQUIPPED WITH DIFFERENT INDICATORS IN TERMS OF INTERNAL BATTERY CHARGE STATE AND EXTERNAL POWER SUPPLY CONNECTION. THE DISPLAY SYMBOLS OF THE POWER SUPPLY ARE LOCATED AT THE FRONT PANEL OF THE DEVICE AND INDICATE THE CHARGE STATUS OF THE INTERNAL BATTERY (A THREE COLORED INDICATOR) AND THE GREEN INDICATOR LIGHTS UP WHEN AN EXTERNAL POWER SOURCE IS CONNECTED. FURTHERMORE, THE DEVICE DISPLAYS THE REMAINING CAPACITY OF THE INTERNAL BATTERY IN INCREMENTS OF 25 % IN THE LOWER RIGHT SECTION OF THE INFORMATION WINDOW (¿BATTERY METER¿) WHILE IN OPERATION. ALARM MESSAGES OF CORRESPONDING PRIORITIES ARE GENERATED IF THE DEVICE DETECTS AN ISSUE WITH THE FUNCTION OF THE INTERNAL BATTERY OR THE NEED TO CHARGE IT. CHARGING THE BATTERY MIGHT BE REQUIRED TO ENSURE THAT THE DEVICE CAN CONTINUOUSLY BE OPERATED WITHOUT CONNECTING IT TO AN EXTERNAL POWER SOURCE. IN ADDITION TO THE LOG FILE ANALYSIS, THE POWER SUPPLY INCLUDING THE POWER CORD WAS SUBJECTED TO AN ENGINEERING TEST (TEST OF PULLING FORCES). THE POWER SUPPLY AS RETURNED FOR INVESTIGATION WAS WITH A POWER CORD OTHER THAN THE ORIGINAL POWER CORD (1844369), WHICH IS INSIDE THE PACKAGING OF THE OXYLOG 3000PLUS AS PART OF THE DEVICE CONFIGURATION. THE POWER CORD AS RETURNED HAS A STRAIGHT CONNECTION LEAD, WHICH HAS A HIGHER POSSIBILITY OF BEING PULLED OUT OF THE SOCKET. THE PULLING FORCES OF A POWER CORD DOES NOT DIRECTLY EXERT FORCE ON THE CONNECTION LEAD WHEN IT IS POSITIONED AT A 90-DEGREE ANGLE; THE ANGLE ALTERS THE DIRECTION OF THE FORCE VECTOR AND REDUCES THE PULLING FORCE ACTING ALONG THE CONNECTION LEAD PREVENTING UNINTENTIONAL UNPLUGGING. THIS ASSESSMENT COULD BE CONFIRMED BY THE PULLING FORCES AS MEASURED DURING THE ENGINEERING TEST. THE PULLING FORCE WHICH HAD TO BE APPLIED FOR PULLING THE ANGLED CONNECTION LEAD OUT OF THE POWER SUPPLY WAS SIGNIFICANTLY HIGHER (APPROX. 3 TIMES) COMPARED TO THE STRAIGHT CONNECTION LEAD. DRÄGER CONSIDERS THE CURRENT PRODUCT DESIGN OF THE POWER SUPPLY INCLUDING THE POWER CORD (1844369) TO BE PROVEN AND SAFE. AN ADDITIONAL RETENTION CLIP IS THEREFORE NOT NECESSARY. IN THE CURRENT EVENT, THE DEVICE REACTED AS SPECIFIED. BASED ON THE INVESTIGATION, THERE WAS NO DEVICE FAILURE FOUND. THE NUMBER OF SIMILAR CASES IS WITHIN THE EXPECTED RANGE OF THE RESPECTIVE RISK ASSESSMENT AND THUS ACCEPTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21343175&pc=BZD | 2518422-2025-101878 | 2024-11-12 05:00 | Death | RESPIRONICS, INC. | 2025-02-10 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DREAMSTATION AUTO CPAP DEVICE. THE CALLER STATED THE DEVICE WAS DISCARDED AND CANNOT BE RETURNED. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21343348&pc=OZD | 1220648-2025-26060 | 2025-01-07 05:00 | Death | ABIOMED, INC. | 2025-02-10 05:00 | OZD | IMPELLA CP | Fluid/Blood Leak |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT WAS ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) THERAPY LESS THAN ONE HOUR PRIOR TO INITIAL IMPELLA MCS. AFTER INSERTION OF THE PUMP, PURGE FLUID LEAK ALARM SOUNDED WITH NO VISIBLE LEAK. IN RESPONSE TO THE ISSUE, THE PURGE CASSETTE, PURGE HOUSING WAS REPLACED, BUT THE ISSUE REMAINED UNRESOLVED. CONSEQUENTLY, APPROXIMATELY FOUR HOURS AFTER START OF INITIAL MCS, THE IMPELLA CP PUMP WAS EXPLANTED AND REPLACED WITH A NEW IMPELLA CP DEVICE. THE IMPELLA TIMING WAS BAILOUT WITH PERCUTANEOUS CORONARY ANGIOPLASTY TO THE LEFT MAIN ARTERY. ONE DAY AFTER THE IMPELLA CP SUPPORT, THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 DEVICE DUE TO THE NEED OF LONG-TERM SUPPORT. THE PATIENT'S HEMODYNAMICS DID NOT IMPROVE, AND THE PATIENT EXPIRED OF MULTI-ORGAN FAILURE SIX DAYS LATER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE OUTCOME GIVEN INITIAL SUPPORT WITH THE PURGE SYSTEM FAILURE (PURGE ALARMS TRIGGER WITH NO VISIBLE LEAK RESULTING IN THE NEED FOR ADDITIONAL IMPELLA INTERVENTION PROCEDURE). HOWEVER, OF NOTE, THE PATIENT RECEIVED UNINTERRUPTED MCS FOR SEVEN DAYS. HEMODYNAMICS DID NOT IMPROVE (MEDICAL HISTORY AND ANY COMORBIDITIES ARE UNKNOWN). PER THE REPORT THE PATIENT EXPIRED FROM MULTI-ORGAN FAILURE DUE TO AN UNKNOWN CAUSE. Manufacturer Narrative: THE IMPELLA PUMP WITH PURGE CASSETTE HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF PURGE LEAK ¿ PUMP HAS BEEN COMPLETED. THE PRODUCT WAS RETURNED AND EVALUATED. THERE WERE NO ABNORMALITIES FOUND UPON VISUAL INSPECTION. UPON RUNNING THE PUMP WITH A TEST PURGE CASSETTE USING DEXTROSE FIVE PERCENT (5%) IN WATER, THE LOW PURGE PRESSURE AND HIGH PURGE FLOW REPRODUCED WITH PURGE PRESSURE AT APPROXIMATELY 200 MMHG AND PURGE FLOW AT APPROXIMATELY 30 MILLILITER PER HOUR. THERE WERE NO LEAKS OBSERVED. THE DATA LOGS ALSO SHOWED LOW PURGE PRESSURE FROM THE BEGINNING WITH PURGE FLOW AT APPROXIMATELY 30 MILLILITERS PER HOUR. THE GLUCOSE CONCENTRATION WAS 5% IN FIELD. NO OTHER ABNORMALITIES WERE SEEN WITH THE DATA LOGS. ALTHOUGH THE PUMP PASSED ALL CHECKS IN MANUFACTURING AND WAS BUILT TO SPECIFICATION, IT IS LIKELY THAT THE TOLERANCE STACK UP OF THE SPECIFIC COMPONENTS OF THIS BUILD LEAD TO LESS RESISTANCE AND HIGHER PURGE FLOW. THE ROOT CAUSE OF THE LOW PURGE PRESSURE WAS MOST LIKELY A DESIGN LIMITATION OF THE PUMP. D.10 THERAPY DATES SHOULD HAVE BEEN LEFT BLANK AS INFORMATION WAS ERRONEOUSLY ENTERED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26060. E.1 REVISED ADDRESS LINE 2 AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26060. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26060 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21344525&pc=QEZ | 1319211-2025-00003 | 2025-01-09 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-10 05:00 | QEZ | ALPHAVAC | Difficult to Advance | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26 YEAR OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT . (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING . THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING . AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 1 OF 3. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT EMBOLIZATION AND SUBSEQUENT EXPIRATION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). THERE WERE REPORTS OF ALPHAVAC CANNULA MALFUNCTION DURING THE PROCEDURE, WHICH WAS DETERMINED TO BE KINK DAMAGE TO CANNULA COILS, LIKELY DUE TO HANDLING DAMAGE DURING PROCEDURE AS A RESULT OF ENGAGING THE TUMOR MATERIAL. THE ALPHAVAC DEVICE WAS BEING USED IN AN OFF-LABEL CAPACITY, I.E. FOR REMOVAL OF TUMOR MATERIAL AND NOT THROMBI OR EMBOLI. THE USE OF THE ALPHAVAC DEVICE IS NOT BELIEVED TO BE A CONTRIBUTING FACTOR TO THE PATIENT'S EXPIRATION. ALPHAVAC WAS BEING USED OFF-LABEL ALONG WITH OTHER SIMILAR ASPIRATION DEVICES (ONO SNARE, PENUMBRA INDIGO). DISTAL EMBOLIZATION OF THROMBUS AND DEATH ARE POTENTIAL/ANTICIPATED PROCEDURAL COMPLICATIONS OF AN ALPHAVAC PROCEDURE; THIS IS CAUTIONED IN THE DEVICE DFU. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. NO CORRECTION IS REQUIRED SINCE THERE WAS NO MANUFACTURING NON-CONFORMANCE OBSERVED DURING EVALUATION OF THE RETURNED ALPHAVAC DEVICES. THE KINKED TUBING OBSERVED WAS LIKELY A RESULT OF HANDLING DAMAGE DURING DEVICE USE TO ENGAGE WITH THE TUMOR MATERIAL. THE PATIENT THROMBOSIS/EMBOLISM AND SUBSEQUENT EXPIRATION IS NOT RELATED TO THE USE OF THE ALPHAVAC DURING THIS OFF-LABEL PROCEDURE. DISTAL EMBOLIZATION OF THROMBUS AND DEATH ARE POTENTIAL/ANTICIPATED PROCEDURAL COMPLICATIONS OF AN ALPHAVAC PROCEDURE; THIS IS CAUTIONED IN THE DEVICE DFU. THIS SERIOUS ADVERSE EVENT HAS BEEN CAPTURED IN THE POST MARKET SURVEILLANCE OF THE ALPHAVAC PRODUCT FAMILY. LABELING REVIEW: INFORMATION RECEIVED THAT THE ALPHAVAC DEVICE AS NOT USED IN ACCORDANCE WITH ITS LABELING, I.E. OFF-LABEL USE FOR REMOVAL OF TUMOR MATERIAL AND NOT THROMBI OR EMBOLI. THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE: THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: DISTAL EMBOLIZATION OF THROMBUS, PULMONARY EMBOLISM, CARDIAC ARREST, DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE: (B)(4) DEVICE 1 OF 3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21344633&pc=CBK | 2518422-2025-101886 | 2025-01-27 05:00 | Death | RESPIRONICS, INC. | 2025-02-10 05:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT ON (B)(6) 2025 AT APPROXIMATELY 10:19 A NURSE AT AN INSTITUTIONAL FACILITY DISCOVERED THE PATIENT WAS NOT BREATHING WHILE ON A TRILOGY EVO DEVICE. THE PATIENT WAS NOT BREATHING AT THE TIME, SO THE NURSE TOUCHED THE RESPIRATOR TO CHECK IF IT WAS WORKING, BUT WAS SO UPSET THAT THEY DID NOT REMEMBER HOW THEY OPERATED IT. THE CAUSE OF DEATH IS REPORTED TO BE RESPIRATORY FAILURE DUE TO THE PROGRESSION OF AMYOTROPHIC LATERAL SCLEROSIS (ALS). NO DIRECT ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED WITHIN THE COMPLAINT RECORD. FURTHER ANALYSIS OF THE DEVICE HAS BEEN REQUESTED TO DETERMINE IF THERE WAS A PRESENCE OF A MALFUNCTION, HUMAN RELATED USE-ERROR, AND/OR DEVICE FULL FUNCTIONALITY. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT ON 27-JAN-2025 AT APPROXIMATELY 1019 A NURSE AT AN INSTITUTIONAL FACILITY DISCOVERED THE PATIENT WAS NOT BREATHING WHILE ON A TRILOGY EVO DEVICE. THE PATIENT WAS NOT BREATHING AT THE TIME, SO THE NURSE TOUCHED THE RESPIRATOR TO CHECK IF IT WAS WORKING, BUT WAS SO UPSET THAT THEY DID NOT REMEMBER HOW THEY OPERATED IT. THE CAUSE OF DEATH IS REPORTED TO BE RESPIRATORY FAILURE DUE TO THE PROGRESSION OF AMYOTROPHIC LATERAL SCLEROSIS (ALS). NO DIRECT ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED WITHIN THE COMPLAINT RECORD. FURTHER ANALYSIS OF THE DEVICE HAS BEEN REQUESTED TO DETERMINE IF THERE WAS A PRESENCE OF A MALFUNCTION, HUMAN RELATED USE-ERROR, AND/OR DEVICE FULL FUNCTIONALITY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, REVIEW OF THE ERROR LOG, CONFIRMED THAT THE TREATMENT OPERATION WAS STOPPED BY PRESSING A BUTTON ON THE DATE AND TIME SPECIFIED. NO ABNORMALITY IN THE DEVICE WAS CONFIRMED BY ANY OTHER INVESTIGATIONS. ADDITIONALLY, A PERIODIC MAINTENANCE 1 WAS PERFORMED. THE DEVICE'S AIR INLET FOAM FILTER AND PARTICLE FILTER WERE REPLACED TO PREVENT FAILURE. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. CONFIRMED THE SOFTWARE VERSION IS 1.06.10.00.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21344663&pc=OZO | 2032227-2025-123685 | 2024-12-26 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-10 05:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Renal Failure; Urinary Tract Infection; Diabetic Ketoacidosis; Constipation | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024, AND THE CAUSE OF DEATH WAS NATURAL CAUSES. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT ANY HEALTH ISSUES OR ILLNESS THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS URINARY TRACT INFECTION, PROBLEM WITH HER BOWEL, DIABETES KETOACIDOSIS, KIDNEYS AND LUNGS FAILING. THE PUMP WAS NOT WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE UNK_SENSOR, MMT-332A, UNOMEDICAL, AND MMT-1715KL. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. NO PRODUCT RETURN IS REQUIRED FOR UNK_SENSOR. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE DEVICE, AND NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21344878&pc=QEZ | 1319211-2025-00004 | 2025-01-09 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-10 05:00 | QEZ | ALPHAVAC | Difficult to Advance | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26 YEAR OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING. THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 2 OF 3. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT EMBOLIZATION AND SUBSEQUENT EXPIRATION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). THERE WERE REPORTS OF ALPHAVAC CANNULA MALFUNCTION DURING THE PROCEDURE, WHICH WAS DETERMINED TO BE KINK DAMAGE TO CANNULA COILS, LIKELY DUE TO HANDLING DAMAGE DURING PROCEDURE AS A RESULT OF ENGAGING THE TUMOR MATERIAL. THE ALPHAVAC DEVICE WAS BEING USED IN AN OFF-LABEL CAPACITY, I.E. FOR REMOVAL OF TUMOR MATERIAL AND NOT THROMBI OR EMBOLI. THE USE OF THE ALPHAVAC DEVICE IS NOT BELIEVED TO BE A CONTRIBUTING FACTOR TO THE PATIENT'S EXPIRATION. ALPHAVAC WAS BEING USED OFF-LABEL ALONG WITH OTHER SIMILAR ASPIRATION DEVICES (ONO SNARE, PENUMBRA INDIGO). DISTAL EMBOLIZATION OF THROMBUS AND DEATH ARE POTENTIAL/ANTICIPATED PROCEDURAL COMPLICATIONS OF AN ALPHAVAC PROCEDURE; THIS IS CAUTIONED IN THE DEVICE DFU. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. CORRECTION/CORRECTIVE ACTIONS: NO CORRECTION IS REQUIRED SINCE THERE WAS NO MANUFACTURING NON-CONFORMANCE OBSERVED DURING EVALUATION OF THE RETURNED ALPHAVAC DEVICES. THE KINKED TUBING OBSERVED WAS LIKELY A RESULT OF HANDLING DAMAGE DURING DEVICE USE TO ENGAGE WITH THE TUMOR MATERIAL. THE PATIENT THROMBOSIS/EMBOLISM AND SUBSEQUENT EXPIRATION IS NOT RELATED TO THE USE OF THE ALPHAVAC DURING THIS OFF-LABEL PROCEDURE. DISTAL EMBOLIZATION OF THROMBUS AND DEATH ARE POTENTIAL/ANTICIPATED PROCEDURAL COMPLICATIONS OF AN ALPHAVAC PROCEDURE; THIS IS CAUTIONED IN THE DEVICE DFU. THIS SERIOUS ADVERSE EVENT HAS BEEN CAPTURED IN THE POST MARKET SURVEILLANCE OF THE ALPHAVAC PRODUCT FAMILY. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE: THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: DISTAL EMBOLIZATION OF THROMBUS. PULMONARY EMBOLISM. CARDIAC ARREST. DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE: (B)(4), DEVICE 2 OF 3. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21344915&pc=QEZ | 1319211-2025-00005 | 2025-01-09 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-10 05:00 | QEZ | ALPHAVAC | Difficult to Advance | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26 YEAR OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING . THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 3 OF 3. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT EMBOLIZATION AND SUBSEQUENT EXPIRATION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). THERE WERE REPORTS OF ALPHAVAC CANNULA MALFUNCTION DURING THE PROCEDURE, WHICH WAS DETERMINED TO BE KINK DAMAGE TO CANNULA COILS, LIKELY DUE TO HANDLING DAMAGE DURING PROCEDURE AS A RESULT OF ENGAGING THE TUMOR MATERIAL. THE ALPHAVAC DEVICE WAS BEING USED IN AN OFF-LABEL CAPACITY, I.E. FOR REMOVAL OF TUMOR MATERIAL AND NOT THROMBI OR EMBOLI. THE USE OF THE ALPHAVAC DEVICE IS NOT BELIEVED TO BE A CONTRIBUTING FACTOR TO THE PATIENT'S EXPIRATION. ALPHAVAC WAS BEING USED OFF-LABEL ALONG WITH OTHER SIMILAR ASPIRATION DEVICES (ONO SNARE, PENUMBRA INDIGO). DISTAL EMBOLIZATION OF THROMBUS AND DEATH ARE POTENTIAL/ANTICIPATED PROCEDURAL COMPLICATIONS OF AN ALPHAVAC PROCEDURE; THIS IS CAUTIONED IN THE DEVICE DFU. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. CORRECTION/CORRECTIVE ACTIONS: NO CORRECTION IS REQUIRED SINCE THERE WAS NO MANUFACTURING NON-CONFORMANCE OBSERVED DURING EVALUATION OF THE RETURNED ALPHAVAC DEVICES. THE KINKED TUBING OBSERVED WAS LIKELY A RESULT OF HANDLING DAMAGE DURING DEVICE USE TO ENGAGE WITH THE TUMOR MATERIAL. THE PATIENT THROMBOSIS/EMBOLISM AND SUBSEQUENT EXPIRATION IS NOT RELATED TO THE USE OF THE ALPHAVAC DURING THIS OFF-LABEL PROCEDURE. DISTAL EMBOLIZATION OF THROMBUS AND DEATH ARE POTENTIAL/ANTICIPATED PROCEDURAL COMPLICATIONS OF AN ALPHAVAC PROCEDURE; THIS IS CAUTIONED IN THE DEVICE DFU. THIS SERIOUS ADVERSE EVENT HAS BEEN CAPTURED IN THE POST MARKET SURVEILLANCE OF THE ALPHAVAC PRODUCT FAMILY. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE: THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: DISTAL EMBOLIZATION OF THROMBUS, PULMONARY EMBOLISM, CARDIAC ARREST, OR DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4), DEVICE 3 OF 3. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21345333&pc=NPT | 2025587-2025-01127 | 2025-02-04 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-10 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Degraded; Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED A PATIENT WITH AN EXISTING 21MM NON-MEDTRONIC (ABBOTT TRIFECTA) VALVE IMPLANTED IN 2011 AND A DEGENERATED 23MM MEDTRONIC EVOLUT R VALVE IMPLANTED IN 2017, THE PATIENT UNDERWENT A TRANSCATHETER AORTIC VALVE REPLACEMENT. PER THE PHYSICIAN, THIS WAS A HIGH-RISK INTERVENTION AND THE PATIENT HAD A DIAMETER LESS THAN 16MM. DURING THE PROCEDURE, IMPLANT WAS ATTEMPTED WITH A 20MM NON-MEDTRONIC (EDWARDS) VALVE. HOWEVER, THE PATIENT DIED DURING THE PROCEDURE. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNCLEAR; HOWEVER, IT WAS LIKELY DUE TO A CORONARY OCCLUSION - A RESULT FROM THE PLACEMENT OF STENTS FOR A CHIMNEY TECHNIQUE. THE PHYSICIAN FURTHER NOTED THE MEDTRONIC VALVE CAN BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE PATIENT'S DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT APPROXIMATELY 7 YEARS AND 2 MONTHS FOLLOWING THE VALVE IMPLANT, THE TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE WAS PERFORMED.Event Description: IT WAS REPORTED A PATIENT WITH AN EXISTING 21MM NON-MEDTRONIC (ABBOTT TRIFECTA) VALVE IMPLANTED IN 2011 AND A DEGENERATED 23MM MEDTRONIC EVOLUT R VALVE IMPLANTED IN 2017, THE PATIENT UNDERWENT A TRANSCATHETER AORTIC VALVE REPLACEMENT. PER THE PHYSICIAN, THIS WAS A HIGH-RISK INTERVENTION AND THE PATIENT HAD A DIAMETER LESS THAN 16MM. DURING THE PROCEDURE, IMPLANT WAS ATTEMPTED WITH A 20MM NON-MEDTRONIC (EDWARDS) VALVE. HOWEVER, THE PATIENT DIED DURING THE PROCEDURE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: NON-MEDTRONIC - ABBOTT BRAND NAME: 21MM TRIFECTA VALVE (SERIAL: UNKNOWN); IMPLANT DATE: 2011 NON-MEDTRONIC - EDWARDS BRAND NAME: 20MM VALVE (SERIAL: UNKNOWN); IMPLANT DATE: ATTEMPTED (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE EXPIRATION DATE, SERIAL NUMBER, AND UDI NUMBER ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 D6A MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21345376&pc=NPT | 2015691-2025-00896 | 2025-01-17 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-02-10 05:00 | NPT | COMMANDER DELIVERY SYSTEM | Burst Container or Vessel; Difficult to Remove | Unspecified Vascular Problem | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST, DURING A TRANSFEMORAL AORTIC VALVE REPLACEMENT PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE, THE DELIVERY BALLOON RUPTURED HORIZONTALLY DURING DEPLOYMENT OF THE VALVE AND THE BALLOON WOULD NOT COME BACK INTO THE SHEATH. THE TEAM HAD TO CUT DOWN ON LEFT FEMORAL ARTERY TO RETRIEVE THE DISTAL END OF THE DELIVERY SYSTEM. PER FOLLOW-UP COMMUNICATION, THE BALLOON/TIP OF DELIVERY SYSTEM DAMAGED THE ARTERY. THE PATIENT DECEASED LATER THAT EVENING.Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST, DURING A TRANSFEMORAL AORTIC VALVE REPLACEMENT PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE, THE DELIVERY BALLOON RUPTURED HORIZONTALLY DURING DEPLOYMENT OF THE VALVE AND THE BALLOON WOULD NOT COME BACK INTO THE SHEATH. THE TEAM HAD TO CUT DOWN ON LEFT FEMORAL ARTERY TO RETRIEVE THE DISTAL END OF THE DELIVERY SYSTEM. PER FOLLOW-UP COMMUNICATION, THE BALLOON/TIP OF DELIVERY SYSTEM DAMAGED THE ARTERY. THE PATIENT DECEASED LATER THAT EVENING. PER THE MEDICAL RECORD, THE PATIENT PRESENTED WITH AORTIC STENOSIS PRIOR TO TAVR. AT THE TIME OF TAVR, A RIGHT FEMORAL BALLOON PUMP WAS PLACED. THE TAVR CASE COMPLEXITY INCREASED WITH RUPTURE OF BALLOON WITH DEPLOYMENT CONSISTENT WITH LIKELY AORTIC CALCIFIC DEBRIS PUNCTURING THE BALLOON UPON INFLATION. THE VALVE WAS INITIALLY UNDER-EXPANDED DUE TO THE RUPTURE, WITH SUBSEQUENT POST DILATATION BALLOON VALVULOPLASTY. THE PATIENT MAINTAINED MILD AORTIC INSUFFICIENCY ALONG THE LEFT CORONARY CUSP BY ARTERIOGRAM. WITH BALLOON RUPTURE, THERE WAS SEPARATION OF THE DISTAL NOSE CONE THAT REMAINED CONTINUOUS WITH THE DEVICE, BUT WOULD NOT RETRACT ADEQUATELY INTO THE SHEATH. ATTEMPTED SHEATH REMOVAL WITH CONTIGUOUS NOSE CONE REMOVAL WAS UNSUCCESSFUL. SIGNIFICANT TRACTION WAS ENCOUNTERED AND HALTED WITH PLANS FOR CUTDOWN AND OPEN REMOVAL. THE PATIENT UNDERWENT COMPLEX VASCULAR REPAIR WITH INTRAOPERATIVE CONSULTATION AND PROCEDURE WITH VASCULAR SURGERY. THE PATIENT DID RECEIVE 2 UNITS OF PACKED RED BLOOD CELLS EMPIRICALLY, GIVEN BLOOD LOSS AND KNOWN UNDERLYING CORONARY ARTERY DISEASE. APPROXIMATELY 30 MINUTES TO AN HOUR LATER SHE BECAME HYPOTENSIVE. A STAT ECHO WAS PERFORMED IN THE ROOM AND NO PERICARDIAI EFFUSION OF SIGNIFICANCE WAS FOUND. AN ECHOCARDIOGRAM WAS REPEATED A COUPLE OF HOURS LATER WHICH AGAIN SHOWED NO EVIDENCE OF TAMPONADE. THE PATIENT WAS STARTED ON MULTIPLE PRESSORS. THE PATIENT SEEMED TO BE SUFFERING FROM PROFOUND VASOPLEGIA OF UNCLEAR ETIOLOGY. HER CONDITION CONTINUED TO DETERIORATE AND EVENTUALLY SHE WAS ON MULTIPLE PRESSER THERAPY. SHE WAS PLACED ON DO NOT RESUSCITATED STATUS BUT THE MEDICAL TEAM CONTINUED FULL ACTIVE TREATMENT AS THE VASOPLEGIA MIGHT RESOLVE AND SHE MIGHT BE ABLE TO RECOVER TO BE ABLE TO HAVE BYPASS SURGERY AT A LATER TIME. UNFORTUNATELY, AFTER THE ADMINISTRATION OF ANGIOTENSIN II, SHE HAD AN ABRUPT DROP IN BLOOD PRESSURE THAT NIGHT AND DECEASED. AFTER THE DEVICES WERE RETURNED, THE DISTAL TIP OF THE RETURNED SHEATH WAS SPLIT AS OBSERVED PRIOR TO DECONTAMINATION. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED, FOLLOWING COMPLETION OF THE ENGINEERING EVALUATION. B4, G3, G6, H2, AND H3 HAVE BEEN UPDATED. H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS, AND H11 HAVE BEEN CORRECTED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS RETURNED FOR EVALUATION. THE DELIVERY SYSTEM WAS RETURNED WITH THE LOCKED POSITION BETWEEN WARNING MARKER WITH 0 FLEX AND 25% FINE ADJUST. THE DISTAL BALLOON UMBRELLA WAS OVER THE DISTAL NOSE TIP. THE BALLOON WAS BURST RADIALLY AND LONGITUDINALLY. NO BALLOON PIECES WERE MISSING. THE PROXIMAL BALLOON SHAFT WAS OBSERVED TO HAVE BEEN CUT COMPLETELY THROUGH THE GUIDEWIRE LUMEN (GWL) AT DEFAULT MARKER (PROXIMAL TO THE HANDLE) AND PROXIMAL SECTION OF THE BALLOON SHAFT WITH THE Y-CONNECTOR WAS NOT RETURNED. A DISTAL BALLOON REMNANT REMAINED ATTACHED WITH THE DISTAL TIP SUBASSEMBLY (NOSE CONE, DISTAL SHOULDER, GWL) AND WAS OBSERVED TO HAVE BEEN REMOVED FROM THE BALLOON SHAFT SUBASSEMBLY PRIOR TO RETURN, LIKELY TO PERFORM RECAPTURE OF THE DISTAL END OF THE BURST BALLOON DURING PROCEDURE. APPROXIMATELY 10.5 INCHES FROM THE DISTAL HANDLE FLEX SHAFT WAS CUT THROUGH TO THE BALLOON SHAFT AND THE BALLOON SHAFT CRACKED BRAIDING WAS EXPOSED. LACERATION MARKS WERE OBSERVED, LIKELY DUE TO CUSTOMER ATTEMPTING TO CUT THE FLEX SHAFT. THE FULL LOADER ASSEMBLY AND A 14F ESHEATH+ WERE INSERTED OVER THE FLEX SHAFT. STRETCHING WAS OBSERVED ON THE BALLOON SPRING COIL. THE INFLATION BALLOON SINGLE WALL THICKNESS WAS MEASURED ALONG THE EDGES OF THE BURST LOCATION. A THICKNESS DEVIATING FROM THE SPECIFICATION COULD BE INDICATIVE OF AN ISSUE DURING MANUFACTURING OF THE COMPONENT, AS A THIN WALL COULD CONTRIBUTE TO THE OBSERVED BURST. ALL MEASUREMENTS MET THE SPECIFICATION. DUE TO THE NATURE OF THE COMPLAINT, NO APPLICABLE FUNCTIONAL TESTING WAS PERFORMED. THE COMPLAINT WAS CONFIRMED THROUGH VISUAL EXAMINATION. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, BALLOON BURST EVENTS DURING VALVE DEPLOYMENT HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO CALCIFICATION IN THE LANDING ZONE OR OVER-INFLATION OF THE BALLOON. IN THIS CASE, THE BALLOON BURST WAS CONFIRMED BASED ON EVALUATION OF THE RETURNED DEVICE. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THE SINOTUBULAR JUNCTION WAS NARROW WITH SEVERE CALCIFICATION. THE PRESENCE OF CALCIFICATION CAN CREATE A CHALLENGING ANATOMY FOR BALLOON INFLATION. WHILE THE BALLOONS ARE SUFFICIENTLY DESIGNED AND TESTED TO ENSURE THE BURST PRESSURE IS AT OR ABOVE THE RATED BURST PRESSURE, CALCIFIED NODULES CAN COMPROMISE THE STRUCTURE OF THE BALLOON WALL VIA FOLLOWING MECHANISMS SUCH AS PUNCTURE, LOCAL OVERSTRETCHING, OPEN CELL IMPINGEMENT, OR STRESS CONCENTRATION. THE COMPLAINT FOR THE INABILITY TO WITHDRAW SYSTEM THROUGH SHEATH' WAS CONFIRMED BASED ON EVALUATION OF THE RETURNED DEVICE. AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (ALTERED BALLOON PROFILE) LIKELY CONTRIBUTED TO THE EVENT. AS THE BALLOON WAS BURST, THE ALTERED BALLOON PROFILE COULD HAVE MADE IT MORE SUSCEPTIBLE TO CATCH ON THE DISTAL TIP OF THE SHEATH OR PATIENT VASCULATURE, WHICH WOULD HAVE THEN LED TO THE EXPERIENCED RETRIEVAL DIFFICULTY. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION OR PRODUCT RISK ASSESSMENT IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION FOR THIS REPORT IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED, DUE TO MEDICAL RECORDS RECEIVED. B4, D9, G3, AND H2 HAVE BEEN UPDATED. B5, B7, H6: HEALTH EFFECT IMPACT, H10, AND H11 HAVE BEEN CORRECTED. THIS IS ONE OF TWO MANUFACTURER REPORTS REGARDING THIS EVENT. THE INVESTIGATION FOR THIS REPORT IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21345418&pc=NKM | 2135147-2025-00651 | 2025-01-16 05:00 | Death | ABBOTT MEDICAL | 2025-02-10 05:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH GRADE 4 FUNCTIONAL MITRAL REGURGITATION (MR) FOR A MITRACLIP PROCEDURE. ALL STEPS WERE PERFORMED PER INSTRUCTIONS FOR USE (IFU). DURING THE PROCEDURE, THE STEERABLE GUIDE CATHETER (SGC) WAS INSERTED AND THEN THE CLIP DELIVERY SYSTEM (CDS). DURING THE POSITIONING OF THE DEVICE IN THE LEFT ATRIUM (THE CLIP NEVER CROSSED THE MITRAL VALVE AND NEVER DESCENDED INTO THE VENTRICLE); THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH HYPOTENSION. THERE WAS NO SIGN OF PERICARDIAL EFFUSION DETECTED DURING TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) AND TRANSTHORACIC ECHOCARDIOGRAM (TTE). IT WAS DECIDED TO HALT THE PROCEDURE. THE MITRACLIP DEVICE WAS REMOVED FROM THE ANATOMY. ONCE REMOVED, RESUSCITATION MANEUVERS WERE PERFORMED ACCORDING TO GUIDELINES. AFTER MORE THAN 30 MINUTES OF RESUSCITATION MANEUVERS, WITHOUT OBTAINING ANY RESPONSE FROM THE PATIENT, THE PHYSICIANS DECIDED TO DISCONTINUE. THE PATIENT'S DEATH WAS CONFIRMED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR REPORTED DEATH, HYPOTENSION, AND CARDIAC ARREST. DEATH, HYPOTENSION, AND CARDIAC ARREST ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS RESULT OF CASE SPECIFIC CIRCUMSTANCE AS RESUSCITATION MANEUVERS WERE PERFORMED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21345581&pc=DSQ | 2916596-2025-00500 | 2023-12-21 05:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Multiple Organ Failure; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS CHRONICALLY ILL AND DEVELOPED PROGRESSIVE RIGHT VENTRICULAR FAILURE AND MULTIPLE SYSTEM ORGAN FAILURE LEADING TO DEATH. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL AND INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MANUFACTURER REFERENCE NUMBER: 2916596-2025-01498. Manufacturer Narrative: SECTION D4: DEVICE MANUFACTURE DATE CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS UPDATED WITH ALL OF THE CURRENT INFORMATION AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21345931&pc=PNJ | 2017865-2025-12745 | 2025-01-21 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-10 05:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P150035 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-12746. IT WAS REPORTED THAT DURING AVEIR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE, THE PATIENT EXPERIENCED SUDDEN DROP IN BLOOD PRESSURE. IT WAS ELECTED TO REMOVE THE LP SYSTEM AND ABANDON THE PROCEDURE. THE PATIENT WAS LATER PRONOUNCED DECEASED. IT WAS NOTED THAT THE PHYSICIAN STATED THAT THE PATIENT'S DEATH WAS NOT DUE TO OR RELATED TO THE PROCEDURE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21345950&pc=PNJ | 2017865-2025-12746 | 2025-01-21 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-10 05:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P150035 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-12745. IT WAS REPORTED THAT DURING AVEIR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE, THE PATIENT EXPERIENCED SUDDEN DROP IN BLOOD PRESSURE. IT WAS ELECTED TO REMOVE THE LP SYSTEM AND ABANDON THE PROCEDURE. THE PATIENT WAS LATER PRONOUNCED DECEASED. IT WAS NOTED THAT THE PHYSICIAN STATED THAT THE PATIENT'S DEATH WAS NOT DUE TO OR RELATED TO THE PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21346238&pc=DSQ | 2916596-2025-00744 | 2025-01-24 05:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Post Operative Wound Infection; Multiple Organ Failure; Skin Infection | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. NO OTHER INFORMATION WAS PROVIDED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD BEEN DIAGNOSED WITH CHRONIC DRIVELINE/PUMP INFECTION ALONG WITH BACTEREMIA. THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED AND IT WOULD NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE ONSET OF THE REPORTED INFECTION; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (DRIVELINE, PUMP POCKET, LOCALIZED), MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. VARIOUS SECTIONS OF THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ABOUT PREVENTING INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21346489&pc=NIP | 2024168-2025-01469 | 2019-04-01 04:00 | Death | ABBOTT VASCULAR | 2025-02-10 05:00 | NIP | SUPERA | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P120020 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH SUMMARY ARTICLE THAT THE PURPOSE WAS TO COMPARE THE SAFETY AND EFFICACY OF SUPERA INTERWOVEN NITINOL STENT (IWNS) AND NON-ABBOTT DRUG ELUTING STENT (DES) IN SEVERELY CALCIFIED FEMOROPOPLITEAL (FP) LESIONS. 257 CONSECUTIVE PATIENTS WHO UNDERWENT ENDOVASCULAR THERAPY USING EITHER IWNS (N = 123) OR DES (N = 134) FOR FP LESIONS WITH PERIPHERAL ARTERIAL CALCIUM SCORING SYSTEM (PACSS) GRADE 3 OR 4 SEVERE CALCIFICATION BETWEEN APRIL 2018 AND DECEMBER 2021 AT EIGHT CARDIOVASCULAR CENTERS IN JAPAN. PROPENSITY SCORE (PS) MATCHING EXTRACTED 138 MATCHED PATIENTS WITH NO REMARKABLE INTERGROUP DIFFERENCE IN PATIENT 1- AND LESION CHARACTERISTICS. THE YEAR PRIMARY PATENCY RATES IN THE MATCHED POPULATION WERE NOT SIGNIFICANTLY DIFFERENT BETWEEN THE IWNS AND DES GROUPS (85.4% VS. 89.8%, P = 0.320). A SIGNIFICANT INTERACTION BETWEEN THE STENTS USED AND THE NUMBER OF BELOW-THE- KNEE (BTK) RUNOFF VESSELS WAS OBSERVED (INTERACTION P = 0.048). THE HAZARD RATIO FOR RESTENOSIS WAS 2.68 0.51¿ (95% CONFIDENCE INTERVAL, 14.2) IN THE GROUP WITH NO BTK RUNOFF, FAVORING DES. PATIENTS EXPERIENCED RESTENOSIS, TARGET LESION REVASCULARIZATION (TLR), AMPUTATION AND DEATH. NO THROMBOTIC OCCLUSION WAS OBSERVED AFTER EVT WITH IWNS. IT WAS FOUND THAT 1 -YEAR PRIMARY PATENCY WAS NOT SIGNIFICANTLY DIFFERENT BETWEEN TREATMENTS USING SUPERA IWNS AND ELUVIA DES. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "COMPARATIVE OUTCOMES OF SUPERA INTERWOVEN NITINOL VERSUS ELUVIA FLUOROPOLYMER-BASED DRUG-ELUTING STENTS FOR THE TREATMENT OF SEVERELY CALCIFIED FEMOROPOPLITEAL ARTERY LESIONS: RESULTS OF THE ELDORADO STUDY." Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A PRODUCTION RECORD REVIEW AND A REVIEW OF THE COMPLAINT HANDLING DATABASE WAS NOT PERFORMED BECAUSE THE PART AND LOT NUMBER WAS NOT REPORTED. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE SUPERA PERIPHERAL STENT SYSTEMS INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EFFECT OF PERIPHERAL PERCUTANEOUS INTERVENTION. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT(S). LITERATURE ATTACHMENT: ARTICLE TITLE ¿COMPARATIVE OUTCOMES OF SUPERA INTERWOVEN NITINOL VERSUS ELUVIA FLUOROPOLYMER-BASED DRUG-ELUTING STENTS FOR THE TREATMENT OF SEVERELY CALCIFIED FEMOROPOPLITEAL ARTERY LESIONS: RESULTS OF THE ELDORADO STUDY¿ B3: DATE OF EVENT IS ESTIMATED AS 04/01/2019. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21346518&pc=NPT | 2025587-2025-01134 | 2025-02-03 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-10 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE SUCCESSFUL TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE USING A 26 MM EVOLUTFX, THE PATIENT DEVELOPED SEVERE HYPOTENSION APPROXIMATELY 40 MINUTES AFTER THE CONCLUSION OF THE PROCEDURE. DESPITE THE MEDICAL TEAM'S EFFORTS, THE PATIENT WAS UNABLE TO BE STABILIZED AND SUBSEQUENTLY DIED.Event Description: IT WAS REPORTED THAT FOLLOWING THE SUCCESSFUL TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING A 26 MM EVOLUTFX, THE PATIENT DEVELOPED SEVERE HYPOTENSION APPROXIMATELY 40 MINUTES AFTER THE CONCLUSION OF THE PROCEDURE. DESPITE THE MEDICAL TEAM'S EFFORTS, THE PATIENT WAS UNABLE TO BE STABILIZED AND SUBSEQUENTLY DIED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE TAVI PROCEDURE WAS URGENT AND WAS PERFORMED WITH GOOD ACUTE RESULTS. IT WAS FURTHER REPORTED THAT UPON LEAVING THE OPERATING ROOM, THE PATIENT EXPERIENCED ACUTE HYPOTENSION AND DESATURATION, PROMPTING THE INITIATION OF ADVANCED RESUSCITATION MANEUVERS, INCLUDING ENDOTRACHEAL INTUBATION. AN ECHOCARDIOGRAM RULED OUT PERICARDIAL EFFUSION. THE EVOLUT FX VALVE WAS NOTED TO BE IN AN ORTHOTOPIC POSITION, AND THEN PULSELESS ELECTRICAL ACTIVITY WAS OBSERVED. THORACO-ABDOMINAL ANGIOGRAPHY SHOWED NO EVIDENCE OF DISSECTION OR ACTIVE BLEEDING. AFTER TWENTY-FIVE MINUTES OF RESUSCITATION EFFORTS, THE PATIENT EXPERIENCED REFRACTORY VENTRICULAR FIBRILLATION (VF) UNRESPONSIVE TO DEFIBRILLATION SHOCK. FOLLOWING A COLLEGIAL CONSULTATION, RESUSCITATION MANEUVERS WERE SUSPENDED AND THE PATIENT SUBSEQUENTLY DIED. ACCORDING TO THE IMPLANTING PHYSICIAN, SUBACUTE CORONARY OCCLUSION DUE TO HEAVY CALCIFICATIONS OF THE LEAFLETS AND HEAVY CALCIUM AT NON-CORONARY-RIGHT CUSP RAPHE WAS THE CAUSE OF THE HYPOTENSION, WHILE THE CAUSE OF DEATH WAS DUE TO ELECTROMECHANICAL DISSOCIATION AND IRREVERSIBLE VF. ADDITIONALLY, THE PHYSICIAN STATED THAT THE EVOLUT FX VALVE AND DELIVERY CATHETER SYSTEM CAN BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH AND THEY THINK IT WAS A PROCEDURE-R ELATED COMPLICATION. IT IS UNKNOWN IF AN AUTOPSY WAS REQUESTED OR PERFORMED. Manufacturer Narrative: UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: D-EVOLUTFX-2329; PRODUCT LOT NUMBER: UNKNOWN; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.+- Manufacturer Narrative: ADDITIONAL INFORMATION: SECTION B2 - ADDITIONAL OUTCOMES. SECTION B5 - SECOND PARAGRAPH. SECTION H6 - ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349137&pc=DSQ | 2916596-2025-00641 | 2025-01-22 05:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED ON (B)(6) 2024 PASSED AWAY IN THE HOSPITAL ON (B)(6) 2025. THE PATIENT HAD RIGHT SIDED FAILURE LEADING TO MULTIORGAN FAILURE. NO EQUIPMENT WAS RETURNING. ADDITIONAL INFORMATION PROVIDED THAT THE PATIENT'S DEATH WAS CONSIDERED TO BE A PROGRESSION OF THEIR UNDERLYING CARDIAC DISEASE. THE PATIENT'S RIGHT HEART FAILURE DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HM3 LVAS. SECTION 6 "PATIENT CARE AND MANAGEMENT" DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. SECTION 6 ALSO STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP.". NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349238&pc=DSQ | 2916596-2025-00848 | 2025-01-15 05:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Tachycardia; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO VENTRICULAR TACHYCARDIA (VT). WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED AND WOULD BE RETURNED.Event Description: THE PATIENT HAD A HISTORY OF VENTRICULAR TACHYCARDIA (VT) PRIOR TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: NO DEVICE-RELATED ISSUES WITH THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) WERE REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE-RELATED ISSUES WERE REPORTED, SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349336&pc=LYJ | 1644487-2025-00145 | 2025-01-05 05:00 | Death | CYBERONICS - HOUSTON | 2025-02-10 05:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS NOTED THAT THE PATIENT DIED UNEXPECTEDLY ON (B)(6) 2025 AND THE CAUSE REMAINS UNKNOWN AT THIS TIME. THE PATIENT WAS RECEIVING VNS THERAPY AT THE TIME OF THEIR DEATH. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349416&pc=DSQ | 2916596-2025-00697 | 2024-07-15 04:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICLE (RV) FAILURE AND MULTI-SYSTEM ORGAN FAILURE (MSOF) ON (B)(6) 2024. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED OR RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RIGHT VENTRICULAR FAILURE, MULTI SYSTEM ORGAN FAILURE, AND THE PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN FURTHER INFORMATION FROM THE CUSTOMER REGARDING THE REPORTED EVENTS; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿), ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349425&pc=DSQ | 3007042319-2025-00481 | 2025-02-02 05:00 | Death | HEARTWARE, INC. | 2025-02-10 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Electrical /Electronic Property Problem; Pumping Stopped; Use of Device Problem; Battery Problem; Connection Problem; Material Twisted/Bent; Power Problem; Pumping Problem | Cardiac Arrest; Hypoxia; Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) WAS DRIVING THE CAR WHEN THE PATIENT RECEIVED A CONTROLLER FAULT ALARM DUE TO THE INTERNAL BATTERY. THE PATIENT AND A FAMILY MEMBER ATTEMPTED TO CHANGE THE CONTROLLER. DURING THE ATTEMPTS THE PATIENT SPOUSE SAID THEY COULDN'T GET THE DRIVELINE (DL) TO SEAT ALL THE WAY IN BECAUSE THE DL COVER WAS NOT PULLED BACK ALL THE WAY. IT WAS NOT STUCK THEY JUST HAD PANICKED BECAUSE IT WASN'T SEATED. THE CONTROLLER ALSO EXHIBITED AN ELECTRICAL FAULT ALARM DUE TO THE DRIVELINE NOT BEING SEATED PROPERLY. THE PATIENT WAS NOT SUCCESSFUL EXCHANGING THE CONTROLLER, AND THE VAD WAS OFF FOR APPROXIMATELY 40 MINS. EMERGENCY MEDICAL SERVICE (EMS) ARRIVED AND RESTARTED THE VAD USING THE CONTROLLER THAT HAD THE FAULT ALARM. THE PATIENT SUSTAINED AN ANOXIC INJURY AND BECAME UNRESPONSIVE. THE CONTROLLER FAULT ALARM WAS PERMANENTLY SILENCED. THE PATIENT WAS INTUBATED DUE TO THE INABILITY TO SUPPORT THEIR OWN AIRWAY AND WAS ON ANTICOAGULANT AND ACETYLSALICYLIC ACID (ASA) THERAPY, WHICH WAS THERAPE UTIC WITH A LAB VALUE OF 2.3. A COMPUTED TOMOGRAPHY (CT) BRAIN/HEAD SCAN CONFIRMED THE ANOXIC INJURY. THE PATIENT ALSO EXPERIENCED NEUROLOGICAL DYSFUNCTION AND WAS ADMINISTERED INTRAVENOUS (IV) FLUIDS AND WAS PLACED ON A VENTILATOR. THE PATIENT SUBSEQUENTLY PASSED AWAY AFTER THE VAD WAS STOPPED AT THE TIME OF WITHDRAWAL AND SUPPORT.Event Description: IT WAS FURTHER REPORTED THAT THE LOSS OF POWER WAS SUSPECTED DUE TO DOUBLED DISCONNECTED WHEN CHANGING POWERS. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. ADDITIONAL PRODUCTS: D1: CONTROLLER D4: (B)(6) D1: CONTROLLER D4: (B)(6) INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - (B)(6) H3: YES D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - (B)(6) H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: THE PUMP (B)(6) AND TWO (2) CONTROLLERS (B)(6) WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED TWO (2) CONTROLLER POWER-UP EVENTS WITH ASSOCIATED MOTOR START EVENTS LOGGED ON (B)(6) 2025 AT 10:42:02 AND AT 11:58:47, INDICATING LOSSES OF POWER TO THE CONTROLLER. THE DATA POINT RECORDED PRIOR TO THE FIRST LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 34% RELATIVE STATE OF CHARGE (RSOC) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 73% RSOC. THE DATA POINT RECORDED AFTER THE FIRST LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE DATA POINT PRIOR TO THE SECOND LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 97% RELATIVE STATE OF CHARGE (RSOC) AND NO POWER SOURCE WAS CONNECTED TO POWER PORT TWO (2). THE DATA POINT RECORDED AFTER THE SECOND LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE CONTROLLER WAS WITHOUT POWER FOR SIXTEEN (16) SECONDS AND SEVENTEEN (17) MINUTES AND THREE (3) SECONDS, RESPECTIVELY. NO ANOMALIES WERE OBSERVED LEADING UP TO THE LOSSES OF POWER. REVIEW OF THE ALARM LOG FILE REVEALED ONE (1) CONTROLLER FAULT ALARM WAS LOGGED AT 11:16:26, INDICATING AN ISSUE WITH THE INTERNAL BATTERY. THIS WAS FOLLOWED BY FOUR (4) VAD DISCONNECT ALARMS LOGGED AT 11:21:39, 11:41:01, 11:46:43, AND 11:58:55, INDICATING A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER, LIKELY DURING TROUBLESHOOTING. THE VAD DISCONNECT ALARM CLEARED AT 12:01:13, INDICATING THAT THE DRIVELINE WAS CONNECTED TO THE CONTROLLER. ADDITIONALLY, LOG FILES REVEALED ONE (1) ELECTRICAL FAULT ALARM WAS LOGGED AT 12:01:13, INDICATING THAT THE REAR STATOR WAS NOT CONNECTED, CAUSING THE PUMP TO RUN ON A SINGLE STATOR. FURTHER REVIEW OF THE ALARM LOG FILE REVEALED (1) ADDITIONAL VAD DISCONNECT ALARM WAS LOGGED AT 12:01:19, INDICATING A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER, LIKELY DUE TO TROUBLESHOOTING. LOG FILE ANALYSIS ALSO REVEALED ONE (1) LOW FLOW ALARM WAS LOGGED AT 12:11:36 AND TWO (2) ADDITIONAL CONTROLLER FAULT ALARMS WERE LOGGED AT 12:33:10 AND 18:40:57 DUE TO AN ISSUE WITH THE INTERNAL BATTERY. IN ADDITION, LOG FILES REVEAL ED THAT THE CONTROLLER HAD BEEN IN USE FOR MORE THAN TWO (2) YEARS. VISUAL EVIDENCE PROVIDED BY THE SITE INVOLVING (B)(6) REVEALED BENT PINS WITHIN THE SERIAL PORT OF THE CONTROLLER. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) COULD NOT BE PERFORMED SINCE LOG F ILES WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED ELECTRICAL FAULT ALARM, VAD DISCONNECT ALARMS, LOW FLOWS, LOSSES OF POWER, CONTROLLER FAULT ALARMS, AND BENT PIN EVENTS WERE CONFIRMED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER THE VAD WAS STOPPED AT THE TIME OF WITHDRAWAL OF SUPPORT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. THE MOST LIKELY ROOT CAUSE OF THE ELECTRICAL FAULT ALARM CAN BE ATTRIBUTED, BUT NOT LIMITED, CAN BE ATTRIBUTED, BUT NOT LIMITED, TO CONTAMINATION BY FOREIGN MATERIAL OF THE DRIVELINE CONNECTOR AND/OR A MARGINAL DRIVELINE CONNECTION AS DESCRIBED IN THE EVENT DETAILS. A POSSIBLE ROOT CAUSE OF THE LOSS OF POWER EVENTS CAN BE ATTRIBUTED TO THE DOUBLE DISCONNECTION OF POWER SOURCES, AS DESCRIBED IN THE EVENT DETAILS. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. BASED ON AN INVESTIGATION CONDUCTED UNDER CAPA PR00384004, THE ROOT CAUSE OF BENT SOCKET PINS WITHIN THE SERIAL PORT CONNECTORS IS ATTRIBUTED TO MISALIGNMENT DURING CONNECTION ATTEMPTS BETWEEN THE METAL RECEPTACLES ON THE CONTROLLER AND THE PLASTIC CONNECTOR PLUGS IN THE POWER SOURCE AND DATA CABLES. THE MISALIGNMENT RESULTS IN WEAR ON THE CONNECTOR PLUGS THAT CAN LEAD TO CONTACT BETWEEN THE CONNECTOR PLUG AND SOCKET PINS FROM THE CONTROLLER. THE SOCKET PINS MAY NOT WITHSTAND APPLIED FORCES FROM SUBSEQUENT MISALIGNED CONNECTIONS, CAUSING THE PINS TO BEND. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION, CARDIOPULMONARY ARREST, RESPIRATORY DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF NEUROLOGICA L DYSFUNCTION EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: 407652 LEAD IMPLANTED (B)(6) 2015, DDPB3D4 ICD, IMPLANTED (B)(6) 2023. H6: THE CODES PRESENT IN SECTION. H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0 D4: MODEL#: 1420 / CATALOG#: 1420/ EXPIRATION DATE: 28-FEB-2022 / SERIAL#: (B)(6) D9: NO H3: NO H4: MFG DATE: 09-FEB-2021 H5: NO ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0 D4: MODEL#: 1420 / CATALOG#: 1420/ EXPIRATION DATE: 30-SEP-2018 / SERIAL#: (B)(6), D9: NO, H3: NO, H4: MFG DATE: 30-SEP-2017, H5: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349450&pc=DSQ | 2916596-2025-00699 | 2024-10-01 04:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICLE (RV) FAILURE ON (B)(6) 2024. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED OR RETURNED. Manufacturer Narrative: CORRECTION - ADDED H6 CODE REGARDING HEMODYNAMIC INSTABILITY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RIGHT VENTRICLE FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349785&pc=DSQ | 2916596-2025-00747 | 2025-01-24 05:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. IT WAS REPORTED THE PATIENT¿S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED, AND AN AUTOPSY WILL NOT BE PERFORMED.Event Description: THERE WERE NO DEVICE ISSUES AT THE TIME OF THE OUTCOME. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT NO AUTOPSY WAS PERFORMED, AND HEARTMATE 3 LVAS SERIAL NUMBER (B)(6) WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU), REVISION D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21349884&pc=NGV | 2124215-2025-07145 | 2025-01-30 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-10 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. THE 31MM CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN A PATIENT ON (B)(6) 2025 AND THE PATIENT WAS DISCHARGED THE SAME DAY. THE PATIENT CALLED THE HOSPITAL COORDINATOR THE NEXT DAY REGARDING TAKING BLOOD THINNERS. THERE WERE NO ABNORMAL SYMPTOMS REPORTED AND THE PATIENT WAS REPORTED TO BE DOING FINE. ON (B)(6) 2025 THE PATIENT WAS FOUND DECEASED. THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21350120&pc=DSQ | 2916596-2025-00942 | 2024-12-17 05:00 | Death | THORATEC CORPORATION | 2025-02-10 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WOULD NOT BE PERFORMED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD A FALL AT HOME WHICH RESULTED IN INTRACRANIAL HEMORRHAGE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND PATIENT OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, UNDER "ANTICOAGULATION", CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21350974&pc=OZD | 1220648-2025-26076 | 2025-01-16 05:00 | Death | ABIOMED, INC. | 2025-02-10 05:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY IN CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED EPISODES OF TORSADES DE POINTES WITH DEFIBRILLATION AND DEVICE REPOSITIONING, AND PURGE PRESSURE ISSUE. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT WAS ADMITTED AND NEWLY DIAGNOSED WITH HEART FAILURE WITH REDUCED EJECTION FRACTION OF 10% AND WAS NOTED LIKELY IN CARDIOGENIC SHOCK. INPATIENT DIURESIS WAS STARTED, AND THE PATIENT WAS TRANSFERRED TO OUT FOR ADVANCED THERAPIES EVALUATION WITH PLANS WERE IN PROGRESS FOR AN IMPELLA 5.5 PLACEMENT AND WORKUP FOR DURABLE LEFT VENTRICULAR ASSIST DEVICE VERSUS HEART TRANSPLANT. THE PATIENT WAS SEDATED AND INTUBATED ON ARRIVAL, AND THE IMPELLA 5.5 DEVICE WAS PLACED VIA RIGHT AXILLARY ARTERY WITHOUT COMPLICATION. PERFORMANCE LEVEL STARTED AT P-2 AND GRADUALLY INCREASED TO P-8, ACHIEVING FLOWS OF 4.8 L/MIN. HEMODYNAMICS WERE STABLE THROUGHOUT PROCEDURE. THE PATIENT WAS EXTUBATED AT END OF CASE AND TRANSFERRED TO UNIT ON MCS WITH MULTIPLE DRIPS INFUSING (EPINEPHRINE 4 MCG/MIN, CARDENE 5 MG/HR. MILRINONE 0.25 MCG/KG/MIN). APPROXIMATELY 14 DAYS LATER, THE PATIENT WAS AWAKE, ALERT AND SITTING IN THE CHAIR WHEN EPISODE OF SUSTAINED TORSADES WAS EXPERIENCED. THE PATIENT WAS DEFIBRILLATED A TOTAL OF SEVEN TIMES THIS DAY AND THE IMPELLA 5.5 PUMP WAS REPOSITIONED. THE PATIENT CONTINUED ON SUPPORT, WAS PLACED ON VENOATERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AND WAS BEING WORKED UP FOR HEART TRANSPLANT LAST NOTED AS A STATUS 2. THE FOLLOWING DAY, HIGH PURGE PRESSURE WAS ENCOUNTERED; PURGE FLOW WENT MISSING AND AFTER A MINUTE RETURNED AT 3.4ML/HR. TISSUE PLASMINOGEN ACTIVATOR WAS ADDED, PERFORMANCE LEVEL WENT BACK UP TO P-2 AND THEN RAMPED BACK UP NEEDED SUPPORT LEVEL. FOLLOWING, 8 DAYS LATER, THE PATIENT WOULD EXPIRE; CARE WAS WITHDRAWN. ADDITIONAL DETAILS SURROUNDING THE PATIENT'S DEMISE WAS NOT PROVIDED. IMPELLA RELATED EVENTS ARE UNABLE TO BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S OUTCOME. IT IS OF NOTE, HOWEVER, THE PATIENT HAD SEVERE HEART FAILURE (EJECTION FRACTION 10%), SUBSEQUENTLY REQUIRED ADDITIONAL SUPPORT FROM ECMO, PENDING A TRANSPLANT AND CONSIDERING COULD LIKELY BE THE PRIMARY FOR THE DEMISE OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ Manufacturer Narrative: THE INVESTIGATION OF ARRYTHMIA AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL REPORT, THE PATIENT HAD AN EPISODE OF SUSTAINED TORSADES, AND THE IMPELLA WAS REPOSITIONED. THE NEXT DAY, THE PURGE FLOW DROPPED TO 0, AND TISSUE PLASMINOGEN ACTIVATOR WAS GIVEN PER PROTOCOL. ADDITIONALLY, THE PURGE CASSETTE WAS REPLACED AFTER REPORTED LOW PURGE VOLUME ALARM AND AN AIR IN PURGE ALARM. THE DATA LOG SHOWED A GRADUAL RISE IN PURGE PRESSURE, WHICH WAS RESOLVED WITH A GRADUAL DECREASE TREND OVER SIX HOURS. ALSO, THE AIR IN THE PURGE SYSTEM ALARM WAS OBSERVED AFTER THE COMPLETED BAG CHANGE PROCEDURE. THE DOCTOR INITIATED THE DE-AIR PROCEDURE, BUT IT DID NOT RESOLVE THE AIR IN THE PURGE CASSETTE ALARM, AND THE PURGE CASSETTE HAD TO BE REPLACED. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON DATA LOG REVIEW AND THE PROVIDED CLINICAL DETAILS. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE ARRYTHMIA WAS NOT DETERMINED. B.2 REVISED AS AN INCORRECT SELECTION WAS REPORTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26076. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26076 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21351104&pc=DXF | 2124215-2025-07214 | 2025-01-21 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-10 05:00 | DXF | VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Bradycardia | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. IT WAS REPORTED THAT VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE WAS SELECTED FOR A PULSE FIELD ABLATION (PFA) PROCEDURE. THE PHYSICIAN COMPLETED AN OFF-LABEL CTI ABLATION AND POSTERIOR WALL ABLATION. NITRO AND NEO WAS GIVEN PRIOR TO ABLATING. THERE WERE NO ISSUES WITH TRANSSEPTAL, AND THE PATIENT WAS STABLE. AFTER ABLATING, THE PHYSICIAN CHECKED A HIS WHICH WAS 105. THE CASE WAS CONCLUDED, AND ALL CATHETERS WERE WITHDRAWN OUT OF THE BODY. THE PATIENT WAS CLOSED, AND ANESTHESIA WAS TOLD TO WAKE THE PATIENT UP. THE PATIENT WOULD NEED TO COME BACK LATER FOR A PACEMAKER. IT WAS THEN REPORTED THAT THE PATIENT WAS HAVING BRADYCARDIA EPISODES, A TRANSVENOUS PACING WIRES WERE PLACED AND AN ARTERIAL LINE. AN ULTRASOUND WAS PERFORMED AND RULED OUT ANY PERICARDIAL EFFUSION. THE PATIENT HAD A LOW EJECTION FRACTION (EF) PRIOR TO ABLATION. A SWAN CATHETER WAS PLACED, AND THE PATIENT WAS SENT TO INTENSIVE CARE UNIT (ICU) WHERE WENT INTO VENTRICULAR TACHYCARDIA AND DID NOT MAKE IT. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. THERE WAS NO DEVICE MALFUNCTIONS REPORTED, NOR COMPLICATIONS DURING THE WHOLE PROCEDURE. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21351254&pc=MVK | 3008642652-2025-01485 | 2024-11-02 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-10 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 12:30:43 ON (B)(6) 2024. AT 23:24:42 ON (B)(6) 2024 AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 23:25:26, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 23:44:29 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21351258&pc=MVK | 3008642652-2025-01486 | 2025-01-11 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-10 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 04:46:58 ON (B)(6) 2025. AT 05:59:19, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 120 BPM WITH PVC¿S. THE RHYTHM THEN DEGRADED TO VF. AT 05:59:57, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF/FINE VF WITH CPR/MOTION ARTIFACT/VARYING AMPLITUDES. THE ELECTRODE BELT WAS DISCONNECTED AT 06:12:23 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21341036&pc=QEZ | 1317056-2025-00021 | 2025-01-09 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-09 05:00 | QEZ | ALPHAVAC | Difficult to Advance | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26-YEAR-OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT ((B)(4)) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION ((B)(4)). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. DAVE ELISON PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE ((B)(4)). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING. THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY, THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO, THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 3 OF 3. Manufacturer Narrative: CORRECTION: INITIAL MEDWATCH WAS INADVERTENTLY SUBMITTED UNDER AN INCORRECT MFR REPORT NUMBER. THE INITIAL WILL BE RESUBMITTED UNDER THE CORRECT MFR REPORT NUMBER, 1319211-2025-00005. Manufacturer Narrative: THE DEVICES WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3. (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3. (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21341037&pc=QEZ | 1317056-2025-00020 | 2025-01-09 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-09 05:00 | QEZ | ALPHAVAC | Difficult to Advance | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26-YEAR-OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT ((B)(4)) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION ((B)(4)). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. DAVE ELISON PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE ((B)(4)). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING. THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY, THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO, THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 2 OF 3. Manufacturer Narrative: CORRECTION: INITIAL MEDWATCH WAS INADVERTENTLY SUBMITTED UNDER AN INCORRECT MFR REPORT NUMBER. THE INITIAL WILL BE RESUBMITTED UNDER THE CORRECT MFR REPORT NUMBER, 1319211-2025-00004. . Manufacturer Narrative: THE DEVICES WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3. (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3. (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21341038&pc=QEZ | 1317056-2025-00019 | 2025-01-09 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-09 05:00 | QEZ | ALPHAVAC | Difficult to Advance | Pulmonary Embolism | K211081 |  | 1 | Event Description: AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26-YEAR-OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING. THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY, THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO, THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 1 OF 3. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3. (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21330274&pc=FRN | 2016493-2025-61537 | 2025-01-13 05:00 | Death | CAREFUSION SD | 2025-02-07 05:00 | FRN | ALARIS SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT A TECH CALLED IN TO REPORT A PATIENT INCIDENT WHERE IT WAS NOTED THAT THE PATIENT PASSED AWAY. AT THE TIME, THE TECH DIDN'T HAVE ALL THE DEVICES INVOLVED AND SAID THEY WOULD CALL BACK WITH THE OTHER DEVICE INFORMATION AND WHETHER THEY WANTED THE INCIDENT INVESTIGATED. HOWEVER, IT WAS NOTED THAT THE TECH NEVER CALLED BACK. AFTER NUMEROUS ATTEMPTS OF REACHING OUT TO THE TECH, THEY NEVER RESPONDED.Event Description: IT WAS REPORTED THAT A PATIENT "PASSED AWAY" BUT NO INFORMATION ON THE DEVICE WAS PROVIDED AT THE TIME. ALTHOUGH MULTIPLE ATTEMPTS HAVE BEEN MADE, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED TO THE MANUFACTURER TO DATE AND NO DEVICE HAS BEEN RETURNED. Manufacturer Narrative: THE ACTUAL DATE OF EVENT IS UNKNOWN. THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: CORRECTION: DESCRIBE EVENT OR PROBLEM ADDITIONAL INFORMATION: IMDRF ANNEX B CODE AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. ROOT CAUSE: UNABLE TO DETERMINE THE ROOT CAUSE. ALTHOUGH MULTIPLE REQUESTS HAVE BEEN MADE, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED, AND NO PRODUCT HAS BEEN RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21330878&pc=PNJ | 9612164-2025-00656 | 2025-01-31 05:00 | Death | MEDTRONIC IRELAND | 2025-02-07 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | P150033 |  | 1 | Event Description: IT WAS REPORTED DURING IMPLANT THAT THE PATIENT BECAME HYPOTENSIVE AFTER THE INTRODUCER WAS INITIALLY PLACED AND INTRA-ABDOMINAL BLEEDING WAS NOTED. MEDICATION WAS REQUIRED TO IMPROVE THE BLOOD PRESSURE. CATECHOLAMINES WERE PRODUCED AND THE PATIENT APPEARED IN BETTER CONDITION FOR SEVERAL MINUTES. THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) WAS PLACED AND THE PATIENT BECAME HYPOTENSIVE AGAIN. AN ECHO WAS PERFORMED AND BLOOD WAS OBSERVED IN THE RIGHT HIP. THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21330881&pc=DYB | 9612164-2025-00657 | 2025-01-31 05:00 | Death | MEDTRONIC IRELAND | 2025-02-07 05:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | K132030 |  | 1 | Event Description: IT WAS REPORTED DURING IMPLANT THAT THE PATIENT BECAME HYPOTENSIVE AFTER THE INTRODUCER WAS INITIALLY PLACED AND INTRA-ABDOMINAL BLEEDING WAS NOTED. MEDICATION WAS REQUIRED TO IMPROVE THE BLOOD PRESSURE. CATECHOLAMINES WERE PRODUCED AND THE PATIENT APPEARED IN BETTER CONDITION FOR SEVERAL MINUTES. THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) WAS PLACED AND THE PATIENT BECAME HYPOTENSIVE AGAIN. AN ECHO WAS PERFORMED AND BLOOD WAS OBSERVED IN THE RIGHT HIP. THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21331580&pc=MVK | 3008642652-2025-01413 | 2025-01-12 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-07 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR TREATMENTS FROM THE LIFEVEST. THE DEVICE WAS STARTED UP AT 11:37:58 ON (B)(6) 2025. AT 20:29:43, AN ARRHYTHMIA WAS DETECTED. ECG IS NOT AVAILABLE. BETWEEN 20:31:04 AND 20:58:15, THE PATIENT RECEIVED FOUR TREATMENTS FROM THE LIFEVEST. THE RHYTHM AT THE TIME OF THESE TREATMENTS IS UNAVAILABLE. THE POST SHOCK RHYTHM OF THESE TREATMENTS IS UNAVAILABLE. THE ELECTRODE BELT WAS DISCONNECTED AT 21:16:04 ON (B)(6) 2025. PER REVIEW OF THE HOLTER DATA, PATIENT IS LAST SEEN IN SINUS RHYTHM @ 70 BPM WHEN THE HOLTER DATA ENDS AT 16:01:37 ON (B)(6) 2025. REPORTING OUT OF AN ABUNDANCE OF CAUTION DUE TO THE ECG NOT BEING AVAILABLE AT THE TIME OF THE TREATMENTS. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21331856&pc=DSQ | 2916596-2025-00515 | 2025-01-19 05:00 | Death | THORATEC CORPORATION | 2025-02-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Mechanical Problem; Pumping Stopped; Infusion or Flow Problem | Apnea; Loss of consciousness; Convulsion/Seizure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A QUESTIONABLE PUMP STOP, AND THE REASON WAS UNKNOWN LOG FILES WOULD BE SUBMITTED FOR SUSPICIONS OF DRIVELINE DISCONNECT OR A SHORT TO SHIELD. LOG FILES WERE SUBMITTED FOR REVIEW WHICH CAPTURED LOW FLOW EVENTS. IT WAS NOTED BY TECH SERVICES THAT THE PATIENT MAY HAVE BEEN SYMPTOMATIC AND A COMPUTED TOMOGRAPHY (CT) ANGIOGRAPHY OR ECHOCARDIOGRAM (ECHO) WAS PERFORMED TO CONFIRM PRESENCE OF AN OBSTRUCTION IN THE VENTRICULAR ASSIST DEVICE (VAD) CIRCUIT. CONTINUED MONITORING OF THE PATIENT WAS RECOMMENDED. IT WAS ALSO NOTED A CONTROLLER CLOCK CORRUPT MESSAGE WAS SHOWING IN THE LOG FILES DUE TO THE BACKUP BATTERY NOT BEING INSTALLED. WHEN THE BATTERY WAS INSTALLED, THE CLOCK AND DATES WERE RE-SET. IT WAS ADDITIONALLY REPORTED THAT A SYSTEM CONTROLLER EXCHANGE WAS PERFORMED AND LOG FILES FROM PRIOR TO THE EXCHANGE WERE SUBMITTED FOR REVIEW. LOG FILES CAPTURED A DRIVELINE DISCONNECT EVENT ON (B)(6) 2025 AT 05:22:44. PRIOR TO THE DRIVELINE DISCONNECT THERE WERE 3 LOW POWER ADVISORIES DUE TO NORMAL BATTERY DEPLETION. POWER WAS DISCONNECTED FROM THE SYSTEM CONTROLLER LEADS ON (B)(6) 2025 AT 06:43:52 AND NO EXTERNAL POWER EVENTS WERE CAPTURED. THE BACKUP BATTERY (EBB) POWERED THE CONTROLLER FROM 06:43:52 TO 14:08:00. AFTER THE 14:08:00 TIME STAMP THE EBB WAS DRAINED AND THE CONTROLLER CLOCK WAS RESET TO THE DEFAULT TIME STAMP. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD SEIZURE-LIKE ACTIVITY, FOLLOWED BY A PERIOD OF APNEA AND UNRESPONSIVENESS. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION (CPR) IN THE NURSING FACILITY, FOLLOWED BY 3-4 MORE EPISODES OF CPR IN THE OUTSIDE HOSPITAL (OSH) WITH UNKNOWN TOTAL CPR TIME. THE PATIENT WAS REPORTED TO HAVE "PUMP STOP" ALARMS, "DRIVELINE DISCONNECT" ALARMS, AND "LOW FLOW" ALARMS. THE PATIENT'S SPOUSE EXCHANGED THE SYSTEM CONTROLLER, AND ALL ALARMS REPORTEDLY RESOLVED EXCEPT FOR "LOW FLOW" ALARMS AND "REPLACE BACKUP BATTERY" ALARMS. THE SYSTEM CONTROLLER EXCHANGE WAS PERFORMED MORE THAN HOUR INTO THE PATIENT'S "EPISODE". A CT SCAN OF THE HEAD AND AN ECHO WAS PERFORMED DUE TO CONTINUED LOW FLOW ALARMS AT THE OSH IN ADDITION TO THE PATIENT BEING INTUBATED WITH PRESSOR SUPPORT. AT THE TIME OF THE ECHO THE LOW FLOW ALARMS HAD RESOLVED. THE PATIENT HAD A HISTORY OF EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO) WITH STENT PLACEMENT (MFR # 2916596-2024-02939). A CHEST CT ANGIOGRAM WAS ALSO PERFORMED WHICH SHOWED STABLE APPEARANCE OF LVAD OUTFLOW TRACT GRAFT AND STENT, WITH SOME MURAL THROMBUS WITHIN THE GRAFT SIMILAR TO THE PRIOR EXAM. NO HIGH-GRADE STENOSIS WAS IDENTIFIED. SOME MASS EFFECT UPON THE RIGHT VENTRICLE BY THE GRAFT WAS STABLE. THE PATIENT WAS TRANSPORTED TO THE IMPLANTING CENTER VIA LIFE FLIGHT. ONCE AT IMPLANTING CENTER, IT WAS DETERMINED THAT PATIENT'S SYSTEM CONTROLLER DID NOT HAVE AN EBB INSTALLED. THIS WAS REPLACED WITH A NEW EBB. LOGFILES WERE OBTAINED AND THE PATIENT WAS TAKEN TO CATH LAB WHERE AN IMPELLA RP FLEX WAS PLACED FOR RIGHT VENTRICULAR SUPPORT. THE PATIENT REMAINED IN CRITICAL CONDITION. BUT DID NOT APPEAR TO BE NEUROLOGICALLY INTACT. THE HEAD CT WAS NEGATIVE. OBSTRUCTION OF OUTFLOW GRAFT FROM EOGO DOES NOT APPEAR TO BE THE INCITING ISSUE. LOGFILES HAD ONLY BEEN PARTIALLY HELPFUL BECAUSE THE ORIGINAL SYSTEM CONTROLLER SHOWED THAT IT RAN ON EBB UNTIL THE BATTERY FULLY DEPLETED AND THEN THE TIME REVERTED TO (B)(6) 2000. FOLLOWING THE SYSTEM CONTROLLER EXCHANGE, THE BACK-UP SYSTEM CONTROLLER SHOWED SIGNS OF BEING TAMPERED WITH AND DID NOT CONTAIN ITS EBB. THIS ALSO RESULTED IN INACCURATE TIME CODES, IN ADDITION TO THE MULTIPLE STORIES RECEIVED FROM TWO OTHER FACILITIES AND A LIFE FLIGHT TRANSPORT TEAM. THE PATIENT'S PRESSORS WERE DISCONTINUED ON (B)(6) 2025 AND THE IMPELLA WAS REMOVED ON (B)(6) 2025. THE PATIENT REMAINED INTUBATED AND UNRESPONSIVE, WITH NO GAG REFLEX AND PINPOINT PUPILS BILATERALLY. THE PATIENT HAS NO PURPOSEFUL OR NON-PURPOSEFUL MOVEMENT. GIVEN THE NEUROLOGICAL FINDINGS, THE PATIENT'S PROGNOSIS WAS POOR.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS OF NEUROLOGICAL DYSFUNCTION AND RIGHT HEART FAILURE, AS WELL AS THE REPORTED PATIENT OUTCOME, COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED A DRIVELINE DISCONNECT ALARM AND SUBSEQUENT PUMP STOP EVENT. ALTHOUGH A SPECIFIC CAUSE FOR THIS FINDING COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION, THE ACCOUNT INDICATED THAT THE PATIENT'S SPOUSE PERFORMED A SYSTEM CONTROLLER EXCHANGE. REVIEW OF THE LOG FILES ALSO CONFIRMED MULTIPLE LOW FLOW HAZARD ALARMS FOLLOWING THE SYSTEM CONTROLLER EXCHANGE; HOWEVER, A SPECIFIC CAUSE FOR THESE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED. THE SYSTEM CONTROLLER EVENT LOG FILES FROM (B)(6) CONTAINED RELEVANT EVENTS FROM (B)(6) 2025, PER THE TIMESTAMPS. ON (B)(6) 2025, A DRIVELINE DISCONNECT ALARM AND SUBSEQUENT PUMP STOP WAS CAPTURED. THIS EVENT APPEARED CONSISTENT WITH THE DRIVELINE BEING PHYSICALLY DISCONNECTED. THE DRIVELINE WAS DISCONNECTED FOR THE REMAINDER OF THE FILE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP OPERATED AT THE SET SPEED THROUGHOUT THE ENTIRETY OF THE FILES WHILE THE DRIVELINE WAS CONNECTED. THE SYSTEM CONTROLLER EVENT LOG FILES FROM (B)(6) CONTAINED EVENTS FROM (B)(6) 2025. ADDITIONAL EVENTS PRIOR TO THIS TIMEFRAME WERE ALSO CAPTURED; HOWEVER, THE SYSTEM CONTROLLER CLOCK WAS CORRUPT DURING THESE EVENTS. EVENTS ASSOCIATED WITH THE DRIVELINE CONNECTION TO (B)(6) WERE OVERWRITTEN BY NEWER INCOMING EVENTS, PER DESIGN. DUE THE CORRUPT TIMESTAMPS AND OVERWRITTEN EVENTS, THE EXACT LENGTH OF TIME IN WHICH THE PUMP WAS STOPPED DUE TO THE DRIVELINE BEING DISCONNECTED FROM (B)(6) IS UNKNOWN. MULTIPLE, TRANSIENT LOW FLOW HAZARD ALARMS WERE CAPTURED PRIOR TO THE SYSTEM CONTROLLER CLOCK BEING SYNCHRONIZED ON (B)(6) 2025. ADDITIONAL, TRANSIENT LOW FLOW HAZARD ALARMS WERE CAPTURED ON (B)(6) 2025. A DRIVELINE DISCONNECT ALARM WAS ALSO CAPTURED ON (B)(6) 2025, CONSISTENT WITH THE REMOVAL OF DEVICE SUPPORT FOLLOWING THE PATIENT'S DEATH. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP OPERATED AT THE SET SPEED THROUGHOUT THE ENTIRETY OF THE FILES WHILE THE DRIVELINE WAS CONNECTED. THE ACCOUNT INDICATED THAT AN AUTOPSY WAS NOT PERFORMED. IT WAS ALSO NOTED THAT THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. CHAPTER 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED) AND RIGHT HEART FAILURE. CHAPTER 2 OF THE IFU, "SYSTEM OPERATIONS" (UNDER "SYSTEM CONTROLLER WARNINGS AND CAUTIONS"), STATES "CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION." ADDITIONALLY, THIS CHAPTER PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. CHAPTER 4 OF THE IFU, ¿HEARTMATE TOUCH¿ COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. SECTION 2 OF THE PATIENT HANDBOOK, ¿HOW YOUR HEART PUMP WORKS¿, ALSO ADVISES THE USER TO ¿CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR OFTEN TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP WILL STOP. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION¿. THE PATIENT HANDBOOK ALSO EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. CHAPTER 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS ALL SYSTEM ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD AND DRIVELINE DISCONNECTED ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21331949&pc=NVN | 2649622-2025-03688 | 2025-01-19 05:00 | Death | MPRI | 2025-02-07 05:00 | NVN | CAPSURE FIX MRI | Pacing Problem; Under-Sensing | No Clinical Signs, Symptoms or Conditions | P090013 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS EXHIBITED UNDERSENSING WHICH RESULTED IN INAPPROPRIATE RV PACING. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED WITH CAUSE OF DEATH UNKNOWN. Manufacturer Narrative: CORRECTED: B2 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: W1DR01 IPG; IMPLANT DATE: (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21331963&pc=NVN | 2649622-2025-03689 | 2025-01-19 05:00 | Death | MPRI | 2025-02-07 05:00 | NVN | CAPSURE FIX MRI | Pacing Problem; Under-Sensing | No Clinical Signs, Symptoms or Conditions | P090013 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS EXHIBITED UNDERSENSING WHICH RESULTED IN INAPPROPRIATE RV PACING. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED WITH CAUSE OF DEATH UNKNOWN. Manufacturer Narrative: CORRECTED: B2 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: W1DR01 IPG; IMPLANT DATE: (B)(6) 2020 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21332055&pc=MHX | 1218950-2025-000086 | 2025-01-09 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-02-07 05:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Cardiac Arrest | K183387 |  | 1 | Event Description: IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHETHER AN ALARM WAS ACTIVELY ACKNOWLEDGED RELATED TO A PATIENT DEATH. IT WAS INDICATED THE ALARMS FOR 'ECG LEADS OFF' WERE SILENCED, WHICH LED TO A DELAY IN CARE. THE PATIENT WAS FOUND AND SUCCESSFULLY RESUSCITATED BUT LATER PASSED AWAY; HOWEVER, THE CAUSE OF DEATH REMAINS UNKNOWN. NO OTHER CLINICAL INFORMATION WAS PROVIDED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL SPECIALIST REVIEWED THE LOGS PROVIDED. THE CLINICAL AUDIT LOG PROVIDES A CHRONOLOGICAL RECORD OF THE ALARMS AND ACTIONS. THE CLINICAL AUDIT LOG SHOWS FREQUENT SPO2 INOPS ALARM DURING THE TIME IN QUESTION, SUCH AS SPO2 SENSOR OFF, SPO2 NO PULSE, SPO2 PULSE? AND SPO2 NOISY SIGNAL. THESE ALARMS WERE ACKNOWLEDGED AT THE PIC IX. INOPS ARE TECHNICAL ALARMS, THEY INDICATE THAT THE MONITOR CANNOT MEASURE OR DETECT ALARM CONDITIONS RELIABLY. IF AN INOP INTERRUPTS MONITORING AND ALARM DETECTION (FOR EXAMPLE, LEADS OFF), THE MONITOR PLACES A QUESTION MARK IN PLACE OF THE MEASUREMENT NUMERIC AND AN AUDIBLE INDICATOR TONE WILL BE SOUNDED. THESE INOPS APPEAR IN THE CLINICAL AUDIT LOG. CANNOT ANALYZE ECG INOPS ARE ALSO SHOWN IN THE LOG STARTING AT 07:05:45 AND TRANSITED TO !!ECG LEADS OFF AT 7:33:52. CANNOT ANALYZE ECG INOP IS GENERATED IF 20 OF THE LAST 30 SECONDS ARE CLASSIFIED AS EITHER NOISY OR QUESTIONABLE. DURING THIS INOP/TECHNICAL CONDITION, THE ARRHYTHMIA ANALYSIS CONTINUES AND AN ALARM WILL BE ANNOUNCED IF AN ALARM CONDITION IS MET. HOWEVER, THE CANNOT ANALYZE ECG INOP/TECHNICAL ALARM INDICATES THAT THE EFFECTIVENESS OF THE ARRHYTHMIA MONITORING FOR THE PATIENT IS COMPROMISED AND A QUICK RESPONSE TO THIS ALARM IS IMPORTANT.!!! ECG LEADS OFF INDICATES THAT THE REQUIRED ECG LEADS ARE NOT AVAILABLE. AT 7:37:02, ALL ALARMS WERE PAUSED AT THE MONITOR. IN ADDITION, TESTS CARRIED OUT ON SITE CONFIRMED THAT THE DEVICE WAS FUNCTIONING WITHIN MANUFACTURER'S SPECIFICATIONS AND ARE READY FOR CLINICAL USE. BASED ON A REVIEW OF THE INFORMATION PROVIDED AND THE TESTING CONDUCTED ON SITE, THE PIC IX DEVICE WAS FUNCTIONING AS DESIGNED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. A4: WEIGHT GIVEN BY CUSTOMER AS OVERWEIGHT. REPORTING INSTITUTION PHONE #: (B)(6). REPORTER PHONE #: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21332333&pc=OZD | 1220648-2025-26048 | 2025-01-20 05:00 | Death | ABIOMED, INC. | 2025-02-07 05:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Fluid/Blood Leak; Increase in Pressure | Arrhythmia; Stroke/CVA; Septic Shock; Thrombocytopenia; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION WAS INITIALLY IMPLANTED WITH AN IMPELLA CP AND ESCALATED TO AN IMPELLA 5.5 DEVICE. THREE DAYS AFTER START OF SUPPORT A SMALL LEAK WAS NOTED AT DISTAL END OF THE PURGE SIDE ARM WHERE IT CONNECTS TO THE PATIENT SIDE OF THE DRIVELINE. PURGE FLOW AND PURGE PRESSURE ARE WITHIN NORMAL LIMITS. THE LEAK IS BEING MONITORED. THE PATIENT CONTINUES ON MECHANICAL CIRCULATORY SUPPORT WITH THE DEVICE AND IS NOTED TO BE IN STABLE CONDITION. THERE WAS NO REPORT OR INDICATION OF PATIENT HARM AS A RESULT OF THIS EVENT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE CLINICAL SITE THAT THE PATIENT WAS ON IMPELLA 5.5 SUPPORT FOR OVER TWO MONTHS. THE PATIENT WAS NOT A CANDIDATE FOR ADVANCED THERAPIES. WHILE ON IMPELLA SUPPORT THE PATIENT EXPERIENCED ARRHYTHMIAS, THROMBOCYTOPENIA, AND SEPTIC SHOCK WAS SUSPECTED REQUIRING LEVO TO BE INCREASED. THE PATIENT HAD A STROKE THAT OCCURRED AFTER STARTING ON AN ANTICOAGULANT. THE IMPELLA ALSO ENCOUNTERED SOME SUCTIONS AND P-LEVEL WAS LOWER. THE IMPELLA HAD HIGH PURGE PRESSURE ALARMS THAT WERE RESOLVED WITH LYSIS THERAPY. THE PATIENT EXPIRED WHILE ON IMPELLA SUPPORT. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS THE DEVICE REMAINS IMPLANTED, SUPPORTING THE PATIENT AT THE TIME OF THE MDR WRITING, AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF PURGE LEAK ¿ PUMP HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOG REVIEW CONFIRMED THAT THE PURGE FLOW AND PURGE PRESSURE REMAINED WITHIN NORMAL LIMITS. THE ROOT CAUSE OF THE PUMP PURGE LEAK WAS MOST LIKELY A DAMAGED SIDEARM BUT THE EXACT LOCATION OF THE LEAK IS UNKNOWN WITHOUT THE RETURNED PRODUCT. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK; SECTION: CLEANING ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿ A.5 REVISED RACE AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26048. D.4 REVISED UNIQUE IDENTIFIER (UDI) # AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26048. H.10 ADDED INSTRUCTIONS FOR USE AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26048. Manufacturer Narrative: B1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B2 REVISED TO REFLECT DEATH AND REQUIRED INTERVENTION BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B5 REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. D6B EXPLANT DATE ADDED H1 TYPE OF EVENT REVISED TO DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE, AND MEDICAL DEVICE PROBLEM CODE REVISED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21332567&pc=OZD | 1220648-2025-26050 | 2025-01-17 05:00 | Death | ABIOMED, INC. | 2025-02-07 05:00 | OZD | IMPELLA CP | Restricted Flow rate; Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED LIMB ISCHEMIA AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT WAS FLU POSITIVE AND WITH ELEVATED TROPONINS. INITIALLY DIAGNOSED AS NON-ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND ADMITTED TO THE TELEMETRY UNIT. THE PATIENT CONVERTED TO SINUS RHYTHM SPONTANEOUSLY, AND THREE DAYS LATER ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION WAS IDENTIFIED AND CONFIRMED. THE PATIENT UNDERWENT A LEFT HEART CATHETERIZATION. UPON ENGAGING CORONARIES (LEFT MAIN, LEFT ANTERIOR DESCENDING [LAD], LEFT CIRCUMFLEX [LCX] AND RIGHT CORONARY ARTERIES), THE PATIENT BECAME HYPOTENSIVE, LEVO STARTED, AND DECISION WAS MADE TO PLACE IMPELLA CP. AFTER THE IMPELLA CP WAS IMPLANTED, THE PATIENT'S BLOOD PRESSURE "SIGNIFICANTLY" IMPROVED, AND THE PHYSICIAN PROCEEDED WITH THE PERCUTANEOUS CORONARY INTERVENTION. CARDIOTHORACIC SURGERY WAS CONSULTED FOR LAD AND LCX TARGETS. WHILE IN CATHETERIZATION LAB, THE PHYSICIAN INITIALLY PREFERRED TO LEAVE IN THE PEEL-AWAY SHEATH. THE SHEATH SHOT SHOWED NO FLOW DISTALLY. PEEL-AWAY SHEATH WAS EXCHANGED FOR REPOSITIONING SHEATH. A SHEATH SHOT WAS REPEATED THROUGH RE-ACCESS PORT AND SHOWED FLOW. HOWEVER, THE FLOW WAS SLUGGISH. THE PATIENT WAS IMMEDIATELY TAKEN TO THE OPERATING ROOM. CORONARY ARTERY BYPASS GRAFT SURGERY (CABG) X2 WAS PERFORMED. THE PATIENT WAS ON MULTIPLE VASOACTIVE DRUGS. THE PHYSICIAN INSPECTED THE RIGHT LEG, SAW MOTTLING AND FELT COOL TEMPERATURE. ABIOMED SUPPORT RECOMMENDED TRANSITIONING TO IMPELLA 5.5 SINCE CHEST WAS STILL OPEN AND THERE WAS OPTION FOR DIRECT INSERTION. HOWEVER, THE PHYSICIAN OPTED TO PERFORM AN EXTERNAL FEM-FEM BYPASS. AN 8 FR. RETROGRADE ON THE LEFT FEMORAL SUPPLYING BLOOD 6 FR ANTEGRADE VIA RIGHT USING MALE TO MALE CONNECTOR. INITIALLY RIGHT FOOT HAD A PALPABLE PULSE. AS THE PHYSICIAN COMPLETED THE CABG, THE PATIENT CONTINUED TO BE HYPOTENSIVE WHICH RESULTED IN INCREASED USE OF VASOACTIVE MEDICATIONS. THE RIGHT VENTRICLE FUNCTION WAS NOTED TO BE IMPAIRED. AN IMPELLA RP WAS UNAVAILABLE. DOBUTAMINE AND MILRINONE WERE ADDED FOR SUPPORT. THE IMPELLA CP POSITION WAS CONFIRMED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM, AND THE PATIENT IN CRITICAL CONDITION WAS TRANSFERRED TO THE UNIT. APPROXIMATELY SIX DAYS LATER THE RIGHT FOOT WAS NOTED WITHOUT PULSE. FEM-FEM BYPASS IN PLACE. THE NEXT DAY POSITIVE POPLITEAL PULSE ON RIGHT EXTREMITY, NO DORSALIS PEDIS PULSE WAS NOTED. FIVE DAYS FOLLOWING AN IMPELLA RP FLEX WAS INSERTED. THE PATIENT NOW ON BIPELLA SUPPORT. THE NEXT DAY CONTINUOUS RENAL REPLACEMENT THERAPY WAS NOTED. ADDITIONALLY, NO FLOW IN THE RIGHT LOWER EXTREMITY WAS NOTED AND THE PHYSICIAN WAS UNABLE TO FLOW ABOVE P-5 WITHOUT SUCTION. IT WAS STATED THAT THE PATIENT CONTINUED TO PROGRESSIVELY DECLINE WITH NO PROGRESS RESULTING IN MULTIPLE ORGAN DYSFUNCTION (MOD) SYNDROME. FAMILY ULTIMATELY DECIDED TO WITHDRAW ALL LIFE-SUSTAINING MEASURES. ALL THERAPIES WERE STOPPED, AND THE PATIENT EXPIRED. THE PATIENT NEVER FULLY RECOVERED FROM THE CARDIAC EVENT. THE PATIENT WAS IN A POOR STATE RESULTING TO MOD AND EXPIRED. THE PATIENT HAD SIGNIFICANT LIMB ISCHEMIA PRIOR TO THE INSERTION OF THE RP FLEX DEVICE. LIMB ISCHEMIA WAS TREATED WITH FEM-FEM BYPASS SO THE IMPLANT WAS ALREADY QUITE SEVERE. THE IMPELLA RELATED EVENT FOR THE LIMB ISCHEMIA IS THE IMPELLA CP. AND THE IMPELLA CP RELATED EVENTS/PERIODS OF LIMB ISCHEMIA WITH SURGICAL PROCEDURE FEM-FEM BYPASS CANNOT BE EXCLUDED AS POSSIBLY CONTRIBUTING TO THE OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B.1 REVISED TO REFLECT ADVERSE EVENT ONLY AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26050. B.2 REVISED SELECTION AS INCORRECT SELECTIONS WERE ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26050. B.7 ADDED INFORMATION THAT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26050. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26050 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 HEALTH EFFECT - IMPACT CODE FOR DEATH WAS ADDED AS IT WAS INADVERTENTLYOMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26050. CODE 1248 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26050. A NEW CODE WAS ADDED TO MEDICAL DEVICE PROBLEM CODE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21332694&pc=GWM | 9617601-2025-00153 | 2023-02-27 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-02-07 05:00 | GWM | UNKNOWN EXTERNAL DRAINAGE AND MONITORING SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: KANEEZ ZAHRA, RICARDO A. DOMINGO, MARION T. TURNBULL, CHRISTAN D. SANTOS, SARAH H. PEACOCK, DANIEL A. JACKSON, RABIH G. TAWK, JASON L. SIEGEL, AND WILLIAM DAVID FREEMAN. SAFETY AND TOLERABILITY OF CONCENTRATED LNTRAVENTRICULAR NICARDIPINE FOR POOR-GRADE ANEURYSMAL SUBARACHNOID HEMORRHAGE-RELATED VASOSPASM. JOURNAL OF PERSONALIZED MEDICINE 13 (2023). DOI: 10.3390/ JPM13030428. ABSTRACT: OBJECTIVE: TO REPORT THE PRELIMINARY SAFETY, TOLERABILITY, AND CEREBRAL SPINAL FLUID (CSF) SAMPLING UTILITY OF SERIAL INJECTIONS OF CONCENTRATED INTRAVENTRICULAR NICARDIPINE (IVN) IN THE TREATMENT OF ANEURYSMAL SUBARACHNOID HEMORRHAGE (ASAH). METHODS: WE REPORT THE CLINICAL, RADIOGRAPHIC, AND LABORATORY SAFETY AND TOLERABILITY DATA OF A RETROSPECTIVE CASE SERIES FROM A SINGLE ACADEMIC MEDICAL CENTER. ALL PATIENTS WITH ASAH DEVELOPED VASOSPASM DESPITE ENTERAL NIMODIPINE AND RECEIVED SERIAL INJECTIONS OF CONCENTRATED IVN (2.5 MG/ML). CSF INJECTION SAFETY, TOLERABILITY, AND UTILITY ARE DEFINED AND REPORTED. RESULTS: A TOTAL OF 59 DOSES OF CONCENTRATED IVN WERE ADMINISTERED TO THREE PATIENTS WITH POOR-GRADE SAH. IN CASE 1, A 33-YEAR-OLD MAN WITH MODIFIED FISHER SCALE (MFS) GRADE 4 AND HUNT-HESS SCALE (HH) SCORE 4 RECEIVED 26 DOSES; IN CASE 2, A 36-YEAR-OLD WOMAN WITH MFS GRADE 4 AND HH SCORE 5 RECEIVED 13 DOSES; AND IN CASE 3, A 70-YEAR-OLD WOMAN WITH MFS GRADE 3 AND HH SCORE 4 RECEIVED 20 DOSES. NO MAJOR SAFETY OR TOLERABILITY EVENTS OCCURRED. TWO PATIENTS WERE DISCHARGED TO A REHABILITATION FACILITY, AND ONE DIED AFTER DISCHARGE FROM THE HOSPITAL. CONCLUSIONS: A CONCENTRATED 4 MG IVN DOSE (2.5 MG/ML) IN A 1.6 ML INJECTION APPEARS RELATIVELY SAFE AND TOLERABLE AND POTENTIALLY OFFERS A SECOND-LINE STRATEGYFOR TREATING REFRACTORY VASOSPASM IN POOR-GRADE SAH WITHOUT COMPROMISING INTRACRANIAL PRESSURE OR CEREBRAL PERFUSION PR ESSURE. REPORTED EVENT - A 36-YEAR-OLD WOMAN WITH A PAST MEDICAL HISTORY OF SICKLE CELL ANEMIA, CHRONIC BACK PAIN, AND HEADACHES PRESENTED WITH ASAH (MFS GRADE 4, HH SCORE 4) (FIGURE 4). ON ARRIVAL, HER GCS SCORE WAS 3 {EL, VI, ML), AND SHE HAD AN UNREACTIVE 7 MM RIGHT PUPIL. HEAD CT AND CTA REVEALED DIFFUSE ASAH BLOOD, BLOOD IN THE FOURTH VENTRICLE, AND OBSTRUCTIVE HYDROCEPHALUS WITH A BASILAR TIP ANEURYSM. THE EPICENTER OF THE ASAH WAS PREDOMINATELY AROUND THE MIDBRAIN, SUGGESTING THE BASILAR ANEURYSM AS THE SOURCE OF BLEEDING (FIGURE 2). THE PATIENT WAS INTUBATED, AND SHE UNDERWENT CEREBRAL ANGIOGRAPHY WITH STENT-ASSISTED COILING OF THE BASILAR TIP ANEURYSM. BILATERAL EVDS WERE PLACED, AND SHE WAS STARTED ON 60 MG OF NIMODIPINE EVERY 4 H ENTERALLY VIA NASOGASTRIC TUBE. HOWEVER, DUE TO HYPOTENSION, THE PATIENT'S NIMODIPINE DOSE WAS REDUCED TO 30 MG EVERY 2 H. ON POST-ASAH DAY 5, SHE WAS NOTED WITH VASOSPASM ON TCDS (LEFT PROXIMAL MCA 24.2 CM/S AND RIGHT MCA 24.8 CM/S). CTA REVEALED ML STENOSIS, AND CT PERFUSION SHOWED WATERSHED DISTRIBUTION INFARCTS. THE PATIENT WAS STARTED ON CONCENTRATED IVN {2.5 MG/ ML) EVERY 8 H, WITH TOTAL VOLUMES AND PRE
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21332708&pc=DSQ | 2916596-2025-00884 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "PERI-IMPLANT RENAL REPLACEMENT THERAPY AND MORTALITY IN A CONTEMPORARY MULTICENTER REGISTRY OF HEARTMATE 3 LVAD RECIPIENTS" THAT HEARTMATE 3 MAY BE ASSOCIATED WITH RENAL DYSFUNCTION AND DEATH. THE STUDY AIMED TO EXAMINE THE ASSOCIATION BETWEEN EARLY NEED RENAL REPLACEMENT THERAPY (RRT) AND LONG-TERM SURVIVAL IN A CONTEMPORARY, MULTICENTER COHORT OF PATIENTS IMPLANTED WITH A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE STUDY INCLUDED CONSECUTIVE PATIENTS IMPLANTED WITH A HEARTMATE 3 LVAD AT 6 U.S. CENTERS BETWEEN 2015-2022 WITH DATA AVAILABLE. EARLY RRT WAS DEFINED AS RRT AT THE TIME OF LVAD IMPLANTATION OR WITHIN 1 MONTH OF IMPLANT. PATIENTS WHO DIED WITHIN 1 MONTH (N=28) WERE EXCLUDED. 496 PATIENTS (AGE 55 ± 13 YEARS, 26% FEMALE, 40% BLACK, 39% WHITE, 45% INTERMACS PROFILE 1 OR 2) WITH MEAN FOLLOW-UP OF 2.2 ± 1.6 YEARS WERE STUDIED. THIRTY-SIX (7.3%) PATIENTS MET CRITERIA FOR EARLY RRT. THERE WERE NO SIGNIFICANT DIFFERENCES IN AGE, SEX, RACE, OR INTERMACS PROFILE BETWEEN EARLY RRT AND NON-EARLY RRT PATIENTS. KAPLAN-MEIER SURVIVAL ANALYSIS SHOWED THAT EARLY RRT PATIENTS HAD SIGNIFICANTLY LOWER SURVIVAL AT 2-YEARS (65%) COMPARED TO NON-EARLY RRT PATIENTS (85%) (P=0.0038 FOR OVERALL DIFFERENCE DURING FOLLOW-UP. Manufacturer Narrative: CORRECTION: E2, E3, G2 MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2022 AS PATIENTS WERE IMPLANTED BETWEEN 2015 AND 2022. AUTHOR INFORMATION: GENUARDI, M., ZHAO, M., MOSS, N., RADAKRISHNAN, A., KILIC, A., WELCH, B., SAEED, O., MENDAPARA, P., JAWAID, A., SHAH, S., MARECKI, G., MOIN, D., CHAWLA, S., MEHTA, M., ATLURI, P., & VIDULA, H. (2024). PERI-IMPLANT RENAL REPLACEMENT THERAPY AND MORTALITY IN A CONTEMPORARY MULTICENTER REGISTRY OF HEARTMATE 3 LVAD RECIPIENTS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S396. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1280. UNIVERSITY OF PENNSYLVANIA, PHILADELPHIA, PA; MOUNT SINAI HOSPITAL, NEW YORK, NY; MEDICAL UNIVERSITY OF SOUTH CAROLINA, CHARLESTON, SC; MONTEFIORE MEDICAL CTR, NEW YORK, NY; UT SOUTHWESTERN, DALLAS, TX; NORTHWELL, NEW YORK, NY; NORTHWELL HEALTH, NEW YORK, NY; NYU LANGONE HEALTH, NEW YORK, NY; ROBERT WOOD JOHNSON, NEW BRUNSWICK, NJ; TUFTS UNIVERSITY, BOSTON, MA. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21333528&pc=MHX | 1218950-2025-000070 | 2025-01-11 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-02-07 05:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Ventricular Fibrillation | K183387 |  | 1 | Event Description: IT WAS REPORTED THAT A (B)(6) Y/O MALE PATIENT WAS BEING MONITORED WHEN INSTANCES OF V-FIB DID NOT ALARM ON THE MONITOR AND THE PATIENT PASSED AWAY. THE PATIENT HAD AN IMPLANTED ICD/PACEMAKER, WHICH ATTEMPTED TO PACE UNSUCCESSFULLY. THE EVENT OCCURRED IN THE MED SURG UNIT AND THE SPEAKERS WERE TESTED FOR FUNCTIONALITY. THE REMAINING QUESTIONS WERE RESPONDED TO WITH ASKU (ASKED BUT UNKNOWN), NOT ALL CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. AT THIS TIME IT CANNOT BE DETERMINED IF THIS DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. Manufacturer Narrative: DIAGNOSTIC/FUNCTIONAL TESTING WAS PERFORMED PER A PHILIPS FIELD SERVICE ENGINEER (FSE) ONSITE. THE FSE COLLECTED LOG FILES AND DATA FOR THE EVENT AND VERIFIED THE INTELLIVUE MX40 AND THE PATIENT INFORMATION CENTER IX (PIC IX) ENTERPRISE SYSTEM FOR PROPER OPERATION. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION TO THE PHILIPS SENIOR PRODUCT SUPPORT ENGINEER AND PHILIPS CLINICAL PRODUCT SPECIALIST TO VERIFY IF THE (B)(6) MX40 SN: (B)(6) PATIENT INFORMATION CENTER IX FUNCTIONED AS DESIGNED AND DID ALARM. THE INTELLIVUE MX40 AND THE PATIENT INFORMATION CENTER IX (PIC IX) DEVICE WERE CONFIRMED TO HAVE ALARMED FOR THE PATIENT¿S ECG CHANGES (BOTH YELLOW AND RED LEVEL ALARMS). THE EQUIPMENT WAS WORKING AS DESIGNED, CONFIGURED, AND OPERATED. THE UNACKNOWLEDGED ASYSTOLE ALARM SUPPRESSED VF/VT DETECTION. NO PRODUCT MALFUNCTION WAS DETECTED. THE RESULTS OF THE CLINICAL AUDIT LOG REVIEW SHOW: THE AUDIT LOGS SHOW THAT FOR THE 1-HOUR TIMEFRAME ON (B)(6)2025, FROM 13:00:00 UNTIL 14:00:00: A RED LEVEL CARDIAC ALARM WAS ACTIVE FOR 0:41:36 (41 MINUTES AND 36 SECONDS: ASYSTOLE FOR 40 MINUTES AND 13 SECONDS, VENTRICULAR FIBRILLATION/TACHYCARDIA FOR 1 MINUTE AND 23 SECONDS). A YELLOW LEVEL CARDIAC ALARM (IRREGULAR HR OR PAUSE) WAS ACTIVE FOR 15 MINUTES. THE AUDIT LOGS ALSO REVEAL: ISSUES IMPACTING ECG ANALYSIS: CANNOT ANALYZE ECG INOPERATIVE CONDITION FOR 27 MINUTES AND 29 SECONDS. LL LEAD OFF INOPERATIVE CONDITION FOR 20 MINUTES AND 26 SECONDS. ECG LEADS OFF (NO MONITORING POSSIBLE) FOR 3 MINUTES AND 26 SECONDS. BASED ON THE INFORMATION AVAILABLE AND THE RESULT OF ADDITIONAL ANALYSIS CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS DUE TO USER KNOWLEDGE. THE REPORTED PROBLEM WAS NOT CONFIRMED. ANALYSIS WAS PERFORMED AND INVESTIGATION HAS BEEN COMPLETED. INFORMATION WAS PROVIDED TO THE CUSTOMER OF THE ANALYSIS FINDINGS AND CONFIRMED THE DEVICE WAS WORKING TO ITS SPECIFICATION. THERE IS NO DEVICE MALFUNCTION AND THE REPORTED ISSUE WAS DUE TO USER KNOWLEDGE. THE DEVICE REMAINS AT THE CUSTOMER SITE. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. - PLEASE SEE SPLIT CASE FOR INTELLIVUE MX40 (PRODUCT PROBLEM) REPORTED UNDER MFR REPORT NUMBER 1218950-2025-000075. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21333638&pc=NKM | 2135147-2025-00591 | 2018-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-02-07 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "HYPERACUTE EFFECTS OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR ON LEFT VENTRICULAR VOLUMES AND FUNCTIONS" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE IMMEDIATE EFFECTS OF TEER ON LV VOLUMES AND FUNCTIONS, AS WELL AS THEIR INFLUENCE ON MID-TERM OUTCOMES, USING HIGH-RESOLUTION 3D TRANSESOPHAGEAL ECHOCARDIOGRAPHY. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THE ACUTE DECLINE IN LV FUNCTION FOLLOWING TEER CORRELATES WITH THE DEGREE OF MR REDUCTION, WITH GREATER IMPACTS OBSERVED IN CIRCUMFERENTIAL FUNCTION AND PATIENTS WITH HIGHER BASELINE LVEF. RELATIVE LVEF REDUCTION IS A CRITICAL ECHOCARDIOGRAPHIC PREDICTOR OF MORTALITY. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NICOLAS BRUGGER AT INSELSPITAL, BERN UNIVERSITY HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL NICOLAS.BRUGGER@INSEL.CH]. THE TIME FRAME OF THE STUDY WAS JANUARY 2018 TO DECEMBER 2020. A TOTAL OF 80 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 80 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED CARDIOGENIC SHOCK, DIABETES, HYPERTENSION, PREVIOUS STROKE. (B)(4) UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: HYPERACUTE EFFECTS OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR ON LEFT VENTRICULAR VOLUMES AND FUNCTIONS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING CARDIOGENIC SHOCK, DIABETES, HYPERTENSION, PREVIOUS STROKE. COMPLICATIONS REPORTED INCLUDED DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21333873&pc=QZI | 2124215-2025-06512 | 2024-12-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-07 05:00 | QZI | FARAWAVE | Off-Label Use | Respiratory Failure; Ischemia Stroke; Thrombosis/Thrombus |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A FARAWAVE PULSED FIELD ABLATION CATHETER, THE PATIENT WAS DISCHARGED THE SAME DAY. HOWEVER, THE PATIENT THEN ARRIVED TO THE HOSPITAL WITH AN ISCHEMIC STROKE WHERE CLOT WAS FROM BRAIN. INTUBATION WAS REQUIRED DUE TO AGITATION, THE PATIENT HAD A SEVERE CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND PASSED AWAY FROM THE COMPLICATIONS DUE TO INTUBATION. THE USE OF THE FARAWAVE CATHETER TO PERFORM ABLATION ON THE POSTERIOR WALL IS CONSIDERED AN OFF-LABEL USE. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS AS IT WAS DISPOSED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334119&pc=NPS | 2135147-2025-00606 | 2024-12-19 05:00 | Death | ABBOTT MEDICAL | 2025-02-07 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "ACCURACY OF LONE THREE-DIMENSIONAL TRANSTHORACIC ECHOCARDIOGRAPHY IN TRICUSPID VALVE FUNCTION AND GEOMETRY ASSESSMENT: IMPLICATION FOR PREOPERATIVE EVALUATION OF TRANSCATHETER TRICUSPID VALVE THERAPIES" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO COMPARE DIFFERENT IMAGING MODALITIES AS THREE-DIMENSIONAL (3D) TRANSTHORACIC ECHOCARDIOGRAPHY (TTE), TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TOE) AND ECG-GATED CARDIAC CT FOR THE DEFINITION OF TRICUSPID VALVE (TV) FUNCTION AND GEOMETRY ASSESSMENT.. DEVICES MENTIONED INCLUDE TRICLIP, TWO OTHER NON-ABBOTT DEVICES. THE ARTICLE CONCLUDED THAT THE PRESENT SERIES SHOWS HOW A LONE 3D TTE HAS A GOOD RELIABILITY IN THE DEFINITION OF TV FUNCTION, LEAFLET CHARACTERISTICS, AND GEOMETRY WHEN COMPARED WITH SECOND LEVEL IMAGING MODALITIES AND MAY BE SAFELY USED TO SELECT OPTIMAL CANDIDATES FOR COMPLEX TTVI. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NOEMI BRUNO AT SAN CAMILLO FORLANINI HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL: BRUNO@UNIROMA1.IT.]. THE TIME FRAME OF THE STUDY WAS UNKNOWN. A TOTAL OF 21 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH THREE RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, ATRIAL FIBRILLATION, PREVIOUS STROKE. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: B2 - DATE OF DEATH WAS ESTIMATED AS (B)(6) 2021 TO END OF STUDY RANGE. SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, ATRIAL FIBRILLATION, PREVIOUS STROKE.. COMPLICATIONS REPORTED INCLUDED DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2 - DATE OF DEATH WAS ESTIMATED AS (B)(6) 2021 TO END OF STUDY RANGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334205&pc=BZD | 2518422-2025-101848 | 2024-12-30 05:00 | Death | RESPIRONICS, INC. | 2025-02-07 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OF THE REP DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY AND THE PATIENT'S WIFE NEEDED A RETURN LABEL TO SEND BACK THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334366&pc=DSQ | 2916596-2025-00493 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-02-07 05:00 | DSQ | THORATEC HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS THE PATIENT NEVER RECOVERING AFTER IMPLANT. THE PATIENT WAS ON RIGHT SUPPORT AFTER IMPLANT HOWEVER IT WAS NOT POSSIBLE TO WEAN THEM AS THEIR CONDITION WAS TOO SICK. THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. IT WAS REPORTED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE RIGHT FAILURE DID NOT EXIST PRIOR TO IMPLANT, HOWEVER A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334383&pc=BZD | 2518422-2025-101847 | 2024-12-23 05:00 | Death | RESPIRONICS, INC. | 2025-02-07 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE DREAMSTATION AUTO CPAP DEVICE WAS NO LONGER AVAILABLE TO RETURN AND HAD BEEN DONATED. ACCORDING TO THE PATIENT'S DAUGHTER, HER FATHER PASSED AWAY LAST YEAR AND STOPPED USING THE DEVICE PRIOR TO PASSING. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. THE PATIENT'S DAUGHTER REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334384&pc=BZD | 2518422-2025-101850 | 2024-12-30 05:00 | Death | RESPIRONICS, INC. | 2025-02-07 05:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OF THE REMSTAR AUTO A-FLEX DEVICE HAD PASSED AWAY AND THE PATIENT'S WIFE NEEDED A RETURN LABEL TO SEND BACK THE DEVICE. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334603&pc=OJX | 2649622-2025-03745 | 2025-01-26 05:00 | Death | MPRI | 2025-02-07 05:00 | OJX | ATTAIN PERFORMA S MRI SURESCAN | Failure to Capture | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEFT VENTRICULAR (LV) LEAD EXHIBITED POSSIBLE INTERMITTENT LOSS OF CAPTURE. THE LV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED, VIA FOLLOW UP, THE PATIENT HAD PREVIOUSLY PRESENTED WITH VENTRICULAR TACHYCARDIA (VT), NON-ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION (NSTEMI), AND CONGESTIVE HEART FAILURE WITH MULTIPLE APPROPRIATE SHOCKS IN THE EMERGENCY ROOM. THE PATIENT UNDERWENT A PERCUTANEOUS CORONARY INTERVENTION/STENTING OF THE MID LEFT ANTERIOR DESCENDING (LAD) VEIN. THE PATIENT CONTINUED TO HAVE VT AND WAS TREATED WITH ANTI-ARRHYTHMIC MEDICATION AND TWO VT ABLATIONS, HOWEVER CONTINUED TO HAVE RECURRENT VT REQUIRING SHOCKS. A CODE WAS CALLED DUE TO VT AND ADVANCED CARDIAC LIFE SUPPORT WAS PROVIDED AND RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) BUT CONTINUED TO ARREST. ANOTHER CODE WAS CALLED, HOWEVER, THE PATIENT PASSED AWAY. IT WAS NOTED THAT LV LEAD CAPTURE WAS UNABLE TO BE MEASURED DUE TO COMPETING FAST RHYTHM AND WAS NOT A PRIMARY CONCERN FOR THE PHYSICIAN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21334675&pc=NPS | 2135147-2025-00615 | 2020-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-07 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "PREDICTIVE VALUE OF CT-BASED AND AI-RECONSTRUCTED 3D-TAPSE IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO ANALYZE WHETHER CT-CAPTURED TRICUSPID ANNULAR PLANE SYSTOLIC EXCURSION (TAPSE) PROVIDES ADDITIONAL VALUE IN PREDICTING OUTCOMES AFTER TRANSCATHETER TRICUSPID VALVE INTERVENTION (TTVI). DEVICES MENTIONED INCLUDE TRICLIP, PASCAL, AND CARDIOBAND. THE ARTICLE CONCLUDED THAT : POSTERIOR INDEXED ANTERIOR AND POSTERIOR (ITAPSE) IS AN INDEPENDENT PREDICTOR OF CARDIOVASCULAR OUTCOMES AMONG PATIENTS UNDERGOING TTVI. FURTHERMORE, CT-MEASURED TAPSE HAS INCREMENTAL VALUE AND REFINES RISK STRATIFICATION FOR CLINICAL OUTCOMES IN PATIENTS UNDERGOING TTVI. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TANJA KATHARINA RUDOLPH AT RUHR-UNIVERSITÄT BOCHUM, MEDIZINISCHE FAKULTÄT OWL (UNIVERSITÄT BIELEFELD), BAD OEYNHAUSEN, GERMANY WITH CORRESPONDING EMAIL *TRUDOLPH@HDZ-NRW.DE*. THE TIME FRAME OF THE STUDY WAS MAY 2020 AND JUNE 2023. A TOTAL OF 75 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH LESS THAN 20% RECEIVED AN ABBOTT DEVICE; EXACT AMOUNT IS UNKNOWN. THE MEAN PATIENT AGE WAS 77 ± 8 YEARS, AND 61% WERE FEMALE. COMORBIDITIES INCLUDED RENAL IMPAIRMENT, PACEMAKER, CARDIAC DECOMPENSATION, CORONARY ARTERY DISEASE, DIABETES MELLITUS, HISTORY OF STROKE, COPD, ATRIAL FIBRILLATION, HISTORY OF CARDIAC SURGERY, HEART FAILURE, MINERALOCORTICOID RECEPTOR ANTAGONIST, ACE INHIBITOR, BETABLOCKERS, DOAC, VITAMIN-K ANTAGONIST, DOSAGE OF LOOP FUROSEMIDE OR TORASEMIDE, TRICUSPID REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE "PREDICTIVE VALUE OF CT-BASED AND AI-RECONSTRUCTED 3D-TAPSE IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR." THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING RENAL IMPAIRMENT, PACEMAKER, CARDIAC DECOMPENSATION, CORONARY ARTERY DISEASE, DIABETES MELLITUS, HISTORY OF STROKE, COPD, ATRIAL FIBRILLATION, HISTORY OF CARDIAC SURGERY, HEART FAILURE, MINERALOCORTICOID RECEPTOR ANTAGONIST, ACE INHIBITOR, BETABLOCKERS, DOAC, VITAMIN-K ANTAGONIST, DOSAGE OF LOOP FUROSEMIDE OR TORASEMIDE, TRICUSPID REGURGITATION. COMPLICATIONS REPORTED INCLUDED HOSPITALIZATION, HEART FAILURE, DEATH, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, RIGHT CORONARY ARTERY PERFORATION, STENT IMPLANTATION, PERIPROCEDURAL DIALYSIS, TRANSIENT ATRIOVENTRICULAR BLOCK, BLEEDING, BLOOD TRANSFUSION.; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21335094&pc=NPT | 2135147-2025-00616 | 2025-01-14 05:00 | Death | ABBOTT MEDICAL | 2025-02-07 05:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device; Improper or Incorrect Procedure or Method | Low Blood Pressure/ Hypotension; Vascular Dissection; Pericardial Effusion | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 35MM NAVITOR TITAN VISION WAS CHOSEN FOR PRE-PLANNED IMPLANTATION UTILIZING A LARGE FLEXNAV DELIVERY SYSTEM. PATIENT WAS CONSIDERED HIGH RISK WITH MULTIPLE COMORBIDITIES. SUFFICIENT CALCIUM WAS PRESENT FOR ANCHORING. AORTIC VALVE ANGLE WAS 38 DEGREES. PRE-DILATATION WAS PERFORMED WITH 26MM TRUE BALLOON (25.5MM MINIUMN ANNULUS DIAMETER). THE NAVITOR WAS AT A DEPTH OF 5MM AT 80% DEPLOYMENT. AT RELEASE, THE VALVE WAS AT A DEPTH OF 5MM ON NON-CORONARY CUSP (NCC). POST DEPLOYMENT, THE NAVITOR MIGRATED TOWARDS THE AORTA. THERE WAS NO ENTANGLEMENT WITH DELIVERY SYSTEM AND NO BLOCKAGE OF THE CORONARIES. THE IMPLANTER THOUGHT THE MIGRATION WAS RELATED TO BACKWARDS TENSION ON THE DEVICE. THE NAVITOR WAS SNARED UPWARDS AND A REPLACEMENT 29MM NON-ABBOTT VALVE (EDWARDS SAPIEN) WAS IMPLANTED. POST IMPLANT OF THE NON-ABBOTT VALVE, A LARGE CIRCUMFERENTIAL PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE WAS NOTED. OPEN HEART SURGERY TO REPAIR ANNULAR RUPTURE AT AORTIC ROOT WAS PERFORMED. THE PATIENT DIED DURING THE SURGICAL INTERVENTION ON 14 JANUARY 2025. CAUSE OF DEATH WAS REPORTED TO BE THE ANNULAR RUPTURE. IT WAS THOUGHT THE ANNULAR RUPTURE OCCURRED DUE TO THE INFLATION OF THE NON-ABBOTT VALVE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF MIGRATION, IMPROPER OR INCORRECT PROCEDURE OR METHOD, VASCULAR DISSECTION, PERICARDIAL EFFUSION, HYPOTENSION, AND DEATH WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE CASE WAS REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. BASED ON ALL AVAILABLE INFORMATION, A CAUSE FOR THE REPORTED MIGRATION COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED VASCULAR DISSECTION, PERICARDIAL EFFUSION, HYPOTENSION, AND DEATH APPEAR TO BE RELATED TO PROCEDURAL CONDITIONS (DEPLOYMENT OF NON-ABBOTT BALLOON-EXPANDABLE DEVICE). THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGERY WERE THE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES. THE REPORTED IMPROPER OR INCORRECT PROCEDURE OR METHOD IS RELATED TO THE USER SNARING THE DEVICE AFTER DEPLOYMENT. THIS DEVIATION FROM THE IFU DID NOT CAUSE OR CONTRIBUTE TO ANY ISSUES AND WAS PERFORMED FOR TREATMENT OF THE MIGRATION. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21335268&pc=FNM | 0001831750-2025-00331 | 2025-01-13 05:00 | Death | STRYKER MEDICAL-KALAMAZOO | 2025-02-07 05:00 | FNM | ISOAIR DOM 84 IN TYPE B PLUG | Adverse Event Without Identified Device or Use Problem | Pressure Sores |  |  | 1 | Event Description: IT WAS REPORTED THROUGH A CLINICAL SURVEY THAT A 54-YEAR-OLD, OVERWEIGHT, MALE PATIENT WAS EMERGENTLY ADMITTED TO THE INTENSIVE CARE UNIT WITH ¿CARDIAC RESUSCITATION WITH RSOC (RETURN OF SPONTANEOUS CIRCULATION).¿ THE PATIENT PRESENTED WITH A BRADEN SCORE OF 12 AND A STAGE 2 PRESSURE INJURY INVOLVING THE SACRAL AREA. DURING THE TREATMENT PERIOD, THE PRESSURE INJURY DID NOT IMPROVE AND PROGRESSED TO A STAGE 3 PRESSURE INJURY. ADDITIONALLY, IT WAS REPORTED THAT TWO NEW STAGE 2 PRESSURE INJURIES DEVELOPED ON THE HEELS, CONTRIBUTING TO A DETERIORATION IN THE PATIENT¿S HEALTH, RESULTING IN PROLONGED HOSPITALIZATION AND ULTIMATELY, DEATH. THE SURVEY RESPONDENT NOTED THAT THERE WAS NO DEVICE MALFUNCTION, AND THE PRODUCT WAS BEING USED IN ACCORDANCE WITH THE INDICATIONS FOR USE. SEVERAL COMORBID CONDITIONS AND MEDICATIONS ARE NOTED, WHICH MAY IMPEDE WOUND HEALING. GIVEN THE PREVIOUS INFORMATION AND THE MULTI-FACTORIAL NATURE OF PRESSURE INJURY DEVELOPMENT IT CANNOT BE CONCLUSIVELY DETERMINED THAT THE PRODUCT HAD A CAUSAL RELATIONSHIP OR CONTRIBUTION TO THE PATIENT¿S OUTCOME. NONETHELESS, WE ARE CHOOSING TO FILE THIS EVENT OUT OF AN ABUNDANCE OF CAUTION. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL ATTEMPTS ARE BEING MADE TO OBTAIN MORE INFORMATION.Event Description: NO NEW INFORMATION. Manufacturer Narrative: MULTIPLE ATTEMPTS HAVE BEEN MADE TO GATHER FURTHER INFORMATION, WITH NO RESPONSE RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21335298&pc=BTI | 2518422-2025-013638 | 2023-11-21 05:00 | Death | RESPIRONICS, INC. | 2025-02-07 05:00 | BTI | ESSENCE RENTAL WITH SSD 230V50HZ UK | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K042655 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE THAT A USER OF THE ESSENCE RENTAL HAD STATES PATIENT HAS PASSED AWAY AND LOST THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21335637&pc=DSQ | 2916596-2025-00553 | 2025-01-20 05:00 | Death | THORATEC CORPORATION | 2025-02-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Loss of Range of Motion; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD ISCHEMIC STROKE ONE DAY POST IMPLANT ON (B)(6) 2025 WITH SYMPTOM OF RIGHT HEMIPLEGIA. A COMPUTED TOMOGRAPHY (CT) WAS PERFORMED TO CONFIRM THE STROKE WHICH ESTABLISHED INFRACT INVOLVING THE ENTIRE MIDDLE CEREBRAL ARTERY (MCA) TERRITORY. THERE WAS ASSOCIATED EFFACEMENT OF THE SULCI AND LEFT LATERAL VENTRICLE, RIGHTWARD MIDLINE SHIFTS OF 1CM, MILD UNCAL HERNIATION, AND MINIMAL SUBFALCINE LINE HERNIATION. NO HYDROCEPHALUS WAS NOTED. THE PATIENT WAS STARTED ON COUMADIN (WARFARIN) BUT WAS SUBTHERAPEUTIC. THE FAMILY REQUESTED TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THE DEATH WAS DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21335816&pc=MKJ | 1220908-2025-00413 | 2025-01-15 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-07 05:00 | MKJ | X SERIES | Unable to Obtain Readings | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 71-YEAR-OLD FEMALE PATIENT, THE FIRST EMS CREW ATTACHED THEIR PADS TO THIS DEVICE AND THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. COMPLAINANT INDICATED THE CLINICIAN OBTAINED A NEW SET OF PADS TO CONTINUE TREATING THE PATIENT BUT THE DEVICE WAS STILL UNABLE TO OBTAIN AN ECG SIGNAL VIA PADS. THE CREW UTILIZED A 4-LEAD PLACEMENT TO MONITOR THE PATIENT'S HEART RHYTHM FOR THE REMAINDER OF THE CALL. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. A REVIEW OF DEVICE LOGS FOUND THE DEVICE PROMPTED DEFIB CABLE ID "MULTIFUNCTION-PADS" MESSAGE. NO ECG SIGNAL WAS DETECTED FOR THREE MINUTES DESPITE MULTIPLE TROUBLESHOOTING ATTEMPTS, INCLUDING POWER CYCLING AND RECONNECTING THE MULTIFUNCTION CABLE (MFC). THE OPERATOR WAS ABLE TO OBTAIN AN ECG SIGNAL IN LEAD II VIEW BUT ENCOUNTERED A PADS LEAD FAULT WITH NO SIGNAL WHEN SWITCHING BACK TO PADS VIEW. THE DEVICE PASSED ALL FUNCTIONAL TESTING WITH TEST PADS AND MFC CABLE. THE CUSTOMER'S ACCESSORIES WERE NOT RETURNED FOR EVALUATION. THE CUSTOMER IS ADVISED TO ASSESS THEIR ACCESSORIES FOR COMPATIBILITY AND FUNCTIONALITY. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BEFORE THE UDI REQUIREMENTS. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21335822&pc=DSQ | 2916596-2025-00947 | 2024-06-01 04:00 | Death | THORATEC CORPORATION | 2025-02-07 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding; Sepsis; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿NINE YEARS OF CONTINUOUS FLOW LVAD (HEARTMATE 3): SURVIVAL AND LVAD-RELATED COMPLICATIONS BEFORE AND AFTER HOSPITAL DISCHARGE¿ THAT HEARTMATE 3 (HM3) IMPLANT IS ASSOCIATED WITH BLEEDING, INFECTION, DRIVELINE INFECTION, HEMOLYSIS, NEUROLOGICAL DYSFUNCTION, STROKE, RIGHT VENTRICULAR (RV) FAILURE, MULTISYSTEM ORGAN FAILURE (MSOF), GASTROINTESTINAL (GI) BLEEDING, ACUTE KIDNEY INJURY (AKI), THE NEED FOR ADDITIONAL SURGERY, THE NEED FOR TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) SUPPORT, PROLONGED HOSPITALIZATION, AND DEATH. TO EVALUATE SURVIVAL RATES, LEFT VENTRICULAR ASSIST DEVICE (LVAD) RELATED COMPLICATIONS OF THE HM3 LVAD, AND CAUSES OF HOSPITAL RE-ADMISSIONS AFTER HOSPITAL DISCHARGE, A RETROSPECTIVE REVIEW OF ALL PATIENTS IMPLANTED WITH A HM3 BETWEEN SEPTEMBER 2015 AND JUNE 2024 WAS CONDUCTED, EXCLUDING ALL PATIENT WHO RECEIVED A HEART WARE VENTRICULAR ASSIST DEVICE (HVAD) OR HEARTMATE II (HMII) LVAD. OVERALL, 48 PATIENTS (AGE: 56.1 ± 10.6 YEARS, 40 MALE (83.3%)) WERE ENROLLED IN THE STUDY. THE MOST COMMON INDICATION FOR LVAD IMPLANT WAS ISCHEMIC CARDIOMYOPATHY IN 25 PATIENTS (52.1%), FOLLOWED BY DILATATIVE CARDIOMYOPATHY IN 19 PATIENTS (39.6%) AND THEN MYOCARDITIS IN 4 PATIENTS (8.3%). 35 PATIENTS UNDERWENT LVAD IMPLANTATION AS A BRIDGE TO TRANSPLANTATION (72.9%), 5 AS BRIDGE TO DECISION (10.4%), 1 AS BRIDGE TO RECOVERY (2.1%), AND 7 AS DESTINATION THERAPY (14.6%). AT THE TIME OF LVAD IMPLANTATION, 16 PATIENTS EACH WERE AN INTERAGENCY REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT PROFILE 1 AND 4 (33.3%), FOLLOWED BY 9 PROFILE 3 PATIENTS (18.8%) AND 7 PROFILE 2 PATIENTS (14.6%). REGARDING PRE-IMPLANTATION RV FUNCTION, 18 PATIENTS (37.5%) HAD NORMAL FUNCTION, 15 (31.2%) HAD MILDLY IMPAIRED FUNCTION, 12 (25%) MODERATELY IMPAIRED, AND 3 (6.2%) SEVERELY IMPAIRED, WITH A MEAN TRICUSPID ANNULAR PLANE SYSTOLIC EXCURSION (TAPSE) VALUE OF 16.4 ± 4.7 MM. THE AVERAGE LENGTH OF STAY (LOS) WAS 27.21 ± 22.7 DAYS FOR THE INTENSIVE CARE UNIT (ICU) AND 50 ± 36.1 DAYS FOR THE HOSPITAL OVERALL. PERIOPERATIVELY, 28 (58.3%) PATIENTS DEVELOPED AN AKI, 23 (47.9%) PATIENTS DEVELOPED AN INFECTION (MAINLY PNEUMONIA), 5 (10.4%) PATIENTS SUFFERED A STROKE, AND 5 (10.4%) PATIENTS SUFFERED FROM GI BLEEDING. POSTOPERATIVE BLEEDING LED TO 19 (39.6%) PATIENTS REQUIRING A RE-THORACOTOMY. RVAD SUPPORT DUE TO DEVELOPED OR WORSENED RV FAILURE WAS REQUIRED BY 6 (12.5%) PATIENTS. PATIENTS WHO SUFFERED AN AKI HAD A STRONG TREND TO LONGER IN-HOSPITAL STAYS COMPARED TO THOSE WITHOUT AN AKI (P = 0.07). THERE WERE 7 (14.6%) IN-HOSPITAL DEATHS, 4 (8.3%) FROM RV FAILURE, AND 1 (2.1%) EACH FROM CEREBROVASCULAR STROKE, ABDOMINAL BLEEDING, AND MSOF. OF THE 41 PATIENTS WHO WERE DISCHARGED, THERE WERE 11 (26.8%) DEATHS, WITH 4 (9.8%) DUE TO SEPSIS, 2 (4.9%) DUE TO RV FAILURE, 2 (4.9%) DUE TO NON-LVAD CAUSES (1 LIVER CARCINOMA AND 1 PANCREATIC CARCINOMA), 2 (4.9%) DUE TO UNKNOWN CAUSES, AND 1 (2.1%) DUE TO CEREBROVASCULAR STROKE. THE OVERALL COHORT HAD A MORTALITY RATE OF 18 (37.5%). OF THE 41 DISCHARGES, HOSPITAL READMISSIONS WERE MAINLY CAUSED BY DRIVELINE INFECTION IN 4 (9.8%) PATIENTS, FOLLOWED BY BLEEDING COMPLICATIONS IN 3 (7.3%) PATIENTS. OF THE PATIENTS WHO DEVELOPED A DRIVELINE INFECTION, 2 (4.9%) UNDERWENT A SURGICAL RELOCATION; THE OTHER PATIENTS WERE TREATED CONSERVATIVELY. 1 (2.4%) PATIENT EACH SUFFERED SPONTANEOUS THORACIC BLEEDING, GI BLEEDING, AND POSTOPERATIVE BLEEDING AFTER SURGICAL DEBRIDEMENT OF A DRIVELINE INFECTION. NONE OF THESE COMPLICATIONS WERE LETHAL. DURING FOLLOW-UP, THERE WERE NO THROMBOSIS EVENTS REPORTED, AND NO NEED FOR A LVAD REPLACEMENT. 1 (2.4%) PATIENT UNDERWENT A HEART TRANSPLANTATION, AND 26 (63.4%) WERE STILL ON LVAD SUPPORT. THE MEAN FOLLOW-UP TIME WAS 1,263 ± 852 DAYS. THE RECORDED 1-, 2-, AND 5-YEAR SURVIVAL RATES WERE 83%, 785, AND 65%, RESPECTIVELY. THE CONTINUOUS FLOW LVAD (HM3) WAS ASSOCIATED WITH AN EXCELLENT IN-HOSPITAL SURVIVAL RATE IN PATIENTS WITH END-STAGE HEART FAILURE, WITH AN 85.4% SUCCESSFUL DISCHARGE RATE. Manufacturer Narrative: E1 - REPORTER NAME: CORRECTION. SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DEPARTMENT OF THORACIC AND CARDIOVASCULAR SURGERY, TÜBINGEN UNIVERSITY HOSPITAL, 72076 TÜBINGEN, GERMANY MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE ORGAN FAILURES/DYSFUNCTIONS (INCLUDING RIGHT HEART FAILURE), SEPSIS, STROKE, BLEEDING, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. ADDITIONALLY, THIS SECTION STATES THAT SUDDEN RIGHT VENTRICULAR FAILURE MAY DEVELOP DURING OR SHORTLY AFTER PUMP IMPLANTATION AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: BOBURG, R. S., MARINOS, S. L., BAUMGAERTNER, M., RUSTENBACH, C. J., SALEWSKI, C., DOLL, I., BERGER, R., SCHLENSAK, C., & RADWAN, M. (2024). NINE YEARS OF CONTINUOUS FLOW LVAD (HEARTMATE 3): SURVIVAL AND LVAD-RELATED COMPLICATIONS BEFORE AND AFTER HOSPITAL DISCHARGE. JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE, 11(10), 301. HTTPS://DOI.ORG/10.3390/JCDD11100301 . NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21336814&pc=JWH | 1038671-2025-00754 | 2025-01-29 05:00 | Death | EXACTECH, INC. | 2025-02-07 05:00 | JWH | UNKNOWN KNEE COMPONENTS | Device Dislodged or Dislocated | Joint Dislocation |  |  | 1 | Event Description: AS REPORTED, THE PARTY STATED HER LATE FAMILY MEMBER'S KNEE REPLACEMENT CAME OUT THE SOCKET THREE TIMES, WHICH LED TO ITS AMPUTATION AND THEIR EVENTUAL DEATH. INFORMATION PROVIDED INDICATES THE PATIENT WAS REVISED. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: THE REPORTED EVENT WAS UNABLE TO BE CONFIRMED DUE TO LIMITED INFORMATION RECEIVED FROM THE CUSTOMER. NO DEVICE WAS RETURNED FOR EVALUATION; FURTHER, PHOTOGRAPHS AND/OR RADIOGRAPH IMAGES WERE NOT PROVIDED FOR REVIEW. OPERATIVE NOTES AND/OR MEDICAL RECORDS WERE NOT PROVIDED FOR REVIEW OF USAGE/ TECHNIQUE. A REVIEW OF MANUFACTURING DATA WAS UNABLE TO BE PERFORMED AS THE LOT INFORMATION OF THE PRODUCT INVOLVED IN THE EVENT WAS NOT AVAILABLE. A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AS THE NECESSARY INFORMATION TO ADEQUATELY INVESTIGATE THE REPORTED EVENT WAS NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE CORRECTED: B3: DATE OF EVENT UNKNOWN. E1 H6: COMPONENT CODE, TYPE OF INVESTIGATION. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21336900&pc=EZL | 9610711-2025-00027 | 2023-10-01 04:00 | Death | COLOPLAST A/S | 2025-02-07 05:00 | EZL | FOLYSIL SILICONE CATHETER | Insufficient Information | Drug Resistant Bacterial Infection |  |  | 1 | Event Description: ACCORDING TO THE AVAILABLE INFORMATION, THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM IN (B)(6) 2023 AND WAS SUBSEQUENTLY DISCHARGED TO A HOME HEALTH FACILITY. A FOLEY CATHETER WAS UTILIZED EITHER IN THE EMERGENCY ROOM OR THE HOME HEALTH FACILITY. AFTER RECEIVING THE FOLEY CATHETER, THE PATIENT REPORTEDLY DEVELOPED MRSA INFECTION AND PASSED AWAY FOUR WEEKS LATER. NO KNOWN FACTS ESTABLISH WHICH COMPANY MANUFACTURED THE FOLEY CATHETER. Manufacturer Narrative: THE INVESTIGATION IS NOT YET COMPLETE, A FOLLOW UP REPORT WILL BE SUBMITTED ON COMPLETION OF THE INVESTIGATION. B2: DATE IS UNKNOWN B3: ESTIMATED DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337009&pc=NKM | 2135147-2025-00632 | 2025-01-09 05:00 | Death | ABBOTT MEDICAL | 2025-02-07 05:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Mitral Valve Stenosis; Ventricular Fibrillation; Heart Failure/Congestive Heart Failure | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+, EJECTION FRACTION (EF) OF 25%, AND A MEAN PRESSURE GRADIENT OF 1MMHG. TWO MITRACLIP NTW CLIPS WERE DEPLOYED ON THE MITRAL VALVE. A THIRD MITRACLIP NT WAS THEN INSERTED AND DEPLOYED ON THE MITRAL VALVE, REDUCING MR TO A GRADE OF 1 AND THE EF WAS REDUCED TO 20%. HOWEVER, AFTER DEPLOYMENT, THE MEAN PRESSURE GRADIENT INCREASED TO 7MMHG. ON (B)(6) 2025, POST PROCEDURE, THE PATIENT WAS REPORTED TO BE STABLE. ON (B)(6) 2025, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF) ARREST AND DIED.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+, EJECTION FRACTION (EF) OF 25%, AND A MEAN PRESSURE GRADIENT OF 1MMHG. TWO MITRACLIP NTW CLIPS WERE DEPLOYED ON THE MITRAL VALVE. A THIRD MITRACLIP NT WAS THEN INSERTED AND DEPLOYED ON THE MITRAL VALVE, REDUCING MR TO A GRADE OF 1 AND THE EF WAS REDUCED TO 20%. HOWEVER, AFTER DEPLOYMENT, THE MEAN PRESSURE GRADIENT INCREASED TO 7MMHG. ON (B)(6) 2025, POST PROCEDURE, THE PATIENT WAS REPORTED TO BE STABLE. ON (B)(6) 2025, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF) ARREST AND DIED. Manufacturer Narrative: THE CLIP REMAINS IN THE PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED MITRAL STENOSIS, HEART FAILURE, VENTRICULAR FIBRILLATION, AND CARDIAC ARREST. THE REPORTED DEATH APPEARS TO BE RELATED TO THE HEART FAILURE. MITRAL STENOSIS, HEART FAILURE, VENTRICULAR FIBRILLATION, CARDIAC ARREST, HEART FAILURE, AND DEATH ARE LISTED IN THE INSTRUCTIONS FOR USE IS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED HOSPITALIZATION WAS A RESULT OF CASE SPECIFIC CIRCUMSTANCE AS THE PATIENT REMAINED HOSPITALIZED DUE TO THE PATIENT EFFECTS. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337053&pc=NKM | 2135147-2025-00631 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-07 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "EFFECT OF RENAL IMPAIRMENT ON CLINICAL OUTCOMES AFTER MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI-CENTER STUDY, TO EVALUATE THE EFFECT OF RENAL IMPAIRMENT ON CLINICAL OUTCOMES AFTER MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER). DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT RENAL IMPAIRMENT WAS ASSOCIATED WITH AN INCREASED INCIDENCE OF MAJOR ADVERSE EVENTS, AND DIALYSIS WAS THE STRONGEST INDEPENDENT PREDICTOR OF POOR CLINICAL OUTCOMES AFTER M-TEER IN ASIAN-PACIFIC PATIENTS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS (B)(6). THE TIME FRAME OF THE STUDY WAS (B)(6) 2018 AND (B)(6) 2021. A TOTAL OF (B)(4) PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION, HEART FAILURE, HYPERTENSION, DIABETES, CURRENT SMOKER, ATRIAL FIBRILLATION, CARDIAC RESYNCHRONIZATION THERAPY, CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, PRIOR PCI/CABG, PERIPHERAL ARTERY DISEASE, STROKE, COPD, RENAL FAILURE, DIALYSIS. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED CARDIOVASCULAR DEATH, NON-CARDIOVASCULAR DEATH, HEART FAILURE HOSPITALIZATION, STROKE, INFECTION, BLEEDING. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DEGENERATIVE MITRAL REGURGITATION, FUNCTIONAL MITRAL REGURGITATION, HEART FAILURE, HYPERTENSION, DIABETES, CURRENT SMOKER, ATRIAL FIBRILLATION, CARDIAC RESYNCHRONIZATION THERAPY, CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, PRIOR PCI/CABG, PERIPHERAL ARTERY DISEASE, STROKE, COPD, RENAL FAILURE, DIALYSIS. COMPLICATIONS REPORTED INCLUDED CARDIOVASCULAR DEATH, NON-CARDIOVASCULAR DEATH, HEART FAILURE HOSPITALIZATION, STROKE, INFECTION, BLEEDING; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE WAS ESTIMATED TO THE END OF STUDY RANGE. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLED " EFFECT OF RENAL IMPAIRMENT ON CLINICAL OUTCOMES AFTER MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR ".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337247&pc=MKJ | 1220908-2025-00552 | 2025-01-16 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-07 05:00 | MKJ | R SERIES | Unable to Obtain Readings; Use of Device Problem | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 14-YEAR-OLD MALE PATIENT, THE DEVICE HAD "UNUSUAL ACTIVITIES". THIS IS ALL THE INFORMATION AVAILABLE AT THIS TIME. COMPLAINANT INDICATED THAT THE CLINICIAN CONTINUED TO USE THIS DEVICE WITH A NEW SET OF PADS TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED; HOWEVER, THEY DO NOT BELIEVE THIS WAS DEVICE ASSOCIATED.Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 14-YEAR-OLD MALE PATIENT, THE DEVICE DISPLAYED AN "ACCESSORY ERROR 6" MESSAGE. THIS ERROR IS RELATED TO THE SOFTWARE BEING UNABLE TO COMMUNICATE WITH THE CPR PUCK AND IS A NON-CRITICAL ERROR THAT DOES NOT IMPACT THE DEVICE'S ABILITY TO DELIVER DEFIBRILLATION THERAPY. COMPLAINANT INDICATED THAT THE CLINICIAN CONTINUED TO USE THIS DEVICE WITH A NEW SET OF PADS TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED; HOWEVER, THEY DO NOT BELIEVE THIS WAS DEVICE ASSOCIATED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THIS SUPPLEMENTAL MEDWATCH REPORT IS REPORTING THE EVALUATION OF THE DEVICE AND CORRECTING INFORMATION SUBMITTED ON THE INITIAL MEDWATCH REPORT. PLEASE REFERENCE SECTION B5 AND H6 (DEVICE PROBLEM CODE). THIS REPORT WAS INADVERTENTLY SUBMITTED. REPORTS OF THIS NATURE HAVE NO POTENTIAL FOR CLINICAL IMPACT. THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE LOGS. HOWEVER, THE DEVICE WAS PUT THROUGH EXTENSIVE FUNCTIONAL TESTING INCLUDING STRESS TESTING USING THE CUSTOMER'S RETURNED ACCESSORIES WITHOUT DUPLICATING THE REPORT. THE DEFIB PADS USED DURING THE EVENT WERE NOT RETURNED FOR EVALUATION. A RETAIN SAMPLE EVALUATION WAS PERFORMED ON THE ONESTEP CPR A/A LOT 4324A AND THE PADS WERE DETERMINED TO BE ASSEMBLED TO SPECIFICATION. THE ANALOG BOARD WAS REPLACED AS A PRECAUTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337670&pc=LWS | 3004209178-2025-02339 | 2024-12-09 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-07 05:00 | LWS | VISIA AF MRI S VR SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY "A FEW DAYS" AFTER A SURGERY OCCURRED. IT WAS NOTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS "TURNED ON". ADDITIONAL INFORMATION AND A CAUSE OF DEATH WAS REQUESTED, BUT NOT RECEIVED.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY "A FEW DAYS" AFTER A SURGERY OCCURRED. IT WAS NOTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS "TURNED ON". ADDITIONAL INFORMATION AND A CAUSE OF DEATH WAS REQUESTED, BUT NOT RECEIVED. Manufacturer Narrative: CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337688&pc=DSQ | 2916596-2025-00696 | 2024-10-19 04:00 | Death | THORATEC CORPORATION | 2025-02-07 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Bacterial Infection; Purulent Discharge; Hemorrhage/Bleeding; Septic Shock; Fungal Infection; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿FIRST EVIDENCE OF A HEARTMATE 3 DRIVELINE INFECTION BY RHIZOPUS ARRHIZUS: A CASE REPORT¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH INFECTION AND DEATH. THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY AND HAD AN IMPLANTED CARDIAC DEFIBRILLATOR. THE PATIENT HAD MULTIPLE IMMUNOSUPPRESSIVE RISK FACTORS INCLUDING TYPE 2 DIABETES, CHRONIC KIDNEY DISEASE, AND ADVANCED HEART FAILURE. THE PATIENT DEVELOPED UROSEPTIC SHOCK BEFORE THE HM3 IMPLANTATION AND WAS THUS TREATED WITH MEROPENEM FOR A PERIOD OF 10 DAYS. AFTER THE INFECTION WAS TREATED THE HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS IMPLANTED. DUE TO INTRAOPERATIVE RIGHT VENTRICULAR (RV) FAILURE, A RV CENTRIMAG WAS IMPLANTED. ANTIBIOTIC PROPHYLAXIS WITH DAPTOMYCIN AND CEFTAZIDIME-AVIBACTAM WAS THEN PRESCRIBED DUE FOR A OXACILLINASE-48 (OXA-48) BACTERIAL INFECTION. FOR 2 DAYS AFTER SURGERY BLEEDING THROUGH CHEST TUBES PERSISTED. 2 ADDITIONAL SURGICAL INTERVENTIONS WERE REQUIRED TO CONTAIN BLEEDING. LAB TESTS REVEALED PROGRESSIVE INCREASE IN LEUKOCYTE COUNT AND C-REACTIVE PROTEIN (203.0MG/DL), IN ADDITION TO ELEVATED PROCALCITONIN LEVELS(1.08NG/DL). TO ADDRESS THE INFECTION, ANIDULAFUNGIN, DAPTOMYCIN AND CEFTAZIDIME-AVIBACTAM WERE PRESCRIBED. 10 DAYS POST IMPLANT ANTIMICROBIAL DRUGS WERE STOPPED BECAUSE OF CLINICAL IMPROVEMENT. 23 DAYS AFTER HM3 IMPLANTATION PURULENT DRAINAGE WAS OBSERVED IN STERNOTOMY WOUND. THIS WAS TREATED WITH TEICOPLANIN, CEFTAZIDIME-AVIBACTAM AND ANIDULAFUNGIN. LATER THE SAME PURULENT DRAINAGE WAS DETECTED AT THE HM3 DRIVELINE INSERTION SITE AND A SKIN ULCER APPEARED ABOVE THE DRIVELINE PATHWAY. CULTURES FROM THESE SITES IDENTIFIED R ARRHIZUS. THIS LED TO ANIDULAFUNGIN TO BE REPLACED WITH INTRAVENOUS AMPHOTERICIN B AND ORAL POSACONAZOLE. A SOFT TISSUE ULTRASOUND SHOWED PHLEGMONOUS CHANGES AROUND THE DRIVELINE WITHOUT A CLEAR ABSCESS. R ARRHIZUS WAS NOT FOUND IN BLOOD, URINE, OR BRONCHIAL ASPIRATE CULTURES. A TRANSESOPHAGEAL ECHOCARDIOGRAM WAS PERFORMED REVEALING VEGETATION CONSISTENT WITH IMPLANTABLE CARDIAC DEFIBRILLATOR-RELATED ENDOCARDITIS. PATIENT PASSED AWAY DUE TO SEPTIC SHOCK 33 DAYS AFTER HM3 IMPLANTATION. THE ARTICLE CONCLUDED THAT EARLY RECOGNITION AND TREATMENT OF FUNGAL INFECTIONS IN LVAD PATIENTS IS ESSENTIAL AND PROMPT INVESTIGATION OF NONSPECIFIC SYMPTOMS THAT COULD POINT TO INFECTION IS ESSENTIAL. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN; DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF PUBLICATION 19OCT2024 SINCE THE DATE OF DATA COLLECTION WAS NOT PROVIDED. AUTHOR INFORMATION: VILA-GARCÍA, J., MARCO CLEMENT, I., POVEDA PINEDO, I. D., LOECHES YAGÜE, B., CID MENÉNDEZ, A., & MORENO, R. (2024). FIRST EVIDENCE OF A HEARTMATE 3 DRIVELINE INFECTION BY RHIZOPUS ARRHIZUS: A CASE REPORT. THE CANADIAN JOURNAL OF CARDIOLOGY, 40(12), 2637¿2639. HTTPS://DOI.ORG/10.1016/J.CJCA.2024.10.013. LA PAZ UNIVERSITY HOSPITAL, HOSPITAL LA PAZ INSTITUTE FOR HEALTH RESEARCH- IDIPAZ, MADRID, SPAIN. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). H6: ADDITIONAL HEALTH EFFECT- CLINICAL CODE: ENDOCARDITUS-E0610, ULCER-E2339, HEMODYNAMIC INSTABILITY-E2343. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTIONS E2 AND E3: CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL INFORMATION. THE INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MFR#2916596-2023-08404.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337731&pc=EOQ | 3014447948-2025-00002 | 2025-01-13 05:00 | Death | AURIS HEALTH, INC. | 2025-02-07 05:00 | EOQ | MONARCH ENDOSCOPY PLATFORM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K211493 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT CODED DURING PROCEDURE, THE PHYSICIAN HAD CONDUCTED 4-BIOPSIES AND ON THE 5TH BIOPSY AN ABNORMAL AMOUNT OF BLOOD WAS COMING THROUGH THE ET TUBE. THE SCOPE AND THE PATIENT INTRODUCER WERE REMOVED AND THEY WERE ABLE TO STABILIZE THE PATIENT IN THE ROOM. THE PATIENT WAS MOVED TO ICU WHERE THEY LATER CODED AND PASSED AWAY. THE PHYSICIAN STATED THAT THE MONARCH DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21337741&pc=NPT | 2025587-2025-01097 | 2025-02-03 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-07 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Device Damaged by Another Device; Device Dislodged or Dislocated; Patient Device Interaction Problem | Cardiac Arrest; Cardiogenic Shock; Foreign Body In Patient; Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUT R 29 MM, THE VALVE DISLODGED AFTER DEPLOYMENT. SUBSEQUENTLY, A SECOND VALVE WAS IMPLANTED. THE PATIENT EXPERIENCED A SHOCK AND SUBSEQUENTLY DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE IMPLANT, NO PRE-IMPLANT BALLOON DILATATION WAS PERFORMED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC DISSECTION. THE PATIENT'S SHOCK EXPERIENCED WAS CARDIOGENIC SHOCK DUE TO CARDIAC ARREST.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUT R 29 MM, THE VALVE DISLODGED AFTER DEPLOYMENT. SUBSEQUENTLY, A SECOND VALVE WAS IMPLANTED. THE PATIENT EXPERIENCED A SHOCK AND SUBSEQUENTLY DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE IMPLANT, NO PRE-IMPLANT BALLOON DILATATION WAS PERFORMED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC DISSECTION. THE PATIENT'S SHOCK EXPERIENCED WAS CARDIOGENIC SHOCK DUE TO CARDIAC ARREST. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PATIENT'S NATIVE AORTIC VALVE WAS HEAVILY CALCIFIED WHICH ADDED STRESS TO THE FIRST EVOLUT VALVE AND RESULTED IN DISLODGEMENT TO THE ASCENDING AORTA. DURING ADVANCEMENT OF THE DELIVERY CATHETER SYSTEM (DCS) WITH THE LOADED SECOND VALVE, THE FIRST VALVED MOVED MULTIPLE TIMES AND RESULTED IN AN ASCENDING AORTIC DISSECTION WITH SUBSEQUENT CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS INITIATED AND PERFORMED FOR MORE THAN ONE HOUR. AFTERWARD, THE PATIENT WAS PREPARED FOR OPEN HEART SURGERY TO REPAIR THE AORTIC DISSECTION. THE PATIENT WAS PLACED ON CARDIOPULMONARY BYPASS; HOWEVER, THE PATIENT'S HEART DID NOT RESPOND TO THE BYPASS PROCEDURE. SUBSEQUENTLY, THE PATIENT'S DEATH WAS PRONOUNCED AND A TIME OF DEATH WAS DECLARED. PER THE PHYSICIAN, THE MEDTRONIC VALVES AND THE SECOND MEDTRONIC DCS CANNOT BE EXCLUDED AS CONTRIBUTING FACTORS TO THE AORTIC DISSECTION AND SUBSEQUENT DEATH. Manufacturer Narrative: UPDATED DATA: H6. ADDITIONAL CODES. PRODUCT ANALYSIS: AS THE EVENT REPORTED AN ADVERSE EVENT WITH A POSSIBLE FAILURE OF A DEVICE, LABELING, OR PACKAGING TO MEET ANY OF ITS SPECIFICATIONS, AN INVESTIGATION WAS REQUIRED. AS THE DEVICE REMAINED IN THE PATIENT AT THE TIME THIS INVESTIGATION WAS COMPLETED, NO RETURN PRODUCT ANALYSIS WAS ABLE TO BE PERFORMED. THE REPORTED EVENT IS UNCONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. NO POTENTIAL MANUFACTURING ISSUES THAT MAY BE RELATED TO THIS EVENT WERE IDENTIFIED. THE DEVICE MET ALL MANUFACTURING SPECIFICATIONS UPON RELEASE. A MEDICAL SAFETY ASSESSMENT WAS PERFORMED FOR THIS EVENT. THE EXCERPT OF THE MEDICAL SAFETY SUBJECT MATTER EXPERT (SME) REVIEW REPORT FOLLOWS: BASED ON THE SEQUENCE OF EVENTS AND MEDICAL EXPERTISE, THE EVENT OF VALVE DISLODGEMENT (TO THE ASCENDING AORTA), REQUIRING PLACEMENT OF A 2ND EVOLUT R VALVE, IS LIKELY RELATED TO THE DEVICE THAT COULD NOT SEAT WELL IN THE ANNULUS LIKELY DUE TO SIGNIFICANT AORTIC VALVE CALCIFICATION AS A CONTRIBUTING FACTOR. THE EVENT OF AORTIC DISSECTION, LEADING TO CARDIOGENIC SHOCK AND INITIATION OF LIFE-SAVING MEASURES, IS LIKELY RELATED TO THE 2ND DCS CROSSING THROUGH THE FIRST EVOLUT R VALVE CAUSING THE VALVE TO "MOVE" AND LIKELY RESULTING IN THE AORTIC DISSECTION PER SITE REPORTING. THE EVENT OF DEATH IS LIKELY RELATED TO COMPLICATIONS FROM THE AORTIC DISSECTION. NO ADDITIONAL TECHNICAL ASSESSMENTS WERE PERFORMED. CONCLUSION: THE VALVE INSTRUCTIONS FOR USE (IFU) LISTS DISLODGEMENT, DEATH, CARDIAC ARREST, DISSECTION, AND HEART FAILURE AS POTENTIAL RISKS ASSOCIATED WITH THE TREATMENT PROCEDURES. THERE IS NO IDENTIFIED INADEQUACY WITH THE IFU IN RELATION TO THIS EVENT. THIS EVENT WILL CONTINUE TO BE MONITORED AND TRENDED. BASED ON THE AVAILABLE INFORMATION AND INVESTIGATION ACTIVITIES, IT IS UNABLE TO BE DETERMINED IF THE MEDTRONIC DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. THEREFORE, IT IS UNABLE TO BE DETERMINED IF THE EVENT IS RELATED TO THE DEVICE. POTENTIAL FACTORS THAT CAN INFLUENCE A DISLODGED VALVE INCLUDE TENSION APPLIED ON THE DELIVERY CATHETER SYSTEM (DCS) DURING POSITIONING, CALCIFICATION LEVELS IN THE NATIVE VESSEL, COMPLIANCE OF THE AORTA AND NATIVE VESSELS, INADEQUATE DEPLOYMENT DUE TO IMPROPER SIZING BETWEEN VALVE AND ANNULUS, IRREGULAR PATIENT ANATOMY, INCOMPLETE FRAME EXPANSION, AND A NUMBER OF OTHERS. DISLODGE EVENTS ARE TYPICALLY NOT RELATED TO A DEVICE MALFUNCTION. THE EVENT OF VALVE DISLODGEMENT (TO THE ASCENDING AORTA), REQUIRING PLACEMENT OF A 2ND EVOLUT R VALVE, IS LIKELY RELATED TO THE DEVICE THAT COULD NOT SEAT WELL IN THE ANNULUS LIKELY DUE TO SIGNIFICANT AORTIC VALVE CALCIFICATION AS A CONTRIBUTING FACTOR. UNFORTUNATELY, A RELATIONSHIP OF THE REPORTED CARDIAC ARREST TO THE DEVICE OR PROCEDURE COULD NOT BE DETERMINED AND A CONCLUSIVE CAUSE COULD NOT BE DETERMINED FROM THE LIMITED INFORMATION AVAILABLE. UNFORTUNATELY, A RELATIONSHIP OF THE REPORTED HEART FAILURE TO THE DEVICE OR PROCEDURE COULD NOT BE DETERMINED WITH THE LIMITED INFORMATION AVAILABLE, THOUGH IT IS LIKELY RELATED TO PATIENT CONDITION. THE COMPLAINT INVESTIGATION DETERMINED THAT THIS EVENT IS FORESEEN IN RISK MANAGEMENT AND IS INCLUDED IN MONITORING/TRENDING. VASCULAR COMPLICATIONS, SUCH AS DISSECTION, CAN OCCUR DURING ANY PERCUTANEOUS INTERVENTION. THESE EVENTS ARE TYPICALLY RELATED TO PATIENT FACTORS (TORTUOUS ANATOMY, CALCIFICATION, ETC.), AND/OR PROCEDURAL EFFECTS (SHEATH USED, USER TECHNIQUE, GUIDEWIRE MANAGEMENT, ETC.). UNFORTUNATELY, THIS COULD NOT BE CONFIRMED BY MEDTRONIC WITHOUT A COPY OF THE ECHOCARDIOGRAM OR IMAGING FILM FOR REVIEW, THE FAILURE MECHANISM OF THE DEVICE CANNOT BE ESTABLISHED, AND THE CONCLUSIVE CAUSE OF THE REPORTED DISSECTION CANNOT BE DETERMINED. A PROCEDURE- OR VALVE-RELATED DEATH IS AN INHERENT RISK WHEN THE PATIENT CONDITION IS SUCH THAT A TRANSCATHETER AORTIC VALVE IS NEEDED TO SUSTAIN CARDIAC FUNCTION, AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC DISSECTION. MEDICAL SAFETY REVIEW INDICATES THE EVENT OF DEATH IS LIKELY RELATED TO COMPLICATIONS FROM THE AORTIC DISSECTION. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED AND WITH VALVE REMAINED IMPLANTED, A CONCLUSIVE ASSESSMENT OF THE RELATIONSHIP BETWEEN THE EVENT AND THE DEVICE COULD NOT BE REACHED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2 B5 H6 ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21339234&pc=NPT | 2015691-2025-00873 | 2025-01-15 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-07 05:00 | NPT | COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM A FIELD CLINICAL SPECIALIST THAT A PATIENT UNDERWENT TMVR IN THE NATIVE MITRAL POSITION CASE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE. AS REPORTED, THIS WAS A MITRAL ANNULAR CALCIFICATION CASE. THE 29MM S3UR VALVE WAS IMPLANTED SUCCESSFULLY. THEN ECHO SHOWED INCREASING PERICARDIAL EFFUSION. ANGIO SHOWED WHAT LOOKED TO BE A LV WIRE PERF (THE WIRE USED WAS NOT MANUFACTURED BY EDWARD'S). THE DELIVERY SYSTEM WASN'T BELIEVED TO BE THE WIRE AT THE TIME OF THE PERFORATION, HOWEVER THEY DO NOT KNOW IF IT WAS. THERE WERE NO ALLEGATIONS OF ANY EDWARDS DEVICE FAILURES WERE MADE. IT WAS THOUGHT THAT THE WIRE PERF HAPPENED WHEN THEY WERE TESTING THE PACING OVER THE WIRE. INITIALLY THE LV PERF WAS TREATED WITH A PERICARDIOCENTESIS. SURGERY EVALUATED BUT NOTHING FURTHER WAS DONE. THE PATIENT WAS NOT ALIVE AT THE END OF THE PROCEDURE. THE FAMILY DID NOT WANT TO PROCEED WHEN THEY HEARD WHAT THE SURGICAL TEAM HAD TO SAY ABOUT FURTHER TREATMENT. Manufacturer Narrative: PER THE REPORT, THE DEVICE WAS USED IN AN OFF-LABEL IMPLANTATION IN THE NATIVE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL THV PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A THV PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PROCEDURE RELATED FACTORS (PERFORATION DURING TESTING OF THE PACING OVER THE WIRE) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21339900&pc=NPT | 2025587-2025-01106 | 2024-11-28 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-07 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Dyspnea; Dizziness | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 YEARS FOLLOWING THE IMPLANT OF THE BIOPROSTHETIC TRANSCATHETER VALVE THE PATIENT EXPERIENCE SHORTNESS OF BREATH AND LIGHTHEADEDNESS. LATER THAT SAME DAY THE PATIENT UNEXPECTEDLY DIED WHILE AT HOME. THE CAUSE OF DEATH WAS REPORTED AS CARDIAC. AN AUTOPSY WAS NOT PERFORMED. THE PHYSICIAN INDICATED THE DEATH WAS NOT RELATED TO THE MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21320251&pc=KFM | 8010762-2025-0000054 | 2025-01-31 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-02-06 05:00 | KFM | HEART LUNG MACHINE | Electrical /Electronic Property Problem | Adult Respiratory Distress Syndrome; No Clinical Signs, Symptoms or Conditions | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN INDIA. IT WAS REPORTED THAT THE ERROR MESSAGE ¿LAMPTEST¿ WAS DISPLAYED ON THE ROTAFLOW CONSOLE. IT WAS NOTICED DURING PATIENT TREATMENT AND THE DEVICE WAS IMMEDIATELY EXCHANGED WITH A BACKUP DEVICE. DURING SERVICE THE ERROR MESSAGE ¿HEAD ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE. NO HARM TO ANY PERSON HAS BEEN REPORTED. DUE TO THE REPORTED EXCHANGE OF THE DEVICE DURING PATIENT TREATMENT AND DUE TO THAT THE REPORTED FAILURE "HEAD ERROR" COULD LEAD TO THE PUMP STOP A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). Event Description: NEW INFORMATION HAS BEEN PROVIDED BY THE SSU DATED ON 2025-03-20 THAT THE PATIENT PASSED AWAY. COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THIS EVENT OCCURRED ON THE INDIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO ROTAFLOW, WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#(B)(4), PRIMARY DI NUMBER HAS BEEN USED FOR ROTAFLOW WITH CATALOG NUMBER 701043291. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321024&pc=CBK | 9611500-2025-00063 | 2025-01-28 05:00 | Death | DRÄGERWERK AG & CO. KGAA | 2025-02-06 05:00 | CBK | EVITA 4 | Unexpected Therapeutic Results; Human-Device Interface Problem | Insufficient Information | K980642 |  | 1 | Event Description: IT WAS REPORTED THAT THE EVITA 4 WENT OUT DURING PREPARATION FOR TRANSFER TO HELICOPTER TRANSPORT AND WOULD NOT HAVE CONTINUED TO VENTILATE. IT WAS REPORTED THAT DUE TO THE SITUATION (INCLUDING A CHANGE IN VITAL SIGN MONITORING), THE DETERIORATION IN VITAL SIGNS WAS NOT NOTICED IMMEDIATELY AND THE PATIENT PASSED AWAY. THE DOCTORS RESPONSIBLE COULD NEITHER CONFIRM NOR COMPLETELY RULE OUT WHETHER THE FAILURE OF THE VENTILATOR WAS THE CAUSE OF DEATH. ACCORDING TO THE DOCTORS, THE VENTILATOR DISPLAY WAS BLACK. THE DEVICE COULD BE RESTARTED IMMEDIATELY AND FUNCTIONED WITHOUT ANY PROBLEMS. IT WAS NOT POSSIBLE TO SAY FOR CERTAIN WHETHER THE DEVICE HAD ALARMED THE MALFUNCTION. THE DEVICE HAS NO UDI AS AN UDI WAS NOT REQUIRED AT THE TIME THE DEVICE WAS MANUFACTURED. Manufacturer Narrative: THE LOG FILE OF THE AFFECTED DEVICE WAS ANALYSED REGARDING THE REPORTED EVENT. BASED ON THE LOG FILE ANALYSIS THE REPORTED EVENT COULD BE TRACED. THE DEVICE WAS OPERATED IN BIPAP ASB VENTILATION MODE DURING THE EVENT. DURING VENTILATION, THE DEVICE ISSUED THE ALARMS ¿MV LOW!!!¿, ¿MV HIGH!!!¿, ¿TIDAL VOLUME HIGH!!!¿, ¿FIO2 HIGH!!!¿ AND ¿ETCO2 HIGH!!!¿ OUTPUT. AT 12:13 P.M. THE DEVICE WAS SWITCHED TO STANDBY MODE, WHICH WAS ALARMED AS SPECIFIED WITH ¿STANDBY ACTIVATED!!!¿ AND THEN CONFIRMED WITH THE "ALARM RESET" BUTTON. AT 12:19 P.M., THE DEVICE WAS SWITCHED BACK INTO OPERATION AND AGAIN EMITTED THE VENTILATION-RELEVANT ALARM MESSAGES. THE DEVICE WAS SWITCHED BACK TO STANDBY MODE AND SWITCHED OFF BY THE USER AT 12:45 P.M.. THE ANALYSIS OF THE LOGFILE COULD NOT FIND ANY INDICATION OF A DEVICE MALFUNCTION. IF THE "STANDBY" HARDKEY IS PRESSED CONTINUOUSLY FOR 3 SECONDS BY THE USER DURING OPERATION, EVITA 4 SWITCHES TO "STANDBY" MODE AND ISSUES A VISUAL AND AUDIBLE ALARM WITH THE WARNING "STANDBY ACTIVATED !!!". THIS ALARM REMAINS ACTIVE UNTIL THE USER CONFIRMS THE ALARM BY PRESSING THE "ALARM RESET" HARDKEY. IN STANDBY MODE, THE PATIENT CAN BREATHE SPONTANEOUSLY VIA THE EMERGENCY BREATHING DEVICE. THE DEVICE HAS BEHAVED AS SPECIFIED AND ALERTED THE USER ACOUSTICALLY AND VISUALLY TO THE CHANGE TO STANDBY MODE. THE NUMBER OF SIMILAR CASES, RELATED TO THE SAME ROOT CAUSE, IS WITHIN THE EXPECTED RANGE OF THE RESPECTIVE RISK ASSESSMENT AND THUS ACCEPTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321126&pc=OZD | 1220648-2025-26023 | 2025-01-18 05:00 | Death | ABIOMED, INC. | 2025-02-06 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Defective Alarm | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED WHILE SUPPORTING A PATIENT IMPLANTED WITH AN IMPELLA 5.5 DEVICE, THE AUTOMATED IMPELLA CONTROLLER (AIC) ALARM NOISE WAS VERY LOW AND MAKING A DIFFERENT SOUND THAN NORMAL. A SILENT HUM WAS OBSERVED RATHER THAN THE NUMBER BEEP THE AIC MAKES FOR WHITE INFORMATIONAL ALARM. THE STATED PLAN IS TO RETURN THE AIC FOR REPAIR AFTER THE PATIENT COMPLETES THE IMPELLA 5.5 MECHANICAL CIRCULATORY SUPPORT. THERE WAS NO HARM TO THE PATIENT AS A RESULT OF THIS EVENT WITH NOT REPORT OR INDICATION OF UNINTERRUPTED SUPPORT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. PER MEDICAL SAFETY OFFICER REVIEW THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE AUTOMATED IMPELLA CONTROLLER (AIC) AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE CONSOLE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - BUZZER/ ALARM INAUDIBLE HAS BEEN COMPLETED. THE CONSOLE LOGS FROM THE REPORTED DAY OF EVENT DID NOT CONTAIN ANY INFORMATION RELEVANT TO THE REPORTED ISSUE. THE CONSOLE WAS RECEIVED FOR INVESTIGATION WHERE IT WAS EXTENSIVELY TESTED. WE WERE UNABLE TO REPRODUCE ANY AUDIBLE ANOMALY, INCLUDING ONE THAT WOULD MATCH SPECIFIC PROBLEM DESCRIPTION. THE CAUSE OF THE ISSUE WAS NOT DETERMINED SINCE ISSUE COULD NOT BE REPRODUCED. B.5 ADDITIONAL INFORMATION WAS RECEIVED, AND ABIOMED'S MEDICAL SAFETY OFFICER REVIEW WAS COMPLETED. B.1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26023 TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM. B.2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26023 TO REFLECT DEATH AND DATE OF DEATH. E.1 FAX NUMBER WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26023. H.1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26023 TO REFLECT DEATH. H.6 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26023 TO REFLECT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321162&pc=DWF | 8010762-2025-0000053 | 2024-11-29 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-02-06 05:00 | DWF | CANNULA & CATHETER | Disconnection | Air Embolism; Hypovolemic Shock | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT VENOUS CANNULA WAS DISCONNECTED DURING PROCEDURE. THE PATIENT HAS BEEN IN PALLIATIVE CARE UNDER BIVENTRICULAR SUPPORT SINCE 23/11. ON 24/11, THE CANNULA INITIALLY PLACED IN THE RIGHT FEMORAL WAS REPLACED BY A CANNULA PLACED IN THE LEFT FEMORAL DUE TO SUSPECTED NERVE COMPRESSION BY THE FIRST CANNULA. ON 29/11, ACCIDENTAL VENOUS DECANNULATION OCCURRED WITH HEMORRHAGIC SHOCK AND AIR ENTRAINMENT IN THE CIRCUIT. THE VENOUS CANNULA WAS FIXED WITH ONLY ONE SUTURE POINT. THE REMOVABLE FIXATION DEVICE PROVIDED WITH THE CANNULA IS RARELY USED AS IT IS DEEMED INSUFFICIENT. PATIENT DIED DURING TREATMENT. AN ANALYSIS MEETING WAS ORGANIZED BY THE RELATED DEPARTMENT OF HOSPITAL. IT WAS DECIDED THAT THE CANNULA FIXATION PROTOCOL WILL BE DRAFTED BY SURGEONS, AND THE MONITORING PROCEDURES WILL BE DETAILED BY PARAMEDICAL PROFESSIONALS. COMPLAINT #(B)(4).Event Description: COMPLAINT # (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: IT WAS REPORTED THAT VENOUS CANNULA WAS DISCONNECTED DURING PROCEDURE. THE PATIENT HAS BEEN IN PALLIATIVE CARE UNDER BIVENTRICULAR SUPPORT SINCE 23/11. ON 24/11, THE CANNULA INITIALLY PLACED IN THE RIGHT FEMORAL WAS REPLACED BY A CANNULA PLACED IN THE LEFT FEMORAL DUE TO SUSPECTED NERVE COMPRESSION BY THE FIRST CANNULA. ON 29/11, ACCIDENTAL VENOUS DECANNULATION OCCURRED WITH HEMORRHAGIC SHOCK AND AIR ENTRAINMENT IN THE CIRCUIT. THE VENOUS CANNULA WAS FIXED WITH ONLY ONE SUTURE POINT. THE REMOVABLE FIXATION DEVICE PROVIDED WITH THE CANNULA IS RARELY USED AS IT IS DEEMED INSUFFICIENT. PATIENT HAS EXPIRED DUE TO ADVERSE EVENT. SAMPLE INVESTIGATION COULD NOT BE PERFORMED SINCE THE PRODUCT WAS DISCARDED BY CUSTOMER. MORE INFORMATION HAS BEEN REQUESTED FROM CUSTOMER BY MEDICAL AFFAIRS IN ORDER TO INVESTIGATE THE ISSUE, HOWEVER IT WAS REPORTED BY GETINGE SALES & SERVICE REPRESENTATIVE THAT THE CUSTOMER IS NOT ABLE TO PROVIDE FURTHER INFORMATION. MEDICAL REVIEW HAS BEEN PERFORMED ON 2025-02-26. THE DESCRIBED DECANNULATION OF THE HLS CANNULA (EXPLAINED AS ACCIDENTAL) RESULTED IN HAEMORRHAGIC SHOCK WHICH LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. IT WAS REPORTED IN THE COMPLAINT THAT THE VENOUS CANNULA WAS FIXED WITH A SINGLE STITCH AND THAT THE FIXATION DEVICE PROVIDED WITH THE CANNULA WAS NOT USED AS IT IS DEEMED INSUFFICIENT. THE INSTRUCTION FOR USE OF THE HLS CANNULA THAT WAS PLACED ON THE PATIENT EXPLAINS THE FOLLOWING REGARDING WARNINGS AND PRECAUTIONS. IFU, HLS CANNULAE, G-139, V05 5.3 SAFETY INSTRUCTIONS FOR CANNULAE -IF THE PATIENT IS REPOSITIONED OR TRANSPORTED, THERE IS A RISK OF DECANNULATION CAUSED BY STRAIN ON THE TUBING AND MECHANICAL DAMAGE. THE GREATEST CARE SHOULD THEREFORE BE EXERCISED WHEN CARRYING OUT THESE MEASURES. -ENSURE THAT THERE IS NO STRAIN ON THE CANNULA. -AVOID SUBJECTING THE CANNULA TO MECHANICAL LOADS. IFU, HLS CANNULAE, G-139, V05 7 APPLICATION -PERFORM A CAREFUL VISUAL INSPECTION OF THE DEVICE BEFORE USE. IN PARTICULAR, ENSURE THERE IS NO DAMAGE TO THE MATERIAL, CRACKS, BURRS OR FISSURES. CAUTION! ATTACH THE CANNULA SECURELY TO THE VESSEL OR SURROUNDING TISSUE TO PREVENT ACCIDENTAL DISLOCATION DURING THE EXTRACORPOREAL CIRCULATION. CAUTION! IF A SUTURE IS DIRECTLY LOOPED AROUND THE WIRE-REINFORCED SECTION, THE CANNULA CAN BE CUT INTO, KINKED OR DAMAGED. SECURE THE CANNULA AROUND THE RIGID BARBED CONNECTORS IN SUCH A WAY AS TO ENSURE THAT IT IS NOT KINKED AT THE INSERTION SITE. -ATTACH THE CANNULA SECURELY TO THE VESSEL. FOR THIS PURPOSE, USE THE REMOVABLE SKIN ATTACHMENT ACCESSORY SUPPLIED OR ELSE A SLIT AT THE END OF THE CANNULA FOR THE SUTURE MATERIAL. THE REMOVABLE SKIN ATTACHMENT ACCESSORY HAS HOLES TO TAKE THE SUTURE MATERIAL. CAUTION! CHECK THE CONNECTION BETWEEN CANNULA AND TUBING SET REGULARLY. IN ABSENCE OF MORE DETAILS, A PROBABLE ROOT CAUSE OF THE REPORTED DECANNULATION CANNOT BE DETERMINED. FURTHER, WITHOUT FURTHER INFORMATION REGARDING THE EVENT, POSSIBLE ROOT CAUSES CANNOT BE PROPOSED. BASED ON THE INVESTIGATION RESULTS, COMPLAINT COULD BE CONFIRMED HOWEVER PRODUCT RELATED INFLUENCES ARE UNLIKELY. BESIDES, THE LOT NUMBER OF THE AFFECTED HLS CANNULAE WAS NOT PROVIDED BY CUSTOMER. HOWEVER, THE FOLLOWING BATCH NUMBERS WERE FOUND VIA SAP THAT ARE SOLD TO CHU BORDEAUX - HOPITAL PELLEGRIN WITHIN THE SHELF LIFE: 3000434622, 3000432796, 3000427010, 3000423314, 3000414511, 3000404875, 3000401849, 3000367020, 3000353387, 3000332654, 3000304537. THE PRODUCTION HISTORY RECORDS (DHR) OF THE AFFECTED BE-PVL 2155#BE-HLS CANNULA 21F VL WITH THE ABOVE LOT #S WERE REVIEWED ON 2025-01-29. ACCORDING TO THE DHR RESULT, THE PRODUCT BE-PVL 2155#BE-HLS CANNULA 21F VL PASSED THE DEFINED MANUFACTURING AND FINAL RELEASE SPECIFICATIONS. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED FAILURE AND PRODUCT AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321296&pc=BZD | 2518422-2025-101789 | 2024-10-24 04:00 | Death | RESPIRONICS, INC. | 2025-02-06 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. THE EXECUTOR OF THE ESTATE CONTACTED US, MENTIONING THAT SHE HAD RECEIVED A LETTER FROM PHILIPS REGARDING A REPLACEMENT DEVICE. HOWEVER, SINCE THE PATIENT HAD ALREADY PASSED AWAY, SHE WAS UNCERTAIN WHETHER THE DEVICE WAS IN THEIR POSSESSION. SHE STATED THAT SHE WOULD FOLLOW UP IF THE DEVICE IS LOCATED SO THAT IT CAN BE RETURNED. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS INVOLVED IN THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321616&pc=DWF | 3013876692-2025-00008 | 2024-11-29 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-02-06 05:00 | DWF | CANNULA & CATHETER | Disconnection | Air Embolism; Hypovolemic Shock | K102532 |  | 1 | Event Description: IT WAS REPORTED THAT VENOUS CANNULA WAS DISCONNECTED DURING PROCEDURE. THE PATIENT HAS BEEN IN PALLIATIVE CARE UNDER BIVENTRICULAR SUPPORT SINCE 23/11. ON 24/11, THE CANNULA INITIALLY PLACED IN THE RIGHT FEMORAL WAS REPLACED BY A CANNULA PLACED IN THE LEFT FEMORAL DUE TO SUSPECTED NERVE COMPRESSION BY THE FIRST CANNULA. ON 29/11, ACCIDENTAL VENOUS DECANNULATION OCCURRED WITH HEMORRHAGIC SHOCK AND AIR ENTRAINMENT IN THE CIRCUIT. THE VENOUS CANNULA WAS FIXED WITH ONLY ONE SUTURE POINT. THE REMOVABLE FIXATION DEVICE PROVIDED WITH THE CANNULA IS RARELY USED AS IT IS DEEMED INSUFFICIENT. PATIENT DIED DURING TREATMENT. AN ANALYSIS MEETING WAS ORGANIZED BY THE RELATED DEPARTMENT OF HOSPITAL. IT WAS DECIDED THAT THE CANNULA FIXATION PROTOCOL WILL BE DRAFTED BY SURGEONS, AND THE MONITORING PROCEDURES WILL BE DETAILED BY PARAMEDICAL PROFESSIONALS. COMPLAINT # (B)(4).Event Description: COMPLAINT- # (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: IT WAS REPORTED THAT VENOUS CANNULA WAS DISCONNECTED DURING PROCEDURE. THE PATIENT HAS BEEN IN PALLIATIVE CARE UNDER BIVENTRICULAR SUPPORT SINCE (B)(6). ON (B)(6), THE CANNULA INITIALLY PLACED IN THE RIGHT FEMORAL WAS REPLACED BY A CANNULA PLACED IN THE LEFT FEMORAL DUE TO SUSPECTED NERVE COMPRESSION BY THE FIRST CANNULA. ON (B)(6), ACCIDENTAL VENOUS DECANNULATION OCCURRED WITH HEMORRHAGIC SHOCK AND AIR ENTRAINMENT IN THE CIRCUIT. THE VENOUS CANNULA WAS FIXED WITH ONLY ONE SUTURE POINT. THE REMOVABLE FIXATION DEVICE PROVIDED WITH THE CANNULA IS RARELY USED AS IT IS DEEMED INSUFFICIENT. PATIENT HAS EXPIRED DUE TO ADVERSE EVENT. SAMPLE INVESTIGATION COULD NOT BE PERFORMED SINCE THE PRODUCT WAS DISCARDED BY CUSTOMER. MORE INFORMATION HAS BEEN REQUESTED FROM CUSTOMER BY MEDICAL AFFAIRS IN ORDER TO INVESTIGATE THE ISSUE, HOWEVER IT WAS REPORTED BY GETINGE SALES & SERVICE REPRESENTATIVE THAT THE CUSTOMER IS NOT ABLE TO PROVIDE FURTHER INFORMATION. MEDICAL REVIEW HAS BEEN PERFORMED ON (B)(6)2025. THE DESCRIBED DECANNULATION OF THE HLS CANNULA (EXPLAINED AS ACCIDENTAL) RESULTED IN HAEMORRHAGIC SHOCK WHICH LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. IT WAS REPORTED IN THE COMPLAINT THAT THE VENOUS CANNULA WAS FIXED WITH A SINGLE STITCH AND THAT THE FIXATION DEVICE PROVIDED WITH THE CANNULA WAS NOT USED AS IT IS DEEMED INSUFFICIENT. THE INSTRUCTION FOR USE OF THE HLS CANNULA THAT WAS PLACED ON THE PATIENT EXPLAINS THE FOLLOWING REGARDING WARNINGS AND PRECAUTIONS. IFU, HLS CANNULAE, G-139, V05 5.3 SAFETY INSTRUCTIONS FOR CANNULAE -IF THE PATIENT IS REPOSITIONED OR TRANSPORTED, THERE IS A RISK OF DECANNULATION CAUSED BY STRAIN ON THE TUBING AND MECHANICAL DAMAGE. THE GREATEST CARE SHOULD THEREFORE BE EXERCISED WHEN CARRYING OUT THESE MEASURES. -ENSURE THAT THERE IS NO STRAIN ON THE CANNULA. -AVOID SUBJECTING THE CANNULA TO MECHANICAL LOADS. IFU, HLS CANNULAE, G-139, V05 7 APPLICATION -PERFORM A CAREFUL VISUAL INSPECTION OF THE DEVICE BEFORE USE. IN PARTICULAR, ENSURE THERE IS NO DAMAGE TO THE MATERIAL, CRACKS, BURRS OR FISSURES. CAUTION! ATTACH THE CANNULA SECURELY TO THE VESSEL OR SURROUNDING TISSUE TO PREVENT ACCIDENTAL DISLOCATION DURING THE EXTRACORPOREAL CIRCULATION. CAUTION! IF A SUTURE IS DIRECTLY LOOPED AROUND THE WIRE-REINFORCED SECTION, THE CANNULA CAN BE CUT INTO, KINKED OR DAMAGED. SECURE THE CANNULA AROUND THE RIGID BARBED CONNECTORS IN SUCH A WAY AS TO ENSURE THAT IT IS NOT KINKED AT THE INSERTION SITE. -ATTACH THE CANNULA SECURELY TO THE VESSEL. FOR THIS PURPOSE, USE THE REMOVABLE SKIN ATTACHMENT ACCESSORY SUPPLIED OR ELSE A SLIT AT THE END OF THE CANNULA FOR THE SUTURE MATERIAL. THE REMOVABLE SKIN ATTACHMENT ACCESSORY HAS HOLES TO TAKE THE SUTURE MATERIAL. CAUTION! CHECK THE CONNECTION BETWEEN CANNULA AND TUBING SET REGULARLY. IN ABSENCE OF MORE DETAILS, A PROBABLE ROOT CAUSE OF THE REPORTED DECANNULATION CANNOT BE DETERMINED. FURTHER, WITHOUT FURTHER INFORMATION REGARDING THE EVENT, POSSIBLE ROOT CAUSES CANNOT BE PROPOSED. BASED ON THE INVESTIGATION RESULTS, COMPLAINT COULD BE CONFIRMED HOWEVER PRODUCT RELATED INFLUENCES ARE UNLIKELY. BESIDES, THE LOT NUMBER OF THE AFFECTED HLS CANNULAE WAS NOT PROVIDED BY CUSTOMER. HOWEVER, THE FOLLOWING BATCH NUMBERS WERE FOUND VIA SAP THAT ARE SOLD TO (B)(6) WITHIN THE SHELF LIFE: 3000434622, 3000432796, 3000427010, 3000423314, 3000414511, 3000404875, 3000401849, 3000367020, 3000353387, 3000332654, 3000304537. THE PRODUCTION HISTORY RECORDS (DHR) OF THE AFFECTED BE-PVL 2155#BE-HLS CANNULA 21F VL WITH THE ABOVE LOT #S WERE REVIEWED ON (B)(6)2025. ACCORDING TO THE DHR RESULT, THE PRODUCT BE-PVL 2155#BE-HLS CANNULA 21F VL PASSED THE DEFINED MANUFACTURING AND FINAL RELEASE SPECIFICATIONS. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED FAILURE AND PRODUCT AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321857&pc=OZD | 1220648-2025-26033 | 2025-01-18 05:00 | Death | ABIOMED, INC. | 2025-02-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS), AND THE FOLLOWING START OF SUPPORT, THE PATIENT DEVELOPED LIMB ISCHEMIA AND EXPIRED. THE PATIENT WAS IN A HEMODIALYSIS RUN AND STARTED TO FEEL WEEK; SYSTOLIC BLOOD PRESSURE WAS IN THE 40'S. HEPARIN AND LEVO DRIPS WERE STARTED, AND THE PATIENT WAS TRANSFERRED TO THE HOSPITAL. A LEFT HEART CATHETERIZATION WAS PERFORMED FOR CONCERN OF AN ACUTE MYOCARDIAL INFARCTION. "EXTENSIVE" LEFT ANTERIOR DESCENDING ARTERY DISEASE WAS NOTED AND VESSEL SIZE TOO SMALL TO CORRECT. ADDITIONALLY, THERE WAS DIFFUSE DISEASE IN RIGHT CORONARY ARTERY. THE LEFT VENTRICULAR END-DIASTOLIC PRESSURE WAS 25 AND THE PATIENT'S HYPOTENSION WAS WORSENING. THE DECISION MADE TO PLACE IMPELLA CP VIA LEFT FEMORAL ARTERY WHICH WAS SUCCESSFULLY COMPLETED. UPON ARRIVAL TO THE UNIT, LEFT PEDAL PULSE WAS NOT PALPABLE. THE NURSE ATTEMPTED TO FIND THE PEDAL PULSE USING A DOPPLER WITH NO SUCCESS. THE PHYSICIAN WAS INFORMED, IT WAS NOTED THE PATIENT HAD SEVERE VASCULAR DISEASE, AND, HOWEVER, DUE TO THE PATIENT'S DEGREE OF CARDIOGENIC SHOCK, THE PATIENT NEEDED THE TEMPORARY MCS. REPORTEDLY THE PATIENT WOULD EXPIRE THE NEXT DAY. NO FURTHER DETAILS SURROUNDING THE DEATH WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED LIMB ISCHEMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B.2 REVISED SELECTIONS AS DISABILITY WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26033. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26033 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4614 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26033.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321860&pc=OZD | 1220648-2025-26034 | 2025-01-18 05:00 | Death | ABIOMED, INC. | 2025-02-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT, IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION, IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT DEVELOPED LIMB ISCHEMIA AND WOULD EVENTUALLY EXPIRE. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL VIA AMBULANCE AND CODED DURING TRANSPORT. CARDIOPULMONARY RESUSCITATION WAS PERFORMED WITH RETURN OF SPONTANEOUS CIRCULATION AND THE PATIENT WAS INTUBATED. THE INITIAL ANGIOGRAM SHOWED SMALL DIAMETER OF RIGHT FEMORAL ARTERY AND THERE WAS UNCERTAINTY IF THE IMPELLA SHEATH WOULD FIT. DUE TO PATIENT'S DECLINE IN STATUS THE DECISION WAS MADE TO QUICKLY PLACE AN IMPELLA CP DEVICE. PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED TO THE LEFT CIRCUMFLEX ARTERY. THERE WAS ANOTHER LESION IN THE LEFT ANTERIOR DESCENDING ARTERY LEFT UNTREATED AS THE PHYSICIAN WAS UNABLE TO PASS GUIDEWIRE DUE TO SEVERE CALCIFICATION. THE PLAN WAS TO COOL THE PATIENT AND BRING THE PATIENT BACK FOR TREATMENT AT A LATER DATE. AT END OF CASE THE PEEL-AWAY SHEATH WAS EXCHANGED FOR THE REPOSITIONING SHEATH. THE TEAM WAS UNABLE TO DOPPLER ANY PULSES IN RIGHT FOOT. LOWER LEG WAS COLD, AND BOTTOM OF RIGHT FOOT WAS MOTTLED. DECISION WAS MADE TO PLACE AN EXTERNAL FEM-FEM BYPASS. A 6F ANTEGRADE ACCESS IN RIGHT SUPERFICIAL FEMORAL ARTERY BELOW IMPELLA SHEATH AND 6F RETROGRADE ACCESS IN LEFT FEMORAL ARTERY. SHEATHS WERE CONNECTED AND FLOW WAS NOTED. THE PATIENT WAS TRANSFERRED TO THE UNIT FOR MONITORING, ON SUPPORT, BUT WOULD EXPIRE TWO DAYS LATER. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED LIMB ISCHEMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE LIMB ISCHEMIA COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21321870&pc=BZD | 2518422-2025-101798 | 2024-12-06 05:00 | Death | RESPIRONICS, INC. | 2025-02-06 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE REPORTED RECEIVING A LETTER REGARDING THE RETURN OF HER HUSBAND'S DEVICE, BUT SINCE HE PASSED AWAY TWO YEARS AGO, SHE IS UNSURE OF THE DEVICE'S WHEREABOUTS OR WHAT HAPPENED TO IT. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21322223&pc=JDI | 1038671-2025-00696 |  | Death | EXACTECH, INC. | 2025-02-06 05:00 | JDI | UNKNOWN HIP COMPONENTS | Insufficient Information | Pain; Ambulation Difficulties |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT'S HIP WAS EXPERIENCING ISSUES POST INITIAL OPERATION. IT WAS STATED THAT THE PATIENT WAS UNABLE TO WALK AND SUFFERED PAIN POST-OPERATION. PATIENT HAD ALSO PASSED AWAY THREE MONTHS LATER. THE REPORTED INFORMATION DOES NOT INDICATE INITIAL PROCEDURE DATE OR IF A REVISION WAS PERFORMED PRIOR TO THE DEATH OF THE PATIENT. INFORMATION REGARDING THE NATURE OF THE PATIENT'S DEATH WAS NOT PROVIDED. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: THE REPORTED EVENT WAS UNABLE TO BE CONFIRMED DUE TO LIMITED INFORMATION RECEIVED FROM THE CUSTOMER. NO DEVICE WAS RETURNED FOR EVALUATION; FURTHER, PHOTOGRAPHS AND/OR RADIOGRAPH IMAGES WERE NOT PROVIDED FOR REVIEW. OPERATIVE NOTES AND/OR MEDICAL RECORDS WERE NOT PROVIDED FOR REVIEW OF USAGE/ TECHNIQUE. A REVIEW OF MANUFACTURING DATA WAS UNABLE TO BE PERFORMED AS THE LOT INFORMATION OF THE PRODUCT INVOLVED IN THE EVENT WAS NOT AVAILABLE. A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AS THE NECESSARY INFORMATION TO ADEQUATELY INVESTIGATE THE REPORTED EVENT WAS NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21322247&pc=OZD | 1220648-2025-26039 | 2025-01-21 05:00 | Death | ABIOMED, INC. | 2025-02-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PUMP PLACED TO SUPPORT FOR A PATIENT CODING AND IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE PUMP WAS PLACED AND ON 2ND DAY OF SUPPORT AN ACCESS SITE BLEED WAS NOTED. THE BLEED WAS TREATED BY THREE UNITS OF BLOOD AND PRESSURE BEING HELD WHILE MEDICAL THERAPY AND FEMOSTOP PLACED. FAMILY DECIDED TO WITHDRAW CARE AND PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS DETERMINED TO BE ANTICOAGULATION MANAGEMENT BECAUSE ANTICOAGULATION WAS SUPRATHERAPEUTIC AND HEPARIN WAS DISCONTINUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21322401&pc=BZD | 2518422-2025-101805 | 2024-12-12 05:00 | Death | RESPIRONICS, INC. | 2025-02-06 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S HUSBAND CONTACTED US, EXPRESSING HIS INTENTION TO RETURN THE DEVICE BUT MENTIONED HE WAS UNFAMILIAR WITH HOW TO USE FEDEX. HE WAS REFERRED TO THE DME FOR FURTHER ASSISTANCE. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21322752&pc=MSX | 1218950-2025-000084 | 2024-10-31 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-02-06 05:00 | MSX | CAREEVENT | No Audible Alarm | Insufficient Information | K161164 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THE ALARMS DID NOT RING ON THE SMARTPHONE AND REQUESTED A LOG REVIEW. THE LOGS INDICATED THAT THE ALARMS OCCURRED ON THE MONITOR AND CENTRAL STATION AND WERE TRANSMITTED TO THE SMARTPHONE. ADDITIONAL INFORMATION IS REQUESTED FOR FURTHER CLARIFICATION. THE DEVICE WAS IN CLINICAL USE AT TIME OF EVENT AND A PATIENT DIED. Manufacturer Narrative: THE FOLLOWING FUNCTIONAL TESTS WERE PERFORMED: PHILIPS WAS NOTIFIED ON (B)(6) 2025, BUT ALL THE LOGS FROM THE EVENT DATE (31 OCT 2024) HAD BEEN OVERWRITTEN AND THE ONLY LOGS THAT COULD BE RETRIEVED WAS THE ONE COLLECTED BY THE CUSTOMER. THE FIELD SERVICE ENGINEER (FSE) PROVIDED AN ANALYSIS OF THOSE LOGS. FOR THE EVENT TABLE FOR (B)(6) 2024 BETWEEN 08:00 AND 09:00. 223 EVENTS, 0 EVENTS NOT TRANSMITTED, 22 ESCALATIONS. THE PHILIPS FIELD SERVICE ENGINEER, SUPPORT (B)(6), PHILIPS TECHNICAL MANAGER IN CHARGE AND THE CUSTOMER SENIOR BIOMEDICAL TECHNICIAN CONDUCTED PERFORMANCE VERIFICATION TESTING ONSITE. THE CUSTOMER'S ALARM CONFIGURATION OF CAREEVENT WAS TESTED AND VERIFIED. IN ADDITION, THE REVIEW STATED THAT THE "AUTO REMOVE ENDED EVENTS" FEATURE WAS UNCHECKED. AS CONFIGURED, THIS SETTING AUTOMATED THE DELETION OF COMPLETED EVENTS IN THE PHONE HISTORY, RESULTING IN THE DISAPPEARANCE OF SOME EVENTS FROM THE HISTORY. THIS COULD EXPLAIN THE CUSTOMER'S OBSERVATIONS. AFTER RESTARTING THE VIRTUAL SERVER AND CHECKING THE POINTS STATED IN THE SERVICE MANUAL, IT SHOWED THAT THE CAREEVENT SYSTEM WAS OPERATING PER SPECIFICATIONS. ALARMS THAT ARE SET TO BE SENT TO PHONES ARE BEING SENT. ESCALATIONS ARE TRIGGERED AFTER THE CONFIGURED TIME OF ONE MINUTE ON THE DEFINED PHONES. USERS DESIGNATED FOR DELEGATION ARE NOTIFIED AT THE SAME TIME AS FIRST-LEVEL USERS. THE SYSTEM STOPS SENDING NOTIFICATIONS IF USERS ARE LOGGED OUT OR IF THE PHONE IS RUNNING OUT OF BATTERY. LOGOUT/RECONNECT NOTIFICATIONS ARE GENERATED ON THE PHONE AND ARE FOUND IN THE HISTORY. NO MALFUNCTION WAS OBSERVED DURING THE TESTS. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE CUSTOMER WAS PROVIDED WITH INFORMATION FROM REVIEW. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: E1; REPORTER INSTITUTION PHONE NUMBER (B)(6). A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21323278&pc=GDW | MW5165696 | 2025-01-29 05:00 | Death | ETHICON ENDO-SURGERY, LLC | 2025-02-06 05:00 | GDW | ETHICON ECHELON FLEX POWERED VASCULAR STAPLER |  | Exsanguination; Vascular Dissection |  |  | 1 | Event Description: AN 80 Y/O FEMALE PATIENT HAD SCHEDULED PROCEDURE FOR LEFT VIDEO ASSISTED THORACIC SURGERY WITH WEDGE RESECTION, POSSIBLE THORACOTOMY WITH LOBECTOMY, LEFT THORACIC NODE DISSECTION. SURGEON NOTED MALFUNCTION WITH STAPLER DEVICE, EXCESSIVE BLEEDING OCCURRED, AND PATIENT EXSANGUINATED IN THE OPERATING ROOM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21323515&pc=DSQ | 2916596-2025-00908 | 2022-10-01 04:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿CLINICAL OUTCOMES OF HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SUPPORT WITH A BRIDGE TO TRANSPLANT VS A DESTINATION THERAPY STRATEGY: A SINGLE-CENTRE RETROSPECTIVE COHORT¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH INFECTION, STROKE, BLEEDING, TRANSPLANTATION, RENAL DYSFUNCTION, AND DEATH. THIS SINGLE-CENTER RETROSPECTIVE STUDY EVALUATED 71 PATIENTS WHO WERE IMPLANTED WITH HM3 BETWEEN NOV2015 AND OCT2022. THE MEDIAN FOLLOW-UP DURATION WAS 18.3 MONTHS. THE STUDY AIMED TO COMPARE THE PROFILE AND CLINICAL OUTCOMES OF PATIENTS SUPPORTED WITH A HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) AT THEIR INSTITUTION, WITH A FOCUS ON WHETHER THE PATIENTS' PREIMPLANTATION STRATEGY WAS BRIDGE TO TRANSPLANT (BTT) OR DESTINATION THERAPY (DT). THE AVERAGE AGE OF PATIENTS WAS 58 YEARS. 63 OF THE 71 PATIENTS WERE MALE, AND 8 WERE FEMALE. OUT OF THE TOTAL 71 EVALUATED PATIENTS, A TOTAL OF 13 PATIENTS PASSED AWAY. 18 PATIENTS RECEIVED TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) SUPPORT (P=0.13); 12 PATIENTS EXPERIENCE RENAL DYSFUNCTION REQUIRING TEMPORARY DIALYSIS(P=0.08). 35 PATIENTS EXPERIENCED BLEEDING (P=0.06), OF WHICH 23 REQUIRED SURGICAL INTERVENTION (P=0.17), AND 14 OF THE BLEEDS WERE SPECIFICALLY GASTROINTESTINAL (GI) (P=0.2), 2 PATIENTS EXPERIENCED POST TRAUMATIC INTRACRANIAL BLEEDING (P=0.08). 8 PATIENTS EXPERIENCED EITHER AN ISCHEMIC STROKE OR A TRANSIENT ISCHEMIC ATTACK (TIA) (P=0.19). 1 PATIENT DIED DUE TO THE STROKE, AND 1 ADDITIONAL PATIENT EXPERIENCED A PERMANENT DISABILITY DUE TO THE STROKE. 50 PATIENTS EXPERIENCED SOME SORT OF VENTRICULAR ASSIST DEVICE (VAD) RELATED INFECTION (P=0.45), OF WHICH 48 OF THE PATIENTS EXPERIENCED SOME SORT OF DRIVELINE CULTURE INFECTION, 10 EXPERIENCED A DRIVELINE SURGICAL INFECTION, AND 2 EXPERIENCED A PUMP INFECTION. 4 MORE PATIENTS HAD VAD RELATED MEDIASTINITIS INFECTION (P=0.31). 29 PATIENTS EXPERIENCED A HEART TRANSPLANT, AND 1 FINAL PATIENT HAD THEIR LVAD EXPLANTED AS PART OF THEIR RECOVERY FROM PERIPARTUM CARDIOMYOPATHY. THE STUDY CONCLUDED THAT ALTHOUGH HM3 LVAD DEMONSTRATED FAVORABLE SURVIVAL IN PATIENTS WITH ADVANCED HEART FAILURE, PATIENTS IN THE DT GROUP HAD A POORER PROGNOSIS COMPARED TO THOSE IN THE BTT GROUP. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. UNIVERSITY HOSPITAL OF LAUSANNE, SWITZERLAND. KIKOÏNE, J., NOWACKA, A., SCHUKRAFT, S., ABDURASHIDOVA, T., YERLY, P., TOZZI, P., LTAIEF, Z., ROSNER, L., HULLIN, R., & KIRSCH, M. (2024). CLINICAL OUTCOMES OF HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SUPPORT WITH A BRIDGE TO TRANSPLANT VS A DESTINATION THERAPY STRATEGY: A SINGLE-CENTRE RETROSPECTIVE COHORT. SWISS MEDICAL WEEKLY, 154, 3529. HTTPS://DOI.ORG/10.57187/S.3529. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21323541&pc=BZD | 2518422-2025-101814 | 2022-12-23 05:00 | Death | RESPIRONICS, INC. | 2025-02-06 05:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S HUSBAND CONTACTED US WANTING TO RETURN THE DEVICE AS HIS WIFE PASSED AWAY 2 YEARS AGO. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21323555&pc=DSQ | 3007042319-2025-00468 | 2024-12-20 05:00 | Death | HEARTWARE, INC. | 2025-02-06 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT, WHO HAD A VENTRICULAR ASSIST DEVICE IMPLANT, DIED OF MULTIPLE ORGAN FAILURE. NO ISSUES WITH THE VENTRICULAR ASSIST DEVICE SYSTEM WERE REPORTED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING MULTIPLE ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, MULTI-ORGAN FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ###INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ### MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21323994&pc=DSQ | 2916596-2025-00279 | 2025-01-06 05:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Tachycardia; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO THERAPY RESISTANT VENTRICULAR TACHYCARDIA'S (VT'S). THE PATIENT HAD A HISTORY OF ARRHYTHMIA PRE-LVAD. THE OUTCOME AND ARRHYTHMIA WERE NOT CONSIDERED DEVICE OR THERAPY RELATED.Event Description: THE PATIENT HAD A HISTORY OF VENTRICULAR TACHYCARDIA'S (VT'S) PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD). Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THERE WERE NO DEVICE-RELATED ISSUES WITH (B)(6), REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. IT WAS DETERMINED THAT THE REPORTED INFORMATION DID NOT MEET THE REQUIREMENTS OF A COMPLAINT. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21324315&pc=BZD | 2518422-2025-101815 | 2024-12-27 05:00 | Death | RESPIRONICS, INC. | 2025-02-06 05:00 | BZD | BIPAP SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A BIPAP AUTIO BIFLEX. THE PATIENT'S SON IN LAW CONTACTED US NEEDING TO RETURN. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21324483&pc=DSQ | 2916596-2025-00471 | 2024-07-01 04:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿ASPIRIN AND HEMOCOMPATIBILITY AFTER LVAD IMPLANTATION IN PATIENTS WITH ATHEROSCLEROTIC VASCULAR DISEASE: A SECONDARY ANALYSIS FROM THE ARIES-HM3 RANDOMIZED CLINICAL TRIAL¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE RELATED TO HEMOCOMPATIBILITY RELATED ADVERSE EVENTS AND DEATH. IN THIS TRIAL, IT WAS EXPLORED WHETHER COMMON CONDITIONS REQUIRING ASPIRIN (PRIOR PERCUTANEOUS CORONARY INTERVENTION [PCI], CORONARY ARTERY BYPASS GRAFTING [CABG], STROKE, OR PVD) IN RECIPIENTS OF THE HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) WOULD AFFECT OUTCOMES DIFFERENTIALLY WHEN ASPIRIN WAS ELIMINATED. THE STUDY INCLUDED 589 PATIENTS WITH A MEAN AGE OF 57.1 YEARS, AND 133 WERE FEMALES. THE PRIMARY END POINT CONSISTED OF A COMPOSITE OF SURVIVAL FREE OF NON-SURGICAL MAJOR HEMOCOMPATIBILITY-RELATED ADVERSE EVENTS (HRAES) (I.E., STROKE, PUMP THROMBOSIS, MAJOR NON-SURGICAL BLEEDING, AND ARTERIAL PERIPHERAL THROMBOEMBOLISM) AT 12 MONTHS. THROMBOTIC EVENTS WERE RARE, WITH NO DIFFERENCES BETWEEN ASPIRIN AND PLACEBO IN PATIENTS WITH AND WITHOUT VASCULAR DISEASE (P FOR INTERACTION = 0.77). ASPIRIN TREATMENT WAS ASSOCIATED WITH A HIGHER RATE OF NONSURGICAL MAJOR BLEEDING EVENTS IN THE GROUP WITH PRIOR VASCULAR CONDITION HISTORY COMPARED WITH THOSE WITHOUT ASPIRIN (RATE RATIO FOR PLACEBO COMPARED WITH ASPIRIN, 0.52; 95% CI,0.35-0.79). THE SURVIVAL FREE FROM DEATH AND NON-SURGICAL HRAE ANALYSIS REPORTED AS SIGNIFICANT REDUCTION IN HAZARD RATIO (HR) AT 24 MONTHS BETWEEN ASPIRIN AND PLACEBO FOR PATIENTS WITH PRIOR INDICATION FOR ASPIRIN (HR,0.63; 95% CI, 0.41-0.98; P=0.04) BUT NOT IN PATIENTS WITHOUT SUCH INDICATION (HR,0.79; 95% CI, 0.54 1.15; P=0.22). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC DEVICE AND PATIENT INFORMATION ARE UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JUL2024 SINCE THE DATE OF DATA ANALYSIS WAS CONDUCTED FROM APRIL TO JULY 2024. AUTHOR INFORMATION: FINN GUSTAFSSON ET AL., ASAIO JOURNAL70.11: E150-E152. DOI: 10.1097/MAT.0000000000002201. RIGSHOSPITALET, UNIVERSITY OF COPENHAGEN, COPENHAGEN, DENMARK (GUSTAFSSON). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21324603&pc=DTK | 1820334-2025-00117 |  | Death | COOK INC | 2025-02-06 05:00 | DTK | GUNTHER TULIP FEMORAL VENA CAVA FILTER SET | Inadequate Filtration Process; Device Tipped Over; Detachment of Device or Device Component; Unintended Movement | Pulmonary Embolism; Abdominal Pain; Chest Pain; Hemorrhage/Bleeding; Pain; Perforation of Vessels |  |  | 1 | Event Description: ADDITIONAL INFORMATION WAS PROVIDED WHICH INDICATES THE PATIENT HAS SUBSEQUENTLY EXPIRED AND ALLEGATION OF WRONGFUL DEATH AGAINST THE PRODUCT HAS BEEN RECEIVED AT THIS TIME. THIS IS A FOLLOW-UP REPORT FOR THE INITIAL REPORT SUBMITTED WITH REFERENCE NUMBER 1820334-2019-00222 ON 31JAN2019. THE SUBSEQUENT FOLLOW-UP DATES INCLUDE: FOLLOW-UP #1 - SENT ON 13MAR2019, FOLLOW-UP #2 - SENT ON 18APR2019, FOLLOW-UP #3 - SENT ON 19JUL2019, FOLLOW-UP #4 - SENT ON 30OCT2019, FOLLOW-UP #5 - SENT ON 08AUG2022. THIS REPORT IS FOLLOW-UP #6. Manufacturer Narrative: THIS IS A FOLLOW-UP REPORT FOR THE INITIAL REPORT SUBMITTED WITH REFERENCE NUMBER 1820334-2019-00222 ON 31JAN2019. THE SUBSEQUENT FOLLOW-UP DATES INCLUDE: FOLLOW-UP #1 - SENT ON 13MAR2019, FOLLOW-UP #2 - SENT ON 18APR2019, FOLLOW-UP #3 - SENT ON 19JUL2019, FOLLOW-UP #4 - SENT ON 30OCT2019, FOLLOW-UP #5 - SENT ON 08AUG2022. THIS REPORT IS FOLLOW-UP #6. B2 (DATE OF DEATH): THIS IS THE DATE THAT COOK WAS NOTIFIED OF THE PATIENT¿S DEATH. HOWEVER, THE EXACT DATE OF THE PATIENT¿S DEATH IS UNKNOWN AT THIS TIME. CORRECTED INFORMATION: D4 (MODEL #). ADDITIONAL INFORMATION: B1, B2, B5, H1 AND H6. INVESTIGATION: INVESTIGATION IS REOPENED DUE TO ADDITIONAL INFORMATION PROVIDED. THE REPORTED ALLEGATIONS HAVE BEEN FURTHER INVESTIGATED BASED ON THE INFORMATION PROVIDED TO DATE. THE FOLLOWING ALLEGATIONS HAVE BEEN INVESTIGATED: DEATH. UNKNOWN IF THE REPORTED DEATH IS DIRECTLY RELATED TO THE FILTER AND UNABLE TO IDENTIFY A CORRESPONDING FAILURE MODE AT THIS POINT IN TIME. (B)(4) DEVICES IN LOT. NO RELEVANT NOTES ON WORK ORDER FOR DEVICE. NO OTHER COMPLAINTS ON LOT. PRODUCT IS MANUFACTURED AND INSPECTED ACCORDING TO SPECIFICATIONS. NO EVIDENCE TO SUGGEST THAT THIS DEVICE WAS NOT MANUFACTURED ACCORDING TO SPECIFICATIONS AND NOTHING INDICATES THAT THE FILTER DID NOT PERFORM AS INTENDED, E.G. INTENDED FOR THE PREVENTION OF RECURRENT PULMONARY EMBOLISM (PE) VIA PLACEMENT IN THE VENA CAVA. COOK WILL REOPEN ITS INVESTIGATION IF FURTHER INFORMATION IS RECEIVED WARRANTING SUPPLEMENTATION IN ACCORDANCE WITH 21 C.F.R. 803.56. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW-UP MEDWATCH REPORT WILL BE SUBMITTED IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED, THAT A DEATH OR SERIOUS INJURY OCCURRED, OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21325059&pc=QBJ | 3004753838-2025-031952 | 2024-12-27 05:00 | Death | DEXCOM, INC. | 2025-02-06 05:00 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED BY A HEALTHCARE PROVIDER VIA MAUDE (MW5164151) AND BY THE FOOD AND DRUG ADMINISTRATION THAT A PATIENT PASSED AWAY ON (B)(6) 2024. THERE IS REPORTER INFORMATION TO GATHER ADDITIONAL INFORMATION AS THE HCP DID NOT WISH TO FOLLOW UP AND THERE IS NO PATIENT INFORMATION AVAILABLE. PER THE MAUDE REPORT, THE PATIENT ALSO USES A FREESTYLE LIBRE 3 SENSOR. NO DATA OR PRODUCT WAS PROVIDED FOR INVESTIGATION. THE ALLEGATION WAS UNDETERMINED. THE PROBABLE CAUSE COULD NOT BE DETERMINED.Event Description: THERE IS NO REPORTER INFORMATION TO GATHER ADDITIONAL INFORMATION AS THE HCP DID NOT WISH TO FOLLOW UP AND THERE IS NO PATIENT INFORMATION AVAILABLE. Manufacturer Narrative: (B)(4). B5 DESCRIBE EVENT OR PROBLEM - CORRECTION TO THE FOLLOWING STATEMENT IN THE INITIAL MDR, "THERE IS REPORTER INFORMATION TO GATHER ADDITIONAL INFORMATION AS THE HCP DID NOT WISH TO FOLLOW UP AND THERE IS NO PATIENT INFORMATION AVAILABLE." H2 CORRECTION. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21325159&pc=DSQ | 2916596-2025-00685 | 2024-10-31 04:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THERE WERE NO DEVICE ISSUES, AND THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. THE CAUSE OF DEATH IS UNKNOWN, BUT NOT THOUGHT TO BE DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21326454&pc=DSQ | 2916596-2025-00750 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. THE PUMP OPERATED AS EXPECTED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21326740&pc=CBF | MW5165704 |  | Death | SECHRIST INDUSTRIES, INC. | 2025-02-06 05:00 | CBF | CHAMBER, HYPERBARIC |  | Insufficient Information |  |  | 1 | Event Description: THE COMPLAINANT ALLEGES THAT A CHILD WAS KILLED AS THE RESULT OF A HYPERBARIC CHAMBER FIRE. I DO NOT KNOW THE SPECIFIC DEVICE INVOLVED IN THE ADVERSE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21326830&pc=FNM | 0001831750-2025-00329 | 2023-11-30 05:00 | Death | STRYKER MEDICAL-KALAMAZOO | 2025-02-06 05:00 | FNM | ISOAIR DOM 84 IN TYPE B PLUG | Adverse Event Without Identified Device or Use Problem | Pressure Sores |  |  | 1 | Event Description: IT WAS REPORTED THROUGH A CLINICAL SURVEY THAT A 67-YEAR-OLD, UNDERWEIGHT, FEMALE PATIENT WAS EMERGENTLY ADMITTED TO THE INTENSIVE CARE UNIT WITH ¿SEPSIS FROM PRESSURE ULCERS.¿ THE PATIENT PRESENTED WITH A BRADEN SCORE OF 11 AND 4 PRESSURE INJURIES AT THE ONSET OF TREATMENT. ONE STAGE 3 AND ONE STAGE 4 FOR PRESSURE INJURY INVOLVING THE HEEL AREA. ADDITIONALLY, THE PATIENT PRESENTED WITH TWO UNSTAGEABLE PRESSURE INJURIES INVOLVING THE SACRAL AND TROCHANTERIC AREAS. DURING THE TREATMENT PERIOD, THE TWO PRESSURE INJURIES INVOLVING THE HEELS IMPROVED BY 40% REDUCTION IN VOLUME. HOWEVER, TWO NEW STAGE 2 PRESENTING IN THE ELBOW AREA, REPORTEDLY CONTRIBUTING TO A DETERIORATION IN THE PATIENT¿S HEALTH, RESULTING IN PROLONGED HOSPITALIZATION AND ULTIMATELY, DEATH. THE SURVEY RESPONDENT NOTED THAT THERE WAS NO DEVICE MALFUNCTION, AND THE PRODUCT WAS BEING USED IN ACCORDANCE WITH THE INDICATIONS FOR USE. SEVERAL COMORBID CONDITIONS AND MEDICATIONS ARE NOTED WHICH MAY IMPEDE WOUND HEALING. GIVEN THE PREVIOUS INFORMATION AND THE MULTI-FACTORIAL NATURE OF PRESSURE INJURY DEVELOPMENT IT CANNOT BE CONCLUSIVELY DETERMINED THAT THE PRODUCT HAD A CAUSAL RELATIONSHIP OR CONTRIBUTION TO THE PATIENT¿S OUTCOME. NONETHELESS, WE ARE CHOOSING TO FILE THIS EVENT OUT OF AN ABUNDANCE OF CAUTION. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL ATTEMPTS ARE BEING MADE TO OBTAIN MORE INFORMATION.Event Description: NO NEW INFORMATION. Manufacturer Narrative: MULTIPLE ATTEMPTS HAVE BEEN MADE TO GATHER FURTHER INFORMATION, WITH NO RESPONSE RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21326927&pc=GDW | 3005075853-2025-01040 | 2025-01-29 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-06 05:00 | GDW | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Failure to Fire; Difficult to Open or Close; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K141952 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE WHERE THE STAPLER WAS STUCK ON THE PULMONARY ARTERY AND COULD NOT BE OPENED. REP WALKED THE SURGEON THROUGH THE STEPS OVER THE PHONE BUT THEY WERE STILL HAVING TROUBLE. SALES REP THEN DROVE TO THE FACILITY WHERE HE WAS TOLD THE PATIENT PASSED PRIOR TO HIS ARRIVAL. SURGEON STATED TO THE REP THAT HE JUST COULDN'T OPEN IT AND THE PATIENT BLED OUT. SURGEON DESCRIBED THE SITUATION TO THE REP THAT HE FIRED THE DEVICE AND IT SOUNDED LIKE IT FIRED HALFWAY. SURGEON TRIED TO OPEN BUT COULDN'T OPEN THE DEVICE. REP ASKED THE SURGEON IF HE USED THE REVERSE BUTTON, SURGEON REPLIED THAT HE DID. SURGEON SAID HE ENGAGED THE MANUAL OVERRIDE HOWEVER HE WAS OPERATING ALONE. REP INFORMED THE SURGEON THAT WHEN ENGAGING THE MANUAL OVERRIDE SOMEONE NEEDS TO HOLD ON TO THE SHAFT TO ENGAGE BECAUSE ITS A PRETTY HARD RATCHET TO RATCHET THE KNIFE BLADE BACK. SURGEON SAID HE TRIED THE MANUAL OVERRIDE BUT WAS UNABLE TO GET IT BACK. THERE WAS A LOT OF BLEEDING. A SECOND STAPLER WAS USED TO STAPLE UNDERNEATH WHEN THEY COULDN'T OPEN THE FIRST STAPLER. IT WAS DESCRIBED BY THE SURGICAL TECH THAT AFTER THE SECOND STAPLING THERE WAS BLOOD EVERYWHERE AND THE PASSED VERY QUICKLY AFTER THAT. THE SURGEON WALKED THE REP TO THE OR WHERE HE WAS ABLE TO TAKE A LOOK AT THE STAPLER, THE MANUAL OVERRIDE HAD BEEN ENGAGED BECAUSE THE CAP WAS OFF AND THE WRENCH WAS UP. Manufacturer Narrative: (B)(4). DATE SENT 2/6/2025. D4 BATCH # UNK. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WAS THE INDICATION FOR THE PROCEDURE? WHAT VESSEL WAS THE DEVICE BEING FIRED ON WHEN THE ISSUE OCCURRED, MAIN PA OR PV? WAS IT A LOBECTOMY OR PNEUMONECTOMY? WAS THERE DIFFICULTY CLOSING THE DEVICE? WAS THERE ANY UNUSUAL SOUND HEARD? HOW FAR IS IT BELIEVED THAT THE DEVICE MAY HAVE CUT AND/OR STAPLED? WHY IS IT BELIEVED THAT A DIFFICULT TO OPEN ISSUE MAY HAVE LED TO VESSEL BLEEDING? WHAT EXACT TROUBLESHOOTING STEPS DID THE TEAM TAKE TO OPEN THE DEVICE? WHEN ATTEMPTING TO OPEN, DID THE SURGEON SQUEEZE THE CLOSING TRIGGER, THEN SIMULTANEOUSLY PRESS THE ANVIL RELEASE BUTTON ON EITHER SIDE OF THE INSTRUMENT? WHILE PRESSURE IS STILL ON THE ANVIL RELEASE BUTTON, SLOWLY RELEASE THE CLOSING TRIGGER? WHEN USING THE MANUAL OVERRIDE, DID THE SURGEON RATCHET THE MANUAL OVERRIDE BACK AND FORTH MULTIPLE TIMES IN AN ATTEMPT TO RETURN THE KNIFE BLADE TO THE HOME POSITION? DID THE SURGEON ATTEMPT TO FORCE THE CLOSING TRIGGER UPWARD (AWAY FROM THE HANDLE) UNTIL THE CLOSING TRIGGER RETURNS TO ITS ORIGINAL POSITION CAN ENABLE THE OPENING OF THE JAWS? IS IT POSSIBLE THAT THERE WERE CLIPS NEAR THE AREA WHERE THE DEVICE WAS FIRED? WHAT IS THE SURGEON¿S EXPERIENCE WITH USING PVS DEVICE? IS THE LOT AND/OR BATCH NUMBER OF THE CARTRIDGE USED WITH THE DEVICE AVAILABLE? IS THE SURGEON INTERESTED IN SPEAKING ETHICON MEDICAL PERSONNEL? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 2/19/2025. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE INDICATION FOR THE PROCEDURE? LOBECTOMY. WHAT VESSEL WAS THE DEVICE BEING FIRED ON WHEN THE ISSUE OCCURRED, MAIN PA OR PV? PA. WAS IT A LOBECTOMY OR PNEUMONECTOMY? LOBECTOMY. WAS THERE DIFFICULTY CLOSING THE DEVICE? UNKNOWN. WAS THERE ANY UNUSUAL SOUND HEARD? UNKNOWN. HOW FAR IS IT BELIEVED THAT THE DEVICE MAY HAVE CUT AND/OR STAPLED? HALF WAY WHY IS IT BELIEVED THAT A DIFFICULT TO OPEN ISSUE MAY HAVE LED TO VESSEL BLEEDING? YES. WHAT EXACT TROUBLESHOOTING STEPS DID THE TEAM TAKE TO OPEN THE DEVICE? ENGAGED MANUAL OVERRIDE. SURGEON OPERATED ALONE WITHOUT PA, FA OR RESIDENT. WHEN ATTEMPTING TO OPEN, DID THE SURGEON SQUEEZE THE CLOSING TRIGGER, THEN SIMULTANEOUSLY PRESS THE ANVIL RELEASE BUTTON ON EITHER SIDE OF THE INSTRUMENT? UNKNOWN. WHILE PRESSURE IS STILL ON THE ANVIL RELEASE BUTTON, SLOWLY RELEASE THE CLOSING TRIGGER? UNKNOWN. WHEN USING THE MANUAL OVERRIDE, DID THE SURGEON RATCHET THE MANUAL OVERRIDE BACK AND FORTH MULTIPLE TIMES IN AN ATTEMPT TO RETURN THE KNIFE BLADE TO THE HOME POSITION? UNKNOWN. DID THE SURGEON ATTEMPT TO FORCE THE CLOSING TRIGGER UPWARD (AWAY FROM THE HANDLE) UNTIL THE CLOSING TRIGGER RETURNS TO ITS ORIGINAL POSITION CAN ENABLE THE OPENING OF THE JAWS? UNKNOWN. IS IT POSSIBLE THAT THERE WERE CLIPS NEAR THE AREA WHERE THE DEVICE WAS FIRED? YES. WHAT IS THE SURGEON¿S EXPERIENCE WITH USING PVS DEVICE? SINCE LAUNCH. IS THE LOT AND/OR BATCH NUMBER OF THE CARTRIDGE USED WITH THE DEVICE AVAILABLE? UNKNOWN. IS THE SURGEON INTERESTED IN SPEAKING ETHICON MEDICAL PERSONNEL? MAYBE. Manufacturer Narrative: (B)(4). DATE SENT: 4/8/2025. D4: BATCH # 874C51. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON ENDO-SURGERY FOR EVALUATION. VISUAL INSPECTION WAS CONDUCTED ON THE RETURNED DEVICE. VISUAL ANALYSIS OF THE RETURNED SAMPLE DETERMINED THAT ONE PVE35A DEVICE WAS RETURNED WITH NO APPARENT DAMAGE, WITH THE CLOSING TRIGGER AND ANVIL IN CLOSED POSITION, AND WITH A RELOAD LOADED ON THE DEVICE. THE RELOAD WAS RECEIVED PARTIALLY FIRED 1/10. THE CONDITION IDENTIFIED DURING THE INVESTIGATION OF THE RELOAD RECEIVED HAS BEEN CORRELATED TO THE MANUFACTURING PROCESS. THE MANUAL OVERRIDE DOOR OUT OF POSITION AND MISSING; THE OVERRIDE LEVER WAS UP WHICH DENOTES THAT THE KNIFE WAS MANUALLY RETURNED TO HOME POSITION. IT SHOULD BE NOTED AFTER THE MANUAL OVERRIDE SYSTEM IS USED THE INSTRUMENT IS DISABLED AND CANNOT BE USED FOR ANY SUBSEQUENCE FIRINGS. IN ADDITION, THE ARTICULATION MECHANISM WAS NOTED TO BE DAMAGED AS WOULD NOT HOLD THE ARTICULATION SETTING. THE CT SCAN ANALYSIS SHOWED THE DISTAL CHANNEL RETAINER WAS NOTED TO BE DAMAGED. THE DAMAGE TO THE DISTAL CHANNEL RETAINER IS POTENTIALLY RELATED TO A MANUFACTURING PROCESS. BASED ON THE INFORMATION CURRENTLY AVAILABLE, A PRODUCT ISSUE WAS IDENTIFIED DURING THE INVESTIGATION OF THE SAMPLE RECEIVED. THIS PRODUCT ISSUE WILL BE ADDRESSED THROUGH ETHICON ENDO SURGERY¿S QUALITY SYSTEM. IN ORDER TO VERIFY THE CONDITION OF THE INTERNAL COMPONENTS, THE DEVICE WAS DISASSEMBLED. UPON DISASSEMBLING, IT WAS OBSERVED THAT THE SLIDE LEVER IS PAST THE FULLY FIRED POSITION AND THE TAB THAT RIDES IN THE DRIVE BAR SLOT IS BROKEN. THIS DAMAGED NOTED WOULD NOT HAVE PREVENTED THE INSTRUMENT FROM OPENING, HOWEVER NO FUNCTIONAL TEST WAS PERFORMED DUE THE CONDITION OF THE DEVICE. THE DAMAGE TO THE TAB WAS MOST LIKELY CAUSED BY APPLYING A VERY LARGE FORCE TO THE REVERSE BUTTON SWITCH WHEN ATTEMPTING TO OPEN THE INSTRUMENT AFTER THE DRIVE BAR HAD ALREADY RETURNED TO THE HOME POSITION. PLEASE REFERENCE THE INSTRUCTION FOR USE FOR MORE INFORMATION. THE EVENT DESCRIBED OF DIFFICULT TO OPEN COULD NOT BE CONFIRMED AS THE DEVICE OPENED AND CLOSED WITHOUT ANY DIFFICULTIES NOTED. THE INSTRUMENT UNCLAMPED NORMALLY WHEN THE UNCLAMP LOCKOUT WAS PRESSED WHILE SQUEEZING THE CLAMPING TRIGGER. ALTHOUGH NO CONCLUSION COULD BE REACH ON THE CAUSE OF THE REPORTED EVENT OF HEMOSTASIS INTERVENTION, THE INSTRUCTIONS FOR USE CONTAIN THE FOLLOWING CAUTION: TISSUE THICKNESS SHOULD BE CAREFULLY EVALUATED BEFORE FIRING ANY STAPLER. HOLDING THE JAWS IN PLACE FOR 15 SECONDS AFTER CLOSING AND PRIOR TO FIRING MAY RESULT IN BETTER COMPRESSION AND STAPLE FORMATION. WHEN POSITIONING THE STAPLER ON THE APPLICATION SITE, ENSURE THAT NO OBSTRUCTIONS SUCH AS CLIPS, STENTS, GUIDE WIRES, ETC. ARE WITHIN THE INSTRUMENT JAWS. DEVICE HISTORY REVIEW A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER 874C51, AND NO NON-CONFORMANCES WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: OPERATION WAS BUMPED BACK AND WAS SUPPOSED TO HAVE ANOTHER SURGEON. REMOVAL OF CANCER IN THE UPPER LUNG AND THE TISSUE WAS FRAIL. STAPLE LINE FAILED. JUST BEFORE LAST STAPLE, TECHS WERE CHANGED AND DID THE STAPLER BEFORE THEY LEFT THE ROOM. PUT STAPLER ON THE LARGEST ARTERY GOING TO THE UPPER LOBE. CLAMPED DOWN AND STAPLE LINE WAS FINE. FIRED DOWN AND JAMMED, REVERSE STAPLER WOULD NOT WORK. OPENED CLAMP, GOT ANOTHER STAPLER AND CLAMPED AND WAS FIRED, MASSIVE BLEED OCCURRED. STARTED TO PACK AND EXTENDED INCISION. RE-STAPLED THE OTHER PART OF LUNG AND PATIENT PASSED. DISTAL TOP WAS ON A MAJOR ARTERY AND THERE WAS A FEAR OF TEARING MAIN PA. IT WAS HARD TO SQUEEZE THE HANDLE. WHEN THE STAPLER WAS CLOSED, WAS THERE ANY BLEEDING? NO. PROXIMAL WITH ANOTHER PVS AND FIRED, WAS THE FIRING OKAY? YES. WHEN RELEASED, IS THAT WHEN THE MASSIVE BLEED OCCURRED? YES. WAS THE THUMB BUTTON PUSHED? YES. WHAT ATTEMPTS TO OPEN DEVICE? THUMB MULTIPLE TIMES. WAS THE MANUAL OVERRIDE USED? YES. FROM THE OPENING, DID YOU FIRST SQUEEZE THE CLOSING TRIGGER AND THEN PRESS THE ANVIL RELEASE BUTTON? YES, FIRST THING TRIED. Manufacturer Narrative: (B)(4). DATE SENT: 5/12/2025. D4: BATCH # 874C51. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON ENDO-SURGERY FOR EVALUATION. VISUAL INSPECTION WAS CONDUCTED ON THE RETURNED DEVICE. VISUAL ANALYSIS OF THE RETURNED SAMPLE DETERMINED THAT ONE PVE35A DEVICE WAS RECEIVED WITH NO APPARENT DAMAGE, WITH THE CLOSING TRIGGER AND ANVIL IN THE CLOSED POSITION, AND WITH A RELOAD LOADED ON THE DEVICE. THE RELOAD WAS RECEIVED PARTIALLY FIRED 1/10TH IN THE LOCK-OUT POSITION WITH THE REMAINING STAPLES INSIDE THE RELOAD POCKETS. IT CANNOT BE CONFIRMED WHY THE RELOAD LOCKED OUT; HOWEVER, IT IS POSSIBLE THAT THE ONE-PIECE SLED COMPONENT WITHIN THE RELOAD MAY HAVE BEEN MOVED OUT OF POSITION DURING THE LOADING PROCESS LEADING TO THE INADVERTENT LOCKOUT. PLEASE REFER TO THE INSTRUCTIONS FOR USE FOR MORE INFORMATION. THE MANUAL OVERRIDE DOOR WAS NOT PRESENT ON THE DEVICE AND WAS NOT RETURNED FOR ANALYSIS; THE OVERRIDE LEVER WAS UP WHICH DENOTES THAT THE KNIFE WAS MANUALLY RETURNED TO HOME POSITION. IT SHOULD BE NOTED AFTER THE MANUAL OVERRIDE SYSTEM IS USED THE INSTRUMENT IS DISABLED AND CANNOT BE USED FOR ANY SUBSEQUENCE FIRINGS. THE DEVICE WAS CT SCANNED, AND THE ARTICULATION LOCK TOOTH WAS NOTED TO BE BROKEN. IT IS POSSIBLE THAT AN OUTSIDE THE NORMAL USE FORCE WAS APPLIED ON THE DISTAL JAW CREATING A TORQUE ON THE ARTICULATION COMPONENTS AND BREAKING THE ARTICULATION BAR. THIS CONDITION PREVENTS THE DEVICE FROM HOLDING THE ARTICULATION SETTING AND DOES NOT AFFECT THE OPENING MECHANISM. AN ATTEMPT WAS MADE TO OPEN THE INSTRUMENT, AND IT OPENED WITHOUT DIFFICULTY WHEN THE ANVIL RELEASE BUTTON WAS ACTUATED WHILE SQUEEZING THE CLOSING TRIGGER. PLEASE NOTE THAT TO OPEN THE JAWS, SQUEEZE THE CLOSING TRIGGER, THEN SIMULTANEOUSLY PRESS THE ANVIL RELEASE BUTTON ON EITHER SIDE OF THE INSTRUMENT. WHILE PRESSURE IS STILL ON THE ANVIL RELEASE BUTTON, SLOWLY RELEASE THE CLOSING TRIGGER. FOR FURTHER INSTRUCTIONS PLEASE REFER TO THE INSTRUCTIONS FOR USE. TO VERIFY THE CONDITION OF THE INTERNAL COMPONENTS, THE DEVICE WAS DISASSEMBLED, AND THE SLIDE LEVER WAS NOTED TO BE PAST THE FULLY FIRED POSITION AND THE TAB THAT RIDES IN THE DRIVE BAR SLOT WAS BROKEN. THE DAMAGE TO THE TAB WAS MOST LIKELY CAUSED BY APPLYING AN EXCESS OF FORCE TO THE REVERSE BUTTON SWITCH WHEN ATTEMPTING TO OPEN THE INSTRUMENT AFTER THE DRIVE BAR HAD ALREADY RETURNED TO THE HOME POSITION. THIS DAMAGE FOUND WOULD NOT HAVE PREVENTED THE INSTRUMENT FROM OPENING. NO FUNCTIONAL TEST WAS PERFORMED DUE TO THE CONDITION OF THE DEVICE. THE EVENT DESCRIBED OF DIFFICULT TO OPEN COULD NOT BE CONFIRMED AS THE DEVICE OPENED AND CLOSED WITHOUT ANY DIFFICULTIES NOTED. THE INSTRUMENT OPENED NORMALLY WHEN THE UNCLAMP LOCKOUT WAS PRESSED WHILE SQUEEZING THE CLAMPING TRIGGER. DEVICE HISTORY REVIEW A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER 874C51, AND NO NON-CONFORMANCES WERE IDENTIFIED. Manufacturer Narrative: (B)(4) DATE SENT: 5/12/2025. H9: 1527736-04/22/2025-001-C. Manufacturer Narrative: (B)(4). DATE SENT: 7/21/2025. D4: BATCH # 874C51. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON ENDO-SURGERY FOR EVALUATION. VISUAL INSPECTION WAS CONDUCTED ON THE RETURNED DEVICE. VISUAL ANALYSIS OF THE RETURNED SAMPLE DETERMINED THAT ONE PVE35A DEVICE WAS RECEIVED WITH NO APPARENT DAMAGE, WITH THE CLOSING TRIGGER AND ANVIL IN THE CLOSED POSITION, AND WITH A RELOAD LOADED ON THE DEVICE. THE RELOAD WAS RECEIVED PARTIALLY FIRED 1/10TH IN THE LOCK-OUT POSITION WITH THE REMAINING STAPLES INSIDE THE RELOAD POCKETS. IT CANNOT BE CONFIRMED WHY THE RELOAD LOCKED OUT; HOWEVER, IT IS POSSIBLE THAT THE ONE-PIECE SLED COMPONENT WITHIN THE RELOAD MAY HAVE BEEN MOVED OUT OF POSITION DURING THE LOADING PROCESS LEADING TO THE INADVERTENT LOCKOUT. PLEASE REFER TO THE INSTRUCTIONS FOR USE FOR MORE INFORMATION. THE MANUAL OVERRIDE DOOR WAS NOT PRESENT ON THE DEVICE AND WAS NOT RETURNED FOR ANALYSIS; THE OVERRIDE LEVER WAS UP WHICH DENOTES THAT THE KNIFE WAS MANUALLY RETURNED TO HOME POSITION. IT SHOULD BE NOTED AFTER THE MANUAL OVERRIDE SYSTEM IS USED THE INSTRUMENT IS DISABLED AND CANNOT BE USED FOR ANY SUBSEQUENCE FIRINGS. THE DEVICE WAS CT SCANNED, AND THE ARTICULATION LOCK TOOTH WAS NOTED TO BE BROKEN. IT IS POSSIBLE THAT AN OUTSIDE THE NORMAL USE FORCE WAS APPLIED ON THE DISTAL JAW CREATING A TORQUE ON THE ARTICULATION COMPONENTS AND BREAKING THE ARTICULATION BAR. THIS CONDITION PREVENTS THE DEVICE FROM HOLDING THE ARTICULATION SETTING AND DOES NOT AFFECT THE OPENING MECHANISM. AN ATTEMPT WAS MADE TO OPEN THE INSTRUMENT, AND IT OPENED WITHOUT DIFFICULTY WHEN THE ANVIL RELEASE BUTTON WAS ACTUATED WHILE SQUEEZING THE CLOSING TRIGGER. PLEASE NOTE THAT TO OPEN THE JAWS, SQUEEZE THE CLOSING TRIGGER, THEN SIMULTANEOUSLY PRESS THE ANVIL RELEASE BUTTON ON EITHER SIDE OF THE INSTRUMENT. WHILE PRESSURE IS STILL ON THE ANVIL RELEASE BUTTON, SLOWLY RELEASE THE CLOSING TRIGGER. FOR FURTHER INSTRUCTIONS PLEASE REFER TO THE INSTRUCTIONS FOR USE. TO VERIFY THE CONDITION OF THE INTERNAL COMPONENTS, THE DEVICE WAS DISASSEMBLED, AND THE SLIDE LEVER WAS NOTED TO BE PAST THE FULLY FIRED POSITION AND THE TAB THAT RIDES IN THE DRIVE BAR SLOT WAS BROKEN. THE DAMAGE TO THE TAB WAS MOST LIKELY CAUSED BY APPLYING AN EXCESS OF FORCE TO THE REVERSE BUTTON SWITCH WHEN ATTEMPTING TO OPEN THE INSTRUMENT AFTER THE DRIVE BAR HAD ALREADY RETURNED TO THE HOME POSITION. THIS DAMAGE FOUND WOULD NOT HAVE PREVENTED THE INSTRUMENT FROM OPENING. NO FUNCTIONAL TEST WAS PERFORMED DUE TO THE CONDITION OF THE DEVICE. THE EVENT DESCRIBED OF DIFFICULT TO OPEN COULD NOT BE CONFIRMED AS THE DEVICE OPENED AND CLOSED WITHOUT ANY DIFFICULTIES NOTED. THE INSTRUMENT OPENED NORMALLY WHEN THE UNCLAMP LOCKOUT WAS PRESSED WHILE SQUEEZING THE CLAMPING TRIGGER. DEVICE HISTORY REVIEW: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER 874C51, AND NO NON-CONFORMANCES WERE IDENTIFIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21326983&pc=DSQ | 2916596-2025-00520 | 2025-01-08 05:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiovascular Insufficiency; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT WAS FOUND UNRESPONSIVE IN THEIR CHAIR. THERE WERE NOT ANY ALARMS. THE PATIENT HAD A MILD COLD AND HAD RECENT DENTAL WORK DONE BUT WAS NOT OVERTLY SICK. DEATH CERTIFICATE AND CORONER'S OFFICE STATED ATHEROSCLEROTIC CARDIOVASCULAR DISEASE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE WAS OPERATING AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, OF THIS DOCUMENT LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21327386&pc=FRN | 9615050-2025-00063 | 2025-01-01 05:00 | Death | ICU MEDICAL COSTA RICA LTD. | 2025-02-06 05:00 | FRN | PLUM 360 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K161469 |  | 1 | Event Description: A COMPLAINT WAS RECEIVED WITH REGARDS TO A PLUM 360 THAT THERE WAS A DEATH WITHOUT KNOWN PRODUCT PROBLEM. IT WAS REPORTED THAT THE PATIENT HAD EXPIRED. ADDITIONALLY, IT WAS STATED THAT THEY DO NOT BELIEVE THE PUMP IS THE PROBLEM, BUT THE PATIENT'S FAMILY TOOK A PICTURE OF THE PUMP AND THEY QUESTIONED WHAT THEY SAW ON THE PUMP TO IMPLY THE CLINICIAN MAY NOT HAVE DONE SOMETHING RIGHT/CORRECT. IT WAS NOTED THAT THE EVENT OCCURRED DURING INFUSION. Manufacturer Narrative: THIS IS A GENERAL COMPLAINT, AND NO DEVICE WAS RETURNED TO THE SERVICE HUB FOR ANALYSIS AND EVALUATION. THEREFORE, NO VISUAL INSPECTION AND FUNCTIONAL TESTING COULD BE PERFORMED TO DETERMINE IF THE DEVICE PLAYED A ROLE IN THE ADVERSE EVENT. A 12-MONTH REVIEW WAS NOT ABLE TO BE CONDUCTED AS NO SERIAL NUMBER WAS PROVIDED TO ICU MEDICAL. PUMP EVENT LOGS WERE NOT PROVIDED. CONCLUSION: IN CONCLUSION, SINCE THE DEVICE WAS NOT RETURNED TO THE SERVICE HUB FOR ANALYSIS AND EVALUATION, NO VISUAL INSPECTION AND FUNCTIONAL TESTING COULD BE PERFORMED TO DETERMINE IF THE DEVICE HAD A ROLE IN THE ADVERSE EVENT. Manufacturer Narrative: B2: DATE OF DEATH: THE DATE OF DEATH IS UNKNOWN, AND 01 JAN 2025 HAS BEEN USED AS A PLACEHOLDER. B3: DATE OF EVENT: THE DATE OF EVENT IS UNKNOWN, AND 01 JAN 2025 HAS BEEN USED AS A PLACEHOLDER. D4 AND H4 THE SERIAL#, AND MANUFACTURE DATE ARE UNKNOWN. THE INVESTIGATION IS PENDING COMPLETION. UPON COMPLETION OF THE INVESTIGATION A SUPPLEMENTAL MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21327495&pc=MSX | 3009156722-2025-00003 | 2025-01-13 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-02-06 05:00 | MSX | TD60 TELEMETRY TRANSMITTER | Device Alarm System | Insufficient Information | K220058 |  | 1 | Event Description: NO PRODUCT MALFUNCTION WAS REPORTED. IT WAS REPORTED THAT ON (B)(6) 2025, AT 20:15, THE PATIENT IN ROOM (B)(6), CONNECTED TO THE TD60 TELEMETRY MONITOR (ID TELE 22), WAS EXPERIENCING RHYTHM CHANGES AND MULTIPLE PVCS. THE PATIENT WAS ALSO REPORTED TO BE OFF TELE 45 MINUTES TO AN HOUR BEFORE ASYSTOLE. THE PATIENT EXPIRED.Event Description: NO PRODUCT MALFUNCTION WAS REPORTED. IT WAS REPORTED THAT ON (B)(6) 2025, AT 20:15, THE PATIENT IN ROOM (B)(6), CONNECTED TO THE TD60 TELEMETRY MONITOR (ID TELE 22), WAS EXPERIENCING RHYTHM CHANGES AND MULTIPLE PVCS. THE PATIENT WAS ALSO REPORTED TO BE OFF TELE 45 MINUTES TO AN HOUR BEFORE ASYSTOLE. THE PATIENT EXPIRED. Manufacturer Narrative: THE MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND ECG STRIPS FOR INVESTIGATION. THE TD60 WAS MISPLACED BY THE USER AND WAS NOT AVAILABLE FOR PHYSICAL EVALUATION. THE USER FACILITY CONFIRMED THAT THE TIME OF PATIENT DISCONNECTION WAS 19:13:58, AND THE TIME OF PATIENT DEATH WAS APPROXIMATELY 21:46. LOG ANALYSIS SHOWS THAT DURING THE INCIDENT, THE DEVICE ISSUED APPROPRIATE ALARMS, AND THE DMS GENERATED EXPECTED SOUNDS. THE BENEVISION DMS IN USE WITH TD60 PERFORMED PER SPECIFICATIONS; NO PRODUCT MALFUNCTION OCCURRED. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND ECG STRIPS FOR INVESTIGATION. THE TD60 WAS MISPLACED BY THE USER AND WILL NOT AVAILABLE FOR PHYSICAL EVALUATION. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21327496&pc=MSX | 2221819-2025-00003 | 2025-01-13 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-02-06 05:00 | MSX | TD60 TELEMETRY TRANSMITTER | Device Alarm System | Insufficient Information |  |  | 1 | Event Description: NO PRODUCT MALFUNCTION WAS REPORTED. IT WAS REPORTED THAT ON (B)(6), 2025, AT 20:15, THE PATIENT IN ROOM 432, CONNECTED TO THE TD60 TELEMETRY MONITOR (ID TELE 22), WAS EXPERIENCING RHYTHM CHANGES AND MULTIPLE PVCS. THE PATIENT WAS ALSO REPORTED TO BE OFF TELE 45 MINUTES TO AN HOUR BEFORE ASYSTOLE. THE PATIENT EXPIRED.Event Description: NO PRODUCT MALFUNCTION WAS REPORTED. IT WAS REPORTED THAT ON (B)(6) 2025, AT 20:15, THE PATIENT IN ROOM (B)(6), CONNECTED TO THE TD60 TELEMETRY MONITOR (ID TELE 22), WAS EXPERIENCING RHYTHM CHANGES AND MULTIPLE PVCS. THE PATIENT WAS ALSO REPORTED TO BE OFF TELE 45 MINUTES TO AN HOUR BEFORE ASYSTOLE. THE PATIENT EXPIRED. Manufacturer Narrative: THE MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND ECG STRIPS FOR INVESTIGATION. THE TD60 WAS MISPLACED BY THE USER AND WAS NOT AVAILABLE FOR PHYSICAL EVALUATION. THE USER FACILITY CONFIRMED THAT THE TIME OF PATIENT DISCONNECTION WAS 19:13:58, AND THE TIME OF PATIENT DEATH WAS APPROXIMATELY 21:46. LOG ANALYSIS SHOWS THAT DURING THE INCIDENT, THE DEVICE ISSUED APPROPRIATE ALARMS, AND THE DMS GENERATED EXPECTED SOUNDS. THE BENEVISION DMS IN USE WITH TD60 PERFORMED PER SPECIFICATIONS; NO PRODUCT MALFUNCTION OCCURRED. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND ECG STRIPS FOR INVESTIGATION. THE TD60 INVOLVED IN THE OCCURRENCE WAS LOST, AND MINDRAY WILL NOT BE ABLE TO TEST IT. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21327677&pc=DSQ | 2916596-2025-00752 | 2021-03-01 05:00 | Death | THORATEC CORPORATION | 2025-02-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "HOSPITAL READMISSIONS IN PATIENTS SUPPORTED WITH DURABLE CENTRIFUGAL-FLOW LEFT VENTRICULAR ASSIST DEVICES" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH DEATH. THERE WERE FIFTEEN PATIENTS WHO DIED PRIOR TO HOSPITAL DISCHARGE FOLLOWING LVAD IMPLANTATION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DIVISION OF CARDIOVASCULAR MEDICINE, DEPARTMENT OF INTERNAL MEDICINE, UNIVERSITY OF UTAH HEALTH AND SCHOOL OF MEDICINE, SALT LAKE CITY, UT; GEORGE E. WAHLEN DEPARTMENT OF VETERANS AFFAIRS MEDICAL CENTER, SALT LAKE CITY, UT; NORA ECCLES HARRISON CARDIOVASCULAR RESEARCH AND TRAINING INSTITUTE, UNIVERSITY OF UTAH, SALT LAKE CITY, UT; DIVISION OF CARDIOTHORACIC SURGERY, DEPARTMENT OF SURGERY, UNIVERSITY OF UTAH HEALTH AND SCHOOL OF MEDICINE, SALT LAKE CITY, UT; DEPARTMENT OF HEALTH POLICY & ORGANIZATION, SCHOOL OF PUBLIC HEALTH, THE UNIVERSITY OF ALABAMA AT BIRMINGHAM, BIRMINGHAM, AL; AND THE ARNOLD SCHOOL OF PUBLIC HEALTH, UNIVERSITY OF SOUTH CAROLINA, COLUMBIA, SC. KYRIAKOPOULOS, C. P., SELZMAN, C. H., GIANNOUCHOS, T. V., MYLAVARAPU, R., SIDERIS, K., ELMER, A., VANCE, N., HANFF, T. C., KAGAWA, H., STEHLIK, J., DRAKOS, S. G., & GOODWIN, M. L. (2024). HOSPITAL READMISSIONS IN PATIENTS SUPPORTED WITH DURABLE CENTRIFUGAL-FLOW LEFT VENTRICULAR ASSIST DEVICES. JOURNAL OF CLINICAL MEDICINE, 13(10), 2869. HTTPS://DOI.ORG/10.3390/JCM13102869 . MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. CENTRIFUGAL-FLOW LEFT VENTRICULAR ASSIST DEVICES (CF-LVADS) HAVE IMPROVED MORBIDITY AND MORTALITY FOR LVAD RECIPIENTS. HOSPITAL READMISSIONS, HOWEVER, REMAIN COMMON AND AFFECT QUALITY OF LIFE AND SURVIVAL. WE SOUGHT TO IDENTIFY FACTORS ASSOCIATED WITH READMISSIONS ON LVAD SUPPORT. HOSPITAL READMISSIONS FOLLOWING LVAD IMPLANTATION REMAIN A COMMON AND VEXING CLINICAL ISSUE ADVERSELY AFFECTING PATIENT QUALITY OF LIFE AND SURVIVAL, AS WELL AS INCREASING HEALTHCARE COSTS AND RESOURCE UTILIZATION [1¿4]. IT HAS BEEN REPORTED THAT UP TO 71% OF PATIENTS ARE REHOSPITALIZED WITHIN ONE YEAR FOLLOWING LVAD IMPLANTATION. DATA WERE COLLECTED TWO TO FOUR MONTHS FOLLOWING LVAD IMPLANTATION, WITH AT LEAST TWO CONSECUTIVE WEEKS BEING REQUIRED AS A MINIMUM FOR A PATIENT TO BE CONSIDERED AS BEING ON A SPECIFIC PHARMACOLOGIC AGENT THERAPY. PATIENTS RECEIVING A CF-LVAD BETWEEN FEBRUARY 2011 AND MARCH 2021 WERE PROSPECTIVELY EVALUATED. HOSPITAL READMISSIONS WERE DICHOTOMIZED INTO HF/LVAD-RELATED OR NON-HF/LVAD-RELATED. MULTIVARIABLE LOGISTIC REGRESSION MODELS AUGMENTED USING THE LEAST ABSOLUTE SHRINKAGE AND SELECTION OPERATOR METHOD FOR VARIABLE SELECTION WERE USED TO ESTIMATE ASSOCIATIONS BETWEEN =1 HF/LVAD-RELATED READMISSIONS AND CLINICAL VARIABLES. THE PRIMARY OUTCOME OF INTEREST WAS THE INCIDENCE OF ANY HF/LVAD-RELATED READMISSION UP TO THREE YEARS POST-LVAD IMPLANTATION HOSPITAL DISCHARGE. THE SECONDARY OUTCOME WAS THE HAZARD OF HF/LVAD-RELATED READMISSIONS WHILE ON LVAD SUPPORT OVER THE THREE YEARS POST-LVAD IMPLANTATION HOSPITAL DISCHARGE. A MODEL WAS CREATED ACCURATELY IDENTIFYING PATIENTS WITH A HIGH LIKELIHOOD OF READMISSIONS AFTER LVAD IMPLANT. MOST VARIABLES ASSOCIATED WITH READMISSION CAN BE IDENTIFIED PRE-OPERATIVELY. THE FINDINGS CAN GUIDE THE IMPLEMENTATION OF THERAPEUTIC STRATEGIES TO MITIGATE READMISSION RISK AND FURTHER IMPROVE OUTCOMES. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21328708&pc=DYB | 2015691-2025-00836 | 2025-01-14 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-06 05:00 | DYB | EDWARDS ESHEATH PLUS | Difficult to Insert | Vascular Dissection | K200258 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT A SUBCLAVIAN TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION. PRIOR TO THE PROCEDURE, A PROCTOR REVIEW RECOMMENDED A CAROTID APPROACH DUE TO A TORTUOUS RIGHT SUBCLAVIAN BUT THE PHYSICIAN PROCEEDED WITH SUBCLAVIAN. DURING THE PROCEDURE, THE ESHEATH WAS POINTING LATERAL WHEN THE DILATOR WAS REMOVED. THE TEAM THEN ADVANCED THE PIGTAIL CATHETER FOR ANGIOGRAM AND PERFORMED ANGIO IN THE DISSECTION PLANE. THE PATIENT HAD AN ASCENDING AORTIC DISSECTION UP TO THE INNONIMATE. VASCULAR SURGERY WAS CALLED TO CONSULT; HOWEVER, THE PATIENT WAS NOT A CANDIDATE. THE SURGEON CALLED ANOTHER FACILITY WITH VASCULAR SURGERY TO TAKE OVER CARE. THE PATIENT WAS TRANSFERRED TO PACU AWAITING TRANSFER IN STABLE CONDITION; HOWEVER, THE PATIENT EXPIRED IN THE ICU OVERNIGHT. AS PER FOLLOW-UP WITH THE FCS, THE OPERATOR ALSO PULLED BOTH THE DILATOR AND THE WIRE AFTER THE EDWARDS ESHEATH WAS INTRODUCED. THE SHEATH POPPED UP WITHOUT THE SUPPORT OF THE WIRE WHICH CAUSED THE DISSECTION. THERE WAS NO ALLEGATION THE ESHEATH MALFUNCTIONED. THE PERCEIVED ROOT CAUSE OF THE DEATH IS DUE TO THE AORTIC DISSECTION. THE DEVICES WERE DISCARDED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2 AND H11 HAS BEEN UPDATED. THE COMPLAINT FOR DIFFICULTY INTRODUCING SHEATH WAS UNABLE TO BE CONFIRMED AS NO DEVICE AND/OR RELEVANT IMAGERY WERE PROVIDED. THE REPORTED EVENT DOES NOT ALLEGE A MALFUNCTION THAT COULD BE RELATED TO AN EDWARDS MANUFACTURING DEFICIENCY. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. AS REPORTED, "DURING THE PROCEDURE, THE ESHEATH WAS POINTING LATERAL WHEN THE DILATOR WAS REMOVED. THE TEAM THEN ADVANCED THE PIGTAIL CATHETER FOR ANGIOGRAM AND PERFORMED ANGIO IN THE DISSECTION PLANE. THE PATIENT HAD AN ASCENDING AORTIC DISSECTION UP TO THE INNONIMATE." ADDITIONAL FOLLOW-UP CLARIFIED, "THE OPERATOR ALSO PULLED BOTH THE DILATOR AND THE WIRE AFTER THE EDWARDS ESHEATH WAS INTRODUCED. THE SHEATH POPPED UP WITHOUT THE SUPPORT OF THE WIRE AND CAUSED THE DISSECTION." DURING TAVR, A GUIDEWIRE IS USED TO FACILITATE SAFE SHEATH PLACEMENT AS WELL AS TO KEEP THE SHEATH ORIENTATION SECURE AND PARALLEL WITH RESPECT TO PATIENT-S VASCULATURE. RETRIEVAL OF THE GUIDEWIRE ALONG WITH THE INTRODUCER, WITHOUT HOLDING THE SHEATH IN PLACE, COULD CAUSE THE SHEATH TO DISLOCATE AND DISSECT THE VESSEL, AS REPORTED. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT USE ERROR (DEVICE MANIPULATION TECHNIQUE) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21328796&pc=OZD | 1220648-2025-26044 | 2025-01-12 05:00 | Death | ABIOMED, INC. | 2025-02-06 05:00 | OZD | IMPELLA CP | Crack; Pumping Stopped |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK UNDERGOING AN IMPELLA CP DEVICE IMPLANT FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED A PUMP STOP AND WOULD LATER EXPIRE. THE PATIENT WAS ADMITTED WITH DISTAL LEFT MAIN, LEFT ANTERIOR DESCENDING AND CIRCUMFLEX ARTERIES DISEASE AND UNDER SURGICAL CONSULT WAS PLACED ON AN INTRA-AORTIC BALLOON PUMP (IABP) AND SENT TO THE UNIT. THE PATIENT BEGAN TO FURTHER DECOMPENSATE AND WAS ESCALATED TO AN IMPELLA CP FROM AN IABP. THE RIGHT FEMORAL ARTERY WAS PREPPED AND THE IABP WAS REMOVED. PRIMING STEPS WERE BEING FOLLOWED WHEN THE AUTOMATED IMPELLA CONTROLLER (AIC) SKIPPED THE PRIMING STEP AND WENT STRAIGHT TO THE IMPELLA START MODE/INITIAL PLACEMENT SCREEN. ON AIC PERFORMANCE LEVEL WAS P-0 WITH THE IMPELLA STOPPED RETROGRADE FLOW FOLLOWED BY IMPELLA STOPPED MOTOR CURRENT HIGH. THE TEAM ATTEMPTED TO RE-PRIME WITH DE-AIRING PROCESS; NO FLUID WAS SEEN COMING AT AROUND MOTOR AREA. AT THIS POINT, THE PATIENT UNSTABLE AS IABP WAS ALREADY REMOVED. THEREFORE, A NEW IMPELLA CP DEVICE WAS OPENED, AND THE PHYSICIAN PROCEEDED WITH IMPLANT WHICH WAS COMPLETED WITHOUT INCIDENT. ADDITIONALLY, THE 14FR 13CM PEEL-AWAY SHEATH HAD A CRACKED AT THE HUB AND THEREFORE, IT WAS REPLACED BY A NEW ONE. AFTER ONE DAY ON THE REPLACEMENT IMPELLA CP SUPPORT, IT WAS NOTED LACTATE CAME DOWN BUT DID NOT COMPLETELY CLEAR. THEREFORE, A PLAN WAS MADE IT TO ESCALATE TO THE IMPELLA 5.5 TO UNDERGO PROTECTED PERCUTANEOUS CORONARY INTERVENTION WITH NO PLAN FOR CORONARY ARTERY BYPASS GRAFT SURGERY. THE PATIENT WAS BROUGHT TO THE OPERATING AND PERIPHERAL VENO-ARTERIAL (VA) CANNULATED. THE PHYSICIAN LOST ACCESS FOR LEFT FEMORAL ARTERY EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SITE. A RETROPERITONEAL BLEED OCCURRED WHICH WAS MANAGED BY SURGICAL CUTDOWN AND REPAIRED. IMPELLA CP WAS REMOVED AND DILATED UP TO USE AS ARTERIAL ACCESS FOR VA. NEXT RIGHT AXILLARY ARTERY EXPOSED, 10MM GRAFT ANASTOMOSIS DONE. AXILLARY KIT WAS USED, VENTRICULAR ACCESS WAS DONE WITH A JR4 CATHETER. THE IMPELLA 5.5 DEVICE WAS ADVANCED AND SLOWLY RAMPED COMING ON ECMO FLOWS. IMPELLA WAS SECURED TO CHEST. NEXT PERIPHERAL VA DECANNULATED. MASSIVE BLOOD TRANSFUSION WAS REQUIRED. THE PATIENT RECEIVED 7 UNITS OF PACKED RED BLOOD CELLS, 1 UNITS OF PLATELETS, AND 1 UNITS OF FRESH FROZEN PLASMA. THE PATIENT CONTINUED TO CODE ON THE UNIT, MASSIVE BLOOD PRODUCTS TRANSFUSION GOING; THE PATIENT WAS BLEEDING INTERNALLY PER THE PLAN OF CARE UPDATES. AFTER APPROXIMATELY THREE DAYS ON IMPELLA 5.5 SUPPORT THE PATIENT WOULD EXPIRE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA CP DEVICE INITIALLY IMPLANTED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THERE WAS AN INITIAL DELAY IN GETTING THE PATIENT ON APPROPRIATE SUPPORT. ONCE SUPPORT WAS MET, THE PATIENT RECEIVED IMPELLA MCS FOR APPROXIMATELY 4 DAYS. PATIENT HAD PROCEDURAL COMPLICATIONS RELATED TO NON-ABIOMED DEVICE(S) WHICH MOST LIKELY RESULTED IN THE PATIENT'S DEMISE, HOWEVER, (WHICH INCLUDED INTERNAL BLEEDING, MASSIVE BLOOD TRANSFUSION REQUIREMENTS). Manufacturer Narrative: THE IMPELLA PUMP WITH THE INTRODUCER HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE PUMP STOP AND INTRODUCER DAMAGE HAS BEEN COMPLETED. THE PUMP AND SHORT INTRODUCER WERE RETURNED FOR INVESTIGATION. THE RETURNED PUMP HAD THE RED LUMEN AS IT WAS ASSEMBLED. THE PUMP WAS PRIMED AND OPERATED ACCORDING TO THE SPECIFICATIONS AFTER REMOVING THE RED LUMEN. NO OTHER PHYSICAL DAMAGE WAS NOTED ON THE RETURNED PUMP. THE INTRODUCER WAS INVESTIGATED, AND IT WAS FOUND THAT ONE SIDE OF THE ORANGE HUB WAS DETACHED AND MISSING FROM THE RETURNED PACKAGE. IT WAS NOT POSSIBLE TO VERIFY THE APPLIED ADHESIVE DUE TO THE MISSING COMPONENT. NO OTHER DAMAGE WAS NOTED TO THE INTRODUCER. LOGS PROVIDED EVIDENCE THAT THE CASE START WAS SUCCESSFULLY COMPLETED AND THE PUMP WAS SET TO P9 WHEN THE PUMP STOP-MOTOR CURRENT HIGH ALARM WAS REPORTED. THE CAUSE OF THE PUMP DID NOT START ISSUE WAS A RED LUMEN REMOVAL ISSUE. RED LUMEN SHOULD BE REMOVED BEFORE STARTING OR IMPLANTING THE PUMP. THE CAUSE OF THE INTRODUCER DAMAGE ISSUE WAS NOT DETERMINED SINCE SUFFICIENT PRODUCT WAS NOT RETURNED. B.2 REVISED DATE OF DEATH AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26044. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26044 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21329021&pc=OJX | 2649622-2025-03640 | 2025-01-21 05:00 | Death | MPRI | 2025-02-06 05:00 | OJX | ATTAIN PERFORMA S MRI SURESCAN | Appropriate Term/Code Not Available | Intracranial Hemorrhage | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BRAIN CLOT AND LATER EXPIRED. THE PATIENT WAS FOUND ON THE GROUND BY A FAMILY MEMBER AND WAS TAKEN TO THE EMERGENCY DEPARTMENT. THE PATIENT WAS HOSPITALIZED, AND IT WAS DISCOVERED THAT THE PATIENT HAD A BRAIN BLEED. IT WAS NOTED THAT THE ORDER OF EVENTS WAS UNCERTAIN. CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM REPROGRAMMING AND A DO-NOT-RESUSCITATE (DNR) ORDER WERE REQUESTED. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21329022&pc=DTB | 2649622-2025-03639 | 2025-01-21 05:00 | Death | MPRI | 2025-02-06 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Intracranial Hemorrhage | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BRAIN CLOT AND LATER EXPIRED. THE PATIENT WAS FOUND ON THE GROUND BY A FAMILY MEMBER AND WAS TAKEN TO THE EMERGENCY DEPARTMENT. THE PATIENT WAS HOSPITALIZED, AND IT WAS DISCOVERED THAT THE PATIENT HAD A BRAIN BLEED. IT WAS NOTED THAT THE ORDER OF EVENTS WAS UNCERTAIN. CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM REPROGRAMMING AND A DO-NOT-RESUSCITATE (DNR) ORDER WERE REQUESTED. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21329026&pc=NIK | 9614453-2025-00572 | 2025-01-21 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-06 05:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Intracranial Hemorrhage | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BRAIN CLOT AND LATER EXPIRED. THE PATIENT WAS FOUND ON THE GROUND BY A FAMILY MEMBER AND WAS TAKEN TO THE EMERGENCY DEPARTMENT. THE PATIENT WAS HOSPITALIZED, AND IT WAS DISCOVERED THAT THE PATIENT HAD A BRAIN BLEED. IT WAS NOTED THAT THE ORDER OF EVENTS WAS UNCERTAIN. CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM REPROGRAMMING AND A DO-NOT-RESUSCITATE (DNR) ORDER WERE REQUESTED. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21329170&pc=NPT | 2015691-2025-00841 | 2012-11-01 04:00 | Death | EDWARDS LIFESCIENCES | 2025-02-06 05:00 | NPT | SAPIEN XT TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P130009 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN TURKEY. THE FOLLOWING MEDICAL ARTICLE WAS REVIEWED, ''INCIDENCE AND PREDICTORS OF PERMANENT PACEMAKER IMPLANTATION AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION WITH A BALLOON-EXPANDABLE BIOPROSTHESIS IN PATIENTS WITH BICUSPID AORTIC VALVES'', CORRESPONDING AUTHOR HAKAN SUYGUN ET AL., THROUGH REVIEW OF THE ARTICLE, THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE: ONE PATIENT WITH AN IMPLANT OF SAPIEN XT VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH UNDERWENT PERCUTANEOUS INTERVENTION DUE TO CORONARY OCCLUSION AFTER TAVI, 3 DAYS LATER THE PATIENT DIED. IN THIS STUDY A TOTAL OF 62 PATIENTS ARE STUDIED WITH THE PURPOSE OF EVALUATING THE PREDICTORS AND INCIDENCE OF PPMI IN BICUSPID PATIENTS USING A BALLOON-EXPANDABLE (BE) TAVI DEVICE. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO UPDATE H10. Manufacturer Narrative: THE DATE OF THE EVENT IS UNKNOWN; HOWEVER, THE ARTICLE PROVIDED THE FOLLOWING DATE RANGE ''FROM NOVEMBER 2012 UNTIL JANUARY 2018''. THEREFORE, 1ST NOVEMBER ON 2012 WAS USED AS THE OCCURRENCE DATE. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, DESPITE MULTIPLE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THIS EVENT. DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. AT THE TIME OF THIS REPORT, THE INFORMATION AVAILABLE DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF FOUR MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21329222&pc=NPT | 2015691-2025-00842 | 2025-01-16 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-06 05:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Cardiac Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN FRANCE, THIS WAS A CASE WITH A 29MM SAPIEN 3 VALVE, IN AORTIC POSITION BY TRANSCAROTID APPROACH. DURING THE PROCEDURE, THE FIRST VALVE EMBOLIZED INTO THE AORTA AND THE PATIENT HAD A CARDIAC ARREST. A SECOND VALVE WAS IMPLANTED TO SECURE THE FIRST VALVE IN THE ASCENDING AORTA. A THIRD VALVE WAS DEPLOYED IN THE CORRECT POSITION (90/10). AFTER A FEW MINUTES, THE PATIENT WAS DECLARED AS DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF VALVE EMBOLIZATION WAS UNKNOWN. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. BASED ON THE INFORMATION RECEIVED, THE ROOT CAUSE OF THE EVENT REMAINS INDETERMINABLE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21329268&pc=NPT | 2015691-2025-00845 | 2024-03-28 04:00 | Death | EDWARDS LIFESCIENCES | 2025-02-06 05:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Fluid/Blood Leak |  | P140031 |  | 1 | Event Description: THROUGH THE REVIEW OF THE MEDICAL ARTICLE ''PRESSURE RECOVERY IN TRANSCATHETER AORTIC VALVE REPLACEMENT" CORRESPONDING AUTHOR PRIYANKA RATIBHAN SINGH, HASAN MIRZA, MOHSIN S. MUGHAL ET AL. THE FOLLOWING EVENT WAS IDENTIFIED: A 69-YEAR-OLD FEMALE, PREVIOUSLY WITH A HISTORY OF AORTIC VALVE REPLACEMENT (AVR) USING A NON-EW BOVINE PERICARDIAL TISSUE VALVE (25 MM), PRESENTED WITH SYMPTOMATIC SEVERE AS AND WORSENING CARDIOMYOPATHY (LVEF 40% FROM 75%). DUE TO HER PROHIBITIVE SURGICAL RISK, SHE WAS REFERRED FOR TRANSCATHETER AVR (TAVR). THE PRE TAVR-ECHOCARDIOGRAM REVEALED A MEAN TRANSVALVULAR GRADIENT OF 37 MMHG, AN AORTIC VALVE AREA (AVA) OF 0.8 CM2, AND A PEAK VELOCITY OF 4.1 M/S, ACCOMPANIED BY MODERATE AORTIC REGURGITATION. IT WAS A CASE OF VALVE IN VALVE PROCEDURE WITH A 23 MM SAPIEN 3 VALVE IN THE NON-EW VALVE IN THE AORTIC POSITION. AFTER VALVE DEPLOYMENT, THE INTRAPROCEDURAL MEAN GRADIENT ON ECHO REMAINED BETWEEN 20 AND 25 MMHG. BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED UNDER RAPID PACING TO FRACTURE THE SURGICAL VALVE. UNFORTUNATELY, THIS WAS ASSOCIATED WITH HEMODYNAMIC COMPROMISE AND AORTIC INSUFFICIENCY, LEADING TO THE PATIENT'S DEMISE DESPITE AGGRESSIVE HEMODYNAMIC SUPPORT, INCLUDING EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORT. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CORRECTED H.6 INVESTIGATION CONCLUSIONS AND INVESTIGATION FINDINGS. ADDED NEW INFORMATION TO H.6 TYPE OF INVESTIGATION. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, CENTRAL REGURGITATION EVENTS POST-DEPLOYMENT WITH OR WITHOUT REPORT RELATED TO LEAFLET MOBILITY FOR THE S3 AND S3U VALVES HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PROCEDURAL FACTORS (MALPOSITION, SLOW RECOVERY OF BLOOD FLOW, LEAFLET IMPINGEMENT DUE TO CALCIFICATION OR GUIDEWIRE, OVER-INFLATION, POST-DILATION, AND/OR UNDER-EXPANSION). THE COMPLAINT WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF RELEVANT IMAGERY AND/OR MEDICAL RECORDS. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (POST-DILATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT ''BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED UNDER RAPID PACING TO FRACTURE THE SURGICAL VALVE. UNFORTUNATELY, THIS WAS ASSOCIATED WITH HEMODYNAMIC COMPROMISE AND AORTIC INSUFFICIENCY.'' DEPLOYING THE VALVE USING MORE THAN THE PRESCRIBED VOLUME OR ATTEMPTING A SECOND INFLATION MAY RESULT IN INABILITY FOR THE VALVE LEAFLETS TO PROPERLY FUNCTION. THE INCREASE IN VALVE DIAMETER DUE TO OVERINFLATION OR POST DILATION MAY PREVENT PROPER MOTION OF THE THV LEAFLETS RESULTING IN THV REGURGITATION. ADDITIONALLY, PER THE TRAINING MANUAL, CENTRAL REGURGITATION IS AN ASSOCIATED RISK OF POST-DILATION. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. NO IFU/LABELING/TRAINING MANUAL INADEQUACIES WERE IDENTIFIED. THEREFORE, NO CORRECTIVE OR PREVENTATIVE ACTION NOR PRODUCT RISK ASSESSMENT IS REQUIRED AT THIS TIME. Manufacturer Narrative: THE DATE OF DEATH, IMPLANT DATE AND SERIAL NUMBER ARE UNKNOWN. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21332531&pc=QBJ | MW5165729 |  | Death | ABBOTT DIABETES CARE INC. | 2025-02-06 05:00 | QBJ | FREESTYLE LIBRE 2 SENSOR |  | Insufficient Information |  |  | 1 | Event Description: THE PATIENT HAS PASSED AWAY. CAUSE AND DATE OF DEATH NOT REPORTED. NO ADDITIONAL INFORMATION REPORTED OR AVAILABLE REGARDING THIS EVENT. PAE REPORTED BY HCP, WHO DOES NOT CONSENT TO FOLLOW UP. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309784&pc=PNJ | 2017865-2025-12330 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-05 05:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT IN CARDIAC ARREST. CPR (CARDIOPULMONARY RESUSCITATION) WAS PERFORMED AND MEDICATION WAS GIVEN. THE PATIENT WAS THEN IN VENTRICULAR FIBRILLATION AND SHOCK THERAPY WAS DELIVERED. THE PATIENT THEN BECAME BRADYCARDIC AND WAS INTUBATED. THE PATIENT WAS UNSTABLE AND EVENTUALLY PASSED AWAY. THE OFFICIAL CAUSE OF DEATH WAS UNKNOWN. THERE WAS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE PATIENT'S DEATH WAS CAUSED OR CONTRIBUTED TO BY THE LEADLESS PACEMAKER. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR ANALYSIS DUE TO DECEASED PATIENT. VISUAL INSPECTION OF THE DEVICE FOUND NO ANOMALY WITH THE HELIX. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309794&pc=OAE | 3008452825-2025-00050 | 2025-01-06 05:00 | Death | ST. JUDE MEDICAL | 2025-02-05 05:00 | OAE | TACTIFLEX¿ ABLATION CATHETER, SENSOR ENABLED¿ | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | P220013 |  | 1 | Event Description: ON (B)(6) 2025, A CRYO REDO PULMONARY VEIN ISOLATION PROCEDURE WAS PERFORMED. FOLLOWING THE PROCEDURE THERE WERE COMPLICATIONS WITH A CIRCULAR PERICARDIAL EFFUSION AND THE PATIENT EXPIRED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH NO ALLEGED PERFORMANCE ISSUES WITH ANY ABBOTT DEVICES. A PERICARDIAL TAMPONADE WAS EXCLUDED WITH ULTRASOUND DIRECTLY AFTER THE PROCEDURE IN THE LAB. HOWEVER, 4 HOURS AFTER THE PROCEDURE, THE PATIENT BECAME HYPOTENSIVE IN THE WARD. AN ECHOCARDIOGRAM AND CT SHOWED A CIRCULAR PERICARDIAL EFFUSION. DURING TREATMENT OF THE CIRCULAR PERICARDIAL EFFUSION THERE WERE COMPLICATIONS, A PERICARDIAL PUNCTURE WAS PERFORMED, HOWEVER DURING THIS THERE WAS A HEMOTHORAX. THE CAUSE IS NOT SPECIFIED. LATER, ON (B)(6) THE PATIENT PASSED AWAY DUE TO THE COMPLICATIONS. Manufacturer Narrative: THE LOT NUMBER OF THE DEVICE WAS CONFIRMED TO BE 10570508. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS; HOWEVER, THE LOG FILES FROM THE TACTISYS QUARTZ SYSTEM WERE RETURNED. THE LOG FILE ANALYSIS CONCLUDED THAT THE CATHETER PERFORMED AS INTENDED. THE OPTICAL FIBERS MET SPECIFICATIONS WHILE INSIDE THE PATIENT, THE RECORDED TEMPERATURES INDICATED COOLING DURING RF ABLATION, AND CONTACT FORCE MEASUREMENTS WERE DISPLAYED THROUGHOUT THE DURATION OF THE LOG FILES. FORCE MEASUREMENTS WERE RECORDED DURING ABLATION THAT EXCEEDED THE RECOMMENDED VALUES IN THE TACTIFLEX SE ABLATION CATHETER INSTRUCTIONS FOR USE (IFU); HOWEVER, DUE TO UNKNOWN PROCEDURAL CONDITIONS WE ARE UNABLE TO CONCLUSIVELY DETERMINE THE CAUSE OF THE REPORTED EVENT. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. DUE TO UNKNOWN PROCEDURAL CONDITIONS, WE ARE UNABLE TO CONCLUSIVELY DETERMINE THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION AND SUBSEQUENT PATIENT DEATH REMAINS UNKNOWN. PER THE IFU, CARDIAC TAMPONADE AND PERICARDIAL EFFUSION ARE KNOWN RISKS DURING THE USE OF THIS DEVICE. Manufacturer Narrative: MODEL/LOT NUMBER INFORMATION WAS NOT ABLE TO BE OBTAINED FOR THIS DEVICE. THEREFORE, FULL UDI INFORMATION(D4) AND 510K(G3) ARE NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21311790&pc=BZD | 2518422-2025-101723 | 2024-12-30 05:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR A RETURN LABEL SO THE DEVICE CAN BE RETURNED. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS BEEN RETURNED FOR INVESTIGATION, BUT THE EVALUATION IS NOT YET COMPLETE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312315&pc=NIK | 2017865-2025-12342 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-05 05:00 | NIK | UNIFY ASSURA CRT-D RF HV | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF THE DEATH WAS UNKNOWN. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS EXPLANTED. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS AT EOS WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312373&pc=DQY | 2029214-2025-00285 | 2025-01-25 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-05 05:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Ischemia; Thromboembolism; Ischemia Stroke; Respiratory Insufficiency | K180715 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT A REACT CATHETER PATIENT¿S CLOT PUSHED DISTALLY, PASSING FROM M2 TO M3-M4. PATIENT DETAILS: MRS NOT AVAILABLE AND NIHSS SCORE 21. THE EVENT WAS ASSESSED AS SAE, CAUSALLY RELATED TO PROCEDURE AND POSSIBLY RELATED TO STUDY DEVICE UTILIZED. LOCATION OF PROXIMAL FACE OF CLOT 1: PRE-PROCEDURE MTICI SCORE 0 AND FINAL POST-PROCEDURE MTICI SCORE 0. ANCILLARY DEVICES: ASPIRATION CATHETER 3 MAX, STENT RETRIEVER TREVO 3X20, STENT RETRIEVER NIMBUS ADDITIONAL INFORMATION RECEIVED THAT A POTENTIAL COMPLAINT OF FAILED MECHANICAL THROMBECTOMY ATTEMPTED FOR A CLOT OCCLUSION MCA, M2 - L. A TOTAL OF 7 PASSES PERFORMED WITH REACT -68 AND OTHER NON-MEDTRONIC MT DEVICES WITH MTICI 0 AFTER EACH PASS. THE ACCESSED VESSEL WAS THE FEMORAL ARTERY WITH A CLOT LOCATION IN THE MCA, M2-L. ADDITIONAL INFORMATION RECEIVED REPORTED LESION CHARACTERISTICS (LOCATION, SIZE, TYPE, SIDE, DIMENSION, ETC.): CLOTPUSHED DISTALLY, PASSING FROM M2 TO M3-M4. THE PATIENT VESSEL TORTUOSITY WAS UNKNOWN. FINAL POST-PROCEDURE MTICI SCORE WAS 0. POST-PROCEDURE 24H NIHSS 22 AND OUTCOME OF THE EVENT IS NOT REPORTED. IT WAS UNKNOWN IF THE DEVICE WAS PREPARED AS INDICATED IN THE IFU (INSPECTION, HYDRATION, ETC.).Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED A SURGICAL PROCEDURE WAS REQUIRED. THE EVENT WAS CAUSED BY THE THROMBECTOMY. IT HAD A CAUSAL RELATIONSHIP TO THE PROCEDURE. THE EVENT WAS NOT RELATED TO DISEASE UNDERSTUDY.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED NO HOSPITALIZATION, TREATED WITH SURGICAL PROCEDURE, NO OTHER ACTIONS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT HAD A PRESENCE OF SMALL ISCHEMIC SPOTS ON THE 24-HOUR CONTROL POST PROCEDURE. THERE WAS NO TREATMENT OR OTHER ACTION TAKEN. THE EVENT WAS RECOVERED/RESOLVED ON (B)(6) 2025. THE EVENT WAS NOT A RECURRENT OR NEW STROKE. THE EVENT WAS CAUSALLY RELATED TO THE DISEASE UNDER STUDY AND NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. THE EVENT DID NOT RESULT FROM A DEVICE DEFICIENCY. THE SITE ASSESSED THE EVENT AS NOT RELATED TO THE DEVICES OR PROCEDURE, HOWEVER, THE SPONSOR ASSESSED AS POSSIBLY RELATED TO THE PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED MORE PASSES WERE NEEDED BUT NOT ANOTHER SURGERY.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED ON (B)(6) 2024 REPEAT PULMONARY COMPLICATIONS LED TO DEATH. ASSESSED AS POSSIBLY RELATED TO STUDY PROCEDURE. ISCHEMIC SPOTS ASSESSED AS CAUSAL TO STUDY PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT PASSED AWAY AT NEUROVASCULAR INTENSIVE CARE UNIT (HEAD OF DEPARTMENT: DR. (B)(6)) AT (B)(6) HOSPITAL. THE DEATH WAS ASSESSED NOT RELATED TO THE MEDTRONIC DEVICE OR THERAPY, CAUSAL RELATIONSHIP TO AN UNDERLYING CONDITION OR DISEASE. THE PATIENT WAS ADMITTED ON (B)(6) 2025 WITH AN EXTENSIVE LEFT-SIDED SYLVIAN INFARCTION. DESPITE ATTEMPTED REVASCULARIZATION, HE HAD NOT RECOVERED AND RETAINED MASSIVE RIGHT HEMIPLEGIA, MUTISM, AND HAD TO BE FED BY NASOGASTRIC TUBE. PULMONARY COMPLICATIONS RECURRED IN THIS DIABETIC PATIENT WITH KNOWN CARDIOLOGICAL CONDITIONS, LEADING TO HIS DEATH ON (B)(6) 2025. THE PULMONARY COMPLICATION WAS A RESPIRATORY INFECTION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312534&pc=PNJ | 9612164-2025-00609 | 2025-02-04 05:00 | Death | MEDTRONIC IRELAND | 2025-02-05 05:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY POST LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT. NO FURTHER INFORMATION AVAILABLE AT THIS TIME.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENTS DEATH WAS NOT RELATED TO THE PERFORMANCE OF THE LEADLESS IPG. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312566&pc=OZD | 1220648-2025-26000 | 2025-01-18 05:00 | Death | ABIOMED, INC. | 2025-02-05 05:00 | OZD | IMPELLA CP | Failure to Advance | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PUMP BEING PLACED TO SUPPORT A PATIENT ADMITTED IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE PATIENT ARRESTED, VENTRICULAR FIBRILLATION ARRESTED AND HAD CARDIOPULMONARY RESUSCITATION PRIOR TO HOSPITAL ADMISSION. THE IMPELLA CP SHEATH WAS PLACED AND DURING THE DELIVERY, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION AGAIN AND WAS DEFIBRILLATED TWICE AND WENT ASYSTOLIC. THE PATIENT EXPIRED BEFORE PUMP WAS DELIVERED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, THE PATIENT EXPIRED BEFORE PUMP DELIVERED VIA THE SHEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER: SECTION: DIRECT AORTIC INSERTION: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312658&pc=QFG | 3013756811-2025-24957 | 2025-01-10 05:00 | Death | TANDEM DIABETES CARE | 2025-02-05 05:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Hypoglycemia | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6)2025. PER THE CONTACT, THE CAUSE OF DEATH WAS DIABETES-RELATED HYPOGLYCEMIA. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO ADDITIONAL INFORMATION COULD BE OBTAINED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312809&pc=BTR | MW5165649 | 2025-01-20 05:00 | Death | WELL LEAD MEDICAL CO. LTD. | 2025-02-05 05:00 | BTR | PARKER FLEX TIP ET TUBE 8 FRENCH |  | Cardiac Arrest; Low Oxygen Saturation; Diminished Pulse Pressure |  |  | 1 | Event Description: PATIENT WAS SAT UP IN BED FOR A CHEST X-RAY. ET TUBE HAD BEEN IN PLACE SINCE (B)(6) 2025. SUDDEN CHANGE IN OXYGENATION AND VENTILATION NOTED. OBVIOUS CUFF LEAK. RN AND RT INJECTED 3 CC OF AIR WITH IMPROVEMENT AND THEN RAPID SOFTENING OF THE PILOT BALLOON. RECURRENCE OF OBVIOUS CUFF LEAK. UNABLE TO MAINTAIN CUFF PRESSURE. TUBE EMERGENTLY REPLACED. AFTER REINTUBATION, CARDIAC DYSFUNCTION WITH LOSS OF PULSE. CODE CALLED. PILOT LINE WAS CUT DURING EXTUBATION AND REINTUBATION. POST PROCEDURE EXAMINATION OF THE BALLOON SHOWED BALLOON PARTIALLY INFLATED. UTILIZING NEEDLE AND SYRINGE STAFF UNABLE TO INFLATE CUFF. PRESSURE TO CUFF ON THE OUTSIDE OF THE CUFF DOES NOT CREATE A LEAK. UNABLE TO TEST PILOT LINE OR INJECTION PORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21312989&pc=CAW | 2518422-2025-101734 | 2025-01-02 05:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | CAW | EVERFLO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K061261 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING AN EVERFLO PORTABLE OXYGEN CONCENTRATOR. THE USER'S HUSBAND CALLED IN TO STATE HIS WIFE HAD PASSED AWAY AND SHE HAD AN EVERFLO DEVICE THAT HAD ORIGINALLY BELONGED TO HER MOTHER AND WAS UNSURE WHAT TO DO WITH THE DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21313272&pc=BTI | 2518422-2025-011952 | 2023-08-15 04:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | BTI | ESSENCE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K042655 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH AN ESSENCE DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS MISSING AND UNAVAILABLE FOR RETURN TO THE MANUFACTURER FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21313399&pc=BZD | 2518422-2025-011969 | 2023-11-27 05:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A PHILIPS CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS MISSING AND UNAVAILABLE FOR RETURN TO THE MANUFACTURER FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21313617&pc=NPT | 2025587-2025-00997 | 2024-05-15 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-05 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA; Muscle Weakness; Pneumonia; Dysphasia; Obstruction/Occlusion; Thrombosis/Thrombus; Swelling/ Edema | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT ON DIRECT ORAL ANTICOAGULANTS FOR PRE-EXISTING CHRONIC ATRIAL FIBRILLATION, THE PATIENT WAS FOUND AT HOME WITH RIGHT HEMIPLEGIA, APHASIA, AND CO-EXISTENCE. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL BY AMBULANCE WHERE A LEFT INTERNAL CAROTID ARTERY OCCLUSION WAS OBSERVED. SUBSEQUENTLY, A CARDIOGENIC CEREBRAL EMBOLISM WAS NOTED. THE THROMBOSIS WAS RECOVERED AND REVASCULARIZATION WAS DONE. DESPITE THESE INTERVENTIONS, THE PATIENT'S CONSCIOUSNESS DID NOT IMPROVE. EXTENSIVE CEREBRAL INFARCTION AND CEREBRAL EDEMA WAS NOTED. THE PATIENT DEVELOPED PNEUMONIA AND PASSED AWAY APPROXIMATELY 10 AFTER BEING HOSPITALIZED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21313740&pc=DSQ | 2916596-2025-00637 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE, REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21313942&pc=NPT | 2025587-2025-01001 | 2024-06-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-05 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem | Renal Failure; Obstruction/Occlusion; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DUE TO ACUTE CORONARY SYNDROME AND CONGESTIVE HEART FAILURE (CHF). PERCUTANEOUS CORONARY INTERVENTION ON RIGHT CORONARY ARTERY #2 WAS PERFORMED. DESPITE INTERVENTION, THE PATIENT'S RENAL FAILURE WORSENED. IT WAS NOTED THAT THE CHF WAS DIFFICULT TO MANAGE. THE PATIENT PASSED AWAY APPROXIMATELY 1 MONTH AND 1 WEEK LATER. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21314268&pc=KRA | 2124215-2025-01827 | 2024-12-30 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-05 05:00 | KRA | EKOS ENDOVASCULAR DEVICE, 106X12CM | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED FOLLOWING EKOS THERAPY. TWO EKOS ENDOVASCULAR DEVICES WERE SELECTED FOR USE TO TREAT A BILATERAL CASE OF PULMONARY EMBOLISM. AFTER THREE HOURS OF EKOS THERAPY, THE PATIENT PULLED OUT THE EKOS CATHETERS BY HIMSELF AND FELL UNCONSCIOUS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS ORDERED, WHICH REVEALED INTRACEREBRAL BLEEDING. THE PATIENT EXPIRED THE FOLLOWING DAY, (B)(6) 2024. THE OFFICIAL CAUSE OF DEATH HAS YET TO BE DEFINED. IT WAS UNKNOWN WHETHER THE EKOS DEVICES OR PROCEDURE CAUSED OR CONTRIBUTED TO THE PATIENT EXPIRING. BOSTON SCIENTIFIC HAS BEEN UNABLE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE PATIENT OUTCOME TO DATE, DESPITE GOOD FAITH EFFORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21314271&pc=KRA | 2124215-2025-01825 | 2024-12-30 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-05 05:00 | KRA | EKOS ENDOVASCULAR DEVICE, 106X12CM | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED FOLLOWING EKOS THERAPY. TWO EKOS ENDOVASCULAR DEVICES WERE SELECTED FOR USE TO TREAT A BILATERAL CASE OF PULMONARY EMBOLISM. AFTER THREE HOURS OF EKOS THERAPY, THE PATIENT PULLED OUT THE EKOS CATHETERS BY HIMSELF AND FELL UNCONSCIOUS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS ORDERED, WHICH REVEALED INTRACEREBRAL BLEEDING. THE PATIENT EXPIRED THE FOLLOWING DAY ON (B)(6) 2024. THE OFFICIAL CAUSE OF DEATH HAS YET TO BE DEFINED. IT WAS UNKNOWN WHETHER THE EKOS DEVICES OR PROCEDURE CAUSED OR CONTRIBUTED TO THE PATIENT EXPIRING. BOSTON SCIENTIFIC HAS BEEN UNABLE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE PATIENT OUTCOME TO DATE, DESPITE GOOD FAITH EFFORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21314319&pc=DSQ | 2916596-2025-00501 | 2014-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿INTERMACS ANALYSIS OF IMPACT OF SIGNIFICANT POST-LVAD MITRAL REGURGITATION ON OUTCOMES¿ THAT HEARTMATE 3 (HM3) IMPLANT IS ASSOCIATED WITH OUTCOMES SUCH AS POST IMPLANT MITRAL VALVE REGURGITATION (PI-MR) AND DEATH. THE INTERMACS DATABASE WAS QUERIED TO IDENTIFY 7,385 PATIENTS RECEIVING PRIMARY HEARTMATE 3 IMPLANT FROM 2014-2021. PATIENTS UNDERGOING CONCOMITANT MITRAL VALVE PROCEDURES WERE EXCLUDED. SIGNIFICANT PIMR FOLLOWING LVAD IMPLANT WAS DEFINED AS MODERATE/SEVERE MITRAL REGURGITATION (MR) POST-IMPLANT ON 1 OR 3-MONTH ECHOCARDIOGRAM. THOSE NOT SURVIVING 3 MONTHS (N=682) AND/OR WITHOUT ECHOCARDIOGRAM RESULTS AVAILABLE (N=781) WERE EXCLUDED. SURVIVAL AND READMISSION ANALYSIS USING LOG-RANK STATISTICS WERE CONDITIONAL ON 3 MONTH SURVIVAL. THE STUDY COHORT CONSISTED OF 363 PATIENTS WITH SIGNIFICANT PI-MR AND 2,744 PATIENTS WITHOUT SIGNIFICANT PI-MR. OF THOSE WITH PREOPERATIVE SIGNIFICANT MR, 17% (N=292) HAD PERSISTENT SIGNIFICANT PI-MR. FOR THOSE WITH MILD OR LESS PRE-IMPLANT MR, 5% (N=71) DEVELOPED NEW SIGNIFICANT PI-MR. PATIENTS WITH SIGNIFICANT PI-MR HAD WORSE SURVIVAL AT 2 YEARS, CONDITIONAL UPON SURVIVING 3 MONTHS (80% VS 87%, P=0.024) AND FREEDOM FROM ALL-CAUSE READMISSIONS (27% VS 36%, P=0.03). SIGNIFICANT PI-MR ALSO PREDICTED WORSE CONDITIONAL MORTALITY BASED ON LOG-RANK STATISTICS (P=0.012) AND INCREASED READMISSIONS (P=0.004). THOSE PATIENTS WITH SIGNIFICANT PI-MR WERE MORE LIKELY TO HAVE RECEIVED CONCOMITANT TRICUSPID VALVE SURGERY (17% VS 10%, P < 0.001). SIGNIFICANT RESIDUAL MR FOLLOWING LVAD ADVERSELY IMPACTS 2-YEAR SURVIVAL AND ALL-CAUSE READMISSIONS. THESE DATA SUGGEST THAT STRATEGIES TO REDUCE THE OCCURRENCE OF PI-MR MAY IMPROVE DURABLE LVAD OUTCOMES. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2014 AS THE IMPLANT DATES OCCURRED BETWEEN 2014-2021. AUTHOR INFORMATION: PIENTA, M., PEGUES, J., CASCINO, T., COWGER, J., ROSENBAUM, A., HAWKINS, R., COLVIN, M., AARONSON, K., YANG, J., LIKOSKY, D., PAGANI, F., & TANG, P. (2024). INTERMACS ANALYSIS OF IMPACT OF SIGNIFICANT POST-LVAD MITRAL REGURGITATION ON OUTCOMES. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S136¿S137. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.277. CARDIAC SURGERY, UNIVERSITY OF MICHIGAN- ANN ARBOR, ANN ARBOR, MI; UNIVERSITY OF MICHIGAN, ANN ARBOR, MI; INTERNAL MEDICINE, HENRY FORD HOSPITALS, DETROIT, MI; 4INTERNAL MEDICINE, MAYO CLINIC, ROCHESTER, MN; UNIVERSITY OF MICHIGAN-ANN ARBOR, ANN ARBOR, MI; 6UNIVERSITY OF MICHIGAN HEALTH SYSTEM, ANN ARBOR, MN; INTERNAL MEDICINE, UNIV OF MICHIGAN MED CTR, ANN ARBOR, MI; CARDIAC SURGERY, UNIVERSITY OF MICHIGAN, ANN ARBOR, MI; MAYO CLINIC, ROCHESTER, MN. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21314561&pc=FNM | 0001831750-2025-00327 | 2023-08-23 04:00 | Death | STRYKER MEDICAL-KALAMAZOO | 2025-02-05 05:00 | FNM | ISOAIR DOM 84 IN TYPE B PLUG | Adverse Event Without Identified Device or Use Problem | Pressure Sores |  |  | 1 | Event Description: IT WAS REPORTED THROUGH A CLINICAL SURVEY THAT A 72-YEAR-OLD, OVERWEIGHT, FEMALE PATIENT WAS ADMITTED TO THE HOSPITAL WITH UROSEPSIS AND A STAGE 2 PRESSURE INJURY INVOLVING THE SACRAL AREA. THE PATIENT PRESENTED WITH A BRADEN SCORE OF 6. DURING THE TREATMENT PERIOD, THE PRESSURE INJURY DID NOT IMPROVE AND PROGRESSED TO A STAGE 3 INJURY. ADDITIONALLY, TWO NEW STAGE 2 PRESSURE INJURIES DEVELOPED IN THE ISCHIUM AND HEELS, REPORTEDLY CONTRIBUTING TO A SIGNIFICANT DETERIORATION IN THE PATIENT¿S HEALTH, RESULTING IN PROLONGED HOSPITALIZATION AND ULTIMATELY, DEATH. THE SURVEY RESPONDENT NOTED THAT THERE WAS NO DEVICE MALFUNCTION, AND THE PRODUCT WAS BEING USED IN ACCORDANCE WITH THE INDICATIONS FOR USE. SEVERAL COMORBID CONDITIONS AND MEDICATIONS ARE NOTED WHICH MAY IMPEDE WOUND HEALING. GIVEN THE PREVIOUS INFORMATION AND THE MULTI-FACTORIAL NATURE OF PRESSURE INJURY DEVELOPMENT IT CANNOT BE CONCLUSIVELY DETERMINED THAT THE PRODUCT HAD A CAUSAL RELATIONSHIP OR CONTRIBUTION TO THE PATIENT¿S OUTCOME. NONETHELESS, WE ARE CHOOSING TO FILE THIS EVENT OUT OF AN ABUNDANCE OF CAUTION. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL ATTEMPTS ARE BEING MADE TO OBTAIN MORE INFORMATION.Event Description: NO NEW INFORMATION. Manufacturer Narrative: MULTIPLE ATTEMPTS HAVE BEEN MADE TO GATHER FURTHER INFORMATION, WITH NO RESPONSE RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21315272&pc=MVK | 3008642652-2025-01348 | 2025-01-06 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-05 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 11:50:29 ON (B)(6) 2025. AT 17:33:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 17:34:27, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 17:35:34, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY. AT 17:36:02, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 17:36:34, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION ARTIFACT POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 17:45:39 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21315280&pc=MVK | 3008642652-2025-01349 | 2025-01-04 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-05 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 09:12:42 ON (B)(6) 2025. AT 11:27:46, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM WITH VARYING RATES/AMPLITUDE. THE DEVICE PROPERLY DETECTED VT. VARYING RATES/AMPLITUDES PREVENTED THE DEVICE FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 11:31:14 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21315307&pc=MVK | 3008642652-2025-01350 | 2025-01-04 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-05 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 12:03:46 ON (B)(6) 2025. AT 18:32:50 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 19:12:57 ON (B)(6) 2025. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21315511&pc=NPT | 2025587-2025-01015 | 2024-08-16 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-05 05:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED APPROXIMATELY 9 YEARS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED AT AN INPATIENT HOSPICE FACILITY. THE CAUSE OF DEATH WAS NOT KNOWN. THE PHYSICIAN REPORTED THE EVENT WAS NOT RELATED TO THE MEDTRONIC PRODUCTS OR THE IMPLANT PROCEDURE. HOWEVER, IT WAS ALSO REPORTED THE DEATH WAS RELATED. AN AUTOPSY WAS NOT PERFORMED. AS REPORTED, NO FURTHER RECORDS OR INFORMATION WERE AVAILABLE.Event Description: ADDITIONAL INFORMATION RECEIVED, CONFIRMED THE PHYSICIAN INDICATED THE DEATH WAS NOT RELATED TO THE MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. UPDATED DATA: B5, H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21315785&pc=DSQ | 2916596-2025-00568 | 2025-01-20 05:00 | Death | THORATEC CORPORATION | 2025-02-05 05:00 | DSQ | HEARTMATE II LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Abscess; Fever; Multiple Organ Failure; Drug Resistant Bacterial Infection | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. IT WAS NOT PROVIDED WHETHER THIS WAS DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE CAUSE OF DEATH WAS A CHRONIC MRSA INFECTION OF THE PUMP CAUSING CONTINUING WASTING AND MULTIORGAN FAILURE AT THE END. THE INFECTION WAS RELATED TO THE DEVICE. THE DEVICE OPERATED AS EXPECTED.Event Description: THE APPROXIMATE ONSET DATE OF INFECTION WAS (B)(6) 2023 (MFR #: 2916596-2025-02115). THE PATIENT HAD SYMPTOMS OF FEVERS AND ABSCESS AT THE DRIVELINE EXIT SITE. THE CAUSAL PATHOGEN WAS METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA). THE INFECTION WAS TREATED WITH REPEATED SURGICAL REVISIONS AND ANTIBIOTICS WERE GIVEN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21316024&pc=CBK | 2518422-2025-101766 | 2025-01-31 05:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT ON VENTILATION THERAPY USING A TRILOGY EVO VENTILATOR HAS DIED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE CUSTOMER REPORTED THE PATIENT WENT TO BED ON (B)(6) 2025 AT 22:00. AT 01:00 ON (B)(6) 2025 THE OBSTRUCTION ALARM OCCURRED, AND THE PATIENT WAS TRANSPORTED BY AMBULANCE DUE TO THEIR BREATHS STOPPING. THE PATIENT PASSED AWAY DUE TO RESPIRATORY ARREST. ON (B)(6) AT AROUND 08:30, THE POLICE DEPARTMENT'S CRIMINAL INVESTIGATION ACTION DIVISION CONTACTED THE CALL CENTER. THEY INQUIRED ABOUT THE CONDITIONS UNDER WHICH THE OBSTRUCTION ALARM WAS TRIGGERED AT THE TIME OF THE PATIENT¿S DEATH. THE BRANCH MANAGER RESPONDED TO THEM THAT THE OBSTRUCTION COULD HAVE BEEN CAUSED BY ANY OF THE FOLLOWING: BLOCKED EXHALATION PORT, NO POSTMORTEM RESPIRATION OR MUCUS PLUGGING IN THE CANNULA. THIS IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE BRANCH MANAGER RESPONDED TO THEM THAT THE OBSTRUCTION COULD HAVE BEEN CAUSED BY ANY OF THE FOLLOWING. BLOCKED EXHALATION PORT, NO POSTMORTEM RESPIRATION OR MUCUS PLUGGING IN THE CANNULA. LISTED AS THE REASON FOR USING THE DEVICE IS INTRACTABLE EPILEPSY, SEVERE RESPIRATORY FAILURE. FURTHER GOOD FAITH EFFORTS AND INFORMATION GATHERING ARE REQUIRED TO DISPOSITION DEVICE CAUSE AND/OR CONTRIBUTION TO THE PATIENT HARMS INCURRED. THE PATIENT PASSED AWAY DUE TO RESPIRATORY ARREST WAS REVIEWED BY THE PMS CLINICAL EXPERT. THIS EVENT IS ASSESSED AS POSSIBLY RELATED TO THE PRODUCT PROBLEM WITH THE DEVICE, USE ERROR, USE OF THE DEVICE. Manufacturer Narrative: THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER AND EVALUATION AND TESTING COMPLETED. THE REPORTED ISSUE WAS NOT DUPLICATED DURING AN OPERATIONAL CHECK. UPON CHECKING THE ERROR LOG, THE RECORD OF "OBSTRUCTION" ALARM OCCURRING WAS OBSERVED, HOWEVER, THERE WERE NO ERRORS INDICATING ABNORMALITIES OF THE DEVICE. TESTING AND INSPECTION WERE COMPLETED AND CONFIRMED THAT THERE WERE NO ABNORMALITIES. IT WAS DETERMINED THERE WERE NO ABNORMALITIES OF THE DEVICE AND THE ISSUE WAS CAUSED BY AN EXTERNAL FACTOR. FINAL TESTING, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING CONFIRMED THE UNIT OPERATED PROPERLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21318085&pc=DXF | 2124215-2025-06203 | 2025-01-14 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-05 05:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion; Retroperitoneal Hemorrhage | K150709 |  | 1 | Event Description: A PERICARDIAL EFFUSION (PE) OCCURRED. A PATIENT DEATH WAS REPORTED. IT HAS BEEN REPORTED THAT A VERSACROSS CONNECT LAAC ACCESS SOLUTION KIT WAS SELECTED FOR USE FOR A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE. THE PHYSICIAN PERFORMED THE TRANSSEPTAL PUNCTURE. THERE WERE NO DIFFICULTIES WITH TRANSSEPTAL PUNCTURE. RF WAS APPLIED TWICE WITH NO EXCESSIVE FORCE, AND NO DAMAGE NOTED UPON REMOVAL. FOLLOWING TRANSSEPTAL, THE PHYSICIAN NOTICED A DECREASE IN THE SYSTOLIC BLOOD PRESSURE. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SCAN REVEALED A CIRCUMFERENTIAL PE. PERICARDIOCENTESIS WAS THEN PERFORMED TO MANAGE THE EFFUSION. PATIENT'S VITALS WERE THEN MONITORED AND THE PHYSICIAN DECIDED TO MOVE FORWARD WITH THE WATCHMAN PLACEMENT. PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY RE-OCCURRING EFFUSIONS AND A PERICARDIAL DRAIN WAS LEFT IN PLACE. POST-PROCEDURE, IT WAS NOTED THAT THE PATIENT'S DRAIN HAD CLOTTED OFF. THIS RESULTED IN THE PATIENT BEING BROUGHT BACK INTO THE LAB WHERE A LARGER PERICARDIAL DRAIN WAS PLACED TO DRAIN ADDITIONAL FLUID. THE PATIENT REMAINS IN STABLE CONDITION AND IT WAS BEING MONITORED OVERNIGHT. THE PATIENT PASSED AWAY OVERNIGHT, WITHIN 24HRS. THE PHYSICIAN MENTIONED THAT THE AUTOPSY REVEALED CAUSE OF DEATH TO BE RETROPERITIONEAL BLEED. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IT WAS MENTIONED THAT THE PATIENT HAD A 40 YEAR HISTORY OF IMMUNOSUPPRESSION, AND TWO RENAL TRANSPLANTS PERFORMED. A LARGE PERICARDIAL EFFUSION RESULTING IN HEMODYNAMIC COMPROMISE WAS REPORTED. THE PHYSICIAN DOES NOT BELIEVE THE VERSACROSS DEVICES MALFUNCTION OR CONTRIBUTED TO THE ADVERSE EVENT. HOWEVER, THE PHYSICIAN CONFIRMED THAT THE PATIENT HAS SINCE BEEN PRONOUNCED DECEASED. THE BLEEDING FROM THE EFFUSION DID STOP BUT THEY WERE NOT SURE IF IT WAS DUE TO A CLOT COMPRESSING THE HEART. THE PATIENT WAS ALSO SAID TO HAVE ALL KINDS OF OTHER ISSUES WITH SEVERE HYPERKALEMIA AND THEN ASPIRATED.Event Description: A PERICARDIAL EFFUSION (PE) OCCURRED. A PATIENT DEATH WAS REPORTED. IT HAS BEEN REPORTED THAT A VERSACROSS CONNECT LAAC ACCESS SOLUTION KIT WAS SELECTED FOR USE FOR A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE. THE PHYSICIAN PERFORMED THE TRANSSEPTAL PUNCTURE. THERE WERE NO DIFFICULTIES WITH TRANSSEPTAL PUNCTURE. RF WAS APPLIED TWICE WITH NO EXCESSIVE FORCE, AND NO DAMAGE NOTED UPON REMOVAL. FOLLOWING TRANSSEPTAL, THE PHYSICIAN NOTICED A DECREASE IN THE SYSTOLIC BLOOD PRESSURE. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SCAN REVEALED A CIRCUMFERENTIAL PE. PERICARDIOCENTESIS WAS THEN PERFORMED TO MANAGE THE EFFUSION. PATIENT'S VITALS WERE THEN MONITORED AND THE PHYSICIAN DECIDED TO MOVE FORWARD WITH THE WATCHMAN PLACEMENT. PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY RE-OCCURRING EFFUSIONS AND A PERICARDIAL DRAIN WAS LEFT IN PLACE. POST-PROCEDURE, IT WAS NOTED THAT THE PATIENT'S DRAIN HAD CLOTTED OFF. THIS RESULTED IN THE PATIENT BEING BROUGHT BACK INTO THE LAB WHERE A LARGER PERICARDIAL DRAIN WAS PLACED TO DRAIN ADDITIONAL FLUID. THE PATIENT REMAINS IN STABLE CONDITION AND IT WAS BEING MONITORED OVERNIGHT. THE PATIENT PASSED AWAY OVERNIGHT, WITHIN 24HRS. THE PHYSICIAN MENTIONED THAT THE AUTOPSY REVEALED CAUSE OF DEATH TO BE RETROPERITONEAL BLEED. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IT WAS MENTIONED THAT THE PATIENT HAD A 40 YEAR HISTORY OF IMMUNOSUPPRESSION, AND TWO RENAL TRANSPLANTS PERFORMED. A LARGE PERICARDIAL EFFUSION RESULTING IN HEMODYNAMIC COMPROMISE WAS REPORTED. THE PHYSICIAN DOES NOT BELIEVE THE VERSACROSS DEVICES MALFUNCTION OR CONTRIBUTED TO THE ADVERSE EVENT. HOWEVER, THE PHYSICIAN CONFIRMED THAT THE PATIENT HAS SINCE BEEN PRONOUNCED DECEASED. THE BLEEDING FROM THE EFFUSION DID STOP BUT THEY WERE NOT SURE IF IT WAS DUE TO A CLOT COMPRESSING THE HEART. THE PATIENT WAS ALSO SAID TO HAVE ALL KINDS OF OTHER ISSUES WITH SEVERE HYPERKALEMIA AND THEN ASPIRATED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE SINCE IT WAS DISPOSED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED TO CORRECT THE INITIAL MDR IN BLOCK(S) PATIENT CODES (F10 AND H6), IN SECTIONS F - FOR USE BY UF/IMPORTER AND H - DEVICE MANUFACTURERS ONLY. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE SINCE IT WAS DISPOSED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21318491&pc=LGW | 3006630150-2025-00454 | 2025-01-22 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-02-05 05:00 | LGW | STIMULATOR, SPINAL-CORD, TOTALLY IMPLANTED FOR PAIN RELIEF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P030017 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PREPARATION FOR THE TRIAL PHASE OF THE SPINAL CORD STIMULATOR (SCS) IMPLANT PROCEDURE THE PATIENT CODED ON THE TABLE, THE PATIENT WAS UNDER LOCAL ANESTHESIA HOWEVER THIS WAS PRIOR TO ANY OF THE DEVICES BEING INTRODUCED. THE PROCEDURE WAS ABORTED AND THE PATIENT TRANSPORTED TO VIA AMBULANCE. IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY. IT WAS STATED THAT THE PATIENT DID HAVE PREEXISTING CARDIAC ISSUES. NO OTHER INFORMATION HAS BEEN PROVIDED TO THE MANUFACTURER. Manufacturer Narrative: REPORT WAS FILED ON AN UNK PRODUCT AND SPECIFIC PRODUCT IDENTIFIERS ARE MISSING SUCH AS SUSPECT MEDICAL DEVICE BATCH, LOT, SERIAL NUMBER, MANUFACTURING SITE INFORMATION AS THERE WAS NO DEVICE INVOLVED IN THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21318605&pc=KRD | 9615728-2025-00004 | 2022-05-01 04:00 | Death | MERIT MEDICAL ROISSY-EN-FRANCE | 2025-02-05 05:00 | KRD | HEPASPHERE¿.30-60 ¿M.25 MG.VIAL.STERILE.GAMMA. | Adverse Event Without Identified Device or Use Problem | Liver Damage/Dysfunction |  |  | 1 | Event Description: ON JANUARY 22, 2025, A CER LITERATURE SEARCH FOR AN IMPLANTABLE EMBOLIC WAS COMPLETED. THE CASE ALLEGES THAT AN 82-YEAR-OLD WOMAN WITH A LARGE HEPATOCELLULAR CARCINOMA (HCC) WHO UNDERWENT ELECTIVE TACE DUE TO THE HIGH SURGICAL RISK ASSOCIATED WITH HER TUMOR SIZE IN (B)(6) 2022. THE PATIENT UNEXPECTEDLY EXPERIENCED LIVER RUPTURE 20 HOURS POST-PROCEDURE, LEADING TO ACUTE SURGICAL INTERVENTION, AND EVEN THOUGH THE BLEEDING WAS SUCCESSFULLY STOPPED, THE PATIENT DEVELOPED PROGRESSIVE MULTI-ORGAN FAILURE AND EXPIRED A FEW DAYS LATER. PÉ]OVÁ, M., BENKO, [ A RARE EARLY-ONSET FATAL COMPLICATION AFTER TRANSARTERIAL CHEMOEMBOLIZATION: A CASE REPORT AND REVIEW OF THE LITERATURE. CURRENT ONCOLOGY, 31(4), 1961-1970.] Manufacturer Narrative: CER ARTICLE / LITERATURE SEARCH FOR AN EMBOLIC WAS COMPLETED. THE SUSPECT MEDICAL DEVICE IS NOT EXPECTED TO RETURN FOR EVALUATION. THE COMPLAINT COULD NOT BE CONFIRMED. THE ROOT CAUSE COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE HISTORY AND COMPLAINT DATABASE COULD NOT BE PERFORMED SINCE THE LOT NUMBER WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21318635&pc=BZD | 2518422-2025-101780 | 2024-10-29 04:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN EMAIL RECEIVED BY A CUSTOMER REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21318764&pc=OZO | 2032227-2025-121516 | 2022-06-16 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-05 05:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2022. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS HYPERGLYCEMIA AND DIABETES KETOACIDOSIS. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2022, PRIOR TO PASSING. THE INSULIN PUMP WILL NOT BE RETURNED FOR PRODUCT ANALYSIS. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21318775&pc=BZD | 2518422-2025-012154 | 2024-06-18 04:00 | Death | RESPIRONICS, INC. | 2025-02-05 05:00 | BZD | REMSTAR AUTO | Degraded | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO ALLEGATION REPORTED BY THE PATIENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS NO ALLEGATION REPORTED BY THE PATIENT. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED THAT PATIENT HAS PASSED AWAY AND HIS DEVICE DISCARDED. PLEASE TAKE HIM OFF YOUR CONTACT LISTS. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. AFTER DEVICE RETURN AND COMPLETION OF EVALUATION, A FOLLOW-UP REPORT WILL BE FILED. BOX B1 WAS UPDATE TO BOTH. (PREVIOUS IT WAS PRODUCT PROBLEM). BOX B2 WAS UPDATED TO DEATH. (PREVIOUSLY IT WAS BLANK). BOX H1 WAS CHANGED FROM PRODUCT PROBLEM TO DEATH. BOX H6- PATIENT OUTCOME CODE GRID, HEALTH IMPACT CODE, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID WAS UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21301755&pc=NVZ | 3008973940-2025-00811 |  | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-02-04 05:00 | NVZ | ADAPTA | Appropriate Term/Code Not Available | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) WAS EXPLANTED FOR AN UNKNOWN REASON. FOLLOW-UP WAS ATTEMPTED TO OBTAIN THE RELEVANT INFORMATION; HOWEVER, NO ADDITIONAL DETAILS ARE AVAILABLE AT THIS TIME. THIS EVENT IS BEING CONSERVATIVELY REPORTED IN AN ABUNDANCE OF CAUTION IN RESPONSE TO THE ADVISORY DESCRIBING THE POTENTIAL FOR DEVICE CIRCUIT ERROR TO OCCUR WHILE THE DEVICE IS PROCESSING AN ATRIAL-SENSED EPISODE.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTION: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 1882TC COMPETITOR LEAD; THERAPY DATE: (B)(6) 2009. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21302509&pc=KDI | 3003464075-2025-00036 | 2025-01-11 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-02-04 05:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | K170469 |  | 1 | Event Description: A REPORT WAS RECEIVED ON 14 JAN 2025 FROM THE CAREGIVER OF A 68 YEAR OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING HISTORY OF HEART ATTACK, MULTI VESSEL CORONARY ARTERY DISEASE, HYPERTENSION AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT LOST BLOOD (AMOUNT AND SOURCE OF BLOOD LOSS NOT PROVIDED) AND EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 20 JAN 2025 FROM THE HOME THERAPY NURSE (HTN) WHO STATED EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, THEY ARRIVED ON SCENE AND MEDICAL INTERVENTION WAS ATTEMPTED (NOS). NO FURTHER DETAILS OF THE EVENT WERE PROVIDED. Manufacturer Narrative: THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21302524&pc=DSP | 2249723-2025-0000523 | 2025-01-16 05:00 | Death | DATASCOPE CORP. | 2025-02-04 05:00 | DSP | CARDIOSAVE HYBRID W/ E/F PLUG | Electrical /Electronic Property Problem; Overheating of Device | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K112372 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025 CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) WAS USED ON A PATIENT IMMEDIATELY AFTER THE BATTERY MAINTENANCE PROGRAM WAS COMPLETED. DEVICE HAD AN 'OVERTEMPERATURE IN THE SYSTEM' ALARM AND ERROR CODE #3, HOWEVER; ON (B)(6) 2025 DEVICE SHOWED AN ALARM MESSAGE OF "OVERTEMPERATURE IN THE SYSTEM". THE USER ENDED THE IABP THERAPY BECAUSE THE PATIENT'S OVERALL THERAPY HAD BEEN STOPPED. THE PATIENT HAS SINCE DIED, BUT ACCORDING TO THE USER, THERE IS NO DIRECT CONNECTION TO THE IABP ALARM SITUATIONS. Manufacturer Narrative: DUE TO CHARACTER LIMITATION, BELOW FIELD ARE ADDED FROM E. EVENT SITE TELEPHONE- (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS - B4, G3, G6, H2, H3, H6 ( TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSION), H11. A GETINGE FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO EVALUATE THE IABP UNIT AND TROUBLESHOOTING REVEALED A LEAKY DRIVE MANIFOLD (POWER-UP TEST FAILS CODE #___ #3 (RESERVOIR VACUUM)) AND A BLOCKED CASE FAN IN THE DEVICE. DRIVE MANIFOLD RENEWED. CASE FAN CLEANED AND BLOCKAGE CAUSED BY INTERNAL CABLE REMOVED. THE TWO PROBLEMS TOGETHER MAY HAVE CAUSED THE OVERHEATING. FUNCTIONAL TEST, INSPECTION AND TEST ENDURANCE RUN CARRIED OUT, DEVICE OK. DEVICE IS READY FOR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21302711&pc=LJT | 3006260740-2025-00310 | 2025-01-29 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-02-04 05:00 | LJT | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED INFECTION AND LATER PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6)1900 FOR THE MDR SUBMISSION REQUIREMENT. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INFECTION AND PATIENT DEATH, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED INFECTION AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3. H11: D4 (UNIQUE IDENTIFIER (UDI)#). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INFECTION AND PATIENT DEATH, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED INFECTION AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3, H6 (COMPONENT, METHOD). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21302868&pc=MHX | 9610816-2025-000097 | 2025-01-02 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-02-04 05:00 | MHX | INTELLIVUE MP5 | Defective Alarm | Insufficient Information | K150310 |  | 1 | Event Description: IT WAS REPORTED ST ANALYSIS NEEDED TO BE ADJUSTED FOR EACH NEW EMERGENCY ROOM PATIENT; HOWEVER, CONSULTATION WITH BIOMED INDICATED ST ANALYSIS DEFAULT WAS SET TO OFF. A PATIENT DEATH OCCURRED AFTER ST ELEVATION WAS NOTICED BY STAFF, WHICH MAY HAVE BEEN RELATED TO NOT BEING ABLE TO HEAR ALARMS. IT REMAINS UNKNOWN WHETHER ST ANALYSIS WAS TURNED ON AT THE TIME OF THE EVENT. THE DEVICE WAS IN USE ON A PATIENT. THERE WAS A REPORT OF PATIENT OR USER HARM. Manufacturer Narrative: ADDITIONAL INFORMATION WAS PROVIDED INDICATING THERE WAS NO ALLEGATION OF HARM RELATED TO THE DEVICE. THE PATIENT DEATH WAS REVEALED IN A CONVERSATION AS THE REASON FOR WANTING TO MAKE THE CONFIGURATION CHANGE. GOOD FAITH EFFORT WAS MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE EVENT DETAILS; HOWEVER, NO RESPONSE WAS RECEIVED FROM THE CUSTOMER. ONSITE SERVICE WAS CANCELED WITH NOTES INDICATING IT WAS NOT NEEDED; THEREFORE, ADDITIONAL FUNCTIONAL TESTING WAS NOT ABLE TO BE PERFORMED. BASED ON THE INFORMATION AVAILABLE, THE CAUSE OF THE REPORTED PROBLEM MAY HAVE BEEN RELATED TO SETTINGS/CONFIGURATION; HOWEVER, THIS WAS NOT ABLE TO BE VERIFIED. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE ENGINEER PROVIDED THEIR ANALYSIS FINDINGS HOWEVER WE ARE UNABLE TO CONFIRM THE FINAL DISPOSITION OF THE DEVICE BECAUSE THE CUSTOMER DID NOT RESPOND TO REQUESTS FOR ADDITIONAL INFORMATION. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. Manufacturer Narrative: A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21302952&pc=DSQ | 2916596-2025-00661 | 2021-03-01 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿EFFICACY AND SAFETY OF SGLT2 INHIBITORS IN PATIENTS WITH A LVAD¿ THAT HEARTMATE 3 MAY BE ASSOCIATED WITH WORSENING HEART FAILURE EVENTS AND DEATH. THE RETROSPECTIVE STUDY EVALUATED 66 LVAD PATIENTS INITIATED ON A SODIUM-GLUCOSE COTRANSPORTER 2 (SGLT2) INHIBITOR FROM MARCH 2021 TO JULY 2022. MOST PATIENTS WERE BLACK (64%) MALES (90%) SUPPORTED BY A HEARTMATE3 DEVICE (73%). SIMILAR TO BENCHMARKS, A TOTAL OF 23 PATIENTS (34.8%) EXPERIENCED A WORSENING HEART FAILURE EVENT, INCLUDING 4 DEATHS, WITH A MEDIAN TIME TO EVENT AT 90 DAYS AFTER SGLT2 INITIATION. EIGHTEEN PATIENTS (27.3%) DISCONTINUED THERAPY MOST COMMONLY DUE TO HYPOVOLEMIA, DIZZINESS, OR HYPOTENSION. ADDITIONALLY, MEAN LOOP DIURETIC DOSE DECREASED BY AN AVERAGE OF 50MG FUROSEMIDE EQUIVALENTS AT 1 YEAR. MEAN ATRIAL PRESSURE (MAP) WAS SIGNIFICANTLY DECREASED FROM BASELINE TO 1 YEAR (P < 0.05). SGLT2 INHIBITORS WERE WELL TOLERATED, REDUCED BLOOD PRESSURE AND DIURETIC BURDEN, IMPROVED GUIDELINE-DIRECTED MEDICAL THERAPY (GDMT) UTILIZATION, AND DID NOT INCREASE HEART FAILURE HOSPITALIZATIONS COMPARED TO NATIONAL AVERAGES. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE SAME AS PUBLISHED DATE SINCE DATE OF DATA COLLECTION WAS NOT PROVIDED. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. PREHEIM, E., STRAW, L., & MARDIS, A. (2024). EFFICACY AND SAFETY OF SGLT2 INHIBITORS IN PATIENTS WITH A LVAD. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S647. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1027 PRISMA HEALTH, COLUMBIA, SC. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, WERE NOT PROVIDED. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303071&pc=MVK | 3008642652-2025-01252 | 2024-12-26 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-04 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 09:56:05 ON (B)(6) 2024. AT 05:52:47 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 40 BPM WITH MOTION ARTIFACT. AT 05:53:39, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS IDIOVENTRICULAR RHYTHM @ 30 BPM WITH PVC¿S AND MOTION ARTIFACT. THE POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM WITH MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 06:21:54 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303160&pc=DSQ | 2916596-2025-00613 | 2025-01-19 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Fall; Hemorrhage/Bleeding | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY WITH A HEAD BLEED AFTER A FALL ON (B)(6) 2025. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE FALL WAS NOT DEVICE RELATED. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED AS THE PATIENT WAS AT GOAL AT THE TIME. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED BLEED AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS, INCLUDING BLEEDING AND DEATH. THE IFU, UNDER ¿ANTICOAGULATION¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303265&pc=DSQ | 2916596-2025-00559 | 2017-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT HEARTMATE 3 (HM3) IMPLANT WAS ASSOCIATED WITH OUTCOMES SUCH AS RIGHT HEART FAILURE, HEMODYNAMIC INSTABILITY, TRICUSPID REGURGITATION AND DEATH. THE ROLE AND BENEFIT OF CONCOMITANT TRICUSPID VALVE REPAIR (TVR) DURING VENTRICULAR ASSIST DEVICE (VAD) IMPLANT REMAINED UNCLEAR. PREVIOUSLY PUBLISHED DATA UTILIZED HISTORICAL VADS. OUTCOMES OF TVR WITH CONTEMPORARY LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT REMAINED UNKNOWN. A CENTER¿S OUTCOMES WERE ANALYZED WITH A FOCUS ON THE ROLE OF TVR IN MITIGATING RIGHT VENTRICLE (RV) DYSFUNCTION AND IMPROVING LEFT VENTRICLE (LV) UNLOADING POST LVAD. A SINGLE CENTER, RETROSPECTIVE ANALYSIS OF CONSECUTIVE PATIENTS (2017-2022) IMPLANTED WITH A HEARTMATE3 LVAD WAS PERFORMED. THE SURGICAL STRATEGY WAS TO PERFORM TRICUSPID VALVE (TV) ANNULOPLASTY WHEN EITHER OF 2 CONDITIONS WERE MET: GREATER THAN MODERATE TR, OR TV ANNULAR DIAMETER OF GREATER THAN 4 CM. BASELINE DATA WERE COMPARED TO 1 YEAR DATA. PATIENTS RECEIVING TVR HAD A HIGHER BASELINE RISK PROFILE FOR RIGHT VENTRICULAR FAILURE BY RIGHT HEART CATHETERIZATION AND ECHOCARDIOGRAPHY. RIGHT VENTRICULAR ASSIST DEVICE (RVAD) UTILIZATION WAS INFREQUENT, WITH A TREND TOWARD HIGHER USE IN THE NO TVR GROUP. POST OPERATIVE INOTROPES USE WAS SIMILAR BETWEEN GROUPS, WITH LONGER DURATION OF USE IN THOSE WITHOUT TVR. SURVIVAL AT 1 YEAR WAS SIMILAR BETWEEN GROUPS, AS WERE HEMODYNAMIC PROFILES. LEFT VENTRICULAR DIMENSION SHOWED GREATER IMPROVEMENT IN PATIENTS WITHOUT TVR. IN CONTEMPORARY PATIENTS RECEIVING VAD THERAPY, ALGORITHMIC USE OF CONCOMITANT TVR APPEARS TO COMPENSATE FOR INCREASED BASELINE RISK, SUGGESTING EITHER A SALUTARY EFFECT OF TVR, OR ALLOWING MORE OPTIMAL POST IMPLANT MANAGEMENT. Manufacturer Narrative: SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT PATIENT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. BRANDT, A., CHUNG, E., EMANI, S., MEHAL, A., STOKES, J., GUPTA, S., BARTONE, C., ANSWINI, G., & EGNACZYK, G. (2024). OUTCOMES OF CONCOMITANT TRICUSPID VALVE REPAIR WITH CONTEMPORARY VENTRICULAR ASSIST DEVICE IMPLANT. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S288. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1236. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303306&pc=OZD | 1220648-2025-25974 | 2025-01-17 05:00 | Death | ABIOMED, INC. | 2025-02-04 05:00 | OZD | IMPELLA CP | Infusion or Flow Problem | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA CP SUPPORT, THERE WAS A PRESSURE SATURATION LOW ALARM. INSTRUCTIONS WERE GIVEN TO OPTIMIZE POSITION AND PARALLEL DURING CALL COULD SEE SATURATED FLOW AND INCREASING MC >1050MA. THE DOCTOR WAS ADVISED TO ESCALATE TO IMPELLA 5.5 SUPPORT DUE TO A RISK OF A PUMP STOP. THE PATIENT SUBSEQUENTLY PASSED AWAY ON SUPPORT THE FOLLOWING DAY. IT IS UNKNOWN IF THE COMPLICATION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF SATURATED FLOW HAS BEEN COMPLETED. THE PRODUCT WAS NOT RETURNED FOR EVALUATION. DATA LOGS WERE REVIEWED AND THE LOGS OF FULL CASE SHOW AT TIME OF ¿PLACEMENT SIGNAL LOW¿ ALARMS, MOTOR CURRENT IS VARIABLE AT TIME OF ALARMS WITH MULTIPLE SPIKES IN PLACEMENT SIGNAL. AT BEGINNING OF CASE, SATURATED FLOW IS OBSERVED WITH A MAX FLOW OF APPROXIMATELY 4.3 WHICH IS HIGHER THAN THE SET P-LEVEL RANGE (P-9) AT THAT TIME. REAL TIME (RT) LOG AT TIME OF ISSUE SHOWS MULTIPLE SPIKES IN PLACEMENT SIGNAL WITH ¿IMPELLA IN AORTA¿ AND ¿IMPELLA IN VENTRICLE¿ ALARMS PRIOR TO ¿PLACEMENT SIGNAL LOW¿ ALARMS. AT BEGINNING OF ALARMS OCCURRING, PUMP FLOWS BECAME LOW WITH PUMP RUNNING AT AROUND P-3. TWO LARGE MOTOR CURRENT SPIKES WITH SPEED DEVIATIONS ARE OBSERVED AT TIME WITH LOW PUMP FLOW. LOW PUMP FLOWS SEEMS TO RESOLVE SLIGHTLY AND ONCE FLOWS ARE WITHIN A NORMAL RANGE FOR THE P-LEVEL, THE FLOWS BECOME FULLY SATURATED FOR A FEW INSTANCES WITH NO PULSATILITY. CLINICAL DETAILS NOTED THAT SATURATED FLOWS WERE OBSERVED AND AN INCREASE IN MOTOR CURRENT (MC) TO GREATER THAN 1050MA WAS SEEN. PUMP WAS REPOSITIONED FOR OPTIMAL POSITIONING. THE ROOT CAUSE OF THE SATURATED FLOW WAS MOST LIKELY DUE TO BIOMATERIAL. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303332&pc=QAU | 1037905-2025-00092 | 2024-10-01 04:00 | Death | WILSON-COOK MEDICAL INC | 2025-02-04 05:00 | QAU | HEMOSPRAY ENDOSCOPIC HEMOSTAT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation; Abdominal Distention; Gastrointestinal Hemorrhage | K200972 |  | 1 | Event Description: COOK ENDOSCOPY WAS NOTIFIED OF THIS EVENT VIA A CLINICAL LITERATURE ARTICLE. THIS ARTICLE WAS PUBLISHED IN 2024. PLEASE SEE BELOW FOR RELEVANT EXCERPTS OF THIS ARTICLE. PERFORATION DUE TO CO2 INSUFFLATION GASTROINTESTINAL PERFORATION HAS BEEN REPORTED AS A RARE ADVERSE EVENT OF TC-325 [HEMOSPRAY] USE. A 90-YEAR-OLD WOMAN WITH A HISTORY OF STROKE (ON ASPIRIN) WAS SENT TO THE EMERGENCY DEPARTMENT FROM SUB-ACUTE REHABILITATION FOR HYPOTENSION, LETHARGY, AND DARK STOOL. UPON ARRIVAL, HER VITAL SIGNS WERE TEMPERATURE 36.4C, BLOOD PRESSURE 76/45 MMHG, HEART RATE 143 BPM, OXYGEN SATURATION 94% ON A NON-REBREATHER. HER LABS SHOWED: HEMOGLOBIN 4.4 G/DL, PLATELETS 210 X 103 PLATELETS/UL, BLOOD UREA NITROGEN 51 MG/DL, CREATININE 0.64 MG/DL, PROTHROMBIN TIME 14.5 SECONDS, ACTIVATED PARTIAL THROMBOPLASTIN TIME 27.3 SECONDS. SHE RECEIVED INTRAVENOUS FLUID RESUSCITATION AND 3 UNITS OF PACKED RED BLOOD CELLS WITH STABILIZATION OF HER VITAL SIGNS. THE FOLLOWING DAY, ESOPHAGOGASTRODUODENOSCOPY SHOWED ACTIVE BLEEDING IN THE PRESENCE OF SEVERE EROSIVE ESOPHAGITIS IN THE HIATAL HERNIA SAC AND 4 NONOBSTRUCTING OOZING LINEAR GASTRIC ULCERS IN THE GASTRIC FUNDUS AND LESSER CURVATURE OF THE GASTRIC BODY. A FORREST IB LINEAR FUNDIC ULCER, 10MM IN LARGEST DIMENSION, WAS NOTED AND SUCCESSFULLY APPROXIMATED WITH 2 HEMOSTATIC CLIPS. GIVEN THE PERSISTENCE OF ACTIVE BLEEDING FROM AN UNCLEAR SOURCE THE PATIENT WAS INTUBATED FOR AIRWAY PROTECTION AND TC-325 [HEMOSPRAY} WAS APPLIED BROADLY IN THE PROXIMAL STOMACH AND HERNIA SAC. DURING THIS PHASE OF THE PROCEDURE THE PATIENT ACUTELY DEVELOPED A DISTENDED, TYMPANITIC ABDOMEN. HER VENTILATION WORSENED REQUIRING HIGHER POSITIVE END-EXPIRATORY PRESSURE, AND SHE EXPERIENCED INTRAPROCEDURAL CARDIAC ARREST. ADVANCED CARDIAC LIFE SUPPORT MEASURES WERE INITIATED; NEEDLE DECOMPRESSION WAS PERFORMED FOR TENSION CAPNOPERITONEUM WHICH IMMEDIATELY IMPROVED HER VENTILATION. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE PATIENT LATER EXPIRED AND AUTOPSY WAS DECLINED. GASTRIC PERFORATION AND TENSION CAPNOPERITONEUM FOLLOWING HEMOSTATIC POWDER SPRAY APPLICATION BRIAN MCSTEEN, MD THE AMERICAN JOURNAL OF GASTROENTEROLOGY VOLUME 119 - SUPPLEMENT - OCTOBER 2024 WWW.AMJASTRO.COM Manufacturer Narrative: INVESTIGATION EVALUATION: A PRODUCT EVALUATION WAS NOT PERFORMED IN RESPONSE TO THIS REPORT BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT PROVIDED TO COOK FOR EVALUATION. THE REPORT COULD NOT BE CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE CONDUCTED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. INVESTIGATION CONCLUSION: WE COULD NOT CONDUCT A COMPLETE INVESTIGATION BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT RETURNED FOR EVALUATION. THERE IS NO EVIDENCE TO SUGGEST THE REPORT IS DUE TO A DEVICE FAILURE. THE IFU PRECAUTIONS "ENSURE GASTROINTESTINAL LUMEN IS NOT DISTENDED BECAUSE HEMOSPRAY ADDS VOLUME IN EXCESS OF INSUFFLATION VOLUMES DURING PROCEDURE. CLOSELY MONITOR BOWEL DISTENSION AND BALANCE INSUFFLATION AND HEMOSPRAY VOLUMES AS NECESSARY." ADDITIONALLY, THE IFU PRECAUTIONS "LIKE OTHER MODALITIES, HEMOSPRAY MAY NOT BE EFFECTIVE FOR ALL TYPES OF BLEEDS. GASTROINTESTINAL BLEEDING MAY EXACERBATE EXISTING COMORBIDITIES, INCREASING THE POTENTIAL FOR ADVERSE EVENTS INCLUDING PATIENT MORTALITY." PRIOR TO DISTRIBUTION, ALL HEMOSPRAY ENDOSCOPIC HEMOSTAT ARE SUBJECTED TO A VISUAL INSPECTION TO ENSURE DEVICE INTEGRITY. CORRECTIVE ACTION: BASED ON THE QUALITY ENGINEERING RISK ASSESSMENT NO CORRECTIVE ACTION IS WARRANTED AT THIS TIME. A REVIEW OF THE COMPLAINT HISTORY WAS CONDUCTED. THIS REPORT REPRESENTS AN UNUSUAL OCCURRENCE. BASED ON THIS REVIEW, THE LIKELIHOOD OF THIS TYPE OF REPORT IS RARE. QUALITY ASSURANCE WILL CONTINUE TO MONITOR FOR COMPLAINT TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303665&pc=OZD | 1220648-2025-25979 | 2024-12-16 05:00 | Death | ABIOMED, INC. | 2025-02-04 05:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Increase in Pressure; Malposition of Device |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE) WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DURING SUPPORT POSITIONING ISSUE AND HIGH PURGE PRESSURE WERE ENCOUNTERED. THE PATIENT WOULD EVENTUALLY EXPIRE AFTER SIX DAYS OF SUPPORT. THE PATIENT WITH AN EXTENSIVE MEDICAL HISTORY WAS DEEMED HEART FAILURE WITH REDUCED EJECTION FRACTION OF 20%. AN ECHOCARDIOGRAM SHOWED THE LEFT VENTRICLE SEVERELY DILATED, GLOBAL HYPOKINESIS, AND SMALL MURAL MASS. LEFT VENTRICULAR END-DIASTOLIC DIAMETER WAS 6.6CM. THE PATIENT WITH ISCHEMIC CARDIOMYOPATHY COMPLICATED BY CARDIOGENIC SHOCK AND INOTROPE DEPENDENT WITH CLASS IV HEART FAILURE. THE PATIENT WAS PLANNED FOR AN IMPELLA 5.5 DEVICE AS A BRIDGE TO HEART TRANSPLANT OR BRIDGE TO LEFT VENTRICULAR ASSIST DEVICE. THE IMPELLA 5.5 DEVICE WAS IMPLANTED, AND POOR POSITIONING AND HIGH PURGE PRESSURE ALARM WERE NOTED TO HAVE OCCURRED APPROXIMATELY SIX DAYS AFTER START OF SUPPORT. THERE WERE ALSO PERSISTENT INTERMITTENT SUCTION ALARMS AND MOTOR CURRENT WAS FLAT, AND AORTA/LEFT VENTRICLE DECOUPLED. THE PATIENT WAS SLIDING AND IN MORE SHOCK. THE PATIENT WAS TAKEN FOR AN IMPELLA 5.5 DEVICE EXCHANGED AND VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) PLACEMENT. HOWEVER, THE PATIENT CONTINUED DOING POORLY, AND SUPPORT WAS REMOVED THE SAME DAY OF THE NEW 5.5 INSERTION. THE CAUSE OF DEATH IS UNKNOWN. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS A POSSIBLE CONTRIBUTING FACTOR IN THE OVERALL OUTCOME. HOWEVER, IT SHOULD BE NOTED THE PATIENT HAD A "POOR" HEART FUNCTION, LESS OPTIONS AVAILABLE (NOT A CANDIDATE FOR PERCUTANEOUS CORONARY INTERVENTION OR CORONARY ARTERY BYPASS GRAFT SURGERY), AND DID NOT FULLY RECOVER FROM THE CARDIAC EVENT. Manufacturer Narrative: THE IMPELLA PUMP WITH THE DATA STICK AND PURGE CASSETTE HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE FAILURE MODE "POSITIONING ISSUES" WAS CHANGED TO "LOW PUMP FLOWS" BECAUSE LOW PUMP FLOW WAS THE HIGHER SEVERITY ISSUE OF THE TWO RESULTING PUMP PERFORMANCE ISSUES OF THE SAME ROOT CAUSE. THE PUMP WAS RETURNED AND INSPECTED. THERE WAS BIOMATERIAL ON THE IMPELLER AND IN THE OUTFLOW CAGE. THE PUMP WAS SOAKED AND THE BIOMATERIAL WAS CLEARED. THE PUMP WAS RUN IN THE LAB AND THE LOW PUMP FLOW DID NOT REPRODUCE. THE DATA LOGS SHOWED A MOTOR CURRENT SPIKE CORRESPONDING WITH A SPEED DEVIATION, SUCTION ALARMS, AND A SHIFT IN PLACEMENT SIGNAL. THE LOGS SHOWED A SHIFT IN PURGE PRESSURE AT THE TIME OF THE SPIKE BEFORE THE PURGE PRESSURE VALUES ARE ELEVATED LATER ON AND HIGH PURGE PRESSURE ALARM IS TRIGGERED. THE HIGH PURGE PRESSURE RESOLVED AFTER ABOUT 40 MINUTES. THE POSITIONING ISSUES REPORTED BY CLINICAL ARE MOST LIKELY DUE TO BIOMATERIAL IMPACTING OPTICAL SENSOR AND CAUSING FALSE IMPELLA POSITION IN AORTA ALARMS. THE SUCTION ALARMS ARE ALSO A RESULT OF INGESTED BIOMATERIAL. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS DETERMINED TO BE DUE TO BIOMATERIAL. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS DETERMINED TO BE BIOMATERIAL. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ H.6 A NEW MEDICAL DEVICE PROBLEM CODE WAS ADDED TO REFLECT FAILURE MODE CHANGES NOTED IN THE INVESTIGATION RESULTS. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE WERE REVISED FROM THE INITIAL SUBMISSION OF MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25979 IN ACCORDANCE WITH UPDATED PROCEDURES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21303885&pc=NGV | 2124215-2025-05376 | 2025-01-19 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-04 05:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Dysphagia/ Odynophagia; Intracranial Hemorrhage |  |  | 1 | Event Description: THE PATIENT WAS ENROLLED IN THE WATCHMAN CHAMPION-AF STUDY ON (B)(6) 2022 WITH PATIENT IDENTIFIER (B)(6). IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN ACUTE CEREBRAL VASCULAR ACCIDENT (CVA). A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 27 MM WATCHMAN FLX LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT WITH COMPLETE LAA SEAL AND DEPLOYED DEVICE DIAMETER OF 23 MM. THERE WERE NO COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. THE PATIENT WAS DISCHARGED ON A MEDICATION REGIMEN OF ASPIRIN (81MG /DAY) AND APIXABAN (2.5MG/DAY). ON (B)(6) 2025, 1054 DAYS POST INDEX PROCEDURE, THE PATIENT ARRIVED AT THE EMERGENCY DEPARTMENT (ED) VIA EMERGENCY MEDICAL SERVICES (EMS) DUE TO COMPLAINT OF APHASIA, GARBLED SPEECH, RIGHT FACIAL DROOP ALONG WITH DYSARTHRIA THAT STARTED APPROXIMATELY 30 MINUTES PRIOR TO ARRIVAL. THE PATIENT DENIED ANY ILLNESS OR RECENT TRAUMA. ON (B)(6) 2025, COMPUTED TOMOGRAPHY (CT) OF THE BRAIN WAS PERFORMED WITHOUT CONTRAST IN RESPONSE TO THE EVENT WHICH REVEALED NO ACUTE INTRACRANIAL ABNORMALITY. A CT ANGIOGRAM OF THE HEAD AND NECK WERE PERFORMED WHICH REVEALED MODERATE ATHEROSCLEROTIC PLAQUE OF THE BILATERAL CAROTID. CT PERFUSION SHOWED ABNORMALITY OF LATERAL, POSTERIOR, AND FRONTAL LOBE. CT CEREBRAL PERFUSION ANALYSIS REVEALED A SMALL REGION OF POSSIBLE PERFUSION ABNORMALITY IN THE LEFT POSTERIOR FRONTAL LOBE OF QUESTIONABLE CLINICAL SIGNIFICANCE, NOT MEETING THE THRESHOLD FOR TRUE INFRACT. BASED ON THE DIAGNOSTIC FINDINGS, THE PATIENT WAS DIAGNOSED WITH ACUTE CVA. THE PATIENT WAS ON A MEDICATION REGIME OF ASPIRIN (81 MG /DAY) AT THE TIME OF THE EVENT. IN RESPONSE TO EVENT, THE PATIENT WAS GIVEN TNK-TPA-TREATMENT (TENECTEPLASE). ON (B)(6) 2025, POST TNK TPA-TREATMENT, THE PATIENT WAS NOTED TO HAVE DEVELOPED A HEADACHE FOR WHICH REPEATED CT BRAIN WITHOUT CONTRAST WAS PERFORMED. IMAGING REVEALED AN INTERVAL DEVELOPMENT OF EXTENSIVE PARENCHYMAL HEMORRHAGE IN THE RIGHT TEMPORAL LOBE WITH SCATTERED SUBARACHNOID HEMORRHAGE AND INTRAVENTRICULAR HEMORRHAGE WITHIN THE RIGHT LATERAL VENTRICLE AND IN THE 3RD AND 4TH VENTRICLE. NO MIDLINE SHIFTS OR HYDROCEPHALUS NOTED. GREY-WHITE MATTER DIFFERENTIATION WAS PRESERVED. NO LOCALIZED ACUTE INFRACT WAS NOTED. THE PATIENT WAS ADMINISTERED WITH INTRAVENOUS (IV)/INTRAMUSCULAR (IM) AND ORAL MEDICATION CHANGE/ADJUSTMENT WAS DONE. ON (B)(6) 2025, CT BRAIN WITHOUT CONTRAST REVEALED SIMILAR EXTENT OF RIGHT TEMPORAL INTRAPARENCHYMAL HEMORRHAGE. NO MIDLINE SHIFTS OR BRAIN HERNIATION. AN INCREASED DIFFUSE SUBARACHNOID HEMORRHAGE INCREASED INTRAVENTRICULAR HEMORRHAGE NOW WITH NEAR COMPLETE OPACIFICATION OF THE RIGHT LATERAL VENTRICLES AND SIMILAR INVOLVEMENT OF THE THIRD VENTRICLE AND FOURTH VENTRICLE WERE NOTED. ON THE SAME DAY, AN MRI OF BRAIN DEMONSTRATED THERE WAS EITHER A LARGE RIGHT TEMPORAL INTRAPARENCHYMAL HEMORRHAGE OR EXTENSIVE HEMORRHAGIC TRANSFORMATION OF A RIGHT TEMPORAL INFARCT. THIS WAS ASSOCIATED WITH INTRAVENTRICULAR BLOOD AND SIGNIFICANT SUBARACHNOID BLOOD. THERE WAS NO OBVIOUS MIDLINE SHIFT OR HERNIATION. A SMALL CONTRALATERAL LEFT FRONTAL LOBE STROKE WAS NOTED WHICH DID NOT DEMONSTRATE HEMORRHAGIC TRANSFORMATION. A BEDSIDE SWALLOW STUDY WAS ALSO PERFORMED TO TEST THE PATIENT ABILITY TO SWALLOW FOOD, WHICH REVEALED THE PATIENT HAVING MODERATE TO SEVERE DYSPHAGIA. THE PATIENT WAS ADVISED FOR LEVEL 4 DIET (PUREED/ SPEECH THERAPY). ON (B)(6) 2024, THE PATIENT WAS ORIENTED TO PERSON, PLACE AND TIME AND NO FOCAL DEFICITS WAS NOTED. THE PATIENT WAS STARTED ON A MEDICATION REGIME OF CRESTOR DUE TO STROKE. ON (B)(6) 2025, THE PATIENT WAS STARTED ON APIXABAN (2.5MG/DAY) FOR STROKE. ON (B)(6) 2025, THE PATIENT WAS DISCHARGED TO A REHABILITATION/ SKILLED NURSING FACILITY.Event Description: THE PATIENT WAS ENROLLED IN A CLINICAL STUDY ON (B)(6) 2022. IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN ACUTE CEREBRAL VASCULAR ACCIDENT (CVA). A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 27 MM WATCHMAN FLX LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT WITH COMPLETE LAA SEAL AND DEPLOYED DEVICE DIAMETER OF 23 MM. THERE WERE NO COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. THE PATIENT WAS DISCHARGED ON A MEDICATION REGIMEN OF ASPIRIN (81MG /DAY) AND APIXABAN (2.5MG/DAY). ON (B)(6) 2025, 1054 DAYS POST INDEX PROCEDURE, THE PATIENT ARRIVED AT THE EMERGENCY DEPARTMENT (ED) VIA EMERGENCY MEDICAL SERVICES (EMS) DUE TO COMPLAINT OF APHASIA, GARBLED SPEECH, RIGHT FACIAL DROOP ALONG WITH DYSARTHRIA THAT STARTED APPROXIMATELY 30 MINUTES PRIOR TO ARRIVAL. THE PATIENT DENIED ANY ILLNESS OR RECENT TRAUMA. ON (B)(6) 2025, COMPUTED TOMOGRAPHY (CT) OF THE BRAIN WAS PERFORMED WITHOUT CONTRAST IN RESPONSE TO THE EVENT WHICH REVEALED NO ACUTE INTRACRANIAL ABNORMALITY. A CT ANGIOGRAM OF THE HEAD AND NECK WERE PERFORMED WHICH REVEALED MODERATE ATHEROSCLEROTIC PLAQUE OF THE BILATERAL CAROTID. CT PERFUSION SHOWED ABNORMALITY OF LATERAL, POSTERIOR, AND FRONTAL LOBE. CT CEREBRAL PERFUSION ANALYSIS REVEALED A SMALL REGION OF POSSIBLE PERFUSION ABNORMALITY IN THE LEFT POSTERIOR FRONTAL LOBE OF QUESTIONABLE CLINICAL SIGNIFICANCE, NOT MEETING THE THRESHOLD FOR TRUE INFRACT. BASED ON THE DIAGNOSTIC FINDINGS, THE PATIENT WAS DIAGNOSED WITH ACUTE CVA. THE PATIENT WAS ON A MEDICATION REGIME OF ASPIRIN (81 MG /DAY) AT THE TIME OF THE EVENT. IN RESPONSE TO EVENT, THE PATIENT WAS GIVEN TNK-TPA-TREATMENT (TENECTEPLASE). ON (B)(6) 2025, POST TNK TPA-TREATMENT, THE PATIENT WAS NOTED TO HAVE DEVELOPED A HEADACHE FOR WHICH REPEATED CT BRAIN WITHOUT CONTRAST WAS PERFORMED. IMAGING REVEALED AN INTERVAL DEVELOPMENT OF EXTENSIVE PARENCHYMAL HEMORRHAGE IN THE RIGHT TEMPORAL LOBE WITH SCATTERED SUBARACHNOID HEMORRHAGE AND INTRAVENTRICULAR HEMORRHAGE WITHIN THE RIGHT LATERAL VENTRICLE AND IN THE 3RD AND 4TH VENTRICLE. NO MIDLINE SHIFTS OR HYDROCEPHALUS NOTED. GREY-WHITE MATTER DIFFERENTIATION WAS PRESERVED. NO LOCALIZED ACUTE INFRACT WAS NOTED. THE PATIENT WAS ADMINISTERED WITH INTRAVENOUS (IV)/INTRAMUSCULAR (IM) AND ORAL MEDICATION CHANGE/ADJUSTMENT WAS DONE. ON (B)(6) 2025, CT BRAIN WITHOUT CONTRAST REVEALED SIMILAR EXTENT OF RIGHT TEMPORAL INTRAPARENCHYMAL HEMORRHAGE. NO MIDLINE SHIFTS OR BRAIN HERNIATION. AN INCREASED DIFFUSE SUBARACHNOID HEMORRHAGE INCREASED INTRAVENTRICULAR HEMORRHAGE NOW WITH NEAR COMPLETE OPACIFICATION OF THE RIGHT LATERAL VENTRICLES AND SIMILAR INVOLVEMENT OF THE THIRD VENTRICLE AND FOURTH VENTRICLE WERE NOTED. ON THE SAME DAY, AN MRI OF BRAIN DEMONSTRATED THERE WAS EITHER A LARGE RIGHT TEMPORAL INTRAPARENCHYMAL HEMORRHAGE OR EXTENSIVE HEMORRHAGIC TRANSFORMATION OF A RIGHT TEMPORAL INFARCT. THIS WAS ASSOCIATED WITH INTRAVENTRICULAR BLOOD AND SIGNIFICANT SUBARACHNOID BLOOD. THERE WAS NO OBVIOUS MIDLINE SHIFT OR HERNIATION. A SMALL CONTRALATERAL LEFT FRONTAL LOBE STROKE WAS NOTED WHICH DID NOT DEMONSTRATE HEMORRHAGIC TRANSFORMATION. A BEDSIDE SWALLOW STUDY WAS ALSO PERFORMED TO TEST THE PATIENT ABILITY TO SWALLOW FOOD, WHICH REVEALED THE PATIENT HAVING MODERATE TO SEVERE DYSPHAGIA. THE PATIENT WAS ADVISED FOR LEVEL 4 DIET (PUREED/ SPEECH THERAPY). ON (B)(6) 2024, THE PATIENT WAS ORIENTED TO PERSON, PLACE AND TIME AND NO FOCAL DEFICITS WAS NOTED. THE PATIENT WAS STARTED ON A MEDICATION REGIME OF CRESTOR DUE TO STROKE. ON (B)(6) 2025, THE PATIENT WAS STARTED ON APIXABAN (2.5MG/DAY) FOR STROKE. ON (B)(6) 2025, THE PATIENT WAS DISCHARGED TO A REHABILITATION/ SKILLED NURSING FACILITY. IT WAS FURTHER REPORTED THAT ON (B)(6) 2025, THE PATIENT PRESENTED TO EMERGENCY ROOM (ER) WITH THE COMPLAINT OF NOSEBLEED AND VAGINAL BLEEDING WHILE AT THE NURSING HOME. IN ER, UPON EVALUATION THERE WAS NO NASAL OR VAGINAL BLEEDING NOTED. HOWEVER, THE PATIENT WAS FOUND TO BE TACHYCARDIAC AND HAD ELEVATED SODIUM. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR EVALUATION AND FURTHER MANAGEMENT OF HYPERNATREMIA. UPON ADMISSION, BRAIN IMAGING, LAB TEST AND EKG WAS PERFORMED. ON (B)(6) 2025, THE PATIENT WAS NOTED WITH ENCEPHALOPATHY AND NEUROLOGY WAS CONSULTED. THE PATIENT WAS RECOMMENDED FOR AN ELECTROENCEPHALOGRAM (EEG) TO EXCLUDE SUBCLINICAL SEIZURES, BRAIN MRI TO EVALUATE FOR NEW STROKES, CHEST X-RAY AND BLOOD WORK UP. THE SITE CONFIRMED MRI WAS ORDERED DURING THIS ADMISSION, BUT IT WAS NOT DONE. ON (B)(6) 2025, THE PATIENT BECOME VERY LETHARGIC, NOT RESPONSIVE, REFUSED FOOD, WATER, OR MEDICATION AND SUBSEQUENTLY DIED DUE TO POST STROKE ENCEPHALOPATHY (1080 DAYS POST IMPLANT PROCEDURE) AT THE HOSPITAL. AT THE TIME OF THE DEATH, THE SUBJECT WAS ON ASPIRIN (81 MG /DAY) AND APIXABAN (5 MG/DAY). Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H1 TYPE OF REPORTABLE EVENT UPDATED TO DEATH. H6 DEATH IMPACT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304288&pc=PNJ | 9612164-2025-00578 |  | Death | MEDTRONIC IRELAND | 2025-02-04 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion; Thrombosis/Thrombus |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPACT OF OMITTING THE INTRAVENOUS (IV) HEPARIN BOLUS OUTCOMES OF LEADLESS PACEMAKER IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE CARDIAC PERFORATION REQUIRING EMERGENT SURGICAL REPAIR OF THE RIGHT VENTRICLE. THERE WERE PATIENTS WHO EXPERIENCED ACCESS SITE HEMATOMA REQUIRING INTERVENTION, PSEUDOANEURYSM, CARDIAC PERFORATION, PERICARDIAL EFFUSION REQUIRING PERICARDIOCENTESIS, THROMBUS FORMATION DURING THE IMPLANT PROCEDURE AND REHOSPITALIZATION. SOME PATIENTS REQUIRED CONVERSION TO TRANSVENOUS DUAL-CHAMBER PERMANENT PACEMAKER (DC PPM) SYSTEMS OR CARDIAC RESYNCHRONIZATION THERAPY (CRT) SYSTEMS DURING FOLLOW-UP. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/80 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPACT OF OMITTING THE INTRAVENOUS HEPARIN BOLUS ON OUTCOMES OF LEADLESS PACEMAKER IMPLANTATION. JOURNAL OF CARDIOVASCULAR ELECTROPHYSIOLOGY 2024; 35:1212¿1216. DOI: 10.1111/JCE.16284. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304339&pc=DSQ | 3007042319-2025-00178 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED PULMONARY ISSUE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS AN UNSPECIFIED PULMONARY ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304347&pc=DSQ | 3007042319-2025-00214 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304375&pc=DSQ | 3007042319-2025-00237 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE PATIENT HAD BEEN PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304382&pc=DSQ | 3007042319-2025-00251 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304397&pc=DSQ | 3007042319-2025-00236 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE C CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304398&pc=DSQ | 3007042319-2025-00257 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED COVERT CENTRAL NERVOUS SYSTEM INJURY NEUROLOGICAL DYSFUNCTION DIAGNOSED BY COMPUTERIZED TOMOGRAPHY (CT). THE PATIENT SUBSEQUENTLY EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING AN UNSPECIFIED COVERT CENTRAL NERVOUS SYSTEM INJURY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304400&pc=DSQ | 3007042319-2025-00255 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Renal Failure; Pericardial Effusion; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR ACUTE STAGE 3 RENAL DYSFUNCTION, AN UNSPECIFIED BLEED, PERICARDIAL DRAINAGE REQUIRING SURGICAL INTERVENTION, AND EARLY POST-IMPLANT RIGHT HEART FAILURE REQUIRING INITIATION OR CONTINUATION OF INOTROPIC OR VASOPRESSOR SUPPORT OR INHALED NITRIC OXIDE AFTER 14 DAYS FOLLOWING VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION OR HAVING TO INITIATE SUPPORT WITHIN 30 DAYS OF IMPLANT FOR A DURATION OF AT LEAST 14 DAYS. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING ACUTE STAGE 3 RENAL DYSFUNCTION AND AN UNSPECIFIED BLEED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304402&pc=DSQ | 3007042319-2025-00238 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE CARDIOMYOPATHY. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS END STAGE CARDIOMYOPATHY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304420&pc=DSQ | 3007042319-2025-00193 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Renal Failure; Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR RENAL DYSFUNCTION. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304422&pc=DSQ | 3007042319-2025-00185 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304423&pc=DSQ | 3007042319-2025-00263 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304440&pc=DSQ | 3007042319-2025-00277 |  | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED MAJOR INFECTION. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304447&pc=DSQ | 3007042319-2025-00287 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304466&pc=DSQ | 3007042319-2025-00290 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304472&pc=DSQ | 3007042319-2025-00307 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304509&pc=DSQ | 3007042319-2025-00346 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE CARDIOMYOPATHY. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS END STAGE CARDIOMYOPATHY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304529&pc=DSQ | 3007042319-2025-00361 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED MAJOR INFECTION. THE PATIENT SUBSEQUENTLY EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING AN UNSPECIFIED MAJOR INFECTION AND WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304571&pc=DSQ | 3007042319-2025-00400 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT.MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304593&pc=DSQ | 3007042319-2025-00425 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD WAS DECOMMISSIONED AND EXPLANTED. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF VAD SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304594&pc=DSQ | 3007042319-2025-00424 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Respiratory Failure; Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR SEPSIS, HEPATIC DYSFUNCTION, RESPIRATORY FAILURE, AND A GASTROINTESTINAL INFECTION REQUIRING INTRAVENOUS MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING SEPSIS, HEPATIC DYSFUNCTION, RESPIRATORY FAILURE, AND A GASTROINTESTINAL INFECTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304595&pc=DSQ | 3007042319-2025-00427 |  | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Brain Injury | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED BLEED AND A TRAUMATIC BRAIN INJURY DIAGNOSED BY COMPUTERIZED TOMOGRAPHY (CT). THE PATIENT SUBSEQUENTLY EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING AN UNSPECIFIED BLEED AND A TRAUMATIC BRAIN INJURY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304601&pc=DSQ | 3007042319-2025-00423 |  | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hypervolemia; Multiple Organ Failure; Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR FLUID OVERLOAD AND AN UNSPECIFIED OVERT CENTRAL NERVOUS SYSTEM INJURY DIAGNOSED CLINICALLY. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304602&pc=DSQ | 3007042319-2025-00430 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CARDIA ARRHYTHMIA. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304606&pc=DSQ | 3007042319-2025-00431 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED REASON. THE PATIENT SUBSEQUENTLY EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD WAS DECOMMISSIONED AND EXPLANTED. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER THE PATIENT WAS HOSPITALIZED FOR AN UNSPECIFIED REASON. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304614&pc=DSQ | 3007042319-2025-00446 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Respiratory Failure; Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A CEREBRAL HEMORRHAGE DIAGNOSED BY COMPUTERIZED TOMOGRAPHY (CT), RESPIRATORY FAILURE REQUIRING REINTUBATION, AND HEPATIC DYSFUNCTION. THE PATIENT SUBSEQUENTLY EXPIRED AFTER WITHDRAWAL OF VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED WHEN VAD SUPPORT WAS WITHDRAWN AFTER THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE, RESPIRATORY FAILURE, AND HEPATIC DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304634&pc=DSQ | 3007042319-2025-00444 |  | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Ventricular Fibrillation; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A VENTRICULAR ASSIST DEVICE (VAD) THROMBUS, AN UNSPECIFIED DEVICE MALFUNCTION, AND SUSTAINED VENTRICULAR ARRHYTHMIA RESULTING IN CLINICAL COMPROMISE, OR REQUIRING HOSPITALIZATION OR DRUG TREATMENT, DEFIBRILLATION, CARDIOVERSION, ICD THERAPY, OR ARRHYTHMIA ABLATION PROCEDURE. THE VENTRICULAR ASSIST DEVICE (VAD) WAS ELECTIVELY DECOMMISSIONED AND THE PATIENT SUBSEQUENTLY EXPIRED DUE TO THE CARDIAC ARRHYTHMIA. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED UNSPECIFIED VAD MALFUNCTION. THE REPORTED EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO CARDIAC ARRHYTHMIA. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304641&pc=DSQ | 3007042319-2025-00445 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-02-04 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Anemia; Hemolysis; Renal Failure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR MAJOR HEMOLYSIS, ANEMIA, HYPERBILIRUBINEMIA, HEMOGLOBINURIA, LATE RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, ACUTE STAGE 3 RENAL DYSFUNCTION, A SUSPECTED VENTRICULAR ASSIST DEVICE (VAD) THROMBUS, AND ABNORMAL PUMP PARAMETERS REQUIRING INTRAVENOUS MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. THE REPORTED EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING MAJOR HEMOLYSIS, ANEMIA, HYPERBILIRUBINEMIA, HEMOGLOBINURIA, LATE RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, ACUTE STAGE 3 RENAL DYSFUNCTION, A SUSPECTED THROMBUS, AND ABNORMAL PUMP PARAMETERS. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE ABNORMAL PUMP PARAMETERS EVENT INCLUDING BUT NOT LIMITED TO THROMBUS FORMATION/INGESTION, THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q3 2024 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304718&pc=DSQ | 2916596-2025-00642 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WOULD NOT BE EXPLANTED. Manufacturer Narrative: SECTION A: PATIENT AGE AND WEIGHT WERE ESTIMATED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE ACCOUNT WILL NOT COMMUNICATE ANY ADDITIONAL INFORMATION REGARDING THE EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304737&pc=MKJ | 3004123209-2025-00028 | 2025-01-07 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-02-04 05:00 | MKJ | UNKNOWN_MEDICAL - BELFAST_PRODUCT | Unexpected Shutdown | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE POWERED OFF DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE WOULD BE INOPERABLE AND DEFIBRILLATION THERAPY WOULD NOT BE AVAILABLE IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED.Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE POWERED OFF DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE WOULD BE INOPERABLE AND DEFIBRILLATION THERAPY WOULD NOT BE AVAILABLE IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: A STRYKER SERVICE REPRESENTATIVE PERFORMED AN INITIAL EVALUATION OF THE CUSTOMER¿S DEVICE. DURING THE REPORTED PATIENT-INVOLVED EVENT ON THE (B)(6) 2025, THE DEVICE FIRST RECORDED A MANUAL POWER-ON OF 14 SECOND DURATION, DURING WHICH THE DEVICE POWERED OFF SUDDENLY WITHOUT ENTERING THE SHUTDOWN SEQUENCE, WHICH WOULD INDICATE POWER HAD BEEN REMOVED FROM THE DEVICE. THE DEVICE THEN RECORDED A POWER-ON OF 5 MINUTES AND 19 SECONDS, DURING WHICH A PATIENT IMPEDANCE WAS SUCCESSFULLY DETECTED HOWEVER NO SHOCK WAS ADVISED, AS THE PATIENT PRESENTED A NON-SHOCKABLE RHYTHM. THE IMPEDANCE THEN WENT OUT OR RANGE AFTER 2 MINUTES AND 43 SECONDS. THE DEVICE THEN ISSUED ¿CHECK PADS¿ PROMPTS FOR THE REMAINDER OF THE POWER-ON, DURING WHICH TIME THE DEVICE RECORDED 3 ¿WARNING, OFF BUTTON PRESSED¿ PROMPTS, SUGGESTING THE USER WAS ATTEMPTING TO POWER OFF THE DEVICE. THE DEVICE THEN POWERED OFF ISSUING A ¿WARNING LOW BATTERY¿ PROMPT. NO FAULT WAS FOUND ON THE HDF-3500 OR RETURNED PAD-PAK. NO CONTINUITY ISSUE WAS IDENTIFIED ON THE PATIENT OUTPUT CABLES AND NO FAULT FOUND ON THE POGO PINS THAT WOULD HAVE RESULTED IN AN IMPEDANCE DETECTION OR POWER SUPPLY ISSUE. THE DEVICE ACCURATELY MEASURED IMPEDANCES THROUGHOUT THE SPECIFIED RANGE, EVEN UNDER THE STRESS OF ELEVATED TEMPERATURE. THE DEVICE DELIVERED THE SHOCK THERAPY SEQUENCE WITHOUT FAULT USING THE RETURNED PAD-PAK. FURTHERMORE, THE DEVICE DID NOT LOSE POWER DURING TESTING AT HEARTSINE. GIVEN NO FAULT ON THE HDF-3500, THE INVESTIGATION WAS UNABLE TO DETERMINE THE CAUSE OF THE DEVICE POWERING OFF WITH ¿WARNING LOW BATTERY¿ PROMPTS. GIVEN THE DEVICE RECORDING 3 ¿WARNING, OFF BUTTON PRESSED¿ PROMPTS BEFORE SHUTDOWN THIS COULD SUGGEST THE USER WAS ALREADY ATTEMPTING TO POWER OFF THE DEVICE PRIOR TO SHUTDOWN, WHICH COULD SUGGEST THE DEVICE ISSUED THE ¿WARNING LOW BATTERY¿ PROMPTS DUE TO PAD-PAK BEING REMOVED. HOWEVER, THIS COULD NOT BE CONFIRMED. THE CUSTOMER'S DEVICE WILL BE SCRAPPED. THE CUSTOMER WILL RECEIVE A REPLACEMENT DEVICE. Manufacturer Narrative: HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. THE HDF-3500 DEVICE IS NOT AVAILABLE FOR DISTRIBUTION IN THE UNITED STATES BUT IS SIMILAR TO THE SAM 350P AND 450P WHICH IS DISTRIBUTED IN THE UNITED STATES. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. HEARTSINE REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED HEARTSINE WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304922&pc=ILY | MW5165613 | 2024-02-17 05:00 | Death | TESLA BIOHEALING, INC. | 2025-02-04 05:00 | ILY | TESLA MEDBED GENERATOR | Device Markings/Labelling Problem; Therapeutic or Diagnostic Output Failure | Dementia; Tics/Tremor; Unspecified Gastrointestinal Problem |  |  | 1 | Event Description: I CLAIM I PURCHASED YOUR 2 TESLA BIOHEALERS ON (B)(6) 2021, ORDER (B)(4); I CLAIM I HAD SEVERAL CONVERSATIONS WITH (B)(4) REGARDING ALL MY MOM¿S (B)(6) MEDICAL CONDITIONS I WAS HOPEFUL RESULTS WOULD OCCUR. WHEN WE STARTED, SHE HAD ALREADY BEEN THROUGH MICRO-CHRONIC STROKES, ALONG WITH 3 OTHER LARGER STROKES. SHE HAD LOST THE VISION IN HER LEFT EYE. SHE HAD AN ABDOMINAL AORTIC ANEURYSM AT 4.7 CM THAT GREW TO 6.8 CM OVER A 3-YEAR PERIOD WHICH NO ONE WOULD DO SURGERY BECAUSE SHE WAS ALSO DIAGNOSED WITH DEMENTIA IN 2020. THEY FELT THE DEMENTIA WOULD GET WORSE FASTER IF THEY DID THE OPEN SURGERY ON HER AAA. IN 2020, MY MOM ALSO STARTED GETTING TREMORS IN HER RIGHT HAND AND HAVING SOME ISSUES WITH WALKING STABLE ON HER OWN. I CLAIM WE RECEIVED THE MED BEDS WITH NO ISSUE. MY MOM USED THEM FOR SEVERAL MONTHS WITH HOPES OF HER MEMORY COMING BACK SOME. THE BIG HOPE WAS THAT NOTHING MORE WOULD GROW AND THE SMALL ISSUES WOULD GO AWAY. I HAD HOPED HER TREMORS AND BALANCE ISSUES WOULD RESOLVE AND THAT HER SHORT TERM MEMORY WOULD START TO COME BACK. NOTHING EVER HAPPENED. I CLAIM I SPOKE WITH YOU OVER THE PHONE ON OCTOBER 9, 2021. YOU STATED TO PUT ONE DEVICE UNDER HER BED AND ONE DEVICE NEXT TO HER HEAD ON THE NIGHTSTAND. WE DID THAT WITH NO IMPROVEMENT. I CLAIM ON SEPTEMBER OF 2023, I DECIDED TO SEND A RESPONSE VIA EMAIL TO ASK HOW THESE ARE EVEN WORTH THE COST? I RECEIVED A RESPONSE FROM (B)(4) WHO HAD ZERO INPUT INTO THE WORKINGS OF THESE MEDBEDS. NOTHING GOOD HAPPENED TO MY MOM, SHE EVEN HAD STOMACH ISSUES AND HAD HER GALL BLADDER REMOVED (B)(6) 2021 WHICH WAS AFTER WE PURCHASED THE BIOHEALERS. HER AAA NEVER GOT BETTER, IT GOT WORSE. HER DEMENTIA NEVER GOT BETTER IT GOT WORSE. HER PARKINSON'S SYMPTOMS NEVER GOT BETTER EITHER! THEN IN 2023 MY DAD ALSO WAS DIAGNOSED WITH PROSTATE CANCER AND HE WAS SLEEPING IN THE SAME BED. MY MOTHER PASSED AWAY (B)(6) 2024. I CLAIM THE 2 TESLA BIOHEALERS ARE NOTHING BUT FRAUD. I CLAIM I SUBMITTED 2 REVIEWS ON THE WEBSITE, 3/10/23 AND 9/13/23 BOTH INCLUDED MY NEGATIVE REVIEW AND NEITHER ARE CURRENTLY ON YOUR WEBSITE ANY LONGER. I GOT THE JUDGE.ME CONFIRMATION AND CLICKED THE LINK. I CLAIM TODAY YOU CAN SEE ZERO BAD REVIEWS ON YOUR PRODUCT ON YOUR WEBSITE. I CLAIM THIS IS FRAUDULENT AND DECEITFUL FOR YOUR CURRENT AND POTENTIAL CUSTOMERS. I CLAIM I FOUND A LETTER ONLINE LAST YEAR ABOUT A WARNING FROM THE FDA! I SPOKE TO THE REPRESENTATIVE LISTED AT THE BOTTOM AND SHE STATED SHE HAS HAD SEVERAL PHONE CALLS BASED ON THIS FDA LETTER AND THAT THE WARNING HAS NOT BEEN RESOLVED. SEE LINK: TESLA BIOHEALING, INC. - (B)(4)- 08/10/2023 ¦ FDA. I CLAIM YOU ARE SCAMMING PEOPLE WHO JUST WANT TO SAVE THEIR LOVED ONES AND ARE IN A VULNERABLE STATE. I CLAIM I WAS WILLING TO SPEND WHATEVER I NEEDED TO TRY AND HEAL MY MOM. I CLAIM YOU AND YOUR COMPANY TOOK ADVANTAGE OF THE NEED TO HELP HEAL MY MOM AND MY VULNERABLE STATE. I REQUIRE FULL RETURN OF PURCHASED PRICE PAID OF (B)(6). I REQUIRE MY TWO REVIEWS BE PUBLISHED ON YOUR WEBSITE. IF I DO NOT RECEIVE RESPONSES AND ACTIONS TO MY SPECIFIC ASKS AND REQUIREMENTS WITHIN FIFTEEN (15) DAYS, AS REQUESTED SHALL CONSTITUTE YOUR EXPLICIT, BINDING AND COMPLETE TACIT AGREEMENT AS YOUR UNEQUIVOCAL ACCEPTANCE BY YOURSELF AND YOUR ORGANIZATION AND WILL PURSUE ALL APPROPRIATE REMEDIES AVAILABLE TO HOLD YOU AND YOUR COMPANY COMMERCIALLY AND PERSONALLY LIABLE. IT IS MY DUTY AS AN AMERICAN TO KEEP THE PEACE AND UPHOLD THE PUBLIC LAW. AND I OBJECT TO EVIL. REFERENCE REPORT MW5165614. Event Description: ADDITIONAL INFORMATION RECEIVED ON 2/6/2025 FOR REPORT NUMBER MW5165613. REFERENCE REPORT NUMBER MW5165614 AND MW5165614_1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21304929&pc=ILY | MW5165614 | 2024-02-17 05:00 | Death | TESLA BIOHEALING, INC. | 2025-02-04 05:00 | ILY | TESLA MEDBED GENERATOR | Adverse Event Without Identified Device or Use Problem | Inadequate Pain Relief; Unspecified Gastrointestinal Problem |  |  | 1 | Event Description: I CLAIM I PURCHASED YOUR 2 TESLA BIOHEALERS ON (B)(6) 2021, ORDER (B)(4). I CLAIM I HAD SEVERAL CONVERSATIONS WITH (B)(4) REGARDING ALL MY MOM¿S (B)(6) MEDICAL CONDITIONS I WAS HOPEFUL RESULTS WOULD OCCUR. WHEN WE STARTED, SHE HAD ALREADY BEEN THROUGH MICRO-CHRONIC STROKES, ALONG WITH 3 OTHER LARGER STROKES. SHE HAD LOST THE VISION IN HER LEFT EYE. SHE HAD AN ABDOMINAL AORTIC ANEURYSM AT 4.7 CM THAT GREW TO 6.8 CM OVER A 3-YEAR PERIOD WHICH NO ONE WOULD DO SURGERY BECAUSE SHE WAS ALSO DIAGNOSED WITH DEMENTIA IN 2020. THEY FELT THE DEMENTIA WOULD GET WORSE FASTER IF THEY DID THE OPEN SURGERY ON HER AAA. IN 2020, MY MOM ALSO STARTED GETTING TREMORS IN HER RIGHT HAND AND HAVING SOME ISSUES WITH WALKING STABLE ON HER OWN. I CLAIM WE RECEIVED THE MED BEDS WITH NO ISSUE. MY MOM USED THEM FOR SEVERAL MONTHS WITH HOPES OF HER MEMORY COMING BACK SOME. THE BIG HOPE WAS THAT NOTHING MORE WOULD GROW AND THE SMALL ISSUES WOULD GO AWAY. I HAD HOPED HER TREMORS AND BALANCE ISSUES WOULD RESOLVE AND THAT HER SHORT TERM MEMORY WOULD START TO COME BACK. NOTHING EVER HAPPENED. I CLAIM I SPOKE WITH YOU OVER THE PHONE ON OCTOBER 9, 2021. YOU STATED TO PUT ONE DEVICE UNDER HER BED AND ONE DEVICE NEXT TO HER HEAD ON THE NIGHTSTAND. WE DID THAT WITH NO IMPROVEMENT. I CLAIM ON SEPTEMBER OF 2023, I DECIDED TO SEND A RESPONSE VIA EMAIL TO ASK HOW THESE ARE EVEN WORTH THE COST? I RECEIVED A RESPONSE FROM (B)(4) WHO HAD ZERO INPUT INTO THE WORKINGS OF THESE MEDBEDS. NOTHING GOOD HAPPENED TO MY MOM, SHE EVEN HAD STOMACH ISSUES AND HAD HER GALL BLADDER REMOVED (B)(6) 2021 WHICH WAS AFTER WE PURCHASED THE BIOHEALERS. HER AAA NEVER GOT BETTER, IT GOT WORSE. HER DEMENTIA NEVER GOT BETTER IT GOT WORSE. HER PARKINSON'S SYMPTOMS NEVER GOT BETTER EITHER. THEN IN 2023 MY DAD ALSO WAS DIAGNOSED WITH PROSTATE CANCER AND HE WAS SLEEPING IN THE SAME BED. MY MOTHER PASSED AWAY (B)(6) 2024. I CLAIM THE 2 TESLA BIOHEALERS ARE NOTHING BUT FRAUD. I CLAIM I SUBMITTED 2 REVIEWS ON THE WEBSITE, 3/10/23 AND 9/13/23 BOTH INCLUDED MY NEGATIVE REVIEW AND NEITHER ARE CURRENTLY ON YOUR WEBSITE ANY LONGER. I GOT THE JUDGE.ME CONFIRMATION AND CLICKED THE LINK. I CLAIM TODAY YOU CAN SEE ZERO BAD REVIEWS ON YOUR PRODUCT ON YOUR WEBSITE. I CLAIM THIS IS FRAUDULENT AND DECEITFUL FOR YOUR CURRENT AND POTENTIAL CUSTOMERS. I CLAIM I FOUND A LETTER ONLINE LAST YEAR ABOUT A WARNING FROM THE FDA! I SPOKE TO THE REPRESENTATIVE LISTED AT THE BOTTOM AND SHE STATED SHE HAS HAD SEVERAL PHONE CALLS BASED ON THIS FDA LETTER AND THAT THE WARNING HAS NOT BEEN RESOLVED. SEE LINK: TESLA BIOHEALING, INC. - (B)(4)- 08/10/2023 ¦ FDA. I CLAIM YOU ARE SCAMMING PEOPLE WHO JUST WANT TO SAVE THEIR LOVED ONES AND ARE IN A VULNERABLE STATE. I CLAIM I WAS WILLING TO SPEND WHATEVER I NEEDED TO TRY AND HEAL MY MOM. I CLAIM YOU AND YOUR COMPANY TOOK ADVANTAGE OF THE NEED TO HELP HEAL MY MOM AND MY VULNERABLE STATE. I REQUIRE FULL RETURN OF PURCHASED PRICE PAID OF (B)(6). I REQUIRE MY TWO REVIEWS BE PUBLISHED ON YOUR WEBSITE. IF I DO NOT RECEIVE RESPONSES AND ACTIONS TO MY SPECIFIC ASKS AND REQUIREMENTS WITHIN FIFTEEN (15) DAYS, AS REQUESTED SHALL CONSTITUTE YOUR EXPLICIT, BINDING AND COMPLETE TACIT AGREEMENT AS YOUR UNEQUIVOCAL ACCEPTANCE BY YOURSELF AND YOUR ORGANIZATION AND WILL PURSUE ALL APPROPRIATE REMEDIES AVAILABLE TO HOLD YOU AND YOUR COMPANY COMMERCIALLY AND PERSONALLY LIABLE. IT IS MY DUTY AS AN AMERICAN TO KEEP THE PEACE AND UPHOLD THE PUBLIC LAW. AND I OBJECT TO EVIL. REFERENCE REPORT MW5165613. Event Description: ADDITIONAL INFORMATION RECEIVED ON 2/6/2025 FOR REPORT NUMBER MW5165614. REFERENCE REPORT NUMBER MW5165613 AND MW5165613_1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21305077&pc=DSQ | 2916596-2025-00647 | 2021-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE " OUTCOME OF LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION WITH CONCOMITANT AORTIC VALVE PROCEDURE" THAT THE HEARTMATE 3 MAY BE RELATED TO DEATH, CARDIOGENIC SHOCK, IN HOSPITAL STROKES, AND RIGHT HEART FAILURE. SOME PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD) REQUIRE SURGICAL AORTIC VALVE PROCEDURES (AVP) DUE TO MODERATE OR SEVERE AORTIC VALVE REGURGITATION (AR) TO ENSURE SUFFICIENT LVAD FUNCTION. IT IS UNCERTAIN HOW CONCOMITANT AVP DURING LVAD IMPLANTATION IMPACTS OUTCOMES. SO, THE STUDY WAS TO EVALUATE THE CLINICAL RESULTS OF PATIENTS RECEIVING AN LVAD CONCOMITANT WITH AVP IN COMPARISON WITH ISOLATED LVAD IMPLANTATION. 183 PATIENTS WITH ADVANCED HEART FAILURE UNDERWENT LVAD IMPLANTATION BETWEEN 2010-2021. OF THESE, 15 PATIENTS (GROUP 1) UNDERWENT LVAD IMPLANTATION WITH CONCOMITANT AVP AND WERE COMPARED WITH 140 PATIENTS (GROUP 2) WITH ISOLATED LVAD IMPLANTATION. DATA WERE RETROSPECTIVELY ANALYZED. THE RESULTS SHOWED BOTH GROUPS HAD PATIENTS THAT WENT INTO ACUTE CARDIOGENIC SHOCK (INTERMACS PROFILE 1 + 2, GROUP 1 VS. GROUP 2: 41 VS. 33%, P=0.8). INDICATION FOR AVP WAS MODERATE OR SEVERE AR. 13 PATIENTS UNDERWENT AORTIC VALVE REPLACEMENT AND 2 BENTALL PROCEDURE. IMPLANTED PERMANENT VAD DEVICES WERE HVAD AND HEART MATE 3 (HM3) (GROUP 1: HVAD: 93.4%, HM3: 6.6%; GROUP 2: HVAD: 79.9%, HM3: 21.1%, P=0.2). PROCEDURE AND BYPASS TIME WERE SIGNIFICANTLY LONGER IN GROUP 1. IN-HOSPITAL STROKE (0 VS. 1.6%, P=0.8) AND RETHORACOTOMY RATES (0 VS. 0%, P=1.0) WERE LOW IN BOTH GROUPS. RATES OF CONCOMITANT TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) IMPLANTATION (13.3 VS. 22.0%, P=0.7) AND CONSECUTIVE RIGHT HEART FAILURE IN THE FIRST 3 MONTHS (5.5 VS. 5.1%, P=1.0) SHOWED NO SIGNIFICANT DIFFERENCES BETWEEN GROUPS. GROUP 1 PATIENTS HAD A COMPARABLE 30-DAY RATE OF MORTALITY (6.6 VS. 13.6%, P=0.6). IN CONCLUSION, LVAD IMPLANTATION WITH CONCOMITANT AVP IS SAFE AND SHOWS EQUIVALENT RESULTS COMPARED TO ISOLATED LVAD IMPLANTATION. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT WAS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN 2010-2021. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. AUTHOR CITATION: BHADRA, O., ET AL (2024). OUTCOME OF LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION WITH CONCOMITANT AORTIC VALVE PROCEDURE. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S288. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1235. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21305227&pc=DSQ | 2916596-2025-00566 | 2025-01-19 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ALTHOUGH THE CAUSE OF DEATH WAS NOT PROVIDED; IT WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE CAUSE OF DEATH WAS NOT PROVIDED DUE TO PRIVACY REGULATION. THE DEVICE WAS CONFIRMED TO HAVE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21305750&pc=JDI | 1038671-2025-00640 | 2020-06-06 04:00 | Death | EXACTECH, INC. | 2025-02-04 05:00 | JDI | UNKNOWN HIP COMPONENTS | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT HAD AN INITIAL HIP IMPLANTED ON AN UNKNOWN DATE AND UNDERWENT A REVISION PROCEDURE ON (B)(6) 2020. THE PARTY STATED THAT THEIR PARENT HAD A HIP REPLACEMENT THAT CAUSED THEIR DEATH. INFORMATION REGARDING THE TIMING AND NATURE OF THE PATIENT'S DEATH WAS NOT PROVIDED. NO FURTHER INFORMATION AVAILABLE.Event Description: IT WAS REPORTED THAT A PATIENT HAD AN INITIAL HIP IMPLANTED ON AN UNKNOWN DATE AND UNDERWENT A REVISION PROCEDURE APPROXIMATELY 55 MONTHS AGO. THE PARTY STATED THAT THEIR PARENT HAD A HIP REPLACEMENT THAT CAUSED THEIR DEATH. INFORMATION REGARDING THE TIMING AND NATURE OF THE PATIENT'S DEATH WAS NOT PROVIDED. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE CORRECTED: B5. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. Manufacturer Narrative: THE REPORTED EVENT WAS UNABLE TO BE CONFIRMED DUE TO LIMITED INFORMATION RECEIVED FROM THE CUSTOMER. NO DEVICE WAS RETURNED FOR EVALUATION; FURTHER, PHOTOGRAPHS AND/OR RADIOGRAPH IMAGES WERE NOT PROVIDED FOR REVIEW. OPERATIVE NOTES AND/OR MEDICAL RECORDS WERE NOT PROVIDED FOR REVIEW OF USAGE/ TECHNIQUE. A REVIEW OF MANUFACTURING DATA WAS UNABLE TO BE PERFORMED AS THE LOT INFORMATION OF THE PRODUCT INVOLVED IN THE EVENT WAS NOT AVAILABLE. A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AS THE NECESSARY INFORMATION TO ADEQUATELY INVESTIGATE THE REPORTED EVENT WAS NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21305801&pc=NPT | 2135147-2025-00524 | 2025-01-07 05:00 | Death | ABBOTT MEDICAL | 2025-02-04 05:00 | NPT | FLEXNAV DELIVERY SYSTEM | Difficult to Advance | Stroke/CVA; Non specific EKG/ECG Changes; Intracranial Hemorrhage; Obstruction/Occlusion; Embolism/Embolus | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 35MM NAVITOR VISON VALVE WAS CHOSEN FOR IMPLANT USING AN LARGE FLEXNAV DELIVERY SYSTEM. DURING PROCEDURE, THE VALVE WAS LOADED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). DURING THE ATTEMPT TO PLACE THE VALVE OVER THE AORTIC ARCH, RESISTANCE AND DIFFICULTIES WERE ENCOUNTERED, PREVENTING THE TRANSITION. THE NON-ABBOTT GUIDE WIRE FOR A NON-ABBOTT WIRE AND STILL WEREN¿T ABLE TO PULL THE VALVE TO THE INNER CURVE OF THE AORTIC ARCH, SNARING WAS NOT AN OPTION BECAUSE OF SEVERE PERIPHERAL VASCULAR DISEASE WITHOUT SECOND ACCESS. THE PATIENT DEVELOPED FOCAL NEUROLOGICAL DEFICITS (RIGHT SIDED HEMIPARESIS AND APHASIA) DURING THE ATTEMPT TO CROSS THE ARCH WITH THE VALVE. THE PHYSICIAN MENTIONED THERE WAS NO EVIDENT MALFUNCTION OF THE NAVITOR VALVE. THE SYSTEM WITH THE VALVE HAD TO BE REMOVED AND NEW NON-ABBOTT VALVE WAS IMPLANTED WITHOUT ANY ISSUES. IMMEDIATELY AFTER THE PROCEDURE, PATIENT RECEIVED A CEREBRAL COMPUTED TOMOGRAPHY (CT) SCAN AND WAS TRANSFERRED TO NEURO-ANGIOGRAPHY. THE PATIENT HAD AN OCCLUSION OF THE P2 SEGMENT OF THE LEFT A. CEREBRI POSTERIOR. THERE WAS NO POSSIBILITY TO RETRIEVE THE EMBOLUS THEREFORE THE SEGMENT WAS STENTED REGAINING FLOW DISTAL OF THE OCCLUSION. THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) AND TREATED WITH TIROFIBAN TO AVOID STENT THROMBOSIS. UNFORTUNATELY, THE PATIENT DEVELOPED MASSIVE INTRACRANIAL HEMORRHAGE WITHOUT POSSIBLE NEUROSURGICAL COUNTERMEASURES. NEXT DAY, THE PATIENT WAS REPORTED PASSED AWAY. THE CAUSE OF DEATH WAS INTRACRANIAL HEMORRHAGE. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF DIFFICULT ADVANCEMENT THROUGH PATIENT ANATOMY, DEATH, STROKE, OCCLUSION, INTRACRANIAL HEMORRHAGE, AND EMBOLISM WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. THE EVENT AND PROVIDED IMAGING WERE REVIEWED BY AN ABBOTT DIRECTOR OF MEDICAL AFFAIRS. THE DEVICE WAS RETURNED FOR ANALYSIS. NO ANOMALIES WERE FOUND DURING VISUAL AND FUNCTIONAL INSPECTION. BASED ON ALL AVAILABLE INFORMATION, THE REPORTED DIFFICULT ADVANCEMENT THROUGH PATIENT ANATOMY APPEARS TO BE RELATED TO CHALLENGING PATIENT ANATOMY. THE REPORTED DEATH, STROKE, AND OCCLUSION APPEAR TO BE CASCADING EVENTS OF THE INTRACRANIAL HEMORRHAGE. THE REPORTED INTRACRANIAL HEMORRHAGE IS RELATED TO THE EMBOLISM. A CAUSE FOR THE REPORTED EMBOLISM COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS THE RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: 1817.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21306065&pc=DSQ | 2916596-2025-00549 | 2024-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE TITLED, ¿CRYOABLATION FOR VENTRICULAR ARRHYTHMIA DURING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION IS SAFE AND FEASIBLE¿, A RETROSPECTIVE COHORT STUDY STATED THAT POST-OPERATIVE VENTRICULAR ARRHYTHMIA (VA) OCCURRED IN UP TO 60% OF LEFT VENTRICULAR ASSIST DEVICE (LVAD) RECIPIENTS, ASSOCIATED WITH INCREASED MORTALITY AND MORBIDITY, INCLUDING RIGHT VENTRICULAR DYSFUNCTION. A SINGLE-CENTER RETROSPECTIVE COHORT STUDY WAS DESIGNED INCLUDING CONSECUTIVE LVAD RECIPIENTS WITH DOCUMENTED PREOPERATIVE VAS. PATIENTS WERE STRATIFIED BY WHETHER THEY HAD CONCOMITANT CRYOABLATION PERFORMED (¿CRYO GROUP¿) OR NOT (¿NON-CRYO GROUP¿). THE PRIMARY OUTCOME WAS THE OCCURRENCE OF POST-OPERATIVE VAS, WHILE SECONDARY OUTCOMES INCLUDED 30-DAY ALL-CAUSE MORTALITY, ANTI-ARRHYTHMIC DRUG (AAD) USE, AND COMPLICATIONS. THE COHORT CONSISTED OF 14 PATIENTS IN THE CRYO GROUP AND 25 PATIENTS IN THE NON-CRYO GROUP, WITH >= 80% OF PATIENTS IN EITHER GROUP RECEIVING HEARTMATE 3. POST-OPERATIVE VAS OCCURRED IN 36% OF PATIENTS IN THE CRYO GROUP AND 56% IN THE NON-CRYO GROUP (P=.22). EARLY POSTOPERATIVE VA (WITHIN 30 DAYS POST-OPERATIVELY) OCCURRED IN 14% OF PATIENTS IN THE CRYO GROUP AND 24% IN THE NON-CRYO GROUP(P=.69). AT SIX MONTHS POST-OPERATIVE, 50% OF SURVIVING PATIENTS IN THE CRYO GROUP AND 76% IN THE NON-CRYO GROUP REMAINED ON AAD (P=.12). THE 1-YEAR SURVIVAL PROBABILITY WAS 71.4% IN THE CRYO GROUP AND 80.0% IN THE NON-CRYO GROUP (P=.25). CONCLUSION: THERE WERE NO STATISTICALLY SIGNIFICANT DIFFERENCES IN THE RATES OF THE PREDEFINED COMPLICATIONS, INCLUDING RE-STERNOTOMY FOR HEMOSTASIS, TRACHEOSTOMY REQUIREMENT AND THROMBOEMBOLIC EVENTS. DURATION OF HOSPITAL STAY (P=0.12) AND INTENSIVE CARE UNIT (ICU) STAY (P=0.07) DID NOT DIFFER SIGNIFICANTLY. THERE WERE NO 30-DAY ALL-CAUSE MORTALITY IN THE CRYO GROUP. FINDINGS SUGGEST THAT THIS COMBINED PROCEDURE IS SAFE AND FEASIBLE. ALTHOUGH LIMITED BY A MODEST SAMPLE SIZE, THERE WAS A NON-STATISTICALLY SIGNIFICANT TREND TOWARDS REDUCED POST-OPERATIVE VAS AND AAD USE FOLLOWING SURGICAL CRYOABLATION. MOVING FORWARD, CRYOABLATION SHOULD BE CONSIDERED AS A TREATMENT OPTION IN LVAD RECIPIENTS WITH PREOPERATIVE VAS. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE SAME AS PUBLISHED DATE SINCE DATE OF DATA COLLECTION WAS NOT PROVIDED SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. AUTHOR CITATION: WONG, C.P., ET AL (2024). CRYOABLATION FOR VENTRICULAR ARRHYTHMIA DURING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION IS SAFE AND FEASIBLE: A RETROSPECTIVE COHORT STUDY, THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, VOLUME 43, ISSUE 4, PAGE S59, HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.122. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS IN USE WERE NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21306079&pc=DSQ | 2916596-2025-00548 | 2025-01-18 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS DECREASED FLOW ON THE HEARTMATE 3 RIGHT VENTRICULAR ASSIST DEVICE (RVAD) UP TO ZERO FLOW WITH HIGH PUMP POWER CONSUMPTION. THERE WAS ZERO FLOW THROUGH THE PUMP CONFIRMED BY ECHOCARDIOGRAM (ECHO). THE RVAD WAS TURNED OFF. THE PATIENT WAS HEMODYNAMICALLY STABLE AT THE TIME OF THE REPORT WITHOUT INOTROPES. THE HOSPITAL WAS NOT PLANNING TO EXCHANGE THE RVAD, AND THE PATIENT REMAINED AT THE HOSPITAL FOR FURTHER MONITORING. IT WAS REQUESTED THAT THE LOG FILES BE READ AND EXPLAINED FOR DIFFERENT READINGS. THE LOG FILES WERE REVIEWED AND CONTAINED ONSET OF INCREASED PUMP POWER/FLOW, HARMONIC COMPENSATION/MOTOR INSTABILITY FAULTS, ELEVATED LEFT VENTRICULAR ASSIST DEVICE (LVAD) TEMPERATURE, HIGH CURRENT FAULTS AS WELL AS VARIOUS ALARMS ASSOCIATED WITH THE LVAD OFF CONDITION. THIS DATA SUGGESTED THAT THE PUMP WAS HAVING DIFFICULTY MAINTAINING THE ROTOR SET SPEED (POSSIBLE OBSTRUCTION). THE PERIODIC LOG FILE ALSO INDICATED THE PUMP STOP TO OCCUR BETWEEN READING OF 2:47AM AND 3:47AM ON (B)(6) 2025. THE EVENT LOG FILE SUBMITTED CONTAINED VARIOUS ALARMS ASSOCIATED WITH THE OBSTRUCTION AND PUMP OFF CONDITION. IT WAS ASKED IF THE DIFFERENT READINGS QUESTION BE CLARIFIED. TWO SETS OF FILES WERE SUBMITTED FOR REVIEW. THEY APPEARED TO BE FROM THE SAME PUMP BUT FROM 2 DIFFERENT DAYS.Event Description: PUMP THROMBOSIS WAS CONFIRMED TO BE THE CAUSE OF THE ISSUES FOR THE RIGHT SIDED PUMP. THE PATIENT HAD RIGHT HEART FAILURE PRIOR TO LVAD IMPLANT. THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE.Event Description: THIS WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: DATE OF DEATH AND DATE OF EVENT CORRECTED. SECTION D6A: DATE OF IMPLANT CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE AND MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING THROMBOEMBOLISM AND DEATH. SECTION 6 OF THE IFU ENTITLED ¿PATIENT CARE AND MANAGEMENT¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 ALSO PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR VALUES. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: THE PATIENT'S RIGHT SIDE PUMP IS CAPTURED UNDER MFR. REPORT # 2916596-2025-00545. THIS REPORT CAPTURES THE PATIENT'S LEFT SIDE PUMP. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21306270&pc=MCX | 3004742232-2025-00028 | 2025-01-21 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-02-04 05:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation of Vessels; Obstruction/Occlusion; Vascular Dissection | P130005 |  | 1 | Event Description: IT WAS REPORTED DURING ADVANCEMENT OF THE VIPERWIRE GUIDE WIRE, A DISTAL DISSECTION OCCURRED. THE PHYSICIAN DECIDED TO CONTINUE WITH THE ORBITAL ATHERECTOMY TREATMENT AS THE DISSECTION WAS NOT NEAR THE TARGET LESION AREA IN THE LEFT ANTERIOR DESCENDING (LAD). LOW SPEED TREATMENTS WERE PERFORMED. POST ORBITAL ATHERECTOMY, LUMINAL GAIN WAS OBSERVED. AFTER A STENT WAS PLACED AND POST-DILATION WAS PERFORMED, IMAGING SHOWED A SMALL PERFORATION IN THE STENTED SEGMENT/PREVIOUS DISSECTED AREA. TO ADDRESS THE PERFORATION, A GRAFTMASTER WAS ATTEMPTED TO BE USED BUT WAS UNABLE TO CROSS THE LESION. A GUIDE EXTENSION CATHETER WAS USED AND AN INJECTION WAS ADMINISTERED TO MANAGE THE PERFORATION, HOWEVER, THIS LED TO SLOW FLOW AND NO RE-FLOW, WHICH RESULTED IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED, HOWEVER, AFTER SEVERAL HOURS IT WAS INDICATED THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, THE ADVERSE EVENT WAS NOT RELATED TO ORBITAL ATHERECTOMY OR THE STENT AND THE CAUSE OF DEATH WAS DUE TO INJECTION WHICH WAS ADMINISTERED TO MANAGE THE PERFORATION. IT WAS CLARIFIED THE VIPERWIRE CAUSED THE DISSECTION AND THE PERFORATION OCCURRED AFTER FORCEFUL DYE INJECTION AFTER POST DILATION AND POST STENT PLACEMENT.Event Description: IT WAS REPORTED DURING ADVANCEMENT OF THE VIPERWIRE GUIDE WIRE, A DISTAL DISSECTION OCCURRED. THE PHYSICIAN DECIDED TO CONTINUE WITH THE ORBITAL ATHERECTOMY TREATMENT AS THE DISSECTION WAS NOT NEAR THE TARGET LESION AREA IN THE LEFT ANTERIOR DESCENDING (LAD). LOW SPEED TREATMENTS WERE PERFORMED. POST ORBITAL ATHERECTOMY, LUMINAL GAIN WAS OBSERVED. AFTER A STENT WAS PLACED AND POST-DILATION WAS PERFORMED, IMAGING SHOWED A SMALL PERFORATION IN THE STENTED SEGMENT/PREVIOUS DISSECTED AREA. TO ADDRESS THE PERFORATION, A GRAFTMASTER WAS ATTEMPTED TO BE USED BUT WAS UNABLE TO CROSS THE LESION. A GUIDE EXTENSION CATHETER WAS USED AND AN INJECTION WAS ADMINISTERED TO MANAGE THE PERFORATION, HOWEVER, THIS LED TO SLOW FLOW AND NO RE-FLOW, WHICH RESULTED IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED, HOWEVER, AFTER SEVERAL HOURS IT WAS INDICATED THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, THE ADVERSE EVENT WAS NOT RELATED TO ORBITAL ATHERECTOMY OR THE STENT AND THE CAUSE OF DEATH WAS DUE TO INJECTION WHICH WAS ADMINISTERED TO MANAGE THE PERFORATION. IT WAS CLARIFIED THE VIPERWIRE CAUSED THE DISSECTION AND THE PERFORATION OCCURRED AFTER FORCEFUL DYE INJECTION AFTER POST DILATION AND POST STENT PLACEMENT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT CARDIAC ARREST, PERFORATION OF VESSELS, OBSTRUCTION, DEATH, AND RELATED SURGICAL INTERVENTION APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS LIKELY THAT THE REPORTED PATIENT CARDIAC ARREST, PERFORATION OF VESSELS, OBSTRUCTION, DEATH, AND RELATED SURGICAL INTERVENTION ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21306468&pc=DSQ | 2916596-2025-00635 | 2025-01-10 05:00 | Death | THORATEC CORPORATION | 2025-02-04 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO STROKE. IT WAS REPORTED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION.Event Description: IT WAS LATER REPORTED THAT PATIENT PASSED DUE TO AN ISCHEMIC STROKE AND SHOWED SYMPTOMS OF LEFT EYE DEVIATION. DEATH WAS SAID TO NOT HAVE BEEN DEVICE RELATED AND THAT THE DEVICE WAS OPERATING AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21306475&pc=OZD | 1220648-2025-25990 | 2025-01-17 05:00 | Death | ABIOMED, INC. | 2025-02-04 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED ISCHEMIA IN THEIR LEFT LOWER EXTREMITY FOLLOWING IMPELLA CP IMPLANT. THEIR LEG BECAME MOTTLED WITH AN ABSENCE OF PULSE. INTERVENTION WAS DISCUSSED, BUT THE PATIENT'S FAMILY EVENTUALLY MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED LIMB ISCHEMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE LIMB ISCHEMIA COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21306766&pc=MHX | 1218950-2025-000066 | 2025-01-09 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-02-04 05:00 | MHX | PATIENT INFORMATION CENTER IX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K183387 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON A PATIENT INFORMATION CENTER IX. IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN OBTAINING THE CLINICAL AUDIT LOGS FOLLOWING A PATIENT INCIDENT. IT WAS INDICATED THE PATIENT DEVELOPED AN SPO2 ISSUE AND PASSED AWAY; HOWEVER, IT WAS ALSO INDICATED THERE WAS NO IMPACT TO THE PATIENT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. Manufacturer Narrative: A PHILIPS TECHNICAL CONSULTANT (TC) SPOKE TO THE CUSTOMER BIOMEDICAL ENGINEER (BIOMED) MANAGER. THE BIOMED CONTACTED THE NURSE/CLINICAL QUALITY IMPROVEMENT SPECIALIST WHO STATED THEY WERE REQUESTING THE LOGS FILES TO CONFIRM THAT THERE WERE NO DELAYS IN NOTIFICATION OR TREATMENT OF THE PATIENT AND TO HOPEFULLY VERIFY THAT THE PATIENT WAS ON CONTINUOUS PULSE OXIMETRY IF POSSIBLE. THE BIOMED LATER CONFIRMED WITH THE TC THAT THERE WAS NO DELAY; THEY JUST WANTED TO SEE WHAT O2 SENSOR WAS READING AT THE MOMENT. BASED ON THE INFORMATION AVAILABLE, THERE WAS NO ALLEGED PRODUCT FAILURE; THE CUSTOMER JUST NEEDED THE LOG FILES FOR REVIEW. THIS COMPLAINT IS DEEMED NOT REPORTABLE WITH PHILIPS DUE THERE IS NO ALLEGATION OF PRODUCT DEFICIENCY OR OTHER INDICATION THAT A PRODUCT PERFORMANCE ISSUE OR MALFUNCTION HAS OCCURRED. THE CUSTOMER HAS REQUESTED ASSISTANCE FROM PHILIPS ONLY. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21307067&pc=MHX | 1218950-2025-000049 | 2025-01-04 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-02-04 05:00 | MHX | PATIENT INFORMATION CENTER IX | No Audible Alarm | Insufficient Information | K153702 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY WHILE CONNECTED TO THE MONITOR AND THE CUSTOMER REQUESTED ASSISTANCE WITH THE ALARM LOGS AND CONFIGURATION SETTINGS. BIOMED INDICATED THE BEDSIDE MONITOR ALARMED AS EXPECTED BUT THE CLINICAL MANAGER WANTED TO REVIEW THE AUDIT LOG, SPECIFICALLY RELATED TO SPO2 SENSOR OFF INOP TECHNICAL ALARMS, SPO2 SENSOR OFF CONFIGURATION, AND ECG AND HEART RATE EXTREME ALARMS. THERE WAS NO INDICATION THE MONITOR CONTRIBUTED TO THE DEATH; HOWEVER, GOOD FAITH EFFORTS ARE BEING PERFORMED FOR CONFIRMATION.Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY WHILE CONNECTED TO THE MONITOR AND THE CUSTOMER REQUESTED ASSISTANCE WITH THE ALARM LOGS AND CONFIGURATION SETTINGS. THE HOSPITAL DID NOT SHARE THE CAUSE OF DEATH AS THEY FELT IT WAS PROTECTED HEALTH INFORMATION (PHI). BIOMED INDICATED THE BEDSIDE MONITOR ALARMED AS EXPECTED BUT THE CLINICAL MANAGER WANTED TO REVIEW THE AUDIT LOG, SPECIFICALLY RELATED TO SPO2 SENSOR OFF INOP TECHNICAL ALARMS, SPO2 SENSOR OFF CONFIGURATION, AND ECG AND HEART RATE EXTREME ALARMS. GOOD FAITH EFFORTS (GFE) CONFIRMED THE DEVICE DID NOT CONTRIBUTE TO THE DEATH SINCE THE ISSUE IN QUESTION WAS SPO2 SENSOR OFF. THE SYSTEM GENERATED EKG ALARMS AND OTHER ALARMED APPROPRIATELY. Manufacturer Narrative: A PHILIPS REMOTE APPLICATION SPECIALIST (RAS) INTERVIEWED THE CUSTOMER, AND THEY REPORTED ON (B)(6) 2024 (TIME 2200 TO (B)(6) 2024 1:30AM) THEIR PATIENT (ICU BED239) PASSED AWAY. THE CUSTOMER STATED THAT THE PATIENT'S BEDSIDE MONITOR DID ALARM AS EXPECTED; HOWEVER, THE CLINICAL MANAGER (CM) HAD INQUIRIES REGARDING THE ALARM LOG AND THE APPLICATION CONFIGURATION SETTINGS. THE CM HAD THE FOLLOWING QUESTIONS: 1) THE CLINICAL AUDIT LOG DOES NOT SHOW THE "SPO2 SENSOR OFF" INOP ALARM EVENT WHY? 2) IS IT POSSIBLE TO CHANGE THE "SPO2 SENSOR OFF" ALARM SEVERITY CONFIG SETTING? 3) HOW IT THE ECG/HR EXTREME ALARM BEHAVIORS CONFIGURED? THE PHILIPS RAS REVIEWED FOLLOWING QUESTIONS AND NOTED A CENTRAL STATION WITH SOFTWARE B.02.X CANNOT LOG A "SENSOR OFF INOP" ALARM AS AN EVENT IN THE CLINICAL AUDIT LOG. THE CLINICAL AUDIT CAN LOG THE FOLLOWING INOPS EVENTS: ALL SERVER YELLOW OR RED INOP MESSAGES. ALL BATTERY INOP MESSAGES. ECG LEADS OFF. TRANSMITTER OFF. NO SPO2T. BATT LOW. THE CENTRAL STATION SERVER WITH SOFTWARE C AND GREATER WILL TRACK ALL THE TECHNICAL INOP ALARMS EVENT ON THE AUDIT LOG. THE RAS EXPLAINED THE SPO2 INOP TECHNICAL ALARMS ARE NOT CONFIGURABLE. THESE ALARMS CANNOT BE CONFIGURED TO YELLOW OR RED ALARM SEVERITY. THE CENTRAL STATION LOCAL SERVER INOPERATIVE SOUND VOLUME CAN BE SET TO BE LOUDER THAN THE CURRENT VOLUME. THE RAS EMAILED REFERENCES TO THE CUSTOMER, AND THEY CONFIRMED THE MONITOR GENERATED THE PATIENT'S HEART RATE LOW ALARMS AND THE ALARMS WERE SILENCED. CENTRAL STATION LOGS WERE REQUESTED BUT COULD NOT BE PROVIDED. THE CUSTOMER NOTED THE SYSTEM WORKED AS DESIGNED SINCE SPO2 SENOR BEING OFF WAS A BLUE ALARM AND DOES NOT RE-ALARM LIKE YELLOW AND RED ALARMS. THE QUESTIONS FROM THE CLINICAL MANGER WERE ANSWERED TO THEIR SATISFACTION AS THEY NOW UNDERSTAND THAT SPO2 SENSOR ALARM IS A BLUE ALARM. NO FURTHER INVESTIGATION OR ACTION IS WARRANTED AT THIS TIME. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21307114&pc=OZD | 1220648-2025-25991 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-02-04 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Sepsis | P140003 |  | 1 | Event Description: PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT AND ESCALATED TO AN IMPELLA 5.5 FOUR DAYS LATER DUE TO THE EXPECTED PROLONGATION OF SUPPORT. APPROXIMATELY AFTER 11 DAYS ON IMPELLA 5.5 SUPPORT, LEAKING FROM PURGE CASSETTE WAS OBSERVED. PER THE CLINICAL ENGINEER LIQUID WAS LEAKING FROM THE PURGE CASSETTE ON THE FOURTH DAY AFTER STARTING USE. CONSEQUENTLY, THE PURGE SET WAS REPLACED. HOWEVER, 16 DAYS AFTER START OF SUPPORT, ABNORMAL NOISE WAS HEARD AROUND THE PURGE CASSETTE INSERTION AREA OF THE AUTOMATED IMPELLA CONTROLLER (USED WITH BOTH THE IMPELLA CP AND IMPELLA 5.5), WHICH PROMPTED REPLACEMENT OF THE PURGE SET. THE ABNORMAL NOISE WAS RESOLVED AFTER REPLACING THE PURGE HOUSING. AT SOME POINT THE PATIENT DEVELOPED SEPSIS AS CASE WRAP-UP NOTED ¿CARDIAC FUNCTION RECOVERS POORLY, AND EVENTUALLY SEPSIS DISAPPEARS.¿ THE TIMEFRAME FOR THIS ADVERSE EVENT AND ITS RELATIONSHIP TO THE IMPELLA PUMP IS UNCLEAR. HOWEVER, THE PATIENT WOULD EVENTUALLY EXPIRE BASED ON REFLECTED CASE NOTE ¿EXPIRED ON SUPPORT,¿ WHICH OCCURRED ON THE 25TH DAY OF SUPPORT. REPORTEDLY THE PATIENT WENT INTO MULTIPLE ORGAN FAILURE CAUSED BY SEPSIS. ANTIBIOTICS WERE ADMINISTERED. NO DETAIL INFORMATION SURROUNDING THE PATIENT'S CONDITION, THE SEPSIS AND DEATH WERE PROVIDED, AND WITH LIMITED INFORMATION, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS A CONTRIBUTING FACTOR IN THE PATIENT¿S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MEDICAL DEVICE REPORT FOR THE AUTOMATED IMPELLA CONTROLLER SERIAL NUMBER (B)(6) WITH DATE OF EVENT 05-JAN-2025 AND IDENTIFIER ENDING IN (B)(6). SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS: ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS HAS BEEN COMPLETED. NOT ENOUGH CLINICAL INFORMATION WAS PROVIDED TO DETERMINE THE ROOT CAUSE OF THIS FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21307129&pc=NPW | 2015691-2025-00738 | 2024-12-13 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-04 05:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device | Cardiac Arrest; Insufficient Information | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE VALVE IN TRICUSPID POSITION THAT WAS COMPLETED WITHOUT ANY REPORTED EVENTS. INTERNAL REVIEW OF PROCEDURAL IMAGING IDENTIFIED THE SEPTAL SIDE OF THE VALVE WAS DEPLOYED TOO VENTRICULAR. ON POST-OPERATIVE DAY (POD) 1, THE TTE INDICATED THE EVOQUE VALVE APPEARED IN THE SAME POSITION AND WAS STABLE. THERE WAS TRACE TR AND NO PVL. THE TV MEAN GRADIENT WAS 1MMHG AT 68BPM. ON POD 2, PATIENT WAS DISCHARGED HOME. ON POD 3, THE PATIENT EXPIRED AT HOME AND WAS IN ASYSTOLE WHEN PARAMEDICS ARRIVED. NO AUTOPSY WILL BE PERFORMED AND NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. THE COMPLAINT FOR VALVE DEPLOYED TOO VENTRICULAR WAS CONFIRMED WITH OBJECTIVE EVIDENCE. THE IMAGING EVALUATION DEMONSTRATED THAT THE SEPTAL SIDE OF THE VALVE WAS DEPLOYED TOO VENTRICULAR. LIMITED IMAGING WAS PROVIDED FOR REVIEW, SO A ROOT CAUSE COULD NOT BE DETERMINED. BASED ON EXTENSIVE COMPLAINT INVESTIGATIONS, THE ROOT CAUSE FOR EVENTS MORE LIKELY DUE TO VARIOUS PROCEDURAL, PATIENT, IMAGING OR A COMBINATION OF THESE FACTORS, SUCH AS DIFFICULT OR LACK OF LEAFLET CAPTURE, LEAFLET PLASTERING (SEPTAL, FIXED LEAFLET TO SEPTAL WALL), AND DEVICE SHADOWING. THESE EVENTS ARE NOT ASSOCIATED WITH MANUFACTURING NON-CONFORMANCES OR DEFICIENCIES. DEVICE MALPOSITION IS A KNOWN POTENTIAL ADVERSE EVENT LISTED IN THE EVOQUE IFU. THE EVOQUE IFU AND TRAINING MANUALS PROVIDES INSTRUCTIONS ON DEVICE USE, WARNINGS, CAUTIONS INTENDED TO MITIGATE THESE EVENTS AND ARE SUFFICIENT TO ADDRESS THESE EVENTS. THESE EVENTS WILL CONTINUE TO BE MONITORED AND COMPLAINTS TRENDING, AND CONTROL LIMITS ARE MANAGED AND ASSESSED. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21307188&pc=DRA | 2124215-2025-05694 | 2025-01-14 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-04 05:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE TO TREAT A PERSISTENT ATRIAL FIBRILLATION A FARADRIVE STEERABLE SHEATH CLEAR WAS SELECTED FOR USE. THE PATIENT WAS INTUBATED AND IN ATRIAL FLUTTER. THE PHYSICIAN MAPPED THE RIGHT ATRIUM AND WENT TO THE LEFT SIDE. AFTER MAPPING, THE PHYSICIAN DECIDED TO INSERT THE VERSACROSS PIGTAIL INTO THE LEFT SUPERIOR PULMONARY VEIN (LSPV) AND THEN EXCHANGE IT FOR THE FARADRIVE SHEATH. AFTER GETTING TRANSEPTAL, THE PHYSICIAN MOVED THE SHEATH BACK AND FORTH A COUPLE TIMES BEFORE PULLING THE SHEATH INTO THE RIGHT ATRIUM AND ADVANCING THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER INTO THE LEFT ATRIUM. DURING THIS PERIOD, THE PATIENT BLOOD PRESSURE DROPPED AND THE ICE IMAGING BECAME FOGGY, A LARGE EFFUSION WAS OBSERVED. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS, BUT AS THE EVENT PROGRESSED, AN OPEN CHEST SURGERY WAS CALLED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED. WHEN THE SURGEON OPENED THE CHEST IT WAS FOUND THAT THE WIRE THROUGH THE APPENDAGE. THE PATIENT HAD PASSED AWAY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21307526&pc=NIK | 9614453-2025-00538 | 2024-09-04 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-04 05:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SPOUSE THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DID NOT TRIGGER AN ALERT. THE PATIENT SUBSEQUENTLY EXPIRED AND THE PATIENT'S SPOUSE INQUIRED WHY THE DEVICE DIDN'T PREVENT THE PATIENT'S DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: 5076-45 LEAD; IMPLANT DATE: (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21308357&pc=NVN | 2017865-2025-12285 | 2024-11-21 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-04 05:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MFR REPORT NUMBER: 2017865-2025-12284. RELATED MFR REPORT NUMBER: 2017865-2025-12286. IT WAS REPORTED THAT A PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER, ATRIAL LEAD, AND RIGHT VENTRICULAR LEAD WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL EXAMINATION DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21308414&pc=FRN | 3004548776-2025-00053 | 2024-08-21 04:00 | Death | FRESENIUS VIAL S.A.S. | 2025-02-04 05:00 | FRN | AGILIA SP MC WIFI CA | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED: SCREW ANCHOR CAME OUT OF BRACKET AND MAINS BOARD KIT. REPLACED KIT WITH NEW ONE. UPGRADED SOFTWARE TO NEW VERSION 2.2D PM COMPLETED USING AGILIA PARTNER VERSION V2R 4.0.3.21072 WI-FI CONFIGURED SQL AND CENTERIUM SERVER CHECKED FOR CONNECTIVITY AND DATA DOWNLOAD: PASS. REPORTING AS A CONSERVATIVE MEASURE. NO ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.Event Description: DEVICE HISTORY RECORD WAS REVIEWED, NO EVENT LINKED WITH THE REPORTED ISSUE WAS FOUND. DEVICE LOG WAS NOT PROVIDED; THEREFORE NO REVIEW COULD BE PERFORMED. THE REPORTED DEVICE AND THE DEFECTIVE SPARE PART WERE NOT RETURNED TO OUR LABORATORY FOR ANALYSIS. AS A RESULT, NO INVESTIGATION COULD BE CONDUCTED, AND THE REPORTED EVENT COULD NOT BE REPRODUCED OR CONFIRMED BY OUR LABORATORY. HOWEVER, A REPAIR REPORT WAS PROVIDED, STATING THAT THE BRACKET AND MAINS BOARD KIT WAS REPLACED, AND THE DEVICE WAS SUBSEQUENTLY OPERATIONAL. BASED ON THE AVAILABLE INFORMATION AND THE FACT THAT THE DEVICE/PART WAS NOT RETURNED, THE ROOT CAUSE OF THE REPORTED ISSUE REMAINS UNKNOWN (NOT RETURNED). THIS COMPLAINT HAS BEEN REGISTERED FOR STATISTICAL MONITORING. IF FURTHER EVIDENCE IS PROVIDED LATER, WE WILL RE-OPEN THE COMPLAINT AND PURSUE THE INVESTIGATION. TO OUR KNOWLEDGE THIS COMPLAINT IS NOT BEING RELATED TO PATIENT SAFETY. THIS COMPLAINT IS CONSIDERED AS: NOT INVESTIGATED. THE TREND IS: NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21308619&pc=NVZ | 2017865-2025-12284 | 2024-11-21 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-04 05:00 | NVZ | ACCENT DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: RELATED MFR REPORT NUMBER: 2017865-2025-12285. RELATED MFR REPORT NUMBER: 2017865-2025-12286. IT WAS REPORTED THAT A PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER, ATRIAL LEAD, AND RIGHT VENTRICULAR LEAD WERE EXPLANTED. Manufacturer Narrative: THE DEVICE WAS RECEIVED FROM THE FIELD WITH BATTERY VOLTAGE OPERATING AT ELECTIVE REPLACEMENT INDICATOR (ERI) LEVEL. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. REVIEW OF THE DEVICE IMAGE INDICATES WAS NORMAL BATTERY DEPLETION. THE IMPLANT DURATION EXCEEDS THE PROJECTED LONGEVITY BASED ON AVERAGE MONTHLY CURRENT INDICATING NORMAL BATTERY DEPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21308661&pc=NVN | 2017865-2025-12286 | 2024-11-21 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-04 05:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MFR REPORT NUMBER: 2017865-2025-12284. RELATED MFR REPORT NUMBER: 2017865-2025-12285. IT WAS REPORTED THAT A PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER, ATRIAL LEAD, AND RIGHT VENTRICULAR LEAD WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21308864&pc=NIK | 3004209178-2025-02054 | 2024-12-25 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-04 05:00 | NIK | AMPLIA MRI QUAD CRT-D SURESCAN | Protective Measures Problem | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENTS FAMILY MEMBER THAT THE PATIENT IS DECEASED. THE FAMILY MEMBER REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS CONTINUOUSLY EMITTING AN ALERT TONE AND EVERY TIME THE PATIENT RAISED THEIR HAND A FLAT LINE FOUND WHICH IS INCREDIBLY DISTRACTING WHILE THE PATIENT WAS IN HOSPICE. THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309055&pc=LOX | 2124215-2025-05048 | 2025-01-12 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-04 05:00 | LOX | EMERGE? | Detachment of Device or Device Component | Bradycardia; Chest Pain; Low Blood Pressure/ Hypotension; Foreign Body In Patient | K113220 |  | 1 | Event Description: IT WAS REPORTED THAT SHAFT BREAK AND THROMBOSIS OCCURRED, RESULTING IN PATIENT DEATH. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT. CORONARY ANGIOGRAPHY REVEALED THAT THE LEFT ANTERIOR DESCENDING (LAD) ARTERY WAS COMPLETELY OCCLUDED FROM THE PROXIMAL SEGMENT, WITH CALCIFICATION VISIBLE DISTAL TO THE OCCLUSION. RE-EXAMINATION ANGIOGRAPHY REVEALED A THROMBUS IN THE PROXIMAL TO MID-SEGMENT OF THE ANTERIOR DESCENDING ARTERY, WITH A 90% STENOSIS IN THE DISTAL SEGMENT. A 2.50MM X 15MM EMERGE BALLOON CATHETER WAS ADVANCED TO TREAT THE LESION DILATED TO 12 ATMOSPHERES FOR 30 SECONDS. WHEN ANOTHER PRE-DILATION NON-BOSTON SCIENTIFIC BALLOON WAS INSERTED, THE PHYSICIAN BELIEVED THAT THE EMERGE BALLOON TIP HAD FRACTURED WITHIN THE GUIDING CATHETER AND THE NON-BSC BALLOON PUSHED THE DEVICE FRAGMENT INTO THE DISTAL LAD. ACUTE THROMBOSIS WAS IDENTIFIED, AND NON-BOSTON SCIENTIFIC STENTS WERE IMPLANTED IN THE LAD. THE PATIENT EXPERIENCED CHEST PAIN AND DISCOMFORT. MEDICATION WAS ADMINISTERED AND ANGIOGRAPHY SHOWED GOOD BLOOD FLOW IN THE LAD. THE PATIENT THEN EXPERIENCED A DROP IN BLOOD PRESSURE AND OXYGEN SATURATION. TRACHEAL INTUBATION WAS PERFORMED, AND AN INTRA-AORTIC BALLOON PUMP WAS INSERTED. ADDITIONAL NON-BOSTON SCIENTIFIC STENTS WERE IMPLANTED, AND POST DILATION WAS PERFORMED WITH A QUANTUM BALLOON. REPEAT ANGIOGRAPHY REVEALED GOOD BLOOD FLOW IN THE LAD, HOWEVER, THE PATIENT'S BLOOD PRESSURE CONTINUED TO DROP AND HEART RATE DECREASED SIGNIFICANTLY. EXTERNAL CHEST COMPRESSIONS WERE PERFORMED, AND ADDITIONAL MEDICATION WAS ADMINISTERED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU), WHERE ONE HOUR LATER WAS CONFIRMED DECEASED.Event Description: IT WAS REPORTED THAT SHAFT BREAK AND THROMBOSIS OCCURRED, RESULTING IN PATIENT DEATH. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT. CORONARY ANGIOGRAPHY REVEALED THAT THE LEFT ANTERIOR DESCENDING (LAD) ARTERY WAS COMPLETELY OCCLUDED FROM THE PROXIMAL SEGMENT, WITH CALCIFICATION VISIBLE DISTAL TO THE OCCLUSION. A 2.50MM X 15MM EMERGE BALLOON CATHETER WAS ADVANCED TO TREAT THE LESION DILATED TO 12 ATMOSPHERES FOR 30 SECONDS. WHEN ANOTHER PRE-DILATION NON-BOSTON SCIENTIFIC BALLOON WAS INSERTED, THE PHYSICIAN BELIEVED THAT THE EMERGE BALLOON TIP HAD FRACTURED WITHIN THE GUIDING CATHETER AND THE NON-BSC BALLOON PUSHED THE DEVICE FRAGMENT INTO THE DISTAL LAD. ACUTE THROMBOSIS WAS IDENTIFIED AND NON-BOSTON SCIENTIFIC STENTS WERE IMPLANTED IN THE LAD. THE PATIENT EXPERIENCED CHEST PAIN AND DISCOMFORT. MEDICATION WAS ADMINISTERED AND ANGIOGRAPHY SHOWED GOOD BLOOD FLOW IN THE LAD. THE PATIENT THEN EXPERIENCED A DROP IN BLOOD PRESSURE AND OXYGEN SATURATION. TRACHEAL INTUBATION WAS PERFORMED, AND AN INTRA-AORTIC BALLOON PUMP WAS INSERTED. ADDITIONAL NON-BOSTON SCIENTIFIC STENTS WERE IMPLANTED, AND POST DILATION WAS PERFORMED WITH A QUANTUM BALLOON. REPEAT ANGIOGRAPHY REVEALED GOOD BLOOD FLOW IN THE LAD, HOWEVER, THE PATIENT'S BLOOD PRESSURE CONTINUED TO DROP AND HEART RATE DECREASED SIGNIFICANTLY. EXTERNAL CHEST COMPRESSIONS WERE PERFORMED, AND ADDITIONAL MEDICATION WAS ADMINISTERED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU), WHERE ONE HOUR LATER WAS CONFIRMED DECEASED. IT WAS FURTHER CONFIRMED THAT RE-EXAMINATION ANGIOGRAPHY REVEALED ACUTE THROMBOSIS AFTER BALLOON PRE-TREATMENT WITH THE NON-BSC BALLOON AND THE EMERGE BALLOON IN THE PROXIMAL TO MID-SEGMENT OF THE LAD, WITH A 90% STENOSIS IN THE DISTAL SEGMENT. THEREFORE, AN EMERGE BALLOON WAS INSERTED INTO THE DISTAL SEGMENT LAD FOR PRE-DILATATION PRIOR TO THE NON-BSC STENT IMPLANTATIONS. Manufacturer Narrative: G4 - PREMARKET: K163174. DEVICE EVALUATED BY MFR: THE DEVICE FG EMERGE MR, US 2.50MM X 15MM WAS RETURNED FOR ANALYSIS. VISUAL AND TACTILE INSPECTION IDENTIFIED MULTIPLE KINKS ALONG THE HYPOTUBE SHAFT. VISUAL, TACTILE AND MICROSCOPIC ANALYSIS IDENTIFIED A BREAK IN THE INNER LUMEN AND BALLOON, 22.8CM FROM THE PORT EXCHANGE WITH THE DISTAL SECTION NOT RETURNED. MICROSCOPIC ANALYSIS OF THE BALLOON PROFILE REVEALED A COMPLETE CIRCUMFERENTIAL TEAR 22.8CM FROM THE PORT EXCHANGE WITH THE DISTAL SECTION NOT RETURNED. THE BUMPER TIP SHOWED NO SIGNS OF DISTAL TIP DAMAGE. THE DISTAL MARKERBAND WAS NOT RETURNED, HOWEVER, NO ISSUES WERE NOTED WITH THE PROXIMAL MARKERBAND. NO OTHER DEVICE ISSUES WERE IDENTIFIED DURING RETURNED PRODUCT ANALYSIS. Manufacturer Narrative: G4 - PREMARKET: K163174.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309418&pc=MAE | 2135147-2025-00544 | 2025-01-10 05:00 | Death | ABBOTT MEDICAL | 2025-02-04 05:00 | MAE | AMPLATZER PICCOLO | Migration or Expulsion of Device | Bradycardia; Low Blood Pressure/ Hypotension; Obstruction/Occlusion; Embolism/Embolus | P020024 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 4 MM X 2 MM AMPLATZER PICCOLO OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING A 4F AMPLATZER TORQVUE DELIVERY SYSTEM. THE PATIENT WAS A PREMATURE INFANT (DOB (B)(6) 2024) WITH A HEMODYNAMICALLY SIGNIFICANT PATENT DUCTUS ARTERIOSUS (PDA). PATIENT WAS REFERRED FOR TRANSCATHETER CLOSURE FOLLOWING THREE (3) PRIOR FAILED ATTEMPTS TO CLOSE THE PDA WITH MEDICAL THERAPY. THE PATIENT WEIGHED 1050 GRAMS AT 4 WEEKS OF LIFE. AN ANGIOGRAM DEMONSTRATED A PDA WITH A MINIMAL DIAMETER OF 2.5 MM, AORTIC END DIAMETER OF 3.3 MM AND A LENGTH OF 11.9 MM. THE 04MM X 02MM AMPALTZER PICCOLO OCCLUDER WAS DEPLOYED INTRA-DUCTAL VIA RIGHT FEMORAL VENOUS ACCESS WITHOUT DIFFICULTY. IMPLANTING TEAM UTILIZED THE PICCOLO PULSE CHECKLIST. HOWEVER, FOLLOWING DEVICE RELEASE THERE WAS EVIDENCE OF PROTRUSION INTO THE LEFT PULMONARY ARTERY (LPA) CAUSING PARTIAL OBSTRUCTION AND A DECISION WAS MADE TO ATTEMPT RETRIEVAL OF THE DEVICE WITH A 5 MM GOOSENECK SNARE AND SUBSEQUENT REPOSITIONING. HOWEVER, DURING THE ATTEMPT TO RETRIEVE THE DEVICE THE DEVICE WAS INADVERTENT DISLODGED INTO THE MAIN PULMONARY ARTERY AND ADDITIONAL ATTEMPTS WERE MADE TO RETRIEVE THE DEVICE WITH A 5 MM GOOSENECK SNARE. IN THE PROCESS OF ATTEMPTING TO RETRIEVE THE DEVICE WITH A SNARE, THE PICCOLO OCCLUDER EMBOLIZED FURTHER INTO THE RIGHT PULMONARY ARTERY (RPA) AND ADDITIONAL ATTEMPTS WERE MADE TO RETRIEVE THE OCCLUDER USING A COMBINATION OF A 4 FRENCH COOK FLEXOR SHEATH (45 CM) AND AN INNER 4 FRENCH TERUMO GLIDE CATHETER. HOWEVER, THE DEVICE COULD NOT BE RETRIEVED, AND THE PATIENT BEGAN TO DEVELOP BRADYCARDIA AND HAD HEMODYNAMIC INSTABILITY. THE PATIENT WAS TAKEN TO SURGERY WHERE THE DEVICE WAS EXPLANTED AND THE PDA LIGATED SURGICALLY. THE PATIENT'S CHEST WAS LEFT OPEN POST SURGERY. THE PATIENT POST-OPERATIVELY INITIALLY WAS IMPROVING BUT SUBSEQUENTLY WORSENED AND A DECISION WAS MADE TO WITHDRAW SUPPORT ON (B)(6) 2025. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT FOLLOWING PICCOLO DEVICE RELEASE THERE WAS EVIDENCE OF PROTRUSION INTO THE LEFT PULMONARY ARTERY (LPA) CAUSING PARTIAL OBSTRUCTION. AN UNSUCCESSFUL ATTEMPT WAS MADE TO RETRIEVAL OF THE DEVICE; HOWEVER, THE DEVICE INADVERTENTLY DISLODGED INTO THE MAIN PULMONARY ARTERY AND ADDITIONAL ATTEMPTS WERE MADE TO RETRIEVE THE DEVICE. THE PICCOLO OCCLUDER EMBOLIZED FURTHER INTO THE RIGHT PULMONARY ARTERY (RPA). AFTER ADDITIONAL UNSUCCESSFUL ATTEMPTS OF DEVICE RETRIEVAL THE PATIENT BEGAN TO DEVELOP BRADYCARDIA AND HAD HEMODYNAMIC INSTABILITY. THE PATIENT WAS TAKEN TO SURGERY WHERE THE DEVICE WAS EXPLANTED AND PDA LIGATED. POST-OPERATIVELY THE PATIENT INITIALLY IMPROVED, BUT SUBSEQUENTLY DETERIORATED AND A DECISION WAS MADE TO WITHDRAW SUPPORT AND THE INFANT EXPIRED ON THE SIXTH POST-OPERATIVE DAY. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE DUCTAL MEASUREMENTS AS REPORTED FROM THE FIELD WAS PERFORMED. PER THE SIZING TABLE IN THE INSTRUCTIONS FOR USE, THE RECOMMENDED SIZE DEVICE WAS USED. INFORMATION FROM FIELD INDICATED THAT THE PICCOLO DEVICE WAS DEPLOYED INTRA-DUCTAL VIA RIGHT FEMORAL VENOUS ACCESS WITHOUT DIFFICULTY. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE MEDICAL REVIEW PERFORMED AT ABBOTT, DUE TO THE PROLONGED ATTEMPTS TO RETRIEVE THE DEVICE ALONG WITH THE SURGERY TO EXPLANT THE DEVICE, THE PATIENT MOST LIKELY HAD COMPLICATIONS FROM SURGERY THAT ULTIMATELY LED TO THE POST-OPERATIVE CLINICAL DECLINE AND FATALITY. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED DEVICE MIGRATION COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED PATIENT EFFECTS BRADYCARDIA AND HYPOTENSION APPEARS TO BE A CASCADING EFFECT FROM PROLONGED ATTEMPTS TO RETRIEVE THE DEVICE. THE CAUSE OF REPORTED PATIENT DEATH IS POSSIBLY RELATED TO THE COMPLICATION FROM SURGERY. THE REPORTED MEDICAL AND SURGICAL INTERVENTION WERE RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS EMBOLIZED DEVICE WAS ATTEMPTED TO BE SNARED AND WAS SURGICALLY REMOVED. THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309481&pc=QFG | 3013756811-2025-23073 | 2025-01-02 05:00 | Death | TANDEM DIABETES CARE | 2025-02-04 05:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DUE TO DIABETIC KETOACIDOSIS. ACCORDING TO THE PHYSICIAN, A REVIEW OF THE PUMP HISTORY REVEALED THAT THE CUSTOMER CHANGED THE CARTRIDGE TWO DAYS PRIOR TO THE EVENT, RESULTING IN BLOOD GLUCOSE (BG) READINGS DISPLAYED AS "HIGH." REPORTEDLY, THE PUMP FUNCTIONED AS INTENDED AND DELIVERED AUTOMATIC BOLUSES IN RESPONSE TO THE HIGH BG LEVELS. THE CUSTOMER ALSO DELIVERED MANUAL OVERRIDE BOLUSES VIA THE PUMP IN AN ATTEMPT TO ADDRESS THE HIGH BG LEVELS; HOWEVER, THE BG VALUES REMAINED HIGH. AS A RESULT OF THE CURRENT FINDINGS, THE PHYSICIAN STATED THAT "THE ALGORITHM IN THE PUMP HAS NOT BEEN THE PROBLEM, BUT PROBABLY THE DELIVERY" AND THAT THE EVENT WAS DUE TO "A VERY UNFORTUNATE USER ERROR." NO ADDITIONAL INFORMATION COULD BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309604&pc=FPA | 9616066-2025-00131 | 2024-12-15 05:00 | Death | SISTEMAS MEDICOS ALARIS S.A. DE C.V. | 2025-02-04 05:00 | FPA | BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | Air/Gas in Device | Brain Injury; Heart Block | K221327 |  | 1 | Event Description: MATERIAL: 2426-0007. BATCH#: 24109004, 24095357, 24109078. IT WAS REPORTED BY THE CUSTOMER THAT 3-4 INCIDENTS WITH AIR GETTING INTO THE LINE. VERBATIM: RCC RECEIVED A COMPLAINT VIA EMAIL. ¿ ADVERSE INCLUSIVE OF OUR TUBING AND PUMPS AND TRANSDUCER ¿ POST PUMP EVENT ¿ 3-4 INCIDENTS WITH AIR GETTING INTO THE LINE ¿ (B)(6) ¿ PART NUMBERS AND LOTS COMING VIA (B)(6) (LISTED BELOW) ¿ BD ALARIS TUBING SETS REMOVED AT THIS TIME ¿ TPN TUBING - NO PORTS. PART # 2426-0007 LOT #: 24109004, 24095357, 24109078. ADDITIONAL INFORMATION RECEIVED ON 16-JAN-2025. EVENT #1: (B)(6)2024 UMBILICAL ARTERIAL (UAC) AND VENOUS (UVC) LINES WERE PLACED ON A 27W PREMATURE INFANT SHORTLY AFTER ARRIVAL TO THE NICU. AN ALARIS IV PUMP, MODEL 8100 WITH ALARIS PC #8015, WAS USED WITH REF# (B)(4) INFUSION SET USED FOR BOTH LINES. THE ARTERIAL LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE PRIMED/FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. THE INFANTS CONDITION ABRUPTLY DETERIORATED WITHIN MINUTES. HEART RATE DROPPED; CYANOSIS WAS NOTED. ACLS MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH SHOWED AIR IN THE LEFT VENTRICLE OF THE HEART. THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2-3 MLS. THIS TUBING WAS NOT SAVED AND IS NOT AVAILABLE FOR TESTING. PATIENT DEATH OCCURRED 8 DAYS LATER. IV PUMP: MODEL 8100 WITH ALARIS PC #(B)(4), IV INFUSION SET: REF# (B)(4) LOT# 24109078 WHAT WAS THE PRIMARY AND SECONDARY CAUSES OF DEATH? COMPLICATIONS OF PREMATURE BIRTH ¿ 27W4DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION. EVENT #2: (B)(6)2024 UAC AND UVC IN USE ON ANOTHER NICU PT. WHO WAS 29W GESTATION. LINES SET UP ON (B)(6)2024 WITH THE SAME CONFIGURATION DESCRIBED IN THE (B)(6)2024 CASE ABOVE. ON (B)(6)2024 AT 1400, THE RN CHECKED PATENCY OF THE UAC LINE BY ASPIRATING AND FLUSHING IT. AT 1500, THE RN NOTED 2ML AIR BUBBLE IN THE TUBING. THIS WAS LOCATED DISTAL TO THE IV PUMP AND PAST THE TRANSDUCER SECTION OF THE LINE IN THE SPACE LEADING TO THE PATIENT¿S CATHETER. SHE WAS ABLE TO REMOVE THE AIR WITH ASPIRATION. NO HARM TO THE PATIENT. THE ENTIRE LINE WAS REPLACED. IV INFUSION PUMP TUBING LOT# BEING RETURNED TO BD FOR EXAMINATION. IV PUMP: MODEL 8100 WITH ALARIS PC #8015 IV INFUSION SET: REF# (B)(4), LOT# 24109078. EVENT #3: (B)(6)2025, UAC AND UVC PLACED ON NEW ADMIT TO NICU. SODIUM ACETATE WITH HEPARIN TO RUN THROUGH THE UAC. THE LINES WERE PRIMED WITH THE TRANSDUCER AND CHECKED FOR AIR; NO AIR WAS FOUND IN THE TUBING AT THAT TIME. ABOUT 10 MINUTES LATER, AIR WAS FOUND TO BE IN THE IV TUBING WAITING TO BE ATTACHED TO THE UVC CATHETER WHERE STARTER TPN D10. THIS LINE WAS UNCLAMPED AND FLUSHED THROUGH A SECOND TIME UNTIL ALL AIR BUBBLES WERE OUT BEFORE BEING CONNECTED TO THE UVC. ONCE PROVIDER ORDERED THE D20 PIGGY-BACK, THE D20 WAS CONNECTED TO THE Y PORT ON THE TPN D10. THE PUMP AT THAT POINT KEPT ALARMING AS OCCLUDED. THE TPN D10 AND D20 LINE WERE DOUBLE CHECKED AND BOTH LINES WERE FOUND TO HAVE AIR IN THE TUBING. THE LINES WERE UNHOOKED FROM THE PATIENT AND FLUSHED AGAIN UNTIL ALL AIR WAS OUT OF THE TUBING. THE TUBING WAS THEN CHECKED AGAIN BY ANOTHER RN AND DR. TO VERIFY NO AIR WAS IN THE TUBING BEFORE BEING RECONNECTED TO THE UVC. A UAC WAS NOT ABLE TO BE ESTABLISHED SO THE SODIUM ACETATE WITH HEPARIN WAS TO BE CONNECTED TO ANOTHER PORT OF THE UVC LINE. THE SODIUM ACETATE WAS LOOKED OVER FOR ANY AIR IN THE TUBING, AS IT HAD BEEN SITTING CLAMPED AND CAPPED FOR ABOUT AN HOUR SINCE BEING PRIMED. AIR WAS FOUND IN THE TUBING OF THE SODIUM ACETATE. THE LINES WERE FLUSHED AGAIN UNTIL ALL AIR BUBBLES WERE OUT. AIR BUBBLES KEPT APPEARING IN THE TUBING NEAR THE PORTS AND RUNNING DOWN THE LINE WHILE FLUIDS WERE RUNNING. STAFF WERE UNABLE TO GET THE LINE TO STOP HAVING AIR BUBBLES. A NEW TUBING/SOLUTIONS SET WAS USED. THE NEW SET OF TUBING/IV BAG OF SODIUM ACETATE DID NOT SEEM TO HAVE ANY PROBLEMS WITH AIR BUBBLES, EVEN AFTER RUNNING FOR A BIT OF TIME. NO HARM OCCURRED TO THE PATIENT. THIS TUBING SET IS BEING RETURNED FOR EXAMINATION. IV PUMP: MODEL 8100 WITH ALARIS PC #(B)(4), IV INFUSION SET: REF# (B)(4); LOT# 24109004 & 2409537Event Description: ADDITIONAL INFORMATION RECEIVED 02/28/2025. EVENT #1: (B)(6) 2024 UMBILICAL ARTERIAL (UAC) AND VENOUS (UVC) LINES WERE PLACED ON A 27W PREMATURE INFANT SHORTLY AFTER ARRIVAL TO THE NICU. AN ALARIS IV PUMP, MODEL 8100 WITH ALARIS PC #8015, WAS USED WITH REF# 2426-0007 INFUSION SET USED FOR BOTH LINES. THE ARTERIAL LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE PRIMED/FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. THE INFANTS CONDITION ABRUPTLY DETERIORATED WITHIN MINUTES. HEART RATE DROPPED; CYANOSIS WAS NOTED. ACLS MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH SHOWED AIR IN THE LEFT VENTRICLE OF THE HEART. THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2-3 MLS. THIS TUBING WAS NOT SAVED AND IS NOT AVAILABLE FOR TESTING. PATIENT DEATH OCCURRED 8 DAYS LATER. IV PUMP: MODEL 8100 WITH ALARIS PC #8015, IV INFUSION SET: REF# 2426-0007 LOT# 24109078. ADDITIONAL INFORMATION RECEIVED FOR EVENT 1 ON 28-FEB-2025. THIS WAS PULLED OVER FROM THE DUPLICATE FILE PR (B)(4) WHICH WAS OPENED IN ERROR. AN INFANT ADMITTED TO NICU AND UMBILICAL ARTERIAL (UAC) AND UMBILICAL VENOUS CATHETERS (UVC) WERE TO BE PLACED. PRIOR TO PLACEMENT, THE UAC AND UVC LINES WERE SPIKED AND HANGING IN THE IV PUMPS. THE PHYSICIAN INSERTED THE UAC AND UVC WITH ASSISTANCE FROM THE RN. FOLLOWING THE DRAWING OF 2 SETS OF LABS AND A X-RAY, THE LINES WERE CONNECTED TO UAC AND UVC. SHORTLY AFTER, THE RN NOTICED THE BEDSIDE MONITOR WENT BLANK AND THE RN BEGAN TO TROUBLESHOOT. THE RN NOTICED THE INFANT WAS VERY PALE WITH SHALLOW CHEST RISE. IMMEDIATELY, THE RN CALLED FOR ASSISTANCE AND INFANT ARRESTED. THE PHYSICIAN NOTICED AIR IN THE UAC TUBING. PROMPTLY, THE RN STOPPED THE LINE AND REMOVED THE AIR. THE PHYSICIAN WAS ABLE TO VISUALIZE AIR ON THE X-RAY.Event Description: IT WAS REPORTED THAT BD ALARIS PUMP MODULE SMARTSITE INFUSION SET HAD AIR IN LINE THE FOLLOWING INFORMATION WAS RECEIVED BY THE INITIAL REPORTER WITH THE FOLLOWING VERBATIM. IT WAS REPORTED BY THE CUSTOMER THAT 3-4 INCIDENTS WITH AIR GETTING INTO THE LINE.Event Description: NO ADDITIONAL INFORMATION. Manufacturer Narrative: NO PRODUCT OR PHOTO WAS RETURNED BY THE CUSTOMER. THE CUSTOMER COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED DUE TO THE PRODUCT NOT BEING RETURNED FOR FAILURE INVESTIGATION. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET. THE ROOT CAUSE OF THIS FAILURE COULD NOT BE IDENTIFIED WITHOUT A FAILURE INVESTIGATION. THIS INCIDENT HAS BEEN ADDED TO OUR DATABASE OF REPORTED INCIDENTS. OUR BUSINESS TEAM REGULARLY REVIEWS THE COLLECTED DATA FOR IDENTIFICATION OF EMERGING TRENDS. YOUR ASSISTANCE IN THIS MATTER HAS BEEN HELPFUL IN TREND IDENTIFICATION AND SUPPORTING OUR COMMITMENT TO CONTINUOUS QUALITY IMPROVEMENT. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED EVENT #1: (B)(6)2024 UMBILICAL ARTERIAL (UAC) AND VENOUS (UVC) LINES WERE PLACED ON A 27W PREMATURE INFANT SHORTLY AFTER ARRIVAL TO THE NICU. AN ALARIS IV PUMP, MODEL 8100 WITH ALARIS PC #8015, WAS USED WITH REF# (B)(4) INFUSION SET USED FOR BOTH LINES. THE ARTERIAL LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE PRIMED/FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. THE INFANTS CONDITION ABRUPTLY DETERIORATED WITHIN MINUTES. HEART RATE DROPPED; CYANOSIS WAS NOTED. ACLS MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH SHOWED AIR IN THE LEFT VENTRICLE OF THE HEART. THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2-3 MLS. THIS TUBING WAS NOT SAVED AND IS NOT AVAILABLE FOR TESTING. PATIENT DEATH OCCURRED 8 DAYS LATER. IV PUMP: MODEL 8100 WITH ALARIS PC #(B)(4), Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED BY THE INITIAL REPORTER 02/28/2025. AN INFANT ADMITTED TO NICU AND UMBILICAL ARTERIAL (UAC) AND UMBILICAL VENOUS CATHETERS (UVC) WERE TO BE PLACED. PRIOR TO PLACEMENT, THE UAC AND UVC LINES WERE SPIKED AND HANGING IN THE IV PUMPS. THE PHYSICIAN INSERTED THE UAC AND UVC WITH ASSISTANCE FROM THE RN. FOLLOWING THE DRAWING OF 2 SETS OF LABS AND A X-RAY, THE LINES WERE CONNECTED TO UAC AND UVC. SHORTLY AFTER, THE RN NOTICED THE BEDSIDE MONITOR WENT BLANK AND THE RN BEGAN TO TROUBLESHOOT. THE RN NOTICED THE INFANT WAS VERY PALE WITH SHALLOW CHEST RISE. IMMEDIATELY, THE RN CALLED FOR ASSISTANCE AND INFANT ARRESTED. THE PHYSICIAN NOTICED AIR IN THE UAC TUBING. PROMPTLY, THE RN STOPPED THE LINE AND REMOVED THE AIR. THE PHYSICIAN WAS ABLE TO VISUALIZE AIR ON THE X-RAY. NO PRODUCT OR PHOTO WAS RETURNED BY THE CUSTOMER. THE CUSTOMER COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED DUE TO THE PRODUCT NOT BEING RETURNED FOR FAILURE INVESTIGATION. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET. THE ROOT CAUSE OF THIS FAILURE COULD NOT BE IDENTIFIED WITHOUT A FAILURE INVESTIGATION. THIS INCIDENT HAS BEEN ADDED TO OUR DATABASE OF REPORTED INCIDENTS. OUR BUSINESS TEAM REGULARLY REVIEWS THE COLLECTED DATA FOR IDENTIFICATION OF EMERGING TRENDS. YOUR ASSISTANCE IN THIS MATTER HAS BEEN HELPFUL IN TREND IDENTIFICATION AND SUPPORTING OUR COMMITMENT TO CONTINUOUS QUALITY IMPROVEMENT. Manufacturer Narrative: THE CUSTOMER REPORTED AIR-IN-LINE, AND RETURNED ONE SAMPLE OF MATERIAL 2426-0007, LOT 24109078. UPON INITIAL VISUAL EXAMINATION, THE SET HAD NO DEFECTS OR ABNORMALITIES. THE SET WAS RUN AT 120ML/HR FOR 30 MINUTES IN ORDER TO REPLICATE AN AGGRESSIVE INFUSION. THERE WAS NO AIR IN LINE, LEAKS, OR OTHER ISSUES OBSERVED DURING THIS TIME. THE COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED. THE ROOT CAUSE FOR THE ISSUE COULD NOT BE FOUND WITHOUT A REPLICATED FAILURE. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21309606&pc=FPA | 9616066-2025-00133 | 2024-12-29 05:00 | Death | SISTEMAS MEDICOS ALARIS S.A. DE C.V. | 2025-02-04 05:00 | FPA | BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | Air/Gas in Device | No Clinical Signs, Symptoms or Conditions | K221327 |  | 1 | Event Description: IT WAS REPORTED THAT BD ALARIS PUMP MODULE SMARTSITE INFUSION SET HAD AIR IN LINE THE FOLLOWING INFORMATION WAS RECEIVED BY THE INITIAL REPORTER WITH THE FOLLOWING VERBATIM . IT WAS REPORTED BY THE CUSTOMER THAT 3-4 INCIDENTS WITH AIR GETTING INTO THE LINE.Event Description: MATERIAL: 2426-0007, BATCH#: 24109004, 24095357, 24109078; IT WAS REPORTED BY THE CUSTOMER THAT 3-4 INCIDENTS WITH AIR GETTING INTO THE LINE. VERBATIM: RCC RECEIVED A COMPLAINT VIA EMAIL. EMAIL(S) ATTACHED. ADVERSE INCLUSIVE OF OUR TUBING AND PUMPS AND TRANSDUCER, POST PUMP EVENT, 3-4 INCIDENTS WITH AIR GETTING INTO THE LINE, (B)(6), PART NUMBERS AND LOTS COMING VIA (B)(6) (LISTED BELOW), BD ALARIS TUBING SETS REMOVED AT THIS TIME, TPN TUBING - NO PORTS. PART # 2426-0007. LOT # : 24109004, 24095357, 24109078. ADDITIONAL INFORMATION RECEIVED ON 16-JAN-2025: EVENT #1: (B)(6) 2024 UMBILICAL ARTERIAL (UAC) AND VENOUS (UVC) LINES WERE PLACED ON A 27W PREMATURE INFANT SHORTLY AFTER ARRIVAL TO THE NICU. AN ALARIS IV PUMP, MODEL 8100 WITH ALARIS PC # (B)(4) , WAS USED WITH REF# 2426-0007 INFUSION SET USED FOR BOTH LINES. THE ARTERIAL LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE PRIMED/FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. THE INFANTS CONDITION ABRUPTLY DETERIORATED WITHIN MINUTES. HEART RATE DROPPED; CYANOSIS WAS NOTED. ACLS MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH SHOWED AIR IN THE LEFT VENTRICLE OF THE HEART. THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2-3 MLS. THIS TUBING WAS NOT SAVED AND IS NOT AVAILABLE FOR TESTING. PATIENT DEATH OCCURRED 8 DAYS LATER. IV PUMP: MODEL 8100 WITH ALARIS PC #(B)(4), IV INFUSION SET: REF# 2426-0007 LOT# 24109078. *******WHAT WAS THE PRIMARY AND SECONDARY CAUSES OF DEATH? COMPLICATIONS OF PREMATURE BIRTH ¿ 27W4DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION. EVENT #2: (B)(6) 2024 UAC AND UVC IN USE ON ANOTHER NICU PT. WHO WAS 29W GESTATION. LINES SET UP ON (B)(6) 2024 WITH THE SAME CONFIGURATION DESCRIBED IN THE (B)(6) 2024 CASE ABOVE. ON (B)(6) 2024 AT 1400, THE RN CHECKED PATENCY OF THE UAC LINE BY ASPIRATING AND FLUSHING IT. AT 1500, THE RN NOTED 2ML AIR BUBBLE IN THE TUBING. THIS WAS LOCATED DISTAL TO THE IV PUMP AND PAST THE TRANSDUCER SECTION OF THE LINE IN THE SPACE LEADING TO THE PATIENT¿S CATHETER. SHE WAS ABLE TO REMOVE THE AIR WITH ASPIRATION. NO HARM TO THE PATIENT. THE ENTIRE LINE WAS REPLACED. IV INFUSION PUMP TUBING LOT# BEING RETURNED TO BD FOR EXAMINATION. IV PUMP: MODEL 8100 WITH ALARIS PC # (B)(4). IV INFUSION SET: REF# 2426-0007, LOT# 24109078. EVENT #3: (B)(6) 2025, UAC AND UVC PLACED ON NEW ADMIT TO NICU. SODIUM ACETATE WITH HEPARIN TO RUN THROUGH THE UAC. THE LINES WERE PRIMED WITH THE TRANSDUCER AND CHECKED FOR AIR; NO AIR WAS FOUND IN THE TUBING AT THAT TIME. ABOUT 10 MINUTES LATER, AIR WAS FOUND TO BE IN THE IV TUBING WAITING TO BE ATTACHED TO THE UVC CATHETER WHERE STARTER TPN D10. THIS LINE WAS UNCLAMPED AND FLUSHED THROUGH A SECOND TIME UNTIL ALL AIR BUBBLES WERE OUT BEFORE BEING CONNECTED TO THE UVC. ONCE PROVIDER ORDERED THE D20 PIGGY-BACK, THE D20 WAS CONNECTED TO THE Y PORT ON THE TPN D10. THE PUMP AT THAT POINT KEPT ALARMING AS OCCLUDED. THE TPN D10 AND D20 LINE WERE DOUBLE CHECKED AND BOTH LINES WERE FOUND TO HAVE AIR IN THE TUBING. THE LINES WERE UNHOOKED FROM THE PATIENT AND FLUSHED AGAIN UNTIL ALL AIR WAS OUT OF THE TUBING. THE TUBING WAS THEN CHECKED AGAIN BY ANOTHER RN AND DR. TO VERIFY NO AIR WAS IN THE TUBING BEFORE BEING RECONNECTED TO THE UVC. A UAC WAS NOT ABLE TO BE ESTABLISHED SO THE SODIUM ACETATE WITH HEPARIN WAS TO BE CONNECTED TO ANOTHER PORT OF THE UVC LINE. THE SODIUM ACETATE WAS LOOKED OVER FOR ANY AIR IN THE TUBING, AS IT HAD BEEN SITTING CLAMPED AND CAPPED FOR ABOUT AN HOUR SINCE BEING PRIMED. AIR WAS FOUND IN THE TUBING OF THE SODIUM ACETATE. THE LINES WERE FLUSHED AGAIN UNTIL ALL AIR BUBBLES WERE OUT. AIR BUBBLES KEPT APPEARING IN THE TUBING NEAR THE PORTS AND RUNNING DOWN THE LINE WHILE FLUIDS WERE RUNNING. STAFF WERE UNABLE TO GET THE LINE TO STOP HAVING AIR BUBBLES. A NEW TUBING/SOLUTIONS SET WAS USED. THE NEW SET OF TUBING/IV BAG OF SODIUM ACETATE DID NOT SEEM TO HAVE ANY PROBLEMS WITH AIR BUBBLES, EVEN AFTER RUNNING FOR A BIT OF TIME. NO HARM OCCURRED TO THE PATIENT. THIS TUBING SET IS BEING RETURNED FOR EXAMINATION. IV PUMP: MODEL 8100 WITH ALARIS PC # (B)(4), IV INFUSION SET: REF# 2426-0007; LOT# 24109004 & 2409537. Event Description: NO ADDITIONAL INFORMATION. Manufacturer Narrative: NO PRODUCT OR PHOTO WAS RETURNED BY THE CUSTOMER. THE CUSTOMER COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED DUE TO THE PRODUCT NOT BEING RETURNED FOR FAILURE INVESTIGATION. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET. THE ROOT CAUSE OF THIS FAILURE COULD NOT BE IDENTIFIED WITHOUT A FAILURE INVESTIGATION. THIS INCIDENT HAS BEEN ADDED TO OUR DATABASE OF REPORTED INCIDENTS. OUR BUSINESS TEAM REGULARLY REVIEWS THE COLLECTED DATA FOR IDENTIFICATION OF EMERGING TRENDS. YOUR ASSISTANCE IN THIS MATTER HAS BEEN HELPFUL IN TREND IDENTIFICATION AND SUPPORTING OUR COMMITMENT TO CONTINUOUS QUALITY IMPROVEMENT. Manufacturer Narrative: THE CUSTOMER REPORTED AIR-IN-LINE, AND RETURNED ONE SAMPLE OF MATERIAL 2426-0007, LOT 24109078. UPON INITIAL VISUAL EXAMINATION, THE SET HAD NO DEFECTS OR ABNORMALITIES. THE SET WAS RUN AT 120ML/HR FOR 30 MINUTES IN ORDER TO REPLICATE AN AGGRESSIVE INFUSION. THERE WAS NO AIR IN LINE, LEAKS, OR OTHER ISSUES OBSERVED DURING THIS TIME. THE COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED. THE ROOT CAUSE FOR THE ISSUE COULD NOT BE FOUND WITHOUT A REPLICATED FAILURE. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED EVENT #3: (B)(6) 2025, UAC AND UVC PLACED ON NEW ADMIT TO NICU. SODIUM ACETATE WITH HEPARIN TO RUN THROUGH THE UAC. THE LINES WERE PRIMED WITH THE TRANSDUCER AND CHECKED FOR AIR; NO AIR WAS FOUND IN THE TUBING AT THAT TIME. ABOUT 10 MINUTES LATER, AIR WAS FOUND TO BE IN THE IV TUBING WAITING TO BE ATTACHED TO THE UVC CATHETER WHERE STARTER TPN D10. THIS LINE WAS UNCLAMPED AND FLUSHED THROUGH A SECOND TIME UNTIL ALL AIR BUBBLES WERE OUT BEFORE BEING CONNECTED TO THE UVC. ONCE PROVIDER ORDERED THE D20 PIGGY-BACK, THE D20 WAS CONNECTED TO THE Y PORT ON THE TPN D10. THE PUMP AT THAT POINT KEPT ALARMING AS OCCLUDED. THE TPN D10 AND D20 LINE WERE DOUBLE CHECKED AND BOTH LINES WERE FOUND TO HAVE AIR IN THE TUBING. THE LINES WERE UNHOOKED FROM THE PATIENT AND FLUSHED AGAIN UNTIL ALL AIR WAS OUT OF THE TUBING. THE TUBING WAS THEN CHECKED AGAIN BY ANOTHER RN AND DR. TO VERIFY NO AIR WAS IN THE TUBING BEFORE BEING RECONNECTED TO THE UVC. A UAC WAS NOT ABLE TO BE ESTABLISHED SO THE SODIUM ACETATE WITH HEPARIN WAS TO BE CONNECTED TO ANOTHER PORT OF THE UVC LINE. THE SODIUM ACETATE WAS LOOKED OVER FOR ANY AIR IN THE TUBING, AS IT HAD BEEN SITTING CLAMPED AND CAPPED FOR ABOUT AN HOUR SINCE BEING PRIMED. AIR WAS FOUND IN THE TUBING OF THE SODIUM ACETATE. THE LINES WERE FLUSHED AGAIN UNTIL ALL AIR BUBBLES WERE OUT. AIR BUBBLES KEPT APPEARING IN THE TUBING NEAR THE PORTS AND RUNNING DOWN THE LINE WHILE FLUIDS WERE RUNNING. STAFF WERE UNABLE TO GET THE LINE TO STOP HAVING AIR BUBBLES. A NEW TUBING/SOLUTIONS SET WAS USED. THE NEW SET OF TUBING/IV BAG OF SODIUM ACETATE DID NOT SEEM TO HAVE ANY PROBLEMS WITH AIR BUBBLES, EVEN AFTER RUNNING FOR A BIT OF TIME. NO HARM OCCURRED TO THE PATIENT. THIS TUBING SET IS BEING RETURNED FOR EXAMINATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289942&pc=FRN | 2016493-2025-60974 | 2024-12-15 05:00 | Death | CAREFUSION SD | 2025-02-03 05:00 | FRN | ALARIS SYSTEM | Break; Corroded; Crack; Failure to Sense; No Audible Prompt/Feedback; Adverse Event Without Identified Device or Use Problem; Misassembly During Maintenance/Repair | Embolism/Embolus | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, AN UMBILICAL ARTERIAL CATHETER (UAC) AND UMBILICAL VENOUS CATHETER (UVC) WERE PLACED ON A 27-WEEK-OLD INFANT SHORTLY AFTER ARRIVAL TO THE NEONATAL INTENSIVE CARE UNIT (NICU). AN INFUSION PUMP WAS USED WITH TUBING REF# (B)(4) FOR BOTH UAC AND UVC LINES. THE UAC LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE REPORTEDLY PRIMED AND FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. IT WAS REPORTED THAT THE INFANT'S CONDITION "ABRUPTLY DETERIORATED WITHIN MINUTES"; "HEART RATE DROPPED"; "CYANOSIS WAS NOTED". ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH REPORTEDLY SHOWED "AIR IN THE LEFT VENTRICLE OF THE HEART." THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2 TO 3ML. THE PATIENT REPORTEDLY DIED 8 DAYS LATER. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "NO AUDIBLE ALARMS SOUNDED, PER STAFF IN THE ROOM. THE LOCATION OF THE AIR WAS DISTAL TO THE PUMP. PUMP FAILURE WAS NOT A CONCERN. PUMPS WERE EVALUATED BY AGILITI BIOMED AND NO DEFICIENCIES WERE FOUND..." THE CUSTOMER (REGISTERED NURSE QUALITY IMPROVEMENT: RL ADMIN LITE & MEDSUN REPORTER) STATED THAT THE PRIMARY AND SECONDARY CAUSES OF DEATH WERE "COMPLICATIONS OF PREMATURE BIRTH ¿ 27 WEEKS AND 4 DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION."Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, AN UMBILICAL ARTERIAL CATHETER (UAC) AND UMBILICAL VENOUS CATHETER (UVC) WERE PLACED ON A 27-WEEK-OLD INFANT SHORTLY AFTER ARRIVAL TO THE NEONATAL INTENSIVE CARE UNIT (NICU). AN INFUSION PUMP WAS USED WITH TUBING REF# 2426-0007 FOR BOTH UAC AND UVC LINES. THE UAC LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE REPORTEDLY PRIMED AND FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. IT WAS REPORTED THAT THE INFANT'S CONDITION "ABRUPTLY DETERIORATED WITHIN MINUTES"; "HEART RATE DROPPED"; "CYANOSIS WAS NOTED". ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH REPORTEDLY SHOWED "AIR IN THE LEFT VENTRICLE OF THE HEART." THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2 TO 3ML. THE PATIENT REPORTEDLY DIED 8 DAYS LATER. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "NO AUDIBLE ALARMS SOUNDED, PER STAFF IN THE ROOM. THE LOCATION OF THE AIR WAS DISTAL TO THE PUMP. PUMP FAILURE WAS NOT A CONCERN. PUMPS WERE EVALUATED BY AGILITI BIOMED AND NO DEFICIENCIES WERE FOUND..." THE CUSTOMER (REGISTERED NURSE QUALITY IMPROVEMENT: RL ADMIN LITE & MEDSUN REPORTER) STATED THAT THE PRIMARY AND SECONDARY CAUSES OF DEATH WERE "COMPLICATIONS OF PREMATURE BIRTH ¿ 27 WEEKS AND 4 DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION." BD RECEIVED ADDITIONAL INFORMATION REGARDING TUBING SET UP: BD TUBING REF 2426-0007, LOT# 24109078 WAS USED. THE IV SOLUTION (BAG) WAS SPIKED USING THE IV INFUSION SETS LISTED. THE IV LINE WAS PRIMED BY GRAVITY THEN THE TUBING WAS PLACED IN THE ALARIS PUMP. UMBILICAL ARTERY CATHETER (UAC) LINES WERE TO BE USED FOR MONITORING BLOOD PRESSURE AND BLOOD SAMPLING. TO ACCOMPLISH THIS, THE DISTAL END OF THE IV TUBING WAS ATTACHED A TRANSDUCER. TRANSDUCER USED WAS ICU MEDICAL (MANUFACTURER), MODEL# 42801-24. AT THE DISTAL END OF THE TRANSDUCER TUBING WERE ONE 3-WAY STOPCOCK WITH A SYRINGE ATTACHED TO ONE PORT WAS ADDED. THIS IS PART OF THE ICU MEDICAL TRANSDUCER SET. THE SET COMES WITH TWO STOPCOCKS, BUT STAFF REMOVES ONE. THE LINE WAS THEN ATTACHED TO END OF THE UAC WHICH LED TO THE BABY.Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, AN UMBILICAL ARTERIAL CATHETER (UAC) AND UMBILICAL VENOUS CATHETER (UVC) WERE PLACED ON A 27-WEEK-OLD INFANT SHORTLY AFTER ARRIVAL TO THE NEONATAL INTENSIVE CARE UNIT (NICU). AN INFUSION PUMP WAS USED WITH TUBING REF# 2426-0007 FOR BOTH UAC AND UVC LINES. THE UAC LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE REPORTEDLY PRIMED AND FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. IT WAS REPORTED THAT THE INFANT'S CONDITION "ABRUPTLY DETERIORATED WITHIN MINUTES"; "HEART RATE DROPPED"; "CYANOSIS WAS NOTED". ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH REPORTEDLY SHOWED "AIR IN THE LEFT VENTRICLE OF THE HEART." THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2 TO 3ML. THE PATIENT REPORTEDLY DIED 8 DAYS LATER. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "NO AUDIBLE ALARMS SOUNDED, PER STAFF IN THE ROOM. THE LOCATION OF THE AIR WAS DISTAL TO THE PUMP. PUMP FAILURE WAS NOT A CONCERN. PUMPS WERE EVALUATED BY AGILITI BIOMED AND NO DEFICIENCIES WERE FOUND..." THE CUSTOMER (REGISTERED NURSE QUALITY IMPROVEMENT: RL ADMIN LITE & MEDSUN REPORTER) STATED THAT THE PRIMARY AND SECONDARY CAUSES OF DEATH WERE "COMPLICATIONS OF PREMATURE BIRTH ¿ 27 WEEKS AND 4 DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION." BD RECEIVED ADDITIONAL INFORMATION REGARDING TUBING SET UP: BD TUBING REF (B)(4), LOT# 24109078 WAS USED. THE IV SOLUTION (BAG) WAS SPIKED USING THE IV INFUSION SETS LISTED. THE IV LINE WAS PRIMED BY GRAVITY THEN THE TUBING WAS PLACED IN THE ALARIS PUMP. UMBILICAL ARTERY CATHETER (UAC) LINES WERE TO BE USED FOR MONITORING BLOOD PRESSURE AND BLOOD SAMPLING. TO ACCOMPLISH THIS, THE DISTAL END OF THE IV TUBING WAS ATTACHED A TRANSDUCER. TRANSDUCER USED WAS ICU MEDICAL (MANUFACTURER), MODEL# 42801-24. AT THE DISTAL END OF THE TRANSDUCER TUBING WERE ONE 3-WAY STOPCOCK WITH A SYRINGE ATTACHED TO ONE PORT WAS ADDED. THIS IS PART OF THE ICU MEDICAL TRANSDUCER SET. THE SET COMES WITH TWO STOPCOCKS, BUT STAFF REMOVES ONE. THE LINE WAS THEN ATTACHED TO END OF THE UAC WHICH LED TO THE BABY. Manufacturer Narrative: ADDITIONAL INFORMATION : UNIQUE IDENTIFIER (UDI) #. CORRECTION : MANUFACTURER NARRATIVE. OMIT : B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE. CORRECTION : DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION : DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, CONCOMITANT MED PROD DATA, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, G, B, C, D CODES, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY : THE REPORTED INCIDENT OF THE PUMP MODULE NOT ALARMING FOR AIR-IN-LINE WAS NOT CONFIRMED THROUGH REVIEW OF THE DEVICE LOGS OR REPLICATED DURING DEVICE TESTING. INSPECTION AND TESTING OF THE PUMP MODULE S/N (B)(6) DID NOT FIND ANY EXISTING MALFUNCTIONS. NO ANOMALIES WERE OBSERVED WITH THE AIR DETECTION SYSTEM. THE SUSPECT DEVICE WAS FOUND TO ALARM FOR AIR-IN-LINE AND ACCUMULATED AIR AS DESIGNED AND WITHIN SPECIFICATIONS. IT WAS OBSERVED THROUGH REVIEW OF THE LOGS THAT THE SUSPECT PUMP MODULE S/N (B)(6) WAS ATTACHED TO THE CONCOMITANT PCU S/N (B)(6) ON THE DATE OF THE REPORTED INCIDENT. THE EVENT LOGS FOR THE PCU WERE OVERWRITTEN, WITH THE LAST RECORDED EVENT DATE BEING (B)(6) 2025. WITHOUT THE PCU EVENT LOGS, AND DUE TO THE LIMITED INFORMATION, THAT THE PUMP MODULE LOGS CONTAIN, DRUG INFORMATION (SUCH AS THE EXACT DRUG PROGRAMMED AND ITS CONCENTRATION), TOTAL VOLUME INFUSED, AND UNIT POWER-OFF TIMES CANNOT BE DETERMINED. THE SUSPECT PUMP MODULE S/N (B)(6) EVENT LOG SHOWS THE UNIT WAS POWERED ON (B)(6) 2024 AT 7:11 PM. THE AIR IN LINE SETTING FOR SINGLE AIR BOLUS WAS SET TO 75MICROL WITH ACCUMULATED AIR ENABLED. BETWEEN 7:12 PM AND 7:25 PM, TWO (2) SAFETY CLAMP OPEN ALARMS WERE OBSERVED. AT 7:26 PM, THE CHANNEL WAS SELECTED, AND AN INFUSION OF AN UNKNOWN DRUG WAS PROGRAMMED AND STARTED AT A RATE OF 1 ML/H AND A VOLUME TO BE INFUSED (VTBI) OF 10 ML. BETWEEN 7:27 AM AND 11:08 PM, FOUR (4) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED. BETWEEN 8:36 PM AND 11:08 PM THE PUMP MODULE DOOR WAS OPENED AND CLOSED IN FIVE (5) INSTANCES. THE DEVICE WAS KEPT IN USE FOR APPROXIMATELY EIGHT (8) DAYS, UNTIL (B)(6) 2024, AT 10:55 AM, AN AIR-IN-LINE ALARM WAS OBSERVED. THREE (3) SECONDS AFTER THE ALARM, THE UNIT WAS CHANNELED OFF. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED PUMP MODULE NOT ALARMING FOR AIR-IN-LINE WAS NOT DETERMINED. NO ANOMALIES WERE OBSERVED WITH THE PUMP MODULE THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULE WAS FOUND TO BE DETECTING AIR-IN-LINE WITHIN SPECIFICATION. IT WAS NOTED DURING THE LOG REVIEW THAT AN AIR-IN-LINE ALARM WAS OBSERVED ON (B)(6) 2024 AT 10:55 AM WHICH INDICATES THAT THE AIR-IN-LINE CIRCUITRY WAS DETECTING AIR AT THAT TIME. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A0401, A040502, A0404, A230502, A070908, C07, C0601, C23, D01, D11, G02017, G0405206, G03005, G0301201 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE. Manufacturer Narrative: OMIT: B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE. CORRECTION: DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, CONCOMITANT MED PROD DATA, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, G, B, C, D CODES, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORTED INCIDENT OF THE PUMP MODULE NOT ALARMING FOR AIR-IN-LINE WAS NOT CONFIRMED THROUGH REVIEW OF THE DEVICE LOGS OR REPLICATED DURING DEVICE TESTING. INSPECTION AND TESTING OF THE PUMP MODULE S/N (B)(6) DID NOT FIND ANY EXISTING MALFUNCTIONS. NO ANOMALIES WERE OBSERVED WITH THE AIR DETECTION SYSTEM. THE SUSPECT DEVICE WAS FOUND TO ALARM FOR AIR-IN-LINE AND ACCUMULATED AIR AS DESIGNED AND WITHIN SPECIFICATIONS. IT WAS OBSERVED THROUGH REVIEW OF THE LOGS THAT THE SUSPECT PUMP MODULE S/N (B)(6) WAS ATTACHED TO THE CONCOMITANT PCU S/N (B)(6) ON THE DATE OF THE REPORTED INCIDENT. THE EVENT LOGS FOR THE PCU WERE OVERWRITTEN, WITH THE LAST RECORDED EVENT DATE BEING (B)(6) 2025. WITHOUT THE PCU EVENT LOGS, AND DUE TO THE LIMITED INFORMATION, THAT THE PUMP MODULE LOGS CONTAIN, DRUG INFORMATION (SUCH AS THE EXACT DRUG PROGRAMMED AND ITS CONCENTRATION), TOTAL VOLUME INFUSED, AND UNIT POWER-OFF TIMES CANNOT BE DETERMINED. THE SUSPECT PUMP MODULE S/N (B)(6) EVENT LOG SHOWS THE UNIT WAS POWERED ON (B)(6) 2024 AT 7:11 PM. THE AIR IN LINE SETTING FOR SINGLE AIR BOLUS WAS SET TO 75MICROL WITH ACCUMULATED AIR ENABLED. BETWEEN 7:12 PM AND 7:25 PM, TWO (2) SAFETY CLAMP OPEN ALARMS WERE OBSERVED. AT 7:26 PM, THE CHANNEL WAS SELECTED, AND AN INFUSION OF AN UNKNOWN DRUG WAS PROGRAMMED AND STARTED AT A RATE OF 1 ML/H AND A VOLUME TO BE INFUSED (VTBI) OF 10 ML. BETWEEN 7:27 AM AND 11:08 PM, FOUR (4) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED. BETWEEN 8:36 PM AND 11:08 PM THE PUMP MODULE DOOR WAS OPENED AND CLOSED IN FIVE (5) INSTANCES. THE DEVICE WAS KEPT IN USE FOR APPROXIMATELY EIGHT (8) DAYS, UNTIL (B)(6) 2024, AT 10:55 AM, AN AIR-IN-LINE ALARM WAS OBSERVED. THREE (3) SECONDS AFTER THE ALARM, THE UNIT WAS CHANNELED OFF. PER LABEL REVIEW, THE ALARIS PCU AND PUMP MODULE TECHNICAL SERVICE MANUAL (V9.33) STATES THE FOLLOWING: THE AIR-IN-LINE DETECTION THRESHOLD SPECIFIES THE AMOUNT OF AIR THAT CAN PASS THROUGH THE DETECTOR IN A SINGLE BOLUS BEFORE AN ALARM SOUNDS. IN NORMAL USE, THE THRESHOLD CAN BE SET AT 50, 75, OR 250MICROL. THE DEFAULT SETTING IS 75MICROL. (IN ANESTHESIA MODE ONLY, THE THRESHOLD VALUE CAN BE SET TO 500MICROL.) WHEN THE ACCUMULATED AIR-IN-LINE OPTION IS SET TO DISABLE, THE SYSTEM SOUNDS AN ALARM ONLY WHEN IT DETECTS AN AIR BOLUS LARGER THAN THE BOLUS SIZE SET AS THE AIR-IN-LINE DETECTION THRESHOLD EITHER 50, 75, OR 250MICROL. (IN ANESTHESIA MODE ONLY, THE THRESHOLD VALUE CAN BE SET TO 500MICROL.) SEE SECTION 2.5.3 AIR-IN-LINE SETTINGS. WHEN ENABLE IS SELECTED, NOT ONLY DOES THE AIR-IN-LINE SYSTEM DETECT AND RESPOND TO A BOLUS SIZE GREATER THAN THE SELECTED AIR-IN-LINE THRESHOLD OF 50, 75, 250, OR 500MICROL, IT ALSO DETECTS AND RESPONDS TO MULTIPLE SMALL BOLUSES OF A PREDETERMINED NUMBER, DEPENDING ON THE BOLUS SIZE SELECTED AS THE AIR-IN-LINE DETECTION THRESHOLD. DESPITE THE IDENTIFICATION OF A THIRD-PARTY PART, IT WAS NOT FOUND TO BE A CONTRIBUTING FACTOR TO THE REPORTED EVENT SINCE THE DEVICE WAS OPERATING AS INTENDED DURING TESTING. A MEDICAL DEVICE SAFETY ALERTMMS-21-4263-US) WAS SENT OUT ON 7 FEBRUARY 2022, WARNING FACILITIES TO ONLY USE BD-APPROVED PARTS WHEN PERFORMING CORRECTIVE MAINTENANCE OR REPAIRS. THE USE OF THIRD-PARTY PARTS CAN AFFECT THE SAFETY AND EFFICACY OF THE BD ALARIS AND ALARIS DEVICES, LEADING TO DEVICE FAILURE, PATIENT INJURY, OR EVEN DEATH. THE DEVICE WAS REPORTEDLY IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). NOTE TO REPAIR CENTER: 8100 PUMP MODULE S/N (B)(6) (SUSPECT). PLEASE REPLACE THE MECHANISM ASSEMBLY AND BEZEL ASSEMBLY AS IT WAS DISASSEMBLED DURING THE INVESTIGATION. MECHANISM ASSEMBLIES CANNOT BE REASSEMBLED SINCE MANUFACTURING PERFORMS SEVERAL TESTS NOT AVAILABLE TO DESIGNATED COMPLAINT HANDLING UNIT (DCHU), REPAIR CENTER OR THE CUSTOMER. THIS MAY BE CHARGED TO DCHU. REPAIR CENTER SHOULD FOLLOW THEIR NORMAL PROCESSING OF THE DEVICE, LOOKING FOR ANY INCIDENTAL ISSUES PRIOR TO RETURNING IT TO THE CUSTOMER. DCHU WILL NOT BE PICKING UP THE COST OF THE INCIDENTAL REPAIRS. IT WAS NOTED THAT THE REAR CASE WAS A THIRD PARTY. 8100 PUMP MODULE S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. 8100 PUMP MODULE S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. 8015 PCU S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. 8015 PCU S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. 8015 PCU S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. 8110 SYRINGE MODULE S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. 8110 SYRINGE MODULE S/N (B)(6) (CONCOMITANT). DEVICE WAS NOT OPENED FOR THIS INVESTIGATION, PLEASE ENSURE PROPER TESTING IS PERFORMED. DCHU IS NOT RESPONSIBLE FOR ANY COST ASSOCIATED WITH INCIDENTAL FINDINGS. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED PUMP MODULE NOT ALARMING FOR AIR-IN-LINE WAS NOT DETERMINED. NO ANOMALIES WERE OBSERVED WITH THE PUMP MODULE THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULE WAS FOUND TO BE DETECTING AIR-IN-LINE WITHIN SPECIFICATION. IT WAS NOTED DURING THE LOG REVIEW THAT AN AIR-IN-LINE ALARM WAS OBSERVED ON (B)(6) 2024 AT 10:55 AM WHICH INDICATES THAT THE AIR-IN-LINE CIRCUITRY WAS DETECTING AIR AT THAT TIME. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A0401, A040502, A0404, A230502, A070908, C07, C0601, C23, D01, D11, G02017, G0405206, G03005, G0301201 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21290976&pc=DTQ | 8010762-2025-0000051 | 2025-01-30 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-02-03 05:00 | DTQ | CARDIOHELP | Infusion or Flow Problem | Tachycardia | K133598 |  | 1 | Event Description: THE EVENT OCCURRED IN AUSTRALIA DURING TREATMENT. IT WAS REPORTED THAT A BACKFLOW PREVENTION MESSAGE OCCURRED ON THE SAME TIME AS THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA. FURTHER A ZERO FLOW MESSAGE APPEARED. DUE TO THE LOW FLOW THE PATIENT REQUIRED CPR. THE LOW FLOW STATE COULD NOT BE RESOLVED. THE CARDIOHELP WAS REPLACED. USING THE EMERGENCY DRIVE RESTORED THE FLOW AND THE FLOW REMAINED ACCEPTABLE ON THE REPLACEMENT CARDIOHELP. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED.Event Description: COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: A GETINGE SERVICE TECHNICIAN WILL INVESTIGATE THE AFFECTED CARDIOHELP. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US. Manufacturer Narrative: IT WAS REPORTED THAT A BACKFLOW PREVENTION MESSAGE OCCURRED ON THE SAME TIME AS PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT), WHICH IS A FAST, ABNORMAL HEART RHYTHM (ARRHYTHMIA). FURTHER A ZERO FLOW MESSAGE APPEARED. DUE TO THE LOW FLOW THE PATIENT REQUIRED CPR (REANIMATION). THE LOW FLOW STATE COULD NOT BE RESOLVED. THE CARDIOHELP WAS EXCHANGED. USING THE EMERGENCY DRIVE RESTORED THE FLOW AND THE FLOW REMAINED ACCEPTABLE ON THE REPLACEMENT CARDIOHELP. ONE OF THE NURSING STAFF HAS MENTIONED THAT THE CONSOLE HAD ¿FROZEN¿. THE PATIENT PASSED AWAY. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 FROM CUSTOMER THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION, WHICH WAS FOUND OUT AFTER THE PATIENT WAS PUT ONTO ECMO. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON (B)(6) 2025 AND (B)(6) 2025. NO FAULT WAS FOUND AND NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOGFILES WERE ANALYZED BY GETINGE LIFE-CYCLE-ENGINEERING ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: THE MESSAGE 'BACKFLOW PREVENTION' WAS TRIGGERED DURING THERAPY. THE USER WAS UNABLE TO START THE PUMP. THE REASON FOR THIS WAS THE ERROR ¿33006 PUMP DISPOSABLE STOP¿ REPORTED SEVEN SECONDS EARLIER. THIS CAUSED THE PUMP TO SWITCH OFF DUE TO A MAGNETIC GAP BETWEEN THE MOTOR AND THE PUMP THAT WAS DETERMINED TO BE TOO LARGE. THE ARTERIAL PRESSURE IN THE BODY CAUSED A NEGATIVE FLOW WHEN THE PUMP STOPPED. THIS TRIGGERED THE 'BACKFLOW PREVENTION' WITH THE 'ZERO FLOW' MODE AFTER 6 SECONDS. BECAUSE THE 33006 ERROR WAS NOT RESET, THE 'ZERO FLOW' MODE WAS UNABLE TO COUNTERACT THE NEGATIVE FLOW BY BLOCKING THE PUMP. THE USER SHOULD HAVE MADE SURE THAT THE HLS IS SECURELY LOCKED IN THE DRIVE. ADDITIONALLY IT WAS REQUIRED TO ADJUST THE 'USER TARGET SPEED' DOWN TO 0 RPM TO RESET THE 33006 ERROR MODE (THE MESSAGE 'PUMP DISPOSABLE ERROR STOP¿ DISAPPEARS FROM THE DISPLAY). THEN EITHER THE 'BACKFLOW PREVENTION' WOULD HAVE BEEN ACTIVATED OR THE PUMP COULD HAVE BEEN CONTROLLED BY THE USER. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: "THE LCE INVESTIGATION REPORT EXPLAINS BACKFLOW PREVENTION APPEARED SECONDARY TO DISPLACEMENT/DISLODGEMENT OF THE HLS MODULE FROM THE PUMP WELL, WHICH ELICITED THE ERROR ¿PUMP DISPOSABLE STOP¿ 6 SECONDS BEFORE THE REPORTED EVENT. FURTHER, THE DESCRIPTION OF A ¿FROZEN¿ SCREEN/SYSTEM MAY BE EXPLAINED BY AN UNRESOLVED THE ALARM/ERROR ¿PUMP DISPOSABLE STOP¿. THE IFU OF THE CARDIOHELP DESCRIBES THE FOLLOWING REGARDING THIS PUMP DISPOSABLE STOP¿ MESSAGE. TO RESOLVE THE ERROR, PROPER SEATING OF THE HLS MODULE IN THE CARDIOHELP PUMP WELL IS REQUIRED WHICH IS ENSURED BY SECURELY LOCKING THE MODULE INTO PUMP DRIVE. ONCE THE HLS MODULE IS PROPERLY SEATED, THE RPMS SHOULD BE RESET TO ZERO (0) TO CLEAR THE ERROR MODE. PRESSING THE ZERO FLOW MODE BUTTON IS REQUIRED TO RESET BACKFLOW PREVENTION AND REESTABLISH FLOW. RPM SHOULD BE INCREASED (AFTER THE ZERO SETTING) TO CREATE REQUISITE TARGET BLOOD FLOW. THE IFU ADDRESSES THIS PROCEDURE IN FOLLOWING MANNER: ¿THE BACKFLOW PREVENTION CAN DETECT AND RESPOND TO A BACKFLOW OF BLOOD. FOR THIS, THE CARDIOHELP-I MONITORS THE BLOOD FLOW, DISPLAYS ANY ALARMS AND ACTIVATES THE ZERO FLOW MODE AUTOMATICALLY TO PREVENT ANY BACKFLOW¿. IT REMAINS UNCLEAR WHAT INITIALLY CAUSED THE ERROR MESSAGE ¿PUMP DISPOSABLE STOP¿. THE INVESTIGATION REPORT EXPLAINS THAT THIS ALARM IS TRIGGERED IF A MAGNETIC GAP BETWEEN THE MOTOR AND THE PUMP IS DETERMINED TO BE TOO LARGE. A POSSIBLE ROOT CAUSE FOR THE REPORTED EVENT THAT MAY BE PROPOSED IS DISLODGMENT OF THE HLS MODULE DURING TRANSPORT AND/OR REPOSITIONING OF THE PATIENT WHICH MAY HAVE EXERTED EXCESS STRESS ON THE TUBES/LINES, LEADING TO DISLODGEMENT OF THE DISPOSABLE. ALTERNATIVELY, THE DISPOSABLE MAY NOT HAVE BEEN FULLY ENGAGED IN THE PUMP RECEIVER WHICH EASILY DISLODGED UPON MINIMAL STRESS APPLIED TO THE TUBES/LINES AND/OR THE HLS MODULE. IN EITHER CASE, THE ERROR MESSAGE ¿PUMP DISPOSABLE STOP¿ WOULD HAVE BEEN ELICITED UPON HLS MODULE DISLODGEMENT. LAST, STRUCTURAL DAMAGE OF THE HLS MODULE IS ANOTHER POSSIBLE ROOT CAUSE (AS A PROPOSED ROOT CAUSE UNDERLYING THE EVENT) WHICH MAY HAVE MANIFESTED AS BENT OR DAMAGED ANCHORING FLANGE(S). DAMAGE TO THE ANCHORING FLANGES MAY HAVE NOT ALLOWED THE HLS SET TO PROPERLY ENGAGE WITH THE PUMP DRIVE OF THE CARDIOHELP HARDWARE THEREBY DISALLOWING THE HLS MODULE TO SECURELY ATTACH TO THE CARDIOHELP PUMP DRIVE. THAT SAID, THIS PROPOSAL COULD BE RULED OUT AS THE HLS SET DID FUNCTION IN A REPLACEMENT CARDIOHELP AND BLOOD FLOW WAS REESTABLISHED. THE CARDIOHELP INVOLVED IN THE COMPLAINT WAS INVESTIGATED BY GETINGE FIELD SERVICE TECHNICIAN (FST). THE INVESTIGATION WAS UNABLE TO IDENTIFY ANY FAULTS OR ISSUES RELATING TO PERFORMANCE WITH THE UNIT. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS PASSED AS DESIGNED. THE CUSTOMER STATED THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM THE PROCEDURE. IT IS ASSUMED THAT THE STATEMENT OF NONRECOVERY WAS PREDICATED ON THE DISPOSITION IO THE PATIENT BEFORE SUPPORT USING CARDIOHELP. UNFORTUNATELY, NO DETAILS OF THE PATIENT CONSTITUTION WERE DISCLOSED BY THE CUSTOMER OR WITHIN THE COMPLAINT NARRATIVE. THAT SAID, THE ABSENCE OR REDUCTION OF BLOOD DURING THE REPORTED EVENT MAY HAVE CONTRIBUTED TO THE REPORTED DETERIORATION OF THE PATIENT. HOWEVER, A DIMINUTION OF PRODUCT PERFORMANCE OR MALPERFORMANCE OF THE PRODUCT COULD NOT BE REPRODUCED DURING INSPECTION AND EVALUATION OF THE CARDIOHELP HARDWARE. FURTHER, EXAMINATION OF THE ASSOCIATED DISPOSABLE WAS UNABLE TO BE PERFORMED DUE TO COMPONENT DISPOSAL; THEREFORE, COMMENTARY PERTAINING TO EITHER PERFORMANCE OR MALPERFORMANCE CANNOT BE ADVANCED IN THIS REVIEW." ACCORDING TO THE INSTRUCTION FOR USE (IFU) CHAPTER "CONNECTING THE COMBINED FLOW/BUBBLE SENSOR" THE BUBBLE MONITORING FUNCTION TEST AND FLOW OFF-SET CALIBRATION HAS TO BE PERFORMED BEFORE EVERY USE. THUS A DEFECTIVE FLOW/BUBBLE SENSOR SHOULD BE DETECTED PRIOR TO USE, DURING PRIMING. IN ADDITION AS THE CARDIOHELP INCLUDES PRESSURE SENSORS AND A VENOUS PROBE IT IS ABLE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN THE IFU CHAPTER "USING THE EMERGENCY DRIVE WITH THE DISPOSABLE HLS RETAINER" IS STATED THAT THE EMERGENCY DRIVE CAN BE USED TO MANUALLY CONTROL THE BLOOD FLOW IN CASE OF A FAILED CARDIOHELP. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON(B)(6) 2025 FOR THE PERIOD OF (B)(6) 2019 TO (B)(6) 2025. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2019-10-09. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. BASED ON THE RESULTS THE REPORTED FAILURE "BACKFLOW PREVENTION" COULD BE CONFIRMED, BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. FURTHER IT WAS STATED BY THE CUSTOMER THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US. Manufacturer Narrative: NEW INFORMATION WAS RECEIVED ON 2025-02-04 THAT THE CUSTOMER STATED THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION, WHICH WAS FOUND OUT AFTER THE PATIENT WAS PUT ONTO ECMO. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-02-03 AND 2025-02-04. NO FAULT WAS FOUND AND NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED BY THE GETINGE SERVICE AND SALES UNIT AND THE BACKFLOW PREVENTION ALARM COULD BE CONFIRMED ON THE DATE OF EVENT. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291019&pc=OZD | 1220648-2025-25958 | 2024-12-31 05:00 | Death | ABIOMED, INC. | 2025-02-03 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Stroke/CVA; Renal Failure | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY, THE PATIENT WAS ON IMPELLA 5.5 SUPPORT. WHILE ON SUPPORT, THE PATIENT ENDURED ATRIAL ARRHYTHMIA AND WAS TREATED WITH INTRAVENOUS AMIODARONE. FURTHER THE PATIENT HAD A CEREBROVASCULAR ACCIDENT. A CAT SCAN ON THE HEAD SHOWED LACUNAR INFARCTS AND CEREBELLAR INFARCT. IT WAS ALSO REPORTED THAT THE PATIENT ENDURED RENAL FAILURE AND TREATED WITH CREATININE HOWEVER LATER THE PATIENT PASSED AWAY BEFORE IT WAS RESOLVED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT, STROKE/CVA, AND RENAL FAILURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291053&pc=MVK | 3008642652-2025-01216 | 2024-12-16 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED NINE APPROPRIATE TREATMENTS AND TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 08:04:12 ON (B)(6) 2024. AT 10:44:30, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 10:46:07, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 10:48:51, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 10:51:22, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS ASYSTOLE WITH NSVT AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 10:51:57, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 10:52:35, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 65 BPM WITH PVCS, BBB AND CPR/MOTION ARTIFACT. AT 10:53:37, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 200 BPM WITH CPR/MOTION ARTIFACT. AT 10:54:10, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VT @ 190 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 120 BPM. AT 10:54:41, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 140 BPM WITH CPR/MOTION ARTIFACT. AT 10:55:14, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 10:55:54, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 10:56:27, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 10:58:17, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 10:59:10, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 10:59:41, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 11:03:39, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 11:04:11, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 30 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 11:04:26 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291056&pc=MIH | 2017233-2025-05787 | 2025-01-04 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-03 05:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Vascular Dissection; Ruptured Aneurysm | P040043 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2017, THIS PATIENT WAS IMPLANTED WITH TWO NON-GORE STENT GRAFTS FOR A THORACIC AORTIC ANEURYSM. ON (B)(6) 2019, DUE TO MIGRATION OF THE PROXIMAL NON-GORE STENT GRAFT AND EXPANSION OF THE DISTAL ANEURYSM, TWO GORE® TAG® CONFORMABLE THORACIC STENT GRAFTS WITH ACTIVE CONTROL SYSTEM (TGMR404015J, TGM454520J) WERE ADDITIONALLY IMPLANTED. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2021, A SURGICAL GRAFT REPLACEMENT OF THE THORACOABDOMINAL AORTA WAS PERFORMED FOR THE REPAIR OF A JUXTA RENAL AND INFRARENAL ABDOMINAL AORTIC ANEURYSM. ON (B)(6) 2024, AN ENDOLEAK OF AN UNSPECIFIED TYPE (SUSPECTED PROXIMAL TYPE I OR TYPE II) WAS OBSERVED IN THE PROXIMAL CTAGAC (TGMR404015J). ADDITIONALLY, A NEW DISSECTION ENTRY (DSINE) ON THE DISTAL CTAGAC (TGM454520J) WAS OBSERVED. AN ANEURYSM EXPANSION WAS ALSO OBSERVED. ADDITIONAL TREATMENT INCLUDING RIBS WAS PLANNED FOR (B)(6) 2025. ON (B)(6) 2025, THE PHYSICIAN INFORMED THE FSA THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. THE PHYSICIAN¿S COMMENT: ¿THE PATIENT WAS FOUND DECEASED IN HIS SLEEP. ALTHOUGH THERE IS A LOW POSSIBILITY, I SUSPECT ANEURYSM RUPTURE MAY HAVE CAUSED IT.¿ Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: DEATH, DISSECTION, RUPTURE OF THE AORTIC VESSEL & SURROUNDING VASCULATURE. Manufacturer Narrative: H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICES NOT AVAILABLE FOR RETURN. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291068&pc=MIH | 2017233-2025-05786 | 2025-01-04 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-03 05:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | Ruptured Aneurysm | P040043 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2017, THIS PATIENT WAS IMPLANTED WITH TWO NON-GORE STENT GRAFTS FOR A THORACIC AORTIC ANEURYSM. ON (B)(6) 2019, DUE TO MIGRATION OF THE PROXIMAL NON-GORE STENT GRAFT AND EXPANSION OF THE DISTAL ANEURYSM, TWO GORE® TAG® CONFORMABLE THORACIC STENT GRAFTS WITH ACTIVE CONTROL SYSTEM (TGMR404015J, TGM454520J) WERE ADDITIONALLY IMPLANTED. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2021, A SURGICAL GRAFT REPLACEMENT OF THE THORACOABDOMINAL AORTA WAS PERFORMED FOR THE REPAIR OF A JUXTA RENAL AND INFRARENAL ABDOMINAL AORTIC ANEURYSM. ON (B)(6) 2024, AN ENDOLEAK OF AN UNSPECIFIED TYPE (SUSPECTED PROXIMAL TYPE I OR TYPE II) WAS OBSERVED IN THE PROXIMAL CTAGAC (TGMR404015J). ADDITIONALLY, A NEW DISSECTION ENTRY (DSINE) ON THE DISTAL CTAGAC (TGM454520J) WAS OBSERVED. AN ANEURYSM EXPANSION WAS ALSO OBSERVED. ADDITIONAL TREATMENT INCLUDING RIBS WAS PLANNED FOR (B)(6) 2025. ON (B)(6) 2025, THE PHYSICIAN INFORMED THE FSA THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. THE PHYSICIAN¿S COMMENT: ¿THE PATIENT WAS FOUND DECEASED IN HIS SLEEP. ALTHOUGH THERE IS A LOW POSSIBILITY, I SUSPECT ANEURYSM RUPTURE MAY HAVE CAUSED IT.¿ Manufacturer Narrative: H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICES NOT AVAILABLE FOR RETURN. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: AORTIC EXPANSION (E.G., ANEURYSM, FALSE LUMEN, LANDING ZONE, LESION), DEATH, RUPTURE OF THE AORTIC VESSEL & SURROUNDING VASCULATURE, ENDOLEAK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291103&pc=DSQ | 2916596-2025-00354 | 2025-01-10 05:00 | Death | THORATEC CORPORATION | 2025-02-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND WAS DUE TO COMORBIDITIES AND GENERAL DECLINE. AN AUTOPSY WAS NOT PERFORMED.Event Description: THE PATIENT PASSED AWAY DUE TO PROGRESSIVE RENAL FAILURE. DIALYSIS WAS NOT PERFORMED, GIVEN COMORBIDITIES. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED RENAL DYSFUNCTION AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PROGRESSIVE RENAL FAILURE. DIALYSIS WAS NOT PERFORMED, GIVEN COMORBIDITIES. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND WAS DUE TO COMORBIDITIES AND GENERAL DECLINE. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING RENAL DYSFUNCTION AND DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291142&pc=BZD | 2518422-2025-010203 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-02-03 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE AND DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291152&pc=BZD | 2518422-2025-010275 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-02-03 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291541&pc=MVK | 3008642652-2025-01214 | 2024-12-30 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing; Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS AND SIX INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 22:57:11 ON (B)(6) 2024. AT 01:55:07 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 01:55:44, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 01:56:31, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 01:58:00, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 02:03:24, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING RATES/AMPLITUDES AND CPR/MOTION ARTIFACT. AT 02:05:48, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH VARYING RATES/AMPLITUDES AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 02:06:17, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 02:06:44, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 02:07:15, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH VARYING RATES/AMPLITUDES AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH VARYING RATES/AMPLITUDES AND CPR/MOTION ARTIFACT. AT 02:08:01, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE ELECTRODE BELT WAS DISCONNECTED AT 02:16:45 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291566&pc=MVK | 3008642652-2025-01212 | 2024-12-30 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 07:34:22 ON (B)(6) 2024. AT 22:03:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 60 BPM DEGRADING TO VF. AT 22:03:50, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 22:07:00, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 22:07:31, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS PACED RHYTHM @ 60 BPM WITH LOW AMPLITUDE CARDIAC SIGNAL. THE ELECTRODE BELT WAS DISCONNECTED AT 22:14:09 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291738&pc=MVK | 3008642652-2025-01213 | 2024-12-27 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 19:14:04 ON (B)(6) 2024. AT 02:46:55 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 110 BPM TRANSITIONING TO VT @ 230 BPM. AT 02:47:33, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM. AT 02:51:51, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 230 BPM DEGRADING TO VF WITH CPR/MOTION ARTIFACT. AT 02:57:12, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 03:59:42 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21291947&pc=LOX | 3002808904-2025-00002 | 2024-12-27 05:00 | Death | KANEKA CORPORATION | 2025-02-03 05:00 | LOX | KANEKA PTCA CATHETER ZINRAI | Material Rupture | Cardiogenic Shock |  |  | 1 | Event Description: THE PATIENT HAD A HISTORY OF AORTIC VALVE STENOSIS. AT A PCI OPERATION FOR 99% OCCLUDED LEFT ANTERIOR DESCENDING AT #6, THE CONCERNED ZINRAI PTCA BALLOON WAS USED FOR PRE-DILATATION, AND BALLOON RUPTURE OCCURRED. AND AFTER THE BALLOON RUPTURE, BLOOD PRESSURE, HEART RATE, AND PULSE PRESSURE DECREASED, RESULTING IN SHOCK VITALITY. SLOW FLOW WAS CONFIRMED, AND IT WAS DETERMINED THAT THE PATIENT COULD NOT MAINTAIN HER CIRCULATORY STATUS, SO AN IABP/ECMO WAS INSERTED AND ASSISTED VENTILATION WAS INITIATED. BIOFREEDOM STENT WAS PLACED AND POST DILATATION WITH SAPPHAIRE NC WAS CONDUCTED FOR SECURING HER BLOOD FLOW. BUT THE PATIENT DIED ON THE NEXT DAY, (B)(6) 2024 BY UNKNOWN CAUSE OF DEATH. A COMMENT FROM A HEALTH PROFESSIONAL: THE PROCEDURE WAS FOR AN ELDERLY, HIGH-RISK PATIENT WITH SEVERE AORTIC STENOSIS AND SEVERE HEART FAILURE. THE BALLOON RUPTURE MIGHT HAVE CAUSED AIR EMBOLIZATION, LEADING TO WORSENING OF BLOOD FLOW. Manufacturer Narrative: THE CONCERNED DEVICE SUBJECT TO THIS REPORTED EVENT, "ZINRAI", AN RX-TYPE PTCA BALLOON CATHETER COMPATIBLE WITH 0.014" GUIDEWIRE (GW), IS NOT DISTRIBUTED IN THE US, HOWEVER, WE INTEND TO REPORT THIS CASE AS THE EVENT OCCURRED ON ONE OF THE SIMILAR DEVICES FOR "RX TAKERU PTCA BALLOON DILATATION CATHETER" DISTRIBUTED IN THE US UNDER 510(K) # K163372. THE DEVICE HISTORY RECORDS (DHR) OF THE DEVICE CONCERNED WAS REVIEWED: THE PRODUCTION LOT, TO WHICH THE DEVICE CONCERNED BELONGS, PASSED ALL IN-PROCESS INSPECTIONS INCLUDING THE SHAFT-PRESSURIZED TEST AND THE BALLOON-WRAPPING TEST FOR EVERY PRODUCT, AND THE FINISHED PRODUCT INSPECTIONS INCLUDING THE SHAFT TENSILE STRENGTH TEST AND THE REPETITIVE BALLOON INFLATION/DEFLATION TEST ON REPRESENTATIVE SAMPLES BASED ON SAMPLING PLAN. NO NONCONFORMITY OR ABNORMALITY IN THE MANUFACTURING PROCESSES OF THE DEVICE CONCERNED WAS FOUND. THE ACTUAL DEVICE WAS NOT RETURNED AND COULD NOT BE INVESTIGATED. WE ASSUME THE CAUSE OF THIS AS FOLLOWS: FACTORS THAT MAY CONTRIBUTE TO RUPTURE OF THE BALLOON INCLUDE, BUT ARE NOT LIMITED TO, THE SITUATION OF THE PATIENT'S VASCULATURE LESIONS SUCH AS CALCIFICATION AND HIGHLY STENOTIC. AFTER THE BALLOON RUPTURE HAD OCCURRED, A PROCEDURE TO SECURE BLOOD FLOW WAS TAKEN BY ASSISTED VENTILATION AND STENT PLACEMENT. THE DIRECT CAUSE OF DEATH IS UNKNOWN, BUT SINCE THERE ARE NO ABNORMALITIES IN THE MANUFACTURING RECORDS, WE ASSUME THAT THE DEATH WAS CAUSED BY THE PATIENT CONDITION. IN THE INSTRUCTIONS FOR USE OF ZINRAI (3283-0) , WE STATE THE POTENTIAL OF KNOWN RISK AS BELOW. [PRECAUTIONS] <IMPORTANT, BASIC PRECAUTIONS> 1. SINCE PTCA PROCEDURE MAY INDUCE DANGEROUS COMPLICATIONS, PERFORM THE PROCEDURE ONLY AFTER HAVING PREPARED FOR EMERGENCY CORONARY ARTERY BYPASS GRAFTING (CABG). 8. IF ANY ABNORMALITY IS OBSERVED DURING BALLOON INFLATION, IMMEDIATELY APPLY NEGATIVE PRESSURE AND DISCONTINUE THE PROCEDURE. 12. THE PHYSICIAN IN CHARGE OF THE PROCEDURE SHOULD DETERMINE THE DURATION AND NUMBER OF BALLOON INFLATIONS BASED ON HIS/HER PAST EXPERIENCES. <PRECAUTIONS DURING USAGE > 2. BEFORE USE, EXPEL THE AIR COMPLETELY FROM THE BALLOON AND BALLOON INFLATION LUMEN BY REPLACING IT WITH THE INFLATION MEDIUM. [IF THE AIR IS NOT COMPLETELY EXPELLED, THE INFLATION OF THE BALLOON MAY NOT BE OBSERVABLE ON FLUOROSCOPY.] 3. USE THE INFLATION MEDIUM TO INFLATE THE BALLOON. AIR OR OTHER GASES SHOULD NEVER BE USED TO INFLATE THE BALLOON. [DEVICE FAILURES AND ADVERSE EVENTS] THE USE OF THE PRODUCT MAY CAUSE THE FOLLOWING DEVICE FAILURE: 1. SEVERE DEVICE FAILURES (1) BALLOON RUPTURE 2. OTHER DEVICE FAILURES (1) LEAKAGE OF THE INFLATION MEDIUM 1. SEVERE ADVERSE EVENTS (1) DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21292086&pc=MVK | 3008642652-2025-01207 | 2024-12-21 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 11:09:10 ON (B)(6) 2024. AT 09:13:22 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 90 BPM DEGRADING TO VF. AT 09:13:57, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 270 BPM DEGRADING TO VF. AT 09:14:31, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY FOR 12 SECONDS TRANSITIONING TO AF @ 50 BPM WITH PVCS. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 09:24:14, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT. AT 09:24:46, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BEATS DEGRADING TO VF. AT 09:25:15, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BEATS DEGRADING TO VF WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE ELECTRODE BELT WAS DISCONNECTED AT 10:22:08 ON (B)(6) 2024. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21292323&pc=MVK | 3008642652-2025-01208 | 2024-11-19 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 17:20:35 ON (B)(6) 2024. AT 23:25:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 80 BPM DEGRADING TO VF. AT 23:26:38, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 00:13:09 ON (B)(6) 2024. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21292544&pc=MVK | 3008642652-2025-01206 | 2025-01-10 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:58:44 ON (B)(6 )2025. AT 00:31:47 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 30 BPM WITH PVCS, BBB, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. AT 00:32:55, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 30 BPM WITH PVCS, BBB, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PVCS, BBB, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION/TACTILE ARTIFACT. AT 08:21:21, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 30 BPM WITH PVCS, BBB, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. AT 08:22:00, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 40 BPM WITH PVCS, BBB, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PVCS, BBB, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 11:32:48 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21293142&pc=DTQ | 3013876692-2025-00007 | 2025-01-30 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-02-03 05:00 | DTQ | CARDIOHELP | Infusion or Flow Problem | Tachycardia | K133598 |  | 1 | Event Description: THE EVENT OCCURRED IN AUSTRALIA DURING TREATMENT. IT WAS REPORTED THAT A BACKFLOW PREVENTION MESSAGE OCCURRED ON THE SAME TIME AS THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA. FURTHER A ZERO FLOW MESSAGE APPEARED. DUE TO THE LOW FLOW THE PATIENT REQUIRED CPR. THE LOW FLOW STATE COULD NOT BE RESOLVED. THE CARDIOHELP WAS REPLACED. USING THE EMERGENCY DRIVE RESTORED THE FLOW AND THE FLOW REMAINED ACCEPTABLE ON THE REPLACEMENT CARDIOHELP. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED.Event Description: COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: A GETINGE SERVICE TECHNICIAN WILL INVESTIGATE THE AFFECTED CARDIOHELP. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US. Manufacturer Narrative: NEW INFORMATION WAS RECEIVED ON 2025-02-04 THAT THE CUSTOMER STATED THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION, WHICH WAS FOUND OUT AFTER THE PATIENT WAS PUT ONTO ECMO. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-02-03 AND 2025-02-04. NO FAULT WAS FOUND AND NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED BY THE GETINGE SERVICE AND SALES UNIT AND THE BACKFLOW PREVENTION ALARM COULD BE CONFIRMED ON THE DATE OF EVENT. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US. Manufacturer Narrative: IT WAS REPORTED THAT A BACKFLOW PREVENTION MESSAGE OCCURRED ON THE SAME TIME AS PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT), WHICH IS A FAST, ABNORMAL HEART RHYTHM (ARRHYTHMIA). FURTHER A ZERO FLOW MESSAGE APPEARED. DUE TO THE LOW FLOW THE PATIENT REQUIRED CPR (REANIMATION). THE LOW FLOW STATE COULD NOT BE RESOLVED. THE CARDIOHELP WAS EXCHANGED. USING THE EMERGENCY DRIVE RESTORED THE FLOW AND THE FLOW REMAINED ACCEPTABLE ON THE REPLACEMENT CARDIOHELP. ONE OF THE NURSING STAFF HAS MENTIONED THAT THE CONSOLE HAD ¿FROZEN¿. THE PATIENT PASSED AWAY. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 FROM CUSTOMER THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION, WHICH WAS FOUND OUT AFTER THE PATIENT WAS PUT ONTO ECMO. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON (B)(6) 2025. NO FAULT WAS FOUND AND NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOGFILES WERE ANALYZED BY GETINGE LIFE-CYCLE-ENGINEERING ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: THE MESSAGE 'BACKFLOW PREVENTION' WAS TRIGGERED DURING THERAPY. THE USER WAS UNABLE TO START THE PUMP. THE REASON FOR THIS WAS THE ERROR ¿33006 PUMP DISPOSABLE STOP¿ REPORTED SEVEN SECONDS EARLIER. THIS CAUSED THE PUMP TO SWITCH OFF DUE TO A MAGNETIC GAP BETWEEN THE MOTOR AND THE PUMP THAT WAS DETERMINED TO BE TOO LARGE. THE ARTERIAL PRESSURE IN THE BODY CAUSED A NEGATIVE FLOW WHEN THE PUMP STOPPED. THIS TRIGGERED THE 'BACKFLOW PREVENTION' WITH THE 'ZERO FLOW' MODE AFTER 6 SECONDS. BECAUSE THE 33006 ERROR WAS NOT RESET, THE 'ZERO FLOW' MODE WAS UNABLE TO COUNTERACT THE NEGATIVE FLOW BY BLOCKING THE PUMP. THE USER SHOULD HAVE MADE SURE THAT THE HLS IS SECURELY LOCKED IN THE DRIVE. ADDITIONALLY, IT WAS REQUIRED TO ADJUST THE 'USER TARGET SPEED' DOWN TO 0 RPM TO RESET THE 33006 ERROR MODE (THE MESSAGE 'PUMP DISPOSABLE ERROR STOP ¿DISAPPEARS FROM THE DISPLAY). THEN EITHER THE 'BACKFLOW PREVENTION' WOULD HAVE BEEN ACTIVATED OR THE PUMP COULD HAVE BEEN CONTROLLED BY THE USER. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: "THE LCE INVESTIGATION REPORT EXPLAINS BACKFLOW PREVENTION APPEARED SECONDARY TO DISPLACEMENT/DISLODGEMENT OF THE HLS MODULE FROM THE PUMP WELL, WHICH ELICITED THE ERROR ¿PUMP DISPOSABLE STOP¿ 6 SECONDS BEFORE THE REPORTED EVENT. FURTHER, THE DESCRIPTION OF A ¿FROZEN¿ SCREEN/SYSTEM MAY BE EXPLAINED BY AN UNRESOLVED THE ALARM/ERROR ¿PUMP DISPOSABLE STOP¿. THE IFU OF THE CARDIOHELP DESCRIBES THE FOLLOWING REGARDING THIS PUMP DISPOSABLE STOP¿ MESSAGE. TO RESOLVE THE ERROR, PROPER SEATING OF THE HLS MODULE IN THE CARDIOHELP PUMP WELL IS REQUIRED WHICH IS ENSURED BY SECURELY LOCKING THE MODULE INTO PUMP DRIVE. ONCE THE HLS MODULE IS PROPERLY SEATED, THE RPMS SHOULD BE RESET TO ZERO (0) TO CLEAR THE ERROR MODE. PRESSING THE ZERO FLOW MODE BUTTON IS REQUIRED TO RESET BACKFLOW PREVENTION AND REESTABLISH FLOW. RPM SHOULD BE INCREASED (AFTER THE ZERO SETTING) TO CREATE REQUISITE TARGET BLOOD FLOW. THE IFU ADDRESSES THIS PROCEDURE IN FOLLOWING MANNER: ¿THE BACKFLOW PREVENTION CAN DETECT AND RESPOND TO A BACKFLOW OF BLOOD. FOR THIS, THE CARDIOHELP-I MONITORS THE BLOOD FLOW, DISPLAYS ANY ALARMS AND ACTIVATES THE ZERO FLOW MODE AUTOMATICALLY TO PREVENT ANY BACKFLOW¿. IT REMAINS UNCLEAR WHAT INITIALLY CAUSED THE ERROR MESSAGE ¿PUMP DISPOSABLE STOP¿. THE INVESTIGATION REPORT EXPLAINS THAT THIS ALARM IS TRIGGERED IF A MAGNETIC GAP BETWEEN THE MOTOR AND THE PUMP IS DETERMINED TO BE TOO LARGE. A POSSIBLE ROOT CAUSE FOR THE REPORTED EVENT THAT MAY BE PROPOSED IS DISLODGMENT OF THE HLS MODULE DURING TRANSPORT AND/OR REPOSITIONING OF THE PATIENT WHICH MAY HAVE EXERTED EXCESS STRESS ON THE TUBES/LINES, LEADING TO DISLODGEMENT OF THE DISPOSABLE. ALTERNATIVELY, THE DISPOSABLE MAY NOT HAVE BEEN FULLY ENGAGED IN THE PUMP RECEIVER WHICH EASILY DISLODGED UPON MINIMAL STRESS APPLIED TO THE TUBES/LINES AND/OR THE HLS MODULE. IN EITHER CASE, THE ERROR MESSAGE ¿PUMP DISPOSABLE STOP¿ WOULD HAVE BEEN ELICITED UPON HLS MODULE DISLODGEMENT. LAST, STRUCTURAL DAMAGE OF THE HLS MODULE IS ANOTHER POSSIBLE ROOT CAUSE (AS A PROPOSED ROOT CAUSE UNDERLYING THE EVENT) WHICH MAY HAVE MANIFESTED AS BENT OR DAMAGED ANCHORING FLANGE(S). DAMAGE TO THE ANCHORING FLANGES MAY HAVE NOT ALLOWED THE HLS SET TO PROPERLY ENGAGE WITH THE PUMP DRIVE OF THE CARDIOHELP HARDWARE THEREBY DISALLOWING THE HLS MODULE TO SECURELY ATTACH TO THE CARDIOHELP PUMP DRIVE. THAT SAID, THIS PROPOSAL COULD BE RULED OUT AS THE HLS SET DID FUNCTION IN A REPLACEMENT CARDIOHELP AND BLOOD FLOW WAS REESTABLISHED. THE CARDIOHELP INVOLVED IN THE COMPLAINT WAS INVESTIGATED BY GETINGE FIELD SERVICE TECHNICIAN (FST). THE INVESTIGATION WAS UNABLE TO IDENTIFY ANY FAULTS OR ISSUES RELATING TO PERFORMANCE WITH THE UNIT. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS PASSED AS DESIGNED. THE CUSTOMER STATED THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM THE PROCEDURE. IT IS ASSUMED THAT THE STATEMENT OF NONRECOVERY WAS PREDICATED ON THE DISPOSITION IO THE PATIENT BEFORE SUPPORT USING CARDIOHELP. UNFORTUNATELY, NO DETAILS OF THE PATIENT CONSTITUTION WERE DISCLOSED BY THE CUSTOMER OR WITHIN THE COMPLAINT NARRATIVE. THAT SAID, THE ABSENCE OR REDUCTION OF BLOOD DURING THE REPORTED EVENT MAY HAVE CONTRIBUTED TO THE REPORTED DETERIORATION OF THE PATIENT. HOWEVER, A DIMINUTION OF PRODUCT PERFORMANCE OR MALPERFORMANCE OF THE PRODUCT COULD NOT BE REPRODUCED DURING INSPECTION AND EVALUATION OF THE CARDIOHELP HARDWARE. FURTHER, EXAMINATION OF THE ASSOCIATED DISPOSABLE WAS UNABLE TO BE PERFORMED DUE TO COMPONENT DISPOSAL; THEREFORE, COMMENTARY PERTAINING TO EITHER PERFORMANCE OR MALPERFORMANCE CANNOT BE ADVANCED IN THIS REVIEW." ACCORDING TO THE INSTRUCTION FOR USE (IFU) CHAPTER "CONNECTING THE COMBINED FLOW/BUBBLE SENSOR" THE BUBBLE MONITORING FUNCTION TEST AND FLOW OFF-SET CALIBRATION HAS TO BE PERFORMED BEFORE EVERY USE. THUS, A DEFECTIVE FLOW/BUBBLE SENSOR SHOULD BE DETECTED PRIOR TO USE, DURING PRIMING. IN ADDITION, AS THE CARDIOHELP INCLUDES PRESSURE SENSORS AND A VENOUS PROBE IT IS ABLE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN THE IFU CHAPTER "USING THE EMERGENCY DRIVE WITH THE DISPOSABLE HLS RETAINER" IS STATED THAT THE EMERGENCY DRIVE CAN BE USED TO MANUALLY CONTROL THE BLOOD FLOW IN CASE OF A FAILED CARDIOHELP. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON (B)(6) 2025 FOR THE PERIOD OF (B)(6) 2019 TO (B)(6) 20250. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2019-10-09. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. BASED ON THE RESULTS THE REPORTED FAILURE "BACKFLOW PREVENTION" COULD BE CONFIRMED BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. FURTHER IT WAS STATED BY THE CUSTOMER THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21293323&pc=MKJ | 3030677-2025-000361 | 2025-01-24 05:00 | Death | PHILIPS GOLDWAY (SHENZHEN) INDUSTRIAL INC. | 2025-02-03 05:00 | MKJ | EFFICIA DFM100 | Incorrect Interpretation of Signal | Ventricular Fibrillation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K110825 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE EFFICIA DFM100 DEFIBRILLATOR MONITOR INDICATING THE DEVICE SUGGESTED SHOCK ON AED WHEN NOT NEEDED. AFTER THE CUSTOMER CHECKED THE ECG STRIP, THE ECG SHOWED VENTRICULAR FIBRILLATION WITH A PACED RHYTHM, THE DEVICE WORKED PROPERLY, AND SHOCK WAS ADVISED TWO TIMES BUT THE STAFF WAITED UNTIL TIMEOUT (AUTO DISCHARGE) OVERALL. SIX OF EIGHT SHOCKS WERE PROVIDED DURING CPR. THE PATIENT DIED. THE EFFICIA DFM100 DEFIBRILLATOR, MODEL# 866199, IS SUBSTANTIALLY SIMILAR TO THE HEARTSTART XL+ DEFIBRILLATOR (MODEL# 861290) AND WILL BE REPORTED IN THE UNITED STATES UNDER DEVICE MODEL# 861290.Event Description: PHILIPS RECEIVED A COMPLAINT ON THE EFFICIA DFM100 DEFIBRILLATOR MONITOR INDICATING THE DEVICE SUGGESTED SHOCK ON AED WHEN NOT NEEDED. AFTER THE CUSTOMER CHECKED THE ECG STRIP, THE ECG SHOWED VENTRICULAR FIBRILLATION WITH A PACED RHYTHM, THE DEVICE WORKED PROPERLY, AND SHOCK WAS ADVISED TWO TIMES BUT THE STAFF WAITED UNTIL TIMEOUT (AUTO DISCHARGE) OVERALL. THE PATIENT DIED.Event Description: PHILIPS RECEIVED A COMPLAINT ON THE EFFICIA DFM100 DEFIBRILLATOR MONITOR INDICATING THE DEVICE SUGGESTED SHOCK ON AED WHEN NOT NEEDED. AFTER THE CUSTOMER CHECKED THE ECG STRIP, THE ECG SHOWED VENTRICULAR FIBRILLATION WITH A PACED RHYTHM. THE DEVICE WORKED PROPERLY, AND SHOCK WAS ADVISED TWO TIMES BUT THE STAFF WAITED UNTIL TIMEOUT (AUTO DISCHARGE) OVERALL. THE PATIENT DIED. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. Manufacturer Narrative: THE CUSTOMER INDICATED THAT 8 SHOCKS WERE DELIVERED AND 3 SHOCKS WERE NOT DELIVERED. THE PATIENT HAD A PACEMAKER. THE CUSTOMER INDICATED THE DEVICE WORKED CORRECTLY AND OPCHECK WAS PERFORMED. THE DEVICE WAS RETURNED TO FULL FUNCTIONALITY. PHILIPS DID NOT EVALUATE THE DEVICE. A PHILIPS CLINICAL EXPERT ANALYZED THE PATIENT EVENT FILE AND FOUND A TOTAL OF 8 SHOCKS WERE DELIVERED, AND 4 SHOCKS WERE DISARMED DUE TO A TIMEOUT/TIME TO AUTO DISARM. TIME TO AUTO DISARM OCCURS WHEN A CHARGE IS AUTOMATICALLY DISARMED WHEN THE SHOCK BUTTON WAS NOT PRESSED WITHIN 30 SECONDS OF CHARGING THE DEVICE. THIS FEATURE IS DESIGNED FOR THE SAFETY OF USERS AND PATIENTS TO AVOID AN UNPLANNED SHOCK. THIS FEATURE IS CONFIGURABLE, ENABLING LONGER TIMES BETWEEN CHARGING AND SHOCK DELIVERY OF 60 OR 90 SECONDS. INFORMATION REGARDING THIS CONFIGURATION IS AVAILABLE IN THE EFFICIA DFM100 INSTRUCTIONS FOR USE (IFU) ON PAGE 148. THE REPORTED PROBLEM WAS CONFIRMED AND THE CAUSE OF THE PROBLEM WAS DUE TO USER ERROR. THE DEVICE PERFORMED ACCORDING TO ITS SPECIFICATION AND DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT DEATH. BASED ON THE AVAILABLE INFORMATION, THE DEVICE PERFORMED AS INTENDED. THE ANALYSIS FOUND NO SIGN OF DEVICE MALFUNCTION. IT HAS BEEN CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. Manufacturer Narrative: ADDITIONAL INFORMATION: THE CUSTOMER INDICATED THAT 8 SHOCKS WERE DELIVERED AND 3 SHOCKS WERE NOT DELIVERED. THE PATIENT HAD A PACEMAKER. THE CUSTOMER INDICATED THE DEVICE WORKED CORRECTLY AND OPCHECK WAS PERFORMED. THE DEVICE WAS RETURNED TO FULL FUNCTIONALITY. THIS INVESTIGATION IS PENDING PATIENT EVENT FILE REVIEW. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21293518&pc=MNS | 2518422-2025-101638 | 2025-01-23 05:00 | Death | RESPIRONICS, INC. | 2025-02-03 05:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21293890&pc=DSQ | 2916596-2025-00476 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-02-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Pumping Stopped; Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DECEASED ON (B)(6) 2025 AT THEIR HOME. THE BATTERIES WERE COMPLETELY DEPLETED, AND THE PUMP WAS OFF. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED LOW POWER ADVISORIES BEGINNING AT 9:16 PM ON (B)(6) 2025, WITH THE BATTERIES BEING COMPLETELY DEPLETED BY 10:31 WHEN NO EXTERNAL POWER ALARMS BEGAN. THE EBB WAS THEN DEPLETED AT 12:03 AM ON (B)(6) 2025 WHICH CAUSED THE SYSTEM CONTROLLER AND PUMP TO SHUTT OFF. THIS WAS INDICATIVE THAT THE PATIENT MAY HAVE BEEN SLEEPING ON BATTERIES AS NOTED BY THE CLINIC. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. PER ADDITIONAL INFORMATION, THE DEVICE WAS OPERATING AS EXPECTED. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294048&pc=KZL | 9617604-2025-00084 | 2024-12-12 05:00 | Death | SMITHS HEALTHCARE MANUFACTURING S.A. DE C.V. | 2025-02-03 05:00 | KZL | LEVEL 1 FAST FLOW TRAUMA FLUID WARMER | Collapse; Fluid/Blood Leak; Insufficient Flow or Under Infusion | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THAT SPIKING THE BLOOD TUBING WAS DIFFICULT AND THE DEVICE HAD A LOW FLOW RATE. THERE WAS LEAKING FROM THE BAG AROUND THE SPIKE AND DRIP CHAMBERS, WHICH HAD COLLAPSED DURING INFUSION. THE PATIENT WAS PART OF A MASSIVE TRANSFUSION PROTOCOL AND HAD RECEIVED MULTIPLE BLOOD PRODUCTS AS WELL AS MEDICATIONS. THOUGH THE PATIENT DID RECEIVE BLOOD AND TREATMENT, THE PATIENT ULTIMATELY PASSED AWAY. THE PRODUCT THAT WAS INITIALLY USED IS STILL AVAILABLE AND DOES HAVE COAGULATED BLOOD THROUGHOUT THE TUBING WHICH WAS UNABLE TO BE FLUSHED. Manufacturer Narrative: THIS COMPLAINT WAS INCORRECTLY FILED UNDER THIS REGISTRATION NUMBER, PLEASE REFERENCE MRN 3012307300-2025-03499-00 FOR DETAILS PERTINENT TO THIS EVENT. Manufacturer Narrative: B2. DATE OF DEATH: UNKNOWN. NO INFORMATION HAS BEEN PROVIDED TO DATE. D4. PRIMARY UDI NUMBER: THE PRIMARY DI# HAS BEEN USED AS THE LOT/SERIAL INFORMATION IS UNKNOWN. H3. REASON DEVICE NOT EVALUATED BY MFG: OTHER; DEVICE WAS NOT RETURNED TO MANUFACTURER. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294459&pc=NPT | 2135147-2025-00505 | 2025-01-09 05:00 | Death | ABBOTT MEDICAL | 2025-02-03 05:00 | NPT | FLEXNAV DELIVERY SYSTEM | Use of Device Problem; Difficult to Advance | Low Blood Pressure/ Hypotension; Perforation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE PATIENT WAS NOTED TO HAVE VERY TORTUOUS VASCULATURE, ABDOMINAL ANEURYSM, SEVERE AORTIC STENOSIS, AORTIC/STJ CALCIUM, AND WAS CONSIDERED HIGH RISK FOR SURGICAL INTERVENTION. DURING THE PROCEDURE, THE PHYSICIAN EXPERIENCED DIFFICULTY ADVANCING THE DELIVERY SYSTEM TO THE AORTIC VALVE. AFTER SEVERAL ATTEMPTS, THE ABDOMINAL ANEURYSM HAD BECOME PERFORATED. ITS BELEIVED THAT CAUSE WAS DUE TO TORTUROUS ANATOMY AND THE FORCE ON THE DS TO ADVANCE IT. SEVERAL COVERED STENTS WERE PLACED IN THE ABDOMINAL AORTA AND ILIAC. CPR AND RAPID BLOOD INFUSION WERE PERFORMED UNTIL THE PATIENTS' BLOOD PRESSURE STABILIZED. A SECOND VALVE WAS LOADED IN A NEW DS AND IT WAS SUCCESSFULLY IMPLANTED. THE PATIENT WAS TAKEN TO THE ICU WHERE THEY EXPIRED LATER THAT DAY. THE PHYSICIAN DOESN'T BELIEVE THE PATIENT EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. THE CAUSE OF DEATH IS CURRENTLY UNKNOWN.Event Description: IT WAS REPORTED THAT ON 09 JAN. 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE PATIENT WAS NOTED TO HAVE VERY TORTUOUS VASCULATURE, ABDOMINAL ANEURYSM, SEVERE AORTIC STENOSIS, AORTIC/STJ CALCIUM, AND WAS CONSIDERED HIGH RISK FOR SURGICAL INTERVENTION. DURING THE PROCEDURE, THE PHYSICIAN EXPERIENCED DIFFICULTY ADVANCING THE DELIVERY SYSTEM TO THE AORTIC VALVE. AFTER SEVERAL ATTEMPTS, THE ABDOMINAL ANEURYSM HAD BECOME PERFORATED. ITS BELIEVED THAT CAUSE WAS DUE TO TORTUROUS ANATOMY AND THE FORCE ON THE DS TO ADVANCE IT. SEVERAL COVERED STENTS WERE PLACED IN THE ABDOMINAL AORTA AND ILIAC. CPR AND RAPID BLOOD INFUSION WERE PERFORMED UNTIL THE PATIENTS' BLOOD PRESSURE STABILIZED. A SECOND VALVE WAS LOADED IN A NEW DS AND IT WAS SUCCESSFULLY IMPLANTED. THE PATIENT WAS TAKEN TO THE ICU WHERE THEY EXPIRED LATER THAT DAY. THE PHYSICIAN DOESN'T BELIEVE THE PATIENT EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. THE CAUSE OF DEATH IS CURRENTLY UNKNOWN. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF USE OF DEVICE PROBLEM (USE-RELATED TISSUE DAMAGE), DIFFICULT ADVANCEMENT THROUGH PATIENT ANATOMY, PERFORATION, HYPOTENSION, AND DEATH WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. THE EVENT AND PROVIDED IMAGING WERE REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. BASED ON ALL AVAILABLE INFORMATION, THE REPORTED USE OF DEVICE PROBLEM (USE-RELATED TISSUE DAMAGE), DIFFICULT ADVANCEMENT THROUGH PATIENT ANATOMY, PERFORATION, HYPOTENSION, AND DEATH APPEAR TO BE RELATED TO CHALLENGING PATIENT ANATOMY (SEVERE AORTO-ILIAC ATHEROSCLEROSIS AND TORTUOSITY). THE REPORTED SURGERY AND UNEXPECTED MEDICAL INTERVENTIONS WERE THE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294653&pc=NIK | 2017865-2025-12144 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-03 05:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12146, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12147, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12148. IT WAS REPORTED THAT THE PATIENT PRESENTED WITH AN INFECTION OF THE DEVICE POCKET. THE IMPLANTABLE CARDIAC DEFIBRILLATOR (ICD), THE RIGHT ATRIAL, RIGHT VENTRICULAR AND LEFT VENTRICULAR LEADS WERE EXPLANTED DUE TO INFECTION. THE PATIENT WAS INTUBATED DUE TO THE INFECTION, AND SUBSEQUENTLY EXPIRED. THERE WAS NO ALLEGATION THAT THE IMPLANT PRODUCTS CAUSED THE INFECTION. HOWEVER, THE PHYSICIAN BELIEVED THAT INFECTION WAS RELATED TO THE IMPLANT PROCEDURE. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ERI WHEN RECEIVED. BASED ON THIS INFORMATION, THE DEVICE WAS FOUND TO COMMUNICATE APPROPRIATELY WITH A MERLIN PROGRAMMER AND HAS NOT REACHED THE ELECTIVE REPLACEMENT INDICATOR (ERI). A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: CORRECTION: B2 - OUTCOMES ATTRIBUTED TO ADVERSE AND REPORT TYPE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294676&pc=NIK | 2017865-2025-12146 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-03 05:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12144, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12147, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12148. IT WAS REPORTED THAT THE PATIENT PRESENTED WITH AN INFECTION OF THE DEVICE POCKET. THE IMPLANTABLE CARDIAC DEFIBRILLATOR (ICD), THE RIGHT ATRIAL, RIGHT VENTRICULAR AND LEFT VENTRICULAR LEADS WERE EXPLANTED DUE TO INFECTION. THE PATIENT WAS INTUBATED DUE TO THE INFECTION, AND SUBSEQUENTLY EXPIRED. THERE WAS NO ALLEGATION THAT THE IMPLANT PRODUCTS CAUSED THE INFECTION. HOWEVER, THE PHYSICIAN BELIEVED THAT INFECTION WAS RELATED TO THE IMPLANT PROCEDURE. Manufacturer Narrative: CORRECTION: B2 - OUTCOMES ATTRIBUTED TO ADVERSE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294743&pc=NVY | 2017865-2025-12147 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-03 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12144, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12146, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12148. IT WAS REPORTED THAT THE PATIENT PRESENTED WITH AN INFECTION OF THE DEVICE POCKET. THE IMPLANTABLE CARDIAC DEFIBRILLATOR (ICD), THE RIGHT ATRIAL, RIGHT VENTRICULAR AND LEFT VENTRICULAR LEADS WERE EXPLANTED DUE TO INFECTION. THE PATIENT WAS INTUBATED DUE TO THE INFECTION, AND SUBSEQUENTLY EXPIRED. THERE WAS NO ALLEGATION THAT THE IMPLANT PRODUCTS CAUSED THE INFECTION. HOWEVER, THE PHYSICIAN BELIEVED THAT INFECTION WAS RELATED TO THE IMPLANT PROCEDURE. Manufacturer Narrative: CORRECTION: B2 - OUTCOMES ATTRIBUTED TO ADVERSE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294752&pc=NVN | 2017865-2025-12148 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-03 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12144, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12146, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12147. IT WAS REPORTED THAT THE PATIENT PRESENTED WITH AN INFECTION OF THE DEVICE POCKET. THE IMPLANTABLE CARDIAC DEFIBRILLATOR (ICD), THE RIGHT ATRIAL, RIGHT VENTRICULAR AND LEFT VENTRICULAR LEADS WERE EXPLANTED DUE TO INFECTION. THE PATIENT WAS INTUBATED DUE TO THE INFECTION, AND SUBSEQUENTLY EXPIRED. THERE WAS NO ALLEGATION THAT THE IMPLANT PRODUCTS CAUSED THE INFECTION. HOWEVER, THE PHYSICIAN BELIEVED THAT INFECTION WAS RELATED TO THE IMPLANT PROCEDURE. Manufacturer Narrative: CORRECTION: B2 - OUTCOMES ATTRIBUTED TO ADVERSE, AND E4 - INITIAL REPORTER ALSO SENT THE REPORT TO FDA. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294753&pc=DSQ | 2916596-2025-00590 | 2025-01-17 05:00 | Death | THORATEC CORPORATION | 2025-02-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER DE-AIRING WAS COMPLETED, WEANING CARDIOPULMONARY BYPASS (CPB) HAD STARTED. THE HEARTMATE 3 (HM3) (B)(6) WAS TURNED ON 3000 ROTATIONS PER MINUTE (RPM), GRADUALLY RAMPED UP TO 4000RPM, THEN CONTINUED TO INCREASE RPM UNTIL 4800 RPM WAS REACHED. IT WAS NOTED ON SYSTEM MONITOR THAT FLOWS WERE RUNNING LOW AT 0.2 LITERS PER MINUTE (LPM) WITH PULSATILITY INDEX (PI) AT 11-13, AND POWER AT 2.7 WATTS. ON ECHOCARDIOGRAM, THE LEFT VENTRICLE (LV) WAS NOT UNLOADING. RPM WAS INCREASED TO 5000 RPM AND FLOWS DECREASED TO 0.0 LPM. PI AND POWER REMAINED THE SAME. TROUBLESHOOTING WAS STARTED. AFTER ECHOCARDIOGRAM ASSESSMENT OF LV AND OUTFLOW GRAFT (OFG) ANASTOMOSIS IT WAS DECIDED THAT THE PUMP HAD POSSIBLE "AIR LOCK". THE HM3 WAS TURNED OFF, OFG WAS CLAMPED AND HM3 UNLOCKED FROM APICAL CUFF. THE SURGEON BACK FLOWED BLOOD THROUGH THE PUMP AND THEN REPOSITIONED THE HM3 BACK TO APICAL CUFF VIA WET-TO-WET CONNECTION. THE POWER MODULE AND SYSTEM MONITOR WERE SWAPPED, THE HM3 WAS TURNED BACK ON AND THE DE-AIRING PROCESS RESTARTED. WEANING CARDIOPULMONARY BYPASS (CPB) WAS RESTARTED, AND THE RPM WAS RAMPED TO 4000. THE FLOW REFLECTED 0.5 LPM. THEN AS RAMPING UP TO 5000 RPM, FLOWS DECREASED TO 0.0 LPM. THE LV WAS FULL AND NOT UNLOADED. IT WAS DECIDED TO SWAP HM3 FOR POSSIBLE PUMP ISSUES. AFTER THE PUMP EXCHANGE AND DE-AIRING OF HM3 (B)(6) THE RPM WAS RAMPED UP TO 4000, FLOWS WERE AT 0.5 LPM, PI WAS 10-11, AND POWER WAS 2.9 WATTS. AN ECHOCARDIOGRAM ASSESSMENT SHOWED BLOOD MOVING INTO THE FLOW. RPM WAS THEN RAMPED TO 5000 RPM AND FLOWS DECREASED TO 0.1-0.2LPM. THE SURGEON SAW SOMETHING ON THE ECHOCARDIOGRAM AND REMOVED THE PUMP TO REMOVE TISSUE IN THE LEFT VENTRICLE (LV) FOR POSSIBLE OBSTRUCTION ISSUES. WHEN THE SURGEON UNLOCKED THE PUMP FROM APICAL CUFF USING THE UNLOCK TOOL, THEY SOMEHOW BENT THE LOCKING MECHANISM AND COULD NOT RE-ATTACH AND LOCK THE HM3 TO THE APICAL CUFF, SO ANOTHER HM3 (B)(6) WAS BROUGHT INTO ROOM TO PLACE IN PATIENT. A RIGHT VENTRICULAR ASSIST DEVICE WAS PLACED. INTRAVENOUS (IV) INOTROPES AND BLOOD/BLOOD PRODUCT TRANSFUSIONS WERE GIVEN. DUE TO THE LONG CARDIOPULMONARY BYPASS, 2 PUMP EXCHANGES, RVAD PLACEMENT, MEDICATIONS, AND BLOOD PRODUCT REQUIREMENTS, THE PATIENT DID NOT SURVIVE THE IMPLANT PROCEDURE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. CHAPTER 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING RIGHT HEART FAILURE AND BLEEDING. CHAPTER 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. ADDITIONALLY, THIS SECTION STATES THAT SUDDEN RIGHT VENTRICULAR FAILURE MAY DEVELOP DURING OR SHORTLY AFTER PUMP IMPLANTATION AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: CORRECTION - SECTION B3 DATE OF DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21295201&pc=GZL | 3002250546-2025-00001 | 2025-01-02 05:00 | Death | FHC, INC. | 2025-02-03 05:00 | GZL | MICROTARGETING¿ ELECTRODE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | K033173 |  | 1 | Event Description: A CASE WAS COMPLETED ON (B)(6) 2025 AT (B)(6) MEDICAL CENTER, (B)(6) WITH DR. (B)(6), WHERE A 57 Y/O PATIENT HAD BOSTON SCIENTIFIC LEADS IMPLANTED BILATERALLY IN STN. MER WAS PERFORMED ON BOTH HEMISPHERES SIMULTANEOUSLY USING FHC 44975Z MICROELECTRODES (A SINGLE TRACK PER SIDE) AND A GUIDELINE 4000 LP+ SYSTEM. OTHER FHC PRODUCTS USED DURING THE CASE WERE STARFIX PLATFORM, 66-IT-03P INSERTION TUBES, 66-ST-12P STYLET. PATIENT RESPONDED WELL TO STIMULATION AND O-ARM IMAGING SHOWED ACCURATE TARGETING. POST-OP CT SHOWED NO ABNORMALITIES AND PATIENT WAS DISCHARGED THE SAME DAY OF SURGERY. THE PATIENT WAS DOING WELL AND HAD GONE BACK TO WORK. FACILITY DID NOT RETAIN LOT# OF PRODUCT USED AND DISPOSED OF PRODUCTS. ON 01/16/2025 FHC, INC. REPRESENTATIVE WAS ATTENDING ANOTHER CASE AT (B)(6) MEDICAL CENTER WITH DR. (B)(6) WHERE DURING THE CASE, THE DR. INFORMED THE FHC REP, THAT THE PATIENT THEY OPERATED ON (B)(6) 2025 WAS FOUND DEAD THIS WEEK AND THAT THEY WERE GOING TO CONDUCT AN AUTOPSY TO DETERMINE THE CAUSE OF DEATH. THE AUTOPSY WAS CONDUCTED ON (B)(6) 2025 WHERE AN INTRACEREBRAL HEMORRHAGE (ICH) AT DBS-LEAD ENTRY SITE AND THALAMUS WAS DETERMINED TO BE CAUSE OF DEATH. SINCE POST-OP CT WAS CLEAR, IT WAS CONSIDERED A DELAYED ICH. Manufacturer Narrative: FHC IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY FHC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. FHC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, FHC, OR ITS EMPLOYEES THAT THE DEVICE, FHC OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. FHC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21295388&pc=DSQ | 2916596-2025-00603 | 2025-01-20 05:00 | Death | THORATEC CORPORATION | 2025-02-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Multiple Organ Failure; Liver Failure; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. THE PATIENT HAD COMPLICATED POST-OPERATIVE COURSE INCLUDING REQUIRING INCREASING VASOPRESSOR SUPPORT, SHOCK LIVER, MULTIDRUG-RESISTANT ORGANISM (MDRO) ENTEROBACTER INFECTION, STROKE, CODE. THE PATIENT'S SPOUSE REQUESTED COMFORT CARE MEASURES. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), INFECTION (LOCAL, DRIVELINE OR PUMP POCKET INFECTION), STROKE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21295903&pc=JAA | 9613445-2025-00002 | 2025-01-04 05:00 | Death | GE HANGWEI MEDICAL SYSTEMS CO., LTD. | 2025-02-03 05:00 | JAA | OPTIMA IGS 320 | Insufficient Cooling | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K122457 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT DURING AN EMERGENT VASCULAR PROCEDURE, THE GE HEALTHCARE SYSTEM DISPLAYED A COOLING FAILURE OF THE CHILLER AND X-RAY BECAME UNAVAILABLE A FEW MINUTES LATER. THE PROCEDURE COULD NOT BE COMPLETED AND THE PATIENT EXPIRED LATER THE SAME DAY. GE HEALTHCARE INVESTIGATION HAS BEEN OPENED AND IS IN PROGRESS. Manufacturer Narrative: GE HEALTHCARE HAS COMPLETED ITS INVESTIGATION. ANALYSIS OF SYSTEM LOGS AND SERVICE RECORDS CONFIRMED THAT THE SYSTEM OPERATES ACCORDING TO SPECIFICATIONS. THE X-RAY LOSS WAS ATTRIBUTED TO INADEQUATE ENVIRONMENTAL CONTROL, WITH ROOM TEMPERATURES FREQUENTLY OUT OF RANGE SINCE SEPTEMBER 2021. THE ROOM TEMPERATURE CONTROLS ARE SPECIFIED IN THE GE HEALTHCARE PRE-INSTALLATION MANUAL. THE PATIENT'S DEATH WAS DETERMINED TO BE AN ISOLATED INCIDENT CAUSED BY INADEQUATE ENVIRONMENTAL CONTROL, AND THE DEVICE DID NOT CONTRIBUTE TO THE EVENT. THE INVESTIGATION IS CLOSED AND ACTION HAS BEEN TAKEN BY THE CUSTOMER TO REPLACE THE A/C MODULE WITHIN THE TECHNICAL ROOM. POST-REPLACEMENT SYSTEM LOGS HAVE BEEN REVIEWED AND ARE WITHIN ACCEPTABLE PARAMETERS. NO FURTHER ACTIONS ARE PLANNED BY GE HEALTHCARE. Manufacturer Narrative: LEGAL MANUFACTURER: HCS BEIJING - WEST AREA OF BUILDING NO.3, NO.1 YONGCHANG NORTH ROAD, BEIJING ECONOMIC AND TECHNOLOGICAL DEVELOPMENT AREA CHINA BEIJING BEIJING, 100176. D4: UDI NOT REQUIRED AS DEVICE IS MANUFACTURED PRIOR TO UDI COMPLIANCE DATE. DEVICE INSTALLATION DATE IS 29-FEB-2016. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21296225&pc=DYB | 2124215-2025-05487 | 2025-01-14 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-03 05:00 | DYB | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation | K183655 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A RE-DO ATRIAL FIBRILLATION ABLATION PROCEDURE, A VERSACROSS ACCESS KIT WAS SELECTED FOR USE. THE PATIENT WAS INTUBATED AND IN ATRIAL FLUTTER. THE PHYSICIAN MAPPED THE RIGHT ATRIUM (RA) WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER. THEN, THE PHYSICIAN CROSSED TO LEFT SIDE USING VERSACROSS SHEATH AND VERSACROSS RF WIRE AND MAPPED LEFT ATRIUM (LA). AFTER MAPPING, THE PHYSICIAN DECIDED TO USE FARAPULSE. THE VERSACROSS RF WIRE WAS PLACED INTO THE LEFT SUPERIOR PULMONARY VEIN (LSPV) AND THEN EXCHANGED FOR THE FARADRIVE SHEATH. ONCE THE SHEATH WAS ACROSS, THE PHYSICIAN PASSED IT BACK AND FORTH A COUPLE OF TIMES BEFORE PULLING IT INTO THE RA AND ADVANCING THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER INTO THE LA. SOMETIME DURING THIS PERIOD, THE PATIENT BLOOD PRESSURE DROPPED, AND ICE IMAGINING BECAME VERY FOGGY. AN EFFUSION CHECK WAS PERFORMED AND CONFIRMED THE PATIENT HAD A LARGE EFFUSION ALREADY. THE PHYSICIAN IMMEDIATELY WORKED TO PERFORM A PERICARDIOCENTESIS. A CARDIOTHORACIC (CT) SURGERY CAME IN AND RESUSCITATION WAS STARTED. WHEN CT SURGEON OPENED THE CHEST, IT WAS FOUND THE VERSACROSS THROUGH THE APPENDAGE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH IS BLOOD LOSS FROM PERFORATED LEFT ATRIAL APPENDAGE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). IN THE PHYSICIAN OPINION, FLOSSING THE SHEATH TO DILATE THE SEPTUM MAY HAVE PULLED THE WIRE AND ADVANCED THE SHEATH AND WIRE THROUGH THE APPENDAGE. THERE WERE NO VERSACROSS MALFUNCTION OR CONTRIBUTION TO PATIENT COMPLICATION REPORTED. THE VERSACROSS DILATOR WAS ALSO PRESENT INSIDE THE SHEATH. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS RF WIRE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21296306&pc=DXF | 2124215-2025-05485 | 2025-01-14 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-03 05:00 | DXF | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation | K150709 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A RE-DO ATRIAL FIBRILLATION ABLATION PROCEDURE, A VERSACROSS ACCESS KIT WAS SELECTED FOR USE. THE PATIENT WAS INTUBATED AND IN ATRIAL FLUTTER. THE PHYSICIAN MAPPED THE RIGHT ATRIUM (RA) WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER. THEN, THE PHYSICIAN CROSSED TO LEFT SIDE USING VERSACROSS SHEATH AND VERSACROSS RF WIRE AND MAPPED LEFT ATRIUM (LA). AFTER MAPPING, THE PHYSICIAN DECIDED TO USE FARAPULSE. THE VERSACROSS RF WIRE WAS PLACED INTO THE LEFT SUPERIOR PULMONARY VEIN (LSPV) AND THEN EXCHANGED FOR THE FARADRIVE SHEATH. ONCE THE SHEATH WAS ACROSS, THE PHYSICIAN PASSED IT BACK AND FORTH A COUPLE OF TIMES BEFORE PULLING IT INTO THE RA AND ADVANCING THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER INTO THE LA. SOMETIME DURING THIS PERIOD, THE PATIENT BLOOD PRESSURE DROPPED, AND ICE IMAGINING BECAME VERY FOGGY. AN EFFUSION CHECK WAS PERFORMED AND CONFIRMED THE PATIENT HAD A LARGE EFFUSION ALREADY. THE PHYSICIAN IMMEDIATELY WORKED TO PERFORM A PERICARDIOCENTESIS. A CARDIOTHORACIC (CT) SURGERY CAME IN AND RESUSCITATION WAS STARTED. WHEN CT SURGEON OPENED THE CHEST, IT WAS FOUND THE VERSACROSS THROUGH THE APPENDAGE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH IS BLOOD LOSS FROM PERFORATED LEFT ATRIAL APPENDAGE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). IN THE PHYSICIAN OPINION, FLOSSING THE SHEATH TO DILATE THE SEPTUM MAY HAVE PULLED THE WIRE AND ADVANCED THE SHEATH AND WIRE THROUGH THE APPENDAGE. THERE WERE NO VERSACROSS MALFUNCTION OR CONTRIBUTION TO PATIENT COMPLICATION REPORTED. THE VERSACROSS DILATOR WAS ALSO PRESENT INSIDE THE SHEATH. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS DILATOR. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21296708&pc=NPT | 2025587-2025-00947 | 2023-03-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-03 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE BLOCKED THE RIGHT CORONARY ARTERY DURING DEPLOYMENT. THE VALVE WAS IMMEDIATELY PULLED UP WITH A SNARE. A SECOND VALVE WAS IMPLANTED IN A DEEPER POSITION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED THAT CARDIOPULMONARY ARREST OCCURRED DUE TO THE RIGHT CORONARY ARTERY OCCLUSION. AFTER 10 MINUTES OF CARDIOPULMONARY ARREST, CARDIOPULMONARY MACHINES WERE INTRODUCED AND THE HEART RATE RESUMED. THE PATIENT SUFFERED FROM A POST-OPERATIVE CEREBRAL INFARCTION, AND NO RECOVERY OF CONSCIOUSNESS WAS OBSERVED. SUBSEQUENTLY, THE PATIENT DIED APPROXIMATELY THREE MONTHS AFTER THE VALVE IMPLANT. THE CAUSE OF DEATH WAS INFECTION DUE TO PNEUMONIA. PER THE PHYSICIAN, THERE WAS NO RELATION BETWEEN THE DEVICE NOR THE PROCEDURE AND THE DEATH.Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE BLOCKED THE RIGHT CORONARY ARTERY DURING DEPLOYMENT. THE VALVE WAS IMMEDIATELY PULLED UP WITH A SNARE. A SECOND VALVE WAS IMPLANTED IN A DEEPER POSITION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED THAT CARDIOPULMONARY ARREST OCCURRED DUE TO THE RIGHT CORONARY ARTERY OCCLUSION. AFTER 10 MINUTES OF CARDIOPULMONARY ARREST, CARDIOPULMONARY MACHINES WERE INTRODUCED AND THE HEART RATE RESUMED. THE PATIENT SUFFERED FROM A POST-OPERATIVE CEREBRAL INFARCTION, AND NO RECOVERY OF CONSCIOUSNESS WAS OBSERVED. SUBSEQUENTLY, THE PATIENT DIED APPROXIMATELY THREE MONTHS AFTER THE VALVE IMPLANT. THE CAUSE OF DEATH WAS INFECTION DUE TO PNEUMONIA. PER THE PHYSICIAN, THERE WAS NO RELATION BETWEEN THE DEVICE NOR THE PROCEDURE AND THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT THE STROKE WAS CONFIRMED WITH MAGNETIC RESONANCE IMAGING (MRI). THE SYMPTOMS PRESENTED 2-3 HOURS AFTER THE VALVE IMPLANT. NO PERMANENT IMPAIRMENT RESULTED FROM THE STROKE. PER THE PHYSICIAN, IT WAS UNKNOWN IF THE STROKE WAS RELATED TO THE VALVE OR DELIVERY CATHETER SYSTEM (DCS). PER THE PHYSICIAN, THE CAUSE OF THE STROKE WAS THE BALLOON DILATION OR CARDIAC MASSAGE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21296906&pc=DSQ | 3007042319-2025-00142 | 2024-12-17 05:00 | Death | HEARTWARE, INC. | 2025-02-03 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT WAS ADMITTED FOR MULTI-SYSTEM ORGAN FAILURE (MSOF) RELATED TO RIGHT HEART FAILURE (RHF). THE PATIENT EXPERIENCED CARDIAC ARREST, AND CARDIOPULMONARY RESUSCITATION WAS UNSUCCESSFUL. THE PATIENT SUBSEQUENTLY DIED FROM MSOF DUE TO RHF. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED, OUT OF SERVICE. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED FROM MULTI-SYSTEM ORGAN FAILURE DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, MULTI-ORGAN FAILURE, CARDIOPULMONARY ARREST, RIGHT VENTRICULAR FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21297333&pc=MVK | 3008642652-2025-01243 | 2025-01-03 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TEN TREATMENTS FROM THE LIFEVEST. THE DEVICE WAS STARTED UP AT 21:25:48 ON (B)(6) 2025. AT 02:29:36 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG IS NOT AVAILABLE. BETWEEN 02:30:06 AND 02:41:08, THE PATIENT RECEIVED TEN TREATMENTS FROM THE LIFEVEST. THE RHYTHM AT THE TIME OF THESE TREATMENTS IS UNAVAILABLE. THE POST SHOCK RHYTHM OF THESE TREATMENTS IS UNAVAILABLE. THE ELECTRODE BELT WAS DISCONNECTED AT 04:33:54 ON (B)(6) 2025. PER REVIEW OF THE HOLTER DATA, PATIENT IS LAST SEEN IN SINUS RHYTHM @ 80 BPM WHEN THE HOLTER DATA ENDS AT 20:47:04 ON (B)(6) 2025. REPORTING OUT OF AN ABUNDANCE OF CAUTION DUE TO THE ECG NOT BEING AVAILABLE AT THE TIME OF THE TREATMENTS. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21297487&pc=NVN | 2649622-2025-03261 | 2025-01-20 05:00 | Death | MPRI | 2025-02-03 05:00 | NVN | SELECTSECURE MRI SURESCAN | Under-Sensing; Decreased Sensitivity | Tachycardia; No Clinical Signs, Symptoms or Conditions | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING PRIOR TO DETECTION OF A VENTRICULAR FIBRILLATION (VF) EP ISODE. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT REPROGRAMMING WAS DONE. A LITTLE OVER A MONTH LATER, THE PATIENT PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT BEFORE THE PATIENT PASSED, THE RV LEAD EXHIBITED A DROP IN THE R WAVE MEASUREMENT.Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING PRIOR TO DETECTION OF A VENTRICULAR FIBRILLATION (VF) EP ISODE. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT REPROGRAMMING WAS DONE. A LITTLE OVER A MONTH LATER, THE PATIENT PASSED AWAY. IT WAS FURTHER REPORTED THAT BEFORE THE PATIENT PASSED, THE RV LEAD EXHIBITED A DROP IN THE R WAVE MEASUREMENT. IT WAS ALSO REPORTED THAT UNDERSENSING OF VENTRICULAR TACHYCARDIA (VT) WAS OBSERVED IN THE PATIENT'S LAST REMOTE TRANSMISSION. Manufacturer Narrative: CORRECTED: B5 AND H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUED FROM D10: 694045 LEAD, 6945-65 LEAD IMPLANTED: (B)(6) 1999. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21297498&pc=DXT | 2520313-2025-00002 | 2025-01-30 05:00 | Death | BAYER MEDICAL CARE INC. | 2025-02-03 05:00 | DXT | STELLANT DUAL CT INJECTOR | Insufficient Information | Air Embolism | K082905 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT AN ALLEGED AIR INJECTION AND SUBSEQUENT PATIENT DEATH OCCURRED WHILE A MEDRAD® STELLANT CT INJECTION SYSTEM (SN (B)(6)) WAS IN USE. A REQUEST FOR ADDITIONAL INFORMATION, INCLUDING AN AUTOPSY REPORT AND/OR CAUSE OF DEATH REPORT, HAS BEEN MADE. HOWEVER, THE HOSPITAL INDICATED THAT A POLICE INVESTIGATION WAS UNDERWAY AND THEREFORE, LIMITED INFORMATION IS AVAILABLE AT THIS TIME.Event Description: THE CUSTOMER REPORTED THAT AN ALLEGED AIR INJECTION AND SUBSEQUENT PATIENT DEATH OCCURRED WHILE A MEDRAD® STELLANT CT INJECTION SYSTEM (SN (B)(6)) WAS IN USE. A REQUEST FOR ADDITIONAL INFORMATION, INCLUDING AN AUTOPSY REPORT AND/OR CAUSE OF DEATH REPORT, HAS BEEN MADE. HOWEVER, THE HOSPITAL INDICATED THAT A POLICE INVESTIGATION WAS UNDERWAY AND THEREFORE, LIMITED INFORMATION IS AVAILABLE AT THIS TIME.Event Description: THE CUSTOMER REPORTED THAT AN ALLEGED AIR INJECTION AND SUBSEQUENT PATIENT DEATH OCCURRED WHILE A MEDRAD® STELLANT CT INJECTION SYSTEM (SN (B)(6)) WAS IN USE. A REQUEST FOR ADDITIONAL INFORMATION, INCLUDING AN AUTOPSY REPORT AND/OR CAUSE OF DEATH REPORT, HAS BEEN MADE. HOWEVER, THE HOSPITAL INDICATED THAT A POLICE INVESTIGATION WAS UNDERWAY AND THEREFORE, LIMITED INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: ON JULY 17, 2025, BAYER MEDICAL CARE WAS INFORMED THAT THE MEDRAD® STELLANT CT INJECTION SYSTEM (SN (B)(6)) HAS NOT BEEN RELEASED FOR EVALUATION. MULTIPLE DOCUMENTED ATTEMPTS HAVE BEEN MADE TO GAIN SPECIFIC INFORMATION RELATED TO THE REPORTED EVENT, INCLUDING DISPOSABLE INFORMATION; HOWEVER, THOSE ATTEMPTS HAVE BEEN UNSUCCESSFUL. THEREFORE, AN INVESTIGATION OF THE EQUIPMENT AND/OR DISPOSABLES CANNOT BE CARRIED OUT. OUR INVESTIGATION CONCLUDED THAT THERE IS INSUFFICIENT EVIDENCE OF PRODUCT MALFUNCTION OR THE FAILURE OF THE DEVICE TO MEET SPECIFICATION. THIS INFORMATION DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE, THE COMPANY, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A REPORTABLE EVENT. Manufacturer Narrative: ON FEBRUARY 28, 2025, BAYER MEDICAL CARE WAS INFORMED THAT THE POLICE HAVE SEALED THE ROOM AT THE HOSPITAL. THEREFORE, AN INVESTIGATION OF EQUIPMENT AND/OR DISPOSABLES CANNOT BE PERFORMED AT THIS TIME. IN THE EVENT THAT ADDITIONAL INFORMATION IS OBTAINED, A FOLLOW-UP REPORT WILL BE SUBMITTED. THIS INFORMATION DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE, THE COMPANY, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A REPORTABLE EVENT. Manufacturer Narrative: A BAYER SERVICE INJECTOR SYSTEM CHECKOUT, AS WELL AS A REQUEST FOR THE RETURN OF THE DISPOSABLES AND/OR LOT NUMBERS USED IN THE REPORTED INCIDENT, HAS BEEN MADE AND IS AWAITING THE CUSTOMER RESPONSE. IN THE EVENT THAT ADDITIONAL INFORMATION IS OBTAINED, A FOLLOW-UP REPORT WILL BE SUBMITTED. THIS INFORMATION DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE, THE COMPANY, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A REPORTABLE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21298152&pc=DSQ | 2916596-2025-00817 | 2014-08-01 04:00 | Death | THORATEC CORPORATION | 2025-02-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT EFFECT OF LEFT VENTRICULAR (LV) SIZE ON CLINICAL OUTCOMES FOLLOWING A DURABLE LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION WITH HEARTMATE 3 (HM3) WAS UNCLEAR. THE CLINICAL OUTCOMES OF PATIENTS AFTER HM3 AS A FUNCTION OF LV SIZE IN A MULTICENTER INTERNATIONAL REGISTRY WAS INVESTIGATED. A RETROSPECTIVE PROPENSITY SCORE-ADJUSTED ANALYSIS OF ADULT PATIENTS FROM THE EUROMACS REGISTRY WHO UNDERWENT A HM3 IMPLANTATION BETWEEN AUG2014 AND FEB2023 WAS PERFORMED. PATIENTS WITH PREOPERATIVE LV END DIASTOLIC DIAMETER (LVEDD) ASSESSMENTS AND WITHOUT PREVIOUS OR CONCOMITANT CARDIAC SURGERIES AND TEMPORARY MECHANICAL CIRCULATORY SUPPORT WERE INCLUDED. PATIENTS WERE DIVIDED INTO THREE GROUPS ACCORDING TO THEIR LVEDD: =65; 65-80 AND =80 MM. RESULTS CONTAINED 860 PATIENTS IDENTIFIED; WITH A MEAN AGE 54.4±12.4 YRS, MEAN BSA 2.0±0.23 M2, AND MEAN BMI 27.3±5.0 KG/M2, 87.9% WERE MALE, 40% HAD ISCHEMIC CARDIOMYOPATHY, THE GROUP WITH A LVEDD =65 MM WAS SIGNIFICANTLY OLDER (56.1±11.0 YRS VS.53.8±13.0 YRS VS. 53.0±12.3 YRS, P=0.015), MORE LIKELY FEMALE (17.9% VS. 11.6% VS. 4.9%, P<0.001) AND HAD SMALLER BSA (1.95±2.03 VS.2.01±0.22 VS. 2.08±0.22, P<0.01) AND BMI (26.6±4.8 VS. 27.4±5.1 VS. 28.2±5.0, P=0.01), COMPARED TO THE GROUPS WITH LVEDD OF 65-80 MM AND =80 MM RESPECTIVELY. THE LVEDD =65 GROUP PRESENTED A HIGHER PREVALENCE OF INTUBATED PATIENTS (9.8% VS. 4.0% VS. 2.9%, P=0.003) AND DIABETES (34.1% VS, 26.1% VS. 20.6%, P=0.01). PREVALENCE OF MODERATE TO SEVERE MITRAL REGURGITATION INCREASED WITH LVEDD (52.7% VS. 73.4% VS. 73%, P=0.004), RESPECTIVELY. THE 1-YEAR SURVIVAL SIGNIFICANTLY INCREASED WITH LVEDD (75.7% VS. 84.6% VS. 89.5%, P=0.003). PATIENTS WITH LVEDD =65 MM HAD A SIGNIFICANTLY HIGHER RISK OF POSTOPERATIVE RIGHT HEART FAILURE (14.7% VS.10.4% VS. 6.8%, P=0.03). THE RISK FOR POSTOPERATIVE STROKE DID NOT DIFFER BETWEEN THE GROUPS (10% VS. 12.8% VS. 9.5%, P=0.4). IN CONCLUSION, SMALL LV SIZE (LVEDD =65 MM) SIGNIFICANTLY WORSENED 1-YEAR MORTALITY AND POST-OPERATIVE RIGHT HEART FAILURE IN ADULTS UNDERGOING HEARTMATE 3 IMPLANTATION. PATIENTS WITH LVEDD =65 MM PRESENTED A HIGHER INCIDENCE OF RIGHT HEART FAILURE. THERAPY OPTIMIZATION FOR PATIENTS WITH A LVEDD =65 MM UNDERGOING A DURABLE LVAD IMPLANTATION IS WARRANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO DEVICES WERE RETURNED FOR THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION WERE NOT PROVIDED AND WERE UNABLE TO BE DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21300064&pc=NPT | 2025587-2025-00955 | 2024-10-16 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-03 05:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Fluid/Blood Leak; Insufficient Information; Patient Device Interaction Problem | Sepsis; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A TRANSCATHETER BIOPROSTHETIC AORTIC VALVE WAS IMPLANTED. AT A UNSPECIFIED TIME, MILD AORTIC INSUFFICIENCY WAS IDENTIFIED. APPROXIMATELY 9 YEARS AND 7 MONTHS FOLLOWING THE VALVE IMPLANT THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT SEPSIS WAS THE CAUSE OF DEATH. AN AUTOPSY WAS NOT PERFORMED.Event Description: IT WAS REPORTED THAT A TRANSCATHETER BIOPROSTHETIC AORTIC VALVE WAS IMPLANTED. AT A UNSPECIFIED TIME, MILD AORTIC INSUFFICIENCY WAS IDENTIFIED. APPROXIMATELY 9 YEARS AND 7 MONTHS FOLLOWING THE VALVE IMPLANT THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT SEPSIS WAS THE CAUSE OF DEATH. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: UPDATED DATA: H6. CONCLUSION: THE VALVE REMAINED IMPLANTED; THEREFORE, ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. ADDITIONALLY, NO IMAGES OF THE IMPLANTATION PROCEDURE OR THE IMPLANTED VALVE WERE SUBMITTED TO MEDTRONIC FOR REVIEW. THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THIS PRODUCT MET ALL MANUFACTURING SPECIFICATIONS FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. REGURGITATION IS A KNOWN POTENTIAL ADVERSE EFFECT PER THE DEVICE INSTRUCTIONS FOR USE (IFU) AND CAN BE CAUSED BY A VARIETY OF FACTORS, INCLUDING VALVE POSITIONING, PATIENT ANATOMY, OR PRESENCE OF PRE-EXISTING PATIENT CONDITIONS. IN THIS CASE, NO ROOT CAUSE CAN BE DETERMINED. PATIENT DEATH IS A POTENTIAL RISK ASSOCIATED WITH THE IMPLANTATION OF A BIOPROSTHESIS VALVE PER THE DEVICE IFU. A PROCEDURE- OR VALVE-RELATED DEATH IS AN INHERENT RISK AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. WITH THE LIMITED INFORMATION AVAILABLE, A CONCLUSIVE CAUSE COULD NOT BE DETERMINED AND A RELATIONSHIP BETWEEN THE DEVICE AND THE DEATH COULD NOT BE ESTABLISHED. NO OTHER TECHNICAL ASSESSMENTS WERE PERFORMED. THERE WAS NO INFORMATION TO SUGGEST A DEVICE QUALITY DEFICIENCY MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. THIS EVENT IS FORESEEN IN RISK MANAGEMENT, IS INCLUDED IN MONITORING/TRENDING, AND NO FURTHER ACTION IS RECOMMENDED AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21300178&pc=NKM | 2135147-2025-00518 | 2025-01-09 05:00 | Death | ABBOTT MEDICAL | 2025-02-03 05:00 | NKM | MITRACLIP | Incomplete Coaptation | Low Blood Pressure/ Hypotension; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED ON 09 JANUARY 2025 THAT A PATIENT PRESENTED WITH GRADE 3-4 MIXED MITRAL REGURGITATION (MR) A LARGE FLAIL, FRAGILE LEAFLETS, AND A PROLAPSED LEAFLET FOR A MITRACLIP PROCEDURE. IMAGING WAS POOR DUE TO PATIENT ANATOMY. THE DEVICE WAS PREPARED PER INSTRUCTIONS FOR USE (IFU). THREE CLIPS WERE IMPLANTED. POST-IMPLANT A SINGLE LEAFLET DEVICE ATTACHMENT (SLDA) WAS OBSERVED ON THE THIRD CLIP. THE CLIP WAS DETACHED FROM THE ANTERIOR LEAFLET AND REMAINED ATTACHED TO THE POSTERIOR LEAFLET. A CLEFT WAS NOTED ON THE A1 AND A2 LEAFLETS. ONE NTW CLIP WAS PLACED MEDIALLY TO STABILIZE THE SLDA. AN ADDITIONAL NTW COULD NOT BE IMPLANTED LATERAL TO THE SLDA CLIP. THE TISSUE ON THE LATERAL SIDE WAS DETERMINED TO BE TOO FRAGILE FOR DEPLOYMENT. THE LEAFLETS WERE DAMAGED FROM THE SLDA AND WERE DIFFICULT TO GRASP. TO AVOID FURTHER DAMAGE, THE CLIP WAS REMOVED AND THE PROCEDURE ENDED WITHOUT FURTHER REDUCING THE MR. THE FINAL MR GRADE WAS 3. PER THE PHYSICIAN, THE SLDA WAS DUE TO PRE-EXISTING PATIENT ANATOMY. ON AN ESTIMATED DATE, (B)(6) 2025 (SEVERAL DAYS POST-PROCEDURE), IT WAS REPORTED THAT THE PATIENT HAD EXPIRED DUE TO HEMODYNAMIC INSTABILITY AND MULTIMORBIDITY. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: ON AN ESTIMATED DATE, (B)(6) 2025 (SEVERAL DAYS POST-PROCEDURE), THE PATIENT HAD HYPOTENSION AND MEDICATION WAS PROVIDED WITH HIGH-DOSE CATECHOLAMINES. IT WAS REPORTED THAT THE PATIENT HAD EXPIRED DUE TO HEMODYNAMIC INSTABILITY AND MULTI-COMORBIDITIES THAT INCLUDED MR, TRICUSPID REGURGITATION (TR), POOR LEFT VENTRICLE EJECTION FRACTION (LVEF), AND POOR LUNG CONDITION. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY DID NOT INDICATE A LOT-SPECIFIC QUALITY ISSUE. THE INVESTIGATION DETERMINED THE REPORTED POOR IMAGE RESOLUTION AND SLDA APPEAR TO BE RELATED TO PATIENT MORPHOLOGY/PATHOLOGY. THE REPORTED PATIENT EFFECT OF TISSUE INJURY APPEARS TO BE RELATED TO THE SLDA. DEATH APPEARS TO BE DUE TO THE UNDERLYING CLINICAL STATE, COMORBIDITIES, AND THE ACUTE NATURE OF THE CLINICAL PRESENTATION. TISSUE INJURY AND DEATH ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND MEDICATION REQUIRED WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES AS ANOTHER CLIP WAS PLACED TO STABILIZE THE SLDA AND CATECHOLAMINE WAS PROVIDED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THIS EVENT WAS FURTHER REVIEWED BY AN ABBOTT MEDICAL AFFAIRS DIRECTOR AND THE REVIEWER STATED ¿NARRATIVE REVIEWED. NO IMAGING AVAILABLE. THE PRESENTATION WAS ACUTE DUE TO A FLAIL MITRAL VALVE LEAFLET. AN URGENT MITRA CLIP PROCEDURE WAS PERFORMED AND AN SLDA WAS REPORTED IN THE FLAIL ANATOMY. THE MITRAL REGURGITATION SEVERITY PRIOR TO PROCEDURE WAS STATED AS 3+ AND WAS NOT ABLE TO BE IMPROVED BY THE MITRA CLIP PROCEDURE. THE PATIENT WAS REPORTED TO HAVE SIGNIFICANT COMORBIDITIES INCLUDING SEVERE LV IMPAIRMENT AND THE PATIENT PASSED AWAY AN ESTIMATED 5 DAYS POST MITRA CLIP PROCEDURE. THE CAUSE OF DEATH APPEARS TO BE DUE TO THE UNDERLYING CLINICAL STATE, COMORBIDITIES, AND THE ACUTE NATURE OF THE CLINICAL PRESENTATION. THE SLDA IS LIKELY RELATED TO THE FLAIL ANATOMY AND NOT PROCEDURAL OR DEVICE RELATED.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21300223&pc=MNS | 2518422-2025-101640 | 2025-01-23 05:00 | Death | RESPIRONICS, INC. | 2025-02-03 05:00 | MNS | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21476724&pc=PNJ | 9612164-2025-01045 | 2024-04-03 04:00 | Death | MEDTRONIC IRELAND | 2025-02-27 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SU RVEILLANCE REGISTRY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENTS DEATH WAS NOT RELATED TO THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21476954&pc=DSQ | 2916596-2025-01109 | 2025-01-13 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fatigue; Syncope/Fainting; Tricuspid Valve Insufficiency/ Regurgitation; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS INITIALLY ADMITTED FOR NOROVIRUS AND DIARRHEA. THE PATIENT WAS FOUND TO HAVE SEVERE RIGHT VENTRICLE (RV) DYSFUNCTION WITH WIDE OPEN TRICUSPID REGURGITATION (TR). THE PATIENT WAS KNOWN TO HAVE HAD MILD RV DYSFUNCTION PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT, BUT THE PATIENT WAS DENIED TRANSPLANT DUE TO CONTINUAL POSITIVE TETRAHYDROCANNABINOL (THC) RESULTS ON DRUG SCREENS WITH INTERMITTENT NON-COMPLIANCE. THE TR WAS NEW AS OF THIS VISIT. THE PATIENT WAS NOTED AS HAVING NORMAL FLOWS AT THE TIME, BUT LOW PULSATILITY INDEX (PI). THE PATIENT WAS ALSO KNOWN TO HAVE AN EXTERNAL OUTFLOW GRAFT OBSTRUCTION (EOGO), IDENTIFIED DURING A HOSPITAL STAY IN (B)(6)2024 (MFR# 2916596-2024-06330), AS WELL AS A KNOWN CHRONIC DRIVELINE INFECTION (MFR# 2916596-2024-06752), AND WAS FOUND TO HAVE A 19TH POSITIVE CULTURE WITH STREPTOCOCCUS ANGINOSUS (S. ANGINOSUS)/CANDIDA ALBICANS (C. ALBICANS)/LACTOBACILLUS/PEPTONIPHILUS ASACCHAROLYTICUS (P. ASACCHAROLYTICUS). THE PATIENT WAS LATER NOTED TO HAVE INCREASED LOW FLOW ALARMS WITH LOW PI AND SYNCOPAL EPISODES AND WAS TRANSFERRED TO ANOTHER HOSPITAL, AT WHICH THEY UNDERWENT A RIGHT HEART CATHETERIZATION (RHC) WITH RAMP. THE PATIENT WAS PLACED ON INTRAVENOUS (IV) INOTROPES (DOBUTAMINE), AND WHILE THE RV REMAINED THE SAME THE TR IMPROVED FROM SEVERE TO MODERATE. PALLIATIVE MEDICINE SAW THE PATIENT MULTIPLE TIMES IN THE HOSPITAL, AND THE PATIENT WAS EVENTUALLY DISCHARGED FOR HOME HOSPICE ON (B)(6) 2025 WITH IV DOBUTAMINE. THE PATIENT STATED THEY FELT RELATIVELY NORMAL FOR AROUND A WEEK. THE PATIENT LATER PULLED OUT THEIR PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) LINE AND HAD TO RETURN TO THE HOSPITAL TO HAVE IT REPLACED. THE PATIENT RETURNED TO THE EMERGENCY ROOM (ER) AGAIN ON (B)(6) 2025 WITH WORSENING SHORTNESS OF BREATH, WEAKNESS/FATIGUE, BOTH SIGNS OF WORSENING AND END STAGE RV FAILURE. THE PATIENT PASSED AWAY FROM RV FAILURE ON (B)(6) 2025. THE PUMP WAS EXPLANTED AND WOULD BE RETURNED. THE RV FAILURE AND DEATH WERE NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE REVEALED EVENTS CONSISTENT WITH THE PATIENT¿S OUTCOME. THE LOG FILES APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: E3: OCCUPATION: CORRECTED. H5: LABELED FOR SINGLE USE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: THE EVALUATION OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), CONFIRMED A THROMBUS FORMATION IN THE OUTFLOW GRAFT AND EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO). REVIEW OF THE SUBMITTED IMAGE CONFIRMED OBSTRUCTION TO THE OUTFLOW GRAFT CANNULA. ADDITIONALLY, A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LVAS AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED DURING THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 3¿ FROM THE PUMP HEADER. THE DISTAL PORTION OF THE PUMP CABLE AND THE MODULAR CABLE WERE NOT RETURNED. THE SEALED OUTFLOW GRAFT WAS RETURNED NOT ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE BEND RELIEF ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS NOT RETURNED. APPROXIMATELY 3¿ OF THE OUTFLOW GRAFT WAS RETURNED WITH THE OUTFLOW GRAFT BEND RELIEF PROPERLY ENGAGED TO THE GRAFT ATTACHMENT HARDWARE. UPON REMOVAL OF THE OUTFLOW GRAFT BEND RELIEF, A LENGTHWISE CREASE WAS OBSERVED ALONG THE PROXIMAL PORTION OF THE GRAFT MATERIAL. AN ACCUMULATION OF TISSUE ON THE EXTERIOR OF THE GRAFT WAS OBSERVED TO HAVE FILLED IN AND FORMED OVER THE CREASE, SUGGESTING THAT THE GRAFT HAD BEEN DISTORTED FOR AN UNDETERMINED AMOUNT OF TIME WHILE (B)(6) WAS SUPPORTING THE PATIENT. THESE FINDINGS ARE CONSISTENT WITH EXTRINSIC OUTFLOW GRAFT OBSTRUCTION DURING SUPPORT. THE LUMEN OF THE SEALED OUTFLOW GRAFT REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS. EXAMINATION OF THE LUMEN OF THE SEALED OUTFLOW GRAFT REVEALED A RED, TISSUE-LIKE THROMBUS. THE THROMBUS OBSTRUCTED A SIGNIFICANT PORTION OF THE OUTFLOW GRAFT ATTACHMENT. ONCE REMOVED FROM THE OUTFLOW GRAFT HARDWARE, THE THROMBUS APPEARED TO BE FIRM AND WELL-STRUCTURED AND RETAINED ITS SHAPE. A SPECIFIC ORIGIN AND LENGTH OF TIME THAT THE THROMBUS WAS WITHIN THE OUTFLOW GRAFT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED OR ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE REVEALED EVENTS CONSISTENT WITH THE PATIENT¿S TRANSPLANT SURGERY. THE LOG FILES APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, DOCUMENT (B)(4), REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THIS IFU LISTS RIGHT HEART FAILURE, DRIVELINE INFECTION, PUMP THROMBOSIS, AND DEATH AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 1 ALSO ADDRESSES ALL PUMP PARAMETERS. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES", OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP", INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, EXPLAINS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. THIS SECTION ALSO STATES THAT IF THE AORTIC INSUFFICIENCY IS NOT ADDRESSED, THE DEVICE WILL NOT BE ABLE TO PROVIDE THE INTENDED FLOW. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21477749&pc=LJT | 3006260740-2025-01062 | 2025-02-20 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-02-27 05:00 | LJT | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Thrombosis/Thrombus; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT ABOUT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT WAS ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION AND THROMBOSIS. REPORTEDLY, THE PATIENT WAS EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION AND THROMBOSIS. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR INFECTION, THROMBOSIS, AND DEATH AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FURTHERMORE, THE CLINICAL CONDITIONS ALLEGED IN THE COMPLAINT CANNOT BE CONFIRMED. A DEFINITIVE ROOT CAUSE FOR THE REPORTED INFECTION. THROMBOSIS, AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. IT IS UNKNOWN WHETHER PATIENT AND/OR PROCEDURAL ISSUES CONTRIBUTED TO THE REPORTED EVENT. AT THIS TIME, A RELATIONSHIP BETWEEN THE PORT AND THE REPORTED INFECTION, THROMBOSIS, AND DEATH CANNOT BE CONFIRMED. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21477756&pc=LJT | 3006260740-2025-01063 | 2025-02-20 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-02-27 05:00 | LJT | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K122899 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT ABOUT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT WAS ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT WAS EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: MANUFACTURING REVIEW: A LOT HISTORY REVIEW COULD NOT BE PERFORMED AS THE LOT NUMBER WAS NOT PROVIDED. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INFECTION AND DEATH AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED INFECTION AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. IT IS UNKNOWN WHETHER PATIENT AND/OR PROCEDURAL ISSUES CONTRIBUTED TO THE REPORTED EVENTS. AT THIS TIME, A RELATIONSHIP BETWEEN THE PORT AND THESE EVENTS CANNOT BE CONFIRMED. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H0: G3. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21477943&pc=BZD | 2518422-2025-102408 | 2025-02-11 05:00 | Death | RESPIRONICS, INC. | 2025-02-27 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT USING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S SPOUSE CONTACTED US WITH CONCERNS ABOUT POTENTIAL CHARGES FOR A REPLACEMENT DEVICE. THE RESPIRATORY THERAPIST CLARIFIED THAT THE DEVICE WAS PROVIDED AT NO COST, AND THE SPOUSE IS WELCOME TO KEEP IT OR DONATE IT, AS IT IS HERS TO RETAIN. NO FURTHER DETAILS WERE PROVIDED AND THERE WAS NO ALLEGATION THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478505&pc=OZD | 1220648-2025-26375 | 2025-02-01 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A/AN (UNKNOWN GENDER, RACE AND ETHNICITY) PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND AFTER THE IMPELLA CP WAS PLACED, THE PATIENT EXPERIENCED BLEEDING. THE PATIENT WAS TAKEN TO THE OPERATING ROOM AND LEFT VENTRICLE PERFORATION WAS FOUND, REQUIRING INTERVENTION. THE PATIENT SUBSEQUENTLY EXPIRED. THE PHYSICIAN SUBSEQUENTLY COMMENTED, THE PERFORATION HAD NOTHING TO DO WITH THE IMPELLA. THE BLEEDING MUST HAVE OCCURRED BEFORE AND IT DID COME FROM THE RIGHT HEART. NO FURTHER DETAILS WERE MADE AVAILABLE AFTER THE PHYSICIAN'S COMMENT AND CONSIDERING THIS, THERE IS STILL NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR EVALUATION. IF NEW INFORMATION OR THE DEVICE IS RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478548&pc=OZD | 1220648-2025-26376 | 2025-01-28 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA 5.5 | Deformation Due to Compressive Stress | Atrial Fibrillation; Tachycardia; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED DELIVERY DIFFICULTIES WITH THE GUIDEWIRE. THE PATIENT WOULD ALSO EXPERIENCE ARRHYTHMIC EVENTS AND SUBSEQUENTLY EXPIRE. THE GUIDEWIRE BENT DURING DELIVERY OF THE IMPELLA 5.5 PUMP. A SECOND GUIDEWIRE WAS OBTAINED, AND IT BENT AS WELL. IT WAS NOTED THAT THE PATIENT WAS OBESE WITH A VERY THICK CHEST THAT MAY HAVE COMPROMISED THE ANATOMY/ANGLES BEING TRAVERSED. UPON USE OF A THIRD GUIDEWIRE, THE IMPLANT WAS SUCCESSFUL, AND THE PATIENT WAS ABLE TO RECEIVE IMPELLA MCS. AFTER PLACEMENT OF THE IMPELLA 5.5 DEVICE, THE PATIENT WAS IMPLANTED WITH A COMPETITOR'S RIGHT-SIDED DEVICE (PROTEK DUO), WHICH WAS ADVANCED TO THE PULMONARY ARTERY TO TREAT RIGHT VENTRICULAR FAILURE. FOLLOWING THE PATIENT EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE AND WAS CARDIOVERTED TWICE. THE PATIENT WAS SENT TO THE UNIT FOR CONTINUED CARE. IT WAS NOTED THAT AT SOME POINT THE PATIENT WAS PLACED ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO) AS WELL. WHILE ON SUPPORT APPROXIMATELY FIVE DAYS LATER, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AND REQUIRED SHOCKS. FOLLOWING THE PATIENT WOULD EXPERIENCE INTERMITTENT BOUTS OF VENTRICULAR TACHYCARDIA EVEN RIGHT BEFORE HIS DEMISE. CARE WAS EVENTUALLY WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT HAD MULTIPLE DEVICES, LEFT SIDE IMPELLA, RIGHT SIDE (PROTEK DUO) AND ECMO, AND GIVEN THE ARRHYTHMIC EVENTS IT IS UNCLEAR WHICH DEVICE MAY HAVE POSSIBLY CAUSED OR CONTRIBUTED TO THE EVENT AND ITS OUTCOME. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA 5.5 PUMP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME.Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED DELIVERY DIFFICULTIES WITH THE GUIDEWIRE. THE GUIDEWIRE BENT DURING DELIVERY OF THE IMPELLA 5.5 PUMP. A SECOND GUIDEWIRE WAS OBTAINED, AND IT BENT AS WELL. IT WAS NOTED THAT THE PATIENT WAS OBESE WITH A VERY THICK CHEST THAT MAY HAVE COMPROMISED THE ANATOMY/ANGLES BEING TRAVERSED. UPON USE OF A THIRD GUIDEWIRE, THE IMPLANT WAS SUCCESSFUL, AND THE PATIENT WAS ABLE TO RECEIVE IMPELLA MCS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: ADDITIONAL DETAILS ON THE COMPLAINT WERE REVIEWED BY MEDICAL SAFETY AND UPDATES TO THE TYPE OF REPORTABLE EVENT AND CORRESPONDING FIELDS WERE REQUIRED. THESE UPDATES HAVE BEEN MADE TO THE FOLLOWING SECTIONS: B1, B2, B5, H1 AND H6 (HEALTH EFFECT AND MEDICAL DEVICE PROBLEM CODES) ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478610&pc=GDW | 3005075853-2025-01540 | 2023-10-17 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-27 05:00 | GDW | ECHELON ENDOPATH STAPLER | Patient Device Interaction Problem | Failure to Anastomose | K051002 |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: LONG-TERM OUTCOMES OF LAPAROSCOPIC SLEEVE GASTRECTOMY AS A REVISIONAL PROCEDURE FOLLOWING ADJUSTABLE GASTRIC BANDING: VARIATIONS IN OUTCOMES BASED ON INDICATION AUTHORS: ANAGI WICKREMASINGHE, YIT LEANG, YAZMIN JOHARI, PREM CHANA, MEGAN ALDERUCCIO, KALAI SHAW, CHERYL LAURIE, PETER NOTTLE, WENDY BROWN, PAUL BURTON. CITATION: OBESITY SURGERY (2023) 33:3722-3739. HTTPS:/ /DOI.ORG/ 10. 1007 /S 11695-023-06886-8. THE AIM OF THIS RETROSPECTIVE ANALYSIS OF PROSPECTIVELY COLLECTED DATA COLLECTED FROM 2006 TO 2021 WAS TO COMPREHENSIVELY DETERMINE THE LONG-TERM OUTCOMES OF SLEEVE GASTRECTOMY AS A REVISIONAL PROCEDURE AFTER LAPAROSCOPIC ADJUSTABLE GASTRIC BANDING (LAGB) ACROSS A RANGE OF MEASURES AND DETERMINE PREDICTORS OF OUTCOMES. A TOTAL OF 1800 PATIENTS WERE INCLUDED IN THE STUDY. THE REVISION LAPAROSCOPIC SLEEVE GASTRECTOMY (RLSG) GROUP CONSISTED OF 600 PATIENTS (87% WERE FEMALES; MEAN AGE WAS 45±19.4 YEARS), AND THE CONTROL GROUP CONSISTED OF 1200 PRIMARY LSG (PLSG) PATIENTS (78.8% WERE FEMALES; MEAN AGE WAS 44.6±16.6 YEARS). THE SLEEVE GASTRECTOMY TECHNIQUE WAS SIMILAR TO THAT PERFORMED IN PRIMARY LSG. THE TWO MAJOR TYPES OF STAPLERS USED WERE ETHICON ECHELON¿ STAPLER-ECHELON FLEX¿ STAPLER WITH GST RELOADS (ETHICON ENDO-SURGERY, INC. , CINCINNATI, OH, USA) AND COVIDIEN ENDO GIA ¿/SIGNIA¿ WITH TRI-STAPLE¿ RELOADS TECHNOLOGY (COVIDIEN/MEDTRONIC, MINNEAPOLIS, MN, USA). REPORTED COMPLICATIONS INCLUDE MORTALITY/DEATH DUE TO SLEEVE LEAK (N=?), STAPLE LINE LEAKS (N=?), STENOSIS (N=?), FISTULA (N=?), THROMBOSIS (N=?), WOUND PROBLEMS (N=?), MAJOR BLEED (N=?), SEPSIS (N=?), REGURGITATION (N=?), FOOD INTOLERANCE (N=?), AND DYSPHAGIA (N=?). IN CONCLUSION, RLSG PROVIDES LONG-TERM WEIGHT LOSS, ALTHOUGH PERI-OPERATIVE COMPLICATIONS ARE SIGNIFICANTLY ELEVATED COMPARED TO PLSG. LONGER-TERM RE-OPERATION RATES ARE ELEVATED COMPARED TO PLSG. FOUR VARIABLES PREDICTED WORSE OUTCOMES: ERODED BAND, MULTIPLE PRIOR BANDS, SEVERE OESOPHAGEAL DYSMOTILITY AND ELEVATED BASELINE WEIGHT. Manufacturer Narrative: (B)(4). DATE SENT: 2/27/2025, D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT/BATCH NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: DOES THE AUTHOR/SURGEON BELIEVE THAT THE ETHICON DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATIONS MENTIONED IN THE ARTICLE? IF YES, PLEASE EXPLAIN. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478617&pc=GDW | 3005075853-2025-01541 | 2023-10-17 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-27 05:00 | GDW | UNKNOWN GST RELOAD | Patient Device Interaction Problem | Failure to Anastomose | K110385 |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: LONG-TERM OUTCOMES OF LAPAROSCOPIC SLEEVE GASTRECTOMY AS A REVISIONAL PROCEDURE FOLLOWING ADJUSTABLE GASTRIC BANDING: VARIATIONS IN OUTCOMES BASED ON INDICATION AUTHORS: ANAGI WICKREMASINGHE, YIT LEANG, YAZMIN JOHARI, PREM CHANA, MEGAN ALDERUCCIO, KALAI SHAW, CHERYL LAURIE, PETER NOTTLE, WENDY BROWN, PAUL BURTON. CITATION: OBESITY SURGERY (2023) 33:3722-3739. HTTPS:/ /DOI.ORG/ 10. 1007 /S 11695-023-06886-8 . THE AIM OF THIS RETROSPECTIVE ANALYSIS OF PROSPECTIVELY COLLECTED DATA COLLECTED FROM 2006 TO 2021 WAS TO COMPREHENSIVELY DETERMINE THE LONG-TERM OUTCOMES OF SLEEVE GASTRECTOMY AS A REVISIONAL PROCEDURE AFTER LAPAROSCOPIC ADJUSTABLE GASTRIC BANDING (LAGB) ACROSS A RANGE OF MEASURES AND DETERMINE PREDICTORS OF OUTCOMES. A TOTAL OF 1800 PATIENTS WERE INCLUDED IN THE STUDY. THE REVISION LAPAROSCOPIC SLEEVE GASTRECTOMY (RLSG) GROUP CONSISTED OF 600 PATIENTS (87% WERE FEMALES; MEAN AGE WAS 45±19.4 YEARS), AND THE CONTROL GROUP CONSISTED OF 1200 PRIMARY LSG (PLSG) PATIENTS (78.8% WERE FEMALES; MEAN AGE WAS 44.6±16.6 YEARS). THE SLEEVE GASTRECTOMY TECHNIQUE WAS SIMILAR TO THAT PERFORMED IN PRIMARY LSG. THE TWO MAJOR TYPES OF STAPLERS USED WERE ETHICON ECHELON¿ STAPLER-ECHELON FLEX¿ STAPLER WITH GST RELOADS (ETHICON ENDO-SURGERY, INC., CINCINNATI, OH, USA) AND COVIDIEN ENDO GIA ¿/SIGNIA¿ WITH TRI-STAPLE¿ RELOADS TECHNOLOGY (COVIDIEN/MEDTRONIC, MINNEAPOLIS, MN, USA). REPORTED COMPLICATIONS INCLUDE MORTALITY/DEATH DUE TO SLEEVE LEAK (N=?), STAPLE LINE LEAKS (N=?), STENOSIS (N=?), FISTULA (N=?), THROMBOSIS (N=?), WOUND PROBLEMS (N=?), MAJOR BLEED (N=?), SEPSIS (N=?), REGURGITATION (N=?), FOOD INTOLERANCE (N=?), AND DYSPHAGIA (N=?). IN CONCLUSION, RLSG PROVIDES LONG-TERM WEIGHT LOSS, ALTHOUGH PERI-OPERATIVE COMPLICATIONS ARE SIGNIFICANTLY ELEVATED COMPARED TO PLSG. LONGER-TERM RE-OPERATION RATES ARE ELEVATED COMPARED TO PLSG. FOUR VARIABLES PREDICTED WORSE OUTCOMES: ERODED BAND, MULTIPLE PRIOR BANDS, SEVERE OESOPHAGEAL DYSMOTILITY AND ELEVATED BASELINE WEIGHT. Manufacturer Narrative: (B)(4) DATE SENT: 2/27/2025 D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT/BATCH NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: DOES THE AUTHOR/SURGEON BELIEVE THAT THE ETHICON DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATIONS MENTIONED IN THE ARTICLE? IF YES, PLEASE EXPLAIN. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478630&pc=MKJ | 2112020-2025-00145 |  | Death | CARDIAC SCIENCE CORPORATION | 2025-02-27 05:00 | MKJ | POWERHEART G5 KIT,SEMI,SLOVENE-INTL ENG | Failure to Deliver Shock/Stimulation; Failure to Power Up | Insufficient Information | P160033 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE FAILED TO DISCHARGE. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION.Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THE PATIENT SUBSEQUENTLY EXPIRED. THE CUSTOMER STATED THEY DO NOT BELIEVE THAT THE PATIENT'S LIFE WOULD HAVE BEEN SAVED HAD THE DEVICE WORKED AS INTENDED AS THE PATIENT WAS ELDERLY AND IN POOR HEALTH. Manufacturer Narrative: THIS SUPPLEMENTAL MEDWATCH REPORT IS REPORTING THE EVALUATION OF THE DEVICE AND CORRECTING INFORMATION SUBMITTED ON THE INITIAL MEDWATCH REPORT. PLEASE REFERENCE SECTION B1, B2, B5 AND H6 (MEDICAL DEVICE PROBLEM CODE & HEALTH EFFECT IMPACT CODE). THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS ATTRIBUTED TO THE USER NOT RESPONDING TO THE LOW BATTERY WARNINGS. A REVIEW OF THE DEVICE LOG INDICATED THE DEVICE BEGAN WARNING OF A LOW BATTERY ON (B)(6) 2024, BUT THERE'S NO EVIDENCE THAT THE USER ADDRESSED THE ISSUE UNTIL (B)(6) 2024. THE BATTERY, INSTALLED IN 2019, WAS NEARLY FIVE YEARS OLD, EXCEEDING ITS FOUR-YEAR STANDBY LIFE. THE DEVICE FUNCTIONED CORRECTLY, PROVIDING VISUAL AND AUDIBLE LOW BATTERY ALERTS AS DESIGNED. TESTING CONFIRMED IT WOULD HAVE BEEN CAPABLE OF DELIVERING SHOCKS IF A GOOD BATTERY HAD BEEN INSTALLED. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478687&pc=OZD | 1220648-2025-26386 | 2025-01-31 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY TREATED WITH AN IMPELLA CP AND EXTRACORPOREAL MEMBRANE OXYGENATION (EMCO) AND THEN ESCALATED TO AN IMPELLA 5.5 DEVICE. AFTER APPROXIMATELY 11 DAYS ON SUPPORT, THE PATIENT EXPERIENCED A RIGHT-SIDED STROKE WHICH WAS CONFIRMED ON COMPUTED TOMOGRAPHY SCAN. THE PATIENT HAD VARIOUS BOUTS OF BLEEDING AND ANTICOAGULATION WAS HELD. THE PLAN WAS TO BRING THE PATIENT BACK TO THE OPERATING ROOM LATER IN THE DAY FOR REMOVAL OF VENOARTERIAL ECMO. HOWEVER, THE FAMILY DECIDED TO PULL SUPPORT OVER THE WEEKEND, AND THE PATIENT EXPIRED TWO DAYS LATER ON A SUNDAY. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478870&pc=OZD | 1220648-2025-26390 | 2025-02-03 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK POST CORONARY ARTERY BYPASS GRAFT SURGERY ONE WEEK PRIOR WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED A LEFT VENTRICLE PERFORATION AND WOULD EXPIRE. FURTHER DETAILS NOTED THE GRAFT FOR DIRECT IMPELLA 5.5 IMPLANT WAS TIED TOO CLOSE TO THE AORTIC VALVE AT 5 CENTIMETERS. RELOCATING THE GRAFT WAS RECOMMENDED. THE PHYSICIAN WANTED TO MAKE THE CURRENT LOCATION WORK AND WAS ADVISED THIS WAS NOT RECOMMENDED AND IF THE DECISION WAS TO MOVE FORWARD WITH THE CURRENT LOCATION, THE IMPELLA DEVICE MUST BE CLEAR OF THE GRAFT. THE PATIENT WAS ABLE TO GO ON FULL IMPELLA SUPPORT AND COME OFF BYPASS. HOWEVER, UPON CLOSING THE STERNUM AND/OR TRIMMING GRAFT, A DECLINE IN THE PATIENT AND FLOWS WAS NOTED. A LEFT VENTRICLE PERFORATION WAS DISCOVERED. THIS LED TO CODING AND MANY DIRECT SHOCKS TO HEART. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE VENTRICLE PERFORATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS NOT DETERMINED AS INSUFFICIENT INFORMATION ON THE CAUSE AND TIMING OF PERFORATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478915&pc=LYJ | 1644487-2025-00218 | 2023-03-16 04:00 | Death | LIVANOVA USA, INC. | 2025-02-27 05:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: THE PATIENT'S MOTHER REPORTS THAT THE PATIENT HAD THE WORST SEIZURE SHE EVER HAD AND STOPPED BREATHING BUT WAS REVIVED. EIGHT MONTHS LATER SHE WAS DEAD FROM YET ANOTHER HORRIBLE SEIZURE. THE PATIENT STARTED TO EXPERIENCE AN INCREASE IN SEIZURES WITHIN DAYS OF UNDERGOING A BATTERY REPLACEMENT. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: ADDITIONAL INFORMATION RECEIVED NOTING THAT THE PHYSICIAN IS UNCERTAIN OF THE CAUSE FOR THE INCREASED SEIZURES. BUT IT WAS NOTED TO BE MOST LIKELY SECONDARY TO EPILEPSY WORSENING AND NOT DUE TO THE VNS. THE INCREASE WAS BELOW PRE-VNS BASELINE LEVELS; OVERALL DECREASE AND THE PATIENT WAS STARTED ON LACOSAMIDE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478948&pc=DYE | 2025587-2025-01509 | 2025-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-27 05:00 | DYE | MOSAIC MITRAL BIOPROSTHETIC HEART VALVE | Insufficient Information | Insufficient Information | P990064 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT ON THE SAME DAY AS IMPLANT OF THIS MOSAIC MITRAL VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE MOSAIC MITRAL VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21479010&pc=OZD | 1220648-2025-26391 | 2025-02-03 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. AFTER TWO DAYS OF SUPPORT, THE PATIENT EXPERIENCED A BRAIN BLEED. HEPARIN WAS IN THE PUMP PURGE. SUBSEQUENTLY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF STROKE/CEREBROVASCULAR ACCIDENT (CVA) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS THIS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21479121&pc=DSQ | 2916596-2025-01308 | 2024-10-01 04:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE ARTICLE ¿INFLUENCE OF THE OUTFLOW GRAFT ANGULAR POSITION ON THE OUTCOMES IN PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE¿ THAT THE ANGULAR POSITION OF THE OUTFLOW GRAFT (OFG) MAY BE ASSOCIATED WITH ADVERSE OUTCOMES, SUCH AS INCREASED CEREBRAL THROMBOEMBOLIC EVENTS, GASTROINTESTINAL (GI) BLEEDING, AORTIC INSUFFICIENCY/REGURGITATION, OR DEATH. THIS SINGLE-CENTER RETROSPECTIVE STUDY AIMED TO EXPLORE THE POTENTIAL IMPACT OF THE ANGULAR POSITION OF THE OFG ON ADVERSE EVENTS IN PEOPLE WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD). HOSPITAL RECORDS WERE REVIEWED RETROSPECTIVELY FOR ALL PATIENTS WHO UNDERWENT A HEARTMATE 3 (HM3) IMPLANTATION BETWEEN JANUARY 2016 AND DECEMBER 2021. TO ALLOW FOR AN ADEQUATE FOLLOW UP PERIOD, PATIENT DATA WERE COLLECTED UNTIL DECEMBER 2022. PATIENTS WERE ELIGIBLE IF THEY WERE 18 YEARS OR OLDER AND UNDERWENT A CONTRAST-ENHANCED COMPUTED TOMOGRAPHY (CT) SCAN OF THE THORAX, INCLUDING THE AORTIC ARCH BRANCHES. AMONG 59 PATIENTS, THE MEDIAN FOLLOW-UP WAS 25 MONTHS, WITH A MEDIAN AGE OF 57 YEARS, AND 67.8% WERE MEN. IT WAS NOTED THAT OUT OF THE INCLUDED PATIENTS, 7 EXPERIENCED A CEREBROVASCULAR ACCIDENT (CVA), 12 EXPERIENCED GI BLEEDING, AND 12 PATIENTS PASSED AWAY. OUT OF THE EXCLUDED PATIENTS, 6 EXPERIENCED A CVA, 5 EXPERIENCED GI BLEEDING, AND 20 PATIENTS PASSED AWAY. BASED ON THE FINDINGS, THE OPTIMAL CUTOFF POINT FOR CVA WAS 107° FOR THE VERTICAL ANGLE, 20° FOR THE HORIZONTAL ANGLE, AND 0.49 FOR THE RELATIVE DISTANCE. THE RISK OF EXPERIENCING A CVA DURING FOLLOW UP IS INCREASED WHEN THE VERTICAL ANGLE IS =107° (P = 0.022). THE OPTIMAL CUT OFF POINT FOR SURVIVAL WAS 98° FOR THE VERTICAL ANGLE, 10° FOR THE HORIZONTAL ANGLE, AND 0.50 FOR THE RELATIVE DISTANCE. WHEN THE HORIZONTAL ANGLE IS =10°, THE RISK OF SURVIVAL IS DECREASED DURING FOLLOW-UP MORE (P = 0.004). THE OPTIMAL CUTOFF POINT FOR GI BLEEDING WAS 106° FOR THE VERTICAL ANGLE, 16° FOR THE HORIZONTAL ANGLE, AND 0.52 FOR THE RELATIVE DISTANCE. SUBSEQUENTLY, A SIGNIFICANTLY INCREASED RISK IN THE OCCURRENCE OF GI BLEEDING DURING FOLLOW-UP WAS FOUND WHEN THE VERTICAL ANGLE WAS =106° (P = 0.049). NO DIFFERENCES BETWEEN CVA, SURVIVAL, OR GI BLEEDING WERE FOUND WHEN DIVIDING THE COHORT ACCORDING TO THE HORIZONTAL ANGLE OR RELATIVE DISTANCE. NO SIGNIFICANT DIFFERENCES WERE FOUND BETWEEN THE VERTICAL ANGLE AND AORTIC VALVE REGURGITATION OR SURVIVAL. ONE PATIENT WHO DEVELOPED A CVA DURING FOLLOW-UP HAD A HISTORY OF EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT. NO PATIENTS WHO DEVELOPED A CVA DURING FOLLOW-UP HAD A HISTORY OF CVA OR ATRIAL FIBRILLATION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. ZIJDERHAND, C. F., PEEK, J. J., SJATSKIG, J., MANINTVELD, O. C., BEKKERS, J. A., BOGERS, A. J. J. C., CALISKAN, K. (2024) INFLUENCE OF THE OUTFLOW GRAFT ANGULAR POSITION ON THE OUTCOMES IN PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE. ASAIO JOURNAL, 70(10), 861-867. HTTP://DX.DOI.ORG/10.1097/MAT.0000000000002189. THORAXCENTER, DEPARTMENT OF CARDIOTHORACIC SURGERY, ERASMUS MC, UNIVERSITY MEDICAL CENTER ROTTERDAM, ROTTERDAM, NETHERLANDS THORAXCENTER, DEPARTMENT OF CARDIOLOGY, ERASMUS MC, UNIVERSITY MEDICAL CENTER ROTTERDAM, ROTTERDAM, NETHERLANDS. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21479288&pc=MVK | 3008642652-2025-02121 | 2025-01-26 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 17:23:02 ON (B)(6) 2025. AT 07:30:53 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION/TACTILE ARTIFACT. AT 07:31:33, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 30 BPM WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 07:48:34 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21479327&pc=DSQ | 2916596-2025-01167 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21479469&pc=LJT | 1317056-2025-00045 | 2022-06-10 04:00 | Death | ANGIODYNAMICS, INC | 2025-02-27 05:00 | LJT | SMART PORT CT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Septic Shock | K062414 |  | 1 | Event Description: ON OR ABOUT MARCH 21, 2021, DECEDENT UNDERWENT PLACEMENT OF AN ANGIODYNAMICS SMARTPORT PRODUCT, MODEL NUMBER H787CT75STSDVI1, LOT NUMBER 5708035. THE DEVICE WAS IMPLANTED BY DR. (B)(6), AT (B)(6), FOR THE PURPOSE OF ONGOING CHEMOTHERAPY. ON OR ABOUT (B)(6) 2022, DECEDENT PRESENTED TO (B)(6) MEDICAL CENTER, ARKANSAS, WITH COMPLAINTS OF WEAKNESS. DECEDENT'S MEDICAL TEAM DETERMINED THAT SHE HAD DEVELOPED SEPTIC SHOCK. ON OR ABOUT (B)(6) 2022, DECEDENT WAS TRANSFERRED VIA AIR AMBULANCE TO (B)(6) MEDICAL CENTER- LITTLE ROCK IN LITTLE ROCK, ARKANSAS. DURING TRANSPORT, DECEDENT WAS INTUBATED. ON OR ABOUT (B)(6) 2022, DECEDENT WAS PRONOUNCED DEAD. Manufacturer Narrative: THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4). Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT SEPSIS INFECTION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). NO PORT DEVICE WAS RETURNED FOR EVALUATION SINCE THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION, DAMAGE OR PERFORMANCE ISSUE DURING USE. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE INDICATED PACKAGING LOT REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. STERILIZATION LOAD RELEASE RECORDS WERE REVIEWED FOR THE PACKAGING LOT IN QUESTION; NO ISSUES WERE OBSERVED. THE PORT WAS IMPLANTED INTO THE PATIENT 81 DAYS PRIOR TO THE PATIENT'S INFECTION. THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING USE. IF THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE IN QUESTION WAS TO CONTRIBUTE TO AN INFECTION IN THE PATIENT IT WOULD HAVE LIKELY OCCURRED SHORTLY AFTER THE IMPLANT DATE AND NOT 81 DAYS LATER. DEVICE DIRECTIONS FOR USE CAUTIONS THAT EACH ACCESS OF THE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE, FOLLOWING THE INSTITUTIONS UNIVERSAL PRECAUTIONS. THERE IS NO INDICATION FROM THE REPORTED COMPLAINT THAT THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE COULD HAVE CONTRIBUTED TO THE PATIENT'S INFECTION. INFECTION IS CAUTIONED IN THE DEVICE DFU AS A POTENTIAL COMPLICATION OF PORT SYSTEM USE. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED IN THE SMARTPORT DEVICE CONTAINS THE FOLLOWING DIRECTIONS AND PRECAUTIONS: CONTRAINDICATIONS: CATHETER INSERTION IN THE SUBCLAVIAN VEIN MEDIAL TO THE BORDER OF THE FIRST RIB, AN AREA WHICH IS ASSOCIATED WITH HIGHER RATES FOR PINCH-OFF. WARNINGS: THE DEVICE IS TO BE IMPLANTED, USED, MAINTAINED, AND REMOVED IN STRICT ACCORDANCE WITH INSTITUTIONAL AND OR CENTERS FOR DISEASE CONTROL (CDC) GUIDELINES OR POLICIES. REUSE OF SINGLE-USE DEVICES CREATES A POTENTIAL RISK OF PATIENT OR USER INFECTIONS. CONTAMINATION OF THE DEVICE MAY LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. DO NOT USE SYRINGES SMALLER THAN 10 ML SYRINGE WHEN ACCESSING THE PORT AS SYSTEM DAMAGE CAN OCCUR. FLUSHING OCCLUDED CATHETERS WITH SMALL SYRINGES CAN CREATE EXCESSIVE PRESSURES WITHIN THE PORT SYSTEM. REUSE OF SINGLE-USE DEVICES CREATES A POTENTIAL RISK OF PATIENT OR USER INFECTIONS. CONTAMINATION OF THE DEVICE MAY LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. ABSENCE OF A BLOOD RETURN OR A POOR BLOOD RETURN CAN BE A SIGN OF A POTENTIAL COMPLICATION SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, FIBRIN FORMATION, THROMBOSIS OR MALPOSITION. THIS SHOULD BE EVALUATED PRIOR TO DEVICE USAGE. A BLOOD RETURN SHOULD BE PRESENT PRIOR TO USAGE OF DEVICE FOR ANY THERAPY OR TESTING. DO NOT ATTEMPT TO MEASURE THE PATIENT'S BLOOD PRESSURE ON THE ARM IN WHICH A PERIPHERAL SYSTEM IS LOCATED, SINCE CATHETER OCCLUSION OR OTHER DAMAGE TO THE CATHETER COULD OCCUR. IF THE PATIENT COMPLAINS OF PAIN, OR THERE IS SWELLING WHEN THE DEVICE IS FLUSHED OR WHEN MEDICATION OR CONTRAST MEDIA IS ADMINISTERED, EVALUATE THE DEVICE FOR INFILTRATION, PROPER NEEDLE PLACEMENT, AND POTENTIAL COMPLICATIONS SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, THROMBOSIS OR MALPOSITION. FAILURE TO ASSESS THESE COMPLAINTS OR OBSERVATIONS CAN LEAD TO DEVICE FAILURE. PRECAUTIONS: CAREFULLY READ AND FOLLOW ALL INSTRUCTIONS PRIOR TO USE. STRICT ASEPTIC TECHNIQUE IS OF PARAMOUNT IMPORTANCE WHEN IMPLANTING ANY DEVICE. POTENTIAL COMPLICATIONS: USE OF AN ANGIODYNAMICS PORT SYSTEM INVOLVES POTENTIAL RISKS NORMALLY ASSOCIATED WITH THE INSERTION OR USE OF ANY IMPLANTED DEVICE OR INDWELLING CATHETER INCLUDING BUT NOT LIMITED TO: EROSION OF VESSEL AND SKIN, IMPLANT REJECTION, INFECTION, INFLAMMATION, NECROSIS OF SCARRING OF SKIN OVER IMPLANT AREA, VESSEL TRAUMA. POST-OPERATIVE CARE: THE INCISION SITE SHOULD BE MONITORED FOR SIGNS OF INFECTION, INFLAMMATION, HEMATOMA, DEVICE ROTATION OR EROSION. ROUTINE WOUND CARE SHOULD BE GIVEN TO THESE SITES. THE SMART PORT CT IMPLANTABLE PORT MAY BE USED IMMEDIATELY AFTER VERIFICATION OF CATHETER PLACEMENT. INSTRUCT PATIENT TO AVOID ANY HEAVY EXERTION OR STRENUOUS ACTIVITY DURING THE FIRST FEW DAYS AFTER SURGERY. GENERAL GUIDELINES: EACH ACCESS OF AN ANGIODYNAMICS SMART PORT CT IMPLANTABLE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE. FOLLOW INSTITUTIONAL UNIVERSAL PRECAUTIONS. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21479560&pc=NPT | 2135147-2025-00983 | 2025-02-03 05:00 | Death | ABBOTT MEDICAL | 2025-02-27 05:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | Bradycardia; Perforation; Cardiac Tamponade | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON 03 FEB. 2025, A 25MM NAVITOR VALVE (SN: (B)(6)) WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM (DS) (LOT: 10441074). DURING THE PROCEDURE, WHEN THE DS WAS INSERTED INTO THE RIGHT FEMORAL ARTERY, THE RADIOPAQUE TIP AND INNER SHAFT HAD BECOME BENT. THE SYSTEM WAS REMOVED FROM THE PATIENT AND THE NAVITOR VALVE WAS LOADED INTO A NEW SMALL FLEXNAV DS (LOT: 10458760). THE NEW SYSTEM WAS THEN INSERTED AND WHILE THE NAVITOR WAS AT 70% DEPLOYMENT, THE PATIENT EXPERIENCED SEVERE BRADYCARDIA FOLLOWED BY CARDIAC ARREST. THE VALVE WAS IMPLANTED AND RESUSCITATION MANEUVERS WERE PERFORMED RETURNING THE PATIENT TO THEIR NORMAL PACING. THE SONOGRAPHER THEN DIAGNOSED A RIGHT VENTRICULAR PERFORATION AND BLEEDING ON THE RIGHT SIDE OF THE HEART LEADING TO CARDIAC TAMPONADE. THE PATIENT EXPERIENCED ANOTHER CARDIAC ARREST AND PASSED AWAY. THE PHYSICIAN ALLEGES THE CAUSED TO BE DUE TO THE TEMPORARY PACEMAKER.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE (SN: (B)(6)) WAS SELECTED FOR IMPLANT USING A SMALL FLEXNAV DELIVERY SYSTEM (DS) (LOT: 10441074). AT THE BEGINNING OF THE PROCEDURE IT WAS NOTED THAT A TEMPORARY PACEMAKER WAS PLACED ON THE RIGHT SIDE OF THE HEART. THE PACEMAKER WAS TURNED OFF AFTER PRE-DILATATION WAS PERFORMED. DURING THE PROCEDURE, WHEN THE DS WAS INSERTED INTO THE RIGHT FEMORAL ARTERY, THE RADIOPAQUE TIP AND INNER SHAFT HAD BECOME BENT. THE SYSTEM WAS REMOVED FROM THE PATIENT AND THE NAVITOR VALVE WAS LOADED INTO A NEW SMALL FLEXNAV DS (LOT: 10458760). THE NEW SYSTEM WAS THEN INSERTED AND WHILE THE NAVITOR WAS AT 70% DEPLOYMENT, THE PATIENT EXPERIENCED SEVERE BRADYCARDIA FOLLOWED BY CARDIAC ARREST. THE VALVE WAS IMPLANTED AND RESUSCITATION MANEUVERS WERE PERFORMED RETURNING THE PATIENT TO THEIR NORMAL PACING. THE SONOGRAPHER THEN DIAGNOSED A RIGHT VENTRICULAR PERFORATION AND BLEEDING ON THE RIGHT SIDE OF THE HEART LEADING TO CARDIAC TAMPONADE. THE PATIENT EXPERIENCED ANOTHER CARDIAC ARREST AND PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU). THE PHYSICIAN ALLEGES THE CAUSED TO BE DUE TO THE TEMPORARY PACEMAKER. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF DEATH, BRADYCARDIA, PERFORATION, AND CARDIAC TAMPONADE WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE EVENT AND IMAGING WAS REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. BASED ON ALL AVAILABLE INFORMATION, CAUSES FOR THE REPORTED BRADYCARDIA, PERFORATION, AND CARDIAC TAMPONADE COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED DEATH APPEARS TO BE RELATED TO THE CARDIAC TAMPONADE. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS THE RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480095&pc=NGV | 2124215-2025-12149 | 2025-02-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-27 05:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Device Dislodged or Dislocated | Cardiac Arrest; Low Blood Pressure/ Hypotension | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 27MM WATCHMAN FLX LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. A FEW HOURS POST PROCEDURE THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT WENT INTO CARDIAC ARREST. ULTRASOUND IMAGING SHOWED THAT THE CLOSURE DEVICE HAD EMBOLIZED OUT OF THE LAA OF THE PATIENT. A SURGICAL THORACOTOMY WAS PERFORMED, AND THE CLOSURE DEVICE WAS REMOVED FROM THE PATIENT. RESUSCITATION ATTEMPTS WERE PERFORMED BUT WERE NOT SUCCESSFUL. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 27MM WATCHMAN FLX LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. A FEW HOURS POST PROCEDURE THE PATIENT'S BLOOD PRESSURE DROPPED AND THE PATIENT WENT INTO CARDIAC ARREST. ULTRASOUND IMAGING SHOWED THAT THE CLOSURE DEVICE HAD EMBOLIZED OUT OF THE LAA OF THE PATIENT. A SURGICAL THORACOTOMY WAS PERFORMED, AND THE CLOSURE DEVICE WAS REMOVED FROM THE PATIENT. RESUSCITATION ATTEMPTS WERE PERFORMED BUT WERE NOT SUCCESSFUL. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED. IT WAS FURTHER REPORTED THAT AFTER THE PATIENT WENT INTO CARDIAC ARREST, THEY RECEIVED TWO ROUNDS OF CARDIAC DEFIBRILLATION, BEDSIDE CARDIOPULMONARY RESUSCITATION (CPR), ENDOTRACHEAL INTUBATION, AND INTRAVENOUS EPINEPHRINE ADMINISTRATION. THE PATIENT'S HEART RATE AND BLOOD PRESSURE TEMPORARILY RECOVERED, BUT CONSCIOUSNESS REMAINED IMPAIRED. EMERGENCY CARDIOTHORACIC CONSULTATION WAS REQUESTED. AN EMERGENCY THORACOTOMY WAS THEN PERFORMED. FOLLOWING SURGERY, THE PATIENT REMAINED UNCONSCIOUS AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT ON VENTILATOR SUPPORT. ONE DAY POST PROCEDURE, THE PATIENT'S HEART RATE DROPPED TO 0 BPM. DESPITE AGGRESSIVE RESUSCITATION, SPONTANEOUS RHYTHM WAS NOT RESTORED, AND DEATH WAS DECLARED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480099&pc=MCX | 2124215-2025-11970 | 2025-01-31 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-27 05:00 | MCX | ROTAPRO | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A 1.50MM ROTAPRO WAS SELECTED FOR USE. A PATIENT REFUSED TO UNDERGO CORONARY ARTERY BYPASS GRAFT SURGERY (CABG), THEREFORE THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE WITH MULTIPLE NON-BOSTON SCIENTIFIC DEVICES. A BALLOON ANGIOPLASTY WAS PERFORMED FOLLOWED BY A WIRE EXCHANGE FOR ROTAWIRE DRIVE GUIDEWIRE. ROTATIONAL ATHERECTOMY WAS PERFORMED SUCCESSFULLY WITH A 1.50MM BURR, AFTER ATHERECTOMY THE ROTA BURR WAS REMOVED AND CONVERTED FOR UNREPORTED BRAND OF WORKHORSE GUIDEWIRE. A NON-BOSTON SCIENTIFIC BALLOON WAS INSERTED AND INFLATED. THE PATIENT WAS TREATED WITH 2 CYCLES OF INTRAVASCULAR LITHOTRIPSY (IVL) TREATMENT. THE PATIENT BECAME UNSTABLE AND THE ANGIOGRAM SHOWED A LEFT MAIN DISSECTION CAUSED BY THE IVL TREATMENT WITH A NON-BOSTON SCIENTIFIC BALLOON. BALLOON TAMPONADE AND A STENT WERE USED TO TREAT THE LEFT MAIN ARTERY DISSECTION. THE PATIENT HAS EVENTUALLY DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480104&pc=DSQ | 2916596-2025-01135 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAD A STROKE IN INTENSIVE CARE UNIT (ICU) POST LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT SURGERY. LOG FILES WERE FULL OF PULSATILITY INDEX (PI) EVENTS ON (B)(6) 2025. THE LOG ALSO CAPTURED INTERMITTENT 10 MOMENTARY LOW FLOW ESTIMATES WITH ELEVATED PI ON (B)(6) 2025 FROM 14:02 TO 14:13. PER HEAD COMPUTED TOMOGRAPHY (CT) ON (B)(6) 2025, A NEW LARGE SUBACUTE WAS APPEARING RIGHT MIDDLE CEREBRAL ARTERY (MCA) DISTRIBUTION INFARCT WITH MILD MASS EFFECT AND 3-MILLIMETER LEFTWARD MIDLINE SHIFT. NO HEMORRHAGE WAS NOTED. THERE WERE NO KNOWN CHANGES IN PATIENT CONDITION THAT MAY HAVE CONTRIBUTED TO STROKE. THE PATIENT WAS NOT ON ANTICOAGULATION AT THE TIME OF THE STROKE. THE PATIENT REMAINED IN ICU AT THIS TIME AND TREATMENT CONTINUED.Event Description: ADDITIONAL INFORMATION STATED THAT THE PATIENT WAS TRANSITIONED TO COMFORT CARE AFTER SUFFERING A STROKE POST IMPLANT. THEY ULTIMATELY PASSED AWAY ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS. A SPECIFIC CAUSE FOR THESE EVENTS AND A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED STROKE AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS ON (B)(6) 2025, SOME OF WHICH RESULTED IN LOW FLOW ALARMS. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE AND DEATH. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7OF THE IFU, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480108&pc=OZD | 1220648-2025-26402 | 2025-02-05 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AFTER IMPELLA CP WAS PLACED IN A 78-YEAR-OLD MALE, STAFF WERE UNABLE TO DOPPLER PULSES ON THE RIGHT LOWER EXTREMITY. A FORMAL DUPLEX STUDY WAS ORDERED OF RIGHT LOWER EXTREMITY (RLE) SHOWING NO FLOW DOWN THE LEG. AT THIS TIME, THE PATIENT HAD POOR PROGNOSIS AND AS A RESULT, THE FAMILY DECIDED TO MAKE PATIENT COMFORT CARE AND WITHDRAW SUPPORTIVE MEASURES. THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED LIMB ISCHEMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B.7 ADDED INFORMATION AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 220648-2025-26402. E.1 FAX NUMBER REPORTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 220648-2025-26402 WAS INCORRECT. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480529&pc=DSQ | 2916596-2025-01112 | 2022-03-01 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "DURABILITY OF TRICUSPID VALVE REPAIR IN PATIENTS UNDERGOING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT HEART FAILURE, RENAL FAILURE, ARRYTHMIA, VISCERAL ISCHEMIA, THROMBOSIS, GASTRO-INTESTINAL BLEEDING, AND DEATH. THIS RETROSPECTIVE STUDY EVALUATED 94 PATIENTS WHO RECEIVED CONCOMITANT TVR (LVAD+TVR). ECHOCARDIOGRAPHIC AND CLINICAL DATA WERE ASSESSED. MEDIAN FOLLOW-UP WAS 2.8 YEARS (7 DAYS¿0.6 YEARS). RESULTS WERE CORRELATED WITH CLINICAL OUTCOMES. LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION (HVAD¿¿MEDTRONIC, HEARTASSIST 5¿ MICROMED CARDIOVASCULAR, HEARTMATE II AND HEARTMATE 3¿ABBOTT) WAS PERFORMED VIA A MEDIAN STERNOTOMY IN ALL 94 PATIENTS. IN ALL PATIENTS, OUTFLOW GRAFTS WERE CONNECTED TO THE ASCENDING AORTA AND TVR WAS PERFORMED USING BAND ANNULOPLASTY ON THE BEATING HEART PRIOR TO LVAD IMPLANTATION. ALL PATIENTS WERE FOLLOWED WITH CLINIC VISITS THREE TO FIVE TIMES PER YEAR. MEDIAN FOLLOW-UP WAS 2.8 ± 2.3 YEARS AND RANGED FROM 7 DAYS TO 10.6 YEARS. PATIENTS WERE ON AVERAGE 60 YEARS OLD AND THE VAST MAJORITY WERE MALE (N = 84, 89.4%). FIFTY-SIX PATIENTS (59.6%) HAD UNDERGONE PREVIOUS PACEMAKER OR ICD IMPLANTATION AND TEN (10.6%) SUFFERED FROM COPD. THIRTY-SEVEN PATIENTS (39.4%) RECEIVED LVAD IMPLANTATION FOR ISCHEMIC CARDIOMYOPATHY AND THE REST FOR DILATED CARDIOMYOPATHY. THIRTY-FIVE PATIENTS (37.3%) PRESENTED IN A PREOPERATIVE CRITICAL STATE DEFINED BY INTERMACS LEVELS 1 AND 2 WITH TWELVE PATIENTS (12.8%) BRIDGED WITH SHORT-TERM MCS. TREATMENT STRATEGY WAS DESTINATION THERAPY IN 40 PATIENTS (42.6%), BRIDGE TO TRANSPLANTATION IN 31 (33%) AND BRIDGE TO DECISION IN 23 (24.5%). 28 PATIENTS (29.8%) PRESENTING WITH MODERATE AND 66 (70%) WITH SEVERE TR. TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) IMPLANTATION WAS NECESSARY IN 2 PATIENTS, DIALYSIS WAS NEEDED DUE TO RENAL FAILURE IN 18 PATIENTS, 7 PATIENTS HAD ARRHYTHMIA, 2 PATIENTS HAD VISCERAL ISCHEMIA, AND 5 PATIENTS HAD A PERIOPERATIVE STROKE 5 PATIENTS. 2 PATIENTS HAD PUMP THROMBOSIS, 8 PATIENTS HAD A STROKE, 9 PATIENTS HAD GASTRO-INTESTINAL BLEEDING, 2 PATIENTS HAD RIGHT HEART FAILURE WHICH WAS CHARACTERIZED BY LEG EDEMA, DYSPNEA AND LEFT VENTRICULAR ASSIST DEVICE (LVAD) LOW FLOWS ALARMS LEADING TO REHOSPITALIZATION, 12 PATIENTS WERE TRANSPLANTED, AND 36 PATIENTS PASSED AWAY. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. FÄRBER, G., SCHWAN, I., KIROV, H., ROSE, M., TKEBUCHAVA, S., SCHNEIDER, U., CALDONAZO, T., DIAB, M., & DOENST, T. (2024). DEPARTMENT OF THORACIC AND CARDIOVASCULAR SURGERY, SAARLAND UNIVERSITY MEDICAL CENTER, HOMBURG/SAAR, GERMANY; DEPARTMENT OF CARDIOTHORACIC SURGERY, JENA UNIVERSITY HOSPITAL, FRIEDRICH-SCHILLER-UNIVERSITY JENA, JENA, GERMANY; DEPARTMENT OF CARDIAC SURGERY, ROTENBURG HEART AND VASCULAR CENTRE, ROTENBURG AN DER FULDA, GERMANY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480581&pc=MFE | 2029214-2025-00528 | 2024-06-07 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-27 05:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: ENDOVASCULAR TREATMENT OF SPETZLER-MARTIN GRADE III ARTERIOVENOUS MALFORMATIONS: A SINGLE-CENTER 12 YEARS¿ EXPERIENCE STRATIFIED BY THE SPETZLER-MARTIN MODIFIED SCALE THE TIME FRAME OF THIS STUDY WAS: JANUARY 2010 TO JANUARY 2022 MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION: MEDTRONIC, MICROVENTION, BALT THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX (NONADHESIVE LIQUID EMBOLIC AGENT) DEATHS OCCURRED IN THE STUDY POPULATION: TOTAL DEATHS REPORTED: 3.4% (3 OUT OF 91 PATIENTS), PROCEDURE-RELATED MORTALITY: 3.3% (3 OUT OF 91 PATIENTS, ALL RUPTURED AVMS), UNRELATED MORTALITY: 1 PATIENT DIED A FEW MONTHS AFTER TREATMENT FOR AN UNRELATED CAUSE THE CAUSES OF DEATH WERE: THE DOCUMENT DOES NOT SPECIFY THE EXACT CAUSES OF DEATH, BUT IT MENTIONS THAT 3 PATIENTS WITH RUPTURED AVMS DIED DUE TO THE SEVERITY OF THE INITIAL PRESENTATION AND CLINICAL COMPLICATIONS (TERMINAL CLINICAL DISEASE BEFORE THE AVM RUPTURE) AMONG PATIENT ADVERSE EVENTS INCLUDED: MINOR COMPLICATIONS (MRS 0-2): 16.5%, MAJOR COMPLICATIONS (MRS >2): 2.2%, MORTALITY RATE: 3.4%, OVERALL TREATMENT MORBIMORTALITY (MRS >2): 3% FOR S1V1E1, 0% FOR S2V1E0, AND 16.7% FOR S2V0E1, TRANSITORY NEUROLOGICAL DEFICITS: 2 PATIENTS (2.2%), PERMANENT PROCEDURE-RELATED MORBIDITY (MRS >2): 2 PATIENTS (1 UNRUPTURED AND 1 RUPTURED AVM), FULL INCIDENCE OF COMPLICATIONS, INCLUDING MINOR TEMPORARY COMPLICATIONS, IS APPROXIMATELY 21%, INDICATING THAT WHILE IMMEDIATE COMPLICATIONS ARE FREQUENT, MOST PATIENTS CAN ACHIEVE A GOOD FINAL NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: UNK-NV-ONYX; PRODUCT ID: UNK-NV-ONYX; G2: CITATION: AUTHORS: NATALIA VASCONCELLOS DE OLIVEIRA SOUZA, TABATA LAMIRAUX, FELIPE VENCATO DA SILVA, VINICIUS MOREIRA LIMA, AYMERIC ROUCHAUD, SUZANA SALEME, CHARBEL MOUNAYER. ENDOVASCULAR TREATMENT OF SPETZLER-MARTIN GRADE III ARTERIOVENOUS MALFORMATIONS: A SINGLE-CENTER 12 YEARS¿ EXPERIENCE STRATIFIED BY THE SPETZLER-MARTIN MODIFIED SCALE. NEUROSURGERY 95:1378¿1387 2024. DOI: 10.1227/NEU.0000000000003016. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480724&pc=NPS | 2135147-2025-00996 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-02-27 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "TRICUSPID REGURGITATION RISK SCORES IN PATIENTS UNDERGOING TRICUSPID VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A (PROSPECTIVE, RETROSPECTIVE) (MULTI VS SINGLE CENTER) STUDY, TO EVALUATE THE PERFORMANCE OF SIX CURRENTLY AVAILABLE RISK SCORES (TRISCORE, WANG SCORE, STS SCORE, TRIO SCORE,TRIVALVE SCORE, LAPAR SCORE) IN A LARGE REAL-WORLD TRICUSPID TRANSCATHETER EDGE TO EDGE REPAIR (T-TEER) COHORT FOR THE PREDICTION OF 1-YEAR SURVIVAL. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND PASCAL. THE ARTICLE CONCLUDED THAT ALTHOUGH PROVIDING PREDICTIVE VALUE IN TERMS OF 1-YEAR MORTALITY, AREA UNDER THE CURVE AND THUS PREDICTIVE ACCURACY WERE COMPARABLY LOW FOR ALL INVESTIGATED TR RISK SCORES. FURTHERMORE, THE EXISTING SCORES DID NOT ALLOW THE IDENTIFICATION OF A SUB-POPULATION OF T-TEER PATIENTS IN WHICH THE PROCEDURE MIGHT BE HIGH RISK OR EVEN FUTILE DUE TO EXCESSIVE 1-YEAR MORTALITY RATES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS LUKAS STOLZ AT GERMAN CENTER FOR CARDIOVASCULAR RESEARCH (DZHK), PARTNER SITE MUNICH HEART ALLIANCE, MUNICH, GERMANY WITH CORRESPONDING EMAIL: LUKAS.STOLZ@MED.UNI-MUENCHEN.DE. THE TIME FRAME OF THE STUDY WAS 2016 TO 2023. SAY NOT CONFIRMED OR REPORTED IF APPLICABLE. A TOTAL OF 385 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED TRICUSPID REGURGITATION AND HEART FAILURE. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE "TRICUSPID REGURGITATION RISK SCORES IN PATIENTS UNDERGOING TRICUSPID VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR". SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING TRICUSPID REGURGITATION AND HEART FAILURE. COMPLICATIONS REPORTED INCLUDED DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480751&pc=DSQ | 2916596-2025-01150 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fever; Septic Shock; Tachycardia; Increased Respiratory Rate; Heart Failure/Congestive Heart Failure; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL AFTER 3 MONTHS OF DISPLACEMENT. DUE TO THE PREVIOUS WAR WITNESSED IN LEBANON, THE PATIENT COULD NOT HAVE APPROPRIATE MEDICAL FOLLOW-UPS DURING THE WAR. THE PATIENT WAS COMPLAINING OF A HIGH FEVER AND TACHYCARDIC ARRYTHMIAS, WITH SIGNS OF RIGHT HEART FAILURE, AND AN INFECTED DRIVELINE EXIT SITE. THE INFECTION PROGRESSED LEADING TO SEPTIC SHOCK AND RESPIRATORY DISTRESS AND UNCONTROLLED VASOPLEGIC SYNDROME WHICH RESULTED IN MULTIPLE ORGAN FAILURE (MOF). THE MOF, WHICH DID NOT EXIST PRIOR TO LEFT VENTRICULAR ASSIST DEVICE THERAPY, RESULTED IN GRADUAL DETERIORATION AND EVENTUAL DEATH. IT WAS NOTED THAT THE PATIENT OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING ARRHYTHMIA, INFECTION, RIGHT HEART FAILURE, MULTI-ORGAN FAILURE, AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. SECTION 6 (UNDER "RIGHT HEART FAILURE") ALSO DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. H6: 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480774&pc=NKM | 2135147-2025-00997 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-02-27 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRICUSPID REGURGITATION RISK SCORES IN PATIENTS UNDERGOING TRICUSPID VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A (PROSPECTIVE, RETROSPECTIVE) (MULTI VS SINGLE CENTER) STUDY, TO EVALUATE THE PERFORMANCE OF SIX CURRENTLY AVAILABLE RISK SCORES (TRISCORE, WANG SCORE, STS SCORE, TRIO SCORE, TRIVALVE SCORE, LAPAR SCORE) IN A LARGE REAL-WORLD TRICUSPID TRANSCATHETER EDGE TO EDGE REPAIR (T-TEER) COHORT FOR THE PREDICTION OF 1-YEAR SURVIVAL. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND PASCAL. THE ARTICLE CONCLUDED THAT ALTHOUGH PROVIDING PREDICTIVE VALUE IN TERMS OF 1-YEAR MORTALITY, AREA UNDER THE CURVE AND THUS PREDICTIVE ACCURACY WERE COMPARABLY LOW FOR ALL INVESTIGATED TR RISK SCORES. FURTHERMORE, THE EXISTING SCORES DID NOT ALLOW THE IDENTIFICATION OF A SUB-POPULATION OF T-TEER PATIENTS IN WHICH THE PROCEDURE MIGHT BE HIGH RISK OR EVEN FUTILE DUE TO EXCESSIVE 1-YEAR MORTALITY RATES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS LUKAS STOLZ AT GERMAN CENTER FOR CARDIOVASCULAR RESEARCH (DZHK), PARTNER SITE MUNICH HEART ALLIANCE, MUNICH, GERMANY WITH CORRESPONDING EMAIL: LUKAS.STOLZ@MED.UNI-MUENCHEN.DE. THE TIME FRAME OF THE STUDY WAS 2016 TO 2023. SAY NOT CONFIRMED OR REPORTED IF APPLICABLE. A TOTAL OF 385 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. COMBORIDITIES INCLUDED TRICUSPID REGURGITATION AND HEART FAILURE. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, OFF LABEL USE. Manufacturer Narrative: B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE "TRICUSPID REGURGITATION RISK SCORES IN PATIENTS UNDERGOING TRICUSPID VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR." SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING TRICUSPID REGURGITATION AND HEART FAILURE. COMPLICATIONS REPORTED INCLUDED DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE WAS ASSOCIATED WITH UTILIZING A MITRACLIP DEVICE ON THE TRICUSPID VALVE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480822&pc=OZD | 1220648-2025-26404 | 2025-02-06 05:00 | Death | ABIOMED, INC. | 2025-02-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP PUMP FOR THE PATIENT ADMITTED IN CARDIOGENIC SHOCK WHEN AFTER ONE DAY OF THE SUPPORT THE SIGNS OF HEMOLYSIS PROMPTED THE TEAM TO REMOVE AND REPLACE THE PUMP WHICH RESOLVED THE HEMOLYSIS. THE PATIENT REMAINED STABLE.Event Description: AFTER ONE DAY OF SUPPORT THE IMPELLA CP WAS EXCHANGED. TWO DAYS AFTER NEW IMPELLA WAS IMPLANTED THE PATIENT'S CONDITION DETERIORATED AND THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE PRODUCT AND DATA WERE NOT RETURNED FOR REVIEW. BASED ON CLINICAL DESCRIPTION, THE ROOT CAUSE OF HEMOLYSIS WAS MOST LIKELY DUE TO PUMP WAS NOT IN THE PROPER POSITION. B2 WAS REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26404 TO REFLECT THE PATIENT OUTCOME OF DEATH AND DATE OF DEATH WAS ADDED. B5 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26404 TO REFLECT THE PATIENT OUTCOME. H1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26404 TO REFLECT THE PATIENT OUTCOME OF DEATH. H6 HEALTH EFFECT - IMPACT CODE REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26404 TO REFLECT THE PATIENT OUTCOME OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21480993&pc=CBK | 9617566-2025-00004 | 2025-02-01 05:00 | Death | BREAS MEDICAL AB | 2025-02-27 05:00 | CBK | VIVO 45LS | Adverse Event Without Identified Device or Use Problem | Ventilator Dependent | K193586 |  | 1 | Event Description: ON FEBRUARY 3RD, INTEGRIS MEDICAL SUPPLY REPORTED TO BREAS THAT A PATIENT HAD DIED WHILE BEING TREATED BY A VIVO 45 LS DEVICE. 'WIFE STATES THAT SHE GOT A TEXT FROM PATIENT STATING VENT WAS MAKING NOISE APPROX. 10ISH AT NIGHT. WHEN SHE DID NOT SEE TEXT OR HEAR ANY ALARMS IN THE NIGHT. PATIENT SLEPT UNATTENDED. WHEN SHE WENT IN APPROX. 4AM NO LIGHTS WERE ON IN THE ROOM AND NO ODD SOUND FROM VENT OR PATIENT. WIFE SAID SHE WENT TO SHOWER AND CAME BACK AT APPROX. 5AM AND PATIENT HAD PASSED. HE WAS HOLDING HIS MASK IN HIS HANDS. SHE SAID PATIENT WAS STILL ABLE TO TAKE MASK OFF AND ON. OUR LAST VENT CHECK WE HAVE DOCUMENTED THAT PATIENT WAS STILL WANTING DISCONNECT ALARM TURN OFF.' Manufacturer Narrative: INVESTIGATION REPORT: OUR REFERENCE: (B)(4). YOUR REFERENCE: N/A. REPORTER: (B)(6). 1 EVENT SUMMARY AND BACKGROUND: WHEN WAS BREAS MADE AWARE OF THE EVENT? FEBRUARY 3RD, 2025. WHEN DID THE EVENT HAPPEN? ON (B)(6) 2025. WHEN WAS THE EVENT DISCOVERED? DURING ACTIVE TREATMENT. PATIENT IMPACT: PATIENT DEATH. DEVICE MODEL AND SERIAL# MODE: VIVO 45 LS, SN: (B)(6). REPORTERS DESCRIPTION OF THE EVENT: "WIFE STATES THAT SHE GOT A TEXT FROM PATIENT STATING VENT WAS MAKING NOISE APPROX. 10ISH AT NIGHT. WHEN SHE DID NOT SEE TEXT OR HEAR ANY ALARMS IN THE NIGHT. PATIENT SLEPT UNATTENDED. WHEN SHE WENT IN APPROX. 4AM NO LIGHTS WERE ON IN THE ROOM AND NO ODD SOUND FROM VENT OR PATIENT. WIFE SAID SHE WENT TO SHOWER AND CAME BACK AT APPROX. 5AM AND PATIENT HAD PASSED. HE WAS HOLDING HIS MASK IN HIS HANDS. SHE SAID PATIENT WAS STILL ABLE TO TAKE MASK OFF AND ON. OUR LAST VENT CHECK WE HAVE DOCUMENTED THAT PATIENT WAS STILL WANTING DISCONNECT ALARM TURN OFF." OTHER RELEVANT BACKGROUND INFORMATION: N/A. 2 INVESTIGATIONS: BREAS RECEIVED THE DEVICE ON FEBRUARY 7TH, 2025. THE INVESTIGATION STARTED ON FEBRUARY 13TH, 2025 AND COMPLETED ON FEBRUARY 19TH, 2025. THE INVESTIGATION INCLUDED THE FOLLOWING ACTIVITIES: FULL LOG FILES ANALYSIS, FULL FUNCTIONAL TESTING. 2.1 INSPECTION: VISUAL CONDITION: MINOR DAMAGE TO FRONT CASE SCREEN. NO SAFETY IMPACT. ACCESSORIES: CLICK-IN BATTERY. AFFIXED LABELS: MULTIPLE INTEGRIS LABELS ON TOP CASING. SYSTEM TIME AND DATE: DEVICE IS ON THE CORRECT DATE. APPROXIMATELY 3 MINUTES BEHIND. FIRMWARE VERSION: FW V5.1.5. USAGE HOURS: 4466.5 HOURS. INTERNAL BATTERY: SOH: 95% AND SOC: 100%. MODE AND SETTING AS RECEIVED: ALARMS AS RECEIVED: PHOTOS: 2.2 ANALYSIS OF LOG FILE. FIGURE 1 / FIGURE 2: IT IS IMPORTANT TO NOTE THE SETTINGS OF THE DEVICE PRIOR TO DISCUSSING THE TREATMENT LOGS. AT THE TIME OF THE EVENT, THE DEVICE HAS NEARLY ALL OF THE ALARMS SET TO "OFF". RELEVANT TO THIS CASE, THE DISCONNECTION ALARM WAS SET TO "OFF". THE LOW-PRESSURE ALARM WAS SET TO 1 CMH2O WHILE AT THE SAME TIME THE PEEP WAS SET TO 5 CMH2O. IN THE EVENT OF A DISCONNECTION, THE DEVICE WOULD CONTINUE TO PROVIDE A MINIMUM PRESSURE OF 5CMH2O, MEANING THE LOW-PRESSURE ALARM SETTING WOULD NEVER BE TRIGGERED. OTHER SETTINGS THAT COULD BE FOUND DURING A DISCONNECTION EVENT ARE LOW MVE AND LOW VTE; BOTH WERE SET TO "OFF". INCLUDED ABOVE ARE SECTIONS OF THE LOG FILES THAT WERE OBTAINED FROM THE DEVICE. THE FIRST IMAGE, REFERRED TO AS "FIGURE 1" FROM THIS POINT FORWARD, IS A SNAPSHOT OF THE TREATMENT DATA WITHIN THE TIMELINE PROVIDED BY THE REPORTER. THE SECOND IMAGE, REFERRED TO AS "FIGURE 2" FROM THIS POINT FORWARD, IS A SNAPSHOT OF THE TREATMENT DATA FROM SHORTLY AFTER THE EVENT OCCURRED. IN FIGURE 1, WE FIND SIGNS OF REGULAR TREATMENT COMING TO A HALT AT APPROXIMATELY 1:48AM AND THEN PRESENT CONSTANT VALUES UNTIL THE DEVICE WAS TURNED OFF. FIGURE 2 SHOWS A MORE DETAILED, BREATH-BY-BREATH VIEW OF THESE CONSTANT VALUES. THESE CONSTANT VALUES WERE MEETING THE TREATMENT SETTINGS SET TO THE DEVICE AT THE TIME. THIS WOULD SUGGEST THE DEVICE WAS NOT CONNECTED TO A PATIENT AT THIS TIME. 2.3 CALIBRATION: AS RECEIVED: N/A. AFTER RECALIBRATION: N/A. 2.4 FUNCTIONS AND ALARMS: DEVICE PASSED ALL VERIFICATION, ALARM, AND BUTTON TESTING WITH NO FAILURES. 3 CAUSES: IS THE ROOT CAUSE CONFIRMED? THE DEVICE FUNCTIONED AS DESIGNED BASED ON THE SETTINGS SET AT THE TIME. IS THE CUSTOMER DESCRIPTION CONFIRMED OR NOT? BASED ON THE LOG FILES FROM THE DEVICE, IT APPEARS THE PATIENT WAS DISCONNECTED FROM THE DEVICE IN THE MIDDLE OF TREATMENT. CUSTOMER DESCRIPTION CONFIRMED. PROBABLE CAUSES WITH RATIONALE. THE DEVICE FUNCTIONED AS DESIGNED BASED ON THE SETTINGS SET AT THE TIME. THE CAUSE WAS NOT FROM THE DEVICE. CAUSES THAT COULD BE EXCLUDED WITH RATIONALE: N/A. 4 RISK ASSESSMENT: THIS EVENT DOES NOT INTRODUCE A NEW RISK. INVESTIGATION INDICATES THAT THE DEVICE PERFORMED ACCORDING TO SPECIFICATION. LOG FILES SHOW THAT THE DEVICE WAS NOT CONNECTED TO THE PATIENT AT THE TIME OF THE EVENT WHICH IS ALIGNED WITH THE REPORT FROM THE CUSTOMER. THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THIS ADVERSE EVENT. 5 CONCLUSION AND RECOMMENDED ACTIONS: THE DEVICE WAS INVESTIGATED, AND AN ANALYSIS WAS MADE WHICH SHOWED THAT THE DEVICE PERFORMED ACCORDING TO SPECIFICATION AND ALARMED ACCORDING TO THE DEVICE SETTINGS. LOG FILES SHOW THAT THE PATIENT WAS NOT CONNECTED TO THE DEVICE AT THE TIME OF DEATH. THE INVESTIGATION CONCLUDES THAT THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21481138&pc=NPT | 2015691-2025-01388 | 2025-01-20 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-27 05:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Calcified | Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE IMPLANT PATIENT REGISTRY, APPROXIMATELY 8 YEARS, 5 MONTHS, AND 19 DAYS POST TAVR WITH A 26MM SAPIEN 3 (S3) VALVE THE VALVE WAS EXPLANTED. AN EDWARDS SURGICAL VALVE WAS IMPLANTED. PER FOLLOW UP THE PATIENT HAD SEVERE AORTIC VALVE STENOSIS WITH A MEAN GRADIENT OF 42.7 MMHG AND A PEAK GRADIENT OF 64.8 MMHG WITH A DI OF 0.18. THE AORTIC VALVE AREA WAS 0.4 CM2. THERE WAS MILD PROSTHETIC VALVE REGURGITATION AND MILD MITRAL ANNULAR CALCIFICATION. THERE WAS MODERATE TO SEVERE MITRAL REGURGITATION AND MILD TRICUSPID REGURGITATION. ESTIMATED PULMONARY ARTERY SYSTOLIC PRESSURE WAS 40 MMHG RAP OF 5 MMHG. PER MEDICAL RECORDS, UNDER GENERAL ANESTHESIA, THE CHEST WAS OPENED WITH AN OSCILLATING SAW AND, UPON ENTERING THE UPPER STERNAL TUBULAR JUNCTION THERE WAS VENOUS BLEEDING FROM THE INNOMINATE VEIN. THE CHEST WAS OPENED FURTHER TO HELP CONTROL THE BLEEDING. THE 26MM S3 VALVE WAS EXPLANTED AND, DEBRIDED OF ITS CALCIUM AS WAS THE VALVE LEAFLETS. AFTER MULTIPLE ATTEMPTS, THE PATIENT WAS UNABLE TO COME OFF THE HEART/LUNG MACHINE. AN INTRA-AORTIC BALLOON PUMP WAS PLACED. DESPITE MULTIPLE ATTEMPTS FOR CARDIAC SUPPORT, THE PATIENTS, HEART SUCCUMBED AND COULD NOT BE RESUSCITATED. THE PATIENT EXPIRED. NO OFFICIAL CAUSE OF DEATH. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER A TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, CALCIFIC DEGENERATION IS A COMMON CAUSE OF BIOPROSTHETIC HEART VALVE FAILURES. MANY FACTORS CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION. THESE INCLUDE PATIENT FACTORS (AGE, DISEASE STATE, PHARMACOLOGICAL INTERVENTION, ETC.), MECHANICAL STRESS RELATED TO THE VALVE'S HEMODYNAMIC PERFORMANCE, AND GLUTARALDEHYDE FIXATION OF TISSUE. OF THESE, THE FIXATION PROCESS IS A RELATIVELY MINOR CONTRIBUTOR TO CALCIFICATION FOR EDWARDS' TISSUE VALVES DUE TO SUBSEQUENT ANTI-CALCIFICATION TREATMENTS DURING MANUFACTURING. THOUGH NUMEROUS STUDIES HAVE BEEN CONDUCTED ON PREVENTIVE CALCIFICATION STRATEGIES IN BIOPROSTHETIC HEART VALVES, THE CAUSES OF CALCIFICATION ARE NOT FULLY UNDERSTOOD AND THERE ARE STILL NO MECHANISMS OR MEDICAL THERAPIES WHICH FULLY PREVENT BIOPROSTHESIS FROM CALCIFYING. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. INVESTIGATION RESULTS INDICATE (CORONARY ARTERY DISEASE, LYMPHOMA, CONGESTIVE HEART FAILURE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, AND EJECTION FRACTION OF 30 PERCENT) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21481159&pc=MIH | 2017233-2025-05868 | 2025-02-04 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Insufficient Information | Insufficient Information | P040043 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM UTILIZING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM AND GORE® TAG® THORACIC BRANCH ENDOPROSTHESES IN ZONE 3. IT WAS REPORTED THAT THE PATIENT TOLERATED THE PROCEDURE WELL. THE COMPLETION ANGIOGRAM SHOWED A SUCCESSFUL CASE WITH SUCCESSFUL TREATMENT OF DISEASED AREA. THE ANEURYSM WAS SEALED. SOMETIME AFTER THE PROCEDURE, AFTER REPRESENTATIVES HAD LEFT THE HOSPITAL, THE PATIENT BECAME UNSTABLE AND ALL LIFESAVING EFFORTS WERE UNSUCCESSFUL. THE PATIENT EXPIRED THAT SAME DAY. Manufacturer Narrative: ADDED E1, G3/G4, H1/H2, H6. Manufacturer Narrative: AS IT IS UNKNOWN WHICH DEVICE(S) IS DIRECTLY IMPLICATED IN THIS DEATH EVENT, THE FOLLOWING DEVICE(S) WILL ALSO BE INCLUDED IN THIS REPORT: CATALOG #TGM282815/ SERIAL # (B)(6)/ UDI # (B)(4); CATALOG #TGM282810/ SERIAL # (B)(6)/ UDI # (B)(4). H6: CODE A26- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF GORE DEVICE CONTRIBUTED TO DEATH. H6: CODE E2401- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF ANY HEALTH EFFECT APPEARS TO HAVE OCCURRED. IT SHOULD BE NOTED THAT, PER THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21481234&pc=MIH | 2017233-2025-05869 | 2025-02-04 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Insufficient Information | Insufficient Information | P210032 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM UTILIZING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM AND GORE® TAG® THORACIC BRANCH ENDOPROSTHESES IN ZONE 3. IT WAS REPORTED THAT THE PATIENT TOLERATED THE PROCEDURE WELL. THE COMPLETION ANGIOGRAM SHOWED A SUCCESSFUL CASE WITH SUCCESSFUL TREATMENT OF DISEASED AREA. THE ANEURYSM WAS SEALED. SOMETIME AFTER THE PROCEDURE, AFTER REPRESENTATIVES HAD LEFT THE HOSPITAL, THE PATIENT BECAME UNSTABLE AND ALL LIFESAVING EFFORTS WERE UNSUCCESSFUL. THE PATIENT EXPIRED THAT SAME DAY. Manufacturer Narrative: AS IT IS UNKNOWN WHICH DEVICE(S) IS DIRECTLY IMPLICATED IN THIS DEATH EVENT, THE FOLLOWING DEVICE(S) WILL ALSO BE INCLUDED IN THIS REPORT: CATALOG #TAC083415A/ SERIAL # (B)(6)/ UDI # (B)(4). H6: CODE A26- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF GORE DEVICE CONTRIBUTED TO DEATH. H6: CODE E2401- USED TO CAPTURE INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF ANY HEALTH EFFECT APPEARS TO HAVE OCCURRED. C1. CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. IT SHOULD BE NOTED THAT, PER THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21482585&pc=QFG | 3019004087-2025-00133 | 2025-01-29 05:00 | Death | BETA BIONIC, INC. | 2025-02-27 05:00 | QFG | ILET BIONIC PANCREAS | Insufficient Information | Hyperglycemia; Hypoglycemia; Diabetic Ketoacidosis; Loss of consciousness | K231485 |  | 1 | Event Description: ON (B)(6)2025, THE MANUFACTURER WAS NOTIFIED BY A CLINICAL DIABETES SPECIALIST (CDS) THAT THE USER WAS ADMITTED TO THE HOSPITAL WITH SUSPECTED DIABETIC KETOACIDOSIS (DKA). THE USER'S DAUGHTER REPORTED THAT A NEIGHBOR FOUND THE USER UNCONSCIOUS IN THEIR YARD BEFORE BEING TRANSPORTED TO THE HOSPITAL. THE USER TESTED POSITIVE FOR KETONES AND WAS DIAGNOSED WITH DKA. THE USER WAS INTUBATED AND TREATED WITH INTRAVENOUS INSULIN AND FLUIDS. ON (B)(6) 2025 THE DAUGHTER CONFIRMED THE USER REMAINED IN THE HOSPITAL UNTIL PASSING AWAY ON (B)(6) 2025.Event Description: ON 30-JAN-2025, IT WAS REPORTED THAT THE USER WAS FOUND UNCONSCIOUS IN THEIR YARD ON (B)(6) 2025. THEY WERE TRANSPORTED TO THE HOSPITAL AND TREATED FOR DKA. THE USER REMAINED IN THE HOSPITAL UNTIL PASSING AWAY 15 DAYS LATER ON (B)(6) 2025. Manufacturer Narrative: BASED UPON THE EVENT INFORMATION AND INVESTIGATION THE USER WAS NOT WEARING THE ILET AT THEIR TIME OF DEATH, A SPECIFIC CAUSE OF DEATH COULD NOT BE DETERMINED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. AVAILABLE ILET LOGS WERE REVIEWED. NO EVIDENCE OF DEVICE MALFUNCTION WAS FOUND IN THE ENGINEERING LOGS. THE ILET WAS PERFORMING AS INTENDED AND REACTED APPROPRIATELY TO AVAILABLE CGM GLUCOSE VALUES AND TRIGGERING CGM GLUCOSE ALERTS UNTIL SENSOR READINGS STOPPED AT 1:56AM ON (B)(6) 2025. SENSOR UNAVAILABLE AND REPLACE SENSOR ALERTS WERE TRIGGERED BUT NOT ACKNOWLEDGED. ALERTS WARNING OF IMPENDING SUSPENSION OF INSULIN DELIVERY WERE TRIGGERED BUT NOT ACKNOWLEDGED. INSULIN DELIVERY STOPPED AT 10:06AM ON (B)(6) 2025. ALERTS WARNING OF IMPENDING LOSS OF POWER WERE TRIGGERED BUT WERE NOT ACKNOWLEDGED. THE ILET RAN OUT OF POWER AT 3:30PM (B)(6) 2025. THE USER WAS REPORTEDLY FOUND UNCONSCIOUS IN THEIR YARD SOMETIME IN THE AFTERNOON, EXACT TIME UNKNOWN. AVAILABLE INFORMATION INDICATES THE USER WAS NOT USING THE ILET AT THE TIME OF DEATH ON (B)(6) 2025. THE REPORTED CAUSE OF DEATH APPEARS RELATED TO UNDERLYING DIABETES. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: AT THE TIME OF THIS REPORT THE ILET EVALUATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT SHALL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21482735&pc=NPT | 2025587-2025-01524 | 2025-01-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-27 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Myocardial Infarction; Tachycardia; Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, VENTRICULAR TACHYCARDIA WAS NOTED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED. THE SAME DAY PERCUTANEOUS CORONARY INTERVENTION ON THE RIGHT CORONARY ARTERY WAS COMPLETED UNDER INSERTION OF "ECCO". SUBSEQUENTLY, CARDIAC COLLAPSE WAS NOTED. THE CARDIAC COLLAPSE CAUSED TRAUMA TO THE INTERNAL THORACIC ARTERY WHICH LED TO IDIOVENTRICULAR RHYTHM. FOLLOWING THE OPERATION, THE PATIENT WAS MEDICALLY MANAGED IN THE INTENSIVE CARE UNIT. APPROXIMATELY 2 DAYS LATER, A CARDIAC ARREST WAS OBSERVED AND THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS ASSOCIATED WITH A RIGHT CORONARY ARTERY MYOCARDIAL INFARCTION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT "ECCO" WAS ACTUALLY ECMO. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4., H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21482793&pc=DSQ | 2916596-2025-01469 | 2015-05-01 04:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE HEARTMATE 3 (HM3) MAY BE RELATED TO RENAL FAILURE, NEUROLOGICAL DYSFUNCTION, HEMOLYSIS, INFECTION, PUMP EXCHANGE, HEART TRANSPLANT, AND DEATH. THIS IS A SINGLE CENTER RETROSPECTIVE STUDY WHICH EVALUATED THE CHARACTERISTICS AND OUTCOMES OF PATIENTS AGED LESS THAN 21 YEARS OLD WITH FONTAN PUMP FAILURE WHO UNDERWENT EITHER PULSATILE VENTRICULAR ASSIST DEVICE (PVAD) (BERLINE HEART EXCOR (BHE)) OR CONTINUOUS FLOW VENTRICULAR ASSIST DEVICE (CFVAD) (HEARTWARE VENTRICULAR ASSIST DEVICE (HVAD) OR HEARTMATE 3 (HM3) IMPLANT AS BRIDGE TO TRANSPLANT (BTT) FROM MAY 2015 TO NOVEMBER 2022. FIFTEEN PATIENTS MET INCLUSION CRITERIA¿5 BHE, 9 HVAD, AND 1 HM3. THE MEDIAN VAD DURATION OF BHE PATIENTS WAS 220 WHEREAS THE HVAD/HM3 PATIENTS WAS 199 (P VALUE WAS 0.95). THE MECHANICAL CIRCULATORY SUPPORT (MCS) CARDIAC INDEX (L/MIN/M^2) AS THE LAST VAD FOLLOW-UP WAS 3.13 FOR BHE PATIENTS AND 3.67 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.16). THE MEDIAN DAYS INTUBATED AFTER VAD IMPLANTATION WAS 4 FOR BHE PATIENTS AND 3 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.94). THE MEDIAN POST-VAD INTENSIVE CARE UNIT (ICU) LENGTH OF STAY WAS 71 DAYS FOR BHE PATIENTS AND 29 DAYS FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.023). THE NUMBER OF PATIENTS REQUIRING EXTRACORPOREAL MEMBRANE OXYGENATION POST VAD WAS 3 FOR BHE PATIENTS AND 0 FOR HVAD/HM3 PATIENTS (P VALUE 0.022). THE NUMBER OF PATIENTS REQUIRING RENAL REPLACEMENT THERAPY ON VAD SUPPORT WAS 2 FOR BHE PATIENTS AND 1 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.51). THE NUMBER OF PATIENTS ON CHRONIC MILRINONE ON MCS SUPPORT WAS 3 FOR BHE PATIENTS AND 0 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.022). THE NUMBER OF PATIENTS ON CHRONIC STEROID USE ON MCS SUPPORT WAS 2 FOR BHE PATIENTS AND 1 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.51). THE NUMBER OF PATIENTS WHO TOLERATED FULL ENTERAL FEEDS WAS 3 FOR BHE PATIENTS AND 8 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.56). THE NUMBER OF PATIENTS WHO EXPERIENCED CHEMICAL PANCREATITIS WAS 4 FOR BHE PATIENTS AND 4 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.28). THE NUMBER OF PATIENTS WHO EXPERIENCED A NEUROLOGICAL DYSFUNCTION AFTER VAD IMPLANTATION WAS 2 FOR BHE PATIENTS AND 7 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.58). THE NUMBER OF PATIENTS WHO EXPERIENCED MAJOR HEMOLYSIS WAS 4 FOR BHE PATIENTS AND 5 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.58). THE NUMBER OF PATIENTS WHO EXPERIENCED A MAJOR INFECTION WAS 4 FOR HBE PATIENTS AND 4 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.28). THE NUMBER OF PATIENTS WHO UNDERWENT A PUMP EXCHANGE WAS 4 FOR BHE PATIENTS AND 1 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.017). THE NUMBER OF PATIENTS WHO SURVIVED TO ORTHOTOPIC HEART TRANSPLANTATION (OHT) WAS 3 FOR BHE PATIENTS AND 6 FOR HVAD/HM3 PATIENTS (P VALUE WAS 1). THE NUMBER OF PATIENTS WHO WERE IMPATIENT AT THE TIME OF OHT WAS 3 FOR BHE PATIENTS AND 1 FOR HVAD/HM3 PATIENTS (P VALUE WAS 0.048). THE NUMBER OF PATIENTS WHO HAD ONGOING VAD SUPPORT WAS 0 FOR BHE PATIENTS AND 1 FOR HVAD/HM3 PATIENTS (P VALUE WAS 1). CONCOMITANT FONTAN COMPLETION WAS PERFORMED IN 5 OUT OF 10 PATIENTS RECEIVING CFVAD. THE PVAD COHORT HAD SIGNIFICANTLY LONGER POST-VAD ICU LENGTH OF STAY (P=0.023) AND WERE LESS LIKELY TO WEAN OFF INOTROPE SUPPORT (P=0.022). PATIENTS ON PVAD WERE MORE LIKELY TO REQUIRE ADDITIONAL OXYGENATOR SUPPORT (P=0.022); 2 DUE TO HEMORRHAGIC SHOCK (1 PULMONARY IN SETTING OF PLASTIC BRONCHITIS AND 1 DUE TO CANNULA DISRUPTION) AND 1 DUE TO REFRACTORY HYPOXEMIA. THE FORMER 2 PATIENTS DIED ON MECHANICAL SUPPORT, WHILE THE THIRD PATIENT REMAINED ON OXYGENATOR AND CONTINUOUS FLOW SUPPORT UNTIL HEART TRANSPLANT (HT). OVERALL MORTALITY FOR THE PVAD AND CFVAD COHORTS WERE 60% (2 PRE-HT AND 1 POST-HT) AND 30% (ALL PRE-HT) RESPECTIVELY. PATIENTS ON PVAD WERE MORE LIKELY TO NEED PUMP EXCHANGE (P=0.017) AND BE INPATIENT AT TIME OF HT (P=0.048). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT IS APPROXIMATE AS THE DATA WAS COLLECTED BETWEEN MAY 2015 AND NOVEMBER 2022. AUTHOR INFORMATION: SEYMOUR, E., H. TUNUGUNTLA, SOSA, S., HOPE, K., SPINNER, J., A. ANKOLA, ADACHI, I., & PURI, K. (2024). PULSATILE VERSUS CONTINUOUS FLOW VENTRICULAR ASSIST DEVICE AS A BRIDGE TO TRANSPLANTATION FOR THE PEDIATRIC FAILING FONTAN. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S628¿S628. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.981. PEDIATRICS-CRITICAL CARE MEDICINE, TEXAS CHILDREN'S HOSPITAL, HOUSTON, TX; PEDIATRICS-CARDIOLOGY, TEXAS CHILDREN'S HOSPITAL, HOUSTON, TX; PEDIATRICS-CRITICAL CARE MEDICINE AND CARDIOLOGY, TEXAS CHILDREN'S HOSPITAL, HOUSTON, TX; SURGERY-CONGENITAL HEART SURGERY, TEXAS CHILDREN'S HOSPITAL, HOUSTON, TX. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A SPECIFIC CAUSE FOR THE PATIENT DEATHS COULD NOT BE CONCLUSIVELY DETERMINED. WHETHER THE PATIENT OUTCOMES WERE DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE STUDY CONCLUDED THAT PATIENTS ON BERLIN HEART EXCOR (BHE) WERE MORE LIKELY TO REQUIRE INOTROPIC SUPPORT, AN OXYGENATOR, AND INTENSIVE CARE COMPARED TO CONTINUOUS FLOW VENTRICULAR ASSIST DEVICE (CFVAD) PATIENTS, SUGGESTING ONGOING SYMPTOMS OF HEART FAILURE DESPITE PULSATILE VENTRICULAR ASSIST DEVICE (PVAD) THERAPY. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483119&pc=DSQ | 2916596-2025-01447 | 2023-03-06 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Septic Shock; Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation; Mitral Valve Insufficiency/ Regurgitation; Tricuspid Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD RIGHT HEART FAILURE (RHF). THE PATIENT HAD RIGHT VENTRICULAR (RV) HYPOKINESIS ON (B)(6) 2023. THE PATIENT HAD DILATED AND HYPOKINETIC RV ON (B)(6) 2024. THE PATIENT HAD A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACED ON (B)(6) 2025. THERE WAS NO DEVICE RELATED ISSUE THAT CONTRIBUTED TO THE RHF, AND THE DEVICE OPERATED AS EXPECTED. THE PATIENT WAS LISTED FOR HEART TRANSPLANT.Event Description: RHF EXISTED PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. A DEVICE RELATED ISSUE WAS CONSIDERED TO CAUSE AND CONTRIBUTE TO RHF, BUT IT WAS UNKNOWN HOW. THE PATIENT WAS ACUTE ON CHRONIC SYSTOLIC HEART FAILURE WITH ISCHEMIC CARDIOMYOPATHY (ICM). THERE WAS AN EJECTION FRACTION OF 20-25%. THE PATIENT WAS ACUTE ON CHRONIC RIGHT-SIDED HEART FAILURE AND WERE ON IMPELLA RP FLEX SINCE (B)(6) 2025. THE PATIENT HAD SEPTIC SHOCK DUE TO PSEUDOMONAS AERUGINOSA BACTEREMIA. THE PATIENT HAD CORONARY ARTERY DISEASE (CAD). THE PATIENT HAD MODERATE AORTIC REGURGITATION AND SEVERE MITRAL REGURGITATION. THEY WERE STATUS POST TRICUSPID ANNULOPLASTY. THEY HAD ACUTE KIDNEY INJURY (AKI) ON CHRONIC KIDNEY DISEASE (CKD). THE PATIENT WAS ANURIC AND IT WAS LIKELY CARDIORENAL WITH/WITHOUT PIGMENT NEPHROPATHY. THE PATIENT HAD ACUTE HYPOXEMIC RESPIRATORY FAILURE AND WERE ON A VENTILATOR. THERE WAS BILATERAL PLEURAL EFFUSIONS STATUS POST RIGHT THORACENTESIS ON (B)(6) 2025. THE PATIENT HAD A HISTORY OF STROKE. THE PATIENT WAS HEMODYNAMICALLY STABLE ON LOW-DOSE VASOPRESSORS AND IMPELLA RP FLEX BUT WITH PERSISTENT SEVERE ELEVATION OF LEFT AND RIGHT HEART PRESSURES, WORSENING RENAL FUNCTION AND ANURIA DESPITE NORMAL CARDIAC OUTPUT. THEY CONTINUED HEARTMATE 3 LVAD SUPPORT AT 5400 REVOLUTIONS PER MINUTE (RPM). THE PATIENT CONTINUED IMPELLA RP FLEX SUPPORT AT P7. SENSOR WAS UNABLE TO DETECT FLOW. THERE WAS EVIDENCE OF HEMOLYSIS. IMPELLA RP FLEX EXPLANTATION NEEDED TO BE CONSIDERED, HOWEVER, THE PATIENT WAS UNLIKELY TO TOLERATE IT FROM AN RV FAILURE STANDPOINT. THE PATIENT WEANED VASOPRESSORS AS TOLERATED. THEY CONTINUED ON BUMEX 2 MG/HOUR. NEPHROLOGY WAS CONSULTED TO DISCUSS HEMODIALYSIS AS AN OPTION. AT THE FAMILY MEETING, IT WAS DISCUSSED THAT THE PATIENT HAD SEPTIC SHOCK DUE TO PSEUDOMONAS BACTEREMIA ON A BACKGROUND OF SEVERE BIVENTRICULAR HEART FAILURE WITH AN UNDER SUPPORTED LEFT VENTRICLE (LV) AND REFRACTORY RV FAILURE DESPITE RV SUPPORT WITH IMPELLA RP FLEX. THE PATIENT HAD WORSENING RENAL FUNCTION AND WOULD REQUIRE HEMODIALYSIS, WHICH THEY WERE UNLIKELY TO TOLERATE. IT WAS LIKELY THAT THE PATIENT'S IMPELLA RP FLEX WAS EXACERBATING THEIR RENAL DYSFUNCTION, HOWEVER, IT WAS UNCLEAR THAT THEY WOULD TOLERATE EXPLANT. IT WAS NOTED THAT AN UPGRADE TO A PROTEK RVAD WOULD BE AN OPTION, BUT IT WAS UNCLEAR WHAT THE EXIT STRATEGY WOULD BE. THE PATIENT WAS INACTIVATED ON THE ORTHOTROPIC HEART TRANSPLANTATION WAITLIST DUE TO SEVERE ILLNESS. THE PATIENT HAD ROUTINE HEMODYNAMICS. THE PATIENT WAS ON ANTIBIOTICS AND HEPARIN. HEART FAILURE MEDICATIONS WERE HELD BECAUSE OF SEPSIS. THE PATIENT CONTINUED ON AMIODARONE AT 0.5 MG/MIN AND ON MEXILETINE 150 MG TWICE A DAY. THE PATIENT HAD STRICT INPUTS AND OUTPUTS AND DAILY WEIGHTS. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: A4 PATIENT WEIGHT NOT PROVIDED. ADDITIONAL H6 CODES: 2041 - RENAL FAILURE. 2484 - RESPIRATORY FAILURE. 2010 - PLEURAL EFFUSION. 1770 - STROKE/CVA. 1886 - HEMOLYSIS. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED OR RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN DETAILS REGARDING THE ONSET DATES OF THE SEVERE MITRAL REGURGITATION, RENAL DISFUNCTION (AKI/CKD), RESPIRATORY FAILURE, INFECTION, AND STROKE; HOWEVER, NO RESPONSE WAS RECEIVED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS (INCLUDING DEATH, RIGHT HEART FAILURE, INFECTION, RESPIRATORY FAILURE, SEPSIS, RENAL DYSFUNCTION, STROKE, HEMOLYSIS, AND CARDIAC ARRHYTHMIA) THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS DOCUMENT ALSO WARNS THE USER THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. ADDITIONALLY, INSTRUCTIONS FOR HOW RESPOND TO INFECTION ARE PROVIDED. SECTION 6 ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. BOTH THE IFU AND PATIENT HANDBOOK CONTAIN VARIOUS SECTIONS REGARDING INFECTION AND HOW TO PREVENT IT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483188&pc=NSA | 3004123209-2025-00046 | 2025-01-05 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-02-27 05:00 | NSA | PACKAGE,350P,PP03,USB,HE,350-BAS-HE-10 | Failure to Sense | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A SUDDEN CARDIAC ARREST EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THIS ISSUE IS PATIENT RELATED; HOWEVER THERE WAS NO ADVERSE EVENT REPORTED.Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A SUDDEN CARDIAC ARREST EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THIS ISSUE IS PATIENT RELATED; HOWEVER THERE WAS NO ADVERSE EVENT REPORTED. Manufacturer Narrative: HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. NO RESPONSE HAS BEEN RECEIVED FROM THE CUSTOMER. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. Manufacturer Narrative: THE CUSTOMER'S DEVICE WAS NOT RETURNED TO HEARTSINE FOR EVALUATION. A CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483286&pc=DQY | 2124215-2025-12158 | 2025-02-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-27 05:00 | DQY | WATCHMAN FXD CURVE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Non specific EKG/ECG Changes; Heart Block | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED UNDER DEEP SEDATION ONLY AND INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS USED INTRAPROCEDURE. VENOUS ACCESS WAS OBTAINED, AND A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) WAS INSERTED ALONG WITH VERSACROSS CONNECT (VCC) TRANSSEPTAL ACCESS SYSTEM. TRANSSEPTAL PUNCTURE (TSP) WAS PERFORMED AND THE WAS AND VCC WAS ADVANCED INTO THE LEFT ATRIUM (LA), THEN THE VCC SYSTEM WAS REMOVED. THE PHYSICIAN STATED THE PATIENT WAS OBSERVED TAKING A BIG SNORE WHEN THE VCC DILATOR WAS REMOVED. A PIGTAIL CATHETER WAS ADVANCED INTO THE LAA AND A CONTRAST SHOT TAKEN. IMMEDIATE ST SEGMENT ELEVATION THEN HEART BLOCK WAS NOTED. AIR EMBOLISM WAS NOTED IN THE RIGHT CORONARY ARTERY (RCA). AN INTERVENTIONAL CARDIOLOGIST WHO WAS PRESENT USED A NON-BOSTON SCIENTIFIC (BSC) PERIPHERAL THROMBECTOMY DEVICE TO CLEAR THE RCA OF THE AIR EMBOLISM. VITAL SIGNS STABILIZED AND THE INDEX PROCEDURE WAS CONTINUED. A 24 MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS INSERTED, DEPLOYED, AND IMPLANTED. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS EXPECTED TO FULLY RECOVER BUT INSTEAD WAS TRANSFERRED TO ANOTHER HOSPITAL FOR HYPERBARIC THERAPY. THE FOLLOWING DAY POST INDEX PROCEDURE, THE PATIENT COMPLETED TWO (2) ROUNDS OF HYPERBARIC THERAPY WITHOUT MUCH CHANGE IN CONDITION. IN THE PHYSICIAN'S OPINION, THE AIR EMBOLISM IS BELIEVED TO HAVE OCCURRED WHEN THE VCC DILATOR WAS REMOVED.Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED UNDER DEEP SEDATION ONLY AND INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS USED INTRAPROCEDURE. VENOUS ACCESS WAS OBTAINED, AND A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) WAS INSERTED ALONG WITH VERSACROSS CONNECT (VCC) TRANSSEPTAL ACCESS SYSTEM. TRANSSEPTAL PUNCTURE (TSP) WAS PERFORMED AND THE WAS AND VCC WAS ADVANCED INTO THE LEFT ATRIUM (LA), THEN THE VCC SYSTEM WAS REMOVED. THE PHYSICIAN STATED THE PATIENT WAS OBSERVED TAKING A BIG SNORE WHEN THE VCC DILATOR WAS REMOVED. A PIGTAIL CATHETER WAS ADVANCED INTO THE LAA AND A CONTRAST SHOT TAKEN. IMMEDIATE ST SEGMENT ELEVATION THEN HEART BLOCK WAS NOTED. AIR EMBOLISM WAS NOTED IN THE RIGHT CORONARY ARTERY (RCA). AN INTERVENTIONAL CARDIOLOGIST WHO WAS PRESENT USED A NON-BOSTON SCIENTIFIC (BSC) PERIPHERAL THROMBECTOMY DEVICE TO CLEAR THE RCA OF THE AIR EMBOLISM. VITAL SIGNS STABILIZED AND THE INDEX PROCEDURE WAS CONTINUED. A 24MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS INSERTED, DEPLOYED, AND IMPLANTED. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS EXPECTED TO FULLY RECOVER BUT INSTEAD WAS TRANSFERRED TO ANOTHER HOSPITAL FOR HYPERBARIC THERAPY. THE FOLLOWING DAY POST INDEX PROCEDURE, THE PATIENT COMPLETED TWO (2) ROUNDS OF HYPERBARIC THERAPY WITHOUT MUCH CHANGE IN CONDITION. IN THE PHYSICIAN'S OPINION, THE AIR EMBOLISM IS BELIEVED TO HAVE OCCURRED WHEN THE VCC DILATOR WAS REMOVED. IT WAS FURTHER REPORTED THE PATIENT DID NOT RECOVER AND DIED TWO (2) DAYS POST INDEX PROCEDURE. THE CAUSE OR DETAILS REGARDING THE DEATH ARE UNKNOWN. IN THE PHYSICIAN'S OPINION, THE VALVE ON THE WAS CONTRIBUTED TO THE INTRODUCTION OF AIR DURING THE INDEX PROCEDURE.Event Description: IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED UNDER DEEP SEDATION ONLY AND INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS USED INTRAPROCEDURE. VENOUS ACCESS WAS OBTAINED, AND A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) WAS INSERTED ALONG WITH VERSACROSS CONNECT (VCC) TRANSSEPTAL ACCESS SYSTEM. TRANSSEPTAL PUNCTURE (TSP) WAS PERFORMED AND THE WAS AND VCC WAS ADVANCED INTO THE LEFT ATRIUM (LA), THEN THE VCC SYSTEM WAS REMOVED. THE PHYSICIAN STATED THE PATIENT WAS OBSERVED TAKING A BIG SNORE WHEN THE VCC DILATOR WAS REMOVED. A PIGTAIL CATHETER WAS ADVANCED INTO THE LAA AND A CONTRAST SHOT TAKEN. IMMEDIATE ST SEGMENT ELEVATION THEN HEART BLOCK WAS NOTED. AIR EMBOLISM WAS NOTED IN THE RIGHT CORONARY ARTERY (RCA). AN INTERVENTIONAL CARDIOLOGIST WHO WAS PRESENT USED A NON-BOSTON SCIENTIFIC (BSC) PERIPHERAL THROMBECTOMY DEVICE TO CLEAR THE RCA OF THE AIR EMBOLISM. VITAL SIGNS STABILIZED AND THE INDEX PROCEDURE WAS CONTINUED. A 24MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS INSERTED, DEPLOYED, AND IMPLANTED. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS EXPECTED TO FULLY RECOVER BUT INSTEAD WAS TRANSFERRED TO ANOTHER HOSPITAL FOR HYPERBARIC THERAPY. THE FOLLOWING DAY POST INDEX PROCEDURE, THE PATIENT COMPLETED TWO (2) ROUNDS OF HYPERBARIC THERAPY WITHOUT MUCH CHANGE IN CONDITION. IN THE PHYSICIAN'S OPINION, THE AIR EMBOLISM IS BELIEVED TO HAVE OCCURRED WHEN THE VCC DILATOR WAS REMOVED. IT WAS FURTHER REPORTED THE PATIENT DID NOT RECOVER AND DIED TWO (2) DAYS POST INDEX PROCEDURE. THE CAUSE OR DETAILS REGARDING THE DEATH ARE UNKNOWN. IN THE PHYSICIAN'S OPINION, THE VALVE ON THE WAS CONTRIBUTED TO THE INTRODUCTION OF AIR DURING THE INDEX PROCEDURE. IT WAS FURTHER REPORTED A MAGNETIC RESONANCE IMAGING (MRI) SCAN PERFORMED POST INDEX PROCEDURE, REVEALED AIR IN THE PATIENT'S BRAIN. Manufacturer Narrative: B5: INFORMATION ADDED. Manufacturer Narrative: B2: DEATH OUTCOME AND DATE OF DEATH ADDED. B5: INFORMATION ADDED. H1: UPDATED FROM SERIOUS INJURY TO DEATH. H6 IMPACT CODE ADDED: DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483519&pc=DXF | 2124215-2025-11927 | 2025-02-06 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-27 05:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Perforation; Retroperitoneal Hemorrhage | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC), A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND A 35MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED USING A WATCHMAN TRUSTEER SHEATH. THERE WERE NO RECAPTURES OF THE CLOSURE DEVICE NEEDED AND WAS IMPLANTED AFTER IT MET RELEASE CRITERIA. THE PROCEDURE WAS COMPLETED WITH NO KNOWN ISSUES. AFTER THE PROCEDURE, WHEN THE PATIENT ARRIVED TO POST ANESTHESIA CARE UNIT (PACU), THE PATIENT CODED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS DISCOVERED THAT THE PATIENT HAD A RETROPERITONEAL BLEED CAUSED BY TWO HOLES AT THE INFRARENAL INFERIOR VENA CAVA (IVC). ONE OF THE HOLES WAS SUTURED AND THE SECOND HOLE (WHICH WAS APPROXIMATELY A CENTIMETER LACERATION) WAS UNABLE TO BE REPAIRED. THE PATIENT HAD SEVERAL UNKNOWN COMORBIDITIES THAT CONTRIBUTED TO HIS CONDITION. THE PHYSICIAN EXPERIENCED NO DIFFICULTIES WITH ACCESS, REMOVAL, OR CLOSURE WITH A NON-BOSTON SCIENTIFIC PRE-CLOSE SUTURE DEVICE. IN THE PHYSICIAN OPINION, THE LACERATIONS ARE FROM AN UNKNOWN CAUSE. THE PATIENT EXPERIENCED NO HEMODYNAMIC CHANGES DURING THE PROCEDURE AND ONLY EXHIBITED ISSUES AFTER LEAVING THE PROCEDURE ROOM. THE OFFICIAL CAUSE OF DEATH DETERMINED WAS HEMORRHAGIC SHOCK. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISCARDED). THE PHYSICIAN MENTIONED HE DID NOT THINK THERE WAS AN ISSUE WITH ACCESS, TRANSSEPTAL, OR ANYTHING AT THE BEGINNING OF THE PROCEDURE.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC), A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND A 35MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED USING A WATCHMAN TRUSTEER SHEATH. THERE WERE NO RECAPTURES OF THE CLOSURE DEVICE NEEDED AND WAS IMPLANTED AFTER IT MET RELEASE CRITERIA. THE PROCEDURE WAS COMPLETED WITH NO KNOWN ISSUES. AFTER THE PROCEDURE, WHEN THE PATIENT ARRIVED TO POST ANESTHESIA CARE UNIT (PACU), THE PATIENT CODED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS DISCOVERED THAT THE PATIENT HAD A RETROPERITONEAL BLEED CAUSED BY TWO HOLES AT THE INFRARENAL INFERIOR VENA CAVA (IVC). ONE OF THE HOLES WAS SUTURED AND THE SECOND HOLE (WHICH WAS APPROXIMATELY A CENTIMETER LACERATION) WAS UNABLE TO BE REPAIRED. THE PATIENT HAD SEVERAL UNKNOWN COMORBIDITIES THAT CONTRIBUTED TO HIS CONDITION. THE PHYSICIAN EXPERIENCED NO DIFFICULTIES WITH ACCESS, REMOVAL, OR CLOSURE WITH A NON-BOSTON SCIENTIFIC PRE-CLOSE SUTURE DEVICE. IN THE PHYSICIAN OPINION, THE LACERATIONS ARE FROM AN UNKNOWN CAUSE. THE PATIENT EXPERIENCED NO HEMODYNAMIC CHANGES DURING THE PROCEDURE AND ONLY EXHIBITED ISSUES AFTER LEAVING THE PROCEDURE ROOM. THE OFFICIAL CAUSE OF DEATH DETERMINED WAS HEMORRHAGIC SHOCK. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISCARDED). THE PHYSICIAN MENTIONED HE DID NOT THINK THERE WAS AN ISSUE WITH ACCESS, TRANSSEPTAL, OR ANYTHING AT THE BEGINNING OF THE PROCEDURE. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS DILATOR. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: CORRECTION TO THE INITIAL MDR IN BLOCK(S) PATIENT CODES F10 AND H6, UNDER SECTIONS F. FOR USE BY UF/IMPORTER AND H. DEVICE MANUFACTURERS ONLY. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS DILATOR. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483707&pc=DRE | 2124215-2025-11929 | 2025-02-06 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-02-27 05:00 | DRE | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Perforation; Retroperitoneal Hemorrhage | K220414 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC), A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND A 35MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED USING A WATCHMAN TRUSTEER SHEATH. THERE WERE NO RECAPTURES OF THE CLOSURE DEVICE NEEDED AND WAS IMPLANTED AFTER IT MET RELEASE CRITERIA. THE PROCEDURE WAS COMPLETED WITH NO KNOWN ISSUES. AFTER THE PROCEDURE, WHEN THE PATIENT ARRIVED TO POST ANESTHESIA CARE UNIT (PACU), THE PATIENT CODED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS DISCOVERED THAT THE PATIENT HAD A RETROPERITONEAL BLEED CAUSED BY TWO HOLES AT THE INFRARENAL INFERIOR VENA CAVA (IVC). ONE OF THE HOLES WAS SUTURED AND THE SECOND HOLE (WHICH WAS APPROXIMATELY A CENTIMETER LACERATION) WAS UNABLE TO BE REPAIRED. THE PATIENT HAD SEVERAL UNKNOWN COMORBIDITIES THAT CONTRIBUTED TO HIS CONDITION. THE PHYSICIAN EXPERIENCED NO DIFFICULTIES WITH ACCESS, REMOVAL, OR CLOSURE WITH A NON-BOSTON SCIENTIFIC PRE-CLOSE SUTURE DEVICE. IN THE PHYSICIAN OPINION, THE LACERATIONS ARE FROM AN UNKNOWN CAUSE. THE PATIENT EXPERIENCED NO HEMODYNAMIC CHANGES DURING THE PROCEDURE AND ONLY EXHIBITED ISSUES AFTER LEAVING THE PROCEDURE ROOM. THE OFFICIAL CAUSE OF DEATH DETERMINED WAS HEMORRHAGIC SHOCK. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISCARDED). THE PHYSICIAN MENTIONED HE DID NOT THINK THERE WAS AN ISSUE WITH ACCESS, TRANSSEPTAL, OR ANYTHING AT THE BEGINNING OF THE PROCEDURE.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC), A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND A 35MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED USING A WATCHMAN TRUSTEER SHEATH. THERE WERE NO RECAPTURES OF THE CLOSURE DEVICE NEEDED AND WAS IMPLANTED AFTER IT MET RELEASE CRITERIA. THE PROCEDURE WAS COMPLETED WITH NO KNOWN ISSUES. AFTER THE PROCEDURE, WHEN THE PATIENT ARRIVED TO POST ANESTHESIA CARE UNIT (PACU), THE PATIENT CODED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS DISCOVERED THAT THE PATIENT HAD A RETROPERITONEAL BLEED CAUSED BY TWO HOLES AT THE INFRARENAL INFERIOR VENA CAVA (IVC). ONE OF THE HOLES WAS SUTURED AND THE SECOND HOLE (WHICH WAS APPROXIMATELY A CENTIMETER LACERATION) WAS UNABLE TO BE REPAIRED. THE PATIENT HAD SEVERAL UNKNOWN COMORBIDITIES THAT CONTRIBUTED TO HIS CONDITION. THE PHYSICIAN EXPERIENCED NO DIFFICULTIES WITH ACCESS, REMOVAL, OR CLOSURE WITH A NON-BOSTON SCIENTIFIC PRE-CLOSE SUTURE DEVICE. IN THE PHYSICIAN OPINION, THE LACERATIONS ARE FROM AN UNKNOWN CAUSE. THE PATIENT EXPERIENCED NO HEMODYNAMIC CHANGES DURING THE PROCEDURE AND ONLY EXHIBITED ISSUES AFTER LEAVING THE PROCEDURE ROOM. THE OFFICIAL CAUSE OF DEATH DETERMINED WAS HEMORRHAGIC SHOCK. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISCARDED). THE PHYSICIAN MENTIONED HE DID NOT THINK THERE WAS AN ISSUE WITH ACCESS, TRANSSEPTAL, OR ANYTHING AT THE BEGINNING OF THE PROCEDURE. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS RF WIRE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: CORRECTION TO THE INITIAL MDR IN BLOCK(S) F10 AND H6 (PATIENT CODES), IN SECTIONS F - FOR USE BY UF/IMPORTER AND H - DEVICE MANUFACTURERS ONLY. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS RF WIRE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483773&pc=MKJ | 3004123209-2025-00049 | 2025-02-19 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-02-27 05:00 | MKJ | UNKNOWN_MEDICAL - BELFAST_PRODUCT | Failure to Sense | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED.Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: THE CUSTOMER'S DEVICE WAS RETURNED FOR THE EVALUATION. ON THE DATE OF THE REPORTED EVENT, THE DEVICE RECORDED 3 MANUAL POWER-ONS OF UP TO 7 MINUTES AND 15 SECONDS DURATION. DURING THE FIRST 2 POWER ONS, THE DEVICE ISSUED THE NORMAL ADVISORY SPEECH PROMPTS TO APPLY THE PADS TO PATIENT. THE DEVICE CONTINUED TO LOG ¿APPLY PADS¿ PROMPTS THROUGHOUT THE EVENT, AS AN IN-RANGE PATIENT IMPEDANCE WAS NOT DETECTED. DURING THE THIRD POWER-ON, THE DEVICE DETECTED AN IN-RANGE IMPEDANCE. THE DEVICE ANALYZED, IT DETERMINED THE RHYTHM TO BE NON-SHOCKABLE AND ADVISED NO SHOCK THROUGHOUT. NO FAULT WAS FOUND ON THE HDF-3500 OR RETURNED PAD-PAK DURING THE INVESTIGATION. THE DEVICE WAS FOUND TO ACCURATELY MEASURE IMPEDANCE THROUGHOUT THE SPECIFIED RANGE, EVEN AFTER THE STRESS OF ELEVATED TEMPERATURE. NO CONTINUITY ISSUE WAS IDENTIFIED ON THE PATIENT OUTPUT CABLES AND NO FAULT FOUND ON THE POGO-PINS THAT WOULD HAVE RESULTED IN AN IMPEDANCE DETECTION ISSUE. THE DEVICE DELIVERED THE FULL SHOCK THERAPY SEQUENCE WITHOUT FAULT USING THE RETURNED PAD-PAK AND THE DEVICE RECORDED ECG DATA ACCURATELY WITHOUT NOISE OR OTHER ABNORMALITIES. INFORMATION FROM THE DEVICE MEMORY INDICATES THAT THE DEVICE HAD SUCCESSFULLY DETECTED IN-RANGE IMPEDANCES ON TWO OCCASIONS PRIOR TO THE REPORTED EVENT. IT ALSO SUCCESSFULLY DETECTED IN RANGE IMPEDANCE BETWEEN THE (B)(6) 2025 AND (B)(6) 2025, WHICH IS AFTER THE REPORTED EVENT. THE INVESTIGATION CAN ONLY SUGGEST THAT THE REPORTED COMPLAINT WAS DUE TO SITUATIONAL FACTORS. E.G., A PAD-PAK FITMENT ISSUE OR INCORRECT PLACEMENT OF PADS OR POORLY ADHERED PADS ON THE PATIENT, AS NO FAULT WAS FOUND WITH THE RETURNED HDF-3500. POTENTIAL USE OF DEVICE ERROR. Manufacturer Narrative: HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. THE HDF-3500 DEVICE IS NOT AVAILABLE FOR DISTRIBUTION IN THE UNITED STATES BUT IS SIMILAR TO THE SAM 350P AND 450P WHICH IS DISTRIBUTED IN THE UNITED STATES. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. NO RESPONSE HAS BEEN RECEIVED FROM THE CUSTOMER. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21483809&pc=DSQ | 2916596-2025-01270 | 2025-02-10 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21484488&pc=QFG | 3004464228-2025-08406 | 2025-02-12 05:00 | Death | INSULET CORPORATION | 2025-02-27 05:00 | QFG | OMNIPOD 5 POD | Insufficient Information | Insufficient Information | K203768 |  | 1 | Event Description: IT WAS REPORTED VIA A SOCIAL MEDIA POST BY THE PATIENT'S FRIEND THAT THE PATIENT'S PUMP MALFUNCTIONED AND CONTRIBUTED TO THE PATIENT PASSING AWAY. NO OTHER INFORMATION WAS PROVIDED AS TO WHAT OMNIPOD DEVICE THE PATIENT WAS USING, CIRCUMSTANCES RELATED TO THE PATIENT'S DEATH, TREATMENT OR MEDICAL INTERVENTIONS REQUIRED. THE PATIENT'S FRIEND WAS ASKED TO CONTACT INSULET WITH FURTHER INFORMATION. IF ADDITIONAL INFORMATION BECOMES AVAILABLE A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED PATIENT DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485148&pc=ODG | 3002808148-2025-03205 | 2024-06-26 04:00 | Death | SHIRAKAWA OLYMPUS CO., LTD. | 2025-02-27 05:00 | ODG | EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Sepsis | K093395 |  | 1 | Event Description: OLYMPUS REVIEWED THE FOLLOWING LITERATURE TITLED "ENDOSCOPIC ULTRASOUND-GUIDED ENTERO-COLOSTOMY WITH LUMEN-APPOSING METAL STENT AS A RESCUE TREATMENT FOR MALIGNANT INTESTINAL OCCLUSION: A MULTICENTER STUDY." BACKGROUND. SURGERY IS THE FIRST-CHOICE TREATMENT FOR MALIGNANT INTESTINAL OBSTRUCTION (MIO); HOWEVER, MANY PATIENTS ARE DEEMED UNFIT FOR SURGERY. ENDOSCOPIC ULTRASOUND-GUIDED ENTERO-COLOSTOMY (EUS-EC) WITH A LUMEN-APPOSING METAL STENT (LAMS) COULD REPRESENT A NEW TREATMENT OPTION. METHODS. CONSECUTIVE PATIENTS UNDERGOING EUS-EC FOR MIO FROM NOVEMBER 2021 TO SEPTEMBER 2023 AT FOUR EUROPEAN TERTIARY REFERRAL CENTERS WERE RETROSPECTIVELY ENROLLED. MULTIDISCIPLINARY MEETINGS DETERMINED WHETHER PATIENTS WERE UNSUITABLE FOR SURGERY OR COLONIC STENT PLACEMENT, OR REFUSED SURGERY. THE PRIMARY OUTCOME WAS TECHNICAL SUCCESS OF EUS-EC AND SECONDARY OUTCOMES WERE CLINICAL OUTCOME, SAFETY, AND HOSPITAL STAY. RESULTS. 12 PATIENTS WERE ENROLLED (MEDIAN AGE 72.5 [RANGE 42¿85] YEARS; 58.3% FEMALE). COLONIC ADENOCARCINOMA WAS THE PRIMARY TUMOR IN 75.0% OF PATIENTS AND 91.7% HAD STAGE IV DISEASE. TECHNICAL SUCCESS WAS 100%. NO LAMS MISDEPLOYMENT OR OTHER PROCEDURAL ADVERSE EVENTS OCCURRED; THREE PATIENTS (25.0%) HAD SEVERE POST-PROCEDURAL COMPLICATIONS. CLINICAL SUCCESS WAS ACHIEVED IN 10 PATIENTS (83.3%), WITH 5 (50.0%) RESUMING CHEMOTHERAPY AFTER THE PROCEDURE. MEDIAN POST-PROCEDURAL HOSPITAL STAY WAS 9 (1¿20) DAYS AND MEDIAN OVERALL SURVIVAL WAS 47.5 (2¿270) DAYS. CONCLUSIONS. EUS-EC WAS A FEASIBLE TECHNIQUE AND COULD BE CONSIDERED A POSSIBLE ALTERNATIVE TO STANDARD APPROACHES FOR MIO IN HIGHLY SELECTED PATIENTS. TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS: DEATH DUE TO SEVERE SEPSIS (GRADE V) - 2 PATIENTS.Event Description: ADDENDUM: THE STUDY DID NOT SPECIFY WHICH VIDEOSCOPE WAS INVOLVED IN EACH CASE AS A COMPETITOR DEVICE WAS ALSO USED. FURTHERMORE, THE STUDY REPORTED 100% TECHNICAL SUCCESS RATE AND NO PROCEDURAL FAILURES. Manufacturer Narrative: THREE ATTEMPTS WERE PERFORMED TO OBTAIN ADDITIONAL INFORMATION, BUT NO RESPONSE WAS RECEIVED FROM THE AUTHOR. THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE RESULTS OF THE LEGAL MANUFACTURER'S FINAL INVESTIGATION. UPDATED FIELDS: B5, H6. THIS EVENT WAS CONSERVATIVELY REPORTED AS A DEATH AS IT COULD NOT BE COMPLETELY RULED OUT THAT THE OLYMPUS DEVICE MAY HAVE CONTRIBUTED TO THE ADVERSE EVENT. HOWEVER, THERE WERE NO REPORT OF DEVICE MALFUNCTIONS WHICH REDUCED THE LIKELIHOOD OF DEVICE-RELATED HARM. THE DEVICE WAS NOT RETURNED FOR EVALUATION; THEREFORE, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. THE MOST PROBABLE CAUSE WAS NOT ESTABLISHED; THE INVESTIGATION FINDINGS DO NOT LEAD TO A CLEAR CONCLUSION ABOUT THE CAUSE OF THE REPORTED ADVERSE EVENT. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE. Manufacturer Narrative: THIS REPORT IS RELATED TO PATIENT IDENTIFIER 194940 (REPORT FOR THE 1 PATIENT WITH SEPSIS GRADE II DOCUMENTED IN THE LITERATURE). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485535&pc=MIH | 3013164176-2025-02406 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P20004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. DEFECTS AND ADVERSE EVENTS: POSSIBLE DEFECTS AND ADVERSE EVENTS INCLUDE THE FOLLOWING: ADVERSE EVENTS: EMBOLIZATION (MICRO AND MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485619&pc=DYB | 3007284313-2025-03793 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | K160254 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: THE GORE® DRYSEAL FLEX INTRODUCER SHEATH IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. 2. DEVICE DEFECTS AND ADVERSE EVENTS. [OTHER ADVERSE EVENTS]: EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485635&pc=MJN | 3007284313-2025-03791 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MJN | GORE® MOLDING & OCCLUSION BALLOON CATHETER | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | K172567 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: THE GORE® MOLDING & OCCLUSION BALLOON CATHETER IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. 2. DEVICE DEFECTS AND ADVERSE EVENTS [MAJOR ADVERSE EVENTS]: EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485647&pc=MIH | 3013164176-2025-02407 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P020004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: THE GORE® EXCLUDER® AAA ENDOPROSTHESIS IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. <DEFECTS AND ADVERSE EVENTS> POSSIBLE DEFECTS AND ADVERSE EVENTS INCLUDE THE FOLLOWING: ADVERSE EVENTS: EMBOLIZATION (MICRO AND MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485655&pc=MIH | 3007284313-2025-03795 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P020004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: THE GORE® EXCLUDER® AAA ENDOPROSTHESIS IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. <DEFECTS AND ADVERSE EVENTS> POSSIBLE DEFECTS AND ADVERSE EVENTS INCLUDE THE FOLLOWING: ADVERSE EVENTS: EMBOLIZATION (MICRO AND MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485656&pc=DYB | 3007284313-2025-03792 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | K160254 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: THE GORE® DRYSEAL FLEX INTRODUCER SHEATH IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. 2. DEVICE DEFECTS AND ADVERSE EVENTS [OTHER ADVERSE EVENTS]: EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485669&pc=MIH | 3013164176-2025-02408 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P020004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. <DEFECTS AND ADVERSE EVENTS> POSSIBLE DEFECTS AND ADVERSE EVENTS INCLUDE THE FOLLOWING: ADVERSE EVENTS: EMBOLIZATION (MICRO AND MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485744&pc=DYB | 3007284313-2025-03794 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | K160254 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: THE GORE® DRYSEAL FLEX INTRODUCER SHEATH IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. 2. DEVICE DEFECTS AND ADVERSE EVENTS. [OTHER ADVERSE EVENTS]: EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21485805&pc=MIH | 3013164176-2025-02409 | 2025-01-23 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-27 05:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P020004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER®AAA ENDOPROSTHESES. GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (IBE) WAS USED IN THE RIGHT SIDE OF THE PATIENT. A GORE® MOLDING & OCCLUSION BALLOON CATHETER (MOB) AND THREE GORE® DRYSEAL FLEX INTRODUCER SHEATH WERE USED AS ACCESSORY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE DAY AFTER SURGERY, THE PATIENT SUFFERED A CARDIAC ARREST DUE TO HYPERKALEMIA CAUSED BY AN EMBOLISM. DESPITE ATTEMPTS AT RESUSCITATION, THE PATIENT COULD NOT BE RESUSCITATED AND EXPIRED. THE CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HYPERKALEMIA. INFORMATION OF THE AUTOPSY HAS NOT BEEN OBTAINED. THE PHYSICIAN STATED THAT THE EXACT CAUSE OF THE EVENT IS UNKNOWN. HE MENTIONED THAT IT IS POSSIBLE THE CATHETER MANIPULATIONS, SUCH AS EMBOLIZATION OF MULTIPLE VESSELS, INCLUDING THE LUMBAR ARTERIES, ACCESSORY RENAL ARTERIES, AND LEFT INTERNAL ILIAC ARTERY, MAY HAVE CAUSED THE THROMBUS TO SHIFT WITHIN THE ANEURYSM. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: THE GORE® EXCLUDER® AAA ENDOPROSTHESIS IFU WAS REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD. THE FOLLOWING STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. <DEFECTS AND ADVERSE EVENTS> POSSIBLE DEFECTS AND ADVERSE EVENTS INCLUDE THE FOLLOWING: ADVERSE EVENTS: EMBOLIZATION (MICRO AND MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21486039&pc=DSQ | 2916596-2025-01228 | 2025-01-30 05:00 | Death | THORATEC CORPORATION | 2025-02-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THERE WERE NO DEVICE ISSUES AT THE TIME AND THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465918&pc=OZP | 2032227-2025-129042 | 2024-11-15 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-26 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON NOVEMBER 15, 2024 AT HOME. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-399A, MMT-326, MMT-1884. INFORMATION WAS OBTAINED FROM THE REPORTING PERSON AS PER DECEASED REPORTING TROUBLESHOOTING GUIDELINES. THE CAUSE OF DEATH WAS PROBABLE DIABETIC REACTION/METABOLIC ISCHEMIC. THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-399A. NO PRODUCT RETURN IS REQUIRED FOR MMT-326. PRODUCT RETURN IS EXPECTED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21467542&pc=DSQ | 2916596-2025-01173 |  | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿MORTALITY IN RECIPIENTS OF DURABLE LEFT VENTRICULAR ASSIST DEVICES UNDERGOING VENTRICULAR TACHYCARDIA ABLATION¿ THAT HEARTMATE 3 MAY BE ASSOCIATED WITH VENTRICULAR TACHYCARDIA (VT), RIGHT VENTRICULAR FAILURE, TRANSPLANT, DEVICE EXPLANT, AND DEATH. IN THE STUDY VT WAS DEFINED AS SUSTAINED VT (>30 SECONDS) CAPTURED BY IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) INTERROGATION OR DOCUMENTED BY A CARDIOLOGIST. ALL VT ABLATIONS WERE PERFORMED FOR HEMODYNAMICALLY SIGNIFICANT OR SYMPTOMATIC VT THAT PERSISTED DESPITE MEDICAL MANAGEMENT AND ICD REPROGRAMMING. THIS SINGLE-CENTER RETROSPECTIVE COHORT STUDY SOUGHT TO EXAMINE THE IMPACT OF POST-LVAD IMPLANT VT ABLATION ON SURVIVAL. FOLLOWING REMOVAL OF DUPLICATE ENTRIES, ELECTRONIC MEDICAL RECORDS OF 636 PATIENTS AGED 18 YEARS OR OLDER WITH A HISTORY OF LVAD PLACEMENT FROM (B)(6) 2011 TO (B)(6) 2021 (FIGURE 1) WERE REVIEWED. EXCLUSION CRITERIA INCLUDED INITIAL LVAD IMPLANTATION BEFORE (B)(6) 2011 (N =24), LVAD IMPLANTATION AT ANOTHER INSTITUTION (N ¼ 9), LVAD IMPLANTATION THAT FAILED INTRAOPERATIVELY (N = 1), INSUFFICIENT RECORDS (N =16), AND NON-HEARTMATE OR HEARTWARE LVAD IMPLANTS (N = 11). THE STUDY INCLUDED A TOTAL OF 575 PATIENTS, 60 OF WHICH WERE IMPLANTED WITH HEARTMATE 3. OF THE 60 HM3 PATIENTS, 27 EXHIBITED POST-IMPLANT VENTRICULAR TACHYCARDIA. PATIENTS WITH HEARTWARE AND HEARTMATE II WERE ALSO INCLUDED IN THE ANALYSIS OF RESULTS. OUT OF THE 575 OF PATIENTS INCLUDED IN THE STUDY, 84 REQUIRED AN RVAD, 58 REQUIRED A PUMP EXCHANGE, 117 RECEIVED TRANSPLANT, 27 HAD THEIR DEVICE EXPLANTED, 117 REMAINED IMPLANTED AT LAST FOLLOW UP, AND 314 PASSED AWAY. IT WAS NOTED THE PRIMARY OUTCOME WAS ALL-CAUSE MORTALITY. SECONDARY OUTCOMES INCLUDED INCIDENCE OF POST-IMPLANT VT AND OF POST-IMPLANT VT ABLATION. VT AFFECTED 303 (53%) OF LVAD RECIPIENTS IN THE COHORT. OF THE 303 PATIENTS WHO DEVELOPED POST-IMPLANT VT, 47 UNDERWENT VT ABLATION FOR TREATMENT OF REFRACTORY VT. WORSE SURVIVAL WAS SEEN AMONG PATIENTS WITH VT OCCURRENCE WITHIN THE FIRST YEAR OF LVAD IMPLANTATION (HR: 1.62 [95% CI: 1.15-2.27]). AMONG THIS POPULATION, PERFORMANCE OF A CATHETER ABLATION FOR REFRACTORY VT WAS ASSOCIATED WITH IMPROVED SURVIVAL COMPARED WITH THOSE TREATED WITH MEDICAL THERAPY ALONE FOR THE 45 MONTHS FOLLOWING VT ONSET (HR: 0.48 [95% CI: 0.26-0.89]), POSSIBLY DRIVEN BY HIGHER RATES OF EVENTUAL TRANSPLANTATION. THE FINDING THAT SIGNIFICANTLY WORSE SURVIVAL IS ASSOCIATED WITH VT INCIDENCE UP TO 1-YEAR POST-IMPLANTATION CARRIES IMPLICATIONS FOR THE MANAGEMENT AND PREVENTION OF VT IN LVADS, AS WELL AS FOR PROGNOSTICATION OF PATIENTS WHO ARE BEING CONSIDERED FOR CARDIAC TRANSPLANTATION. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2021 AS PATIENTS WERE IMPLANTED FROM (B)(6) 2011 AND (B)(6) 2021. D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. ARTICLE TITLE: ¿MORTALITY IN RECIPIENTS OF DURABLE LEFT VENTRICULAR ASSIST DEVICES UNDERGOING VENTRICULAR TACHYCARDIA ABLATION¿ . AUTHOR INFORMATION: FROM THE DEPARTMENT OF INTERNAL MEDICINE, BAYLOR COLLEGE OF MEDICINE, HOUSTON, TEXAS, USA; DEPARTMENT OF MEDICINE (DIVISION OF CARDIOLOGY), BAYLOR COLLEGE OF MEDICINE, HOUSTON, TEXAS, USA; TEXAS HEART INSTITUTE, HOUSTON, TEXAS, USA; DIVISION OF CARDIOVASCULAR MEDICINE, JACOBS SCHOOL OF MEDICINE AND BIOMEDICAL SCIENCES, UNIVERSITY AT BUFFALO, BUFFALO, NEW YORK, USA; DEPARTMENT OF CARDIOLOGY, MONTEFIORE MEDICAL CENTER, NEW YORK, NEW YORK, USA; AND THE CARDIOVASCULAR RESEARCH INSTITUTE, BAYLOR COLLEGE OF MEDICINE, HOUSTON, TEXAS, USA. LYNCH, P. T., MALOOF, A., BADJATIYA, A., SAFAVI-NAEINI, P., SEGAR, M. W., KIM, J. A., MARASHLY, Q., MOLINA-RAZAVI, J. E., SIMPSON, L., OBERTON, S. B., XIE, L. X., CIVITELLO, A., MATHURIA, N., CHENG, J., RASEKH, A., SAEED, M., RAZAVI, M., NAIR, A., & CHELU, M. G. (2024). MORTALITY IN RECIPIENTS OF DURABLE LEFT VENTRICULAR ASSIST DEVICES UNDERGOING VENTRICULAR TACHYCARDIA ABLATION. JACC. CLINICAL ELECTROPHYSIOLOGY, 10(9), 2049¿2058. HTTPS://DOI.ORG/10.1016/J.JACEP.2024.04.037. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468178&pc=DRA | 2124215-2025-11543 | 2025-02-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-26 05:00 | DRA | FARADRIVE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE TO TREAT AN ATRIAL FIBRILLATION A FARADRIVE WAS SELECTED FOR USE. THE PHYSICIAN COMPLETED TRANSEPTAL PUNCTURE USING ONLY INTRACARDIAC ECHOCARDIOGRAPHY (ICE), THEN THE BLOOD PRESSURE DROPPED. AN EFFUSION WAS NOTICED AND THE PATIENT WAS TAKEN TO SURGERY AS THE DRAIN WAS UNABLE TO CONTROL THE SITUATION. OPEN CHEST SURGERY WAS PERFORMED, BUT A HOLE IN THE LEFT ATRIAL APPENDAGE WAS NOTICED AND THE PATIENT DID NOT MAKE IT. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468316&pc=OZD | 1220648-2025-26335 | 2025-02-01 05:00 | Death | ABIOMED, INC. | 2025-02-26 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THERE WAS CONCERN OF NEW LEFT VENTRICLE THROMBUS, SO BIVAL WAS STARTED. THE PATIENT REFUSED TO ALLOW IMPELLA TIME FOR HEART RECOVERY AND REQUESTED FOR SUPPORT TO BE DISCONTINUED. CARE WAS WITHDRAWN AND SUBSEQUENTLY THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. G1 REPORTING CONTACT EMAIL REVISED ON MANUFACTURER DEVICE REPORT 1220648-2025-26335 IN ACCORDANCE WITH UPDATED PROCEDURES. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468351&pc=DWJ | MW5166814 | 2023-05-05 04:00 | Death | UNK | 2025-02-26 05:00 | DWJ | BAIR HUGGER |  | Pulmonary Embolism; Abscess; Dehydration; Dyspnea; Muscle Weakness; Pain; Renal Failure; Sepsis; Weight Changes; Partial thickness (Second Degree) Burn; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Blister; Swelling/ Edema |  |  | 1 | Event Description: URGENT COMPLAINT REGARDING ADVERSE DRUG REACTIONS AND MEDICAL NEGLIGENCE LED TO DEATH. DEAR FDA I AM FILING THIS FORMAL COMPLAINT REGARDING SEVERE ADVERSE REACTIONS TO MEDICATIONS ADMINISTERED TO MY LATE HUSBAND, (B)(6), AND THE NEGLIGENCE INVOLVED IN HIS TREATMENT AT (B)(6), UNDER THE CARE OF DR. (B)(6) AND (B)(6). THE FOLLOWING CONCERNS REQUIRE IMMEDIATE INVESTIGATION: 1. ADVERSE DRUG REACTIONS AND LACK OF INFORMED CONSENT MY HUSBAND UNDERWENT EMERGENCY SURGERY AFTER MELOXICAM PERFORATED HIS INTESTINE. ON (B)(6) 2023 POST-SURGERY, HE WAS ADMINISTERED INTRAVENOUS ERTAPENEM, HEPARIN, SODIUM CHLORIDE FLUSH, MEROPENEM, VANCOMYCIN, DIPHENHYDRAMINE, AND EPINEPHRINE FOR 10 DAYS AT THE HOSPITAL, WITH A 5-DAY SUPPLY FOR HOME USE. THE RECORDS INDICATE 14 DAYS, THEY BROUGHT ANOTHER 14 DAYS WERE SHIPPED AT HOME MAKING ONE MONTH SUPPLY. LATER WE GOT NOTIFICATION FROM HIS HEALTH INSURANCE THAT THEY CLAIM OF ONE YEAR SUPPLY WAS SENT FOR PAYMENT BUT I TOLD THEM IT WAS 2 WEEKS SENT TO OUR HOME. NO DETAILED EXPLANATION OF POTENTIAL SIDE EFFECTS WAS GIVEN AT THE HOSPITAL BEFORE THE ADMINISTERED THE FIRST DAY. UPON EXPERIENCING SEVERE SYMPTOMS SUCH AS GENERALIZED WEAKNESS, PULMONARY EMBOLISM, DEEP VEIN THROMBOSIS (DVT), SHORTNESS OF BREATH, AND WEIGHT LOSS (FROM 184 LBS TO 110 LBS), WE REPORTED THESE TO (B)(6) THEY FAILED TO ACT, DESPITE ACKNOWLEDGING RESPIRATORY CONCERNS VIA TEXT MESSAGES BY (B)(6) BUT DR. (B)(6) WAS COVERING UP THAT HER ERROR TILL HIS DEATH SHE ENCOUNTERED HIM AT (B)(6) ON (B)(6) 2024 AND (B)(6) ON (B)(6) 2024 THAT HE WAS INTENTIONALLY STARVED AND DEHYDRATED AND OVERDOSED TO DEATH SO WE DONT SUE. 2. PROGRESSIVE DETERIORATION AND MEDICAL NEGLIGENCE ON (B)(6) 2023, MY HUSBAND WAS READMITTED TO THE HOSPITAL DUE TO AN ABSCESS AT THE SURGICAL SITE. IN (B)(6) 2023, HE WAS TAKEN TO THE EMERGENCY ROOM DUE TO EXTREME WEAKNESS AND SWELLING. HE WAS DISCHARGED WITH FIVE DAYS SUPPLY AND ANOTHER TWO WEEK ANTIBIOTIC SUPPLY, DESPITE EVIDENT WORSENING SYMPTOMS SUCH AS LEG SWELLING, BLISTERS, AND DIFFICULTY BREATHING. DR. (B)(6) NOTED THESE SIDE EFFECTS IN HIS ANTIBIOTIC MANAGEMENT RECORDS BUT DR. (B)(6) FAILED TO ADJUST OR DISCONTINUE THE MEDICATION. 3. FINAL HOSPITALIZATION AND FATAL OVERSIGHTON (B)(6) 2024, MY HUSBAND WAS TRANSPORTED BY AMBULANCE DUE TO LEG PAIN. (EMERGENCY MEDICAL TECHNICIAN) EMTS REPORTED HIS CONDITION HAD IMPROVED, YET UPON ARRIVAL, HE WAS TAKEN TO THE ICU, SEDATED, AND PLACED ON A VENTILATOR WITHOUT FOOD OR WATER FOR SIX DAYS. HE WAS ADMINISTERED HIGH DOSES OF HEPARIN, FENTANYL, HYDROCODONE, MORPHINE, AND OTHER OPIOIDS, CAUSING HIS CONDITION TO DETERIORATE FURTHER. HIS LEG SKIN DEVELOPED SEVERE BURNS AND PEELING DUE TO A BAIR HUGGER WARMING DEVICE, YET THIS WAS MISDIAGNOSED AS CELLULITIS. HIS SCROTUM SWELLED SIGNIFICANTLY AND WAS LATER CATEGORIZED AS ACUTE KIDNEY FAILURE. DESPITE EXCESS LUNG FLUID, THE MEDICAL TEAM REFUSED TO DRAIN IT, CITING SERRATIA MARCESCENS-RELATED SEPSIS THAT ALLEGEDLY POSED A RISK TO HOSPITAL EQUIPMENT. HIS HEART BECAME ENLARGED, ULTIMATELY LEADING TO HIS DEATH. 4. HOSPITAL'S DISMISSAL OF CONCERNS I REPORTED THESE SERIOUS MEDICAL CONCERNS TO (B)(6) PATIENT SAFETY OFFICER, (B)(6), IN 2023 AND 2024 BEFORE HIS LAST ADMISSION YET RECEIVED A LETTER STATING THAT THE CASE WAS CLOSED WITHOUT PROPER INVESTIGATION. I INFORMED (B)(6) PHARMACIST (B)(6) AFTER HIS DEATH BUT THEY NEVER GOT RESPONDED AGAIN. REQUEST FOR INVESTIGATION: I URGE THE FDA TO: INVESTIGATE THE ADVERSE REACTIONS AND LACK OF DISCLOSURE REGARDING THE MEDICATIONS LISTED. EXAMINE THE HOSPITAL'S AND (B)(6) FAILURE TO ADDRESS SEVERE COMPLICATIONS. ASSESS THE MISUSE OF SEDATIVES, PAINKILLERS, AND MEDICAL DEVICES LEADING TO FATAL DETERIORATION. ENSURE ACCOUNTABILITY FOR THE MEDICAL NEGLIGENCE INVOLVED. THIS IS A MATTER OF PUBLIC SAFETY, AND I REQUEST A PROMPT RESPONSE. (B)(6). I HAVE MORE PHOTOS BUT UPLOAD ALLOWS 8MB. (B)(6) ALTERED HIS MEDICAL RECORDS CUT ACROSS ALL HOSPITALS AND CONSULTANTS WE HAD BEEN TO COVER UP BECAUSE I CALLED INDEPENDENT AUTOPSY COMPANY BUT THEY COMPROMISED THE AUTOPSY PHYSICIAN. TO COVER HIGH DOSAGE OF FENTANYL AND OTHER FINDINGS. SOME HOSPITALS REFUSED TO PROVIDE MEDICAL RECORD TO AUTOPSY AND INSURANCE COMPANIES OR GIVE PHYSICIAN NOTES TILL DATE. THEY USED A LAWSUIT TO INTIMIDATE ME BECAUSE OF THE PHOTO AND VIDEO EVIDENCE I HAVE. (B)(6) IV COMPANY. NONE BUT FROM SIDE EFFECTS OF THE (B)(6) IV. I HAVE HIS 2 YEARS PHARMACY PRESCRIPTION RECORDS FROM WALMART WALGREENS AND CVS PHARMACY. IF YOU HAVE WHERE I CAN UPLOAD IT NOTIFY ME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468361&pc=DWJ | MW5166815 | 2024-07-15 04:00 | Death | UNK | 2025-02-26 05:00 | DWJ | BAIR HUGGER |  | Cellulitis; Dehydration; Erythema; Fall; Pain; Sepsis; Septic Shock; Superficial (First Degree) Burn; Swelling/ Edema |  |  | 1 | Event Description: (B)(6) NURSE (B)(6) USED A HOT DEVICE TO BURN MY HUSBAND LEG, FOOT, HIP, SWELLING, BLISTERS, REDISH, PEELING AND TAGGED IT CELLULITIS, SERRATIA MARCESCENS LED TO SEPSIS, SEPTIC SHOCK I HAVE PHOTOS AND VIDEOS. HIS BODY AND ROOM WERE NOT CLEANED, HE WAS STARVED AND DEHYDRATED OVERDOSED TO DEATH FOR 6-DAYS HOSPITALIZATION. ON (B)(6) 2024 TOOK MY HUSBAND TO (B)(6) DUE TO WEAKNESS BUT HE CAUGHT COLD AND I CAUGHT FLU AND LEFT BECAUSE THE ROOM WAS SO COLD. HIS ESOPHAGUS WAS DIALTED TOO. HE WAS UNCONDITIONALLY DISCHARGED WITHOUT MEDICAL CONDITIONS ACCORDING TO THEIR MEDICAL RECORDS BUT THEY TRANSPORTED HIM WITH THEIR AMBULANCE TO (B)(6) FOR NO JUST COURSE. HE HAD A SLIP AND FALL IN (B)(6) ON (B)(6) 2024 AND WAS NOT TAKEN CARE OF BUT WAS DISCHARGED ON (B)(6) 2024. HE WENT TO ORTHOPEDIC FOR INJECTION AND REMOVAL OF EFFUSION ON THE KNEE HE FELL. THE KNEE WAS PAINFUL AMBULANCE GAVE HIM FENTANYL 50MCG WHICH HANDLED THE PAIN HE WAS NO LONGER ON PAIN AT THE ER. I DO NOT KNOW WHY THEY TOOK HIM TO ICU TO TORTURE HIM WITH BURN TILL DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468531&pc=BZD | 2518422-2025-102368 | 2025-01-30 05:00 | Death | RESPIRONICS, INC. | 2025-02-26 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING THE DEATH OF A PATIENT USING A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S DAUGHTER REPORTED RETURNING HER MOTHER'S DEVICE; HOWEVER, THE DEVICE IS NOT SHOWING AS RECEIVED IN THE SYSTEM. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468558&pc=OZD | 1220648-2025-26340 | 2025-02-02 05:00 | Death | ABIOMED, INC. | 2025-02-26 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE EXPERIENCED ACCESS SITE BLEEDING, HEMOLYSIS AND SUBSEQUENTLY EXPIRED. THE FOLLOWING DAY AFTER START OF SUPPORT, BLEEDING AT THE ACCESS SITE WAS NOTED AND LOW-PRESSURE BANDAGE WAS APPLIED. IT DOES NOT APPEAR HEPARIN WAS NOT USED IN IV DRIPS AROUND THE TIME OF THE BLEEDING EVENT, AND IT IS, HOWEVER, UNKNOWN IF THE HEPARIN WAS WITHHELD DUE TO THE BLEEDING. ADDITIONALLY, IT WAS NOTED THE PATIENT DEVELOPED HEMOLYSIS, AND THE PERFORMANCE LEVEL WAS REDUCED TO P-4. NOTEDLY, THE PUMP WAS RUNNING WITHOUT ALARMS AND THE POSITION WAS CHECKED. HOWEVER, THE PATIENT'S CONDITION WAS LISTED AS VERY UNWELL. THE FOLLOWING DAY, CHARACTERIZATION OF THE PATIENT'S CONDITION REMAINED THE SAME, AND IT WAS NOTED THE PATIENT ENTERED LIVER AND KIDNEY FAILURE. THE PATIENT WOULD EXPIRE THIS DAY; CARE WAS WITHDRAWN. THE CAUSE OF DEATH IS UNKNOWN, AND THE IMPELLA CP DEVICE ASSOCIATION TO THE EVENT OUTCOME (PATIENT'S DEMISE) IS UNCLEAR. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21468755&pc=MVK | 3008642652-2025-02064 | 2025-01-20 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-26 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 20:33:59 ON (B)(6) 2025. AT 01:27:07 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 01:28:17, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH HB. AT 05:21:11, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 05:21:47, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 05:22:22, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 250 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH NSVT AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 05:22:54, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND NSVT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH NSVT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 05:24:36, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 05:25:14, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 05:30:39 ON (B)(6) 2025. Manufacturer Narrative: INCOMING TESTING OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21469434&pc=OZD | 1220648-2025-26361 | 2025-01-31 05:00 | Death | ABIOMED, INC. | 2025-02-26 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DEVELOPED BLEEDING AT THE IMPELLA ACCESS SITE. HEPARIN WAS HELD AND ONE UNIT OF RED BLOOD CELLS WAS GIVEN. AFTER MUCH DISCUSSION WITH THE FAMILY AND THE PATIENT, A DECISION WAS MADE TO WITHDRAW CARE WITH FAMILY AT HIS SIDE. MEDICAL SUPPORT WAS TURNED OFF AND THE PATIENT PASSED PEACEFULLY WITH FAMILY AT HIS SIDE. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE ANTICOAGULATION MANAGEMENT BECAUSE THE HEPARIN WAS SUPRATHERAPEUTIC AND HEPARIN WAS KEPT ON HOLD FOR 1 HOUR TO CONTROL THE BLEEDING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21469457&pc=QEZ | 1319211-2025-00008 | 2025-01-29 05:00 | Death | ANGIODYNAMICS, INC. | 2025-02-26 05:00 | QEZ | ALPHAVAC | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K240397 |  | 1 | Event Description: AN ANGIODYNAMICS TERRITORY MANAGER REPORTED A PATIENT DEATH OCCURRED, AFTER USE OF AN ALPHAVAC MULTIPURPOSE MECHANICAL ASPIRATOR F18 C85. THE PATIENT ULTIMATELY EXPIRED AS A RESULT OF A 'MASSIVE PULMONARY EMBOLISM.' THE PHYSICIAN STATED THAT THE PATIENT'S EXPIRATION IS IN NO WAY ASSOCIATED WITH THE ALPHAVAC DEVICE BUT DUE TO PATIENT COMORBIDITIES. THE FOLLOWING WAS REPORTED: "35 YEAR OLD MALE PT,WITH FACTOR 5 DEFICIENCY, NON COMPLIANT WITH MEDICATIONS FOR 6 MONTHS, ADMITTED FOR BILATERAL AND SADDLE PE. PT CONSENTED BY PHYSICIAN, TIME OUT STATED, PT PREPPED AND DRAPED IN STERILE FASHION, RIGHT FEMORAL VENOUS ACCESS WITH 26F GORE DRYSEAL. PT ON HEPARIN GTT AND PHYSICIAN GAVE HEPARIN BOLUS, PHYSICIAN STATES "WE DO NOT NEED TO CHECK ACTS". PT PRE PA PRESSURES WERE 65/18/35 O2 98% HR IN 120'S. PHYSICIAN NAVIGATED ALPHAVAC F1885 TO LUNGS AND BEGAN MECHANICAL ASPIRATION. CLOT REMOVED WAS DARK AND BRIGHT RED IN COLOR, GELATINOUS IN TEXTURE. BLOOD WAS FILTERED AND RETURNED BY PHYSICIAN. PHYSICIAN TOOK POST ANGIOGRAMS, CLOT BURDEN WAS DECREASED. ALPHAVAC DEVICE WAS REMOVED FROM BODY AS PHYSICIAN STATED, "I AM FINISHED REMOVING CLOT AND WILL ALLOW HEPARIN TO WORK." PHYSICIAN ASSESSED DEVICE ON BACK TABLE, CLOT PRESENT IN BARREL OF HANDLE AND LEVER WOULD NOT PUSH ALL THE WAY FORWARD WITHOUT MANUAL FORWARD PRESSURE. PHYSICIAN BROKE HANDLE IN HALF TO ASSESS INSIDE OF BARREL. PHYSICIAN INSERTED PRESSURE CATHETER, PTS PRESSURES INCREASED TO 88/18/45. PHYSICIAN TOOK ADDITIONAL ANGIOGRAMS SHOWING INCREASED AMOUNT OF CLOT IN RIGHT LUNG. PT STATED HE WAS ANXIOUS AND WAS TRYING TO MOVE OFF OF THE TABLE. IR LAB NURSE ADMINISTERED IV SEDATION. 5 MINUTES POST SEDATION ADMINISTRATION, PT BECAME UNCONSCIOUS AND CODE BLUE WAS TRIGGERED. CODE TEAM CAME TO EXECUTE CODE. PT EXPIRED. PHYSICIAN STATED, 'THIS WAS NOT RELATED TO ALPHAVAC AND I DID EVERYTHING THAT I COULD.'" Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT EMBOLIZATION AND SUBSEQUENT EXPIRATION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). THERE WERE NO REPORTS OF ALPHAVAC DEVICE MALFUNCTION DURING THE PROCEDURE; THEREFORE, THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE PHYSICIAN STATED THAT THE PATIENT'S EXPIRATION IS IN NO WAY ASSOCIATED WITH THE ALPHAVAC DEVICE BUT DUE TO PATIENT COMORBIDITIES. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE: THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: DISTAL EMBOLIZATION OF THROMBUS. PULMONARY EMBOLISM. CARDIAC ARREST. DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE: (B)(4). Manufacturer Narrative: THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21469633&pc=DSQ | 2916596-2025-00925 | 2023-08-01 04:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH RESEARCH ARTICLE ¿SOCIAL WORK EVALUATION DOMAINS RELATED TO CAREGIVER SUPPORT AND MENTAL HEALTH ARE STRONG PREDICTORS OF RISK OF READMISSION OR DRIVELINE INFECTION AFTER HEARTMATE 3 (HM3)¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH HOSPITAL READMISSION, DRIVELINE INFECTION, TRANSPLANT, DEVICE EXCHANGE, AND DEATH. THIS SINGLE CENTER RETROSPECTIVE STUDY EVALUATED 134 PATIENTS WHO WERE IMPLANTED WITH HM3 BETWEEN JUN2020 AND AUG2023. THE MEDIAN AGE WAS 61, 15% FEMALE, AND 50% INTERMACS 1-2. THE PURPOSE OF THIS STUDY WAS TO ASSESS THE ROLE OF EACH COMPONENT OF A NOVEL SOCIAL WORK (SW) ASSESSMENT TOOL STRUCTURED ACROSS 6 DOMAINS (6D-SW) IN PREDICTING READMISSION OR DRIVELINE INFECTION AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD). 6D-SW DOMAINS INCLUDED: CAREGIVER SUPPORT, SUBSTANCE ABUSE, INSURANCE STATUS, MENTAL HEALTH, HOUSING, AND TRANSPORTATION. THE IMPACT OF EACH DOMAIN ON READMISSION OR DRIVELINE INFECTION 1-YEAR POST-LVAD WAS ASSESSED. AT 1 YEAR, 2 PATIENTS HAD HEART TRANSPLANT, 1 PATIENT HAD DEVICE EXCHANGE, 124 PATIENTS WERE ONGOING LVAD SUPPORT, AND 7 PATIENTS DIED. 50 PATIENTS WERE READMITTED. DRIVELINE INFECTIONS OCCURRED IN 10 PATIENTS. OF ALL 6D-SW DOMAINS, CAREGIVER SUPPORT AND MENTAL HEALTH WERE THE STRONGEST PREDICTORS OF READMISSIONS AND DRIVELINE INFECTIONS 1 YEAR AFTER LVAD, HIGHLIGHTING THE IMPORTANCE OF SW ASSESSMENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT AND DATE OF DEATH HAS BEEN ENTERED AS 01AUG2023 AS DATA WAS COLLECTED BETWEEN JUN2020 AND AUG2023. AUTHOR INFORMATION: LADANYI, A., CAREY, DELUTY, S., SHAHIDI, C., MONDELLINI, G., TIBURCIO, M., MURPHY, B., SAYER, G., KAKU, Y., URIEL, N., TAKEDA, K., PAVOL, M., YUZEFPOLSKAYA, M., & COLOMBO, P. (2024B). SOCIAL WORK EVALUATION DOMAINS RELATED TO CAREGIVER SUPPORT AND MENTAL HEALTH ARE STRONG PREDICTORS OF RISK OF READMISSION OR DRIVELINE INFECTION AFTER HEARTMATE 3 (HM3). THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S23. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.043. COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY; NEW YORK PRESBYTERIAN HOSPITAL, NEW YORK, NY. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21469690&pc=DSQ | 2916596-2025-01197 | 2025-02-08 05:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Chronic Obstructive Pulmonary Disease (COPD); Hypernatremia; Impaired Healing; Hypervolemia; Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ALTHOUGH THE CAUSE OF DEATH WAS UNKNOWN, THE PATIENT HAD FAILURE TO THRIVE, MULTIORGAN FAILURE WITH ENTEROBACTER INFECTION, NON-HEALING WOUNDS, AND RECURRENT VOLUME OVERLOAD/CHRONIC OBSTRUCTIVE PULMONARY DISEASE. THE PUMP WORKING AS INTENDED. AN AUTOPSY NOT PERFORMED.Event Description: THE PATIENT HAD MULTIORGAN FAILURE WITH POOR PROGNOSIS AND A 30-POUND WEIGHT GAIN. THEY DECOMPENSATED REQUIRING INTUBATION. THE PATIENT'S CLINICAL STATUS WAS COMPLEX INVOLVING RECURRING HYPERNATREMIA AND VOLUME OVERLOAD WITH CONGESTION ON CHEST X-RAY (CXR). THEY ALSO HAD HYPOTENSION REQUIRING 2 PRESSORS, AND ENTEROBACTER INFECTION IN THE URINE. THE PATIENT HAD A SEVERE WOUND ON THE LEFT LOWER EXTREMITY (LLE) THAT WAS NOT HEALING. THE FAMILY WAS GATHERED; AND THE DECISION WAS MADE FOR COMFORT MEASURES WITH TERMINAL EXTUBATION AND VAD DISCONTINUATION. THE DEATH WAS NOT DEVICE RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE II PATIENT HANDBOOK, REV. C ARE CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT", ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE HMII LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D4: DEVICE LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. H6: 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21469733&pc=NPT | 2025587-2025-01468 | 2024-02-12 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-26 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 22 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED OF A CEREBRAL INFARCTION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21470111&pc=NPT | 2025587-2025-01469 | 2024-05-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-26 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Dehydration; Low Blood Pressure/ Hypotension; Unspecified Infection; Pulmonary Edema; Septic Shock; Urinary Retention | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS TRANSFERRED FOR CONTINUED REHABILITATION. LOW BLOOD PRESSURE AND LOSS OF APPETITE WERE NOTED AT THE NEW FACILITY LEADING TO A TRANSFER BACK TO THE ORIGINAL HOSPITAL APPROXIMATELY 2 MONTHS AFTER THE VALVE WAS IMPLANTED. DEHYDRATION COMBINED WITH INFECTION CAUSED BY SHOCK WAS SUSPECTED. TREATMENT CONSISTED OF FLUID REPLACEMENT, NOREPINEPHRINE, AND ANTIBIOTIC MEDICATION. FOLLOWING TREATMENT, THE PATIENT WAS ABLE TO COMMUNICATE. ONE WEEK LATER, A DECREASE IN URINE OUTPUT WAS NOTED. THIS LED TO OLIGURIA WITH NO SPONTANEOUS URINE OUTPUT AND A BLOOD PRESSURE THAT REQUIRED HIGH DOSE EPINEPHRINE TO MAINTAINED. PULMONARY EDEMA DEVELOPED DUE TO WORSENING CONGESTION AND THE PATIENT WAS PLACED ON NONINVASIVE POSITIVE PRESSURE VENTILATION TO MAINTAIN BREATHING. DIALYSIS WAS CONSIDERED; HOWEVER, FAMILY DECLINED INVASIVE PROCEDURES. THE PATIENT WENT ON PALLIATIVE CARE. SUBSEQUENTLY, THE PATIENT'S BLOOD PRESSURE GRADUALLY DECREASED. DEATH WAS CONFIRMED APPROXIMATELY 8 DAYS LATER. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AN EXACERBATION OF CHRONIC HEART FAILURE AND SEPTIC SHOCK. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21470254&pc=NPT | 2025587-2025-01470 | 2024-10-02 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-26 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 MONTHS AND 18 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN, HOWEVER, THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) COULD BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH. IT WAS FURTHER REPORTED THAT THE DEATH WAS CAUSED OR CONTRIBUTED TO BY THE PATIENT'S UNDERLYING MEDICAL HISTORY. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21470906&pc=NPT | 2025587-2025-01478 | 2024-02-02 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-26 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Urinary Retention | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 DAY FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PSEUDOANEURYSM WAS OBSERVED IN THE LEFT GROIN. THROMBIN WAS INJECTED AND HEMOSTASIS WAS CONFIRMED. THE FOLLOWING DAY AN ECHOCARDIOGRAM CONFIRMED THAT THE ABNORMAL BLOOD FLOW WHICH HAD BEEN FLOWING INTO THE ANEURYSM FROM THE MAIN TRUNK HAD VANISHED. APPROXIMATELY 3 MONTHS LATER, THE PATIENT DID NOT PRESENT TO THE HOSPITAL FOR A SCHEDULED OUTPATIENT APPOINTMENT. THE PATIENT'S FAMILY WAS CONTACTED AND CONFIRMED THAT THE PATIENT HAD PASSED AWAY THAT MONTH AFTER A SUDDEN DROP IN BLOOD PRESSURE AND URINARY RETENTION. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471189&pc=DSQ | 2916596-2025-01238 | 2025-02-13 05:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471289&pc=DSQ | 2916596-2025-01401 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fever; Low Blood Pressure/ Hypotension; Sepsis; Shock; Tachycardia; Pressure Sores | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO SEPSIS. THE SEPSIS WAS DUE TO PRESSURE SORE. THE PATIENT HAD DEVELOPED A SEVERE PRESSURE SORE, WHICH BECAME INCREASINGLY DIFFICULT TO MANAGE. IN ADDITION, THE PATIENT HAD PREVIOUSLY UNDERGONE PERIPHERAL EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT, AND OVER TIME, THE CANNULATION SITE DEVELOPED AN INFECTION. DESPITE SEVERAL MEDICAL INTERVENTION, THE INFECTION WORSENED AND EVENTUALLY LED TO SEPSIS. THE PATIENT¿S CONDITION CONTINUED TO DETERIORATE, AND UNFORTUNATELY, THEY SUCCUMBED TO SEPTIC SHOCK. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED.Event Description: IT WAS FURTHER REPORTED THAT THE SEPSIS/INFECTION WAS DUE TO A SEVERE PRESSURE SORE LOCATED ON THE ENTIRE BUTTOCKS AREA. THE PATIENT SHOWED CLINICAL SYMPTOMS CONSISTENT WITH SEPSIS, INCLUDING FEVER, INCREASED HEART RATE, LOW BLOOD PRESSURE, AND ALTERED MENTAL STATUS. THE INFECTION WAS TREATED THROUGH A MULTIDISCIPLINARY APPROACH INVOLVING PLASTIC SURGERY FOR DEBRIDEMENT OF THE INFECTED TISSUE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471656&pc=OZD | 1220648-2025-26368 | 2025-02-05 05:00 | Death | ABIOMED, INC. | 2025-02-26 05:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE LEFT VENTRICLE AND AORTA WAVEFORMS WERE SEPARATED. THE STAFF WERE UNABLE TO GO HIGHER THAN P4 WITHOUT SUCTION EVENTS AND THE DEVICE WAS REPOSITIONED SEVERAL TIMES WITH FLUID BOLUS¿S GIVEN. THE FAMILY DECIDED TO PUT PATIENT ON COMFORT CARE AND THE PATIENT EXPIRED SHORTLY AFTER. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE PRODUCT AND DATA WERE NOT RETURNED FOR REVIEW. THE ROOT CAUSE OF LOW FLOW WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. D6A AND D6B WERE INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26368. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471737&pc=NPT | 2025587-2025-01485 | 2025-01-28 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-26 05:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED VIA SOCIAL MEDIA FROM THE PATIENT'S FAMILY MEMBER THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT, THE PATIENT DIED. NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471807&pc=BZD | 2518422-2025-102386 | 2025-02-21 05:00 | Death | RESPIRONICS, INC. | 2025-02-26 05:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: IT WAS REPORTED THAT THE ORIGINAL USER OF THE DREAMSTATION 2 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS THEN GIFTED TO THE DECEASED USER'S NEPHEW WHO REPORTED THAT THE DEVICE WOULD NO LONGER POWER ON WHEN PLUGGED IN. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471871&pc=NPT | 2025587-2025-01488 | 2024-06-09 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-26 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE DEATH WAS NOT RELATED TO THE VALVE OR VALVE IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21471876&pc=OZD | 1220648-2025-26369 | 2025-02-06 05:00 | Death | ABIOMED, INC. | 2025-02-26 05:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT HAD AN ATTEMPTED IMPELLA CP IMPLANT AND THE PATIENT HAD A SMALL SMALL PORCELAIN AORTA WHICH MADE PUSHING THE IMPELLA CP INTO LEFT VENTRICLE IMPOSSIBLE. THE IMPLANT WAS ABORTED AND PROVIDED PROCEED WITH PERCUTANEOUS CORONARY INTERVENTION. SHORTLY AFTER THE PERCUTANEOUS CORONARY INTERVENTION, THE PATIENT WENT INTO A SUSTAINED VENTRICULAR TACHYCARDIA STORM, CARE WAS WITHDRAWN LATER AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, PATIENT HAD BILATERAL TORTUOUS ANEURYSMAL EXTERNAL ILLIACS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472167&pc=DSQ | 2916596-2025-01082 | 2025-02-08 05:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472207&pc=DSQ | 2916596-2025-01141 | 2025-02-11 05:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT NEVER RECOVERED AFTER IMPLANT AND PASSED AWAY. THE DEVICE WAS FUNCTIONING AS EXPECTED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED PATIENT OUTCOME AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE ACCOUNT DECLINED TO PROVIDE ADDITIONAL INFORMATION REGARDING THE PATIENT OUTCOME. BASED ON A REVIEW OF THE PATIENT¿S AVAILABLE RECORDS AND THE REPORTED PATIENT OUTCOME, AN AUTOPSY WAS NOT PERFORMED. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472290&pc=DRE | 3007284006-2025-00037 | 2025-02-06 05:00 | Death | SPECTRANETICS CORPORATION | 2025-02-26 05:00 | DRE | TIGHTRAIL ROTATING DILATOR SHEATH | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Pericardial Effusion | K150360 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD, A RIGHT ATRIAL (RA) LEAD, AND A LEFT VENTRICULAR (LV) LEAD DUE TO BACTEREMIA. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH OF THE LEADS TO PROVIDE TRACTION. USING A SPECTRANETICS TIGHTRAIL SUB-C ROTATING DILATOR SHEATH, SPECTRANETICS GLIDELIGHT LASER SHEATH, AND 13F TIGHTRAIL (LONG), THE RV AND RA LEADS WERE REMOVED SUCCESSFULLY. THEN, THE 13F TIGHTRAIL (LONG) WAS USED ON THE LV LEAD AND ADVANCEMENT WAS MADE DOWN TO THE CORONARY SINUS (CS) OSTIUM, WHERE TRACTION WAS APPLIED, AND THE LV LEAD WAS EXTRACTED. HOWEVER, AN EFFUSION WAS NOTED, AND RESCUE EFFORTS BEGAN, INCLUDING PERICARDIOCENTESIS AND STERNOTOMY. A CS PERFORATION WAS DISCOVERED, BUT UNABLE TO BE REPAIRED, AND THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE 13F TIGHTRAIL IN USE IN THE AREA WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5) PATIENT ETHNICITY - NOT AVAILABLE FROM FACILITY. A6) PATIENT RACE - NOT AVAILABLE FROM FACILITY. B6) RELEVANT TESTS/LABORATORY DATA - NOT AVAILABLE FROM FACILITY. H3) THE TIGHTRAIL WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472435&pc=LWP | 2017865-2025-15195 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-26 05:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-15196 . IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV) LEAD WERE EXTRACTED DUE TO PATIENT DEATH. THE CAUSE OF DEATH WAS INDICATED AS "UNKNOWN". Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO UNKNOWN DEATH. FAILURE EVENT OBSERVED DURING ANALYSIS. THE DEVICE WAS RECEIVED WITH NORMAL TELEMETRY AND OUTPUT. VISUAL INSPECTION OF THE HEADER ATTACHMENT AREA DETECTED A BONDING ANOMALY BETWEEN THE EPOXY HEADER AND TITANIUM CASE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE DEVICE LONGEVITY ASSESSMENT FOUND THE DEVICE WAS OPERATING ABOVE THE ELECTIVE REPLACEMENT INDICATOR (ERI) VOLTAGE CONSISTENT WITH NORMAL USAGE. HYBRID CIRCUITRY WAS TESTED, AND THE RESULTS INDICATED NORMAL CURRENT DRAIN. A MANUFACTURING ANOMALY MAY HAVE OCCURRED, WHICH RESULTED IN THE HEADER BONDING ANOMALY. THE DEVICE IS INCLUDED IN THE ASSURITY AND ENDURITY PACEMAKERS HEADER ANOMALY ADVISORY ISSUED BY ABBOTT ON 15 MARCH 2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472437&pc=NVN | 2017865-2025-15196 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-26 05:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-15195. IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV) LEAD WERE EXTRACTED DUE TO PATIENT DEATH. THE CAUSE OF DEATH WAS INDICATED AS "UNKNOWN". Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTS DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472459&pc=OUT | 2029214-2025-00514 |  | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-26 05:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Collapse; Activation, Positioning or Separation Problem; Activation Failure | Obstruction/Occlusion | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT DURING THE PROCEDURE THE PIPELINES WERE FLAT IN THE PROXIMAL SECTION. SIX MONTHS POST OPERATIVE THE PIPELINE COLLAPSED OCCLUDING THE VESSEL AND THE PATIENT PASSED AWAY. THE PATIENT WAS UNDERGOING SURGERY FOR TREATMENT OF A FUSIFORM, UNRUPTURED ANEURYSM AT THE VERTEBRAL BASILAR JUNCTION WITH A MAX DIA METER OF 20 MM AND A 30 MM NECK DIAMETER. LANDING ZONE: DISTAL: 5 MM AND PROXIMAL: 5.4 MM. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS MODERATE. DUAL ANTIPLATELET TREATMENT WAS ADMINISTERED. THE ANGIOGRAPHIC RESULT POST PROCEDURE SHOWED THE DEVICES WERE OPEN AND THERE WAS STASIS IN THE ANEURYSM. IT WAS REPORTED THAT AT TIME OF IMPLANT IT WAS DECIDED, AFTER DISCUSSION OF OFF-LABEL USAGE, THAT THE PHYSICIAN WOULD TELESCOPE TWO PIPELINE VANTAGE DEVICES 5.5 X40 (B672476) AND A 5X 40 (B677757). STARTING WITH THE 5.5 MM DEVICE. DEVICE WAS PLACED WELL AND COMPLETELY OPEN AFTER MANIPULATING THE DEVICE TO RESOLVE SOME FLATTENING IN THE PROXIMAL 1/3 OF THE DEVICE. HE THEN PROCEEDED TO PLACE THE SECOND DEVICE 5X40 INSIDE OF THE 5.5X40 DEVICE. AGAIN, FLATTENING WAS OBSERVED AT THE PROXIMAL 1.3 OF THE DEVICES. THE PHYSICIAN DECIDED TO DEPLOY THE DEVICE AS IS AND USE A HYPERGLIDE 5X15 MM (B525269) TO OPEN THE FLATTENED PORTION. DEVICE OPENED AFTER BALLOONING AND WAS WELL APPOSED. THE DOCTOR THEN DID A 4 WEEK FOLLOW UP AND BOTH DEVICE WERE WELL APPOSED AND ANEURYSM FILLING HAD SIGNIFICANTLY DECREASED. AT 6 MONTHS THE PATIENT RETURN WITH UNSPECIFIED SYMPTOMS, UPON ANGIO THE INNER DEVICE HAD COLLAPSED TO APPROXIMATELY 1/3 OF ITS DIAMETER IN THE MIDDLE MOST PORTION OF THE DEVICE. THE PATIENT ENDED UP OCCLUDING THE VESSEL AND PASSING AWAY.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PHYSICIAN DID NOT FEEL THAT THE COMPLICATION WAS RELATED TO THE PIPELINE VANTAGE DEVICE. INSTEAD, THE PHYSICIAN FELT LIKE THE PATIENT WAS ABLE TO LIVE FOR ADDITIONAL TIME BECAUSE OF THE PROCEDURE THAT WAS PERFORMED. THE PATIENT HAD UNDERGONE TREATMENT OF A FUSIFORM DISSECTING GIANT VERTEBRAL ARTERY ANEURYSM THAT WAS VERY CHALLENGING TO TREAT WHICH WAS KNOWN OFF-LABEL USE OF THE PIPELINE VANTAGE. THE IMPLANTATION OF THE PIPELINE WENT EVEN BETTER THAN EXPECTED, WITH ONLY MINOR CHALLENGES. THE PATIENT'S CONDITION IMPROVED OVER TIME POST-OPERATIVELY, HOWEVER, AT 6 MONTH FOLLOW UP THE PATIENT REPORTED SYMPTOMS AND UPON ANGIOGRAPHIC IMAGING, THROMBOSIS AROUND THE IMPLANT AND NARROWING OF THE DEVICES WAS IDENTIFIED. THE PATIENT EXPIRED AROUND 9 MONTHS POST-IMPLANT. THE SOURCE OF THE IN-STENT STENOSIS WAS UNCLEAR BUT THE PHYSICIAN DID NOT ATTRIBUTE IT TO THE PIPELINE AS IT WAS NOTED THE PATIENT HAD OTHER CONDITIONS, AND THE PHYSICIAN CONSIDERED IT MORE LIKELY THAT THE PIPELINE DEVICE CAPTURED SOME THROMBUS, RATHER THAN GENERATING IT. THE PHYSICIAN EMPHASIZED MULTIPLE TIMES THAT THE PATIENT DEATH WAS, ON THEIR PART, NOT CONSIDERED TO BE RELATED TO THE PIPELINE DEVICES. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: 104-4515; LOT: B525269; IMPLANT DATE N/A; EXPLANT DATE N/A. PRODUCT ID: PED3-027-550-40 (B672476); IMPLANT DATE N/A; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: . MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472463&pc=OUT | 2029214-2025-00515 |  | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-26 05:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Collapse; Activation, Positioning or Separation Problem; Activation Failure | Obstruction/Occlusion | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT DURING THE PROCEDURE THE PIPELINES WERE FLAT IN THE PROXIMAL SECTION. SIX MONTHS POST OPERATIVE THE PIPELINE COLLAPSED OCCLUDING THE VESSEL AND THE PATIENT PASSED AWAY. THE PATIENT WAS UNDERGOING SURGERY FOR TREATMENT OF A FUSIFORM, UNRUPTURED ANEURYSM AT THE VERTEBRAL BASILAR JUNCTION WITH A MAX DIAMETER OF 20 MM AND A 30 MM NECK DIAMETER. LANDING ZONE: DISTAL: 5 MM AND PROXIMAL: 5.4 MM. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS MODERATE. DUAL ANTIPLATELET TREATMENT WAS ADMINISTERED. THE ANGIOGRAPHIC RESULT POST PROCEDURE SHOWED THE DEVICES WERE OPEN AND THERE WAS STASIS IN THE ANEURYSM. IT WAS REPORTED THAT AT TIME OF IMPLANT IT WAS DECIDED, AFTER DISCUSSION OF OFF-LABEL USAGE, THAT THE PHYSICIAN WOULD TELESCOPE TWO PIPELINE VANTAGE DEVICES 5.5 X40 (B672476) AND A 5X 40 (B677757). STARTING WITH THE 5.5 MM DEVICE. DEVICE WAS PLACED WELL AND COMPLETELY OPEN AFTER MANIPULATING THE DEVICE TO RESOLVE SOME FLATTENING IN THE PROXIMAL 1/3 OF THE DEVICE. HE THEN PROCEEDED TO PLACE THE SECOND DEVICE 5X40 INSIDE OF THE 5.5X40 DEVICE. AGAIN FLATTENING WAS OBSERVED AT THE PROXIMAL 1.3 OF THE DEVICE. THE PHYSICIAN DECIDED TO DEPLOY THE DEVICE AS IS AND USE A HYPERGLIDE 5X15 MM (B525269) TO OPEN THE FLATTENED PORTION. DEVICE OPENED AFTER BALLOONING AND WAS WELL APPOSED. THE DOCTOR THEN DID A 4 WEEK FOLLOW UP AND BOTH DEVICE WERE WELL APPOSED AND ANEURYSM FILLING HAD SIGNIFICANTLY DECREASED. AT 6 MONTHS THE PATIENT RETURN WITH UNSPECIFIED SYMPTOMS, UPON ANGIO THE INNER DEVICE HAD COLLAPSED TO APPROXIMATELY 1/3 OF ITS DIAMETER IN THE MIDDLE MOST PORTION OF THE DEVICE. THE PATIENT ENDED UP OCCLUDING THE VESSEL AND PASSING AWAY.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PHYSICIAN DID NOT FEEL THAT THE COMPLICATION WAS RELATED TO THE PIPELINE VANTAGE DEVICE. INSTEAD, THE PHYSICIAN FELT LIKE THE PATIENT WAS ABLE TO LIVE FOR ADDITIONAL TIME BECAUSE OF THE PROCEDURE THAT WAS PERFORMED. THE PATIENT HAD UNDERGONE TREATMENT OF A FUSIFORM DISSECTING GIANT VERTEBRAL ARTERY ANEURYSM THAT WAS VERY CHALLENGING TO TREAT WHICH WAS KNOWN OFF-LABEL USE OF THE PIPELINE VANTAGE. THE IMPLANTATION OF THE PIPELINE WENT EVEN BETTER THAN EXPECTED, WITH ONLY MINOR CHALLENGES. THE PATIENT'S CONDITION IMPROVED OVER TIME POST-OPERATIVELY, HOWEVER, AT 6 MONTH FOLLOW UP THE PATIENT REPORTED SYMPTOMS AND UPON ANGIOGRAPHIC IMAGING, THROMBOSIS AROUND THE IMPLANT AND NARROWING OF THE DEVICES WAS IDENTIFIED. THE PATIENT EXPIRED AROUND 9 MONTHS POST-IMPLANT. THE SOURCE OF THE IN-STENT STENOSIS WAS UNCLEAR BUT THE PHYSICIAN DID NOT ATTRIBUTE IT TO THE PIPELINE AS IT WAS NOTED THE PATIENT HAD OTHER CONDITIONS, AND THE PHYSICIAN CONSIDERED IT MORE LIKELY THAT THE PIPELINE DEVICE CAPTURED SOME THROMBUS, RATHER THAN GENERATING IT. THE PHYSICIAN EMPHASIZED MULTIPLE TIMES THAT THE PATIENT DEATH WAS, ON THEIR PART, NOT CONSIDERED TO BE RELATED TO THE PIPELINE DEVICES. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: 104-4515 (LOT: B525269); IMPLANT DATE: N/A; EXPLANT DATE: N/A. PRODUCT ID: PED3-027-500-40 (B677757); IMPLANT DATE: N/A; EXPLANT DATE: N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472848&pc=DSQ | 2916596-2025-01140 | 2025-02-10 05:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS NOTED THAT THE PATIENT WAS SUPPORTED WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND A TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE LVAD WORKED AS EXPECTED. THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: LIMITED PATIENT INFORMATION PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21472999&pc=DSQ | 3007042319-2025-00565 | 2025-02-20 05:00 | Death | HEARTWARE, INC. | 2025-02-26 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Electrical /Electronic Property Problem; Pumping Stopped; Material Integrity Problem; Pumping Problem; Material Split, Cut or Torn | Disseminated Intravascular Coagulation (DIC); Unspecified Infection; Hematuria; Hypervolemia | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT WAS ADMITTED DUE TO A SUICIDE ATTEMPT BY INTENTIONALLY CUTTING THROUGH THE DRIVELINE (DL). THIS CAUSED CABLE/WIRE DAMAGE, THAT WAS ASSOCIATED WITH CONTROLLER ALARMS, HIGH WATT ALARM, ELECTRICAL FAULT ALARMS, VAD DISCONNECT, AND VAD STOP ALARM. IT WAS NOTED THAT THE PATIENT URINE COLOR WAS DARK, AND A BLOOD TEST WILL BE PERFORMED. IT WAS ALSO REPORTED THAT A REVIEW OF LOG FILE DATA REVEALED A MOTOR START EVENT WITH PARAMETER OUTSIDE OF THE TYPICAL RANGE. THE VAD, THE DL AND THE CONTROLLER REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT SERVICING FOR THE DL REPAIR HAS NOT BEEN PERFORMED AND THE DATE HAS NOT BEEN SET YET. IT WAS ALSO REPORTED THAT THE PATIENT RECEIVED LASIX FOR VOLUME OVERLOAD AND ANTIBIOTICS FOR INFECTION. THE LABS SHOWED AN INCREASE IN WHITE BLOOD CELLS DUE TO THE INFECTION AND INTERNATIONAL NORMALIZED RATIO (INR) OF 5.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY DUE TO DISSEMINATED INTRAVASCULAR COAGULATION (DIC). Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE DRIVELINE CABLE ASSOCIATED WITH THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND ONE (1) CONTROLLER ((B)(6)) WERE NOT RETURNED FOR EVALUATION AS THEY REMAIN IN USE. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE EVENT LOG FILE REVEALED A MOTOR START EVENT RECORDED ON (B)(6) 2025 AT 23:28:23 IN WHICH THE MOTOR START PARAMETERS WERE TYPICAL. REVIEW OF THE ALARM AND EVENT LOG FILES REVEALED ONE (1) VAD STOPPED ALARM, ONE (1) ELECTRICAL FAULT ALARM, AND ONE (1) REACTIVATION EVENT LOGGED ALL SIMULTANEOUSLY ON (B)(6) 2025 AT 23:28:03. THE VAD STOPPED ALARM INDICATED THAT THE MOTOR STATORS FAILED. THE ELECTRICAL FAULT ALARM AND REACTIVATION EVENT WERE LOGGED DUE TO AN OVERCURRENT TRIP CONDITION IN THE REAR AND FRONT STATORS. FURTHERMORE, ONE (1) VAD DISCONNECT ALARM WAS LOGGED AT 23:28:04, INDICATING A PHYSICAL DISCONNECTION OF THE DRIVELINE, FOLLOWED BY ONE (1) ELECTRICAL FAULT ALARM LOGGED AT 23:28:16 DUE TO THE FRONT STATOR NOT CONNECTED, CAUSING THE PUMP TO RUN ON A SINGLE STATOR (REAR-STATOR ONLY). A SUCCESSFUL MOTOR START EVENT WAS RECORDED AT 23:28:23 AND ONE (1) HIGH WATT ALARM WAS SUBSEQUENTLY LOGGED AT 23:28:28 DUE TO INCREASED POWER CONSUMPTION REQUIRED TO RUN ON A SINGLE STATOR. AS A RESULT, THE REPORTED EVENTS WERE CONFIRMED. BASED ON A HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF ATYPICAL MOTOR START PARAMETERS MAY BE ATTRIBUTED, BUT NOT LIMITED, TO OUTER SHROUD CONTACT THAT CREATES MORE FRICTION AT THE HOUSING TO IMPELLER INTERFACE DURING PUMP STARTUP. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE MOST LIKELY ROOT CAUSE OF THE DRIVELINE DAMAGE MAY BE ATTRIBUTED TO THE REPORTED CUTTING OF THE DRIVELINE BY THE PATIENT, AS DESCRIBED IN THE EVENT DETAILS. THE MOST LIKELY ROOT CAUSE OF THE ELECTRICAL FAULT AND REACTIVATION AND VAD STOPPED EVENTS CAN BE ATTRIBUTED TO THE FORCE EXERTED TO THE COMPONENTS/CONNECTIONS WHEN THE DRIVELINE CABLE WAS CUT BY THE PATIENT, AS DESCRIBED IN THE EVENT DETAILS, THUS COMPROMISING THE INTEGRITY OF THE OUTER SHEATH WHICH LIKELY CAUSED SEVERAL WIRES TO COME IN CONTACT WITH EACH OTHER TRIGGERING ELECTRICAL FAULT ALARMS DUE TO A SHORT CIRCUIT. PER THE INSTRUCTIONS FOR USE, DEATH, PSYCHIATRIC EPISODE, AND INFECTION ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE, INCLUDING CARE OF THE DRIVELINE EXIT SITE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCTS D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ CONTROLLER D4: SERIAL #: (B)(6). H3: YES. H6: THE CODES PRESENT IN SECTION. H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. ADDITIONAL PRODUCTS: D1: CONTROLLER. D4: (B)(6). H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION TO D4. UNIQUE IDENTIFIER (UDI)#. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - CONTROLLER D4: MODEL#: 1420 / CATALOG#: 1420 / EXPIRATION DATE: 30-NOV-2020 / SERIAL#: (B)(6). D9: NO. H3: NO. H4: MFG DATE: 05-NOV-2019. H5: NO. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21473087&pc=QFG | 3013756811-2025-38135 | 2025-01-31 05:00 | Death | TANDEM DIABETES CARE | 2025-02-26 05:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K201214 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. PER THE REPORTER, THE CUSTOMER WENT TO THE HOSPITAL AND WAS HAVING ISSUES WITH THE PUMP PRIOR TO THE DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO ADDITIONAL INFORMATION COULD BE OBTAINED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21473214&pc=KDJ | 1416980-2025-00839 |  | Death | VANTIVE US HEALTHCARE LLC | 2025-02-26 05:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS. THE CAUSE OF PERITONITIS WAS NOT REPORTED. IT WAS NOT REPORTED IF THE PATIENT WAS HOSPITALIZED FOR THE EVENT. TREATMENT FOR PERITONITIS WAS NOT REPORTED. IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE PERITONITIS. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF THE PERITONITIS WAS RESOLVED PRIOR TO DEATH. THE PATIENT STOPPED USING BAXTER PRODUCT DUE TO DEATH/ PASSED AWAY. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21473333&pc=LWS | 2124215-2025-11972 | 2025-01-12 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-26 05:00 | LWS | RESONATE EL ICD VR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AT HOME ON (B)(6) 2025. THE PATIENT'S SPOUSE FELT THE DEVICE HAD CAUSED/CONTRIBUTED TO THE PATIENT'S DEATH AND REPORTED THAT HE HAD BEEN SHOCKED SEVEN TIMES, THOUGH THE HEALTHCARE PROVIDER STATED HE HAD BEEN SHOCKED TWICE. THE SPOUSE WAS REFERRED BACK TO THE PATIENT'S CLINIC FOR AN INTERROGATION OF THE DEVICE EPISODE. FOLLOW-UP WITH THE FIELD YIELDED NO FURTHER INFORMATION AT THIS TIME. Manufacturer Narrative: DESPITE GOOD FAITH EFFORT, ADDITIONAL INFORMATION REGARDING THIS EVENT WAS NOT OBTAINED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED SHOULD ADDITIONAL INFORMATION BE RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21473365&pc=DQY | 3005168196-2025-00062 | 2025-01-28 05:00 | Death | PENUMBRA, INC. | 2025-02-26 05:00 | DQY | BENCHMARK 6F 071 DELIVERY CATHETER | Fracture; Physical Resistance/Sticking | Insufficient Information | K212838 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A MEDICAL PROCEDURE IN THE BASILAR ARTERY USING A BENCHMARK 6F 071 DELIVERY CATHETER (BENCHMARK). DURING THE PROCEDURE, IT WAS NOTED THAT THE PHYSICIAN WAS WITHDRAWING THE BENCHMARK AND FELT RESISTANCE. THE PHYSICIAN CONTINUED TO PULL THE BENCHMARK CAUSING IT TO STRETCH AND FRACTURE APPROXIMATELY 30CM FROM THE HUB. AN OPEN SURGERY WAS PERFORMED TO REMOVE THE REMAINING PORTION OF THE BENCHMARK FROM THE ILIAC/AORTA/VERTEBRAL ARTERY. AFTER THE SURGERY, THE PATIENT WAS ADMITTED INTO THE INTENSIVE CARE UNIT (ICU). IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2025 DUE TO COMPLICATIONS OF THE PROCEDURE AND A BRAIN HEMORRHAGE. THE PHYSICIAN ATTRIBUTED THE PROCEDURAL COMPLICATIONS RESULTING IN PATIENT EXPIRATION TO THE FRACTURED BENCHMARK. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. WITHOUT THE RETURN OF THE DEVICE, THE ROOT CAUSE OF THE PROBLEM CANNOT BE DETERMINED. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474771&pc=BZD | 2518422-2025-102399 | 2024-10-30 04:00 | Death | RESPIRONICS, INC. | 2025-02-26 05:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO COMMUNICATIONS RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474907&pc=HFA | 9610617-2025-00272 | 2024-01-14 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | MINI STRAIGHT-FORWARD TELESCOPE 0°,2 MM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Premature Labor; Prolapse; Bowel Perforation | H040005 |  | 1 | Event Description: IT WAS REPORTED A 31-YEAR-OLD PREGNANT FEMALE AT 18W5D WITH MONOCHRONIC DIAMNIOTIC TWINS WHO WAS REFERRED FROM OSH FOR TWIN TO TWIN TRANSFUSION SYNDROME. SHE HAD TTTS STAGE III WITH FINDING OF TWIN A THE RECIPIENT HAVING REVERSED A-WAVE DV DIAGNOSED AT 17 WEEKS BUT STILL SIGNIFICANT MEMBRANE SEPARATION, TWIN B WAS NOTED TO HAVE INTERMITTENT ABSENT END-DIASTOLIC FLOW (EDF) OF THE UMBILICAL ARTERY (UA). THE CHORIOAMNIOTIC SEPARATION IMPROVED BY 18W5D. SHE UNDERWENT LASER PHOTOCOAGULATION AND AMNIOREDUCTION ON 2023. SHE HAD AN ANTERIOR PLACENTA AND WAS FOUND TO HAVE 17 VASCULAR ANASTOMOSES INCLUDING 6 SMALL HAIRS, 2 SMALL R-D, 2 MEDIUM R-D, 7 MEDIUM D-R PRE-OPERATIVE FETAL HEART RATES WAS 142 BPM FOR RECIPIENT (TWIN A) AND 164 BPM FOR DONOR (TWIN B), PRE OBJECTIVE MITRAL VALVE PROLAPSE (MVP) WAS 14.2 CM. POST-OP MVP WAS 14.3 CM FOR THE RECIPIENT. POST OPERATIVE FETAL HEART RATES 142 FOR THE RECIPIENT (TWIN A) AND 148 BPM FOR DONOR (TWIN B). PRE-OPERATIVE MIDDLE CEREBRAL ARTERY (MCA) FOR RECIPIENT TWIN A WAS 19.9 CM/S AND FOR DONOR TWIN B WAS 14.0 CM/S AND POST OPERATIVE MCA FOR RECIPIENT TWIN A WAS 19.5 CM/S AND DONOR TWIN B WAS 14.0 CM/S. POST-OPERATIVELY, WHEN SHE WAS 22 22W4D SHE WAS FOUND TO HAVE CERVICAL SHORTENING AND UNDERWENT A RESCUE MCDONALD CERVICAL CERCLAGE. AT 24W 2024 SHE SPONTANEOUSLY LABORED AND DELIVERED BIRTH TWINS. TWIN B FORMER DONOR PASSED ON DOL7. SECONDARY TO GRADE IV IVH BILATERALLY AND BOWEL PERORATION. TWIN A FORMER RECIPIENT AND AT THE 1 YEAR FOLLOW UP IS DOLING WELL. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFEREBCE MDR: 9610617-2025-00273. Manufacturer Narrative: ADDITIONAL INFORMATION: A PRODUCT INVESTIGATION COULD NOT BE CARRIED OUT BECAUSE THE KARL STORZ PRODUCT INVOLVED WAS NOT RETURNED. THE REPORTING PHYSICIAN WAS CONTACTED AND CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED DID NOT EXHIBIT ANY MALFUNCTION AND WERE NOT CONSIDERED CONTRIBUTORY TO THE EVENT. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO INDICATION THAT THE REFERENCED KARL STORZ PRODUCTS MALFUNCTIONED OR PLAYED A ROLE IN THE REPORTED COMPLICATIONS OR THE PATIENT'S OUTCOME. SHOULD THE PRODUCTS BE RETURNED IN THE FUTURE, A DETAILED INSPECTION WILL BE CONDUCTED, AND THE INVESTIGATION WILL BE UPDATED ACCORDINGLY. CURRENTLY, NO CORRECTIVE ACTION IS DEEMED NECESSARY AS THERE IS NO CRITICAL AND/OR SYSTEMATIC DEVIATION FOUND. AS A CONTINUED ACTION, WE WILL TRACK AND TREND RELATED COMPLAINTS ACCORDING TO OUR TRENDING PROCEDURES. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4). Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474928&pc=HFA | 2020550-2025-00268 | 2023-03-03 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | HOPKINS TELESCOPE 0°, 2 MM, 26 CM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose |  |  | 1 | Event Description: IT WAS REPORTED TO KARL STORZ THAT A PATIENT'S MEMBRANE WAS RUPTURED AFTER SURGERY. THE PATIENT WAS PREGNANT WITH TWINS; HOWEVER, THE PATIENT HAD TWIN-TO-TWIN TRANFUSION SYNDROME (TTTS) STAGE 3, A RARE CONDITION WHERE ONE TWIN RECEIVES TOO MUCH BLOOD WHILE THE OTHER TWIN RECEIVES TOO LITTLE BLOOD. ONE TWIN HAD REVERSED BLOOD FLOW IN THE UMBILICAL ARTERY, WHILE THE OTHER TWIN HAD INTERMITTENT ABSENT BLOOD FLOW IN THE UMBILICAL ARTERY. THE DIFFERENCE IN GROWTH BETWEEN THE TWINS WAS 18%. THE PATIENT ALSO HAD LARGE SUBCHORIONIC HEMATOMA. THE PATIENT UNDERWENT FETOSCOPIC LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS FOR TTTS. THE SURGERY WAS SUCCESSFUL, BUT THERE WERE RISKS INVOLVED, SUCH AS MEMBRANE SEPARATION THAT COULD LEAD TO PREMATURE RUPTURE OF MEMBRANES AFTER THE SURGERY AND PLACENTAL ABRUPTION. WITHOUT RECEIVING FETOSCOPIC LASER PHOTOCOAGULATION SURGERY, THE PATIENT WAS AT RISK OF LOSING ONE OR BOTH TWINS. AFTER THE SURGERY, THE PATIENT WAS FOUND TO HAVE A SIGNIFICANT AMOUNT OF CHORIOAMNIOTIC SEPARATION BUT REMAINED STABLE. LATER THAT AFTERNOON, THE PATIENT RUPTURED HER MEMBRANES BUT REMAINED STABLE. THE PATIENT WAS RESEALED AND RECEIVED LATENCY ANTIBIOTICS. THE TWINS RECOVERED FROM SURGERY WITH CONCORDANT FLUID, BLADDER SEEN, AND NORMAL DOPPLERS. ON THE 22ND WEEK AND 1 DAY OF PREGNANCY, THE PATIENT RECEIVED C-SECTION TO DELIVER THE TWINS BECAUSE THE PATIENT HAD SEVERE PREECLAMPSIA ON THE DAYS PRIOR. BOTH TWINS ARE IN THE NICU AND STABLE. IT WAS CONFIRMED BY THE TREATING PHYSICIAN THAT KARL STORZ DEVICES WERE DID NOT MALFUNCTIONED AND DID NOT CONTRIBUTE TO THE COMPLICATION. THE PHYSICIAN BELIEVED THAT THE COMPLICATION WAS DUE TO PATIENT'S PREVIOUS SUBCHORIONIC HEMATOMA ADDITIONAL INFORMATION RECEIVED ON FEBRUARY 14.2025 VIA MW5165994: TWIN B, FORMER DONOR WAS GROWTH RESTRICTED AND BORN AT 0.4KG. ON DAY OF LIFE 11, HE DEVELOPED NECROTIZING ENTEROCOLITIS WITH PNEUMATOSIS REQUIRING RAPID ESCALATION OF MANAGEMENT OVER THE COURSE OF THE DAY. THIS INCLUDED ADMINISTRATION OF MULTIPLE BLOOD PRODUCTS, VASOPRESSOR SUPPORT, ANTIBIOTICS, VENTILATION TITRATION. HE DEVELOPED SEVERE PULMONARY HEMORRHAGE THAT EVENING RESULTING IN A CODE EVENT. THE NICU TEAM ALONG WITH PARENTS MADE THE JOINT DECISION TO COMPASSIONATELY EXTUBATE. HE PASSED ON DAY OF LIFE 12. THE DATE OF PASSING WAS (B)(6) 2023 THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474930&pc=HFA | 9610617-2025-00268 | 2023-03-03 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | HOPKINS TELESCOPE 0°, 2 MM, 26 CM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose | H040005 |  | 1 | Event Description: IT WAS REPORTED A 29-YEAR-OLD PREGNANT FEMALE WHO WAS DIAGNOSED WITH STAGE III TTTS (POLY-OLY AND ABSENT BLADDER IN TWIN B (DONOR) WITH ABNORMAL DOPPLERS OF INTERMITTENT ABSENT OF THE END DIASTOLIC FLOW OF THE UA.MVP WAS 8.5CM AND 0.5CM RESPECTIVELY FOR THE RECIPIENT AND DONOR. ALL OTHER DOPPLERS WERE NORMAL. THE PATIENT AND HER FAMILY UNDERWENT EXTENSIVE COUNSELING, AND THEY ELECTED TO PROCEED WITH FETOSCOPIC DIRECTED LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSES. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON 2023. THE CASE WAS UNEVENTFUL. SHE HAD A POSTERIOR PLACENTA AND HAD A TOTAL OF 14 VASCULAR CONNECTION IDENTIFIED: 4 SMALL D-R, 2 MEDIUM D-R. 6 MEDIUM R-D. 1 LARGE D-R AND 1 EXTRA LARGE VV CONNECTION. SHE UNDERWENT AN AMNIOREDUCTION AS WELL. PRE-OPERATIVE FETAL HEART RATES WERE 153 BPM FOR THE DONOR AND 144 BPM FOR THE RECIPIENT, UNFORTUNATELY ON POD #1 IN 2023 SHE WAS TO HAVE A DEMISE OF TWIN A. SHE CONTINUED WITH PREGNANCY AND EVENTUALLY DELIVERED TWIN B AT 37W. AS OF NOW THE BABY IS DOING WELL AT ONE AND HALF YEARS. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: 9610617-2025-00269. Manufacturer Narrative: ADDITIONAL INFORMATION: A PRODUCT INVESTIGATION COULD NOT BE CARRIED OUT BECAUSE THE KARL STORZ PRODUCT INVOLVED WAS NOT RETURNED. THE REPORTING PHYSICIAN WAS CONTACTED AND CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED DID NOT EXHIBIT ANY MALFUNCTION AND WERE NOT CONSIDERED CONTRIBUTORY TO THE EVENT. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO INDICATION THAT THE REFERENCED KARL STORZ PRODUCTS MALFUNCTIONED OR PLAYED A ROLE IN THE REPORTED COMPLICATIONS OR THE PATIENT'S OUTCOME. SHOULD THE PRODUCTS BE RETURNED IN THE FUTURE, A DETAILED INSPECTION WILL BE CONDUCTED, AND THE INVESTIGATION WILL BE UPDATED ACCORDINGLY. CURRENTLY, NO CORRECTIVE ACTION IS DEEMED NECESSARY AS THERE IS NO CRITICAL AND/OR SYSTEMATIC DEVIATION FOUND. AS A CONTINUED ACTION, WE WILL TRACK AND TREND RELATED COMPLAINTS ACCORDING TO OUR TRENDING PROCEDURES. THIS EVENT IS FILED UNDER INTERNAL COMPLAINT (B)(4). Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474932&pc=HFA | 9610617-2025-00275 | 2023-06-05 04:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Abdominal Cramps |  |  | 1 | Event Description: IT WAS REPORTED A 32-YEAR-OLD PREGNANT PATIENT AT 23W6D WITH ESTIMATED DATE OF DELIVERY IN 2023 WAS REFERRED FOR CONCERNS OF TWIN-TO-TWIN TRANSFUSION SYNDROME (TTTS). HER US (OBSTETRIC ULTRASOUND) FINDINGS SHOWED TWIN A (DONOR) HAD A VISIBLE BLADDER BUT HAD PREDOMINANTLY ABSENT END DIASTOLIC FLOW IN THE UMBILICAL ARTERY (ALTHOUGH SOME CYCLES DEMONSTRATE PRESERVED END DIASTOLIC FLOW), WITH OCCASIONAL REVERSED END DIASTOLIC FLOW. THE DV (DUCTUS VENOSUS) AND MCA (MIDDLE CEREBRAL ARTERY) WERE OTHERWISE NORMAL. ALL DOPPLERS IN TWIN B (RECIPIENT) WAS ALSO NORMAL AND BLADDER WAS VISIBLE AS WELL. THEREFORE, THESE FINDINGS WERE CONSISTENT WITH TTTS STAGE III. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS IN 2023 WHEN SHE WAS 24 WEEKS PREGNANT. SHE HAD A POSTERIOR PLACENTA WITH VASCULAR ANASTOMOSIS BETWEEN RECIPIENT AND DONOR CORD INSERTIONS. A TOTAL OF 22 VASCULAR COMMUNICATIONS WERE IDENTIFIED INCLUDING 4 SMALL D-R, 5 SMALL R-D, 6 MEDIUM D-R, 1 MEDIUM R-D, 3 LARGE R-D AND 3 EXTRA LARGE D-R CONNECTION, SHE ALSO UNDERWENT AN AMNIOREDUCTION. HER CERVIX WAS SHORT SO THERE WAS ALSO A CERCLAGE PLACED ON THE SAME OPERATIVE DAY. POST-OPERATIVELY THERE WERE NO COMPLICATIONS. ON POST OPERATIVE DAY 5, HER ULTRASOUND SHOWED TWO VIABLE FETUSES WITH NORMALIZING AMNIOTIC FLUID AND NO CRITICALLY ABNORMAL WAVEFORMS WERE IDENTIFIED. HOWEVER, ON REPEAT ULTRASOUND ON POST OPERATIVE DAY 10), SHE WAS FOUND TO HAVE A DONOR DEMISE. WHEN SHE WAS 29 WEEKS 1 DAY PREGNANT, SHE PRESENTED WITH VAGINAL BLEEDING AND CONTRACTIONS. FORMER RECIPIENT WAS DELIVERED AND AS OF NOW IS DOING WELL AND IS 1.5 YEARS OF AGE. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFEREBCE MDR: 9610617-2025-00274. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: A PRODUCT INVESTIGATION COULD NOT BE CARRIED OUT BECAUSE THE KARL STORZ PRODUCT INVOLVED WAS NOT RETURNED. THE REPORTING PHYSICIAN WAS CONTACTED AND CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED DID NOT EXHIBIT ANY MALFUNCTION AND WERE NOT CONSIDERED CONTRIBUTORY TO THE EVENT. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO INDICATION THAT THE REFERENCED KARL STORZ PRODUCTS MALFUNCTIONED OR PLAYED A ROLE IN THE REPORTED COMPLICATIONS OR THE PATIENT'S OUTCOME. SHOULD THE PRODUCTS BE RETURNED IN THE FUTURE, A DETAILED INSPECTION WILL BE CONDUCTED, AND THE INVESTIGATION WILL BE UPDATED ACCORDINGLY. CURRENTLY, NO CORRECTIVE ACTION IS DEEMED NECESSARY AS THERE IS NO CRITICAL AND/OR SYSTEMATIC DEVIATION FOUND. AS A CONTINUED ACTION, WE WILL TRACK AND TREND RELATED COMPLAINTS ACCORDING TO OUR TRENDING PROCEDURES. THIS EVENT IS FILED UNDER INTERNAL COMPLAINT ID (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474933&pc=HFA | 2020550-2025-00274 | 2023-06-05 04:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | HOPKINS TELESCOPE 0°, 2 MM, 26 CM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Abdominal Cramps |  |  | 1 | Event Description: IT WAS REPORTED A 32-YEAR-OLD PREGNANT PATIENT AT 23W6D WITH ESTIMATED DATE OF DELIVERY IN 2023 WAS REFERRED FOR CONCERNS OF TWIN-TO-TWIN TRANSFUSION SYNDROME (TTTS). HER US (OBSTETRIC ULTRASOUND) FINDINGS SHOWED TWIN A (DONOR) HAD A VISIBLE BLADDER BUT HAD PREDOMINANTLY ABSENT END DIASTOLIC FLOW IN THE UMBILICAL ARTERY (ALTHOUGH SOME CYCLES DEMONSTRATE PRESERVED END DIASTOLIC FLOW), WITH OCCASIONAL REVERSED END DIASTOLIC FLOW. THE DV (DUCTUS VENOSUS) AND MCA (MIDDLE CEREBRAL ARTERY) WERE OTHERWISE NORMAL. ALL DOPPLERS IN TWIN B (RECIPIENT) WAS ALSO NORMAL AND BLADDER WAS VISIBLE AS WELL. THEREFORE, THESE FINDINGS WERE CONSISTENT WITH TTTS STAGE III. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS IN 2023 WHEN SHE WAS 24 WEEKS PREGNANT. SHE HAD A POSTERIOR PLACENTA WITH VASCULAR ANASTOMOSIS BETWEEN RECIPIENT AND DONOR CORD INSERTIONS. A TOTAL OF 22 VASCULAR COMMUNICATIONS WERE IDENTIFIED INCLUDING 4 SMALL D-R, 5 SMALL R-D, 6 MEDIUM D-R, 1 MEDIUM R-D, 3 LARGE R-D AND 3 EXTRA LARGE D-R CONNECTION, SHE ALSO UNDERWENT AN AMNIOREDUCTION. HER CERVIX WAS SHORT SO THERE WAS ALSO A CERCLAGE PLACED ON THE SAME OPERATIVE DAY. POST-OPERATIVELY THERE WERE NO COMPLICATIONS. ON POST OPERATIVE DAY 5, HER ULTRASOUND SHOWED TWO VIABLE FETUSES WITH NORMALIZING AMNIOTIC FLUID AND NO CRITICALLY ABNORMAL WAVEFORMS WERE IDENTIFIED. HOWEVER, ON REPEAT ULTRASOUND ON POST OPERATIVE DAY 10), SHE WAS FOUND TO HAVE A DONOR DEMISE. WHEN SHE WAS 29 WEEKS 1 DAY PREGNANT, SHE PRESENTED WITH VAGINAL BLEEDING AND CONTRACTIONS. FORMER RECIPIENT WAS DELIVERED AND AS OF NOW IS DOING WELL AND IS 1.5 YEARS OF AGE. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFEREBCE MDR: 9610617-2025-00275. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474934&pc=HFA | 2020550-2025-00273 | 2024-01-14 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Premature Labor; Prolapse; Bowel Perforation |  |  | 1 | Event Description: IT WAS REPORTED A 31-YEAR-OLD PREGNANT FEMALE AT 18W5D WITH MONOCHRONIC DIAMNIOTIC TWINS WHO WAS REFERRED FROM OSH FOR TWIN TO TWIN TRANSFUSION SYNDROME. SHE HAD TTTS STAGE III WITH FINDING OF TWIN A THE RECIPIENT HAVING REVERSED A-WAVE DV DIAGNOSED AT 17 WEEKS BUT STILL SIGNIFICANT MEMBRANE SEPARATION, TWIN B WAS NOTED TO HAVE INTERMITTENT ABSENT END-DIASTOLIC FLOW (EDF) OF THE UMBILICAL ARTERY (UA). THE CHORIOAMNIOTIC SEPARATION IMPROVED BY 18W5D. SHE UNDERWENT LASER PHOTOCOAGULATION AND AMNIOREDUCTION ON 2023. SHE HAD AN ANTERIOR PLACENTA AND WAS FOUND TO HAVE 17 VASCULAR ANASTOMOSES INCLUDING 6 SMALL HAIRS, 2 SMALL R-D, 2 MEDIUM R-D, 7 MEDIUM D-R PRE-OPERATIVE FETAL HEART RATES WAS 142 BPM FOR RECIPIENT (TWIN A) AND 164 BPM FOR DONOR (TWIN B), PRE OBJECTIVE MITRAL VALVE PROLAPSE (MVP) WAS 14.2 CM. POST-OP MVP WAS 14.3 CM FOR THE RECIPIENT. POST OPERATIVE FETAL HEART RATES 142 FOR THE RECIPIENT (TWIN A) AND 148 BPM FOR DONOR (TWIN B). PRE-OPERATIVE MIDDLE CEREBRAL ARTERY (MCA) FOR RECIPIENT TWIN A WAS 19.9 CM/S AND FOR DONOR TWIN B WAS 14.0 CM/S AND POST OPERATIVE MCA FOR RECIPIENT TWIN A WAS 19.5 CM/S AND DONOR TWIN B WAS 14.0 CM/S. POST-OPERATIVELY, WHEN SHE WAS 22 22W4D SHE WAS FOUND TO HAVE CERVICAL SHORTENING AND UNDERWENT A RESCUE MCDONALD CERVICAL CERCLAGE. AT 24W 2024 SHE SPONTANEOUSLY LABORED AND DELIVERED BIRTH TWINS. TWIN B FORMER DONOR PASSED ON DOL7. SECONDARY TO GRADE IV IVH BILATERALLY AND BOWEL PERORATION. TWIN A FORMER RECIPIENT AND AT THE 1 YEAR FOLLOW UP IS DOLING WELL. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: 9610617-2025-00272. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474935&pc=HFA | 9610617-2025-00273 | 2024-01-14 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Premature Labor; Prolapse; Bowel Perforation | H040005 |  | 1 | Event Description: IT WAS REPORTED A 31-YEAR-OLD PREGNANT FEMALE AT 18W5D WITH MONOCHRONIC DIAMNIOTIC TWINS WHO WAS REFERRED FROM OSH FOR TWIN TO TWIN TRANSFUSION SYNDROME. SHE HAD TTTS STAGE III WITH FINDING OF TWIN A THE RECIPIENT HAVING REVERSED A-WAVE DV DIAGNOSED AT 17 WEEKS BUT STILL SIGNIFICANT MEMBRANE SEPARATION, TWIN B WAS NOTED TO HAVE INTERMITTENT ABSENT END-DIASTOLIC FLOW (EDF) OF THE UMBILICAL ARTERY (UA). THE CHORIOAMNIOTIC SEPARATION IMPROVED BY 18W5D. SHE UNDERWENT LASER PHOTOCOAGULATION AND AMNIOREDUCTION ON 2023. SHE HAD AN ANTERIOR PLACENTA AND WAS FOUND TO HAVE 17 VASCULAR ANASTOMOSES INCLUDING 6 SMALL HAIRS, 2 SMALL R-D, 2 MEDIUM R-D, 7 MEDIUM D-R PRE-OPERATIVE FETAL HEART RATES WAS 142 BPM FOR RECIPIENT (TWIN A) AND 164 BPM FOR DONOR (TWIN B), PRE OBJECTIVE MITRAL VALVE PROLAPSE (MVP) WAS 14.2 CM. POST-OP MVP WAS 14.3 CM FOR THE RECIPIENT. POST OPERATIVE FETAL HEART RATES 142 FOR THE RECIPIENT (TWIN A) AND 148 BPM FOR DONOR (TWIN B). PRE-OPERATIVE MIDDLE CEREBRAL ARTERY (MCA) FOR RECIPIENT TWIN A WAS 19.9 CM/S AND FOR DONOR TWIN B WAS 14.0 CM/S AND POST OPERATIVE MCA FOR RECIPIENT TWIN A WAS 19.5 CM/S AND DONOR TWIN B WAS 14.0 CM/S. POST-OPERATIVELY, WHEN SHE WAS 22 22W4D SHE WAS FOUND TO HAVE CERVICAL SHORTENING AND UNDERWENT A RESCUE MCDONALD CERVICAL CERCLAGE. AT 24W 2024 SHE SPONTANEOUSLY LABORED AND DELIVERED BIRTH TWINS. TWIN B FORMER DONOR PASSED ON DOL7. SECONDARY TO GRADE IV IVH BILATERALLY AND BOWEL PERORATION. TWIN A FORMER RECIPIENT AND AT THE 1 YEAR FOLLOW UP IS DOLING WELL. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: 9610617-2025-00272. Manufacturer Narrative: ADDITIONAL INFORMATION: A PRODUCT INVESTIGATION COULD NOT BE CARRIED OUT BECAUSE THE KARL STORZ PRODUCT INVOLVED WAS NOT RETURNED. THE REPORTING PHYSICIAN WAS CONTACTED AND CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED DID NOT EXHIBIT ANY MALFUNCTION AND WERE NOT CONSIDERED CONTRIBUTORY TO THE EVENT. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO INDICATION THAT THE REFERENCED KARL STORZ PRODUCTS MALFUNCTIONED OR PLAYED A ROLE IN THE REPORTED COMPLICATIONS OR THE PATIENT'S OUTCOME. SHOULD THE PRODUCTS BE RETURNED IN THE FUTURE, A DETAILED INSPECTION WILL BE CONDUCTED, AND THE INVESTIGATION WILL BE UPDATED ACCORDINGLY. CURRENTLY, NO CORRECTIVE ACTION IS DEEMED NECESSARY AS THERE IS NO CRITICAL AND/OR SYSTEMATIC DEVIATION FOUND. AS A CONTINUED ACTION, WE WILL TRACK AND TREND RELATED COMPLAINTS ACCORDING TO OUR TRENDING PROCEDURES. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4). Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474938&pc=HFA | 9610617-2025-00274 | 2023-06-05 04:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | HOPKINS TELESCOPE 0°, 2 MM, 26 CM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Abdominal Cramps | H040005 |  | 1 | Event Description: IT WAS REPORTED A 32-YEAR-OLD PREGNANT PATIENT AT 23W6D WITH ESTIMATED DATE OF DELIVERY IN 2023 WAS REFERRED FOR CONCERNS OF TWIN-TO-TWIN TRANSFUSION SYNDROME (TTTS). HER US (OBSTETRIC ULTRASOUND) FINDINGS SHOWED TWIN A (DONOR) HAD A VISIBLE BLADDER BUT HAD PREDOMINANTLY ABSENT END DIASTOLIC FLOW IN THE UMBILICAL ARTERY (ALTHOUGH SOME CYCLES DEMONSTRATE PRESERVED END DIASTOLIC FLOW), WITH OCCASIONAL REVERSED END DIASTOLIC FLOW. THE DV (DUCTUS VENOSUS) AND MCA (MIDDLE CEREBRAL ARTERY) WERE OTHERWISE NORMAL. ALL DOPPLERS IN TWIN B (RECIPIENT) WAS ALSO NORMAL AND BLADDER WAS VISIBLE AS WELL. THEREFORE, THESE FINDINGS WERE CONSISTENT WITH TTTS STAGE III. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS IN 2023 WHEN SHE WAS 24 WEEKS PREGNANT. SHE HAD A POSTERIOR PLACENTA WITH VASCULAR ANASTOMOSIS BETWEEN RECIPIENT AND DONOR CORD INSERTIONS. A TOTAL OF 22 VASCULAR COMMUNICATIONS WERE IDENTIFIED INCLUDING 4 SMALL D-R, 5 SMALL R-D, 6 MEDIUM D-R, 1 MEDIUM R-D, 3 LARGE R-D AND 3 EXTRA LARGE D-R CONNECTION, SHE ALSO UNDERWENT AN AMNIOREDUCTION. HER CERVIX WAS SHORT SO THERE WAS ALSO A CERCLAGE PLACED ON THE SAME OPERATIVE DAY. POST-OPERATIVELY THERE WERE NO COMPLICATIONS. ON POST OPERATIVE DAY 5, HER ULTRASOUND SHOWED TWO VIABLE FETUSES WITH NORMALIZING AMNIOTIC FLUID AND NO CRITICALLY ABNORMAL WAVEFORMS WERE IDENTIFIED. HOWEVER, ON REPEAT ULTRASOUND ON POST OPERATIVE DAY 10), SHE WAS FOUND TO HAVE A DONOR DEMISE. WHEN SHE WAS 29 WEEKS 1 DAY PREGNANT, SHE PRESENTED WITH VAGINAL BLEEDING AND CONTRACTIONS. FORMER RECIPIENT WAS DELIVERED AND AS OF NOW IS DOING WELL AND IS 1.5 YEARS OF AGE. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: 9610617-2025-00275. Manufacturer Narrative: ADDITIONAL INFORMATION: A PRODUCT INVESTIGATION COULD NOT BE CARRIED OUT BECAUSE THE KARL STORZ PRODUCT INVOLVED WAS NOT RETURNED. THE REPORTING PHYSICIAN WAS CONTACTED AND CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED DID NOT EXHIBIT ANY MALFUNCTION AND WERE NOT CONSIDERED CONTRIBUTORY TO THE EVENT. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO INDICATION THAT THE REFERENCED KARL STORZ PRODUCTS MALFUNCTIONED OR PLAYED A ROLE IN THE REPORTED COMPLICATIONS OR THE PATIENT'S OUTCOME. SHOULD THE PRODUCTS BE RETURNED IN THE FUTURE, A DETAILED INSPECTION WILL BE CONDUCTED, AND THE INVESTIGATION WILL BE UPDATED ACCORDINGLY. CURRENTLY, NO CORRECTIVE ACTION IS DEEMED NECESSARY AS THERE IS NO CRITICAL AND/OR SYSTEMATIC DEVIATION FOUND. AS A CONTINUED ACTION, WE WILL TRACK AND TREND RELATED COMPLAINTS ACCORDING TO OUR TRENDING PROCEDURES. THIS EVENT IS FILED UNDER INTERNAL COMPLAINT ID (B)(4). Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474939&pc=HFA | 2020550-2025-00275 | 2023-06-05 04:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Abdominal Cramps |  |  | 1 | Event Description: IT WAS REPORTED A 32-YEAR-OLD PREGNANT PATIENT AT 23W6D WITH ESTIMATED DATE OF DELIVERY IN (B)(6) 2023 WAS REFERRED FOR CONCERNS OF TWIN-TO-TWIN TRANSFUSION SYNDROME (TTTS). HER US (OBSTETRIC ULTRASOUND) FINDINGS SHOWED TWIN A (DONOR) HAD A VISIBLE BLADDER BUT HAD PREDOMINANTLY ABSENT END DIASTOLIC FLOW IN THE UMBILICAL ARTERY (ALTHOUGH SOME CYCLES DEMONSTRATE PRESERVED END DIASTOLIC FLOW), WITH OCCASIONAL REVERSED END DIASTOLIC FLOW. THE DV (DUCTUS VENOSUS) AND MCA (MIDDLE CEREBRAL ARTERY) WERE OTHERWISE NORMAL. ALL DOPPLERS IN TWIN B (RECIPIENT) WAS ALSO NORMAL AND BLADDER WAS VISIBLE AS WELL. THEREFORE, THESE FINDINGS WERE CONSISTENT WITH TTTS STAGE III. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS IN (B)(6) 2023 WHEN SHE WAS 24 WEEKS PREGNANT. SHE HAD A POSTERIOR PLACENTA WITH VASCULAR ANASTOMOSIS BETWEEN RECIPIENT AND DONOR CORD INSERTIONS. A TOTAL OF 22 VASCULAR COMMUNICATIONS WERE IDENTIFIED INCLUDING 4 SMALL D-R, 5 SMALL R-D, 6 MEDIUM D-R, 1 MEDIUM R-D, 3 LARGE R-D AND 3 EXTRA LARGE D-R CONNECTION, SHE ALSO UNDERWENT AN AMNIOREDUCTION. HER CERVIX WAS SHORT SO THERE WAS ALSO A CERCLAGE PLACED ON THE SAME OPERATIVE DAY. POST-OPERATIVELY THERE WERE NO COMPLICATIONS. ON POST OPERATIVE DAY 5, HER ULTRASOUND SHOWED TWO VIABLE FETUSES WITH NORMALIZING AMNIOTIC FLUID AND NO CRITICALLY ABNORMAL WAVEFORMS WERE IDENTIFIED. HOWEVER, ON REPEAT ULTRASOUND ON POST OPERATIVE DAY 10), SHE WAS FOUND TO HAVE A DONOR DEMISE. WHEN SHE WAS 29 WEEKS 1 DAY PREGNANT, SHE PRESENTED WITH VAGINAL BLEEDING AND CONTRACTIONS. FORMER RECIPIENT WAS DELIVERED AND AS OF NOW IS DOING WELL AND IS 1.5 YEARS OF AGE. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: (B)(4). Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474943&pc=HFA | 2020550-2025-00269 | 2023-03-03 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose |  |  | 1 | Event Description: IT WAS REPORTED A 29-YEAR-OLD PREGNANT FEMALE WHO WAS DIAGNOSED WITH STAGE III TTTS (POLY-OLY AND ABSENT BLADDER IN TWIN B (DONOR) WITH ABNORMAL DOPPLERS OF INTERMITTENT ABSENT OF THE END DIASTOLIC FLOW OF THE UA.MVP WAS 8.5CM AND 0.5CM RESPECTIVELY FOR THE RECIPIENT AND DONOR. ALL OTHER DOPPLERS WERE NORMAL. THE PATIENT AND HER FAMILY UNDERWENT EXTENSIVE COUNSELING, AND THEY ELECTED TO PROCEED WITH FETOSCOPIC DIRECTED LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSES. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON 2023. THE CASE WAS UNEVENTFUL. SHE HAD A POSTERIOR PLACENTA AND HAD A TOTAL OF 14 VASCULAR CONNECTION IDENTIFIED: 4 SMALL D-R, 2 MEDIUM D-R. 6 MEDIUM R-D. 1 LARGE D-R AND 1 EXTRA LARGE VV CONNECTION. SHE UNDERWENT AN AMNIOREDUCTION AS WELL. PRE-OPERATIVE FETAL HEART RATES WERE 153 BPM FOR THE DONOR AND 144 BPM FOR THE RECIPIENT, UNFORTUNATELY ON POD #1 IN 2023 SHE WAS TO HAVE A DEMISE OF TWIN A. SHE CONTINUED WITH PREGNANCY AND EVENTUALLY DELIVERED TWIN B AT 37W. AS OF NOW THE BABY IS DOING WELL AT ONE AND HALF YEARS. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: 9610617-2025-00268. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4) THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474947&pc=HFA | 2020550-2025-00272 | 2024-01-14 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | MINI STRAIGHT-FORWARD TELESCOPE 0°,2 MM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Incompetent Cervix; Premature Labor; Prolapse; Bowel Perforation |  |  | 1 | Event Description: IT WAS REPORTED A 31-YEAR-OLD PREGNANT FEMALE AT 18W5D WITH MONOCHRONIC DIAMNIOTIC TWINS WHO WAS REFERRED FROM OSH FOR TWIN-TO-TWIN TRANSFUSION SYNDROME. SHE HAD TTTS STAGE III WITH FINDING OF TWIN A THE RECIPIENT HAVING REVERSED A-WAVE DV DIAGNOSED AT 17 WEEKS BUT STILL SIGNIFICANT MEMBRANE SEPARATION, TWIN B WAS NOTED TO HAVE INTERMITTENT ABSENT END-DIASTOLIC FLOW (EDF) OF THE UMBILICAL ARTERY (UA). THE CHORIOAMNIOTIC SEPARATION IMPROVED BY 18W5D. SHE UNDERWENT LASER PHOTOCOAGULATION AND AMNIOREDUCTION IN 2023. SHE HAD AN ANTERIOR PLACENTA AND WAS FOUND TO HAVE 17 VASCULAR ANASTOMOSES INCLUDING 6 SMALL HAIRS, 2 SMALL R-D, 2 MEDIUM R-D, 7 MEDIUM D-R PRE-OPERATIVE FETAL HEART RATES WAS 142 BPM FOR RECIPIENT (TWIN A) AND 164 BPM FOR DONOR (TWIN B), PRE-OBJECTIVE MITRAL VALVE PROLAPSE (MVP) WAS 14.2 CM. POST-OP MVP WAS 14.3 CM FOR THE RECIPIENT. POST OPERATIVE FETAL HEART RATES 142 FOR THE RECIPIENT (TWIN A) AND 148 BPM FOR DONOR (TWIN B). PRE-OPERATIVE MIDDLE CEREBRAL ARTERY (MCA) FOR RECIPIENT TWIN A WAS 19.9 CM/S AND FOR DONOR TWIN B WAS 14.0 CM/S AND POST OPERATIVE MCA FOR RECIPIENT TWIN A WAS 19.5 CM/S AND DONOR TWIN B WAS 14.0 CM/S. POST-OPERATIVELY, WHEN SHE WAS 22 22W4D SHE WAS FOUND TO HAVE CERVICAL SHORTENING AND UNDERWENT A RESCUE MCDONALD CERVICAL CERCLAGE. AT 24W 2024 SHE SPONTANEOUSLY LABORED AND DELIVERED BIRTH TWINS. TWIN B FORMER DONOR PASSED ON DOL7. SECONDARY TO GRADE IV IVH BILATERALLY AND BOWEL PERORATION. TWIN A FORMER RECIPIENT AND AT THE 1 YEAR FOLLOW UP IS DOLING WELL. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFEREBCE MDR: 9610617-2025-00273. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474950&pc=HFA | 2020550-2024-00022 | 2023-11-18 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; High Blood Pressure/ Hypertension |  |  | 1 | Event Description: IT WAS REPORTED TO KARL STORZ THAT A PATIENT'S MEMBRANE WAS RUPTURED AFTER SURGERY. THE PATIENT WAS PREGNANT WITH TWINS; HOWEVER, THE PATIENT HAD TWIN-TO-TWIN TRANSFUSION SYNDROME (TTTS) STAGE 3, A RARE CONDITION WHERE ONE TWIN RECEIVES TOO MUCH BLOOD WHILE THE OTHER TWIN RECEIVES TOO LITTLE BLOOD. ONE TWIN HAD REVERSED BLOOD FLOW IN THE UMBILICAL ARTERY, WHILE THE OTHER TWIN HAD INTERMITTENT ABSENT BLOOD FLOW IN THE UMBILICAL ARTERY. THE DIFFERENCE IN GROWTH BETWEEN THE TWINS WAS 18%. THE PATIENT ALSO HAD LARGE SUBCHORIONIC HEMATOMA. THE PATIENT UNDERWENT FETOSCOPIC LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS FOR TTTS. THE SURGERY WAS SUCCESSFUL, BUT THERE WERE RISKS INVOLVED, SUCH AS MEMBRANE SEPARATION THAT COULD LEAD TO PREMATURE RUPTURE OF MEMBRANES AFTER THE SURGERY AND PLACENTAL ABRUPTION. WITHOUT RECEIVING FETOSCOPIC LASER PHOTOCOAGULATION SURGERY, THE PATIENT WAS AT RISK OF LOSING ONE OR BOTH TWINS. AFTER THE SURGERY, THE PATIENT WAS FOUND TO HAVE A SIGNIFICANT AMOUNT OF CHORIOAMNIOTIC SEPARATION BUT REMAINED STABLE. LATER THAT AFTERNOON, THE PATIENT RUPTURED HER MEMBRANES BUT REMAINED STABLE. THE PATIENT WAS RESEALED AND RECEIVED LATENCY ANTIBIOTICS. THE TWINS RECOVERED FROM SURGERY WITH CONCORDANT FLUID, BLADDER SEEN, AND NORMAL DOPPLERS. ON THE 22ND WEEK AND 1 DAY OF PREGNANCY, THE PATIENT RECEIVED C-SECTION TO DELIVER THE TWINS BECAUSE THE PATIENT HAD SEVERE PREECLAMPSIA ON THE DAYS PRIOR. BOTH TWINS ARE IN THE NICU AND STABLE. IT WAS CONFIRMED BY THE TREATING PHYSICIAN THAT KARL STORZ DEVICES DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION. THE PHYSICIAN BELIEVED THAT THE COMPLICATION WAS DUE TO PATIENT'S PREVIOUS SUBCHORIONIC HEMATOMA. ADDITIONAL INFORMATION RECEIVED ON FEBRUARY 14.2025 VIA MW5165994: TWIN B, FORMER DONOR WAS GROWTH RESTRICTED AND BORN AT 0.4KG. ON DAY OF LIFE 11, HE DEVELOPED NECROTIZING ENTEROCOLITIS WITH PNEUMATOSIS REQUIRING RAPID ESCALATION OF MANAGEMENT OVER THE COURSE OF THE DAY. THIS INCLUDED ADMINISTRATION OF MULTIPLE BLOOD PRODUCTS, VASOPRESSOR SUPPORT, ANTIBIOTICS, VENTILATION TITRATION. HE DEVELOPED SEVERE PULMONARY HEMORRHAGE THAT EVENING RESULTING IN A CODE EVENT. THE NICU TEAM ALONG WITH PARENTS MADE THE JOINT DECISION TO COMPASSIONATELY EXTUBATE. HE PASSED ON DAY OF LIFE 12. THE DATE OF PASSING WAS (B)(6) 2023. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474951&pc=HFA | 9610617-2025-00269 | 2023-03-03 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | OPERATING SHEATH, 9 FR. | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose |  |  | 1 | Event Description: IT WAS REPORTED A 29-YEAR-OLD PREGNANT FEMALE WHO WAS DIAGNOSED WITH STAGE III TTTS (POLY-OLY AND ABSENT BLADDER IN TWIN B (DONOR) WITH ABNORMAL DOPPLERS OF INTERMITTENT ABSENT OF THE END DIASTOLIC FLOW OF THE UA.MVP WAS 8.5CM AND 0.5CM RESPECTIVELY FOR THE RECIPIENT AND DONOR. ALL OTHER DOPPLERS WERE NORMAL. THE PATIENT AND HER FAMILY UNDERWENT EXTENSIVE COUNSELING, AND THEY ELECTED TO PROCEED WITH FETOSCOPIC DIRECTED LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSES. SHE UNDERWENT LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS ON 2023. THE CASE WAS UNEVENTFUL. SHE HAD A POSTERIOR PLACENTA AND HAD A TOTAL OF 14 VASCULAR CONNECTION IDENTIFIED: 4 SMALL D-R, 2 MEDIUM D-R. 6 MEDIUM R-D. 1 LARGE D-R AND 1 EXTRA LARGE VV CONNECTION. SHE UNDERWENT AN AMNIOREDUCTION AS WELL. PRE-OPERATIVE FETAL HEART RATES WERE 153 BPM FOR THE DONOR AND 144 BPM FOR THE RECIPIENT, UNFORTUNATELY ON POD #1 IN 2023 SHE WAS TO HAVE A DEMISE OF TWIN A. SHE CONTINUED WITH PREGNANCY AND EVENTUALLY DELIVERED TWIN B AT 37W. AS OF NOW THE BABY IS DOING WELL AT ONE AND HALF YEARS. THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. CROSS REFERENCE MDR: 9610617-2025-00268. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. CROSS REFERENCE COMPLAINT ID: (B)(4). THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: A PRODUCT INVESTIGATION COULD NOT BE CARRIED OUT BECAUSE THE KARL STORZ PRODUCT INVOLVED WAS NOT RETURNED. THE REPORTING PHYSICIAN WAS CONTACTED AND CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED DID NOT EXHIBIT ANY MALFUNCTION AND WERE NOT CONSIDERED CONTRIBUTORY TO THE EVENT. BASED ON THE INFORMATION CURRENTLY AVAILABLE, THERE IS NO INDICATION THAT THE REFERENCED KARL STORZ PRODUCTS MALFUNCTIONED OR PLAYED A ROLE IN THE REPORTED COMPLICATIONS OR THE PATIENT'S OUTCOME. SHOULD THE PRODUCTS BE RETURNED IN THE FUTURE, A DETAILED INSPECTION WILL BE CONDUCTED, AND THE INVESTIGATION WILL BE UPDATED ACCORDINGLY. CURRENTLY, NO CORRECTIVE ACTION IS DEEMED NECESSARY AS THERE IS NO CRITICAL AND/OR SYSTEMATIC DEVIATION FOUND. AS A CONTINUED ACTION, WE WILL TRACK AND TREND RELATED COMPLAINTS ACCORDING TO OUR TRENDING PROCEDURES. THIS EVENT IS FILED UNDER INTERNAL COMPLAINT ID_(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21474956&pc=HFA | 2020550-2024-00021 | 2023-11-18 05:00 | Death | KARL STORZ SE & CO. KG | 2025-02-26 05:00 | HFA | MINI STRAIGHT-FORWARD TELESCOPE 0°,2 MM | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; High Blood Pressure/ Hypertension |  |  | 1 | Event Description: IT WAS REPORTED TO KARL STORZ THAT A PATIENT'S MEMBRANE WAS RUPTURED AFTER SURGERY. THE PATIENT WAS PREGNANT WITH TWINS; HOWEVER, THE PATIENT HAD TWIN-TO-TWIN TRANFUSION SYNDROME (TTTS) STAGE 3, A RARE CONDITION WHERE ONE TWIN RECEIVES TOO MUCH BLOOD WHILE THE OTHER TWIN RECEIVES TOO LITTLE BLOOD. ONE TWIN HAD REVERSED BLOOD FLOW IN THE UMBILICAL ARTERY, WHILE THE OTHER TWIN HAD INTERMITTENT ABSENT BLOOD FLOW IN THE UMBILICAL ARTERY. THE DIFFERENCE IN GROWTH BETWEEN THE TWINS WAS 18%. THE PATIENT ALSO HAD LARGE SUBCHORIONIC HEMATOMA. THE PATIENT UNDERWENT FETOSCOPIC LASER PHOTOCOAGULATION OF PLACENTAL ANASTOMOSIS FOR TTTS. THE SURGERY WAS SUCCESSFUL, BUT THERE WERE RISKS INVOLVED, SUCH AS MEMBRANE SEPARATION THAT COULD LEAD TO PREMATURE RUPTURE OF MEMBRANES AFTER THE SURGERY AND PLACENTAL ABRUPTION. WITHOUT RECEIVING FETOSCOPIC LASER PHOTOCOAGULATION SURGERY, THE PATIENT WAS AT RISK OF LOSING ONE ORF BOTH TWINS. AFTER THE SURGERY, THE PATIENT WAS FOUND TO HAVE A SIGNIFICANT AMOUNT OF CHORIOAMNIOTIC SEPARATION BUT REMAINED STABLE. LATER THAT AFTERNOON, THE PATIENT RUPTURED HER MEMBRANES BUT REMAINED STABLE. THE PATIENT WAS RESEALED AND RECEIVED LATENCY ANTIBIOTICS. THE TWINS RECOVERED FROM SURGERY WITH CONCORDANT FLUID, BLADDER SEEN, AND NORMAL DOPPLERS. ON THE 22ND WEEK AND 1 DAY OF PREGNANCY, THE PATIENT RECEIVED C-SECTION TO DELIVER THE TWINS BECAUSE THE PATIENT HAD SEVERE PREECLAMPSIA ON THE DAYS PRIOR. BOTH TWINS ARE IN THE NICU AND STABLE. IT WAS CONFIRMED BY THE TREATING PHYSICIAN THAT KARL STORZ DEVICES WERE DID NOT MALFUNCTIONED AND DID NOT CONTRIBUTE TO THE COMPLICATION. THE PHYSICIAN BELIEVED THAT THE COMPLICATION WAS DUE TO PATIENT'S PREVIOUS SUBCHORIONIC HEMATOMA ADDITIONAL INFORMATION RECEIVED ON FEBRUARY 14.2025 VIA MW5165994: TWIN B, FORMER DONOR WAS GROWTH RESTRICTED AND BORN AT 0.4KG. ON DAY OF LIFE 11, HE DEVELOPED NECROTIZING ENTEROCOLITIS WITH PNEUMATOSIS REQUIRING RAPID ESCALATION OF MANAGEMENT OVER THE COURSE OF THE DAY. THIS INCLUDED ADMINISTRATION OF MULTIPLE BLOOD PRODUCTS, VASOPRESSOR SUPPORT, ANTIBIOTICS, VENTILATION TITRATION. HE DEVELOPED SEVERE PULMONARY HEMORRHAGE THAT EVENING RESULTING IN A CODE EVENT. THE NICU TEAM ALONG WITH PARENTS MADE THE JOINT DECISION TO COMPASSIONATELY EXTUBATE. HE PASSED ON DAY OF LIFE 12. THE DATE OF PASSING WAS (B)(6) 2023 THE PHYSICIAN WHO REPORTED THIS EVENT WAS CONTACTED. THE PHYSICIAN CONFIRMED THAT THE KARL STORZ PRODUCTS INVOLVED IN THIS EVENT DID NOT MALFUNCTION AND DID NOT CONTRIBUTE TO THE COMPLICATION OR DEATH. Manufacturer Narrative: REPORTED DEVICE WOULD NOT BE RETURNED FOR EVALUATION BECAUSE THE PHYSICIAN CONFIRMED THERE WAS NO MALFUNCTION NOR DID THE DEVICE CONTRIBUTE TO THE ADVERSE EVENT. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21475045&pc=BZD | 2518422-2025-102400 | 2024-11-06 05:00 | Death | RESPIRONICS, INC. | 2025-02-26 05:00 | BZD | DREAMSTATION GO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21475151&pc=DSQ | 2916596-2025-01084 | 2024-05-01 04:00 | Death | THORATEC CORPORATION | 2025-02-26 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Liver Damage/Dysfunction; Renal Failure; Ventricular Fibrillation; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿AN INTRACARDIAC SHUNT IN A PATIENT UNDERGOING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH HEMODYNAMIC INSTABILITY, RIGHT VENTRICULAR (RV) FAILURE, BLEEDING, VENTRICULAR FIBRILLATION, NEUROLOGIC DYSFUNCTION, RENAL AND LIVER DYSFUNCTIONS, AND DEATH. THIS CASE STUDY PRESENTED A 52-YEAR-OLD MAN WITH END-STAGE HEART FAILURE DUE TO NONISCHEMIC CARDIOMYOPATHY WITH NOTED MEDICAL HISTORY OF A BIVENTRICULAR PACEMAKER, IMPLANTABLE CARDIOVERTER DEFIBRILLATOR IMPLANT FOR ATRIAL FIBRILLATION WITH PAST ABLATION, CEREBRAL VASCULAR ACCIDENT, AND INSULIN-DEPENDENT DIABETES WITH DIABETIC RETINOPATHY. THE PATIENT PRESENTED WITH CONCERN FOR AN EJECTION FRACTION OF 10%, WITH INTERROGATION NOTING SEVERELY ELEVATED PULMONARY ARTERY (PA) PRESSURES. THE PATIENT WAS NOT CONSIDERED A CANDIDATE FOR HEART TRANSPLANT GIVEN THE HIGH RISK OF GRAFT FAILURE IN THE SETTING OF ELEVATED PA PRESSURE. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) FOR AGGRESSIVE MEDICAL MANAGEMENT, AND AN IMPELLA 5.5 WAS PLACED SEVERAL DAYS LATER AFTER PA PRESSURES RETURNED TO BASELINE FROM ADMISSION. DUE TO THE CONTINUED ELEVATED PA PRESSURE AND HEART TRANSPLANT INELIGIBILITY, THE DECISION WAS MADE TO IMPLANT AN HM3. PRE-OPERATIVE TRANSTHORACIC ECHOCARDIOGRAM DEMONSTRATED A SEVERELY DILATED LEFT VENTRICLE AND ATRIUM, A POORLY VISUALIZED RV, AND MILD MITRAL REGURGITATION. AFTER 35 DAYS OF INPATIENT CARE AND 18 DAYS OF IMPELLA SUPPORT, THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR) FOR HM3 IMPLANT. A TRANSESOPHAGEAL (TEE) ECHOCARDIOGRAM PERFORMED WHILE IN THE OR REVEALED A NEW ATRIAL SEPTAL DEFECT (ASD) AT THE INFERIOR ASPECT OF THE FOSSA OVALIS. A SURGICAL REPAIR WAS PERFORMED TO REPAIR THE ASD; DENSE PERICARDIAL ADHESIONS WERE NOTED, LIKELY RELATED TO PRIOR EPICARDIAL LEAD PLACEMENT AND REMOVAL, RESULTING IN HEMORRHAGE REQUIRING CARDIOPULMONARY BYPASS (CPB). AFTER CPB INITIATION, THE IMPELLA WAS REMOVED AND THE HM3 INFLOW CANNULA WAS INSERTED. BLEEDING WAS THEN NOTED AROUND THE POSTERIOR OF THE HEART AND WAS TREATED WITH COSEAL. THE ASD WAS REPAIRED, AND THE HM3 OUTFLOW CANNULA WAS INSERTED. AS THE PATIENT WAS WEANED FROM CPB, THERE WERE CONCERNS FOR VASOPLEGIA VERSUS ADRENAL INSUFFICIENCY TREATED WITH METHYLENE BLUE, HYDROXOCOBALAMIN, AND STRESS-DOSE STEROIDS. SERIAL ARTERIAL BLOOD GASES DEMONSTRATED A WORSENING LACTIC ACIDOSIS WITH LACTATE OF 6.6 MMOL/L COMPARED TO PRE-OPERATIVE LACTATE OF 1.1 MMOL/L. RIGHT HEART FUNCTION ON TEE WAS ONLY MILDLY DEPRESSED, AND ASD REPAIR APPEARED SUCCESSFUL; THE CPB WAS WEANED. THE CASE WAS FURTHER COMPLICATED BY AN UNCONTROLLED HEMORRHAGE FROM THE POSTERIOR OF THE HEART OF UNCLEAR ORIGIN. LOCAL INTERVENTIONS WITH TACHOSIL, TISSEEL, SURGICEL FIBRILLAR, AND MULTIPLE ROUNDS OF PACKING WERE PERFORMED AND DESMOPRESSIN, FFP, PLATELETS, CRYOPRECIPITATE, AND FACTOR VII WERE ADMINISTERED. THE HEMORRHAGE WAS THEN DEEMED TO BE UNDER CONTROL, AND THE PATIENT WAS TRANSPORTED TO THE ICU WITH THEIR CHEST LEFT OPEN. OVER THE COURSE OF 30 MINUTES, THE PATIENT REQUIRED INCREASED SUPPORT TO MAIN MEAN ARTERIAL PRESSURE AND WAS DEEMED UNSTABLE. A CHEST EXPLORATION AND TEE NOTED THE RV WAS SEVERELY DILATED. A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS PLACED. THE PATIENT ALSO WENT INTO VENTRICULAR FIBRILLATION AND REQUIRED MULTIPLE ROUNDS OF DEFIBRILLATION. SIX DAYS LATER, THE PATIENT WAS DECANNULATED FROM THE RVAD AND TRANSITIONED TO AN IMPELLA RP. THE PATIENT¿S CHEST WAS CLOSED TWO DAYS LATER. THE PATIENT¿S NEUROLOGIC EXAMINATION REMAINED POOR, AND THEY SUFFERED WORSENING RENAL AND LIVER FUNCTION, COAGULOPATHY, AND WORSENING LACTIC ACIDOSIS. THE PATIENT WAS TRANSITIONED TO COMFORT CARE AND PASSED AWAY SHORTLY AFTER. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN LISTED AS THE DATE OF PUBLICATION (01MAR2024) SINCE THE DATE OF DATA COLLECTION WAS NOT PROVIDED. ADDITIONAL H6 CODE: 4426 - UNSPECIFIED NERVOUS SYSTEM PROBLEM. UNIVERSITY OF NORTH CAROLINA AT CHAPEL HILL, DEPARTMENT OF ANESTHESIOLOGY, CHAPEL HILL, NC SCHEIBER, C. J., TEETER, E. G., & SMELTZ, A. M. (2023). AN INTRACARDIAC SHUNT IN A PATIENT UNDERGOING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. JOURNAL OF CARDIOTHORACIC AND VASCULAR ANESTHESIA, 38(5), 1260¿1264. HTTPS://DOI.ORG/10.1053/J.JVCA.2023.12.030. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: SECTIONS A2 AND A3: PATIENT GENDER AND AGE CORRECTED SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBER OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEM IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, BLEEDING, CARDIAC ARRYTHMIA, NEUROLOGICAL DYSFUNCTION, RENAL DYSFUNCTION, HEPATIC DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AND NEUROLOGICAL DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478303&pc=QZI | MW5166909 | 2025-02-17 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-26 05:00 | QZI | FARAPULSE PULSED FIELD ABLATION SYSTEM |  | Bradycardia; Cardiac Arrest |  |  | 1 | Event Description: DURING AN ATRIAL FLUTTER PROCEDURE, THE PATIENT BECAME BRADYCARDIC, WENT INTO CARDIAC ARREST, AND EXPIRED. THE PROCEDURE BEGAN BY POSITIONING THE INQUIRY CATHETER IN THE CORONARY SINUS AND USING A (B)(6) INTRACARDIAC ECHO PROBE TO MAP THE RIGHT ATRIUM. THE MAP AND GEOMETRY WERE COLLECTED WITH A (B)(6) CIRCULAR CATHETER. THE TYPICAL ATRIAL FLUTTER ABLATION WAS PERFORMED WITH A (B)(6) 8MM RF(RADIOFREQUENCY) ABLATION CATHETER. AFTER MAPPING AND ABLATION IN THE RIGHT ATRIUM WERE COMPLETED, THE (B)(6) CIRCULAR CATHETER WAS MOVED TO THE LEFT ATRIUM VIA TRANSSEPTAL PUNCTURE USING SWARTZ AND AGILIS SHEATHS AND A BRK NEEDLE. GEOMETRY OF THE LEFT ATRIUM AND A VOLTAGE MAP WERE COLLECTED AND SHOWED AN ELECTRICAL SIGNAL IN THE PULMONARY VEINS. ISOLATION OF THE PULMONARY VEINS BEGAN WITH A (B)(6) (FARAPULSE, (B)(6)) CATHETER AND (B)(6) ((B)(6)) EQUIPMENT AND TECHNICAL SUPPORT. IMMEDIATELY AFTER ABLATION OF THE PULMONARY VEINS, A NEW MAP OF THE LEFT ATRIUM WAS COLLECTED USING THE (B)(6) CIRCULAR CATHETER TO CHECK WHETHER THE VEINS WERE ISOLATED. AT THAT MOMENT, AFTER THE GEOMETRY COLLECTION BEGAN ON THE POSTERIOR WALL OF THE LEFT ATRIUM, THE PATIENT BECAME BRADYCARDIC AND WENT INTO CARDIAC ARREST. THE TEAM PERFORMED RESUSCITATION MANEUVERS, BUT AFTER A PERIOD OF ATTEMPTS, THE PATIENT DIED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478308&pc=DTB | MW5166910 |  | Death | ST. JUDE / ABBOTT MEDICAL | 2025-02-26 05:00 | DTB | PERMANENT PACEMAKER ELECTRODE |  | Disseminated Intravascular Coagulation (DIC); Low Blood Pressure/ Hypotension; Unspecified Infection; Vascular Dissection |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478310&pc=NVN | MW5166911 |  | Death | ST. JUDE / ABBOTT MEDICAL | 2025-02-26 05:00 | NVN | PERMANENT PACEMAKER ELECTRODE |  | Disseminated Intravascular Coagulation (DIC); Low Blood Pressure/ Hypotension; Failure of Implant; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478353&pc=NVN | MW5166912 |  | Death | BIOTRONIK SE & CO. KG. | 2025-02-26 05:00 | NVN | PACER LEAD |  | Disseminated Intravascular Coagulation (DIC); Exsanguination; Low Blood Pressure/ Hypotension; Cardiac Perforation; Vascular Dissection |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED, AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21478386&pc=NVN | MW5166913 |  | Death | BIOTRONIK SE & CO. KG. | 2025-02-26 05:00 | NVN | PACER LEAD |  | Disseminated Intravascular Coagulation (DIC); Exsanguination; Low Blood Pressure/ Hypotension; Cardiac Perforation; Vascular Dissection |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21455844&pc=FSA | 1419652-2025-00013 | 2024-12-13 05:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-02-25 05:00 | FSA | MAXI MOVE | Detachment of Device or Device Component | Cardiac Arrest; Head Injury |  |  | 1 | Event Description: THE PATIENT WAS REPORTED TO BE TRANSFERRED FROM A BED TO A CHAIR USING A MAXI MOVE DEVICE AND AN ARJO SLING CLIP (MAA4031M-L-L1). DURING THE FINAL STAGE OF THE TRANSFER, THE CAREGIVER ATTEMPTED TO USE THE HANDSET TO TILT THE PATIENT INTO A SITTING POSITION. HOWEVER, THE HANDSET DID NOT WORK AND ASSISTANCE FROM A NURSE WAS REQUIRED. THE NURSE SUCCESSFULLY ACTIVATED THE HANDSET, BUT WHEN THE SPREADER BAR WAS TILTED, THE SLING BUCKLE CRACKED AND DISCONNECTED. AS A RESULT, THE PATIENT SLIPPED OUT OF THE SLING AND SUSTAINED A HEAD INJURY. IT WAS ALSO REPORTED THAT THE PATIENT SUFFERED A CARDIAC ARREST AND PASSED AWAY. FOLLOWING THE INCIDENT, ARJO CARRIED OUT AN INSPECTION OF THE SLING, WHICH CONFIRMED THAT A COMPONENT OF THE BUCKLE HAD BROKEN. THE SLING IN QUESTION WAS MANUFACTURED IN JUNE 2005 AND HAD BEEN IN USE FOR APPROXIMATELY 20 YEARS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21456922&pc=MVK | 3008642652-2025-02003 | 2025-01-28 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHT INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 09:22:17 ON (B)(6) 2025. AT 07:02:15 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 07:03:32, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 07:03:57, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT. AT 07:04:27, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 07:04:53, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 07:23:26, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION/TACTILE ARTIFACT. AT 07:24:32, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION/TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION/TACTILE ARTIFACT. AT 07:27:02, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 07:27:29, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 07:28:32, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 07:30:23, THE PATIENT RECEIVED THE EIGHTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 07:35:52 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21456991&pc=KFM | 2184009-2025-00304 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | KFM | AFFINITY CP PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132712 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: BBAP40 CENTRIFUGAL BLOOD PUMP MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR). TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C) 20 FR LEFT HEART VENT CATHETER AFFINITY OXYGENATORS. CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ONE CEREBRAL STROKE. SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI) ATRIAL FIBRILLATION. ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES. HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID SPMONQUAN (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0190-ELECTRO MECHANICAL SYSTEMS; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID SPMONQUAN (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0190-ELECTRO MECHANICAL SYSTEMS; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID BBAP40 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0189-DISPOSABLES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 95212 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0189-DISPOSABLES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 95212 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0189-DISPOSABLES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 77420 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 77420 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 77422 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 77422 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 91329C (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 91329C (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21457228&pc=DTZ | 2184009-2025-00315 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DTZ | AFFINITY NT OXYGENATOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K191029 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ¿ BBAP40 CENTRIFUGAL BLOOD PUMP ¿ MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR) ¿ TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C) ¿ 20 FR LEFT HEART VENT CATHETER ¿ AFFINITY OXYGENATORS ¿ CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ¿ ONE CEREBRAL STROKE ¿ SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI) ¿ ATRIAL FIBRILLATION ¿ ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES ¿ HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21457280&pc=DWF | 2184009-2025-00319 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DWF | EOPA ARTERIAL CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K031037 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: BBAP40 CENTRIFUGAL BLOOD PUMP. MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR) . TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C) . 20 FR LEFT HEART VENT CATHETER, AFFINITY OXYGENATORS CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ONE CEREBRAL STROKE SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI) . ATRIAL FIBRILLATION. ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES. HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION; THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: CORRECTION H11: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21457372&pc=DWF | 2184009-2025-00323 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DWF | EOPA ARTERIAL CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K031037 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ¿ BBAP40 CENTRIFUGAL BLOOD PUMP ¿ MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR) ¿ TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C) ¿ 20 FR LEFT HEART VENT CATHETER ¿ AFFINITY OXYGENATORS ¿ CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ¿ ONE CEREBRAL STROKE ¿ SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI) ¿ ATRIAL FIBRILLATION ¿ ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES ¿ HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: CORRECTION H11: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21457428&pc=DWF | 2184009-2025-00325 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DWF | MC2 TWO STAGE VENOUS CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K140165 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: BBAP40 CENTRIFUGAL BLOOD PUMP, MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR), TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C), 20 FR LEFT HEART VENT CATHETER, AFFINITY OXYGENATORS, CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ONE CEREBRAL STROKE, SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI), ATRIAL FIBRILLATION, ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES, HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: CORRECTION H11: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21457464&pc=DTS | 2184009-2025-00327 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DTS | VENT CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K834352 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: BBAP40 CENTRIFUGAL BLOOD PUMP, MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR), TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C), 20 FR LEFT HEART VENT CATHETER, AFFINITY OXYGENATORS, CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ONE CEREBRAL STROKE, SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI), ATRIAL FIBRILLATION, ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES, HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: CORRECTION H11: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21457987&pc=OZD | 1220648-2025-26304 | 2025-01-31 05:00 | Death | ABIOMED, INC. | 2025-02-25 05:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE TEAM WAS UNABLE TO DELIVER THE 14FR SHEATH AND THE TEAM OBSERVED PERIPHERAL ARTERY DISEASE. THE 14FR SHEATH WAS REMOVED AND REPLACED WITH ANOTHER NON-PEEL AWAY SHEATH. NO PATIENT HARM WAS REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER 43 HOURS OF SUPPORT. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: B1 ADVERSE EVENT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. B2 OUTCOME DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. B5 PATIENT OUTCOME WAS INADVERTENTLY OMITTED FROM THE NARRATIVE ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. H1 TYPE OF EVENT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. H6 WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE TEAM WAS UNABLE TO DELIVER THE 14FR SHEATH AND THE TEAM OBSERVED PERIPHERAL ARTERY DISEASE (PAD). THE TEAM REMOVED THE PROVIDED 14FR SHEATH AND REPLACED IT WITH ANOTHER NON-PEEL AWAY SHEATH. A2 PATIENT AGE WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. A3 PATIENT GENDER WAS INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. A5 PATIENT RACE AND ETHNICITY WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. D6B WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. E1 INITIAL REPORTED FIRST AND LAST NAME WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. G1 REPORTING CONTACT ADDRESS WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. G1 MANUFACTURING SITE NAME AND ADDRESS WAS INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26304. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458010&pc=DWF | 2184009-2025-00329 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DWF | CORONARY OSTIAL ANTEGRADE CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131269 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: BBAP40 CENTRIFUGAL BLOOD PUMP. MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR). TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C). 20 FR LEFT HEART VENT CATHETER. AFFINITY OXYGENATORS. CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ONE CEREBRAL STROKE. SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI). ATRIAL FIBRILLATION. ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES. HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458068&pc=EZL | 9610711-2025-00037 | 2024-06-01 04:00 | Death | COLOPLAST A/S | 2025-02-25 05:00 | EZL | FOLYSIL SILICONE CATHETER | Adverse Event Without Identified Device or Use Problem | Sepsis | K233411 |  | 1 | Event Description: ACCORDING TO THE AVAILABLE INFORMATION, THE PATIENT PASSED AWAY IN (B)(6) 2024 DUE TO SEPSIS. THE PATIENT'S FAMILY CONCLUDED THAT THE USE OF THE FOLYSIL CATHETER HAD A DIRECT CORRELATION WITH THE PATIENT'S DEATH. Manufacturer Narrative: THE INVESTIGATION IS NOT YET COMPLETE, A FOLLOW UP REPORT WILL BE SUBMITTED ON COMPLETION OF THE INVESTIGATION. B2: DATE IS UNKNOWN. THE PATIENT DIED IN (B)(6) 2024. B3: ESTIMATED DATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458099&pc=DWF | 2184009-2025-00331 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DWF | CORONARY OSTIAL ANTEGRADE CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131269 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: BBAP40 CENTRIFUGAL BLOOD PUMP, MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR), TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C), 20 FR LEFT HEART VENT CATHETER, AFFINITY OXYGENATORS, CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ONE CEREBRAL STROKE, SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI), ATRIAL FIBRILLATION, ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES, HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458111&pc=DWF | 2184009-2025-00333 | 2025-01-25 05:00 | Death | PERFUSION SYSTEMS | 2025-02-25 05:00 | DWF | DLP CORONARY ARTERY OSTIAL ANTEGRADE CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131269 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REGARDING "VACUUM-ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY". ALL DATA WERE COLLECTED FROM MULTIPLE A SINGLE CENTRE BETWEEN SEPTEMBER 2016 AND SEPTEMBER 2022. DEVICE SERIAL/LOT NUMBERS WERE NOT PROVIDED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ¿ BBAP40 CENTRIFUGAL BLOOD PUMP ¿ MEDTRONIC ELONGATED ONE PIECE ARTERIAL CANNULA (20 OR 22 FR) ¿ TWO-STAGE 29/29 FR VENOUS CANNULA (MC2 91329C) ¿ 20 FR LEFT HEART VENT CATHETER ¿ AFFINITY OXYGENATORS ¿ CORONARY ARTERY OSTIAL CANNULAE (10FR, 12FR OR 14 FR) THE QUANTUM PERFUSION SYSTEM MADE BY SPECTRUM MEDICAL WAS ALSO USED. AMONG ALL PATIENTS, TWO IN-HOSPITAL DEATHS OCCURRED IN THE VACUUM-ASSISTED VENOUS DRAINAGE (VAVD) COHORT. BASED ON THE AVAILABLE INFORMATION, THE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCTS. SPECIFIC CAUSES OF THE DEATHS WERE NOT PROVIDED. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: ¿ ONE CEREBRAL STROKE ¿ SEVERE POSTOPERATIVE ACUTE KIDNEY INJURY (AKI) ¿ ATRIAL FIBRILLATION ¿ ELEVATED POSTOPERATIVE ASPARTATE TRANSAMINASE (AST) VALUES ¿ HIGHER MEDIAN POSTOPERATIVE PEAK HIGH SENSIBILITY TROPONIN I LEVELS BASED ON THE AVAILABLE INFORMATION, THESE ADVERSE EVENTS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: G2: CITATION: AUTHORS: RAFFAELE SILVANO, PIETRO GIORGIO MALVINDI, FRANCESCA MAZZOCCA. VACUUM ASSISTED AND GRAVITATIONAL VENOUS DRAINAGE IN AORTIC VALVE SURGERY: A PROPENSITY-MATCH STUDY. PERFUSION VOL. 40(1) 221¿228 2025. 10.1177/02676591241230610 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458384&pc=OTE | MW5166772 | 2025-02-14 05:00 | Death | SIEMENS HEALTHINEERS AG | 2025-02-25 05:00 | OTE | MAMMOGRAPHY-DIGITAL | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Dyspnea |  |  | 1 | Event Description: SUBJECT DEVELOPED SUDDEN ONSET OF SHORTNESS OF BREATH AND SENT TO ED. SUBJECT FOUND TO HAVE PULMONARY EMBOLISM. PT EXPIRED AFTER A SHORT HOSP STAY. NOT RELATED TO MAMMOGRAM. SUDDEN DEATH. Event Description: ADDITIONAL INFORMATION RECEIVED FROM THE REPORTER ON 02/25/2025 FOR REPORT NUMBER MW5166772. COURSE NUMBER ON WHICH EVENT(S) OCCURRED: 4. TOTAL NUMBER OF COURSES TO DATE: 4.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458387&pc=OZD | 1220648-2025-26309 | 2025-01-28 05:00 | Death | ABIOMED, INC. | 2025-02-25 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PUMP WAS SUPPORTING FOR 20 HOURS WHEN THE PATIENT EXPIRED FROM MULTIPLE ORGAN FAILURE DUE TO INTRAPERITONEAL BLEEDING. THE BLEEDING ORIGINATED AT THE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SITE. THE WIRE USED DURING ECMO INSERTION DAMAGED THE MESENTERIC ARTERY, RESULTING IN INTRAPERITONEAL BLEEDING. THE CONTRIBUTION OF THE HEPARIN IN THE IMPELLA PURGE AND IMPELLA USE AND ANY CONTRIBUTION TO THE DEATH IS NOT KNOWN. THE BLEEDING MADE STABILIZATION OF THE PATIENT DIFFICULT AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458700&pc=MKJ | 2112020-2025-00122 |  | Death | CARDIAC SCIENCE CORPORATION | 2025-02-25 05:00 | MKJ | POWERHEART G5 KIT, AUTO, US ENG/ESP LA | Incorrect Interpretation of Signal | Ventricular Fibrillation | P160033 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 78-YEAR-OLD MALE PATIENT, THE DEVICE PROMPTED "NO SHOCK ADVISED" FOR A RHYTHM CLINICIANS BELIEVED TO BE SHOCKABLE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. INSTEAD, THE DATA FILE OF THE CUSTOMER'S REPORT WAS PROVIDED. REVIEW OF THE DATA FILE SHOWS THE DEVICE FUNCTIONED AS DESIGNED. THE G5 RHYTHMX SOFTWARE HAS A SAFETY FEATURE THAT CONSIDERS THE SHOCK NON-COMMITTED, WHICH MEANS THE DEVICE MAY ADVISE A SHOCK AND THEN CANCEL THE SHOCK BEFORE IT IS DELIVERED. THE DEVICE CONTINUALLY ANALYZES THE RHYTHM AND CAN IDENTIFY CHANGES IN SHOCKABILITY CRITERIA. IF THE PATIENT'S RHYTHM CHANGES TO A NON-SHOCKABLE RHYTHM, THE DEVICE WILL PROMPT TO THE USER THAT THE RHYTHM HAS CHANGED AND THE SHOCK HAS BEEN CANCELLED. THE USER HAS BEEN RECOMMENDED TO REVIEW THE DEVICE'S OPERATOR'S GUIDE. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458792&pc=DRF | 9617601-2025-00208 | 2025-01-13 05:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-02-25 05:00 | DRF | ACHIEVE MAPPING CATHETER - 20 MM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K102588 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY SIX HOURS AFTER A COMPLETED CRYO ABLATION PROCEDURE WHEN THE PATIENT RETURNED TO THE WARD, THE PATIENT'S VITAL SIGNS BECAME UNSTABLE. RESCUE EFFORTS WERE ATTEMPTED BUT THE PATIENT DIED. NO EXACT CAUSE OF DEATH WAS DETERMINED AS NO AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE DOCTOR BELIEVED THAT THE PATIENT HAD LEUKEMIA THAT CAUSED THE POSTOPERATIVE ABDOMINAL BLEEDING. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 2AF283 (LOT: 20522); PRODUCT TYPE: 0624-ARCTIC FRONT CATHETER PRODUCT ID 4FC12 (0012288806); PRODUCT TYPE: 0629-FLEXCATH SHEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21458795&pc=DRA | 3002648230-2025-00127 | 2025-01-13 05:00 | Death | MEDTRONIC CRYOCATH LP | 2025-02-25 05:00 | DRA | FLEXCATH ADVANCE STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K123591 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY SIX HOURS AFTER A COMPLETED CRYO ABLATION PROCEDURE WHEN THE PATIENT RETURNED TO THE WARD, THE PATIENT'S VITAL SIGNS BECAME UNSTABLE. RESCUE EFFORTS WERE ATTEMPTED BUT THE PATIENT DIED. NO EXACT CAUSE OF DEATH WAS DETERMINED AS NO AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE DOCTOR BELIEVED THAT THE PATIENT HAD LEUKEMIA THAT CAUSED POSTOPERATIVE ABDOMINAL BLEEDING. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 2AF283 (LOT: 20522); PRODUCT TYPE: 0624-ARCTIC FRONT CATHETER PRODUCT ID 990063-020 (229207704); PRODUCT TYPE: 0623-ACHIEVE CATHETER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21459183&pc=OZD | 1220648-2025-26310 | 2025-01-31 05:00 | Death | ABIOMED, INC. | 2025-02-25 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DEVELOPED LIMB ISCHEMIA. THE PATIENT WOULD EVENTUALLY EXPIRE. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT THREE DAYS PRIOR TO THE IMPELLA IMPLANT PROCEDURE COMPLAINING OF ABDOMINAL PAIN, INCREASED THIRST FOR TWO DAYS ALONG WITH NAUSEA AND VOMITING. INITIALLY THE PATIENT WAS TREATED FOR DIABETIC KETOACIDOSIS AND SEPTIC SHOCK. THE PATIENT'S EKG SHOWED ST ELEVATION IN THE POSTERIOR LEADS AND THE PATIENT WAS TAKEN TO THE CATH LAB AS A CODE ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AT SUCH TIME. THE FOLLOWING DAY SUPPORT WAS DISCONTINUED AS PATIENT DRIP REQUIREMENTS AND LABS WERE HEADED IN THE RIGHT DIRECTION. THERE WAS AN ATTEMPT TO TRANSFER THE PATIENT OUT, BUT TWO HOSPITALS TURNED DOWN THE ADMISSION AT THIS TIME. THE DECISION WAS MADE, TWO DAYS LATER, TO PUT THE PATIENT ON SUPPORT AGAIN DUE TO WORSENING LABS AND INCREASED DRIP REQUIREMENTS AS WELL AS POSSIBLE TRANSFER FOR HIGHER LEVEL OF CARE. THE IMPELLA CP DEVICE WAS PLACED; HOWEVER, IT WAS NOTED NO INTERVENTION WAS REQUIRED AT THIS TIME. THE PATIENT WAS TRANSFERRED TO THE UNIT ON IMPELLA MCS. THE NEXT DAY THE PATIENT WAS CANNULATED ON ECMO. BILATERAL PULSES WERE ABSENT PER THE NURSE, AND IT WAS UNSURE IF EITHER FROM THE CLAMP DOWN OR COMPROMISED PERFUSION. CONTINUOUS RENAL REPLACEMENT THERAPY WAS STARTED. THE NEXT DAY IT WAS NOTED A REPERFUSION CATHETER WAS PLACED BUT THE VASCULAR SUGGESTED ISCHEMIA DID NOT RESOLVE AFTER SIX HOURS WHICH REQUIRED A FASCIOTOMY WHICH WAS COMPLETED AT BEDSIDE. THE FOLLOWING DAY THE IMPELLA PERFORMANCE LEVEL WAS TURNED DOWN TO P-3 DUE TO SUCTION EVENTS. ALSO, ON THIS DAY, THE PATIENT WAS NOTED TO HAVE POOR FLOW TO THE BILATERAL LOWER EXTREMITIES, THE TEAM SUSPECTED COMPARTMENT SYNDROME. THE FOLLOWING DAY THE IMPELLA WAS HAVING SUCTION EVENTS; PERFORMANCE LEVEL WAS LOWERED TO P-2. IT WAS NOTED ECMO WAS "CHUGGING." THE PATIENT CONTINUED TO BE SUPPORTED WITH ALL THERAPIES (CRRT, ECMO AND IMPELLA). THE VASCULAR SURGEON RECOMMENDED THE RIGHT LOWER EXTREMITY AMPUTATION BUT DECIDED TO WAIT UNTIL PATIENT BECAME STABLE. HOWEVER, THE PATIENT WOULD EXPIRE, IT APPEARS, BEFORE THE AMPUTATION AFTER SEVEN DAYS TOTAL SUPPORT. CARE WAS WITHDRAWN. THE PATIENT WAS VERY ILL AND WAS NOT DOING WELL GOING INTO THE PROCEDURE FOR IMPELLA MCS. THE PATIENT HAD A COMBINATION OF THERAPIES INCLUDING ECMO. THE CAUSE OF DEATH IS UNKNOWN. THERE IS NOT ENOUGH INFORMATION, HOWEVER, TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME WHEN CONSIDERING THE IMPELLA RELATED EVENTS (ISCHEMIA LEADING TO AMPUTATION). Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21459428&pc=LYJ | 1644487-2025-00206 | 2024-09-04 04:00 | Death | CYBERONICS - HOUSTON | 2025-02-25 05:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. AT THE TIME OF THE REPORT, THE CAUSE OF DEATH HAD NOT BEEN CONFIRMED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21459783&pc=OZD | 1220648-2025-26315 | 2025-01-17 05:00 | Death | ABIOMED, INC. | 2025-02-25 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD BLEEDING. THE CLINICAL REPRESENTATIVE COULD NOT CONFIRM THE LOCATION OF THE BLEEDING. IT OCCURRED AFTER THE IMPELLA IMPLANT. ONE UNIT OF BLOOD PRODUCTS WERE GIVEN TO THE PATIENT. THE PROCEDURAL OUTCOME WAS THE PATIENT SURVIVED THE SURGERY.Event Description: ADDITIONAL INFORMATION RECEIVED THAT THE PATIENT DEVELOPED PROFOUND NEUROLOGICAL CHANGES AND A LARGE BRAIN BLEED WAS DISCOVERED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ B2 PATIENT OUTCOME REVISED TO REFLECT DEATH AND DATE OF DEATH WHICH INADVERTENTLY WAS NOT INCLUDED ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26315. B5 REVISED TO REFLECT THE PATIENT OUTCOME WHICH INADVERTENTLY WAS NOT INCLUDED ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26315. H6 REVISED TO REFLECT ADDITIONAL INFORMATION AND PATIENT OUTCOME WHICH INADVERTENTLY WAS NOT INCLUDED ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26315. H10 REVISED TO REFLECT INSTRUCTIONS FOR USE WHICH INADVERTENTLY WAS NOT INCLUDED ON THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-26315. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21459801&pc=KTI | 3011050570-2025-00165 |  | Death | GYRUS ACMI, INC. | 2025-02-25 05:00 | KTI | SINGLE USE ASPIRATION NEEDLE | Material Separation | Foreign Body In Patient; Cough | K163469 |  | 1 | Event Description: IT WAS REPORTED THAT SEVERAL MONTHS AFTER A PATIENT UNDERWENT AN ENDOBRONCHIAL ULTRASOUND (EBUS), A PART OF THE NEEDLE WAS FOUND IN THE PATIENT¿S RIGHT MAIN BRONCHUS. THE PATIENT HAD APPROXIMATELY FOUR ADMISSIONS FOR POST-TREATMENT INFECTIONS. THE PATIENT EXPERIENCED A POST-TREATMENT LOCULATED PNEUMOTHORAX, WHICH WAS MANAGED CONSERVATIVELY, FOLLOWED BY ADMISSION AND TREATMENT WITH A DRAIN FOR EMPYEMA WITH PNEUMOTHORAX. THE PATIENT HAD TWO FURTHER ADMISSIONS, PRESUMED TO BE DUE TO ONGOING EMPYEMA. HOWEVER, THE RADIOLOGIST NOTICED A FOREIGN BODY IN THE PATIENT'S LYMPH NODE. THIS APPEARED BETWEEN THE INITIAL STAGING COMPUTED TOMOGRAPHY (CT) AND THE POST-TREATMENT CT, WITH THE ONLY INTERVENTION THE PATIENT HAD BEING THE EBUS PROCEDURE. THE PATIENT WAS NOTED TO BE COUGHING A LOT DURING THE PROCEDURE, AND IT WAS NOT WELL TOLERATED. A BRONCHOSCOPY WAS PERFORMED BUT WAS UNABLE TO SAFELY REMOVE THE FOREIGN BODY. THE FRAGMENT WAS REMOVED THROUGH THORACIC SURGERY. ON AN UNSPECIFIED DATE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. NO FURTHER INFORMATION WAS PROVIDED Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO OLYMPUS. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. OLYMPUS WILL CONTINUE TO MONITOR THE FIELD PERFORMANCE OF THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21459846&pc=DSQ | 2916596-2025-00999 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿ANTI-FACTOR XA AND ACTIVATED PARTIAL THROMBOPLASTIN TIME STRATEGIES FOR UNFRACTIONATED HEPARIN DOSING AFTER HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION¿ THAT HEARTMATE 3 (HM3) IMPLANT IS ASSOCIATED WITH BLEEDING, GASTROINTESTINAL (GI) BLEEDING, THROMBOEMBOLIC EVENTS, AND DEATH. A RETROSPECTIVE REVIEW OF THE ELECTRONIC MEDICAL RECORD WAS PERFORMED FOR 115 CONSECUTIVE ADULT PATIENTS WHO UNDERWENT HM3 INSERTION BETWEEN 2020 AND 2022 AT OUR INSTITUTION. THREE PATIENTS WERE EXCLUDED AS THEY DID NOT RECEIVE UFH IN THE POST- OPERATIVE PERIOD BUT INSTEAD RECEIVED ARGATROBAN DUE TO HEPARIN- INDUCED THROMBOCYTOPENIA. IN OUR INSTITUTION'S CARDIOTHORACIC IN TENSIVE CARE UNIT, STANDARD PROTOCOL FOR UFH TITRATION AFTER HM3 INSERTION TRANSITIONED IN JUNE 2021 FROM UTILIZING APTT GOAL OF 45¿60 SECONDS TO FXA GOAL OF 0.1¿0.2 U/ML. THESE GOAL RANGES WERE ESTABLISHED THROUGH EXPERT CONSENSUS FROM HEMATOLOGY, PHARMACY, CRITICAL CARE MEDICINE, AND LABORATORY MEDICINE AT OUR INSTITUTION, GUIDED BY HM3 MANUFACTURER INSTRUCTIONS 11 AND EQUIVALENT FXA VALUES FROM THE HOSPITAL LABORATORY FACILITY. HM3 INSERTIONS PERFORMED ON AND AFTER JUNE 17TH, 2021, WERE AFFECTED BY THIS PROTOCOL CHANGE. PATIENTS WERE DIVIDED INTO APTT (N = 53, AGE 59.9 YEARS, 44 MALE) AND FXA (N = 59, AGE 62.8 YEARS, 53 MALE) COHORTS BASED ON WHICH ERA THEY WERE IN. DURING THIS STUDY PERIOD, ALL SURGICAL METHODS AND OTHER PATIENT MANAGEMENT PROTOCOLS OTHERWISE REMAINED IDENTICAL. THE STUDY INVESTIGATED THE DIFFERENCE IN THE RATE OF MAJOR BLEEDING AND THROMBOEMBOLIC EVENTS IN PATIENTS WHERE EITHER ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT) OR ANTI-FACTOR XA (FXA) WAS USED TO MONITOR INTRAVENOUS (IV) UNFRACTIONATED HEPARIN (UFH) LEVELS POST-HM3 IMPLANT. IN THE ADJUSTED (INVERSE PROBABILITY TREATMENT WEIGHTED) COHORTS, 97.5% OF FXA PATIENTS AND 91.05 OF APTT PATIENTS REACHED THERAPEUTIC LEVELS OF THEIR RESPECTIVE LABORATORY VALUES WITH NO SIGNIFICANT DIFFERENCES IN TIME BETWEEN HM3 IMPLANTATION AND UFH INITIATION (3.0 VS. 3.0 DAYS, P = 0.904), TOTAL UFH DURATION (8.0 VS. 7.9 DAYS, P = 0.526), OR MAXIMUM WEIGHT-ADJUSTED UFH DOES (14.7 VS 16.2 U/HR./KG, P = 0.181) RECEIVED. THERE WERE ALSO NO SIGNIFICANT DIFFERENCES IN IN-HOSPITAL MORTALITY (2.5% VS 3.1%, P = 0.360), MAJOR BLEEDING EVENTS (4.2% VS. 9.2, P = 0.360), OR THROMBOEMBOLIC EVENTS (21.85 VS. 10.1%, P = 0.151) BETWEEN THE ADJUSTED FXA AND APTT COHORTS. DEEP VENOUS THROMBOSIS (DVT) WAS THE MOST FREQUENT TYPE OF THROMBOEMBOLIC EVENT IN THE FXA GROUP (21.8% FXA VS 3.3% APTT), ALL OF WHICH WERE INTERNAL JUGULAR OR UPPER EXTREMITY VENOUS THROMBOSES, WHILE EXTRACORPOREAL CIRCUIT THROMBOSIS (PATIENTS ON CONCOMITANT CONTINUOUS VENO-VENOUS HEMOFILTRATION, EXTRACORPOREAL LIFE SUPPORT, OR EXTRACORPOREAL RIGHT VENTRICULAR ASSIST DEVICE) WAS THE MOST FREQUENT IN APTT PATIENTS (0.0% FXA VS 6.4% APTT), FOLLOWED BY INTRACARDIAC (0.0% FXA VS. 2.2% APTT) AND ARTERIAL LIMB (0.0% FXA VS. 1.6% APTT). CHEST (1.2% FXA AND APTT) AND CATHETER SITE (3.0% FXA VS 1.9% APTT) BLEEDING WERE OBSERVED IN BOTH FXA AND APTT GROUPS, WITH GASTROINTESTINAL BLEEDING (0.0% FXA VS 6.2% APTT) ALSO OBSERVED IN THE APTT GROUP. OVERALL, THE STUDY OBSERVED NO SIGNIFICANT DIFFERENCE IN BLEEDING OR THROMBOEMBOLIC EVENTS WITH FXA AND APTT MONITORING STRATEGIES DURING THE IMMEDIATE POST-OPERATIVE PERIOD AFTER HM3 IMPLANTATION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01JAN2022 AS THE IMPLANT DATES OCCURRED BETWEEN 2020 AND 2022. FENG I, KURLANSKY PA, POWLEY TR, HYNDS MA, YANG CG, EISENBERGER A, ET AL. ANTI-FACTOR XA AND ACTIVATED PARTIAL THROMBOPLASTIN TIME STRATEGIES FOR UNFRACTIONATED HEPARIN DOSING AFTER HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. ARTIF. ORGANS. 2024; 00: 1¿11. HTTPS://DOI.ORG/10.1111/AOR.14885. DIVISION OF CARDIOTHORACIC AND VASCULAR SURGERY, DEPARTMENT OF SURGERY, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NEW YORK, USA. DIVISION OF CARDIOTHORACIC AND VASCULAR SURGERY, DEPARTMENT OF SURGERY, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NEW YORK, USA. CENTER OF INNOVATION AND OUTCOMES RESEARCH, DEPARTMENT OF SURGERY, COLUMBIA UNIVERSITY, NEW YORK, NEW YORK, USA. DIVISION OF HEMATOLOGY/ONCOLOGY, DEPARTMENT OF MEDICINE, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NEW YORK, USA. DEPARTMENT OF ANESTHESIOLOGY, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NEW YORK, USA. DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NEW YORK, USA. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21460583&pc=NVN | 2017865-2025-15042 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-15041. IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: ANALYSIS WAS NORMAL. NO ANOMALIES WERE NOTED. THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21460688&pc=PNJ | 2017865-2025-15046 | 2025-02-04 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15044. IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR AVEIR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE. DURING THE PROCEDURE, THE LP WAS REPOSITIONED DUE TO HIGH CAPTURE THRESHOLD AND INADEQUATE R-WAVE AMPLITUDES. THE PATIENT HAD GONE INTO CARDIAC ARREST AND IT WAS ELECTED TO ABANDON THE PROCEDURE. CARDIAC PERFORATION RESULTING IN PERICARDIAL EFFUSION AND CARDIAC TAMPONADE OCCURRED. THE PATIENT HAD DECEASED UPON EXPERIENCING A SECOND INSTANCE OF CARDIAC ARREST DURING RECOVERY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21460697&pc=OZD | 1220648-2025-26319 | 2025-02-02 05:00 | Death | ABIOMED, INC. | 2025-02-25 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE 63-YEAR-OLD FEMALE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT LOST PULSES IN THE RIGHT LOWER EXTREMITY. NO INTERVENTION WAS PERFORMED AS THERE WAS ALREADY A PLAN TO EXPLANT THE PUMP. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21460750&pc=LWP | 2017865-2025-15041 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-15042. IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR ANALYSIS DUE TO DECEASED PATIENT. SAVING DEVICE IMAGE WAS SUCCESSFUL. THE DEVICE WAS TESTED AT NOMINAL SETTINGS. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED, AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21460858&pc=NVZ | 2017865-2025-15049 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSE. THERE WAS NO ALLEGATION THAT THE PACEMAKER CONTRIBUTED TO THE DEATH. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND. CORRECTION: CORRECTING G2 FROM HEALTH PROFESSIONAL AND USER FACILITY TO OTHER: FUNERAL HOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21460923&pc=PNJ | 2017865-2025-15044 | 2025-02-04 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Capturing Problem; Device Sensing Problem | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15046. IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR AVEIR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE. DURING THE PROCEDURE, THE LP WAS REPOSITIONED DUE TO HIGH CAPTURE THRESHOLD AND INADEQUATE R-WAVE AMPLITUDES. THE PATIENT HAD GONE INTO CARDIAC ARREST AND IT WAS ELECTED TO ABANDON THE PROCEDURE. CARDIAC PERFORATION RESULTING IN PERICARDIAL EFFUSION AND CARDIAC TAMPONADE OCCURRED. THE PATIENT HAD DECEASED UPON EXPERIENCING A SECOND INSTANCE OF CARDIAC ARREST DURING RECOVERY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21461012&pc=NIK | 2017865-2025-14961 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NIK | GALLANT DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14963, 2017865-2025-14962. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL. Manufacturer Narrative: CORRECTION: D9 RETURN DATE PREVIOUSLY REPORTED AS FEB 13, 2025, NOW CORRECTED TO FEB 5, 2025. D6B EXPLANT DATE UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21461877&pc=DSQ | 2916596-2025-01065 | 2025-02-06 05:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462302&pc=OZD | 1220648-2025-26321 | 2025-01-28 05:00 | Death | ABIOMED, INC. | 2025-02-25 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DEVELOPED THROMBUS, LIMB ISCHEMIA AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT WAS RECEIVING DIALYSIS AND CODED. CARDIOPULMONARY RESUSCITATION (CPR) WITH MULTIPLE SHOCKS WAS DELIVERED. THE PATIENT WENT TO THE EMERGENCY DEPARTMENT AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS PLACED. COMPUTED TOMOGRAPHY (CT) SCAN REVEALED A THORACIC BLEED AND THREE BROKEN RIBS FROM THE CPR. THE PATIENT WAS TRANSPORTED TO THE CATHETERIZATION LAB AND CATHETERIZATION SHOWED CHRONIC TOTAL OCCLUSION OF THE CIRCUMFLEX AND "SEVERELY" RIGHT CORONARY ARTERY AND A SMALL LEFT ANTERIOR DESCENDING ARTERY WAS OPEN AND POSSIBLY CLAMPED DOWN DUE TO PRESSORS. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE WHICH WAS SUCCESSFULLY PLACED. THE PATIENT WAS NOTED AS VERY SICK AND WAS NOT A CANDIDATE FOR VENTRICULAR ASSIST DEVICE OR TRANSPLANT. DUE TO INTERNAL BLEEDING FROM THE CPR, THREE UNITS OF BLOOD AND VOLUME WAS GIVEN. THE FOLLOWING DAY AFTER START OF IMPELLA MCS, THROMBUS AND LIMB ISCHEMIA WAS NOTED TO THE IMPELLA LEG. AN ATTEMPT TO PLACE AN ANTEGRADE SHEATH WAS NOTED WITH THE POSSIBILITY OF A THROMBECTOMY AS WELL. THE PATIENT HAS A HISTORY OF "SEVERE" PERIPHERAL VASCULAR DISEASE AND MULTIPLE INTERVENTIONS ON BOTH LEGS. THE FOLLOWING DAY, THE ANTEGRADE SHEATH THAT WAS PLACED CLOTTED OFF. SEVERE PERIPHERAL VASCULAR DISEASE WITH NO DISTAL OUTFLOW WAS NOTED. CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED WITH NOTED ISSUE OF CIRCUIT CLOTTING AS WELL. THE NEXT DAY THE UNIT UPDATE NOTED THE PATIENT HAD A GASTROINTESTINAL BLEED. THE PATIENT EXPIRED THE SAME DAY. CT SCAN SHOWED NO NEURAL ACTIVITY AND CARE WAS WITHDRAWN. THE PATIENT WAS VERY ILL AND ON A COMBINATION OF THERAPIES INCLUDING ECMO. THERE WAS A LOT GOING ON WITH THE PATIENT NOW IN ADDITION TO THE THROMBUS AND LIMB ISCHEMIA RELATED TO THE IMPELLA, AND WITH THIS, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR TO THE OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LED TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LED TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462356&pc=FKX | 0008030665-2025-00458 | 2025-02-05 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-02-25 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED. IT WAS REPORTED THAT THE CALL WAS PLACED TO DEACTIVATE THE PATIENT. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO PRODUCT. THE DATE OF DEATH WAS (B)(6) 2025 PER CORONER AFTER BEING DISCOVERED IN A STATE OF DECOMPOSITION. IT WAS UNKNOWN IF THE PATIENT WAS ATTACHED TO THE CYCLER WHEN FOUND. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE CONFIRMED THAT THERE IS STILL NO INFORMATION AVAILABLE AS TO WHETHER OR NOT THE PATIENT WAS ATTACHED TO THE CYCLER WHEN SHE DIED. HOWEVER, THE NURSE INDICATED THERE IS NO ALLEGATIONS AGAINST PRODUCT OR DIALYSIS AS A CAUSE FOR THE PATIENT¿S DEATH. THE NURSE REPORTED THAT THAT THE CORONER LISTED CAUSE OF DEATH FROM NATURAL CAUSES FROM END STAGE RENAL DISEASE.Event Description: A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED. IT WAS REPORTED THAT THE CALL WAS PLACED TO DEACTIVATE THE PATIENT. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO PRODUCT. THE DATE OF DEATH WAS (B)(6) 2025 PER CORONER AFTER BEING DISCOVERED IN A STATE OF DECOMPOSITION. IT WAS UNKNOWN IF THE PATIENT WAS ATTACHED TO THE CYCLER WHEN FOUND. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE CONFIRMED THAT THERE IS STILL NO INFORMATION AVAILABLE AS TO WHETHER OR NOT THE PATIENT WAS ATTACHED TO THE CYCLER WHEN SHE DIED. HOWEVER, THE NURSE INDICATED THERE IS NO ALLEGATIONS AGAINST PRODUCT OR DIALYSIS AS A CAUSE FOR THE PATIENT¿S DEATH. THE NURSE REPORTED THAT THE CORONER LISTED CAUSE OF DEATH FROM NATURAL CAUSES FROM END STAGE RENAL DISEASE. Manufacturer Narrative: CLINICAL REVIEW: BASED ON THE AVAILABLE INFORMATION, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT SUGGESTS A FRESENIUS DEVICE OR PRODUCT ISSUE CAUSED OR CONTRIBUTED TO A SERIOUS PATIENT HARM OR INJURY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D.9., H.3. PLANT INVESTIGATION: THE CYCLER WAS FOUND TO HAVE INSECT INFESTATION, THE CYCLER WILL BE QUARANTINED AND AN INVESTIGATION CANNOT BE PERFORMED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462417&pc=DSQ | 2916596-2025-01118 | 2025-01-30 05:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE PATIENT WISHED TO DISCONTINUE WITH ANY THERAPY, THEREFORE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS STOPPED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AFTER THE LVAD WAS STOPPED, THE PATIENT SUBSEQUENTLY PASSED AWAY ABOUT 1 WEEK LATER. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, OF THIS DOCUMENT LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462483&pc=CAF | 2518422-2025-022331 | 2023-10-31 04:00 | Death | RESPIRONICS, INC. | 2025-02-25 05:00 | CAF | DELUXE 230V 50HZ INT. UK PLUG | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K031413 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A UNIT IS LOST AND THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION OF A SPECIFIC MALFUNCTION OR FAILURE THAT IS RELEVANT TO THE REPORTED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. A PMS CLINICAL EXPERT REVIEW DETERMINED THAT THIS EVENT IS NOT ASSESSABLE FOR INJURY OR RELATEDNESS BASED ON THE INFORMATION PROVIDED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS INCIDENT WAS PREVIOUSLY REPORTED IN MFR NO.: 2518422-2025-022331. AFTER CAREFUL REVIEW, IT HAS BEEN FOUND THAT THIS CASE IS NOT REPORTABLE AS THE DEVICE IS NOT MANUFACTURED BY PHILIPS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462530&pc=DSQ | 2916596-2025-00930 | 2024-08-01 04:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "25 YEARS OF LEFT VENTRICULAR ASSIST DEVICES IN PERTH" THAT HEARTMATE 3 MAY BE ASSOCIATED WITH INFECTION, GASTROINTESTINAL BLEEDING, STROKE, AND DEATH. THE STUDY REVIEWED THE OUTCOMES OF THE WESTERN AUSTRALIAN LVAD PROGRAM. 158 LEFT VENTRICULAR ASSIST DEVICES (LVADS) HAVE BEEN IMPLANTED IN 25 YEARS, OF WHICH 106 WERE NEWER GENERATION DEVICES (INCLUDING HEARTMATE 3). ALMOST ALL PATIENTS (92%) WERE CLASSED AS BRIDGE-TO-TRANSPLANT OR -CANDIDACY/DECISION, WITH THE REMAINDER CLASSED ¿DESTINATION THERAPY¿ (DT). MEAN LV EJECTION FRACTION WAS 20% AND 33% HAD SEVERE RIGHT VENTRICULAR DYSFUNCTION. SIXTY-THREE PERCENT OF PATIENTS HAD A NON-ISCHAEMIC CARDIOMYOPATHY. FOLLOWING LVAD IMPLANT THE MEDIAN LENGTH OF STAY IN INTENSIVE CARE WAS 2 DAYS, AND IN HOSPITAL OVERALL 23 DAYS. NINETY-SIX PERCENT SURVIVED TO HOSPITAL DISCHARGE, AND FOLLOWING DISCHARGE 97% OF DAYS WITH LVAD WERE SPENT AS AN OUTPATIENT. THE MEDIAN NUMBER OF HOSPITAL READMISSIONS WAS 1.6 PER PATIENT PER YEAR. LVAD-ASSOCIATED INFECTION REQUIRING ADMISSION OR INTRAVENOUS ANTIBIOTICS AT ANY TIME POST-IMPLANT OCCURRED IN 36%, SIGNIFICANT GASTROINTESTINAL BLEEDING IN 19%, AND STROKE IN 9%. THE PERCENTAGE OF PATIENTS ALIVE WITH LVAD STILL IN SITU AT 1, 2, AND 5 YEARS WAS 92%, 88%, AND 68%, RESPECTIVELY. ALL DT PATIENTS SURVIVED AT LEAST 4 YEARS, SPENDING 97% OF DAYS WITH LVAD AS AN OUTPATIENT, AND FOUR DT PATIENTS HAVE SURVIVED FOR AT LEAST 10 YEARS. Manufacturer Narrative: SECTION D1: BRAND NAME CORRECTED. SECTION D4: MODEL NUMBER AND CATALOG NUMBER CORRECTED. SECTION E: REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE SAME AS PUBLISHED DATE SINCE DATE OF DATA COLLECTION WAS NOT PROVIDED TITLE: 25 YEARS OF LEFT VENTRICULAR ASSIST DEVICES IN PERTH BAUMWOL, J., SILBERT, B., SHAH, A., DEMBO, L., HAYES, H., MCLEAN, J., & LARBALESTIER, R. (2024). 25 YEARS OF LEFT VENTRICULAR ASSIST DEVICES IN PERTH. HEART LUNG AND CIRCULATION, 33, S242. HTTPS://DOI.ORG/10.1016/J.HLC.2024.06.244 NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462594&pc=DSQ | 2916596-2025-00959 | 2025-02-05 05:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Vasoconstriction; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO SEVERE LIVER FAILURE AFTER LONG-TERM VASOPLEGIA. THE PATIENT DEVELOPED SEVERE LIVER FAILURE AFTER LONG-TERM VASOPLEGIA WITH HIGHEST DOSAGES OF VASOPRESSORS. LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION WAS STRAIGHT FORWARD. THE PATIENT'S INDUCTION STARTED ALREADY WITH CATECHOLAMINES, AND THEY WERE ADDED WITH VASOPRESSORS ENTERING THE OPERATION ROOM PRIOR SKIN CUT. THE SURGERY STARTED ALREADY WITH LESS HEMODYNAMICS AND INCREASED DOSAGES OF THIS THERAPY. OVER THE WHOLE SURGERY THE MEDICAL TREATMENT HAD TO BE INCREASED MORE AND MORE WITH MAXIMUM DOSES OF VASOPRESSORS AND CATECHOLAMINES. THE DEVICE PERFORMED AS EXPECTED, AND NO ISSUES WERE MENTIONED. LIVER FAILURE WAS RELATED TO THE OUTCOME AT LEAST. THE LIVER FAILURE WAS SUSPECTED TO BE RELATED TO THE MASSIVE VASOCONSTRICTION BY THE VASOPRESSORS, CONSEQUENTLY HIGHEST SYSTEMIC RESISTANCE AND THEREFORE LESS PERFUSION OF THE LIVER. THE LVAD WAS RUNNING AS EXPECTED FROM THE BEGINNING ON AND SHOWED NO ISSUES IN OPERATION AND SUPPORT WITH AROUND 4 L/MIN NEARLY THE WHOLE TIME AFTER WENT OFF FROM BYPASS. THE MEDICAL TREATMENT COULD BE REDUCED AROUND 40% DURING THE LAST HOUR OF THE SURGERY UNTIL SKIN SUTURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B3: EVENT DATE CORRECTED. H6: CLINICAL AND IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING HEPATIC DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462615&pc=DSQ | 2916596-2025-01183 | 2023-04-30 04:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Electrical /Electronic Property Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Unspecified Infection; Respiratory Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "RIGHT HEART RESERVE FUNCTION ASSESSED WITH FLUID LOADING PREDICTS LATE RIGHT HEART FAILURE AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION" THAT HEARTMATE 3 MAY BE ASSOCIATED WITH LATE RIGHT HEART FAILURE (LRHF), HEART TRANSPLANT, AND DEATH. THE STUDY REVIEWED THE RECORDS OF ALL PATIENTS WHO UNDERWENT CONTINUOUS-FLOW LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION BETWEEN (B)(6) 2007 AND (B)(6) 2022. PATIENTS AGED GREATER THAN 18 YEARS WHO RECEIVED A CONTINUOUS-FLOW LVAD AND UNDERWENT RIGHT HEART CATHETERIZATION (RHC) WITH A FLUID LOADING TEST 1 MONTH AFTER LVAD IMPLANTATION WERE INCLUDED. HEMODYNAMIC RAMP TESTS FOR LVAD SPEED OPTIMIZATION WERE PERFORMED 1 MONTH AFTER LVAD IMPLANTATION. HEMODYNAMIC AND ECHOCARDIOGRAPHIC MEASUREMENTS WERE REPEATED BY CHANGING THE LVAD PUMP SPEEDS. THEN, THE LVAD PUMP SPEED WAS ADJUSTED BASED ON THE HEMODYNAMIC AND ECHOCARDIOGRAPHIC FINDINGS. IN ADDITION, ¿¿FLUID-LOADING TESTS,¿¿ WITH 10 ML/KG OF SALINE FOR 15 MINUTES, WERE PERFORMED. FOLLOW-UP EXAMINATIONS WERE COMPLETED IN ALL PATIENTS; ALL STUDY PARTICIPANTS WERE FOLLOWED UP UNTIL DEATH, HEART TRANSPLANTATION, OR DEVICE EXPLANTATION FOR RECOVERY UNTIL (B)(6) 2023. BETWEEN (B)(6) 2007 AND (B)(6) 2022, A TOTAL OF (B)(6) PATIENTS AGED GREATER THAN 18 YEARS UNDERWENT CONTINUOUS-FLOW LVAD IMPLANTATION AS EITHER BRIDGE-TO-TRANSPLANT OR DESTINATION THERAPY. THREE PATIENTS DIED WITHIN 1 MONTH AFTER LVAD IMPLANTATION. SIXTY-NINE PATIENTS WERE EXCLUDED BECAUSE POSTOPERATIVE RHC WAS PERFORMED WITHOUT APPROPRIATE TIMING OR SALINE LOADING DUE TO UNSTABLE HEMODYNAMIC CONDITIONS. FINALLY, A TOTAL OF 149 PATIENTS WERE INCLUDED IN THE STUDY (18 OF WHICH WERE IMPLANTED WITH HEARTMATE 3). THE STUDY POPULATION COMPRISED 110 MEN (73.8%) AND 39 WOMEN (26.2%), WITH A MEAN AGE OF 42.2 +/- 11.7 YEARS. DURING THE FOLLOW-UP PERIOD (1245 +/- 597DAYS), 71 PATIENTS RECEIVED HEART TRANSPLANTATION, AND 12 UNDERWENT LVAD EXPLANTATION. TEN PATIENTS WHO DEVELOPED LRHF UNDERWENT HEART TRANSPLANTATION. THIRTY-ONE PATIENTS DEVELOPED LRHF OR DIED (1 OF WHICH WAS IMPLANTED WITH A HEARTMATE 3). IN THE CURRENT STUDY, 18 DEATHS OCCURRED, AND THE CAUSES OF DEATH WERE THE FOLLOWING: CEREBROVASCULAR DISEASE (N = 6); INFECTION (N = 3); HEART FAILURE (N = 3); BLEEDING (N = 1); MULTIPLE ORGAN FAILURE (N = 2); RESPIRATORY FAILURE (N = 2); AND DEVICE TROUBLE (N = 1). Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. AUTHOR INFORMATION: TSUJI, M., KURIHARA, T., ISOTANI, Y., BUJO, C., ISHIDA, J., AMIYA, E., HATANO, M., SHIMADA, A., IMAI, H., KIMURA, M., SHIMADA, S., ANDO, M., ONO, M., & KOMURO, I. (2024). RIGHT HEART RESERVE FUNCTION ASSESSED WITH FLUID LOADING PREDICTS LATE RIGHT HEART FAILURE AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. CANADIAN JOURNAL OF CARDIOLOGY, 40(9), 1745¿1754. HTTPS://DOI.ORG/10.1016/J.CJCA.2024.02.012 UNIVERSITY OF TOKYO, TOKYO, JAPAN THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. A, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. C, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, BLEEDING, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, AS WELL AS DEVICE MALFUNCTIONS, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE IFU AND PATIENT HANDBOOK ALSO INSTRUCT THE USER TO CALL YOUR HOSPITAL CONTACT RIGHT AWAY IF YOU NOTICE A CHANGE IN HOW YOUR PUMP SOUNDS, FEELS, OR WORKS. EVEN SMALL CHANGES SHOULD BE REPORTED. THE HANDBOOK ALSO INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE USER THINKS, FOR ANY REASON, ANY PORTION OF THE EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THE USER IS UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462651&pc=BZD | 2518422-2025-001859 | 2024-08-01 04:00 | Death | RESPIRONICS, INC. | 2025-02-25 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF PHILIPS CPAP DEVICE THAT THE REPORTER STATES HER HUSBAND IS DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER¿S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462812&pc=DSQ | 2916596-2025-01057 | 2024-12-10 05:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. LOG FILES WERE SUBMITTED FOR REVIEW, AND THERE DID NOT APPEAR TO BE ANY PUMP OR POWER ISSUES. IT DID NOT APPEAR THE EQUIPMENT CONTRIBUTED TO THE PATIENT DEATH. THESE ARE THE LOG FILES OF A PATIENT THAT PASSED AWAY. PATIENT WOULD BE PRESENTED FOR A MORBIDITY AND MORTALITY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ANALYSIS OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED RELEVANT EVENTS FROM 09DEC2024 THROUGH 10DEC2024. LOW FLOW HAZARD ALARMS WERE CAPTURED THROUGHOUT THE FILE. THE DRIVELINE WAS DISCONNECTED ON 10DEC2024 AT 00:44:48 AND THE CONTROLLER WAS PUT IN SLEEP MODE, WHICH IS CONSISTENT WITH THE PATIENT¿S EXPIRATION. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21462872&pc=NIM | 2032493-2025-00166 | 2023-12-16 05:00 | Death | MICROVENTION, INC. | 2025-02-25 05:00 | NIM | CASPER X | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombosis/Thrombus |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ACUTE DISSECTION OF THE ICA, AN 8MM X 40MM CASPER X STENT WAS NOT IMPLANTED DUE TO THROMBUS FORMATION. DURING DEPLOYMENT, THE STENT WAS FULLY THROMBOSED AND THEREFORE REMOVED. THROMBUS OCCURRED DURING DEPLOYMENT IN THE PATIENT; THE STENT WAS PARTLY DEPLOYED. IT WAS NOT POSSIBLE TO OPEN THE VESSEL AND TREAT THE DISSECTION. NO FURTHER TREATMENT PROVIDED. THE PATIENT DIED TO MCA INFARCTS TWO DAYS POST PROCEDURE. PRE-PROCEDURE AND AT DISCHARGE MEDICATION: ASA: 300MG, SAPT.Event Description: PLEASE SEE SECTION H11 FOR INVESTIGATION FINDINGS. Manufacturer Narrative: THE PRODUCT REPORTED IN THIS REPORT IS AN EXPORTED DEVICE NOT CLEARED OR APPROVED FOR MARKETING IN THE US. THE COMPLAINT IS BEING REPORTED BASED ON THIS PRODUCT MEETING SIMILAR PRODUCT CRITERIA (SIMILAR TO ROADSAVER / CASPER CAROTID STENT DEVICE ¿ PMA#: P210030). THE LOT NUMBER WAS NOT PROVIDED; THEREFORE, A SEARCH FOR PRODUCTION-RELATED NON-CONFORMANCE'S COULD NOT BE PERFORMED. THE DEVICE WAS IMPLANTED AND THEREFORE NOT AVAILABLE FOR RETURN AND INVESTIGATION BY THE MANUFACTURER. HOWEVER, MEDICAL INFORMATION WAS PROVIDED AND WILL BE REVIEWED. THE INVESTIGATION IS ONGOING AND UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL MDR WILL BE SUBMITTED. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. Manufacturer Narrative: INVESTIGATION FINDINGS: BASED ON A REVIEW OF THE DEVICE¿S RISK DOCUMENTATION, THE REPORTED EVENT DID NOT INDICATE THERE WERE THE PRESENCE OF ANY POTENTIAL OR NEW MANUFACTURING, DESIGN, QUALITY, OR OTHER SYSTEMIC ISSUES, OR NON-CONFORMANCES. THE COMPLAINT CODE IS MONITORED THROUGH THE TRENDING PROCESS; CORRECTIVE ACTION IS DETERMINED, AS NEEDED, THROUGH THIS PROCESS. INVESTIGATIONS OF HISTORIC COMPLAINT FILES WITH SIMILAR COMPLAINT CATEGORY CODING ARE RECORDED IN THE COMPLAINT HANDLING SYSTEM; WITHOUT THE ABILITY TO PERFORM AN ANALYSIS OF THE DEVICE, THIS INVESTIGATION IS LIMITED IN ITS ABILITY TO FURTHER EXAMINE ANY POTENTIAL ROOT CAUSES. BATCH REVIEW: A SEARCH FOR NON-CONFORMANCES ASSOCIATED WITH THIS PART / LOT NUMBER COMBINATION COULD NOT BE PERFORMED AS THIS INFORMATION WAS NOT AVAILABLE AT THE TIME THIS INVESTIGATION WAS PERFORMED. COMPLAINT SYSTEM REVIEW: A SEARCH OF THE COMPLAINT HANDLING SYSTEM COULD NOT BE PERFORMED TO DETERMINE IF OTHER SIMILAR COMPLAINTS EXIST FOR THIS BATCH NUMBER, BECAUSE THE BATCH NUMBER WAS NOT PROVIDED FOR THE PRODUCT ON THIS COMPLAINT FILE. IFU REVIEW (ADDITIONAL INFORMATION CAN BE FOUND IN THE IFU): POTENTIAL COMPLICATIONS: POSSIBLE COMPLICATIONS INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: ARRHYTHMIA, ANEURYSM AND PSEUDOANEURYSM FORMATION, ABRUPT VESSEL CLOSURE, ALLERGIC REACTIONS (INCLUDING TO ANTIPLATELET AGENTS, CONTRAST MEDIUM OR STENT MATERIALS), ARTERIOVENOUS FISTULA, BLEEDING FROM ANTICOAGULATION/ANTIPLATELET MEDICATION, BRADYCARDIA AND HYPOTENSION, CAROTID ARTERY SPASM, CORONARY ISCHEMIA, DEATH, DISSEMINATED INTRAVASCULAR COAGULATION, EMBOLI (AIR, TISSUE, PLAQUE, THROMBUS, DEVICE OR OTHER), EMERGENCY ARTERY BYPASS GRAFT SURGERY, HEMATOMA, HEMORRHAGIC OR EMBOLIC STROKE/TIA, INFECTION AND/OR PAIN AT INSERTION SITE/SEPSIS, INTIMAL TEAR/DISSECTION, MYOCARDIAL INFARCTION (MI), NEW OR WORSE ENCEPHALOPATHY, RENAL FAILURE/INSUFFICIENCY, RESPIRATORY ARREST, STENT MISPLACEMENT, TISSUE NECROSIS, VESSEL INJURY/DISSECTION/PERFORATION/RUPTURE/TRAUMA, VESSEL OCCLUSION OR THROMBOSIS, VESSEL SPASM OR RECOIL. WARNINGS: SHOULD UNUSUAL RESISTANCE BE FELT AT ANY TIME DURING ACCESS OR REMOVAL, THE SHEATH INTRODUCER/GUIDE CATHETER AND CASPER X SYSTEM SHOULD BE REMOVED AS A SINGLE UNIT. APPLYING EXCESSIVE FORCE DURING DELIVERY OR RETRIEVAL OF THE CASPER X SYSTEM CAN POTENTIALLY RESULT IN LOSS OR DAMAGE TO THE DEVICE AND DELIVERY COMPONENTS. IF AN EPD DEVICE IS USED, ALLOW FOR AND MAINTAIN ADEQUATE DISTANCE BETWEEN THE FILTER, THE STENT DELIVERY SYSTEM OR DEPLOYED STENT TO AVOID POTENTIAL ENTANGLEMENT. DO NOT REPOSITION THE CASPER X SYSTEM WITHOUT FULLY RETRIEVING THE DEVICE. THE CASPER X STENT MUST BE RETRIEVED INTO THE DELIVERY SYSTEM AND REDEPLOYED AT THE DESIRED TARGET LOCATION OR REMOVED COMPLETELY FROM THE PATIENT. THE USE OF A GUIDING SHEATH OR GUIDING CATHETER WITH A FIXED HEMOSTASIS VALVE MAY CAUSE THE DISTAL TIP OF THE DELIVERY CATHETER TO DETACH AT THE HEMOSTASIS VALVE UPON REMOVAL IF THE VALVE IS NOT ADEQUATELY OPENED. PRECAUTIONS: EXERCISE CAUTION WHEN CROSSING THE DEPLOYED/DETACHED CASPER X SYSTEM WITH ADJUNCTIVE DEVICES SUCH AS GUIDEWIRES, CATHETERS OR BALLOON CATHETERS TO AVOID DISRUPTING THE DEVICE GEOMETRY AND DEVICE PLACEMENT. DIRECTIONS FOR USE: INTRODUCTION OF THE STENT DELIVERY SYSTEM: ACCESS THE TREATMENT SITE USING THE APPROPRIATE ACCESSORY EQUIPMENT AND INSERT AN EPD USABLE AS A GUIDEWIRE. IF AN EPD IS NOT PREFERRED, USE A SUFFICIENTLY LONG .014¿ GUIDEWIRE AND PLACE IT ACROSS THE LESION. WARNING: ALWAYS USE A GUIDEWIRE OR EPD WHEN ADVANCING OR WITHDRAWING THE DELIVERY SYSTEM. IF NECESSARY, PERFORM A PRE-DILATATION AT THE LESION TARGET SITE USING A STANDARD PTA TECHNIQUE USING THE RECOMMENDED GUIDE SHEATH OR GUIDING CATHETER ON THE OUTER BOX LABEL. THE GUIDE WIRE OR EPD DEVICE SHOULD BE ADVANCED PAST THE TARGET LESION SITE. THEN THE DELIVERY SYSTEM SHOULD BE ADVANCED OVER THE GUIDE WIRE OR EPD TO THE LESION SITE USING FLUOROSCOPY. CAUTION: DO NOT ADVANCE THE DELIVERY SYSTEM AGAINST SIGNIFICANT RESISTANCE. THE CAUSE OF THE RESISTANCE SHOULD BE DETERMINED BY FLUOROSCOPY AND REMEDIAL ACTION TAKEN. WITHDRAW THE SYSTEM AND USE A NEW ONE. SLACK REMOVAL: ENSURE THAT THE DELIVERY SYSTEM CATHETER OUTSIDE THE PATIENT REMAINS FLAT AND STRAIGHT. WARNING: SLACK IN THE DELIVERY SYSTEM CATHETER EITHER OUTSIDE OR WITHIN THE PATIENT MAY RESULT IN MISPLACEMENT OF THE STENT BEYOND THE LESION. STENT DEPLOYMENT: NOTE: THE CASPER X STENT SHOULD BE SIZED TO THE VESSEL ACCORDING TO THE REFERENCE VESSEL DIAMETER RANGE LISTED ON THE PRODUCT LABEL. THE CASPER X STENT FORESHORTENS AS IT EXPANDS TO THE DIAMETER OF THE NATIVE VESSEL. THE STENT FORESHORTENING SHOULD BE TAKEN INTO ACCOUNT WHEN SIZING AND DEPLOYING THE CASPER X CAROTID STENT SYSTEM. IF THE CASPER X STENT POSITIONING IS NOT SATISFACTORY ACROSS THE TARGET LESION, THE STENT MAY BE RECAPTURED AND REPOSITIONED IF IT IS NOT FULLY DEPLOYED. THE IMPLANT MAY BE RECAPTURED UP TO 50% OF ITS DEPLOYMENT FROM THE CATHETER DELIVERY SYSTEM (SEE FIGURE 5 AND 6). RECAPTURING OF STENT CAN BE ACCOMPLISHED BY REMOVING THE SLACK IN THE DELIVERY CATHETER SHAFT AND THEN MAINTAINING THE INNER CATHETER (HANDLE) POSITION AND ADVANCING OR PUSHING FORWARD THE OUTER CATHETER. AFTER RECAPTURING THE STENT INTO THE DELIVERY SYSTEM THEN YOU CAN REDEPLOY THE STENT ACROSS THE LESION, TAKING INTO CONSIDERATION THE FORESHORTENING OF THE STENT. ENSURE THE RHV IS OPEN. THE STENT IS DEPLOYED BY OUTER CATHETER RETRACTION. DEPLOYMENT IS COMPLETE BY MAINTAINING INNER CATHETER POSITION (HOLDING HANDLE) WHILE RETRACTING THE OUTER CATHETER AND ALLOWING THE STENT TO EXPAND. WITHDRAWING STENT DELIVERY SYSTEM: AFTER THE STENT HAS BEEN COMPLETELY DEPLOYED, USING FLUOROSCOPY, WITHDRAW THE DELIVERY SYSTEM CAREFULLY, LEAVING THE GUIDEWIRE OR EPD IN PLACE. PERFORM STANDARD POST-PROCEDURAL ANGIOGRAPHY. POST-DEPLOYMENT STENT DILATATION: IF THE STENT IS NOT COMPLETELY EXPANDED WITHIN THE LENGTH OF THE LESION, POST-DILATION WITH THE BALLOON CATHETER INSIDE THE STENT MAY BE DONE WITH STANDARD PTA TECHNIQUE. THE INFLATED NOMINAL DIAMETER OF THE PTA BALLOON USED FOR POST-DILATION SHOULD BE APPROXIMATELY THE SAME AS THE DIAMETER OF THE REFERENCED NATIVE VESSEL. PROCEDURE/MEDICAL INFORMATION REVIEW: A DETAILED MEDICAL REVIEW OF MEDICAL REPORT DATA WAS PERFORMED. DATA REVIEWED INDICATES THAT THE PATIENT WAS TREATED FOR AN ACUTE DISSECTION LOCATED AT THE ICA. DURING THE COURSE OF PERFORMANCE OF THE PROCEDURE ON INDEX PROCEDURE DAY (B)(6) 2023, IT WAS REPORTED THAT DURING THE DEPLOYMENT THE STENT, THE STENT BECAME FULLY THROMBOSED AND WAS THEREFORE REMOVED. IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2023. BASED ON A REVIEW OF AVAILABLE DATA AND PROFESSIONAL MEDICAL OPINION, THE RELATIONSHIP OF THE EVENT TO THE CASPER X CANNOT BE RULED OUT. NO DEVICE MALFUNCTION WAS REPORTED ASSOCIATED WITH THE CASPER X DEVICE AND/OR NO DEVICE MALFUNCTION WAS REPORTED AS THE CAUSE OF THE REPORTED STENT THROMBOSIS. INVESTIGATION CONCLUSION: A MEDICAL REVIEW OF THE PROVIDED MEDICAL REPORT WAS PERFORMED. DATA REVIEWED INDICATES THAT THE PATIENT WAS TREATED FOR AN ACUTE DISSECTION LOCATED AT THE ICA. DURING THE COURSE OF PERFORMANCE OF THE PROCEDURE ON INDEX PROCEDURE DAY (B)(6) 2023, IT WAS REPORTED THAT DURING THE DEPLOYMENT THE STENT, THE STENT BECAME FULLY THROMBOSED AND WAS THEREFORE REMOVED. IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2023. BASED ON THE REVIEW OF AVAILABLE DATA, THE RELATIONSHIP OF THE EVENT TO THE CASPER X CANNOT BE RULED OUT. NO DEVICE MALFUNCTION WAS REPORTED ASSOCIATED WITH THE CASPER X DEVICE AND/OR NO DEVICE MALFUNCTION WAS REPORTED AS THE CAUSE OF THE REPORTED STENT THROMBOSIS. SUPPLEMENTAL IMAGING WAS UNAVAILABLE FOR REVIEW. WITHOUT THE RETURN AND EVALUATION OF THE PHYSICAL DEVICE, THE INVESTIGATION CANNOT FURTHER EXAMINE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21463319&pc=NIQ | 2124215-2025-09189 | 2025-02-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-25 05:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH MULTIVESSEL CORONARY ARTERY DISEASE AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. FOLLOWING PRE-DILATION, A 32 X 3.50 PROMUS ELITE DRUG-ELUTING STENT (DES) WAS DEPLOYED IN THE LEFT ANTERIOR DESCENDING ARTERY WITH 85% STENOSIS. ANOTHER 38 X 3.00 PROMUS PREMIER DES WAS ALSO DEPLOYED IN THE RIGHT CORONARY ARTERY WITH 90% STENOSIS. FOLLOWING THE SUCCESSFUL PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE CARDIAC CARE UNIT. IN THE EVENING, THE PATIENT EXPERIENCED CHEST PAIN, HYPOTENSION AND CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED IMMEDIATELY BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST AND AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21463320&pc=NIQ | 2124215-2025-09179 | 2025-02-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-25 05:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH MULTIVESSEL CORONARY ARTERY DISEASE AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. FOLLOWING PRE-DILATION, A 32 X 3.50 PROMUS ELITE DRUG-ELUTING STENT (DES) WAS DEPLOYED IN THE LEFT ANTERIOR DESCENDING ARTERY WITH 85% STENOSIS. ANOTHER 38 X 3.00 PROMUS PREMIER DES WAS ALSO DEPLOYED IN THE RIGHT CORONARY ARTERY WITH 90% STENOSIS. FOLLOWING THE SUCCESSFUL PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE CARDIAC CARE UNIT. IN THE EVENING, THE PATIENT EXPERIENCED CHEST PAIN, HYPOTENSION AND CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED IMMEDIATELY BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST AND AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21463374&pc=NKE | 9614453-2025-00783 | 2024-10-14 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-25 05:00 | NKE | SOLARA QUAD CRT-P MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P010015 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S FAMILY THAT THE PATIENTS HEART RATE WAS DECREASING AND THE PATIENT SUBSEQUENTLY PASSED AWAY. ADDITIONAL INFORMATION RECEIVED FROM THE CLINICIAN NOTED THE MANNER OF DEATH WAS NATURAL HOWEVER THE CAUSE WAS NOT PROVIDED. THE CLINICIAN ALSO NOTED THAT THE PATIENT WAS TO UNDERGO AN ATRIOVENTRICULAR (AV) NODE ABLATION BUT PASSED AWAY BEFORE THE PROCEDURE COULD TAKE PLACE. AT THE TIME OF DEATH THE PATIENT HAD AN IMPLANTED CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) DEVICE THAT DOES NOT PROVIDE TACHYARRHYTHMIA THERAPY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21463917&pc=DSQ | 2916596-2025-01048 | 2021-11-30 05:00 | Death | THORATEC CORPORATION | 2025-02-25 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "ASSOCIATION BETWEEN CASELOAD VOLUME AND OUTCOMES IN LEFT VENTRICULAR ASSIST DEVICE IMPLANTATIONS¿ A EUROMACS ANALYSIS" THAT HEARTMATE 3 PATIENTS MAY HAVE EXPERIENCED MAJOR ADVERSE CARDIAC AND CEREBROVASCULAR EVENTS, RIGHT HEART FAILURE, GASTROINTESTINAL BLEEDING, DRIVELINE INFECTION, PUMP THROMBOSIS, MAJOR DEVICE MALFUNCTION EXCLUDING PUMP THROMBOSIS, MAJOR NEUROLOGICAL DYSFUNCTION, MAJOR BLEEDING, READMISSION, OUTFLOW GRAFT OBSTRUCTION, MAJOR DEVICE INFECTION EXCLUDING DRIVELINE, DEVICE EXCHANGE, TRANSPLANT AND DEATH. THIS IS AN OBSERVATIONAL EUROPE-WIDE RETROSPECTIVE ANALYSIS OF ALL PRIMARY CONTINUOUS-FLOW LVAD IMPLANTATIONS RECORDED IN THE EUROPEAN REGISTRY FOR PATIENTS WITH MECHANICAL CIRCULATORY SUPPORT (EUROMACS) REGISTRY BETWEEN 1 JANUARY 2011 AND 31 DECEMBER 2020, WITH CENSOR DATE 30 NOVEMBER 2021. THE EUROMACS REGISTRY INCLUDES 52 CENTRES FROM 18 EUROPEAN COUNTRIES AND OVER 5000 PATIENTS ON LVAD SUPPORT. THE STUDY AIMED TO INVESTIGATE THE ASSOCIATION BETWEEN THE ANNUAL AND CUMULATIVE CASELOAD AND CLINICAL OUTCOMES IN LVAD PATIENTS IN EUROPE. PATIENTS WERE FOLLOWED UP FROM THE TIME OF DEVICE IMPLANT UNTIL DEATH, HEART TRANSPLANTATION, DEVICE EXPLANT, RECOVERY, LOSS TO FOLLOW-UP, OR CENSOR DATE, WHICHEVER OCCURRED FIRST. THE PRIMARY ENDPOINT WAS DEFINED AS 1-YEAR OVERALL SURVIVAL. SECONDARY OUTCOMES INCLUDED SURVIVAL AT 90 DAYS AND 3 YEARS, AS WELL AS DEVICE-RELATED ADVERSE EVENTS (DISABLING STROKE, MAJOR DEVICE INFECTION [EXCLUDING DRIVELINE], PUMP THROMBOSIS OR OUTFLOW GRAFT OBSTRUCTION, MAJOR BLEEDING, RIGHT HEART FAILURE), UNPLANNED REHOSPITALIZATION, AND EVENT-FREE-SURVIVAL. COMPETING RISK ANALYSIS WAS USED FOR SURVIVAL AND DEVICE EXPLANT DUE TO DEVICE EXCHANGE, HEART TRANSPLANTATION, OR MYOCARDIAL RECOVERY. A SUBGROUP ANALYSIS CONSIDERED ONLY PATIENTS SUPPORTED WITH HM3 DEVICES. A TOTAL OF 4802 PATIENTS UNDERWENT PRIMARY LVAD IMPLANTATION IN 35 PARTICIPATING CENTRES, 1522 OF THOSE PATIENTS WERE IMPLANTED WITH HEARTMATE 3. CUMULATIVE FOLLOW-UP AMOUNTED TO 10003 PATIENT-YEARS, WITH A MEDIAN FOLLOW-UP OF 1.54 YEARS (INTERQUARTILE RANGE [IQR] 0.52¿3.15). MEDIAN AGE WAS 57.0 YEARS (IQR 48.0¿63.0), AND THE MAJORITY WERE MALE (82.9%). IN TOTAL, 2147 (44.7%) PATIENTS DIED, AFTER MEDIAN 286 DAYS OF SUPPORT (IQR 48¿860). WHEN CONSIDERING INDIVIDUAL ADVERSE EVENTS, HIGHER ANNUAL CASELOAD CENTRES DOCUMENTED ELEVATED INCIDENCE OF AND RISK FOR PUMP THROMBOSIS OR OUTFLOW GRAFT OBSTRUCTION AS FIRST EVENT (ASHR 1.76, 95%CI 1.18¿2.62, P=0.005), AS WELL AS A RECURRENT EVENT (P<0.001). CENTRES WITH HIGHER ANNUAL CASELOADS ALSO NOTED A REDUCED RISK FOR MAJOR DEVICE MALFUNCTIONS OTHER THAN PUMP THROMBOSIS (SHR 0.50, 95% CI 0.27¿0.90, P=0.022), AS WELL AS REDUCED RISK OF DISABLING STROKE (ASHR0.42, 95%CI 0.20¿0.87, P=0.019). IN THE SUBGROUP ANALYSIS INCLUDING HM3 PATIENTS ONLY, FIRST EVENTS OBSERVED A SIGNIFICANTLY ELEVATED 5.2-FOLD RISK INCREASE FOR PUMP THROMBOSIS OR OUTFLOW GRAFT OBSTRUCTION (P=0.007)IN HIGHER CASELOAD CENTRES, AS WELL AS INCREASED RISKS FOR INITIAL DRIVELINE INFECTIONS AND INITIAL MAJOR BLEEDING EVENTS (ASHR 2.36,P0.005; AND A SHR 1.79, P=0.049, RESPECTIVELY). FIRST EVENTS EXPERIENCED BY HM3 PATIENTS ALSO INCLUDED MAJOR ADVERSE CARDIAC AND CEREBROVASCULAR EVENTS, RIGHT HEART FAILURE, GASTROINTESTINAL BLEEDING, DRIVELINE INFECTION, PUMP THROMBOSIS, MAJOR DEVICE MALFUNCTION EXCLUDING PUMP THROMBOSIS, MAJOR NEUROLOGICAL DYSFUNCTION, MAJOR BLEEDING, AND READMISSION. Manufacturer Narrative: SECTION D1: BRAND NAME CORRECTED. SECTION D4: MODEL NUMBER AND CATALOG NUMBER CORRECTED. SECTION E: REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 30NOV2021 AS DATA WAS COLLECTED UNTIL 30NOV2021. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B2: DATE OF DEATH WAS NOT PROVIDED SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE SAME AS PUBLISHED DATE SINCE DATE OF DATA COLLECTION WAS NOT PROVIDED. ARTICLE TITLE: ASSOCIATION BETWEEN CASELOAD VOLUME AND OUTCOMES IN LEFT VENTRICULAR ASSIST DEVICE IMPLANTATIONS¿ A EUROMACS ANALYSIS. MIHALJ, M., REINEKE, D., JUST, I. A., MULZER, J., CHOLEVAS, N., HOERMANDINGER, C., VEEN, K., LUEDI, M. M., HEINISCH, P. P., POTAPOV, E., GUMMERT, J. F., MOHACSI, P., HAGL, C., FAERBER, G., ZIMPFER, D., DE BY, T. M., MEYNS, B., GUSTAFSSON, F., HUNZIKER, L., . . . SCHOENRATH, F. (2024). ASSOCIATION BETWEEN CASELOAD VOLUME AND OUTCOMES IN LEFT VENTRICULAR ASSIST DEVICE IMPLANTATIONS ¿ A EUROMACS ANALYSIS. EUROPEAN JOURNAL OF HEART FAILURE, 26(11), 2400¿2409. HTTPS://DOI.ORG/10.1002/EJHF.3418. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21464967&pc=DRA | 2124215-2025-11619 | 2025-01-28 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-25 05:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Device Dislodged or Dislocated | Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH CARDIAC TAMPONADE OCCURRED. DURING THE PROCEDURE, THE PHYSICIAN MENTIONED THAT THE FIRST TRANSSEPTAL WAS COMPLETED USING THE VERSACROSS CONNECT WITH THE FARADRIVE SHEATH. THEY THEN BEGAN MAPPING WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER THROUGH THE FARADRIVE, AND ENDED UP LOSING TRANSSEPTAL ACCESS AFTER THE DEVICES WERE DIFFICULT TO VISUALIZE. THUS, A SECOND TRANSSEPTAL WAS PERFORMED. THE ANAESTHESIOLOGIST THEN IDENTIFIED THAT THE PATIENT'S CO2 AND BLOOD PRESSURE WERE FALLING. THE ICE CATHETER WAS THEN USED TO IDENTIFY A PERICARDIAL EFFUSION. PROCEDURE WAS CANCELLED. A PERICARDIAL WINDOW WAS USED TO MANAGE THE EFFUSION AND AN OPEN CHEST AT BESIDE AND THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR RECOVERY. THE PHYSICIAN BELIEVES THE VERSACROSS KIT USED FOR TRANSSEPTAL CONTRIBUTED TO THE EFFUSION. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS.Event Description: IT WAS REPORTED THAT DURING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH CARDIAC TAMPONADE OCCURRED. DURING THE PROCEDURE, THE PHYSICIAN MENTIONED THAT THE FIRST TRANSSEPTAL WAS COMPLETED USING THE VERSACROSS CONNECT WITH THE FARADRIVE SHEATH. THEY THEN BEGAN MAPPING WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER THROUGH THE FARADRIVE, AND ENDED UP LOSING TRANSSEPTAL ACCESS AFTER THE DEVICES WERE DIFFICULT TO VISUALIZE. THUS, A SECOND TRANSSEPTAL WAS PERFORMED. THE ANAESTHESIOLOGIST THEN IDENTIFIED THAT THE PATIENT'S CO2 AND BLOOD PRESSURE WERE FALLING. THE ICE CATHETER WAS THEN USED TO IDENTIFY A PERICARDIAL EFFUSION. PROCEDURE WAS CANCELLED. A PERICARDIAL WINDOW WAS USED TO MANAGE THE EFFUSION AND AN OPEN CHEST AT BESIDE AND THE PATIENT WAS ADMITED TO THE HOSPITAL FOR RECOVERY. THE PHYSICIAN BELIEVES THE VERSACROSS KIT USED FOR TRANSSEPTAL CONTRIBUTED TO THE EFFUSION. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY. THE DATE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THAT DURING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARADRIVE SHEATH CARDIAC TAMPONADE OCCURRED. DURING THE PROCEDURE, THE PHYSICIAN MENTIONED THAT THE FIRST TRANSSEPTAL WAS COMPLETED USING THE VERSACROSS CONNECT WITH THE FARADRIVE SHEATH. THEY THEN BEGAN MAPPING WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER THROUGH THE FARADRIVE, AND ENDED UP LOSING TRANSSEPTAL ACCESS AFTER THE DEVICES WERE DIFFICULT TO VISUALIZE. THUS, A SECOND TRANSSEPTAL WAS PERFORMED. THE ANAESTHESIOLOGIST THEN IDENTIFIED THAT THE PATIENT'S CO2 AND BLOOD PRESSURE WERE FALLING. THE ICE CATHETER WAS THEN USED TO IDENTIFY A PERICARDIAL EFFUSION. PROCEDURE WAS CANCELLED. A PERICARDIAL WINDOW WAS USED TO MANAGE THE EFFUSION AND AN OPEN CHEST AT BESIDE AND THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR RECOVERY. THE PHYSICIAN BELIEVES THE VERSACROSS KIT USED FOR TRANSSEPTAL CONTRIBUTED TO THE EFFUSION. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY. THE DATE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS IN THEIR 30S. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465118&pc=DTB | 2649622-2025-05463 | 2025-01-29 05:00 | Death | MPRI | 2025-02-25 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Asystole; Heart Block | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, THE RIGHT VENTRICULAR (RV) LEAD WAS PLACED SUCCESSFULLY. DURING THE IMPLANT OF THE RIGHT ATRIAL (RA) LEAD, THE PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY AND RECOVERED AFTER RECEIVING CARDIOPULMONARY RESUSCITATION. THE RA LEAD WAS PLACED SUCCESSFULLY. APPROXIMATELY FIVE MINUTES LATER, THE PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY AND ASYSTOLE. AFTER FORTY-FIVE MINUTES OF CARDIOPULMONARY RESUSCITATION AND PACING SUPPORT, THE PATIENT ULTIMATELY PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465124&pc=LWS | 2649622-2025-05464 | 2025-01-29 05:00 | Death | MPRI | 2025-02-25 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Appropriate Term/Code Not Available | Asystole; Heart Block | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, THE RIGHT VENTRICULAR (RV) LEAD WAS PLACED SUCCESSFULLY. DURING THE IMPLANT OF THE RIGHT ATRIAL (RA) LEAD, THE PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY AND RECOVERED AFTER RECEIVING CARDIOPULMONARY RESUSCITATION. THE RA LEAD WAS PLACED SUCCESSFULLY. APPROXIMATELY FIVE MINUTES LATER, THE PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY AND ASYSTOLE. AFTER FORTY-FIVE MINUTES OF CARDIOPULMONARY RESUSCITATION AND PACING SUPPORT, THE PATIENT ULTIMATELY PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465563&pc=NVZ | 2017865-2025-15121 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NVZ | ZEPHYR XL DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS IN BACKUP MODE WHEN RECEIVED, CONSISTENT WITH LOW BATTERY VOLTAGE. ANALYSIS INDICATED THE DEVICE WAS OPERATING AT BELOW END-OF-SERVICE LEVEL AND WAS IMPLANTED BEYOND ITS PROJECTED LONGEVITY. ELECTRICAL, LONGEVITY, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465657&pc=NVN | 2017865-2025-15128 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15123. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15126. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15127. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15123. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15126. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15127. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND. Manufacturer Narrative: CORRECTION: B5 FOR STATEMENT AND G2 FOR FUNERAL HOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465663&pc=NVY | 2017865-2025-15127 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15123. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15126. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15128. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15123. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15126. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15128. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND. Manufacturer Narrative: CORRECTION: B5 FOR STATEMENT AND G2 FOR FUNERAL HOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465670&pc=NIK | 2017865-2025-15123 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | NIK | UNIFY ASSURA CRT-D RF HV | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15126, RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15127, RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15128. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15126. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15127. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15128. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: B5 FOR STATEMENT AND G2 FOR FUNERAL HOME. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS NO COMMUNICATION WHEN RECEIVED. ELECTRICAL, LONGEVITY, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21465671&pc=OJX | 2017865-2025-15126 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-25 05:00 | OJX | QUICKFLEX LV LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15123. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15127. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15128. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15123. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15127. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15128. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DECEASED. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND. Manufacturer Narrative: CORRECTION: B5 FOR STATEMENT AND G2 FOR FUNERAL HOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21444258&pc=MGB | 3010252479-2025-00015 | 2025-02-03 05:00 | Death | ESSENTIAL MEDICAL, INC. | 2025-02-24 05:00 | MGB | MANTA 18F | Break; Malposition of Device | No Clinical Signs, Symptoms or Conditions | P180025 |  | 1 | Event Description: IT WAS REPORTED THAT: "PHYSICIAN BROKE SUTURE LINE DURING DEPLOYMENT AND PUSHED COLLAGEN IN VESSEL. THE OPERATOR PULLED BACK THE MANTA DEVICE AFTER WITH EXCESSIVE FORCE AND THE SUTURE BROKE. THE VESSEL WAS PTA (BALLOON INFLATIONS) AND PLACEMENT OF A BMS (BARE METAL STENT). POST ANGIOGRAM REVEALED PATENT BLOOD FLOW." ADDITIONAL INFORMATION: "DURING AN ECMO ARTERIAL DECANNULATION PROCEDURE. MICRO PUNCTURE AND ULTRASOUND WERE UTILISED TO GAIN ACCESS IN THE LEFT COMMON FEMORAL ARTERY. VESSEL SIZE WAS >7.5MM WITH NO REPORTED CALCIFICATION OR TORTUOSITY. MEASURED DEPTH FOR THE MANTA WAS 5+1=6. POST MANTA DEPLOYMENT TIME TO HAEMOSTASIS WAS 15 MINUTES. SYSTOLIC BP WAS 105/55MMHG. ACT WAS 189. THERE WAS NO REPORTED PATIENT HARM. THEY WERE REPORTED AS FINE POST THE PROCEDURE."Event Description: IT WAS REPORTED THAT: "PHYSICIAN BROKE SUTURE LINE DURING DEPLOYMENT AND PUSHED COLLAGEN IN VESSEL. THE OPERATOR PULLED BACK THE MANTA DEVICE AFTER WITH EXCESSIVE FORCE AND THE SUTURE BROKE. THE VESSEL WAS PTA (BALLOON INFLATIONS) AND PLACEMENT OF A BMS (BARE METAL STENT). POST ANGIOGRAM REVEALED PATENT BLOOD FLOW." ADDITIONAL INFORMATION: "DURING AN ECMO ARTERIAL DECANNULATION PROCEDURE. MICRO PUNCTURE AND ULTRASOUND WERE UTILISED TO GAIN ACCESS IN THE LEFT COMMON FEMORAL ARTERY. VESSEL SIZE WAS >7.5MM WITH NO REPORTED CALCIFICATION OR TORTUOSITY. MEASURED DEPTH FOR THE MANTA WAS 5+1=6. POST MANTA DEPLOYMENT TIME TO HAEMOSTASIS WAS 15 MINUTES. SYSTOLIC BP WAS 105/55MMHG. ACT WAS 189. THERE WAS NO REPORTED PATIENT HARM. THEY WERE REPORTED AS FINE POST THE PROCEDURE." Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). UDI RELATED DATA QUALITY UPDATES ONLY SECTION D.4.-UNIQUE IDENTIFIER (UDI)# CORRECTED TO (B)(4). ONE UNIT EACH OF THE MANTA, SHEATH, DILATOR, AND DEPTH LOCATOR WERE RETURNED FOR EVALUATION, ALL WITH VISIBLE BLOOD PARTICULATES. THE MANTA WAS LOCKED INTO THE SHEATH WITH THE LEVER ACTUATED, BUT NO COMPONENTS (COLLAGEN, LOCK, OR ANCHOR) WERE PRESENT OR RETURNED. ANGIOGRAPHY SHOWED A LOWER-POSITIONED RADIOPAQUE LOCK WITH OCCLUDED FLOW AT THE ACCESS SITE. BALLOON ANGIOPLASTY INDICATED PARTIAL DOG-BONING AND A DISSECTION INFERIOR TO THE ACCESS. THE DEVICE LOT HISTORY RECORD REVIEW FOR LOT NUMBER 73B2400587 MANTA 18F INDICATED NO NON-CONFORMITIES RELATED TO THIS LOT. THE CASE DETAILS FROM THE CLINICAL STAFF WERE REVIEWED. A 72-YEAR-OLD FEMALE PATIENT (306 LBS, 5 FT 4 IN, 52 KG/M2) WITH A HISTORY OF LOW EJECTION FRACTION AND SEVERE MITRAL VALVE DISEASE UNDERWENT ECMO ARTERIAL DECANNULATION IN THE OR. THE MANTA INSTRUCTIONS FOR USE (IFU) INCLUDE A WARNING AGAINST USE IN PATIENTS WITH MARKED OBESITY (BMI >40 KG/M2). AFTER ECMO, AN 18F MANTA WAS DEPLOYED AT THE MEASURED DEPTH. DURING DEPLOYMENT, THE PHYSICIAN APPLIED EXCESSIVE FORCE WHILE PULLING BACK THE DEVICE, CAUSING THE SUTURE LINE TO BREAK AND FORCING THE COLLAGEN INTO THE VESSEL. POST-ANGIO CONFIRMED MALPOSITIONING AND INTRAARTERIAL COLLAGEN DEPLOYMENT, CAUSING OCCLUSION. MANUAL PRESSURE WAS APPLIED, FOLLOWED BY BALLOON ANGIOPLASTY, AND A BARE METAL STENT WAS DEPLOYED. HEMOSTASIS WAS ACHIEVED IN 15 MINUTES, AND THE PATIENT INITIALLY REMAINED STABLE WITH NO HARM, INJURY, OR HEALTH CONSEQUENCES DUE TO THIS EVENT. HOWEVER, LATER THAT EVENING, THE PATIENT DEVELOPED LIMB ISCHEMIA AND HYPERLACTATEMIA. DESPITE VASCULAR CONSULTATION, NO INTERVENTION OCCURRED BEFORE THE PATIENT EXPIRED EARLY IN THE MORNING. THE SUSPECTED ROOT CAUSE OF THE OCCLUSION REQUIRING STENTING WAS USER ERROR DUE TO EXCESSIVE FORCE ON THE DEVICE, LEADING TO SUTURE BREAKAGE AND INTRAARTERIAL COLLAGEN DEPLOYMENT. THE PHYSICIAN BELIEVES THAT THE PATIENT'S DEATH WAS UNRELATED TO THE MANTA CLOSURE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21444526&pc=DSP | 2249723-2025-0000849 | 2025-02-03 05:00 | Death | DATASCOPE CORP. | 2025-02-24 05:00 | DSP | CARDIOSAVE HYBRID W/ E/F PLUG | Inability to Auto-Fill | Insufficient Information | K112372 |  | 1 | Event Description: ON MONDAY, FEBRUARY 3, IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) DISLAYED AUTOFILL FAILURE MESSAGE. THE FAILURE OCCURRED PRIOR TO THE CLINICAL APPLICATION OF THE THERAPY, AND THAT THE FAILURE PREVENTED THE USE OF THE CARDIOSAVE SYSTEM IN A PATIENT. ON TUESDAY, FEBRUARY 4, IT WAS NOTIFIED THAT PATIENT DIED.Event Description: N/A Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: AWARE AND EVENT DATE WAS CONFIRMED WITH THE FSE AND UPDATED. UPDATED FIELDS - B3, B4, D9, G3, G6, H2, H11. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. Manufacturer Narrative: UPDATED FIELDS - B4, G3, G6, H2, H3, H6 ( TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSION), H11 A GETINGE FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO EVALUATE THE IABP UNIT AND WAS ABLE TO CONFIRM THE REPORTED ISSUE. THE INTERNAL BLOCKS ARE ADJUSTED (THE PIM AND DRIVE MANIFOLD BLOCKS WERE RELOCATED TO ACHIEVE BETTER SEALING, AS WELL AS THE INTERNAL CONNECTORS) THE FIBER OPTIC CABLE WAS DETECTED TO BE PINCHED, IT IS REPLACED BY A NEW CABLE. ALL CHECKS ARE PERFORMED ACCORDING TO PROTOCOL. THE CORRECT OPERATION OF THE EQUIPMENT IS VERIFIED. EQUIPMENT SUITABLE FOR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21445173&pc=GAW | 2210968-2025-02009 | 2025-02-13 05:00 | Death | ETHICON INC. | 2025-02-24 05:00 | GAW | PRLNE BLU 36IN 2-0 D/A V-7 MP | Break | Insufficient Information | K133356 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A DISSECTION CARDIAC SURGERY ON THE NIGHT OF (B)(6) 2025 AND PURSE-STRING SUTURE WAS USED. THE SUTURE BROKE, RESULTING IN DEATH ON (B)(6) 2025. ADDITIONAL INFORMATION WAS REQUESTED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 DEVICE NOT RETURNED. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WOULD THE SURGEON LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? WHAT IS THE CAUSE OF DEATH? PLEASE PROVIDE THE PATIENT'S DEMOGRAPHIC INFORMATION INCLUDING AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE NAME OF INDEX SURGICAL PROCEDURE? THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WAS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE? (OPEN, LAPAROSCOPIC OR OTHER)? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE ORIGINALLY TIED (MULTIPLE KNOTS, SQUARE KNOT, ETC.)? WHAT INSTRUMENTS WERE USED IN THIS PROCEDURE TO HANDLE THE SUTURE? PLEASE DESCRIBE ANY MEDICAL/SURGICAL INTERVENTION REQUIRED FOR THIS SUTURE EVENT INCLUDING DATES AND RESULTS. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? WHAT SYMPTOMS DID THE PATIENT EXPERIENCE FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? LOT NUMBER? WILL PRODUCT BE RETURNED? IF SO, PLEASE PROVIDE THE RETURN DATE AND TRACKING INFORMATION. SURGEON¿S NAME? Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: RECEIVED INFORMATION AS FOLLOWING FROM SALES REP VIA PHONE TODAY. AFTER INVESTIGATION, THE PATIENT'S SPECIFIC GENDER AND AGE WERE UNKNOWN. ON (B)(6) 2025, THE PATIENT UNDERWENT DISSECTION SURGERY IN DEPARTMENT OF CARDIAC SURGERY IN (B)(6) (THE SPECIFIC SURGICAL NAME AND PATIENT'S DIAGNOSIS WERE UNKNOWN). THE OPERATOR USED THE SUTURE (THE SPECIFIC CODE WAS UNKNOWN, AND THE SUSPECTED CODE WAS W8977) FOR PURSE-STRING SUTURING (THE SPECIFIC SUTURING TISSUE LAYER WAS UNKNOWN) DURING SURGERY. NO DEVICE FAILURE AND PATIENT INJURY EVENT WERE REPORTED DURING SURGERY. SUTURE BREAKAGE OCCURRED AFTER SURGERY (THE SPECIFIC EVENT WAS UNKNOWN). THE PATIENT DIED AT ABOUT 9:00 AM ON (B)(6) 2025. THE FOLLOWING INFORMATION WAS REQUESTED, BUT UNAVAILABLE: WOULD THE SURGEON LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? WHAT IS THE CAUSE OF DEATH? PLEASE PROVIDE THE PATIENT'S DEMOGRAPHIC INFORMATION INCLUDING AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE NAME OF INDEX SURGICAL PROCEDURE? THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WAS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE? (OPEN, LAPAROSCOPIC OR OTHER)? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE ORIGINALLY TIED (MULTIPLE KNOTS, SQUARE KNOT, ETC.)? WHAT INSTRUMENTS WERE USED IN THIS PROCEDURE TO HANDLE THE SUTURE? PLEASE DESCRIBE ANY MEDICAL/SURGICAL INTERVENTION REQUIRED FOR THIS SUTURE EVENT INCLUDING DATES AND RESULTS. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? WHAT SYMPTOMS DID THE PATIENT EXPERIENCE FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? LOT NUMBER? WILL PRODUCT BE RETURNED? IF SO, PLEASE PROVIDE THE RETURN DATE AND TRACKING INFORMATION. SURGEON¿S NAME? --RECEIVED INFORMATION AS FOLLOWING FROM SALES REP VIA PHONE TODAY. PLEASE REFER TO THE INFORMATION PROVIDED, OTHER INFORMATION REQUESTED IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21445282&pc=OZD | 1220648-2025-26276 | 2025-01-29 05:00 | Death | ABIOMED, INC. | 2025-02-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT SUFFERED BILATERAL LOSS OF DOPPLER PULSES FOLLOWING IMPELLA CP IMPLANT. IT WAS NOTED THAT THE PATIENT ALREADY HAD PERIPHERAL ARTERY DISEASE AND WAS ISCHEMIC WHILE ON INTRA-AORTIC BALLOON PUMP SUPPORT PRIOR TO IMPELLA IMPLANT. FOLLOWING IMPLANT, THE ISCHEMIA PERSISTED UNTIL THE FAMILY MADE THE DECISION TO WITHDRAW CARE. THE PATIENT PASSED AWAY SHORTLY AFTER. IT IS UNKNOW IF SUPPORT CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21445397&pc=OZD | 1220648-2025-26277 | 2025-01-27 05:00 | Death | ABIOMED, INC. | 2025-02-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED REGARDING A STUDY PATIENT THAT EXPERIENCED VENTRICULAR TACHYCARDIA, REFRACTORY TO LIDOCAINE, AND SUCCESSFULLY TREATED WITH DIRECT CURRENT CARDIOVERSION TO SINUS TACHYCARDIA, ALBEIT WITH FREQUENT ECTOPY. THE PATIENT PROGNOSIS WAS POOR. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26277 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G2 REPORT SOURCE; STUDY WAS MISSING FROM MANUFACTURER DEVICE REPORT 1220648-2025-26277. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21445440&pc=OZD | 1220648-2025-26281 | 2025-01-29 05:00 | Death | ABIOMED, INC. | 2025-02-24 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE AS ELECTIVE PLACEMENT FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT WAS UNDERGOING A LEFT VENTRICULAR THROMBECTOMY AND AN AORTIC VALVE REPLACEMENT (AVR) SURGERY. AFTER THE SURGERY, THE IMPELLA WAS DIRECTLY INSERTED VIA THE AORTA. THE LEFT VENTRICULAR PERFORATION OCCURRED WHILE THE CATHETER WAS INSERTED INTO THE LEFT VENTRICLE, AND ITS POSITION WAS ADJUSTED. AFTER REPAIR OF THE ADVERSE EVENT, THERE WAS NO FLOW WAS OBTAINED EVEN AFTER IMPELLA 5.5 WAS STARTED. REPAIR WAS ATTEMPTED AGAIN; HOWEVER, BLEEDING COULD NOT BE CONTROLLED. THE IMPELLA 5.5 WAS EXPLANTED IN THE OPERATING ROOM. THE PATIENT EXPIRED AS A RESULT OF THE ADVERSE EVENT. THE PHYSICIAN FURTHER COMMENTED THE FOLLOWING: THE REASON THE IMPELLA WAS INSERTED USING A DIRECT AORTA WAS THE CHEST WAS OPENED AFTER AVR, AND THE BLOOD VESSEL DIAMETER WAS NOT LARGE ENOUGH TO INSERT 5.5. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21445760&pc=FKX | 0008030665-2025-00437 | 2025-01-08 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-02-24 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT CONTACT REPORTED TO FRESENIUS THAT THE PATIENT EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS REPORTED THAT THE PATIENT'S DEATH WAS NOT PD RELATED. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. ADDITIONAL INFORMATION WAS OBTAINED DURING FOLLOW-UP WITH THE PATIENT'S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). PER THE PDRN, THE PATIENT EXPIRED ON (B)(6) 2025. THE PATIENT HAD COMPLETED TREATMENT AND WAS STILL ATTACHED TO THE CYCLER. IT WAS EXPLAINED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF CARDIOVASCULAR DISEASE WITH PREVIOUS STROKE (UNRELATED TO PD THERAPY) AND WAS DECLINING IN OVERALL HEALTH. THE NURSE INDICATED THAT THE PATIENT EXPIRED FROM CARDIAC ARREST WHICH WAS ATTRIBUTED TO THE PATIENT¿S PRE-EXISTING CARDIOVASCULAR COMORBID CONDITIONS. THE PDRN STATED THERE WERE NO ISSUES WITH THE PD TREATMENT AND STATED SEVERAL TIMES THAT THE DEATH WAS NOT RELATED TO PRODUCT OR PD THERAPY. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE.Event Description: A PERITONEAL DIALYSIS (PD) PATIENT CONTACT REPORTED TO FRESENIUS THAT THE PATIENT EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS REPORTED THAT THE PATIENT¿S DEATH WAS NOT PD RELATED. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. ADDITIONAL INFORMATION WAS OBTAINED DURING FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). PER THE PDRN, THE PATIENT EXPIRED ON (B)(6) 2025. THE PATIENT HAD COMPLETED TREATMENT AND WAS STILL ATTACHED TO THE CYCLER. IT WAS EXPLAINED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF CARDIOVASCULAR DISEASE WITH PREVIOUS STROKE (UNRELATED TO PD THERAPY) AND WAS DECLINING IN OVERALL HEALTH. THE NURSE INDICATED THAT THE PATIENT EXPIRED FROM CARDIAC ARREST WHICH WAS ATTRIBUTED TO THE PATIENT¿S PRE-EXISTING CARDIOVASCULAR COMORBID CONDITIONS. THE PDRN STATED THERE WERE NO ISSUES WITH THE PD TREATMENT AND STATED SEVERAL TIMES THAT THE DEATH WAS NOT RELATED TO PRODUCT OR PD THERAPY. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXITS BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH FROM CARDIAC ARREST. HOWEVER, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER OR ANY ISSUES WITH PD TREATMENT. THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED PRE-EXISTING CARDIAC COMORBIDITIES AND STROKE (UNRELATED TO DIALYSIS) WITH REPORT OF DECLINE IN OVERALL HEALTH. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT'S PRE-EXISTING CARDIAC COMORBID CONDITIONS AND STROKE PLACE THAT PATIENT AT HIGHER RISK OF MORTALITY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE CYCLER WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE DUE MISSING DOOR LOGO. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT ANY FAILURES OR PROBLEMS. NO FLUID LEAKS IN THE TEST CASSETTE OCCURRED DURING THE TREATMENT TEST. SYSTEM AIR LEAK, TEACH PUMP, AND VALVE ACTUATION TESTS WERE PERFORMED AND PASSED. AN INTERNAL VISUAL OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21445861&pc=QJZ | 3012172416-2025-00021 | 2024-11-05 05:00 | Death | XENIOS AG | 2025-02-24 05:00 | QJZ | NOVALUNG CONSOLE (LINUX BOARD) | Crack; Material Integrity Problem; Insufficient Information | Cardiac Arrest | K191407 |  | 1 | Event Description: A USER FACILITY REPORTED THAT A PATIENT CAME TO THE CARDIAC CATHETERIZATION LAB FOLLOWING A VENTRICULAR TACHYCARDIA AND VENTRICULAR ECTOPIC ARREST FOR INITIATION OF EXTRACORPOREAL CARDIOPULMONARY RESUSCITATION. THE NOVALUNG DEVICE WAS PREPPED TO HAND UP THE PATIENT CIRCUIT INTO THE STERILE FIELD TO START SUPPORT. WHEN HANDED UP TO THE STERILE FIELD, THE P2 PRESSURE CABLE BROKE AND WAS NOT DETECTED ON THE CONSOLE SCREEN. AN ATTEMPT WAS MADE TO PULL THE CORD AND SWITCH TO ANOTHER CABLE; HOWEVER, THE CABLE WAS STUCK IN THE DEVICE. WITHIN A FEW MINUTES OF ECMO SUPPORT THE P3 PRESSURE CABLE ALSO WAS NOT READING A PRESSURE, AND IT WAS FOUND TO BE IN THE XLUNG KIT, BUT NOT ABLE TO READ A PRESSURE. THIS CABLE ALSO REMAINED IN THE DEVICE. ECMO SUPPORT WAS NOT STOPPED DUE TO THE URGENCY OF SUPPORT AT THAT TIME. THE PATIENT WAS HOOKED UP TO BLOOD PRESSURE AND EKG MONITORS TO ENSURE BEST CARE PRACTICES WERE FOLLOWED. ALSO, THE NOVALUNG CONSOLE WAS READING A TOTAL FLOW TO PATIENT WITH RPMS. THERE WAS NO COLOR CHANGE THAT WAS NOTED ON ARTERIAL LIMB OF CIRCUIT UNTIL THE OXYGEN LEVEL WAS INCREASED TO MAXIMUM SUPPORT OF 100% OXYGEN. UPON ARRIVAL TO THE ICU, THE PATIENT PRESSURE CABLES WERE SWITCHED, AND ONCE CHANGED, THE CLINICIANS WERE ABLE TO NOTE THE NEGATIVE PRESSURE (P1) WAS -400 MMHG. ULTIMATELY, THE PATIENT EXPIRED WITHIN 5 HOURS OF ARRIVAL. THE DISPOSABLE WAS SENT BACK TO FRESENIUS DUE TO THE OXYGENATOR NOT FUNCTIONING (DEVICE AND DATE SHIPPED WERE NOT REPORTED). ARTERIAL BLOOD GASES WERE DRAWN ON THE POST-OXYGENATOR SIDE WITH AN FIO2 OF 100% AND THE HIGHEST YIELD WAS AN ARTERIAL OXYGEN OF 87 MMHG. AN ARTERIAL OXYGEN DIRECTLY FROM THE PATIENT WAS 89 MMHG WHICH INDICATED THE PATIENT'S LUNGS WERE DOING THE WORK AS IT WAS HIGHER THAN THE DEVICE.Event Description: A USER FACILITY REPORTED THAT A PATIENT CAME TO THE CARDIAC CATHETERIZATION LAB FOLLOWING A VENTRICULAR TACHYCARDIA AND VENTRICULAR ECTOPIC ARREST FOR INITIATION OF EXTRACORPOREAL CARDIOPULMONARY RESUSCITATION. THE NOVALUNG DEVICE WAS PREPPED TO HAND UP THE PATIENT CIRCUIT INTO THE STERILE FIELD TO START SUPPORT. WHEN HANDED UP TO THE STERILE FIELD, THE P2 PRESSURE CABLE BROKE AND WAS NOT DETECTED ON THE CONSOLE SCREEN. AN ATTEMPT WAS MADE TO PULL THE CORD AND SWITCH TO ANOTHER CABLE; HOWEVER, THE CABLE WAS STUCK IN THE DEVICE. WITHIN A FEW MINUTES OF ECMO SUPPORT THE P3 PRESSURE CABLE ALSO WAS NOT READING A PRESSURE, AND IT WAS FOUND TO BE IN THE XLUNG KIT, BUT NOT ABLE TO READ A PRESSURE. THIS CABLE ALSO REMAINED IN THE DEVICE. ECMO SUPPORT WAS NOT STOPPED DUE TO THE URGENCY OF SUPPORT AT THAT TIME. THE PATIENT WAS HOOKED UP TO BLOOD PRESSURE AND EKG MONITORS TO ENSURE BEST CARE PRACTICES WERE FOLLOWED. ALSO, THE NOVALUNG CONSOLE WAS READING A TOTAL FLOW TO PATIENT WITH RPMS. THERE WAS NO COLOR CHANGE THAT WAS NOTED ON ARTERIAL LIMB OF CIRCUIT UNTIL THE OXYGEN LEVEL WAS INCREASED TO MAXIMUM SUPPORT OF 100% OXYGEN. UPON ARRIVAL TO THE ICU, THE PATIENT PRESSURE CABLES WERE SWITCHED, AND ONCE CHANGED, THE CLINICIANS WERE ABLE TO NOTE THE NEGATIVE PRESSURE (P1) WAS -400 MMHG. ULTIMATELY, THE PATIENT EXPIRED WITHIN 5 HOURS OF ARRIVAL. THE DISPOSABLE WAS SENT BACK TO FRESENIUS DUE TO THE OXYGENATOR NOT FUNCTIONING (THE DATE AND DESTINATION OF THE SHIPPED PRODUCT WERE NOT REPORTED RESULTING IN NO RECORD OF RETURN OR EVALUATION). ARTERIAL BLOOD GASES WERE DRAWN ON THE POST-OXYGENATOR SIDE WITH AN FIO2 OF 100% AND THE HIGHEST YIELD WAS AN ARTERIAL OXYGEN OF 87 MMHG. AN ARTERIAL OXYGEN DIRECTLY FROM THE PATIENT WAS 89 MMHG WHICH INDICATED THE PATIENT'S LUNGS WERE DOING THE WORK AS IT WAS HIGHER THAN THE DEVICE. NO FURTHER INFORMATION WAS PROVIDED BY THE USER FACILITY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, D9 PLANT INVESTIGATION: NO PARTS WERE RETURNED FOR EVALUATION, NO MACHINE LOG FILES PROVIDED, AND NO SERIAL NUMBERS WERE PROVIDED BY THE REPORTING FACILITY. FOUR SIMILAR COMPLAINTS REPORTING A FUNCTIONALITY ISSUE HAVE BEEN REPORTED IN THE LAST 24 MONTHS. THE REASON FOR THE ABSENCE OR INCORRECT VALUES BEING DISPLAYED ON THE CONSOLE MAY BE ATTRIBUTED TO A DEFECTIVE CABLE, A FAULTY CONNECTION BETWEEN THE CABLE AND SENSOR OR A DEFECT IN THE SENSOR ITSELF. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21447661&pc=MIH | 3005580113-2025-00009 |  | Death | WILLIAM A. COOK AUSTRALIA, PTY LTD | 2025-02-24 05:00 | MIH | ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT PROXIMAL BODY | Adverse Event Without Identified Device or Use Problem |  |  |  | 1 | Event Description: PATIENT EXPIRED. A SCAN CONFIRMED A GRAFT SEAL SEPARATION OF THE ZFEN-P-2-34-137-R AND ZFEN-D-12-28-91-C. APPROXIMATELY 1 WEEK PRIOR TO THIS OCCURRENCE, THE PATIENT EXPERIENCED A SIGNIFICANT FALL DOWN STAIRS WHICH RESULTED IN HOSPITAL ADMISSION. IT IS UNKNOWN AT THIS TIME IF THE PATIENT WAS RELEASED HOME PRIOR TO THIS OCCURRENCE, OR WAS STILL ADMITTED INTO THE HOSPITAL FROM THE FALL. ACCORDING TO THE INITIAL REPORTER THE DATE OF EVENT WAS STATED TO BE "AROUND CHRISTMAS 2024".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21447686&pc=DSQ | 3007042319-2025-00552 | 2025-02-14 05:00 | Death | HEARTWARE, INC. | 2025-02-24 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT SUBSEQUENTLY EXPIRED, DUE TO A DRIVELINE (DL) INFECTION. THE DL HAD REACHED THE CHEST AND HEART TISSUES AND THE PATIENT COULD NOT BE SAVED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THE PATIENT EXPIRED DUE TO DRIVELINE INFECTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DRIVELINE INFECTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE, INCLUDING CARE OF THE DRIVELINE EXIT SITE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D4OUS : D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21447808&pc=DSQ | 2916596-2025-00845 | 2024-05-01 04:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿APPLICATIONS OF THE IMPELLA RP FLEX: SINGLE CENTER EXPERIENCE¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT VENTRICULAR FAILURE, PNEUMONIA AND DEATH. IMPELLA RP FLEX WAS EVALUATED FOR HEMODYNAMIC AND QUALITY MEASURES. 8 PATIENTS WERE IMPLANTED WITH THE RP FLEX: 1 FOR RIGHT VENTRICULAR (RV) PROPHYLAXIS IN MITRAL SURGERY, 1 FOR ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK, 3 FOR PERI-DURABLE VENTRICULAR ASSIST DEVICE FOR RIGHT VENTRICULAR DYSFUNCTION, 1 FOR POST CARDIOTOMY SHOCK FOLLOWING WEANING FROM EXTRACORPOREAL MEMBRANE OXYGENATION, 1 FOR BI-IMPELLA SUPPORT POST CORONARY ARTERY BYPASS GRAFT AND 1 FOR CARDIOMYOPATHY AND CARDIOGENIC SHOCK (CGS) WITH BIVENTRICULAR FAILURE. 3 OF THE 8 PATIENTS HAD REDUCED EJECTION FRACTION REQUIRING CONCOMITANT LEFT VENTRICULAR SUPPORT WITH A HM3. 4 OUT OF 8 PATIENTS HAD PRE-EXISTING GROUP 2 PULMONARY HYPERTENSION WITH MEAN PULMONARY ARTERIAL PRESSURE (PAP) OF 43 MMHG. MEAN PULMONARY ARTERY PULSATILITY INDEX AT BASELINE WAS 1.45, AND MEAN PAP FOR ALL PATIENTS WAS 37 MMHG. MEAN AGE WAS 63.4 YEARS AND 4 PATIENTS WERE FEMALE. 6 OUT OF 8 PATIENTS WERE ABLE TO PARTICIPATE IN PHYSICAL THERAPY WITHIN THE FIRST WEEK POST OPERATIVELY AND WERE OUT OF BED TO A CHAIR IN 2.7 DAYS. NONE OF THE PATIENTS REQUIRED RENAL REPLACEMENT THERAPY DURING DEVICE SUPPORT BUT 1 PATIENT RECEIVED RENAL REPLACEMENT THERAPY 3 WEEKS POST EXPLANT FOR FLUID MANAGEMENT. ONLY 1 PATIENT HAD PROLONGED INTUBATION DUE TO PNEUMONIA. MEAN DAYS ON RP FLEX SUPPORT WAS 4.8 DAYS (1-8 DAYS). NONE OF THE PATIENT¿S REQUIRED LONGSTANDING INOTROPES AND 4 OUT OF 8 PATIENTS REQUIRED SILDENAFIL AT DISCHARGE, AND 2 OUT OF 8 REQUIRED SILDENAFIL AT 30 DAYS POST DISCHARGE. 6 PATIENTS WERE SUCCESSFULLY DISCHARGED WITH MEAN POSTOPERATIVE LENGTH OF STAY (LOS) OF 21.7 DAYS, 13-DAY (8-16 DAYS) MEAN LOS FOR PATIENTS WITH PROPHYLACTIC RP FLEX AND MEAN LOS FOR RIGHT VENTRICULAR RESCUE PATIENTS OF 22.5 DAYS (8-53 DAYS). ONE PATIENT EXPIRED DUE TO BIVENTRICULAR FAILURE AND CGS WHILE 1 PATIENT REMAINED HOSPITALIZED. THE RP FLEX IS AN INTRACORPOREAL RV ASSIST DEVICE THAT ALLOWS FOR RV UNLOADING IN THOSE PATIENTS WHO REQUIRE SHORT TERM RV SUPPORT. IT IS BEST USED EARLY OR IN A PROPHYLACTIC SETTING, ALLOWING FOR STABLE HEMODYNAMICS IN CASES THAT WOULD OTHERWISE REQUIRE HIGH DOSE PRESSORS AND PROLONGED ICU CARE. APPROPRIATE USE OF THE RP FLEX OFFERS PATIENTS WHO OTHERWISE BE TOO HIGH RISK A CHANCE FOR RECOVERY. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF PUBLICATION (01MAY2024) AS THE DATE OF DATA COLLECTION WAS NOT PROVIDED. DATE OF DEATH IS UNKNOWN. AUTHOR INFORMATION: MODY, K., WALTERS, P., STEWART, S., SILVA, C., LANDERS, D., DUDIY, Y., BATSIDES, G., & ANDERSON, M. (2024). APPLICATIONS OF THE IMPELLA RP FLEX: SINGLE CENTER EXPERIENCE. ASAIO JOURNAL, 70, 159. HTTPS://DOI.ORG/10.1097/MAT.0000000000002073. HACKENSACK UNIVERSITY MEDICAL CENTER, HACKENSACK, NJ. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), DOCUMENT IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21447906&pc=OAE | 3002648230-2025-00125 | 2024-01-01 05:00 | Death | MEDTRONIC CRYOCATH LP | 2025-02-24 05:00 | OAE | ARCTIC FRONT ADVANCE® CARDIAC CRYOABLATION CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100010 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SINUS RHYTHM MAINTENANCE AFTER REPEAT CATHETER ABLATION. THE PRIMARY ENDPOINT WAS A COMPOSITE OF ALL-CAUSE MORTALITY, HEART FAILURE (HF) HOSPITALIZATIONS, OR STROKES WITHIN THREE YEARS AFTER THE INITIAL OR REPEAT CATHETER ABLATION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE IS NO ALLEGATION OF ABLATION PROCEDURE-DEATH RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THERE WERE PATIENTS WHO EXPERIENCED HF HOSPITALIZATIONS AND STROKES WITHIN THREE YEARS OF THE ABLATION PROCEDURES. HF HOSPITALIZATIONS WERE DEFINED AS AN ADMISSION TO A HOSPITAL FOR MORE THAN 24 HOURS DUE TO WORSENING HF, AND REQUIRED INTRAVENOUS MEDICATION FOR HF (INCLUDING DIURETICS, VASODILATORS, OR INOTROPIC AGENTS) OR AN INCREASE IN THE ORAL DIURETIC TREATMENT FOR HF. PROCEDURAL COMPLICATIONS INCLUDED ONE CARDIAC TAMPONADE, ONE MAJOR BLEEDING EVENT, GASTRIC HYPOMOTILITY, AND PHRENIC NERVE PALSY. THE STATUS OF THE CATHETERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IS SINUS RHYTHM MAINTENANCE AFTER REPEAT CATHETER ABLATION EFFECTIVE IN PATIENTS WITH ATRIAL FIBRILLATION AND HEART FAILURE WITH PRESERVED EJECTION FRACTION? JOURNAL OF CARDIOVASCULAR ELECTROPHYSIOLOGY. 2024; 35:2452¿2459. DOI: 10.1111/JCE.16464. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21447952&pc=DSQ | 2916596-2025-00896 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Bacterial Infection; Sepsis; Lethargy; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RENAL FAILURE AND SEPSIS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE SOURCE OF INFECTION WAS DRIVELINE (SERRATIA). POSITRON EMISSION TOMOGRAPHY (PET) SCAN DEMONSTRATED UPTAKE IN THE REGION OF THE APICAL COMPONENT OF THE DEVICE EXTENDING INTO THE LINGULA. SERRATIA MARCESCENS WAS IDENTIFIED AND HAD COLONIZED FOR 2 YEARS. THE PATIENT EXPERIENCED SUPRATHERAPEUTIC INTERNATIONAL NORMALIZED RATIO (INR), DIARRHEAL ILLNESS, AND A C-REACTIVE PROTEIN (CRP) THAT WAS GREATER THAN 200. FOR TREATMENT, 'ID INPUT' AND MEROPENEM WERE PERFORMED. THE RENAL DYSFUNCTION WAS UNDERLYING PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND WAS NOT DETERMINED TO BE RELATED TO DEVICE. BASELINE ESTIMATED GLOMERULAR FILTRATION RATE (EGFR) WAS 34. EGFR WAS 27 AT TIME OF ADMISSION AND THE PATIENT HAD AN ACUTE KIDNEY INJURY ON BACKGROUND OF CHRONIC KIDNEY DISEASE RELATED TO SEPSIS. SYMPTOMS INCLUDED DIARRHEAL ILLNESS FOR 8 DAYS PRIOR TO ADMISSION, LETHARGY, GENERALIZED ABDOMEN PAIN (FECAL PCR NEGATIVE), AND EPISODES OF DELIRIUM AT END STAGE. PATIENT WAS PALLIATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: CORRECTION: B3: DATE OF EVENT. MANUFACTURER¿S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT'S INFECTION AND SEPSIS COULD NOT BE CONCLUSIVELY DETERMINED. ADDITIONALLY, DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS INFECTION, SEPSIS, AND DEATH AS ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448073&pc=DSQ | 2916596-2025-00954 | 2024-10-10 04:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Post Traumatic Wound Infection; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT HEART FAILURE ON (B)(6) 2024. THE PATIENT HAD GASTROINTESTINAL (GI) BLEEDING, STERNAL WOUND INFECTION AND FAILURE TO THRIVE. THE PATIENT WAS ADMITTED TO HOSPICE. THE PATIENT WAS INOTROPE DEPENDENT PRIOR TO LEFT VENTRICULAR ASSIST DEVICE SURGERY AFTER CORONARY ARTERY BYPASS GRAFTING. A DEVICE RELATED ISSUE DID NOT CONTRIBUTE TO THE RIGHT HEART FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(4) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS INFECTION, BLEEDING, AND DEATH, AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 OF THE IFU, (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SEVERAL SECTIONS OF THE HEARTMATE 3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448214&pc=LWP | 2017865-2025-14928 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-14930, RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-14931. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO THE PATIENT'S DEATH. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. THE RESULTS OF THE ELECTRICAL AND STRESS/ENVIRONMENTAL TESTS WERE PERFORMED TO INDICATE THAT THE PACER IS EXHIBITING NORMAL DEVICE CHARACTERISTICS. A LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448231&pc=NVN | 2017865-2025-14930 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-14928, RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-14931. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN ONE PORTION FOR ANALYSIS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448236&pc=NVN | 2017865-2025-14931 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-14928, RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-14930. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448238&pc=MHX | 21448238 | 2025-02-10 05:00 | Death | NIHON KOHDEN | 2025-02-24 05:00 | MHX | NIHON KOHDEN | Insufficient Information | Respiratory Arrest |  |  | 1 | Event Description: A 58 Y/O MALE BROUGHT TO ED FOR C/O SOB, EPIGASTRIC PAIN, CHEST DISCOMFORT AND NAUSEA ON (B)(6)2025 AT 1700. PT BEING TREATED FOR CHEST PAIN, PNA, PLEURAL EFFUSION, AND SEVERE SEPSIS. IN ED, PATIENT WAS PLACED ON CARDIAC MONITOR (NIHON KOHDEN). CXR RESULTS: PNEUMONIC INFILTRATION LEFT LUNG WITH MODERATE TO LARGE LEFT SIDED PLEURAL EFFUSION. CTA PULMONARY WITH CONTRAST RESULTS: MODERATE PERICARDIAL EFFUSION. MODERATE LOCULATED PLEURAL EFFUSION ON LEFT. PULMONARY VENOUS CONGESTION. PNEUMONIC INFILTRATION THROUGHOUT LEFT LUNG. ANASARCA. EKG WITH ST DEPRESSION TO LOWER LEADS. LABS AT 1814: WBC 16.8, HGB 7.1, HCT 26.3, PLT 426, BNP 156, LACTATE 5.2, TROP 15. ABG AT 1956: PH 7.42, HCO3 12.1, PCO2 18.3. PT PLACED ON HFNC. LAST RECORDED HR 127 AND RR 21 BY NIHON KOHDEN AT 2002. VITAL SIGNS AT 2006: TEMP 37.5, HR 117, BP 159/131, SPOX 99% ON HFNC. AT 2024, SPO2 DROPPED TO 72% AND ALARM WARNING ALERT TRIGGERED. PT WAS FOUND PULSELESS AROUND 2040 WITH LEADS OFF. EVENT AND ALARM LOGS EXTRACTED. NO SUBSEQUENT ALARMS TRIGGERED AFTER 2024 UNTIL 2043 (19 MIN GAP). BASED ON DISCUSSION WITH BIOMED, NIHON KOHDEN ALARM SHOULD CONTINUE TO TRIGGER EVERY 2 MINUTES UNLESS IT WAS "PAUSED". IF SILENCE (WHICH CAN BE DONE AT BEDSIDE OR CENTRAL STATION), THE CLINICAL ALARMS WILL OVERRIDE THE SILENCE AND CONTINUE TO ALERT WHICH DID NOT OCCUR. PT WAS FOUND PULSELESS AROUND 2040. MEDICAL DEVICE WAS SEQUESTERED. BIOMED PERFORMED FUNCTIONAL AND OPERATIONAL TEST ON THE BSM 1700 AND G5. THIS INCLUDED ALL PARAMETER TESTING, ALARM TESTING, ELECTRICAL SAFETY, AND PHYSICAL INSPECTION WITH NO DAMAGE OR DEFECTS. IN ADDITION, BIOMED VERIFIED BSM 1700 AND G5 COMMUNICATION BACK TO CENTRAL STATION. RESULTS WERE THAT ALL PARAMETERS WERE FUNCTIONAL AND OPERATIONAL WITHIN MANUFACTURER'S SPECIFICATION. TESTING ALSO PERFORMED ON CENTRAL STATION. VERIFIED ALARMS, CONNECTIVITY TO SERVER AND COMMUNICATION TO THE PATIENT MONITOR INVOLVED IN THE INCIDENT. ALL TEST PASSED. EVENT LOG AND ALARM LOG REVIEWED AND ANALYZED. HOWEVER, LOG REVIEW IS INCOMPLETE WITHOUT FURTHER VENDOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448255&pc=MHX | 21448255 | 2025-02-10 05:00 | Death | NIHON KOHDEN | 2025-02-24 05:00 | MHX | NIHON KOHDEN | Insufficient Information | Respiratory Arrest |  |  | 1 | Event Description: A 58 Y/O MALE BROUGHT TO EMERGENCY DEPARTMENT FOR C/O SOB, EPIGASTRIC PAIN, CHEST DISCOMFORT AND NAUSEA ON (B)(6) 2025 AT 1700. PT BEING TREATED FOR CHEST PAIN, PNA, PLEURAL EFFUSION, AND SEVERE SEPSIS. IN ED, PATIENT WAS PLACED ON CARDIAC MONITOR (NIHON KOHDEN). CXR RESULTS: PNEUMONIC INFILTRATION LEFT LUNG WITH MODERATE TO LARGE LEFT SIDED PLEURAL EFFUSION. CTA PULMONARY WITH CONTRAST RESULTS: MODERATE PERICARDIAL EFFUSION. MODERATE LOCULATED PLEURAL EFFUSION ON LEFT. PULMONARY VENOUS CONGESTION. PNEUMONIC INFILTRATION THROUGHOUT LEFT LUNG. ANASARCA. EKG WITH ST DEPRESSION TO LOWER LEADS. LABS AT 1814: WBC 16.8, HGB 7.1, HCT 26.3, PIT 426, BNP 156, LACTATE 5.2, TROP 15. ABG AT 1956: PH 7.42, HCO3 12.1, PCO2 18.3. PT PLACED ON HFNC. LAST RECORDED HR 127 AND RR 21 BY NIHON KOHDEN AT 2002. VITAL SIGNS AT 2006: TEMP 37.5, HR 117, BP 159/131, SPOX 99% ON HFNC. AT 2024, SPO2 DROPPED TO 72% AND ALARM WARNING ALERT TRIGGERED. PT WAS FOUND PULSELESS AROUND 2040 WITH LEADS OFF. EVENT AND ALARM LOGS EXTRACTED. NO SUBSEQUENT ALARMS TRIGGERED AFTER 2024 UNTIL 2043 (19 MIN GAP). BASED ON DISCUSSION WITH BIOMED, NIHON KOHDEN ALARM SHOULD CONTINUE TO TRIGGER EVERY 2 MINUTES UNLESS IT WAS "PAUSED". IF SILENCE (WHICH CAN BE DONE AT BEDSIDE OR CENTRAL STATION), THE CLINICAL ALARMS WILL OVERRIDE THE SILENCE AND CONTINUE TO ALERT WHICH DID NOT OCCUR. PT WAS FOUND PULSELESS AROUND 2040. MEDICAL DEVICE WAS SEQUESTERED. BIOMED PERFORMED FUNCTIONAL AND OPERATIONAL TEST ON THE BSM 1700 AND G5. THIS INCLUDED ALL PARAMETER TESTING, ALARM TESTING, ELECTRICAL SAFETY, AND PHYSICAL INSPECTION WITH NO DAMAGE OR DEFECTS. IN ADDITION, BIOMED VERIFIED BSM 1700 AND G5 COMMUNICATION BACK TO CENTRAL STATION. RESULTS WERE THAT ALL PARAMETERS WERE FUNCTIONAL AND OPERATIONAL WITHIN MANUFACTURER'S SPECIFICATION. TESTING ALSO PERFORMED ON CENTRAL STATION. VERIFIED ALARMS, CONNECTIVITY TO SERVER AND COMMUNICATION TO THE PATIENT MONITOR INVOLVED IN THE INCIDENT. ALL TEST PASSED. EVENT LOG AND ALARM LOG REVIEWED AND ANALYZED. HOWEVER, LOG REVIEW IS INCOMPLETE WITHOUT FURTHER VENDOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448266&pc=MIH | 3005580113-2025-00012 |  | Death | WILLIAM A. COOK AUSTRALIA, PTY LTD | 2025-02-24 05:00 | MIH | ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT PROXIMAL BODY | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: PATIENT EXPIRED. A SCAN CONFIRMED A GRAFT SEAL SEPARATION OF THE ZFEN-P-2-34-137-R AND ZFEN-D-12-28-91-C. APPROXIMATELY 1 WEEK PRIOR TO THIS OCCURRENCE, THE PATIENT EXPERIENCED A SIGNIFICANT FALL DOWN STAIRS WHICH RESULTED IN HOSPITAL ADMISSION. IT IS UNKNOWN AT THIS TIME IF THE PATIENT WAS RELEASED HOME PRIOR TO THIS OCCURRENCE, OR WAS STILL ADMITTED INTO THE HOSPITAL FROM THE FALL. ACCORDING TO THE INITIAL REPORTER THE DATE OF EVENT WAS STATED TO BE "AROUND CHRISTMAS 2024."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448267&pc=MKJ | 1220908-2025-00748 | 2025-02-12 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-24 05:00 | MKJ | X SERIES | Failure to Deliver Shock/Stimulation | Ventricular Fibrillation | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT FIRE SERVICES HAD SUCCESSFULLY DELIVERED ENERGY TO THE PATIENT BEFORE DISCONNECTING THE CPR STAT-PADZ FROM THEIR AED AND CONNECTING THEM TO THIS X SERIES DEVICE. COMPLAINANT INDICATED THAT THE CLINICIAN CONNECTED THE SAME CPR STAT-PADZ TO THE ORIGINAL AED TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL CANADA FOR EVALUATION. THE CUSTOMER'S REPORT WAS NOT REPLICATED. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING USING A KNOWN GOOD TEST BATTERY WITHOUT DUPLICATING THE REPORT. REVIEW OF THE DEVICE LOG FOUND THE DEVICE WAS POWERED ON VIA BATTERY POWER ONLY AND IMMEDIATELY IDENTIFIED LOW BATTERY AND BATTERY CALIBRATION REQUIRED PROMPTS. A SHOCK ADVISED PROMPT WAS GIVEN AND THE DEVICE BEGAN CHARGING; HOWEVER, THE CHARGE WAS DISARMED WHEN THE USER REMOVED THE PADS. THE DEVICE WAS UNABLE TO CHARGE TO THE TARGET ENERGY. THE BATTERY USED DURING THE EVENT WAS RETURNED (SN (B)(6)) AND THE BATTERY LOG RECORDED THE BATTERY WAS PLACED INTO THE CHARGER THE DAY BEFORE THE EVENT WITH 9% CHARGE, BUT DID NOT CHARGE. THE BATTERY RECORDED IT WAS BETWEEN 6% AND 9% AT THE START OF THE EVENT. AS THE DEVICE WAS CHARGING FOR THE SHOCK ADVISED PROMPT THE BATTERY DEPLETED TO 3% BEFORE THE DEVICE SHUT DOWN DUE TO A LOW BATTERY. THE BATTERY WAS RETURNED TO ZOLL AND WAS ABLE TO CHARGE SUCCESSFULLY. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448272&pc=PNJ | 9612164-2025-00995 | 2025-01-27 05:00 | Death | MEDTRONIC IRELAND | 2025-02-24 05:00 | PNJ | MICRA¿ VR2 | Failure to Capture; Capturing Problem | Bradycardia; Cardiac Arrest; Loss of consciousness | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT FIVE DAYS AFTER THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT WAS FOUND UNRESPONSIVE, AND HIS HEART RATE WAS NOTED TO BE FORTY BEATS PER MINUTE. IT WAS NOTED THAT THE THRESHOLDS WERE RISING STEADILY AFTER THE IMPLANT AND LOSS OF CAPTURED WAS SUSPECTED. THE PATIENT WAS CODED AND EVENTUALLY PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21448275&pc=MHX | 21448275 | 2025-02-10 05:00 | Death | NIHON KOHDEN | 2025-02-24 05:00 | MHX | NIHON KOHDEN | Insufficient Information | Respiratory Arrest |  |  | 1 | Event Description: 58 Y/O MALE BROUGHT TO EMERGENCY DEPARTMENT FOR C/O SOB, EPIGASTRIC PAIN, CHEST DISCOMFORT AND NAUSEA ON (B)(6) 2025 AT 1700. PT BEING TREATED FOR CHEST PAIN, PNA, PLEURAL EFFUSION, AND SEVERE SEPSIS. IN ED, PATIENT WAS PLACED ON CARDIAC MONITOR (NIHON KOHDEN). CXR RESULTS: PNEUMONIC INFILTRATION LEFT LUNG WITH MODERATE TO LARGE LEFT SIDED PLEURAL EFFUSION. CTA PULMONARY WITH CONTRAST RESULTS: MODERATE PERICARDIAL EFFUSION. MODERATE LOCULATED PLEURAL EFFUSION ON LEFT. PULMONARY VENOUS CONGESTION. PNEUMONIC INFILTRATION THROUGHOUT LEFT LUNG. ANASARCA. EKG WITH ST DEPRESSION TO LOWER LEADS. LABS AT 1814: WBC 16.8, HGB 7.1, HCT 26.3, PLT 426, BNP 156, LACTATE 5.2, TRAP 15. ABG AT 1956: PH 7.42, HCO312.1, PCO218.3. PT PLACED ON HFNC. LAST RECORDED HR 127 AND RR 21 BY NIHON KOHDEN AT 2002. VITAL SIGNS AT 2006: TEMP 37.5, HR 117, BP 159/131, SPOX 99% ON HFNC. AT 20 SPO2 DROPPED TO 72% AND ALARM WARNING ALERT TRIGGERED. PT WAS FOUND PULSELESS AROUND 2040 WITH LEADS OFF. EVENT AND ALARM LOGS EXTRACTED. NO SUBSEQUENT ALARMS TRIGGERED AFTER 2024 UNTIL 2043 (19 MIN GAP). BASED ON DISCUSSION WITH BIOMED, NIHON KOHDEN ALARM SHOULD CONTINUE TO TRIGGER EVERY 2 MINUTES UNLESS IT WAS "PAUSED". IF SILENCE (WHICH CAN BE DONE AT BEDSIDE OR CENTRAL STATION), CLINICAL ALARMS WILL OVERRIDE THE SILENCE AND CONTINUE TO ALERT WHICH DID NOT OCCUR. PT WAS FOUND PULSELESS AROUND 2040. MEDICAL DEVICE WAS SEQUESTERED. BIOMED PERFORMED FUNCTIONAL AND OPERATIONAL TEST ON THE BSM 1700 AND G5. THIS INCLUDED ALL PARAMETER TESTING, ALARM TESTING, ELECTRICAL SAFETY, AND PHYSICAL INSPECTION WITH NO DAMAGE OR DEFECTS. IN ADDITION, BIOMED VERIFIED BSM 1700 AND G5 COMMUNICATION BACK TO CENTRAL STATION. RESULTS WERE THAT ALL PARAMETERS WERE FUNCTIONAL AND OPERATIONAL WITHIN MANUFACTURER'S SPECIFICATION. TESTING ALSO PERFORMED ON CENTRAL STATION. VERIFIED ALARMS, CONNECTIVITY TO SERVER AND COMMUNICATION TO THE PATIENT MONITOR INVOLVED IN THE INCIDENT. ALL TEST PASSED. EVENT LOG AND ALARM LOG REVIEWED AND ANALYZED. HOWEVER, LOG REVIEW IS INCOMPLETE WITHOUT FURTHER VENDOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21449114&pc=CBK | 2518422-2025-102291 | 2025-02-20 05:00 | Death | RESPIRONICS, INC. | 2025-02-24 05:00 | CBK | TRILOGY EVO | Insufficient Information | Respiratory Failure | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT DIED WHILE ON VENTILATION THERAPY WITH A TRILOGY EVO VENTILATOR. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY AN UNINTENTIONAL USER ERROR BASED ON THE REPORTED STATEMENT THAT THE PATIENT CAME HOME FROM THE HOSPITAL AND WAS ON THE UNIT AND THE FAMILY SAID THEY HAD NO IDEA HOW TO USE OR WORK THE DEVICE (VENTILATOR). THE PATIENT PASSED AWAY. DEVICE EVALUATION AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, DEVICE CAUSE OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE INFORMATION PROVIDED. THE CAUSE OF THE PATIENT¿S DEATH IS CURRENTLY UNKNOWN. INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION HAS BEEN COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21449633&pc=NIQ | 2124215-2025-11030 | 2025-02-10 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-24 05:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | Vascular Dissection; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT A DISSECTION OCCURRED AND THE PATIENT DIED. THE 90% STENOSED TARGET LESION, LOCATED IN THE SEVERELY TORTUOUS AND MILDLY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY (LAD), WAS PRE-DILATED WITH A 3.00 X 8 MM NON-COMPLIANT (NC) BALLOON. A 3.00 X 28 MM PROMUS PREMIER WAS THEN DEPLOYED AT 12 ATM AND POST DILATED WITH A 3.00 X 8 MM NC BALLOON AT 12 ATM. A MAJOR, FLOW LIMITING CORONARY ARTERY DISSECTION WAS NOTED. THE PHYSICIAN ATTEMPTED TO COVER THE DISSECTION, HOWEVER, THE PATIENT COLLAPSED AND DIED DESPITE CARDIOPULMONARY RESUSCITATION (CPR) ATTEMPTS. THE OFFICIAL CAUSE OF DEATH WAS HEART FAILURE. PER THE ASSESSMENT OF THE PHYSICIAN, THE DISSECTION AND DEATH WERE DUE TO OVER DILATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21449715&pc=LOX | 2124215-2025-11060 | 2025-01-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-24 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Extravasation; Low Blood Pressure/ Hypotension; Perforation; Pericardial Effusion | K113220, |  | 1 | Event Description: IT WAS REPORTED THAT PERORATION AND DEATH OCCURRED. ON (B)(6) 2025, THE PATIENT PRESENTED WITH STABLE ANGINA. THE SAME DAY THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION USING CUTTING BALLOON ANGIOPLASTY IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD). THE LESION WAS PREDILATED WITH A 15 MM X 3.00 MM WOLVERINE CORONARY CUTTING BALLOON AT THE LAD. INTRAVASCULAR ULTRASOUND (IVUS) WAS PERFORMED, AND ADDITIONAL DILATATION WAS PERFORMED WITH A 3.50 MM X 15 MM EMERGE BALLOON AND A 4.00 MM X 8 MM EMERGE BALLOON. AN ELLIS CLASS 3 PERFORATION IN THE PROXIMAL LAD OCCURRED WHICH WAS CONSIDERED TO HAVE BEEN CAUSED BY THE 3.50 X 15MM EMERGE BALLOON. THE PERFORATION WAS COVERED INITIALLY WITH A 4.00 X 20 MM NON-BOSTON SCIENTIFIC (NON-BSC) STENT AND PERICARDIOCENTESIS PERFORMED. A 3.00 X 20 MM NON-BSC STENT WAS IMPLANTED DISTAL TO THE INITIAL COVER STENT IN VIEW OF PERSISTENT EXTRAVASATION. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED DUE TO PROFOUND HYPOTENSION. THE PATIENT WAS INTUBATED, AND VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WAS INSERTED. THE LEFT MAIN (LM) WAS RE-ENGAGED VIA THE LEFT FEMORAL ARTERY (LFA) USING AN 8FR NON-BSC GUIDING CATHETER. IT WAS SEEN THAT THE PERFORATION WAS SEALED, AND THE FIRST DIAGONAL WAS LOST BUT THERE WAS SATISFACTORY IVUS AND ANGIOGRAPHIC RESULT. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SHOWED RESTRICTION AND A CALCIFIED AORTIC VALVE. BALLOON VALVULOPLASTY WAS PERFORMED WITH A 20 MM NON-BSC BALLOON. A REPERFUSION SHEATH 6FR WAS INSERTED TO THE RIGHT SUPERFICIAL FEMORAL ARTERY (SFA). THE PATIENT WAS ON ECMO WITH LOW FLOWS AND CHUGGING. THE PERICARDIAL DRAIN HAD BEEN REMOVED. THERE WAS CONCERN FOR RIGHT-SIDED CREPITUS AND A CHEST X RAY WAS OBTAINED. THE PATIENT HAD WORSENING PRESSOR REQUIREMENTS STATUS POST 7 LITERS OF INTRAVENOUS FLUID PROMPTING TRANSFUSION OF MULTIPLE BLOOD PRODUCTS. THE PATIENT SUFFERED A CARDIAC ARREST AND RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED AFTER 1 ROUND OF CPR, EPI AND BICARB. BEDSIDE CARDIAC ULTRASOUND SHOWED MODERATE PERICARDIAL EFFUSION. THE NEXT DAY, THE PATIENT CONDITION WAS FURTHER COMPLICATED BY ECMO CHUGGING AND ELEVATED CENTRAL VENOUS PRESSURE (CVP). ADDITIONALLY, THE PATIENT EXPERIENCED AN ACUTE ISCHEMIC STROKE OF THE LEFT OCCIPITAL AND CEREBELLAR LOBES WITHOUT THROMBOLYSIS OR THROMBECTOMY, AND DEVELOPED THROMBOCYTOPENIA, LEADING TO THE DISCONTINUATION OF HEPARIN. ON (B)(6) 2025, THE PATIENT SUBSEQUENTLY UNDERWENT REPEAT PERICARDIAL DRAINAGE WITH PIGTAIL PLACEMENT AND SWAN GANZ INSERTION IN CATHETERIZATION LAB. ON (B)(6) 2025, THE PATIENT NOTED TO HAVE WORSENING PRESSOR REQUIREMENTS, WORSENING LACTATE, WORSENING RENAL FUNCTION, AND DEVELOPED SUSTAINED VENTRICULAR TACHYCARDIA. THE PATIENT WAS STARTED ON AMIODARONE. THE PATIENT WAS WEANED OFF SEDATION WITH NO NEUROLOGIC RESPONSE NOTED OVER COURSE OF 24 HOURS. ON (B)(6) 2025, THE PATIENT WAS NOTED TO BE ASYSTOLE ON MONITOR. THE ECMO CIRCUIT WAS CLAMPED, AND THE VENTILATOR STOPPED. THE SUBJECT WAS PRONOUNCED AS DECEASED. Manufacturer Narrative: G4 PMA/510(K): K163174.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21449798&pc=OZD | 1220648-2025-26292 | 2025-01-25 05:00 | Death | ABIOMED, INC. | 2025-02-24 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A/AN (UNKNOWN AGE) PATIENT WITH POST-CARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). PRIOR TO IMPELLA MCS IT WAS NOTED THE PATIENT WAS ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO), INOTROPES/VASOPRESSORS AND RESPIRATORY (BILEVEL POSITIVE AIRWAY PRESSURE [BIPAP] OR CONTINUOUS POSITIVE AIRWAY PRESSURE [CPAP]) SUPPORT. APPROXIMATELY THREE DAYS AFTER START OF THE IMPELLA MCS A CALL FROM PERFUSIONIST WAS RECEIVED DUE TO AN OBSERVED SHORT BREAKDOWN FROM AUTOMATED IMPELLA CONTROLLER (AIC) WITH INDEPENDENT RESTART AND ALARM "CONSOLE FAILURE¿ AND NOTED ¿NOW PUMP WORKS WELL.¿ IT WAS RECOMMENDED THAT THE AIC BE TAKEN OUT OF SERVICE, WHICH APPEARS TO HAVE OCCURRED AFTER THE IMPLANT EXPERIENCE (EXACT DATE IS UNCLEAR, UNKNOWN). THERE WERE NO ADVERSE EFFECTS OR HARM TO THE PATIENT BECAUSE OF THE AIC CONSOLE FAILURE BASED ON THE REPORTED INFORMATION FROM THE REPORTER. THE PATIENT CONTINUED ON THE IMPELLA 5.5 MCS. THE PATIENT WOULD EXPIRE THREE DAYS LATER, HOWEVER, AS CARE WAS WITHDRAWN. ADDITIONAL INFORMATION WAS SOUGHT REGARDING DETAILS AROUND THE PATIENT¿S DEMISE AND THE FOLLOWING, ALTHOUGH LIMITED, WAS NOTED: REGARDING THE AIC, AFTER ITS INDEPENDENT REBOOT, THE PUMP MOTOR RESUMED AT THE PREVIOUS FLOW LEVEL, AND IT WAS STATED THIS DID NOT HAVE ANYTHING TO DO WITH THE PATIENT¿S DEMISE. IT WAS NOTED THAT THERAPY WAS STOPPED (CARE WAS WITHDRAWN) BECAUSE OF ¿OTHER DISEASE/COMORBIDITIES, WHICH WERE NOT HEART RELATED,¿ AND ¿THERE WAS NO WAY FOR A GOOD OUTCOME.¿ THESE OTHER ISSUES WERE NOTED AS MULTI-ORGAN FAILURE, MASSIVE CEREBRAL HEMORRHAGE AND INTESTINAL ISCHEMIA. IT IS UNKNOWN, HOWEVER, EXACTLY WHEN THESE ADVERSE EVENTS OCCURRED. CARE WAS WITHDRAWN FOR ADVERSE EXPERIENCES NOT CAUSED BY THE AIC AND GIVEN THE INFORMATION, AT LEAST ONE OF THE ADVERSE EVENTS IS LISTED AS A POTENTIAL OUTCOME WITH USE OF THE IMPELLA PUMP (CEREBRAL VASCULAR ACCIDENT CVA/CEREBRAL HEMORRHAGE). WITH NO ALTERNATIVE CAUSE FOR THE CVA EVENT IDENTIFIED AS THE LIKELY CAUSE, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA 5.5 AS AN ASSOCIATED FACTOR IN THE PATIENT¿S DEMISE. MEDICAL SAFETY REVIEW OF THE EVENT NOTED FOR THE AIC, THERE WAS HEMODYNAMIC SUPPORT MAINTAINED AND REGARDING THE IMPELLA 5.5 DEVICE, HOWEVER, THERE IS NOT ENOUGH DATA TO ASSESS THE CVA IF IMPELLA PUMP RELATED; IT CANNOT BE EXCLUDED FROM THE OUTCOME (DEMISE) BASED ON THE LIMITED, INCOMPLETE DETAILS. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26292 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE DATE OF ONSET FOR THE CEREBRAL HEMORRHAGE IS UNKNOWN AT TIME OF THIS MDR WRITING. PROVISIONAL DATES FOR BOTH THE AWARE AND BEST ESTIMATE EVENT DATES ARE BEING ALIGNED, CONSERVATIVELY CAPTURED AS THE SAME DATES FOR THE AUTOMATED IMPELLA CONTROLLER. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO MANUFACTURING DEVICE REPORTS ASSOCIATED WITH THIS EVENT. REFER TO THE FOLLOWING MEDICAL DEVICE REPORT FOR AUTOMATED IMPELLA CONTROLLER (B)(6) WITH DATE OF EVENT 25-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR IMPELLA 5.5 (B)(6) WITH DATE OF EVENT 25-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). (THIS REPORT) INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: 5.5 WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21450269&pc=DSQ | 2916596-2025-01093 | 2023-11-01 04:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "NEW CHALLENGE TO HM3 WITH LOWERING ANTICOAGULATION REGIMEN: LONG TERM FOLLOW UP SINGLE CENTER EXPERIANCE" THAT HEARTMATE 3 IS ASSOCIATED WITH STROKE, BLEEDING, RENAL DYSFUNCTION, DRIVELINE INFECTION, RIGHT VENTRICULAR DYSFUNCTION, DEVICE THROMBOSIS, AND DEATH. 58 PATIENTS UNDERWENT HEARTMATE 3 (HM3) IMPLANTATION BETWEEN APR2016 AND NOV2023 WITH A MEDIAN FOLLOW UP OF 40 MONTHS (13- 84) TO STUDY THE ANTICOAGULATION REGIMEN AND EXAMINE THE MID-TERM OUTCOMES FOR VENTRICULAR ASSIST DEVICE (VAD) PATIENTS WHO RECEIVED ACETYLSALICYLIC ACID (ASA)/WARFARIN COMPARED TO THOSE WHO RECEIVED WARFARIN ALONE. THIS IS A RETROSPECTIVE CHART REVIEW OF ALL LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS WHO WERE IMPLANTED AT THE CENTER. INFORMED CONSENT WAS TAKEN. 4 WERE EXCLUDED AS THEY WERE NOT UNDER ANY SORT OF ANTICOAGULATION DUE TO COAGULOPATHY. OF THE REMAINING, 30 PATIENTS RECEIVED WARFARIN ALONE AND 24 RECEIVED ASA/WARFARIN. THERE WAS NO DIFFERENCE BETWEEN GROUPS REGARDING BASELINE CHARACTERISTICS. THE ASA/WARFARIN GROUP HAD HIGHER PULMONARY VASCULAR RESISTANCE (PVR) (P= 0.046). EARLY POST-OPERATIVE RESULTS (WITHIN 30 DAYS) SHOWED NO DIFFERENCES IN STROKE (P= 0.36) OR BLEEDING (P= 0.73). THE LONG-TERM OUTCOME, GASTROINTESTINAL (GI) BLEEDING OCCURRED IN 6 PATIENTS IN THE WARFARIN GROUP COMPARED TO 2 IN OTHER GROUP (P= 0.27). 4 PATIENTS IN THE WARFARIN GROUP AND 1 IN THE ASA/WARFARIN GROUP EXPERIENCED STROKE. 5 PATIENTS IN THE WARFARIN GROUP AND 5 IN THE ASA/WARFARIN GROUP EXPERIENCED BLEEDING. 6 PATIENTS IN THE WARFARIN GROUP AND 3 IN THE ASA/WARFARIN GROUP REQUIRED DIALYSIS. 6 PATIENTS IN THE WARFARIN GROUP AND 6 IN THE ASA/WARFARIN GROUP HAD DRIVELINE INFECTIONS. 28 PATIENTS IN THE WARFARIN GROUP AND 23 IN THE ASA/WARFARIN GROUP EXPERIENCED RIGHT VENTRICULAR DYSFUNCTION. DEVICE THROMBOSIS OCCURRED IN 1 PATIENT IN THE ASA/WARFARIN GROUP AND NONE IN THE WARFARIN GROUP. MORTALITY OCCURRED IN 10 PATIENTS: 9 IN THE WARFARIN GROUP AND 1 IN THE OTHER GROUP (P= 0.004). SURVIVAL AT FIRST AND LAST FOLLOW UP WAS 83% AND 79% IN THE WARFARIN GROUP, AND 100%, AND 96% IN THE OTHER GROUP. THIS STUDY FOUND THERE WAS A LONG-TERM SURVIVAL BENEFIT FOR PATIENT WHO RECEIVED BOTH ASA AND WARFARIN. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01NOV202. FERAS H. KHALIEL1, FAISAL ALSHAMDIN2, AHMED K. ALSULBUD1, AHLAM ALSOMALI3, ALWALEED ALTHOBAITI1, MUSAAD ALHUSSEIN4, JEHAD ALBURAIKI4, BANDAR ALAMRO4; 1CARDIAC SURGERY, KFSH&RC-R, RIYADH, SAUDI ARABIA, 2PRINCE MOHAMMAD BIN ABDULAZIZ MEDICAL CITY, ALJOUF, SAUDI ARABIA, 3KFSH&RC-R, RIYADH, SAUDI ARABIA, 4CARDIOLOGY, KFSH&RC-R, RIYADH, SAUDI ARABIA THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21450947&pc=PNJ | 9612164-2025-00996 | 2025-01-02 05:00 | Death | MEDTRONIC IRELAND | 2025-02-24 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Unspecified Infection; Cardiac Tamponade; Cardiac Perforation; Thrombosis/Thrombus; Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRADITIONAL TRANSVENOUS OR LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION IN PATIENTS WITH END STAGE RENAL DISEASE (ESRD). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED THROMBOSIS, DEVICE INFECTION, PERICARDIAL COMPLICATIONS (HEMOPERICARDIUM, CARDIAC TAMPONADE), MYOCARDIAL COMPLICATIONS (MYOCARDIAL INJURY, CARDIAC PERFORATION), UNKNOWN VASCULAR INJURIES, FISTULA, HEMATOMA, RESPIRATORY FAILURE, MECHANICAL VENTILATION, PNEUMOTHORAX, ACUTE KIDNEY INJURY, AND PATIENTS THAT REQUIRED BLOOD TRANSFUSIONS. THERE WERE DEVICES WHICH REQUIRED REVISIONS FOR UNKNOWN REASONS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/72 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LEADLESS PACEMAKER VS. TRANSVENOUS PACEMAKER IN END STAGE KIDNEY DISEASE: INSIGHTS FROM THE NATIONWIDE READMISSION DATABASE. JOURNAL OF CLINICAL MEDICINE. 2025, 14, 202. DOI: 10.3390/JCM14010202 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21451070&pc=NVZ | 2182208-2025-00967 | 2025-01-02 05:00 | Death | MEDTRONIC, INC. | 2025-02-24 05:00 | NVZ | MEDTRONIC IPG | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Unspecified Infection; Cardiac Tamponade; Cardiac Perforation; Thrombosis/Thrombus; Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRADITIONAL TRANSVENOUS OR LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION IN PATIENTS WITH END STAGE RENAL DISEASE (ESRD). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED THROMBOSIS, DEVICE INFECTION, PERICARDIAL COMPLICATIONS (HEMOPERICARDIUM, CARDIAC TAMPONADE), MYOCARDIAL COMPLICATIONS (MYOCARDIAL INJURY, CARDIAC PERFORATION), UNKNOWN VASCULAR INJURIES, FISTULA, HEMATOMA, RESPIRATORY FAILURE, MECHANICAL VENTILATION, PNEUMOTHORAX, ACUTE KIDNEY INJURY, AND PATIENTS THAT REQUIRED BLOOD TRANSFUSIONS. THERE WERE DEVICES WHICH REQUIRED REVISIONS FOR UNKNOWN REASONS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/72 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LEADLESS PACEMAKER VS. TRANSVENOUS PACEMAKER IN END STAGE KIDNEY DISEASE: INSIGHTS FROM THE NATIONWIDE READMISSION DATABASE. JOURNAL OF CLINICAL MEDICINE. 2025, 14, 202. DOI: 10.3390/JCM14010202 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21451204&pc=CBK | 2518422-2025-102292 | 2025-02-20 05:00 | Death | RESPIRONICS, INC. | 2025-02-24 05:00 | CBK | TRILOGY EVO | Insufficient Information | Respiratory Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT DIED WHILE ON VENTILATION THERAPY WITH A TRILOGY EVO VENTILATOR. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY AN UNINTENTIONAL USER ERROR BASED ON THE REPORTED STATEMENT THAT THE PATIENT CAME HOME FROM THE HOSPITAL AND WAS ON THE UNIT AND THE FAMILY SAID THEY HAD NO IDEA HOW TO USE OR WORK THE DEVICE (VENTILATOR). THE PATIENT PASSED AWAY. DEVICE EVALUATION AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, DEVICE CAUSE OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE INFORMATION PROVIDED. THE CAUSE OF THE PATIENT¿S DEATH IS CURRENTLY UNKNOWN. INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION HAS BEEN COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21451396&pc=NVZ | 2017865-2025-14946 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14947, 2017865-2025-14948. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR PACEMAKER SYSTEM. Manufacturer Narrative: THE DEVICE WAS RETURNED WITHOUT REPORTED MALFUNCTION.THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21451451&pc=NVN | 2017865-2025-14948 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14946, 2017865-2025-14947. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR PACEMAKER SYSTEM. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL EXAMINATIONS DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21451502&pc=NVN | 2017865-2025-14947 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14946, 2017865-2025-14948. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR PACEMAKER SYSTEM. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE FOR ANALYSIS. VISUAL EXAMINATION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21452820&pc=DSQ | 2916596-2025-01364 | 2025-02-18 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY, THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME BEING DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE HEALTHCARE PROVIDER.Event Description: THE PATIENT WAS FOUND UNRESPONSIVE AT HOME. Manufacturer Narrative: CORRECTION: SECTION D4, D6, H4. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE II LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION G6 (TYPE OF REPORT): CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION D4: DEVICE LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21452921&pc=DSQ | 2916596-2025-01010 | 2025-02-02 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO CARDIAC ARREST. THE PATIENT WAS AT HOME WHEN EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED FOR UNRESPONSIVENESS, AND CARDIOPULMONARY RESUSCITATION (CPR) INITIATED BY EMS. THE PATIENT WAS TRANSPORTED TO EMERGENCY ROOM AND SUSTAINED MULTIPLE ROUNDS OF CPR AND ADVANCED CARDIAC LIFE SUPPORT (ACLS). THE PUMP WAS SHOWING LOW FLOWS AT 09:17 WITH PUMP STOPPING IN EMERGENCY ROOM AFTER THE PATIENT PASSED AWAY. LOG FILES CAPTURED INTERMITTENT LOW FLOW FLAGS AND LOW FLOW ALARMS BETWEEN (B)(6) 2025 AT 9:57:51 AND (B)(6) 2025 AT 10:07:14. IT WAS ALSO NOTED THAT THERE WERE LOW POWER HAZARD AND MULTIPLE OTHER ASSOCIATED ALARM TYPES ON (B)(6)2025 BETWEEN 16:57:40 AND 16:58:24 AND AGAIN ON (B)(6) 2025 BETWEEN 21:35:46 AND 21:36:30. THIS WAS CAUSED BY POWER TO THE MOBILE POWER UNIT (MPU) BEING BRIEFLY INTERRUPTED. THERE WAS NO INTERRUPTION IN PUMP SUPPORT DURING THESE EVENTS. THE LOG ENDED WITH POWER BEING REMOVED AND THE DRIVELINE DISCONNECTED. THE DEVICE WAS OPERATING AS INTENDED AS THERE WERE NO PUMP SPEED ISSUES IN THE LOG FILES. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION D6B - DATE OF EXPLANT: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THOUGH THIS EVALUATION. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025, PER THE TIMESTAMPS. INTERMITTENT LOW FLOW EVENTS WERE CAPTURED ON (B)(6) 2025, WITH A SEVERAL OF THESE EVENTS ASSOCIATED WITH TRANSIENT LOW FLOW HAZARD ALARMS. OF NOTE, PULSATILITY INDEX (PI) VALUES APPEARED TO BE ELEVATED DURING THE LOW FLOW EVENTS. TOWARD THE END OF THE FILE ON (B)(6) 2025, FLOW CONTINUED TO DECREASE RESULTING IN A SUSTAINED LOW FLOW ALARM BEGINNING AT 08:55:31, WITH FLOW REACHING TO 0.0 LITER PER MINUTE (LPM) BY 10:06:46. THE PUMP WAS DISCONNECTED FROM EXTERNAL POWER AT 10:07:08, FOLLOWED BY THE DISCONNECTION OF THE DRIVELINE AT 10:07:14 RESULTING IN A PUMP STOP. THESE EVENTS APPEAR CONSISTENT WITH REPORTED EVENT OF THE PATIENT'S EXPIRATION. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED AT THE STORED PATIENT SPEED WHILE THE DRIVELINE WAS CONNECTED. WHETHER THE PATIENT'S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. A LIMITED AUTOPSY WAS TO BE PERFORMED; HOWEVER, NO RESULTS WERE PROVIDED. THE PUMP WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. THIS SECTION ALSO NOTES THAT, IN GENERAL, THE MAGNITUDE OF THE PULSATILITY INDEX (PI) VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E., THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21452990&pc=NPT | 2025587-2025-01424 | 2025-02-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-24 05:00 | NPT | EVOLUT PRO PLUS VALVE | Difficult or Delayed Positioning; Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypoxia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 22 MM BALLOON. THE INITIAL VALVE WAS LOADED INTO THE DELIVERY CATHETER SYSTEM (DCS) AND AN INFOLD TO NODE 2 WAS IDENTIFIED. THE VALVE WAS ADVANCED OVER A MEDTRONIC GUIDEWIRE VIA THE FEMORAL ARTERY APPROACH. WHILE SLOWLY RELEASING THE VALVE AT A DEPTH OF 5 MM ON THE NON-CORONARY CUSP, THE GUIDE IN THE LEFT VENTRICLE WAS SLIGHTLY WITHDRAWN. IT WAS REPORTED THAT TWO RECAPTURES WERE MADE PRIOR TO IMPLANT DUE TO A DISLODGE ON THE FIRST ATTEMPT AND A DEEP IMPLANT DEPTH ON THE SECOND ATTEMPT. FOLLOWING THE VALVE IMPLANT AT A DEPTH OF 8 MM, THE VALVE DISLODGED. A SNARE WAS USED TO HOLD THE DISLODGED VALVE WHILE A SECOND VALVE AND DCS WERE ADVANCED AND THE SECOND VALVE WAS IMPLANTED DEEPER. HOWEVER, THE FIRST VALVE ALSO MOVED DEEPER CAUSING A LEFT CORONARY ARTERY OBSTRUCTION. SUBSEQUENTLY, THE PATIENT DIED DURING THE PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE IMPLANT OF THE FIRST VALVE, THE VALVE FRAME WAS UNDER-EXPANDED DUE TO SEVERE CALCIFICATION OF THE ANNULUS AND LEAFLETS. IMMEDIATELY FOLLOWING THE RELEASE OF THE SECOND VALVE, CARDIORESPIRATORY ARREST OCCURRED. ADVANCED RESUSCITATIVE MEASURES WERE PERFORMED FOR APPROXIMATELY 45 MINUTES; HOWEVER, WERE UNSUCCESSFUL AND THE PATIENT DIED. ADDITIONALLY, IT WAS NOTED THAT FOLLOWING THE INDUCTION OF ANESTHESIA AND THROUGHOUT THE PROCEDURE, THE PATIENT WAS HYPOXEMIC AND HYPOTENSIVE REQUIRING OXYGEN AND THE INFUSION OF INOTROPIC AND VASOCONSTRICTOR MEDICATIONS AT HIGH DOSES. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A2. A4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D-EVPROP23-29 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: EVPROPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025; PRODUCT ID: D-EVPROP23-29 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID: L-EVPROP23-29 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW. FLUOROSCOPIC VALVE INSPECTION OF THE FIRST VALVE WAS NOT CAPTURED THUS IT IS NOT POSSIBLE TO VALIDATE A GOOD LOAD. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANT. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND DEPTH ASSESSMENT WAS PERFORMED. VALVE DEPTH APPEARED TO BE APPROXIMATELY 5-6MM ON THE NON-CORONARY CUSP IN THE CUSP OVERLAP VIEW, RAO 23 CAU 5, AND APPEARED TO BE SLIGHTLY UNDER EXPANDED. DEPTH WAS APPROXIMATELY 4-5MM ON THE LEFT CORONARY CUSP IN THE LAO VIEW. IT WAS REPORTED THAT TWO RECAPTURES WERE MADE PRIOR TO THE IMPLANT BUT THE RECAPTURES WERE NOT SAVED. VALVE DEPTH PRIOR TO FINAL RELEASE WAS APPROXIMATELY 3MM ON THE NON-CORONARY CUSP IN RAO 5 CAU 9 VIEW AND THE VALVE APPEARED TO BE MORE EXPANDED. THE IMPLANTING VIEW WAS DIFFERENT AND DEPTH ASSESSMENT IN THE LAO VIEW PRIOR TO FINAL RELEASE WAS NOT PERFORMED THUS DEPTH ON THE LEFT CORONARY CUSP IS UNKNOWN. DEPTH ASSESSMENT AT THE NCC IS PERFORMED IN A CUSP OVERLAP VIEW, AND IF ACCEPTABLE, NEXT STEP IS TO MOVE TO THE 3-CUSP COPLANAR VIEW AND REMOVE PARALLAX FROM THE INFLOW PORTION OF THE VALVE (ROLL LAO NO GREATER THAN 25°) TO VERIFY DEPTH AT THE LCC. THE VALVE MAY BE RELEASED ONCE THESE STEPS ARE COMPLETED. SOMETIME DURING FINAL RELEASE AND REMOVAL OF THE DELIVERY CATHETER SYSTEM THE VALVE DISLODGED TO THE LEVEL OF THE SINOTUBULAR JUNCTION. BASED ON FLUOROSCOPIC EVIDENCE, IT IS LIKELY THAT THE NOSE CONE INTERACTED WITH THE VALVE FRAME CONTRIBUTING TO THE DISLODGEMENT AS THE NOSE CONE APPEARS TO BE IN VERY CLOSE PROXIMITY/IN CONTACT WITH ONE OF THE PADDLES. EVEN THOUGH THE DISLODGEMENT EVENT WAS NOT CAPTURED ON FLUOROSCOPY, NOSE CONE INTERACTION APPEARS TO BE THE MOST LIKELY CAUSE. OF NOTE, CAUTION IS NEEDED WHILE WITHDRAWING THE DELIVERY CATHETER SYSTEM TO MINIMIZE INTERACTION WITH THE EVOLUT FRAME. A SECOND VALVE WAS LOADED AND CONFIRMED A GOOD LOAD. A SNARE WAS USED TO SECURE ONE OF THE PADDLES, AND DISLODGED VALVE POSITION REMAINED IN THE SINOTUBULAR JUNCTION. DURING THE FIRST DEPLOYMENT ATTEMPT WITH THE SECOND VALVE, AN INFOLD OCCURRED. THE VALVE WAS RECAPTURED AND REPOSITIONED AND THE INFOLD PERSISTED. THE INFOLD IS CHARACTERIZED BY AN INWARD FOLD OR CREASE IN THE VALVE, EXTENDING FROM THE INFLOW. INFOLDING COULD OCCUR DUE A MISLOADED VALVE, ANATOMICAL CHARACTERISTICS SUCH AS CALCIUM, AND RECAPTURES. IF AN INFOLD IS IDENTIFIED, THE VALVE SHOULD BE RECAPTURED AND REMOVED FROM THE BODY. NEW STERILE COMPONENTS MUST BE USED. HOWEVER, THE INFOLDED VALVE WAS RELEASED. THERE IS NO EVIDENCE THAT A POST IMPLANT DILATATION WAS PERFORMED SO THE INFOLD REMAINED UNRESOLVED. FINAL ANGIOGRAM REVEALED ABSENCE OF CORONARY FLOW. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED, AND EVIDENCE SUGGESTS CORONARY OCCLUSION WAS THE MOST PROBABLE CAUSE. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453019&pc=NPT | 2025587-2025-01425 | 2025-02-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-24 05:00 | NPT | EVOLUT PRO PLUS DCS | Difficult or Delayed Positioning; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypoxia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 22 MM BALLOON. THE INITIAL VALVE WAS LOADED INTO THE DELIVERY CATHETER SYSTEM (DCS) AND AN INFOLD TO NODE 2 WAS IDENTIFIED. THE VALVE WAS ADVANCED OVER A MEDTRONIC GUIDEWIRE VIA THE FEMORAL ARTERY APPROACH. WHILE SLOWLY RELEASING THE VALVE AT A DEPTH OF 5 MM ON THE NON-CORONARY CUSP, THE GUIDE IN THE LEFT VENTRICLE WAS SLIGHTLY WITHDRAWN. IT WAS REPORTED THAT TWO RECAPTURES WERE MADE PRIOR TO IMPLANT DUE TO A DISLODGE ON THE FIRST ATTEMPT AND A DEEP IMPLANT DEPTH ON THE SECOND ATTEMPT. FOLLOWING THE VALVE IMPLANT AT A DEPTH OF 8 MM, THE VALVE DISLODGED. A SNARE WAS USED TO HOLD THE DISLODGED VALVE WHILE A SECOND VALVE AND DCS WERE ADVANCED AND THE SECOND VALVE WAS IMPLANTED DEEPER. HOWEVER, THE FIRST VALVE ALSO MOVED DEEPER CAUSING A LEFT CORONARY ARTERY OBSTRUCTION. SUBSEQUENTLY, THE PATIENT DIED DURING THE PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE IMPLANT OF THE FIRST VALVE, THE VALVE FRAME WAS UNDER-EXPANDED DUE TO SEVERE CALCIFICATION OF THE ANNULUS AND LEAFLETS. IMMEDIATELY FOLLOWING THE RELEASE OF THE SECOND VALVE, CARDIORESPIRATORY ARREST OCCURRED. ADVANCED RESUSCITATIVE MEASURES WERE PERFORMED FOR APPROXIMATELY 45 MINUTES; HOWEVER, WERE UNSUCCESSFUL AND THE PATIENT DIED. ADDITIONALLY, IT WAS NOTED THAT FOLLOWING THE INDUCTION OF ANESTHESIA AND THROUGHOUT THE PROCEDURE, THE PATIENT WAS HYPOXEMIC AND HYPOTENSIVE REQUIRING OXYGEN AND THE INFUSION OF INOTROPIC AND VASOCONSTRICTOR MEDICATIONS AT HIGH DOSES. Manufacturer Narrative: UPDATED DATA: B5. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: EVPROPLUS-29, SERIAL/LOT #: (B)(6), UBD: 11-JUL-2026, UDI#: (B)(4) OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID EVPROPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; PRODUCT ID D-EVPROP23-29 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID L-EVPROP23-29 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW. FLUOROSCOPIC VALVE INSPECTION OF THE FIRST VALVE WAS NOT CAPTURED THUS IT IS NOT POSSIBLE TO VALIDATE A GOOD LOAD. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANT. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND DEPTH ASSESSMENT WAS PERFORMED. VALVE DEPTH APPEARED TO BE APPROXIMATELY 5-6MM ON THE NON-CORONARY CUSP IN THE CUSP OVERLAP VIEW, RAO 23 CAU 5, AND APPEARED TO BE SLIGHTLY UNDER EXPANDED. DEPTH WAS APPROXIMATELY 4-5MM ON THE LEFT CORONARY CUSP IN THE LAO VIEW. IT WAS REPORTED THAT TWO RECAPTURES WERE MADE PRIOR TO THE IMPLANT BUT THE RECAPTURES WERE NOT SAVED. VALVE DEPTH PRIOR TO FINAL RELEASE WAS APPROXIMATELY 3MM ON THE NON-CORONARY CUSP IN RAO 5 CAU 9 VIEW AND THE VALVE APPEARED TO BE MORE EXPANDED. THE IMPLANTING VIEW WAS DIFFERENT AND DEPTH ASSESSMENT IN THE LAO VIEW PRIOR TO FINAL RELEASE WAS NOT PERFORMED THUS DEPTH ON THE LEFT CORONARY CUSP IS UNKNOWN. DEPTH ASSESSMENT AT THE NCC IS PERFORMED IN A CUSP OVERLAP VIEW, AND IF ACCEPTABLE, NEXT STEP IS TO MOVE TO THE 3-CUSP COPLANAR VIEW AND REMOVE PARALLAX FROM THE INFLOW PORTION OF THE VALVE (ROLL LAO NO GREATER THAN 25°) TO VERIFY DEPTH AT THE LCC. THE VALVE MAY BE RELEASED ONCE THESE STEPS ARE COMPLETED. SOMETIME DURING FINAL RELEASE AND REMOVAL OF THE DELIVERY CATHETER SYSTEM THE VALVE DISLODGED TO THE LEVEL OF THE SINOTUBULAR JUNCTION. BASED ON FLUOROSCOPIC EVIDENCE, IT IS LIKELY THAT THE NOSE CONE INTERACTED WITH THE VALVE FRAME CONTRIBUTING TO THE DISLODGEMENT AS THE NOSE CONE APPEARS TO BE IN VERY CLOSE PROXIMITY/IN CONTACT WITH ONE OF THE PADDLES. EVEN THOUGH THE DISLODGEMENT EVENT WAS NOT CAPTURED ON FLUOROSCOPY, NOSE CONE INTERACTION APPEARS TO BE THE MOST LIKELY CAUSE. OF NOTE, CAUTION IS NEEDED WHILE WITHDRAWING THE DELIVERY CATHETER SYSTEM TO MINIMIZE INTERACTION WITH THE EVOLUT FRAME. A SECOND VALVE WAS LOADED AND CONFIRMED A GOOD LOAD. A SNARE WAS USED TO SECURE ONE OF THE PADDLES, AND DISLODGED VALVE POSITION REMAINED IN THE SINOTUBULAR JUNCTION. DURING THE FIRST DEPLOYMENT ATTEMPT WITH THE SECOND VALVE, AN INFOLD OCCURRED. THE VALVE WAS RECAPTURED AND REPOSITIONED AND THE INFOLD PERSISTED. THE INFOLD IS CHARACTERIZED BY AN INWARD FOLD OR CREASE IN THE VALVE, EXTENDING FROM THE INFLOW. INFOLDING COULD OCCUR DUE A MISLOADED VALVE, ANATOMICAL CHARACTERISTICS SUCH AS CALCIUM, AND RECAPTURES. IF AN INFOLD IS IDENTIFIED, THE VALVE SHOULD BE RECAPTURED AND REMOVED FROM THE BODY. NEW STERILE COMPONENTS MUST BE USED. HOWEVER, THE INFOLDED VALVE WAS RELEASED. THERE IS NO EVIDENCE THAT A POST IMPLANT DILATATION WAS PERFORMED SO THE INFOLD REMAINED UNRESOLVED. FINAL ANGIOGRAM REVEALED ABSENCE OF CORONARY FLOW. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED, AND EVIDENCE SUGGESTS CORONARY OCCLUSION WAS THE MOST PROBABLE CAUSE. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A2. A4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453416&pc=NVN | 2017865-2025-14963 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14961, 2017865-2025-14962. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM. Manufacturer Narrative: CORRECTION TO MISSING EXPLANT DATE. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453424&pc=NPT | 2025587-2025-01427 | 2024-07-04 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-24 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Stroke/CVA; Low Blood Pressure/ Hypotension; Renal Failure; Ventricular Fibrillation; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, THE PATIENT EXPERIENCED LOSS OF CONSCIOUSNESS AND LOW BLOOD PRESSURE PROMPTING AN EMERGENCY TAVR. DURING THE PROCEDURE, EXPANSION WAS PERFORMED DURING INSERTION OF THE DELIVERY CATHETER SYSTEM (DCS). SUBSEQUENTLY, DAMAGE TO THE DCS WAS OBSERVED. HEMOSTASIS BECAME DIFFICULT SO A REQUEST WAS MADE FOR THE SURGERY DEPARTMENT TO PERFORM A LEFT FEMORAL ARTERIOPLASTY. 1 DAY FOLLOWING THE TAVR PROCEDURE, AN EMERGENCY HEAD MAGNETIC RESONANCE IMAGING (MRI) WAS PERFORMED, REVEALING MULTIPLE CEREBRAL INFARCTIONS. DURING A CATHETER-BASED EXAMINATION, THE PATIENT¿S HEART RATE BEGAN TO DECREASE, AND CARDIAC ARREST OCCURRED. RESUSCITATION WAS PERFORMED, AND THE PATIENT WAS INTUBATED. THE EXAMINATION WAS DISCONTINUED. 3 DAYS FOLLOWING THE TAVR PROCEDURE, HEMODYNAMIC IMPROVEMENT WAS MINIMAL, AND THE PATIENT BEGAN DIALYSIS. VENTRICULAR TACHYCARDIA OCCURRED, AND DEFIBRILLATION WAS PERFORMED. 7 DAYS FOLLOWING THE TAVR PROCEDURE, THE PATIENT SUDDENLY WENT INTO VENTRICULAR FIBRILLATION AND DIED. THE CAUSE OF DEATH WAS REPORTED AS CONGESTIVE HEART FAILURE (CHF) AND END-STAGE RENAL FAILURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT 0012063625); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A UPDATED: D6B, H6, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453457&pc=NVY | 2017865-2025-14962 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-24 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14961, 2017865-2025-14963. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453681&pc=NVZ | 3008973940-2025-01397 | 2024-11-08 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-02-24 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED PASSED AWAY POST OPERATIVELY AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453706&pc=DTB | 2649622-2025-05305 | 2024-11-08 05:00 | Death | MPRI | 2025-02-24 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED PASSED AWAY POST OPERATIVELY AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453735&pc=NVN | 2649622-2025-05306 | 2024-11-08 05:00 | Death | MPRI | 2025-02-24 05:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED PASSED AWAY POST OPERATIVELY AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453839&pc=NVZ | 3008973940-2025-01398 | 2021-08-16 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-02-24 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Cardiac Tamponade | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY, THE PATIENT EXPERIENCED CARDIAC TAMPONADE FROM THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM IMPLANT PROCEDURE. EVENTUALLY THE PATIENT PASSED AWAY FROM THE COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453846&pc=DTB | 2649622-2025-05307 | 2021-08-16 04:00 | Death | MPRI | 2025-02-24 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Cardiac Tamponade | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY, THE PATIENT EXPERIENCED CARDIAC TAMPONADE FROM THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM IMPLANT PROCEDURE. EVENTUALLY THE PATIENT PASSED AWAY FROM THE COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21453860&pc=DTB | 2649622-2025-05309 | 2021-08-16 04:00 | Death | MPRI | 2025-02-24 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Cardiac Tamponade | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY, THE PATIENT EXPERIENCED CARDIAC TAMPONADE FROM THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM IMPLANT PROCEDURE. EVENTUALLY THE PATIENT PASSED AWAY FROM THE COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454278&pc=DSQ | 2916596-2025-00966 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT THE HOME. Manufacturer Narrative: SECTION B: DATE OF DEATH CORRECTED. SECTION D4: UDI NUMBER CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PHYSICIAN TO GET INSTRUCTION ON HOW TO TURN OFF THE SYSTEM CONTROLLER AFTER PATIENT PASSED AWAY. INSTRUCTION WAS GIVEN TO PUT THE CONTROLLER TO SLEEP BY HOLDING THE BATTERY SYMBOL ON THE CONTROLLER FOR 5 SECONDS. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. A1- A4: PATIENT INFORMATION NOT PROVIDED. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454361&pc=DSQ | 2916596-2025-00998 | 2018-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding; Unspecified Infection; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "LONG-TERM QUALITY OF LIFE, PSYCHOLOGICAL DISTRESS, AND CAREGIVER BURDEN IN OCTOGENARIANS WITH LVAD: A SINGLE-CENTER EXPERIENCE" THAT HEARTMATE 3 MAY BE RELATED TO PSYCHOLOGICAL DISTRESS. THIS STUDY ASSESSED THE LONG-TERM QUALITY OF LIFE, AS DEFINED BY BOTH THE 36-ITEM SHORT-FORM HEALTH (SF-36) SURVEY AND THE EUROQOL 5 DIMENSIONS, 5-LEVEL QUESTIONNAIRE (EQ-5D-5L)¿INCLUDING THE VISUAL ANALOG SCALE¿IN OCTOGENARIAN LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS. ADDITIONALLY, THE PSYCHOLOGICAL HEALTH OF OCTOGENARIAN LVAD PATIENTS USING THE PSYCHOLOGICAL GENERAL WELL-BEING INDEX (PGWBI) THROUGH THE 22-ITEM VERSION OF THE ZARIT BURDEN INTERVIEW (ZBI) WAS EVALUATED. ONE OF THE STUDY PATIENTS, MALE, 76 YEARS OF AGE AT TIME OF IMPLANT, EXPERIENCED 8 HOSPITALIZATIONS WHILE IMPLANTED WITH THE HEARTMATE 3 FOR INFECTION, ARRHYTHMIA, HEART FAILURE AND RECURRENT BLEEDING. THE PATIENT PASSED AWAY FROM MULTI-SYSTEM ORGAN FAILURE. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. AUTHOR INFORMATION: JACOPO D¿ANDRIA URSOLEO, MARINA PIERI, FRANCESCO CALVO, SAVINO ALTIZIO, MARIO GRAMEGNA, DOMENICO PONTILLO, SILVIA AJELLO AND ANNA MARA SCANDROGLIO DEPARTMENT OF ANESTHESIA AND INTENSIVE CARE, IRCCS SAN RAFFAELE SCIENTIFIC INSTITUTE, MILAN, LOMBARDIA, ITALY NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: E1; REPORTER EMAIL WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), CARDIAC ARRHYTHMIA, BLEEDING, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION ALSO CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INTERNATIONAL NORMALIZED RATIO (INR) RANGE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454500&pc=MFE | 2029214-2025-00489 | 2023-07-04 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-24 05:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke |  |  | 1 | Event Description: MEDTRONIC LITERATURE REVIEW FOUND A REPORT OF DEATH, NONFATAL HEMORRHAGIC STROKE, SEIZURE, WORSENED MRS SCORE, EDEMA, CYST FORMATION, VESSEL PERFORATION, ACUTE STROKE, AND NEUROLOGICAL DECLINE IN ASSOCIATION WITH ONYX EMBOLIZATION. THE TIME FRAME OF THIS STUDY WAS FROM AUGUST 2011 TO AUGUST 2021. THE PURPOSE OF THIS ARTICLE WAS TO COMPARE THE LONG-TERM OUTCOMES OF STEREOTACTIC RADIOSURGERY (SRS) WITH OR WITHOUT PRIOR EMBOLIZATION IN BRAIN ARTERIOVENOUS MALFORMATIONS (AVMS) (VOLUME =10ML) FOR WHICH SRS IS INDICATED. THE AUTHORS REVIEWED 270 CASES OF PATIENTS TREATED FOR BRAIN ARTERIOVENOUS MALFORMATIONS USING ONYX EMBOLIZATION AND STEREOTACTIC RADIOSURGERY. THERE WERE A TOTAL OF 486 PATIENTS SPLIT INTO TWO GROUPS: 243 PATIENTS RECEIVED E + SRS AND 243 PATIENTS SRS ALONE. OF THE 243 PATIENTS WHO RECEIVED E+SRS, THE AVERAGE AGE WAS 26 YEARS, 127 WERE FEMALE AND 116 WERE MALE. OVERALL, COMBINED EMBOLIZATION + STEREOTACTIC RADIOSURGERY (E+SRS) AND STEREOTACTIC RADIO SURGERY (SRS) ALONE WERE SIMILAR IN PREVENTING LONG-TERM NON-FATAL HEMORRHAGIC STROKE AND DEATH, AS WELL AS IN FACILITATING AVM OBLITERATION. HOWEVER, THE E+SRS STRATEGY WAS SIGNIFICANTLY INFERIOR TO THE SRS ALONE STRATEGY IN TERMS OF NEUROLOGICAL DETERIORATION. THE ARTICLE DOES NOT STATE ANY TECHNICAL ISSUES DURING USE OF THE ONYX. THE FOLLOWING INTRA- OR POST-PROCEDURAL OUTCOMES WERE NOTED IN THE E+SRS GROUP: 2 PATIENTS HAD HEMORRHAGIC STROKES LEADING TO DEATH 8 PATIENTS HAD NONFATAL HEMORRHAGIC STROKES 18 PATIENTS HAD A SEIZURE 39 PATIENTS HAD WORSENED MRS SCORES 17 PATIENTS HAD EDEMA 1 PATIENT HAD A CYST FORMATION 2 PATIENTS HAD VESSEL PERFORATIONS 19 PATIENTS HAD ACUTE STROKES 13 PATIENTS HAD A NEUROLOGICAL DECLINE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-ONYX (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: HENGWEI JIN, ZHIPENG LI, DEZHI GAO, YU CHEN, HEZE HAN, LI MA, DEBIN YAN, RUINAN LI, ANQI LI, HAIBIN ZHANG, KEXIN YUAN, YUKUN ZHANG, YANG ZHAO, XIANGYU MENG, YOUXIANG LI, XIAOLIN CHEN, HAO WANG, SHIBIN. ASSOCIATION OF THE COMBINED STEREOTACTIC RADIOSURGERY AND EMBOLIZATION STRATEGY AND LONG-TERM OUTCOMES IN BRAIN ARTERIOVENOUS MALFORMATIONS WITH A VOLUME =10 ML: A NATIONWIDE MULTICENTER OBSERVATIONAL. JOURNAL OF NEUROINTERVENTIONAL SURGERY 16:548¿554 2023. DOI:10.1136/JNIS-2023-020289 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454507&pc=JDI | 1038671-2025-01177 | 2023-11-09 05:00 | Death | EXACTECH, INC. | 2025-02-24 05:00 | JDI | UNKNOWN HIP COMPONENTS | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information |  |  | 1 | Event Description: AS REPORTED, THE PARTY REPORTED THAT THEIR FAMILY MEMBER UNDERWENT AN IMPLANT PROCEDURE FOR THEIR LEFT HIP, FOLLOWED BY REVISION SURGERY THE SAME YEAR. THEY NOTED THAT THE PATIENT EXPERIENCED COMPLICATIONS WITH THEIR LEFT HIP, LEADING TO ANOTHER SURGERY IN THE SAME YEAR. THE PARTY FURTHER NOTED THAT THEIR FAMILY MEMBER PASSED AWAY AT THE END OF THE YEAR, SUGGESTING THAT THEIR COMPLICATIONS MAY HAVE CONTRIBUTED TO THEIR DEATH. NO FURTHER INFORMATION AVAILABLE.Event Description: REPORT 2 OF 2: IT WAS REPORTED THAT IN EARLY 2023, A PATIENT UNDERWENT AN IMPLANT PROCEDURE FOR THE LEFT HIP, FOLLOWED BY REVISION SURGERY (REPORT 1 OF 2). IT WAS NOTED THAT THE PATIENT EXPERIENCED COMPLICATIONS WITH THE LEFT HIP, LEADING TO ANOTHER SURGERY ON LATE 2023. IT WAS FURTHER NOTED THAT THE PATIENT PASSED AWAY IN LATE 2023, SUGGESTING THAT COMPLICATIONS MAY HAVE CONTRIBUTED TO THE PATIENT'S DEATH. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE CORRECTED: H6 MEDICAL DEVICE PROBLEM CODE, H6 INVESTIGATION CODES-ALL. THE FOLLOWING SECTIONS WERE UPDATED: E2, E3. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE CORRECTED: B5, D6A: IMPLANTED (B)(6) 2023; SPECIFIC DATE UNKNOWN, D6B. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. Manufacturer Narrative: THE REPORTED EVENT WAS UNABLE TO BE CONFIRMED DUE TO LIMITED INFORMATION RECEIVED FROM THE CUSTOMER. NO DEVICE WAS RETURNED FOR EVALUATION; FURTHER, PHOTOGRAPHS AND/OR RADIOGRAPH IMAGES WERE NOT PROVIDED FOR REVIEW. OPERATIVE NOTES AND/OR MEDICAL RECORDS WERE NOT PROVIDED FOR REVIEW OF USAGE/ TECHNIQUE. A REVIEW OF MANUFACTURING DATA WAS UNABLE TO BE PERFORMED AS THE LOT INFORMATION OF THE PRODUCT INVOLVED IN THE EVENT WAS NOT AVAILABLE. A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AS THE NECESSARY INFORMATION TO ADEQUATELY INVESTIGATE THE REPORTED EVENT WAS NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454625&pc=DSQ | 2916596-2025-01175 | 2025-02-04 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE FOLLOWED BY MULTI-ORGAN FAILURE (MOF). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT SUFFERED FROM A PROGRESSIVE RIGHT HEART FAILURE (RHF) ALTHOUGH ALL THERAPEUTIC OPTIONS AROUND WERE OPTIMIZED. THE RHF WAS NOT RELATED TO THE LEFT VENTRICULAR ASSIST DEVICE THERAPY. NO DEVICE ISSUES WERE MENTIONED. MOF IN CONSEQUENCE WAS RELATED TO THE OUTCOME AT LEAST. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS NOTED THAT WHILE THE PATIENT HAD A HISTORY OF RV FAILURE PRIOR TO LVAD, IT WAS GOOD TO HANDLE BEFORE. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454653&pc=DSQ | 2916596-2025-01113 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-02-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Tachycardia; Ischemia Stroke; Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation; Mitral Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR ASSIST DEVICE (RVAD) SUPPORT WAS INITIATED POST LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION.Event Description: THE PATIENT'S COURSE WAS COMPLICATED BY A VENTRICULAR TACHYCARDIA (VT) STORM AND A LARGE PERICARDIAL EFFUSION WITHOUT TAMPONADE. THE PATIENT DEVELOPED RECURRENT VT AS WELL AS WORSENING RIGHT VENTRICULAR (RV) DYSFUNCTION. ON (B)(6) 2025, THE PATIENT HAD WORSENING RENAL FUNCTION DESPITE ADEQUATE DIURESIS AND WAS TAKEN FOR A RIGHT HEART CATHETERIZATION (RHC). THE RHC SHOWED ELEVATED FILLING PRESSURES WITH NORMAL CARDIAC INDEX WHICH WAS INDICATIVE OF RV DYSFUNCTION. THE PLAN WAS TO PERFORM ROBUST DIURESIS WITH RV SUPPORT. THE PATIENT WAS TO START EPINEPHRINE (EPI) 2, RESUME BUMEX (BUMETANIDE) 2 MG/HR DRIP, AND TO TAKE DIURIL (CHLOROTHIAZIDE) 500MG ONCE. A CONSULTATION WAS TO BE DONE WITH ELECTROPHYSIOLOGY (EP) FOR VT MANAGEMENT AND THE PATIENT WAS TO BE MONITORED FOR COMPLICATIONS. AN ECHOCARDIOGRAM WAS DONE SHOWING AN ENLARGED LEFT ATRIUM (LA), AND A SEVERELY ENLARGED RIGHT VENTRICULAR CHAMBER SIZE. THE GLOBAL RV SYSTOLIC FUNCTION WAS SEVERELY REDUCED. THE RIGHT ATRIUM WAS DILATED. THE RIGHT ATRIUM PACING WIRE WAS VISUALIZED. THE AORTIC VALVE SHOWED MILD AOV CUSP THICKENING WITH TRACE AORTIC REGURGITATION. THE MITRAL VALVE SHOWED MILD LEAFLET THICKENING WITH MILD MITRAL REGURGITATION. A SMALL ANTERIOR PERICARDIAL EFFUSION WAS PRESENT IN THE PERICARDIUM. THE PATIENT WAS INTUBATED FOR INCESSANT VT ON (B)(6) 2025 AND PROTEK DUO WAS PLACED WITH PERCUTANEOUS CORONARY INTERVENTION (PCI) TO LEFT MAIN CORONARY ARTERY (LMCA) AND LEFT CIRCUMFLEX ARTERY (LCX). THEY UNDERWENT PROTEK DUO REMOVAL AND 29F RE-INSERTION (B)(6) 2025 WITH ANGIOGRAPHY SHOWING PATENT STENTS AND NO HM3 OUTFLOW THROMBUS. THE PATIENT'S HOSPITAL COURSE WAS FURTHER COMPLICATED BY LEFT MAIN CORONARY ARTERY (MCA) STROKE NOTED ON COMPUTED TOMOGRAPHY (CT) SCAN DONE ON (B)(6) 2025. THE PATIENT ALSO UNDERWENT GANGLION BLOCK BILATERALLY ON (B)(6) 2025. PROTEK DUO WAS CAPPED ON (B)(6) 2025. A LEVITRONIX RVAD WAS IMPLANTED ON (B)(6) 2025. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: DUE TO SYSTEM LIMITATIONS, H6: ADDITIONAL HEALTH EFFECT-CLINICAL CODE: E0619-PERICARDIAL EFFUSION. B2 - OUTCOMES ATTRIBUTED TO ADVERSE UPDATED. B5 NARRATIVE INFORMATION. H1 - TYPE OF REPORTABLE EVENT UPDATED. H6 - ADVERSE EVENT PROBLEM CODES UPDATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: B2: CORRECTION, DEATH DATE PROVIDED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS CANNOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, CARDIAC ARRHYTHMIA, STROKE, RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS DOCUMENT ALSO WARNS THE USER THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21454975&pc=NPT | 2025587-2025-01434 | 2025-01-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-24 05:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem | Cardiac Arrest; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVFXPLUS-29 IN A PATIENT WITH TORTUOUS AND CALCIFIED ANATOMY, A PRE-IMPLANT BALLOON VALVULOPLASTY WAS PERFORMED DUE TO CALCIFICATION. THE VALVE WAS INITIALLY DEPLOYED TO 80% BUT EXHIBITED INFOLDING ALONG ITS ENTIRE LENGTH. IT WAS ALSO NOTED THAT THE IMPLANT DEPTH WAS TOO SHALLOW. SUBSEQUENTLY, THE VALVE WAS RECAPTURED INTO THE DELIVERY CATHETER SYSTEM (DCS) AND WITHDRAWN FROM THE PATIENT. A SECOND VALVE WAS THEN LOADED INTO A NEW DCS BUT WAS UNABLE TO CROSS THE NATIVE VALVE DESPITE EMPLOYING VARIOUS TECHNIQUES. THE PATIENT COULD NOT TOLERATE THE AORTIC INSUFFICIENCY, BECAME ISCHEMIC, AND EXPERIENCED CARDIAC ARREST, ULTIMATELY RESULTING IN DEATH. Manufacturer Narrative: CONTINUATION OF D10. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-EVOLUTFX-2329 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE PRODUCT ID EVFXPLUS-29 (K002740); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21455263&pc=NAY | 2955842-2025-02025 | 2025-01-21 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-02-24 05:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT INJURY OCCURRED WHILE PORTS WERE BEING PLACED FOR A DA VINCI-ASSISTED SACROCOLPOPEXY PROCEDURE. ALL PORTS WERE PLACED AND VISUALIZATION WAS CONFIRMED. BEFORE THE DA VINCI SYSTEM WAS DOCKED TO THE PATIENT, THE ANESTHESIOLOGIST OBSERVED THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING. THE SURGEON DID NOT VISUALIZE ANY BLEEDING AT THE TIME, BUT CONVERTED THE PROCEDURE TO OPEN AND REQUESTED MULTIPLE GENERAL SURGEONS TO ASSIST. THE ROBOTICS COORDINATOR REPORTED THAT THE LOCATION OF THE BLEEDING WAS NEVER FOUND. NO FURTHER INFORMATION WAS PROVIDED REGARDING WHAT OCCURRED AFTER CONVERTING TO OPEN. THE ROBOTICS COORDINATOR STATED THE THEORY WAS THAT THE VERESS NEEDLE USED DURING TROCAR PLACEMENT DAMAGED RETROPERITONEAL ANATOMY. ATTEMPTS WERE MADE TO SPEAK WITH THE SURGEON, HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: NO SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE AS THE DA VINCI SYSTEM WAS NEVER DOCKED / USED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT ACCORDING TO THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, THE PATIENT IN THIS REPORT IS SPECULATED TO HAVE DIED POSSIBLY AS A RESULT OF A VERESS NEEDLE INJURY WHICH IS A THIRD PARTY DEVICE. THE INTUITIVE SURGICAL DA VINCI ROBOT WAS NEVER DOCKED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED AS SAME DAY OF SURGICAL PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21442469&pc=BZD | 2518422-2025-020552 | 2024-06-10 04:00 | Death | RESPIRONICS, INC. | 2025-02-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION DUE TO DEATH HOWEVER NO OTHER MEDICAL OR CLINICAL INFORMATION WERE REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21440489&pc=QLO | 3014015528-2025-00015 | 2024-03-13 04:00 | Death | EXTHERA MEDICAL CORPORATION | 2025-02-22 05:00 | QLO | SERAPH 100 | Lack of Effect | Chest Pain; Cancer |  |  | 1 | Event Description: THIS INFORMATION WAS REPORTED PUBLICLY BY THE NEW YORK TIMES. SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. PATIENT P5 (FEMALE, 42, DIAGNOSED WITH METASTATIC BREAST CANCER PRIOR TO BEING TREATED WITH SERAPH 100) FELT CHEST PAIN AND LEARNED OF A LARGE NEW TUMOR BETWEEN HER RIBS AND LUNGS AFTER SECOND TREATMENT. SHE DIED THREE MONTHS AFTER TREATMENT, APPROXIMATELY IN JUNE 2024. THE ARTICLE STATES THAT THESE TREATMENTS TOOK PLACE OUTSIDE THE UNITED STATES, IN ANTIGUA. Manufacturer Narrative: THIS WAS STATED IN THE NYT, SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN ANTIGUA WHERE THE CLAIMED ADVERSE EVENTS TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21440668&pc=QLO | 3014015528-2025-00014 | 2024-04-13 04:00 | Death | EXTHERA MEDICAL CORPORATION | 2025-02-22 05:00 | QLO | SERAPH 100 | Patient Device Interaction Problem; Lack of Effect | Low Blood Pressure/ Hypotension; Pain; Cancer; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THIS INFORMATION WAS REPORTED PUBLICLY BY THE NEW YORK TIMES. SEE (B)(6). "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. PATIENT P4 (MALE, DIAGNOSED WITH METASTATIC ABDOMINAL CANCER PRIOR TO TREATMENT WITH SERAPH 100) "HEALTH DECLINED PRECIPITOUSLY" AFTER RECEIVING FILTER TREATMENT IN (B)(6) 2024. EXACT DATE OF P4'S TREATMENT UNKNOWN BUT OCCURRED APPROXIMATELY TWO MONTHS BEFORE HIS DEATH ON (B)(6) 2024. THE ARTICLE STATES THAT THESE TREATMENTS TOOK PLACE OUTSIDE THE UNITED STATES, IN (B)(6). THE DEVICE WAS BEING USED OFF-LABEL FOR THE TREATMENT OF CANCER, OUTSIDE OF THE UNITED STATES IN (B)(6). FURTHER INVESTIGATION REVEALED THAT P4 RECEIVED 3 FILTERS IN (B)(6) AROUND LATE (B)(6) 2024 AS OFF-LABEL TREATMENT FOR A LIPOSARCOMA MASS. SEVERAL WEEKS OUT HIS CT SCAN SHOWED 2CM GROWTH OF PELVIC MASS. BETWEEN (B)(6) AND (B)(6), THE TUMOR HAD GROWN BY ~50%. AT THAT TIME P4 WAS DIAGNOSED WITH TUMOR LYSIS SYNDROME (URIC ACID LEVEL INCREASED AMONG OTHER THINGS. IMPLIED TREATMENT TO REDUCE LEVELS OF POTASSIUM AND CALCIUM AS WELL). PER P4'S FATHER ON (B)(6) 2024, "LAST NIGHT [P4]'S HEART WAS ALSO ACTING UP WITH REALLY HIGH BRATS TO COMPENSATE FOR THE REALLY LOW BLOOD PRESSURE. 60/42 AT [ONE] POINT... HE HAD TO GO ON OXYGEN AS WELL LAST NIGHT - AND THEN THAT CHANGED BLOOD PRESSURE AND HEART RATE." P4 ALSO "HAD PAIN IN HIS LEFT LEG". Manufacturer Narrative: THIS WAS STATED IN THE NYT, SEE (B)(6). "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN (B)(6) WHERE THE CLAIMED ADVERSE EVENTS TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21441869&pc=QLO | 3014015528-2025-00017 |  | Death | EXTHERA MEDICAL CORPORATION | 2025-02-22 05:00 | QLO | SERAPH 100 | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: THIS INFORMATION WAS REPORTED PUBLICLY BY THE NEW YORK TIMES. SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. PATIENT P7, NOT OTHERWISE IDENTIFIED, DIED FOLLOWING TREATMENT. OVERALL THE ARTICLE INDICATES THAT P6 WAS ONE OF SEVERAL PATIENTS SEEKING OFF-LABEL TREATMENT FOR CANCER. THE ARTICLE STATES THAT THESE TREATMENTS TOOK PLACE OUTSIDE THE UNITED STATES, IN ANTIGUA. FURTHER INVESTIGATION INDICATES THAT TREATMENT AND ANY SUBSEQUENT ADVERSE EVENT WOULD HAVE OCCURRED SOME TIME IN 2024, DUE TO THE TIMING OF THE ARTICLE AND THE AVAILABILITY OF TREATMENT IN ANTIGUA. Manufacturer Narrative: THIS WAS STATED IN THE NYT, SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN ANTIGUA WHERE THE CLAIMED ADVERSE EVENTS TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21430394&pc=EWL | 8032044-2025-00010 |  | Death | ATOS MEDICAL AB | 2025-02-21 05:00 | EWL | PROVOX2 | Inadequacy of Device Shape and/or Size | Abscess; Sepsis; Osteomyelitis | K971244 |  | 1 | Event Description: THIS IS A CASE REPORT DESCRIBED IN THE LITERATURE PUBLICATION: YOHAN SONG, KRISTA KIYOSAKI, EDWARD J. DAMROSE, EPIDURAL ABSCESS AND PARAPLEGIA: DELAYED SEQUELA OF TRACHEOESOPHAGEAL PUNCTURE, OTOLARYNGOLOGY CASE REPORTS, VOLUME 5, 2017, PAGES 21-23, ISSN 2468-5488, HTTPS://DOI.ORG/10.1016/J.XOCR.2017.10.002. (HTTPS://WWW.SCIENCEDIRECT.COM/SCIENCE/ARTICLE/PII/S2468548817301388). THIS CASE REPORT WAS IDENTIFIED DURING A LITERATURE SEARCH FOR VOICE PROSTHESES CONDUCTED IN ACCORDANCE WITH THE PERIODIC UPDATE OF THE CLINICAL EVALUATION REPORT. THE CASE DESCRIBES A 78-YEAR OLD FEMALE WHO HAD TOTAL LARYNGECTOMY AND PRIMARY PLACEMENT OF A 10 MM PROVOX2 IN OCTOBER 2016. FOUR MONTHS LATER SHE WAS BROUGHT TO THE EMERGENCY DEPARTMENT WITH FOR A TWO-DAY HISTORY OF PROGRESSIVE BILATERAL LOWER EXTREMITY WEAKNESS AND SUBSEQUENT FALL. MR IMAGING OF THE CERVICAL AND THORACIC SPINE SHOWED AN EPIDURAL ABSCESS. THE IMAGING WAS HIGHLY SUSPICIOUS FOR EROSION OF THE VOICE PROSTHESIS THROUGH THE POSTERIOR ESOPHAGUS AT THE LEVEL OF THE ABSCESS. SHE HAD AN OPERATION TO REMOVE THE ABSCESS. POST-OPERATIVELY, THERE WAS NO OBVIOUS PROJECTION OF THE DEVICE THROUGH THE POSTERIOR ESOPHAGUS, AS WAS EXPECTED BY THE IMAGING. THE VOICE PROSTHESIS WAS REMOVED AND EXCHANGED FOR A 4.5MM DEVICE. APPROXIMATELY SEVEN DAYS POSTOPERATIVELY, THE PATIENT DEVELOPED WORSENING DYSPHAGIA. MRI OF THE CERVICAL AND THORACIC SPINE AGAIN SHOWED AN ABSCESS AS WELL AS OSTEOMYELITIS OF THE VERTEBRAL BODIES. GIVEN THE PERSISTENT INFECTION, THE PROGRESSION TO OSTEOMYELITIS, HER IRREVERSIBLE NEUROLOGICAL DEFICITS, THE DIFFICULTY AND UNLIKELY SUCCESS OF FURTHER SURGICAL INTERVENTIONS, THE DECISION WAS MADE TO TRANSITION THE PATIENT TO HOSPICE. THE PATIENT PASSED AWAY FROM SEPSIS SEVEN WEEKS LATER. IN THE DISCUSSION OF THE ARTICLE, THE AUTHORS HYPOTHESIZE THAT HER 10 MM TEP PROSTHESIS MAY HAVE BEEN TOO LONG. FOR THIS REASON, IT WAS EXCHANGED TO A SHORTER 4.5 MM ONE DURING HER HOSPITALIZATION. THE PATIENT ALSO WORE A FENESTRATED LARYNGECTOMY TUBE WITH A HEAT AND MOISTURE EXCHANGER WITHIN THE STOMA (PROVOX LARYTUBE, ATOS MEDICAL, NEW BERLIN, WI, USA). DIGITAL PRESSURE TRANSMITTED FROM THE LARYNGECTOMY TUBE TO THE PROSTHESIS DURING PHONATION MAY HAVE HELPED TO DRIVE THE VOICE PROSTHESIS THROUGH THE POSTERIOR ESOPHAGEAL WALL. PLEASE NOTE THAT THIS IS JUST A SUMMARY OF THE CASE. FOR A FULL CASE DESCRIPTION, PLEASE SEE REFERENCED ARTICLE. Manufacturer Narrative: CONCLUSION PRODUCT DEFECT COULD NOT BE CONFIRMED. IT IS ASSESSED THAT THE COMPLICATIONS WERE DUE TO USAGE. THE PATIENT WAS UNDER SUPERVISION. INVESTIGATION THIS IS A THEORETICAL INVESTIGATION; SAMPLES WERE NOT AVAILABLE AND WAS DERIVED FROM A CASE REPORT IN LITERATURE STUDY. THE PATIENT USED A PROVOX2 10 MM VOICE PROSTHESIS FOR FOUR MONTHS. THE PATIENT WAS WEAKENED AND FELL, LEADING TO DISCOVERY OF ABSCESS IN SPINAL CORE. THE VOICE PROSTHESIS WAS REPLACED WITH A 4,5MM TO TRY TO OPTIMIZE FIT. THE PATIENT PASSED AWAY SOME WEEKS LATER, NOT BEING ABLE TO HEAL FROM INFECTION. THE PATIENT ALSO USED A LARYTUBE. THE CASE REPORT SUGGESTS A HYPOTHESIS THAT THE 10MM PROSTHESIS WAS TOO LONG, EXCESSIVE DIGITAL PRESSURE APPLIED TO THE HEAT-MOISTURE EXCHANGER (HME) AND LARYTUBE CAUSED THE PROSTHESIS TO MOVE WITHIN THE PUNCTURE (BACKWARDS AND FORWARDS) AND PERFORATE AND ERODE THE POSTERIOR ESOPHAGEAL WALL. THE DAMAGE TO THE ESOPHAGEAL WALL COULD DERIVE FROM TRACHEO-ESOPHAGEAL PROTHESIS INSERTION ACCORDING TO THE CASE REPORT. THERE WAS NO SIGN OF DAMAGE TO THE ESOPHAGEAL WALL AFTER FIRST SURGERY. FOLLOW UP QUESTIONS ARE SENT TO THE AUTHORS/HOSPITAL: - DOES THE HOSPITAL ROUTINELY REVIEW THE SIZE OF THE VOICE PROSTHESIS AT EACH REPLACEMENT OR REPLACE LIKE FOR LIKE ? ANSWER: YES, WE ROUTINELY REVIEW THE SIZE OF THE PROSTHESIS AT TIME OF REPLACEMENT. THE PHYSICIAN COULD SEE THE DEVICE IN THE TRACHEA AND IT APPEARED TO BE FLUSHED WITH THE TRACHEAL WALL. PHYSICIAN DID NOT PERCEIVE THAT THE DEVICE WAS BEING PUSHED THROUGH THE POSTERIOR ESOPHAGEAL WALL. - WAS THERE ANYONE FROM ATOS MEDICAL PRESENT WITH TROUBLE SHOOTING? EDUCATIONAL TEAM? ANSWER: NO ONE FROM ATOS WAS INVOLVED, WE DID NOT CONSULT. WE DO NOT ROUTINELY DO THIS. - WAS THE RESPONSIBLE CLINICIAN TRAINED WITH THE ATOS MEDICAL PRODUCTS? ANSWER: YES, THE RESPONSIBLE PHYSICIAN (ME) IS TRAINED WITH ATOS PRODUCTS. - WAS THERE FOLLOW UP MEETINGS WITH THE PATIENT AFTER FIRST SURGERY FOR TRAINING/ADVISING THE PATIENT? ANSWER: YES, THERE WERE MANY MEETINGS WITH THE PATIENT. ADDITIONAL FROM AUTHOR: ¿THE REASON WE THINK FOR THE COMPLICATION WAS THAT THE PATIENT WAS WEARING A FENESTRATED LARYNGECTOMY TUBE WHICH WAS APPLYING PRESSURE ONTO THE ANTERIOR SURFACE OF THE VOICE PROSTHESIS. AS SPEECH BECAME MORE DIFFICULT SHE WOULD PUSH HARDER, DRIVING THE DEVICE FURTHER BACK INTO THE ESOPHAGEAL WALL.¿ ROOT CAUSE USE ERROR, IN THE SENSE THAT IT WAS NOT THE PRODUCT BUT CIRCUMSTANCES WITH PUNCTURE AND USAGE WITH LARYTUBE THAT CAUSED THE INFECTION. DISCUSSION IN THE IFU 10877 THE FOLLOWING IS STATED: ¿DISLODGEMENT OR EXTRUSION OF THE PROVOX2 VOICE PROSTHESIS¿ AND SUBSEQUENT INGESTION, ASPIRATION OR TISSUE DAMAGE MAY OCCUR. TO REDUCE THE RISK¿ : - SELECT PROPER PROSTHESIS SIZE - IF LARYNGECTOMY TUBES¿ ARE USED, CHOOSE SIZE THAT DO NOT EXERT PRESSURE ON THE PROSTHESIS¿ - INSTRUCT THE PATIENT TO CONSULT A PHYSICIAN IMMEDIATELY IF THERE ARE ANY SIGNS OF TISSUE EDEMA AND/OR INFLAMMATION/INFECTION TRACHEAL AND/OR ESOPHAGEAL TISSUE DAMAGE MAY OCCUR IF THE PROSTHESIS IS TOO SHORT, TOO LONG, OR IS PUSHED FREQUENTLY AGAINST THE ESOPHAGEAL WALL BY A TRACHEAL CANNULA, STOMA BUTTON, OR THE PATIENT¿S FINGER. IN PATIENTS UNDERGOING RADIOTHERAPY THIS MAY HAPPEN MORE EASILY. SOLUTION: IF THE PATIENT EXPERIENCES SORENESS OR PAIN IN THE AREA AROUND THE VOICE PROSTHESIS, INSPECT THE TISSUE AROUND THE TE PUNCTURE BY ENDOSCOPY TO AVOID SEVERE DAMAGE.¿ RISK ASSESSMENT RH028-HAZ ED 04, H09: HARM TO PUNCTURE OR ESOPHAGUS OR USED BY PATIENT WITH BLEEDING DISORDER, LEADING TO BLEEDING/HEMORRHAGE (SEVERE BLEEDING 15-30% OF BLOOD VOLUME). CLINICAL SEVERITY = 6, SERIOUS. RH028-HAZ ED 04, H13: CONTAMINATION BY MICROORGANISM LEADING TO INFECTION. CLINICAL SEVERITY = 6. HOWEVER, IN PRODUCT RISK ASSESSMENT (UDFMECA) RH028 PROVOX2 ED08, ITEMS 9.15+9.20 ADDRESS USAGE WITH VP IN COMBINATION WITH LARYTUBE/BUTTON. POINTING TO H15: WRONG SIZE INSERTED (TOO SHORT) LEADING TO NECROSIS. CLINICAL SEVERITY = 6. THE RISK IS CONSIDERED TO BE COVERED IN THE PRODUCT RISK ASSESSMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21431167&pc=LWS | 2649622-2025-05057 | 2025-01-24 05:00 | Death | MPRI | 2025-02-21 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT A LEAD INTEGRITY ALERT (LIA) TRIGGERED ON THE RIGHT VENTRICULAR (RV) LEAD DURING A TREATED VENTRICULAR FIBRILLATION (VF) EPISODE DUE TO HIGH-RATE NON-SUSTAINED EPISODES AND SHORT VENTRICULAR INTERVALS. THE PATIENT PRESENTED TO THE EMERGENCY ROOM AND SUBSEQUENTLY EXPIRED. NO SPECIFIC CAUSE OF DEATH WAS KNOWN. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT OBTAINED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. ANALYSIS OF THE DEVICE MEMORY INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21431239&pc=NVZ | 3004209178-2025-03353 | 2021-12-28 05:00 | Death | MEDTRONIC MED REL MEDTRONIC PUERTO RICO | 2025-02-21 05:00 | NVZ | ADAPTA | Adverse Event Without Identified Device or Use Problem | Septic Shock | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPTIC SHOCK AND THE PATIENT HAD DIED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT LINKED TO THE PERFORMANCE ISSUE OF IMPLANTABLE PULSE GENERATOR (IPG). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. ANALYSIS OF THE DEVICE REVEALED NORMAL BATTERY DEPLETION. THE BATTERY INDICATOR SIGNIFYING THAT IT IS TIME FOR DEVICE REPLACEMENT OCCURRED AFTER EXPLANT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21431291&pc=DSQ | 2916596-2025-00969 | 2025-01-15 05:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21431874&pc=NPT | 2015691-2025-01194 | 2025-02-04 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-21 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY OUR SPANISH AFFILIATES, DURING IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSAXIALLY APPROACH IN A HIGH-RISK PATIENT WITH HIGHLY CALCIFIED PORCELAIN AORTA, AN AORTIC DISSECTION IN THE SINOTUBULAR JUNCTION OCCURRED WITH THE GUIDEWIRE OR COMMANDER DELIVERY SYSTEM DEVICE WHEN IMPLANTING THE VALVE. THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR A COMPLETION TO THE ENGINEERING EVALUATION. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ASCENDING AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS INCLUDING BEING HIGH RISK AND HAVING A HIGHLY CALCIFIED PORCELAIN AORTA CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21432020&pc=BZD | 2518422-2025-102245 | 2021-11-06 04:00 | Death | RESPIRONICS, INC. | 2025-02-21 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED REGARDING THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S MOTHER REPORTED THAT HER DAUGHTER PASSED AWAY ON (B)(6) 2021, AND HAS REQUESTED TO BE REMOVED FROM ALL FURTHER COMMUNICATION. THE MOTHER ALSO INDICATED THAT SHE DOES NOT HAVE THE DEVICE AVAILABLE FOR RETURN. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21432602&pc=PNJ | 9612164-2025-00961 | 2025-02-07 05:00 | Death | MEDTRONIC IRELAND | 2025-02-21 05:00 | PNJ | MICRA¿ AV2 | Pacing Problem; Capturing Problem; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; High Capture Threshold | Endocarditis; Drug Resistant Bacterial Infection | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEADLESS (IPG) EXHIBITED HIGH AND RISING THRESHOLDS AND PACING FAILURE WAS OBSERVED. IT WAS NOTED THAT PLACEMENT ANGLE OF THE LEADLESS IPG WAS WAS CHANGED IN THE X-RAY AFTER THE PACING FAILURE POSSIBLY DUE TO ONE TINE MIGHT HAVE FALLEN OUT CAUSING POOR ELECTRODE CONTACT. THE PHYSICIAN COMMENTED THAT THE TINE DISLODGING COULD BE DUE TO THE FRAGILITY OF THE MYOCARDIAL TISSUE. THE DEVICE WAS REPROGRAMMED. PACING FAILURE CONTINUED, THE DEVICE WAS PROGRAMMED OFF AND A TEMPORARY INSERTION WAS MADE. IT WAS REPORTED THAT THE PATIENT DECEASED AND CAUSE OF DEATH WAS DUE TO VALVE INFECTIOUS ENDOCARDITIS WITH POSITIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS AND HAD COVID 19 INFECTION. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRAVR] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS RISING. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS ELEVATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED H6 THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MC2AVR1] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21432728&pc=JDI | 1818910-2025-02706 |  | Death | DEPUY ORTHOPAEDICS INC US | 2025-02-21 05:00 | JDI | UNK HIP FEMORAL HEAD METAL | Adverse Event Without Identified Device or Use Problem | Pain; Metal Related Pathology |  |  | 1 | Event Description: LITIGATION COMPLAINT RECEIVED: PATIENT UNDERWENT A HIP REPLACEMENT SURGERY ON METAL ON METAL IMPLANT. PATIENT ALLEGED PAIN, METAL TOXICITY, SEVERE INJURIES AND ULTIMATELY DIED. DOI: UNKNOWN DOR; UNKNOWN AFFECTED SIDE: UNKNOWN. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. E3: INITIAL REPORTER IS A LAWYER. H3, H6 INVESTIGATION SUMMARY: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE DEVICE LOT NUMBER IS UNKNOWN, THEREFORE A¿DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21432741&pc=LPH | 1818910-2025-02707 |  | Death | DEPUY ORTHOPAEDICS INC US | 2025-02-21 05:00 | LPH | UNKNOWN HIP FEMORAL STEM | Adverse Event Without Identified Device or Use Problem | Pain; Metal Related Pathology |  |  | 1 | Event Description: LITIGATION COMPLAINT RECEIVED: PATIENT UNDERWENT A HIP REPLACEMENT SURGERY ON METAL ON METAL IMPLANT. PATIENT ALLEGED PAIN, METAL TOXICITY, SEVERE INJURIES AND ULTIMATELY DIED. DOI: UNKNOWN DOR; UNKNOWN AFFECTED SIDE: UNKNOWN. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. E3: INITIAL REPORTER IS A LAWYER. H3, H6 INVESTIGATION SUMMARY: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE DEVICE LOT NUMBER IS UNKNOWN, THEREFORE A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21432747&pc=KWA | 1818910-2025-02705 |  | Death | DEPUY ORTHOPAEDICS INC US | 2025-02-21 05:00 | KWA | UNK HIP ACETABULAR LINER METAL | Adverse Event Without Identified Device or Use Problem | Pain; Metal Related Pathology |  |  | 1 | Event Description: LITIGATION COMPLAINT RECEIVED: PATIENT UNDERWENT A HIP REPLACEMENT SURGERY ON METAL ON METAL IMPLANT. PATIENT ALLEGED PAIN, METAL TOXICITY, SEVERE INJURIES AND ULTIMATELY DIED. DOI: UNKNOWN. DOR: UNKNOWN. AFFECTED SIDE: UNKNOWN. Manufacturer Narrative: PRODUCT COMPLAINT #(B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. E3: INITIAL REPORTER IS A LAWYER. H3, H6: INVESTIGATION SUMMARY: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE DEVICE LOT NUMBER IS UNKNOWN, THEREFORE A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21432804&pc=FRO | 3005920706-2025-00004 |  | Death | DERMA SCIENCES, INC. | 2025-02-21 05:00 | FRO | MEDIHONEY GEL IN TUBE, 1.5 FL OZ FLIPCAP - STERILE | Insufficient Information | Unspecified Infection | K101793 |  | 1 | Event Description: A CUSTOMER REPORTED THAT HER HUSBAND PASSED AWAY ON (B)(6) 2024, DUE TO AN INFECTION WHILE USING THE MEDIHONEY PRODUCT. THE PRODUCT WAS APPLIED TO A SURGICAL SITE FOLLOWING A PROCEDURE RELATED TO DIABETES, DURING WHICH BONE WAS REMOVED. THE WOUND TREATED BY CLEANING THE WOUND AND CHANGE THE BANDAGE. THE INFECTION BEGAN APPROXIMATELY FIVE AND HALF WEEKS BEFORE THE PATIENT'S PASSING. THE CUSTOMER ONLY BECAME AWARE OF THE INFECTION AFTER HER HUSBAND WAS HOSPITALIZED, THE PHYSICIAN INFORMED HER ABOUT IT. PRIOR TO THE HOSPITALIZATION, SHE HAD NOTICED THAT HER HUSBAND WAS GETTING WEAKER, WHICH MIGHT HAVE CAUSED A FALL. ACCORDING TO INFORMATION PROVIDED BY THE CUSTOMER, THE MEDICAL RECORDS AND AUTOPSY REPORT ARE NOT AVAILABLE. Manufacturer Narrative: AN INVESTIGATION HAS BEEN INITIATED BASED ON THE REPORTED INFORMATION. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW-UP REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED FIELDS: B2: PREVIOUSLY: OTHER SERIOUS (IMPORTANT MEDICAL EVENTS) - CORRECTED TO: DEATH. H1: PREVIOUSLY: SERIOUS INJURY - CORRECTED TO: DEATH. UPDATED FIELD - H9. Manufacturer Narrative: UPDATED FIELDS: D4, D9, G3, G6, H2, H3, H6, H11. THE MEDIHONEY GEL (ID 31815) WAS NOT RETURNED FOR EVALUATION AND LOT NUMBER INFORMATION HAS NOT BEEN PROVIDED; THEREFORE, AN EVALUATION OF THE DEVICE COULD NOT BE PERFORMED, AND DEVICE HISTORY RECORD (DHR) COULD NOT BE REVIEWED. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED. THE CUSTOMER INDICATED 3 POSSIBLE LOT NUMBERS FOR PRODUCT ID 31815: LOTS 2328, 2332, 2333. THE REVIEW FOCUSED ON IDENTIFYING POSSIBLE DEVIATIONS OR ANOMALIES RELATED TO POSSIBLE CAUSES OF FLIP CAP - STERILE. NO DEVIATIONS (NCS) OR DOCUMENTATION ERRORS WERE IDENTIFIED IN THE DHR RECORDS. THE INSPECTION AND TESTING RESULTS COMPLIED WITH ESTABLISHED SPECIFICATIONS. ADDITIONAL INFORMATION RECEIVED: PATIENT'S DAUGHTER (B)(6) STATED: "DAD WAS ADMITTED TO THE HOSPITAL IN OCTOBER WITH A BROKEN HIP. HE HAD AN INFECTION IN HIS FOOT THAT HE WAS TREATING WITH MEDI HONEY, AND HE ENDED UP HAVING A METATARSAL BONE REMOVED DUE TO THE INFECTION. HE THEN WENT IN FOR HIP SURGERY AND PASSED AWAY."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21433033&pc=DSQ | 2916596-2025-01181 | 2023-10-01 04:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "OMENTAL WRAPPING FOR REFRACTORY DRIVELINE INFECTION AFTER HEARTMATE 3 IMPLANTATION" THAT DRIVELINE INFECTIONS (DLI) ARE A FREQUENT COMPLICATION OF LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY AND ACCOUNT FOR SIGNIFICANT MORBIDITY AND MORTALITY. THE STUDY RETROSPECTIVELY REVIEWED PATIENTS WHO UNDERWENT HEARTMATE (HM) 3 IMPLANTATION AT THE INSTITUTION BETWEEN NOVEMBER 2014 AND OCTOBER 2023. DURING THE STUDY PERIOD, 336 HM3 LVADS WERE IMPLANTED AT THE CENTER. FIFTY-FOUR PATIENTS (16.1%) DEVELOPED AT LEAST ONE DLI. OF THESE, 6 (11.1%) UNDERWENT INCISION AND DRAINAGE (I&D) WITH OMENTAL WRAPPING FOR REFRACTORY DLI. MEDIAN AGE OF THE COHORT WAS 57.9 YEARS (RANGE: 45.2, 65.8) AT INTERVENTION. ALL PATIENTS HAD UNDERGONE AT LEAST ONE PREVIOUS I&D FOR DLI FOLLOWED BY VACUUM-ASSISTED WOUND CLOSURE, WITH MEDIAN TIME TO INITIAL DLI OF 2.1 YEARS (RANGE: 0.3, 6.2). TWO HAD A HISTORY OF VARIABLE ADHERENCE TO ANTIBIOTIC THERAPY. MEDIAN TIME FROM HM3 IMPLANTATION TO OMENTAL FLAP WAS 3.6 YEARS (RANGE: 0.4, 6.5). COMPUTED TOMOGRAPHY REVEALED THAT INFECTION HAD EXTENDED CLOSE TO THE PUMP IN ALL PATIENTS; 4 (66.7%) PATIENTS¿ WOUND CULTURES YIELDED POSITIVE RESULTS. INTRAOPERATIVELY, IN ADDITION TO EXTENSIVE I&D, A LAPAROSCOPICALLY HARVESTED OMENTAL FLAP WAS WRAPPED AROUND DRIVELINE. SKIN WAS CLOSED PRIMARILY. THERE WAS ONE DEATH IN A PATIENT WHO HAD LEFT THE HOSPITAL AGAINST MEDICAL ADVICE WITHOUT ANTIBIOTIC THERAPY. HE WAS READMITTED ON POSTOPERATIVE DAY 18 WITH SEVERE SEPSIS SECONDARY TO S. AUREUS BACTEREMIA AND DIED ON DAY 31. THE FIVE REMAINING PATIENTS CONTINUE ON THE HM3 WITH NO EVIDENCE OF DLI RECURRENCE ON CHRONIC ANTIBIOTIC SUPPRESSION AT A MEDIAN OF 2.4 MONTHS (RANGE: 0.0, 3.8) AFTER THE PROCEDURE. THE FOLLOWING DETAILS WERE GIVEN FOR 4/6 PATIENTS WHO UNDERWENT INCISION AND DRAINAGE (I&D) WITH OMENTAL WRAPPING: PATIENT 1: 59 YEAR OLD MALE AT IMPLANT. COMORBIDITIES AND MEDICAL HISTORY: HYPERTENSION, DIABETES MELLITUS, CORONARY ARTERY DISEASE, CEREBROVASCULAR ACCIDENT, PERIPHERAL VASCULAR DISEASE, PRIOR CORONARY ARTERY BYPASS GRAFTING. PREVIOUS DRIVELINE INFECTION 6.2 YEARS AFTER HEARTMATE (HM) 3 IMPLANT, RECURRENCE 6.5 YEARS AFTER HM 3 IMPLANT. CULTURE RESULTS FOR BOTH INSTANCES WERE NEGATIVE WITH RARE POLYMORPHONUCLEAR NEUTROPHILS. INFECTION WAS MANAGED WITH OMENTAL FLAP. SUBJECT FOLLOWED UP 2.9 MONTHS AFTER OMENTAL FLAP AND CURRENT STATUS IS ALIVE ON THE DEVICE. PATIENT 2: 54 YEAR OLD MALE AT IMPLANT. COMORBIDITIES AND MEDICAL HISTORY: HYPERTENSION, SUBMASSIVE PULMONARY EMBOLISM, POLYSUBSTANCE USE DISORDER. PREVIOUS DRIVELINE INFECTION 3.5 YEARS AFTER HEARTMATE (HM) 3 IMPLANT, CULTURES WERE POSITIVE FOR METHICILLIN SENSITIVE STAPHYLOCOCCUS AUREUS (MSSA) WITH VARIABLE ADHERENCE TO ANTIBIOTIC THERAPY. DRIVELINE INFECTION RECURRENCE 4.1 YEARS AFTER HM 3 IMPLANT, CULTURES POSITIVE FOR MSSA AND ENTEROBACTER CLOACAE. INFECTION WAS MANAGED WITH OMENTAL FLAP. SUBJECT FOLLOWED UP 1 MONTH AFTER OMENTAL FLAP AND CURRENT STATUS IS DECEASED. PATIENT 3: 41 YEAR OLD MALE AT IMPLANT. COMORBIDITIES AND MEDICAL HISTORY: CEREBROVASCULAR ACCIDENT, ATRIAL FIBRILLATION, IMPLANTABLE CARDIAC DEFIBRILLATOR, POLYSUBSTANCE USE DISORDER, HISTORY OF POOR ADHERENCE PREVENTING LISTING FOR TRANSPLANT. PATIENT HAD 2 PREVIOUS DRIVELINE INFECTIONS 1.6 YEARS AFTER IMPLANT WITH CULTURES POSITIVE FOR S. EPIDERMIDIS. DRIVELINE INFECTION RECURRENCE 4 YEARS AFTER IMPLANT, CULTURES WERE POSITIVE FOR MORGANELLA MORGANIL. INFECTION WAS MANAGED WITH OMENTAL FLAP. SUBJECT FOLLOWED UP 4 MONTHS AFTER OMENTAL FLAP AND CURRENT STATUS IS ALIVE ON THE DEVICE. PATIENT 4: 57 YEAR OLD MALE AT IMPLANT. COMORBIDITIES AND MEDICAL HISTORY: HYPERTENSION, HYPERLIPIDEMIA, DIABETES MELLITUS, CEREBROVASCULAR ACCIDENT, IMPLANTABLE CARDIAC DEFIBRILLATOR. PATIENT HAD 3 PREVIOUS DRIVELINE INFECTIONS. THE FIRST WAS 2.2 YEARS AFTER IMPLANT AND CULTURES WERE POSITIVE FOR P. AERUGINOSA. THE SECOND WAS 2.7 YEARS AFTER IMPLANT AND CULTURES WERE POSITIVE FOR A. XYLOSOXIDANS. THE THIRD WAS 3.1 YEARS AFTER IMPLANT AND CULTURES WERE POSITIVE FOR P. AERUGINOSA AND CANDIDA ORTHOPSILOSIS. THERE WAS VARIABLE ADHERENCE TO ANTIBIOTIC THERAPY. DRIVELINE INFECTION RECURRENCE 3.3 YEARS AFTER IMPLANT, CULTURES WERE NEGATIVE WITH RARE POLYMORPHONUCLEAR NEUTROPHILS. INFECTION WAS MANAGED WITH OMENTAL FLAP. SUBJECT FOLLOWED UP 3.3 MONTHS AFTER OMENTAL FLAP AND CURRENT STATUS IS ALIVE ON THE DEVICE. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01OCT2023 AS PATIENTS WERE IMPLANTED BETWEEN NOVEMBER 2014 AND OCTOBER 2023. AUTHOR INFORMATION: VINOGRADSKY, A., HYNDS, M., KAKU, Y., ARNELL, T., POWERS, A., YUZEFPOLSKAYA, M., COLOMBO, P., SAYER, G., URIEL, N., NAKA, Y., & TAKEDA, K. (2024). OMENTAL WRAPPING FOR REFRACTORY DRIVELINE INFECTION AFTER HEARTMATE 3 IMPLANTATION. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S598¿S599. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.910. BAYLOR SCOTT AND WHITE HEALTH, FORT WORTH, TX. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION E1: REPORTER ESTABLISHMENT NAME CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21433995&pc=OZD | 1220648-2025-26251 | 2025-01-27 05:00 | Death | ABIOMED, INC. | 2025-02-21 05:00 | OZD | IMPELLA CP | Difficult to Remove | Perforation of Vessels | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THERE WAS A VENTRICLE PERFORATION. AFTER INSERTION OF THE IMPELLA CP, THE PATIENT'S PRESSURE AS WELL AS THE PUMP'S FLOWS DECREASED. ADDITIONALLY, LARGE AMOUNTS OF BLOOD WAS DRAINED. SEVERAL ANGIOGRAMS AND A THORACTOMY WERE DONE WHICH SHOWED NO PERFORATION. HOWEVER, AFTER THE PATIENT'S CHEST WAS OPENED AND PLACED ON BYPASS THE DOCTOR IDENTIFIED A HOLE IN THE LEFT VENTRICLE FOR WHICH THEN DECISION TO EXPLANT THE IMPELLA CP WAS MADE. THE PATIENT EXPIRED BEFORE THE IMPELLA COULD BE REMOVED. DURING IMPELLA CP EXPLANT, THERE WERE ADDITIONAL REMOVAL ISSUES ENCOUNTERED AS THE PUMP GOT STUCK AND WAS UNABLE TO BE REMOVED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ ORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434224&pc=DSQ | 2916596-2025-00956 | 2024-10-02 04:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSFERRED FROM THE CLINIC TO HOSPICE, WHERE THEY PASSED AWAY A FEW HOURS LATER. Manufacturer Narrative: B3: EVENT DATE HAS BEEN ESTIMATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE EVENT WAS DETERMINED TO BE SUPPLEMENTAL INFORMATION, AND THE INVESTIGATION FINDINGS WERE REPORTED UNDER MFR# 2916596-2024-06847.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434264&pc=DSQ | 2916596-2025-01033 | 2024-03-26 04:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE REPORTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT NO ADDITIONAL INFORMATION COULD BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434280&pc=NGV | 2124215-2025-10331 | 2024-11-18 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-21 05:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Difficult to Open or Close | Pericardial Effusion; Thrombosis/Thrombus |  |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE. IT WAS REPORTED VIA JOURNAL ARTICLE SERIOUS PATIENT INJURIES OCCURRED. THE AMULET IDE TRIAL USED 1:1 RANDOMIZATION TO COMPARE THE SAFETY AND EFFECTIVENESS OF THE AMULET OCCLUDER TO THE WATCHMAN DEVICE. ELIGIBLE PATIENTS WERE 18 YEARS OF AGE OR OLDER WITH DOCUMENTED PAROXYSMAL, PERSISTENT, OR PERMANENT NVAF AND DEEMED TO BE AT HIGH RISK OF STROKE OR SYSTEMIC EMBOLISM (SE). A TOTAL OF 1,878 PATIENTS WERE RANDOMIZED FROM SEPTEMBER 2016 THROUGH MARCH 2019, AND 1,833 UNDERWENT A DEVICE IMPLANT ATTEMPT (917 AMULET OCCLUDER GROUP PATIENTS AND 916 WATCHMAN DEVICE GROUP PATIENTS). RESULTS: THE 5-YEAR FOLLOW-UP VISIT WAS COMPLETED IN 1,140 PATIENTS (595 AMULET OCCLUDER GROUP AND 545 WATCHMAN DEVICE GROUP). PATIENTS THAT RECEIVED THE WATCHMAN CLOSURE DEVICE WERE REPORTED TO HAVE EXPERIENCED SERIOUS INJURIES. 58 PATIENTS EXPERIENCED AN ISCHEMIC STROKE, 14 PATIENTS EXPERIENCED A PERICARDIAL EFFUSION, 9 PATIENTS HAD A HEMATOMA, 2 PATIENTS HAD A PLEURAL EFFUSION, AN UNKNOWN NUMBER OF PATIENTS EXPERIENCED A SYSTEMIC EMBOLISM, DEVICE RELATED THROMBUS EVENT, AND HAD A DEVICE THAT DID NOT SEAL. IT WAS REPORTED THAT A TOTAL OF 117 PATIENT DEATH OCCURRED. 17 PATIENTS EXPERIENCED FATAL CEREBRAL VASCULAR ACCIDENTS, 8 PATIENTS EXPERIENCED DEVICE RELATED THROMBUS, 31 PATIENTS HAD A DEVICE THAT DID NOT SEAL, 1 PATIENT HAD A PERICARDIAL EFFUSION THAT LED TO A PROCEDURAL PATIENT DEATH, AND THE REMAINING PATIENTS WERE NOT NOTED WITH OTHER PATIENT INJURIES OTHER THAN DEATH. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS (B)(6) 2016 THROUGH (B)(6) 2019. LAKKIREDDY, D., ELLIS, C. R., THALER, D., SWARUP, V., GAMBHIR, A., HERMILLER, J., ... & WINDECKER, S. (2024). 5-YEAR RESULTS FROM THE AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER RANDOMIZED CONTROLLED TRIAL. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434301&pc=NPT | 2015691-2025-01210 | 2025-01-30 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-21 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Obstruction/Occlusion | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS AFFILIATES IN FRANCE, A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) WAS PERFORMED USING A 26MM SAPIEN 3 ULTRA VALVE VIA A TRANSFEMORAL APPROACH. DURING DEPLOYMENT, THE PHYSICIAN DID NOT PERFORM AN AORTOGRAM TO CONFIRM THE VALVE'S POSITION. ALTHOUGH THE VALVE WAS THOUGHT TO BE IN THE ANNULUS, IT EMBOLIZED TO THE LEFT VENTRICLE. THIS CAUSED THE IMPLANTED VALVE TO OCCLUDE THE NATIVE VALVE, RESULTING IN THE PATIENT BECOMING HEMODYNAMICALLY UNSTABLE. A DECISION WAS MADE TO PERFORM SURGICAL INTERVENTION, BUT UNFORTUNATELY, THE PATIENT PASSED AWAY BEFORE THE VALVE COULD BE EXPLANTED. Manufacturer Narrative: CORRECTION TO H6: TYPE OF INVESTIGATION, INVESTIGATION CONCLUSION. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED; NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE REPORTED EVENT OF VALVE EMBOLIZATION WAS UNABLE TO BE CONFIRMED DUE TO THE UNAVAILABILITY OF PROCEDURAL IMAGERY AND MEDICAL RECORDS. THERE WAS NO ALLEGATION OR INDICATION THAT A DEVICE MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ADDITIONALLY, A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. AS REPORTED, "DURING DEPLOYMENT, THE PHYSICIAN DID NOT PERFORM AN AORTOGRAM TO CONFIRM THE POSITION OF THE VALVE. IT WAS THOUGHT THAT THE VALVE WAS IN THE ANNULUS BUT IT EMBOLIZED TO THE LEFT VENTRICLE." AS PER THE TRAINING MANUAL, "POSITION THV WITH BOTTOM OF CENTER MARKER AT BASE OF CUSPS OR SLIGHTLY ABOVE" AND "FOR OPTIMAL RESULTS, PLACE THE ENTIRE CENTER MARKER WITHIN THE INITIAL CENTER MARKER ZONE". IN THIS CASE, IT WAS REPORTED THAT THE THV POSITION INSIDE THAT NATIVE VALVE WAS NOT CHECKED BEFORE DEPLOYMENT. ADDITIONAL INFORMATION INDICATED THAT THE THV EMBOLIZED DIRECTLY TO THE LV AND WAS NEVER POSITIONED IN THE ANNULUS, WHICH COULD CAUSE THE INABILITY FOR THE VALVE TO ANCHOR PROPERLY IN THE TARGET SITE (ANNULUS), CAUSING IT TO EMBOLIZE INTO THE VENTRICLE. AVAILABLE INFORMATION SUGGESTS THAT USE ERROR (VALVE POSITIONING TECHNIQUE) MAY HAVE CONTRIBUTED TO THE EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PROCEDURAL FACTORS (IMPROPER POSITIONING MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434387&pc=NKM | 2135147-2025-00893 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-02-21 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "HOW THE COAPT TRIAL AFFECTED THE SELECTION OF PATIENTS WITH SECONDARY MITRAL REGURGITATION UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE GIOTTO REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE POTENTIAL IMPACT OF THE COAPT TRIAL BY ANALYZING THE TEMPORAL TRENDS OF BASELINE CHARACTERISTICS AND OUTCOME OF SECONDARY MITRAL REGURGITATION (SMR) PATIENTS UNDERGOING MITRACLIP (MC) INCLUDED IN THE GIOTTO REGISTRY. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THE SELECTION PROCESS OF SMR PATIENTS UNDERGOING MC CHANGED OVER TIME, WITH A SIGNIFICANT IMPACT OF THE COAPT PUBLICATION MAINLY ON THE TREND OF NYHA CLASS. AT 1-YEAR, PATIENTS IN THE POST-COAPT GROUP SHOWED A NON-SIGNIFICANT REDUCTION IN THE INCIDENCE OF HOSPITALIZATION FOR HF AND ALL-CAUSE DEATH. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARCO DI MAIO AT UNIVERSITY OF SALERNO INSTITUTION WITH CORRESPONDING EMAIL MDIMAIO@UNISA.IT]. THE TIME FRAME OF THE STUDY WAS JANUARY 2016 TO MARCH 2020. A TOTAL OF 1184 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 1184 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, PREVIOUS CARDIAC SURGERY, HISTORY OF HEART FAILURE. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, PREVIOUS CARDIAC SURGERY, HISTORY OF HEART FAILURE. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "HOW THE COAPT TRIAL AFFECTED THE SELECTION OF PATIENTS WITH SECONDARY MITRAL REGURGITATION UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE GIOTTO REGISTRY." INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434388&pc=OZD | 1220648-2025-26256 | 2025-01-30 05:00 | Death | ABIOMED, INC. | 2025-02-21 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED VENTRICULAR TACHYCARDIA REQUIRING CARDIOVERSION AND A STROKE. THE PATIENT WOULD LATER EXPIRE. THE IMPELLA WAS PLACED, AND THE PATIENT WAS SENT TO THE UNIT ON SUPPORT. THREE DAYS LATER, THE PATIENT EXPERIENCED A VENTRICULAR TACHYCARDIA RUN THAT WAS RESOLVED WITH A SHOCK. AN ECHOCARDIOGRAM WAS ORDERED, AND THE DEVICE WAS REPOSITIONED, PULLED BACK ONE CENTIMETER. APPROXIMATELY SIX DAYS LATER THE PATIENT DEVELOPED "SEVERE" LEFT SIDE FACIAL DROOP. THE PATIENT HAD EXPERIENCED A STROKE NOTED ON A COMPUTED TOMOGRAPHY SCAN. THE PATIENT WAS MOVED TO COMFORT CARE, AND FOUR DAYS LATER THE PATIENT EXPIRED. THE FAMILY WITHDREW CARE; THE PUMP WAS TURNED DOWN TO PERFORMANCE LEVEL P-2, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434500&pc=DSQ | 2916596-2025-01321 | 2025-02-05 05:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434513&pc=OZD | 1220648-2025-26257 | 2025-01-27 05:00 | Death | ABIOMED, INC. | 2025-02-21 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE COMPLAINANT REPORTED THAT A PATIENT BECAME UNRESPONSIVE DURING SUPPORT AND WAS FOUND TO HAVE MULTIPLE SEVERE SUBARACHNOID HEMORRHAGES OF UNKNOWN ETIOLOGY. IMPELLA SUPPORT WAS DISCONTINUED IN RESPONSE TO THE CONDITION AND THE PATIENT PASSED AWAY. IT IS UNKNOWN IF SUPPORT CONTRIBUTED TO THE NOTED CONDITION AND DEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434539&pc=NPT | 2015691-2025-01217 | 2025-01-29 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-21 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Obstruction/Occlusion | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN SWEDEN. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE PROCEDURE, AROUND 2 MINUTES AFTER THE VALVE DEPLOYMENT, THE PATIENT EXPERIENCED SOME PAIN IN THE LEFT ARM AND A ST SEGMENT DEPRESSION WAS NOTICED IN ELECTROCARDIOGRAM. THE SYMPTOMS GOT WORST AND PERCUTANEOUS CORONARY INTERVENTION OPERATORS WERE CALLED. IT WAS OBSERVED THAT A BULKY CALCIFICATION WAS PUSHED DURING DEPLOYMENT AND WAS OBSTRUCTING PARTIALLY THE OSTIUM OF THE LEFT MAIN CORONARY ARTERY. FLOW WAS OBSERVED IN THE CORONARY ARTERY. THE TEAM ATTEMPTED TO INSERT A WIRE FOR IMPLANTING A STENT INTO THE LEFT CORONARY MAIN ARTERY, BUT IT WAS NOT POSSIBLE. FINALLY, THE TEAM DECIDED TO PERFORM A CORONARY ARTERY BYPASS GRAFT, BUT ONCE THE PATIENT WAS PREPARED FOR SURGERY, THE PATIENT WENT TO CARDIAC ARREST AND RESUSCITATION WAS STARTED. FINALLY, THE PATIENT PASSED AWAY.Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN SWEDEN. AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE PROCEDURE, AROUND TWO MINUTES AFTER THE VALVE DEPLOYMENT, THE PATIENT PRESENTED SOME PAIN IN THE LEFT ARM AND A ST SEGMENT DEPRESSION WAS NOTICED ON ELECTROCARDIOGRAM. THE SYMPTOMS GOT WORST AND PERCUTANEOUS CORONARY INTERVENTION OPERATORS WERE CALLED. IT WAS OBSERVED THAT A BULKY CALCIFICATION WAS PUSHED DURING DEPLOYMENT AND WAS OBSTRUCTING PARTIALLY THE OSTIUM OF THE LEFT MAIN CORONARY ARTERY. FLOW WAS OBSERVED IN THE CORONARY ARTERY. THE TEAM ATTEMPTED TO INSERT A WIRE FOR IMPLANTING A STENT INTO THE LEFT CORONARY MAIN ARTERY BUT IT WAS NOT POSSIBLE. FINALLY, THEY DECIDED TO PERFORM A CORONARY ARTERY BYPASS GRAFT, BUT ONCE THE PATIENT WAS PREPARED FOR SURGERY, SHE WENT TO CARDIAC ARREST AND RESUSCITATION WAS STARTED. FINALLY, THE CORONARY ARTERY BYPASS GRAFT WAS DONE AND THE PATIENT RESULT WAS STABLE UNDER ECMO. FINALLY, THE PATIENT PASSED AWAY. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO CORRECT B5. ADDED H6: HEALTH EFFECT - IMPACT CODE AND H6: TYPE OF INVESTIGATION CODES. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT FACTORS (BULKY CALCIFICATION WAS PUSHED DURING DEPLOYMENT AND WAS OBSTRUCTING PARTIALLY THE OSTIUM OF THE LEFT MAIN CORONARY) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434608&pc=DSQ | 2916596-2025-00921 | 2013-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿SAFETY AND HEMODYNAMIC IMPLICATIONS OF TRANSSEPTAL ACCESS FOR VENTRICULAR TACHYCARDIA ABLATION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH VENTRICULAR TACHYCARDIA, VENTRICULAR ANEURYSM, PULMONARY HYPERTENSION, MODERATE OR SEVERE RIGHT VENTRICLE DYSFUNCTION, TRANSPLANTATION, AND DEATH. A TOTAL OF 15 PATIENTS WITH LVAD UNDERWENT 19 VENTRICULAR TACHYCARDIA ABLATION PROCEDURES WITH SINGLE TRANSSEPTAL ACCESS FOR LEFT VENTRICLE (LV) ABLATION WITHOUT ATRIAL SEPTAL DEFECT (ASD) CLOSURE. LVADS INCLUDED HEARTMATE 3 AND HEARTWARE HVAD. 6 HAD PULMONARY HYPERTENSION, AND 7 HAD MODERATE OR SEVERE RIGHT VENTRICLE DYSFUNCTION. POST-PROCEDURE ECHOCARDIOGRAMS (N = 14; MEDIAN TIME TO ECHO 57 DAYS) SHOWED NO INTRACARDIAC SHUNTING OR PERSISTENT ASD. RIGHT HEART CATHETERIZATION (RHC) (PRE: N = 14; POST: N = 8, AT MEDIAN FOLLOW-UP 92.5 DAYS) SHOWED SIMILAR PRE-AND POST-ABLATION RA PRESSURES (11.5 VS. 12 MMHG, P = 0.17) AND WEDGE PRESSURES (20.5 VS. 19.5 MMHG, P = 0.99). SIGNIFICANT SHUNTING WAS ABSENT ON PRE- VERSUS POST-ABLATION ARTERIAL (MEDIAN 95% VS. 99%) AND PA OXYGEN SATURATIONS (58% VS. 62%). NO CLINICALLY SIGNIFICANT RIGHT-LEFT SHUNTING OR HYPOXIA/HYPOXEMIA OCCURRED POST-ABLATION. OVER MEAN FOLLOW-UP OF 581 ± 418 DAYS AFTER LAST ABLATION, 10 PATIENTS REMAINED ALIVE, 1 OF WHOM UNDERWENT HEART TRANSPLANTATION. 5 PATIENTS DIED DURING FOLLOW-UP (4 FROM CONDITIONS UNRELATED TO ABLATION AND 1 FROM RECURRENT VT). AT 1 YEAR AFTER LAST ABLATION, VA RECURRED IN 5 (30%) PATIENTS (FIG. 1A). Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEPARTMENT OF INTERNAL MEDICINE, UNIVERSITY OF MICHIGAN, 1500 E MEDICAL CENTER DRIVE TC 311Q, ANN ARBOR, MI 48109-5853, USA. DIVISION OF CARDIOVASCULAR MEDICINE, UNIVERSITY OF MICHIGAN MEDICAL CENTER, ANN ARBOR, MI 48109-5869, USA. ELECTROPHYSIOLOGY SECTION, DIVISION OF CARDIOVASCULAR MEDICINE, UNIVERSITY OF MICHIGAN, 1425 E. ANN ST. FLOOR 3 RECEPTION C, ANN ARBOR, MI 48103, USA DEPARTMENT OF CARDIOVASCULAR SURGERY, UNIVERSITY OF MICHIGAN, ANN ARBOR, MI 48103, USA. CURTIS, B., DESHMUKH, A., LARSON, J., GHANBARI, H., GHANNAM, M., CRAWFORD, T., LATCHAMSETTY, R., PAGANI, F., BOGUN, F., & LIANG, J. (2024). SAFETY AND HEMODYNAMIC IMPLICATIONS OF TRANSSEPTAL ACCESS FOR VENTRICULAR TACHYCARDIA ABLATION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY, 67(6), P.1319-1321. HTTPS://DOI.ORG/10.1007/S10840-024-01768-0. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434612&pc=NKM | 2135147-2025-00899 | 2014-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-02-21 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "RESIDUAL MITRAL REGURGITATION INTERACTS WITH TRANSMITRAL MEAN PRESSURE GRADIENT TO MODIFY THE ASSOCIATION WITH MORTALITY FOLLOWING TRANSCATHETER EDGE-TO-EDGE" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN INTRAOPERATIVELY MEASURED RESIDUAL MITRAL REGURGITATION (RMR) AND TMPG AND 1-YEAR MORTALITY AMONG PATIENTS UNDERGOING MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR TO FACILITATE DECISIONS ON ADDITIONAL DEVICES. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT NOTH RMR AND TMPG WERE NONLINEARLY ASSOCIATED WITH 1-YEAR MORTALITY. AT LOW LEVELS OF RMR, CHANGES IN TMPG ARE ASSOCIATED WITH ONLY SMALL CHANGES IN THE RISK OF DEATH. CONVERSELY, AT HIGHER LEVELS OF RMR, EVEN SMALL CHANGES IN TMPG ARE ASSOCIATED WITH LARGER CHANGES IN THE ABSOLUTE RISK OF DEATH. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NEAL DUGGAL AT UNIVERSITY OF MICHIGAN INSTITUTION WITH CORRESPONDING EMAIL NEALD@MED.UMICH.EDU]. THE TIME FRAME OF THE STUDY WAS JANUARY 2014 TO DECEMBER 2020. A TOTAL OF 570 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 570 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, MYOCARDIAL INFARCTION, STROKE, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH.Event Description: THE ARTICLE "RESIDUAL MITRAL REGURGITATION INTERACTS WITH TRANSMITRAL MEAN PRESSURE GRADIENT TO MODIFY THE ASSOCIATION WITH MORTALITY FOLLOWING TRANSCATHETER EDGE-TO-EDGE" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN INTRAOPERATIVELY MEASURED RESIDUAL MITRAL REGURGITATION (RMR) AND TMPG AND 1-YEAR MORTALITY AMONG PATIENTS UNDERGOING MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR TO FACILITATE DECISIONS ON ADDITIONAL DEVICES. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT NOTH RMR AND TMPG WERE NONLINEARLY ASSOCIATED WITH 1-YEAR MORTALITY. AT LOW LEVELS OF RMR, CHANGES IN TMPG ARE ASSOCIATED WITH ONLY SMALL CHANGES IN THE RISK OF DEATH. CONVERSELY, AT HIGHER LEVELS OF RMR, EVEN SMALL CHANGES IN TMPG ARE ASSOCIATED WITH LARGER CHANGES IN THE ABSOLUTE RISK OF DEATH. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NEAL DUGGAL AT UNIVERSITY OF MICHIGAN INSTITUTION WITH CORRESPONDING EMAIL NEALD@MED.UMICH.EDU]. THE TIME FRAME OF THE STUDY WAS (B)(6) 2014 TO (B)(6) 2020. A TOTAL OF 570 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 570 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, MYOCARDIAL INFARCTION, STROKE, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: RESIDUAL MITRAL REGURGITATION INTERACTS WITH TRANSMITRAL MEAN PRESSURE GRADIENT TO MODIFY THE ASSOCIATION WITH MORTALITY FOLLOWING TRANSCATHETER EDGE-TO-EDGE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, MYOCARDIAL INFARCTION, STROKE, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434882&pc=NKM | 2135147-2025-00908 | 2019-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-21 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "EVOLUTION OF COAGULATION AND PLATELET ACTIVATION MARKERS AFTER TRANSCATHETER EDGE-TO-EDGE MITRAL VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE DEGREE AND THE TIMING OF COAGULATION AND PLATELET MARKER ACTIVATION AFTER TEER. DEVICES MENTIONED INCLUDE MITRACLIP AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT TEER IS ASSOCIATED WITH AN ACUTE ACTIVATION OF THE COAGULATION SYSTEM, WITH NO INCREASE IN PLATELET ACTIVATION MARKERS. HENCE, THE USE OF DUAL APT IS QUESTIONABLE IN THIS POPULATION. OUR RESULTS RAISE THE HYPOTHESIS THAT THE OPTIMAL ANTITHROMBOTIC THERAPY AFTER TEER COULD BE SHORT-TERM ACT OVER APT. FURTHER LARGER STUDIES ARE WARRANTED. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS JEAN-MICHEL PARADIS AT INSTITUTION WITH CORRESPONDING EMAIL JEAN-MICHEL.PARADIS@CRIUCPQ.ULAVAL.CA]. THE TIME FRAME OF THE STUDY WAS MAY 2019 TO NOVEMBER 2021. A TOTAL OF 46 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT OF PATIENTS RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, PREVIOUS MYOCARDIAL INFARCTION, PREVIOUS STROKE, PREVIOUS TRANSIENT ISCHEMIC ATTACK, HYPERTENSION, DIABETES. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED STROKE, HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH.Event Description: THE ARTICLE "EVOLUTION OF COAGULATION AND PLATELET ACTIVATION MARKERS AFTER TRANSCATHETER EDGE-TO-EDGE MITRAL VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE DEGREE AND THE TIMING OF COAGULATION AND PLATELET MARKER ACTIVATION AFTER TEER. DEVICES MENTIONED INCLUDE MITRACLIP AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT TEER IS ASSOCIATED WITH AN ACUTE ACTIVATION OF THE COAGULATION SYSTEM, WITH NO INCREASE IN PLATELET ACTIVATION MARKERS. HENCE, THE USE OF DUAL APT IS QUESTIONABLE IN THIS POPULATION. OUR RESULTS RAISE THE HYPOTHESIS THAT THE OPTIMAL ANTITHROMBOTIC THERAPY AFTER TEER COULD BE SHORT-TERM ACT OVER APT. FURTHER LARGER STUDIES ARE WARRANTED. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS JEAN-MICHEL PARADIS AT INSTITUTION WITH CORRESPONDING EMAIL JEAN-MICHEL.PARADIS@CRIUCPQ.ULAVAL.CA]. THE TIME FRAME OF THE STUDY WAS MAY 2019 TO NOVEMBER 2021. A TOTAL OF 46 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT OF PATIENTS RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, PREVIOUS MYOCARDIAL INFARCTION, PREVIOUS STROKE, PREVIOUS TRANSIENT ISCHEMIC ATTACK, HYPERTENSION, DIABETES. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED STROKE, HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, PREVIOUS MYOCARDIAL INFARCTION, PREVIOUS STROKE, PREVIOUS TRANSIENT ISCHEMIC ATTACK, HYPERTENSION, DIABETES. COMPLICATIONS REPORTED INCLUDED STROKE, HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: 4582. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE: ARTICLE TITLE: EVOLUTION OF COAGULATION AND PLATELET ACTIVATION MARKERS AFTER TRANSCATHETER EDGE-TO-EDGE MITRAL VALVE REPAIR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21434990&pc=OZD | 1220648-2024-06959 | 2022-09-23 04:00 | Death | ABIOMED, INC. | 2025-02-21 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CPWAS PLACED IN A 74-YEAR-OLD FEMALE PATIENT FOR MECHANICAL CIRCULATORY SUPPORT. OOZING AT THE IMPELLA ACCESS WAS NOTED. DRESSING WAS TAKEN DOWN AND THE REPOSITION SHEATH WAS RE-SUTURED WITH FORWARD TENSION, ANGEL MATCHED AND WAS REDRESSED. NO NEW OOZING WAS NOTED. HEPARIN WAS WITHHELD AFTER BLEEDING WAS AGAIN NOTED FROM THE ACCESS SITES. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT SUBSEQUENTLY EXPIRED AT EXPLANT. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEED COULD NOT BE DETERMINED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THIS REPORT IS BEING FILED AS A PART OF THE RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21435204&pc=DSQ | 2916596-2025-01145 | 2023-04-01 04:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿REOPERATIVE VENTRICULAR ASSIST DEVICE INSERTION FOR PATIENTS WITH HISTORY OF CORONARY ARTERY BYPASS GRAFTING: LESS INVASIVE APPROACH VERSUS FULL STERNOTOMY¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT VENTRICULAR (RV) FAILURE, STROKE, BLEEDING, DEATH, RESPIRATORY FAILURE, AND RENAL DYSFUNCTION. THIS STUDY EVALUATED 27 PATIENTS WITH HISTORY OF CORONARY ARTERY BYPASS GRAFTING (CABG) THAT RECEIVED REOPERATIVE LEFT VENTRICULAR ASSIST DEVICE (LVAD) INSERTION BETWEEN MAR2012 AND APR2023. THE AIM OF THE STUDY WAS TO ANALYZE IF LESS INVASIVE APPROACH FOR LVAD IMPLANTATION MAY IMPROVE OUTCOMES BY MINIMIZING DISSECTION AROUND THE HEART, POTENTIALLY REDUCING BLEEDING, GRAFT INJURY, OR RV FAILURE. 16 PATIENTS UNDERWENT A REDO FULL STERNOTOMY (GROUP FS) AND 11 PATIENTS HAD LESS INVASIVE APPROACH (GROUP LI). THE MEDIAN AGES IN GROUPS FS AND LI WERE 68 AND 70 YEARS, RESPECTIVELY (P=0.83), WITH A MEDIAN DURATION OF 96 [24, 192] MONTHS SINCE THE LAST CARDIAC SURGERY. SIX PATIENTS HAD PREOPERATIVE VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORT (GROUP FS, 18.8% [N=3] VS. GROUP LI, 27.3% [N=3], P=0.662). HEARTMATE 2 WAS EXCLUSIVELY USED IN THE FS GROUP (FS, 56.3% [N=9/16] VS. LI, 0%), WHILE HEARTWARE WAS PREDOMINANTLY USED IN GROUP LI (FS, 31.3%, [N=5/16] VS. LI, 63.6% [N=7/11]. HM3 WAS USED IN 6 PATIENTS (FS, 12.5%, [N=2/16] VS. LI, 36.4% [N=4/11]. LEFT INTERNAL MAMMARY ARTERY GRAFT WAS PATENT IN 23 PATIENTS (85.2%). NONE OF THE PATIENTS HAD INJURY OF BYPASS GRAFT. CONCOMITANT PROCEDURE WAS PERFORMED IN 5 PATIENTS IN THE FS GROUP (AORTIC VALVE, N=3, AND TRICUSPID VALVE, N=2). IN ISOLATED LVAD IMPLANTATION, CARDIOPULMONARY BYPASS TIME WAS COMPARABLE BETWEEN GROUPS (FS, 75 [68, 96] MIN VS. LI, 67 [58, 89] MIN, P=0.33). THE RATE OF SEVERE RV FAILURE THAT REQUIRED MECHANICAL SUPPORT WAS HIGHER IN THE FS GROUPS (FS: 43.8% [N=7] VS. 0%, P=0.02). 1 PATIENT IN THE FS GROUP (6.3%) AND 2 PATIENTS IN THE LI GROUP (18.2%) HAD A STROKE (P=0.55). 4 PATIENTS IN THE FS GROUP (25%) AND 2 PATIENTS IN THE LI GROUP (18.2%) HAD REOPERATION FOR BLEEDING (P=1.00). PATIENTS WHO UNDERWENT CONCOMITANT PROCEDURES IN THE FS GROUP DID NOT DEVELOP RV FAILURE. HOSPITAL MORTALITY (FS: 31.3% [N=5] VS. LI, 9.1% [N=1], P=0.35), THE RATES OF TRACHEOSTOMY (FS: 62.5% [N=10] VS. LI, 27.3% [N=3], P=0.12) AND HEMODIALYSIS (FS, 37.5% [N=6] VS LI, 9.1% [N=1], P=0.18) WERE NUMERICALLY HIGHER IN GROUP FS. THE ONE-YEAR ESTIMATED SURVIVAL WAS NOT SIGNIFICANTLY DIFFERENT (FS: 68.8±11.6% VS. LI, 81.8±11.6%, P=0.85). IT WAS CONCLUDED THAT LESS INVASIVE REOPERATIVE LVAD INSERTION MAY HAVE A POTENTIAL TO REDUCE RV FAILURE AND IMPROVE OUTCOMES IN PATIENTS WITH HISTORY OF CABG COMPARED TO FS APPROACH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT IT WAS REPORTED IN THE RESEARCH ARTICLE ¿REOPERATIVE LEFT VENTRICULAR ASSIST DEVICE INSERTION FOLLOWING PREVIOUS CORONARY ARTERY BYPASS GRAFTING: LESS INVASIVE VERSUS STERNOTOMY APPROACH¿ THAT THE HEARTMATE 3 MAY BE ASSOCIATED WITH RIGHT HEART FAILURE, AORTIC INSUFFICIENCY, STROKE, BLEEDING, INFECTION, MULTIPLE ORGAN FAILURES, TRANSPLANT, AND DEATH. THE STUDY SHOWED THAT A LESS INVASIVE LVAD INSERTION MAY REDUCE POSTOPERATIVE COMPLICATIONS, INCLUDING RV FAILURE, COMPOSITE OUTCOMES, AND IN-HOSPITAL MORTALITY COMPARED TO A FULL- STERNOTOMY APPROACH IN PATIENTS WITH A HISTORY OF PRIOR CABG. THE LESS INVASIVE APPROACH LIMITS DISSECTION AROUND THE HEART AND CAN MINIMIZE THE RISK OF INJURY TO THE RV AND PATENT BYPASS GRAFTS, WHICH LIKELY CONTRIBUTES TO FAVORABLE SHORT- TERM OUTCOMES. FROM MARCH 2012 TO APRIL 2023, RETROSPECTIVE, SINGLE-CENTER ANALYSIS WAS PERFORMED ON 29 PATIENTS WITH A HISTORY OF CABG WHO UNDERWENT RE-OPERATIVE LVAD IMPLANTATION. OF THESE, 16 PATIENTS HAD A FULL STERNOTOMY (GROUP FS), AND 13 PATIENTS RECEIVED A LESS INVASIVE APPROACH (GROUP LI). SINCE MAY 2019, THE LI APPROACH HAS BEEN UTILIZED FOR REDO SURGICAL SCENARIOS. THE MEDIAN AGE WAS 68 YEARS IN GROUP FS AND 70 YEARS IN GROUP LI. THE MEDIAN DURATION FROM THE LAST CARDIAC SURGERY WAS 118 MONTHS 114 MONTHS IN GROUP FS AND 122 MONTHS IN GROUP LI. HEARTMATE II WAS EXCLUSIVELY USED IN THE FS GROUP, WHILE HEARTWARE WAS PREDOMINANTLY USED IN GROUP LI. HEARTMATE 3 WAS USED IN 8 PATIENTS. THE PRIMARY OUTCOMES WERE IN-HOSPITAL MORTALITY AND THE INCIDENCE OF SEVERE RV FAILURE, DEFINED BY THE INTERNATIONAL SOCIETY FOR HEART AND LUNG TRANSPLANTATION GUIDELINES FOR MECHANICAL CIRCULATORY SUPPORT. RV FAILURE WAS CHARACTERIZED BY HEMODYNAMIC PARAMETERS (CENTRAL VENOUS PRESSURE > 15 MMHG, MIXED VENOUS SATURATION < 55%, CARDIAC INDEX < 2.0 L/MIN/ M2), AND THE NEED FOR INOTROPES OR MECHANICAL CIRCULATORY SUPPORT (MCS) INCLUDING EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) OR RVAD. SECONDARY OUTCOMES INCLUDED RE- EXPLORATION FOR BLEEDING, STROKE, HEMODIALYSIS, TRACHEOSTOMY, GASTROINTESTINAL BLEEDING, DEEP STERNAL WOUND INFECTION, 1- YEAR OVERALL SURVIVAL, AND MODERATE OR GREATER AORTIC INSUFFICIENCY DURING FOLLOW- UP. THE INCIDENCE OF SEVERE RV FAILURE REQUIRING MCS WAS LOWER WITH THE LI APPROACH. IN- HOSPITAL MORTALITY (FS: 31.3% [N = 5] VS. LI: 9.1% [N = 1], P = 0.35), TRACHEOSTOMY RATES (FS: 62.5% [N = 10] VS. LI: 27.3% [N = 3], P = 0.12), AND HEMODIALYSIS (FS: 37.5% [N = 6] VS. LI: 9.1% [N = 1], P = 0.18) WERE NUMERICALLY HIGHER WITH THE FS APPROACH. THE RATE OF RE- EXPLORATION FOR BLEEDING WAS COMPARABLE BETWEEN THE TWO GROUPS. ON THE OTHER HAND, THE RATE OF STROKE WAS NUMERICALLY HIGHER IN THE LI GROUP (FS: 6.3% [N = 1], LI: 15.4% [N = 2], P = 0.55). IN THE KAPLAN¿MEIER SURVIVAL CURVE ANALYSIS, ONE- YEAR ESTIMATED SURVIVAL WAS NOT SIGNIFICANTLY DIFFERENT BETWEEN THE GROUPS (FS: 68.8% ± 11.6% VS. LI: 81.8% ± 11.6%, P = 0.85). DURING THE FOLLOW- UP PERIOD, AORTIC INSUFFICIENCY OF MORE THAN MODERATE SEVERITY WAS NOT OBSERVED IN EITHER GROUP. SIX PATIENTS RECEIVED HEART TRANSPLANTS DURING THE FOLLOW- UP PERIOD. THERE WERE NO INJURIES TO THE OUTFLOW GRAFTS DURING RE- ENTRY. ALL PATIENTS WERE ALIVE AT THE TIME OF FOLLOW- UP. COMPOSITE OUTCOMES CONSISTED OF OPERATIVE MORTALITY, STROKE, MAJOR INFECTION, MAJOR BLEEDING, SEVERE RIGHT- SIDED HEART FAILURE, RENAL DYSFUNCTION, RESPIRATORY FAILURE, ARRHYTHMIA, AND MAJOR DEVICE MALFUNCTION AS PREVIOUSLY DEFINED IN THE ANALYSIS OF THE STS INTERMACS DATABASE BY THE MICHIGAN CONGESTIVE HEART FAILURE INVESTIGATORS. THE INCIDENCE OF COMPOSITE OUTCOMES WAS LESS IN THE LI APPROACH (FS: 75.0% [N = 12], LI: 30.8% [N = 4], P = 0.03). THE CAUSE OF IN-HOSPITAL MORTALITY INCLUDED MULTI- ORGAN FAILURE IN TWO PATIENTS, REFRACTORY RV FAILURE IN ONE, SEPSIS IN ONE, RESPIRATORY FAILURE IN ONE, AND REFRACTORY TRACHEAL BLEEDING IN ONE. LENGTH OF HOSPITAL STAY WAS NOT SIGNIFICANTLY DIFFERENT BETWEEN THE FS AND LI APPROACHES (FS: 41 [24, 67] DAYS, LI: 31 [20, 42] DAYS, P = 0.36). IN CONCLUSION, RE-OPERATIVE LESS- INVASIVE LVAD IMPLANTATION WITH LEFT THORACOTOMY AND UPPER HEMI STERNOTOMY IS A SAFE AND EFFECTIVE OPTION FOR PATIENTS WITH A HISTORY OF PRIOR CABG, POTENTIALLY IMPROVING OUTCOMES IN THIS CHALLENGING POPULATION. Manufacturer Narrative: AUTHOR INFORMATION: KAI, M., SHIMAMURA, J., TAVOLACCI, S. C., ELGAR, G., ICHIKAWA, H., LEVINE, E., SPIELVOGEL, D., & OHIRA, S. (2025). REOPERATIVE LEFT VENTRICULAR ASSIST DEVICE INSERTION FOLLOWING PREVIOUS CORONARY ARTERY BYPASS GRAFTING: LESS INVASIVE VERSUS STERNOTOMY APPROACH. ARTIFICIAL ORGANS. HTTPS://DOI.ORG/10.1111/AOR.14984. DIVISION OF THORACIC AND CARDIAC SURGERY, DEPARTMENT OF SURGERY, BETH ISRAEL DEACONESS MEDICAL CENTER, BOSTON, MASSACHUSETTS, USA ¦ DIVISION OF CARDIOTHORACIC SURGERY, DEPARTMENT OF SURGERY, WESTCHESTER MEDICAL CENTER, VALHALLA, NEW YORK, USA ¦ NEW YORK MEDICAL COLLEGE, VALHALLA, NEW YORK, USA ¦ DEPARTMENT OF CARDIOLOGY, WESTCHESTER MEDICAL CENTER, VALHALLA, NEW YORK, USA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01APR2023 AS DATA WAS COLLECTED BETWEEN MARCH 2012 AND APRIL 2023. AUTHOR INFORMATION: ICHIKAWA, H., ELGAR, G., ISATH, A., LEVINE, E., SHIMAMURA, J., KAI, M., & OHIRA, S. (2024). REOPERATIVE VENTRICULAR ASSIST DEVICE INSERTION FOR PATIENTS WITH HISTORY OF CORONARY ARTERY BYPASS GRAFTING: LESS INVASIVE APPROACH VERSUS FULL STERNOTOMY. ASAIO JOURNAL, 70, 139-140. CARDIOTHORACIC SURGERY, WESTCHESTER MEDICAL CENTER, VALHALLA, NY; NEW YORK MEDICAL COLLEGE, VALHALLA, NY; CARDIOLOGY, WESTCHESTER MEDICAL CENTER, VALHALLA, NY; CARDIAC SURGERY, BETH ISRAEL DEACONESS MEDICAL CENTER, BOSTON, MA. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21435369&pc=DTZ | 1124841-2025-00012 | 2025-01-29 05:00 | Death | TERUMO CARDIOVASCULAR SYSTEMS CORPORATION | 2025-02-21 05:00 | DTZ | CAPIOX® FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | Use of Device Problem | No Clinical Signs, Symptoms or Conditions | K151791 |  | 1 | Event Description: THE USER FACILITY REPORTED TO TERUMO CARDIOVASCULAR THAT DURING CARDIOPULMONARY BYPASS A POSSIBLE OXYGENATOR FAILURE OCCURRED. ADDITIONALLY, THE PROCEDURE WAS FOR A BILATERAL LEFT VENTRICULAR ASSISTANCE DEVICE INSERTION. THE OXYGENATOR PRODUCED PAOS PREDOMINANTLY IN THE LOW TO MID 200S WITH FRACTION OF INSPIRED OXYGEN (FIO2S) GRADUALLY INCREASING FROM 70 TO 100%, FLOWING FROM 4.5-5 L/MIN. BLOOD GASES, HEAD SATS, AND CDI MEASUREMENTS SUPPORTED THAT THE PATIENT WAS ADEQUATELY PERFUSED. THE PATIENT, UNFORTUNATELY, HAD A BAD OUTCOME, AND PASSED SHORTLY AFTER THE PROCEDURE. SURGEON DISCUSSED WITH PERFUSION AND ANESTHESIA, AND IT WAS THE SURGEON'S OPINION THAT THE OXYGEN WAS NOT PROPERLY OXYGENATING THE PATIENT. PERFUSION DID NOT AGREE WITH THIS ASSESSMENT.Event Description: ADDITIONAL INFORMATION FROM THE CLINICAL SPECIALIST STATES, THAT THE REVIEW OF THE PUMP RECORD REVEALED THAT THE PARTIAL PRESSURE OXYGEN (PO2) VALUES WERE IN THE 200MMHG RANGE, DESPITE INCREASING FRACTION OF INSPIRED OXYGEN (FIO2) SETTINGS FROM 70 PERCENT TO 100 PERCENT. PO2 VALUES IN THE 200S ARE NOT AT CRITICALLY SUPPRESSED LEVELS TO HARM ACCEPTABLE OXYGENATION VALUES. ACTIVATED CLOTTING TIME (ACT) VALUES WERE ABOVE 480 SECONDS, IMPLYING ADEQUATE HEPARINIZATION OF THE PATIENT. THE HISTORY OF THIS PATIENT WAS A 45-YEAR-OLD THAT HAD SEVERE CARDIOMYOPATHY AND A HISTORY OF MITRAL VALVE REPAIR VIA MITRAL CLIP CATHETERIZATION IN ATTEMPT TO IMPROVE HEART FUNCTION. THIS WAS A VERY SICK PATIENT THAT PRESENTED TO THE OPERATING ROOM (OR) ON A VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) FOR INSERTION OF LEFT AND RIGHT VENTRICULAR ASSIST DEVICES (VADS) IN ATTEMPT TO STABILIZE THE PATIENT'S CRITICAL CONDITION. THE CLINICIAN STATED THAT THE PATIENT WAS ADEQUATELY OXYGENATED, AND THIS REVIEW OF THE PUMP RECORD AGREES WITH THAT ASSESSMENT. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS SUBMITTED TO FDA IN ACCORD WITH APPLICABLE REGULATIONS. UPON FURTHER INVESTIGATION OF THE REPORTED EVENT, THE FOLLOWING INFORMATION IS NEW AND/OR CHANGED: B5 (ADDED DESCRIBE EVENT OR PROBLEM), B7 (ADDED OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS), G3 (DATE RECEIVED BY MANUFACTURER), G6 (INDICATION THAT THIS IS A FOLLOW-UP REPORT), H2 (FOLLOW-UP DUE TO ADDITIONAL INFORMATION AND DEVICE EVALUATION), H3 (DEVICE EVALUATED BY MANUFACTURER), H6 (IDENTIFICATION OF EVALUATION CODES 10, 3331, 213, 67). THE RETURNED SAMPLE WAS VISUALLY INSPECTED UPON RECEIPT, WITH NO ANOMALIES SUCH AS BREAKAGE. IT WAS THEN RINSED AND DRIED, THE AMOUNT OF OXYGEN TRANSFER AND CARBON DIOXIDE GAS REMOVAL WERE MEASURED ACCORDING TO PRODUCT INSPECTION PROCEDURE. IT WAS CONFIRMED TO MEET FACTORY'S SPECIFICATIONS, WITH NO ANOMALY FOUND. THE MANUFACTURING AND INCOMING INSPECTION RECORDS OF THE ACTUAL SAMPLE WERE REVIEWED, WITH NO ANOMALIES FOUND. THE ROOT CAUSE COULD NOT BE DETERMINED AS IT WAS FOUND TO FUNCTION AS INTENDED. ALL AVAILABLE INFORMATION HAS BEEN PLACED ON FILE IN QUALITY MANAGEMENT FOR APPROPRIATE TRACKING, TRENDING, AND FOLLOW-UP. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS SUBMITTED TO FDA IN ACCORD WITH APPLICABLE REGULATIONS ¿ AND AS INDICATED BY TERUMO CARDIOVASCULAR SYSTEMS IN THE INITIAL REPORT SUBMITTED TO THE FDA ON FEBRUARY 21, 2025. UPON FURTHER INVESTIGATION OF THE REPORTED EVENT, THE FOLLOWING INFORMATION IS NEW AND/OR CHANGED: D9 (DEVICE AVAILABILITY - ADDED DATE RETURNED TO MANUFACTURER), G3 (DATE RECEIVED BY MANUFACTURER), G6 (INDICATION THAT THIS IS A FOLLOW-UP REPORT), H2 (FOLLOW-UP DUE TO ADDITIONAL INFORMATION), H3 (EVALUATION IS IN PROGRESS, BUT NOT YET CONCLUDED). A SECOND FOLLOW-UP WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION AND/OR SUBMISSION OF NEW INFORMATION, THUS TCVS REFERENCES CONCLUSIONS CODE 11. ALL AVAILABLE INFORMATION HAS BEEN PLACED ON FILE IN QUALITY MANAGEMENT FOR APPROPRIATE TRACKING, TRENDING, AND FOLLOW-UP. Manufacturer Narrative: TERUMO HAS NOT RECEIVED THE DEVICE FOR EVALUATION; THEREFORE, THE INVESTIGATION HAS YET TO BE COMPLETED. TERUMO PLANS ON SUBMITTING A FOLLOW-UP REPORT WHEN THE INVESTIGATION IS COMPLETE AND WHEN MORE INFORMATION BECOMES AVAILABLE. FOR THIS REASON, TERUMO REFERENCES EVALUATION CONCLUSION CODE 11.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21436605&pc=NPT | 2025587-2025-01382 | 2024-06-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-21 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Arrhythmia; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS AND 2 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY DUE TO HEART FAILURE AND VENTRICULAR ARRHYTHMIA. PER THE PHYSICIAN, THERE IS NO RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE. PER THE PHYSICIAN, IT IS UNKNOWN WHETHER THERE IS A RELATIONSHIP BETWEEN THE DEATH AND THE PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21436686&pc=MKJ | 1220908-2025-00724 | 2025-02-08 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-21 05:00 | MKJ | AED PRO | Failure to Power Up; Unexpected Shutdown | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 57-YEAR-OLD MALE PATIENT, THE DEVICE INAPPROPRIATELY SHUTDOWN AND THEN WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL CANADA FOR EVALUATION. THE CUSTOMER'S REPORT WAS VERIFIED AND ATTRIBUTED TO THE USER NOT MANAGING THE DEVICE IN ACCORDANCE WITH ZOLL RECOMMENDATIONS. THE RETURNED BATTERY WAS IDENTIFIED TO BE DEPLETED. A REVIEW OF THE DEVICE LOG INDICATED THE DEVICE WAS POWERED OFF WITHOUT ELECTRODES ATTACHED. WHICH MEANS IT'S NOT RESCUE READY AND WILL PRESENT A RED X AND AN AUDIBLE ALERT THAT ATTENTION IS NEEDED. WHILE IN THIS STATE THE DEVICE WILL NOT PERFORM AUTOMATED SELF TESTS TO ENSURE THE DEVICE IS CLINICALLY READY. THE AED PRO OPERATOR'S GUIDE (9650-0350-01 REV M) ADVISES USERS TO PRE-CONNECT UNEXPIRED DEFIBRILLATION ELECTRODES WHILE THE DEVICE IS BEING DEPLOYED SO IT IS READY TO BE USED. THE AED PRO OPERATOR'S GUIDE ALSO ADVISES USERS TO KEEP A FULLY CHARGED SPARE BATTERY PACK WITH THE DEVICE AT ALL TIMES, PERIODICALLY CHECK THE READY INDICATOR TO ENSURE THAT IT DISPLAYS A GREEN CHECK, AND BEFORE/AFTER EACH CLINICAL USE. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21436855&pc=OJX | 2649622-2025-05143 | 2025-01-16 05:00 | Death | MPRI | 2025-02-21 05:00 | OJX | ATTAIN PERFORMA MRI SURESCAN | Insufficient Information | Insufficient Information | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND DIED WITHIN ONE MONTH POST UPGRADE TO A CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P). THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DEATH WAS NOT LINKED TO THE IMPLANT SURGERY OR THE PERFORMANCE OF THE CRT-P SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21437037&pc=BZD | 2518422-2025-102275 | 2025-01-16 05:00 | Death | RESPIRONICS, INC. | 2025-02-21 05:00 | BZD | DREAMSTATION | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21437196&pc=NKM | 2135147-2025-00910 | 2015-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-21 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE UTILITY OF PULMONARY ARTERY (PA) COMPLIANCE IN PREDICTING CARDIAC EVENTS AFTER MITRAL VALVE TRANSCATHETER EDGE TO EDGE REPAIR (M-TEER). DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT PA COMPLIANCE ASSESSED WITH INVASIVE CATHETER MEASUREMENTS IN A CONSCIOUS STATE BEFORE M-TEER IN VENTRICULAR SECONDARY MITRAL REGURGITATION (VSMR) MAY PLAY AN IMPORTANT ROLE IN PREDICTING POST-M-TEER CARDIAC EVENTS.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MAKOTO AMAKI AT NATIONAL CEREBRAL AND CARDIOVASCULAR CENTER, OSAKA, JAPAN WITH CORRESPONDING EMAIL AMAKI@NCVC.GO.JP. THE TIME FRAME OF THE STUDY WAS OCTOBER 2015 TO JANUARY 2023. A TOTAL OF 107 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, HEART FAILURE, ISCHEMIC CARDIOMYOPATHY, CARDIAC RESYNCHRONIZATION THERAPY, ANTICOAGULANT MEDICATION, HEART FAILURE MEDICATION (BETA BLOCKER, ACE-I/ARB, MRA), PRIMARY AND SECONDARY MITRAL REGURGITATION, TRICUSPID REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED CARDIAC EVENTS (HEART FAILURE HOSPITALIZATION, CARDIOVASCULAR DEATH, LEFT VENTRICULAR ASSIST DEVICE (LVAD). Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP PROCEDURE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, HEART FAILURE, ISCHEMIC CARDIOMYOPATHY, CARDIAC RESYNCHRONIZATION THERAPY, ANTICOAGULANT MEDICATION, HEART FAILURE MEDICATION (BETA BLOCKER, ACE-I/ARB, MRA), PRIMARY AND SECONDARY MITRAL REGURGITATION, TRICUSPID REGURGITATION. COMPLICATIONS REPORTED INCLUDED CARDIAC EVENTS (HEART FAILURE HOSPITALIZATION, CARDIOVASCULAR DEATH, LEFT VENTRICULAR ASSIST DEVICE (LVAD); THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE IS ESTIMATED B3 EVENT DATE IS ESTIMATED THE ADDITIONAL PATIENT EFFECTS REPORTED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. LITERATURE ATTACHMENT: "REDUCED PULMONARY ARTERY COMPLIANCE IN A CONSCIOUS STATE PREDICTS CARDIAC EVENT AFTER TRANSCATHETER EDGE-TO-EDGE REPAIR IN PATIENTS WITH VENTRICULAR SECONDARY MITRAL REGURGITATION".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21437247&pc=NKE | 9614453-2025-00740 | 2025-01-16 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-21 05:00 | NKE | PERCEPTA QUAD CRT-P MRI SURESCAN | Insufficient Information | Insufficient Information | P010015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND DIED WITHIN ONE MONTH POST UPGRADE TO A CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P). THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DEATH WAS NOT LINKED TO THE IMPLANT SURGERY OR THE PERFORMANCE OF THE CRT-P SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21437494&pc=NPT | 2015691-2025-01236 | 2025-01-27 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-21 05:00 | NPT | EDWARDS COMMANDER DELIVERY, 29MM | Difficult to Remove | Unspecified Vascular Problem | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS FIELD CLINICAL SPECIALIST, DURING THE TAVR PROCEDURE, THERE WERE COMPLICATIONS, INCLUDING MAJOR BLEEDING AND DIFFICULTY USING THE DEVICE. ADDITIONALLY, THERE WAS A MAJOR VASCULAR COMPLICATION AT THE ACCESS SITE. THIS VASCULAR COMPLICATION LED TO THE PATIENT'S DEATH AND NECESSITATED UNPLANNED SURGERY. PER A VOLUNTARY MEDWATCH, THE COMMANDER DELIVERY SYSTEM COULD NOT BE WITHDRAWN THROUGH THE ESHEATH, AND THE SHEATH LAYERS WERE SEPARATED.Event Description: AS REPORTED BY THE EDWARDS TERRITORY MANAGER, AFTER DEPLOYMENT OF A 29MM SAPIEN 3 VALVE IN THE AORTIC POSITION, DIFFICULTY WAS ENCOUNTERED DURING THE WITHDRAWAL OF THE COMMANDER DELIVERY SYSTEM AND THE 16FR ESHEATH+ AT THE LOWER AORTA, JUST ABOVE THE RCIA. THE DELIVERY SYSTEM WAS WEDGED INTO THE EXPANSION SEAM OPENING AT THE END OF THE ESHEATH. ALTHOUGH THERE APPEARED TO BE COAXIAL ALIGNMENT UPON RE-ENTRY INTO THE DISTAL END OF THE ESHEATH, THERE WAS ANGULATION FROM THE AORTA INTO THE RCIA. THERE WAS NO CROSSING DIFFICULTY AND NO RETAINED VOLUME IN THE BALLOON. SEVERAL MINUTES PASSED BETWEEN DEFLATION AND WITHDRAWAL OF THE SYSTEM. THE VASCULAR COMPLICATION THAT LED TO THE PATIENT'S DEATH WAS DAMAGE TO THE RIGHT COMMON FEMORAL ARTERY, AND AN UNSUCCESSFUL ATTEMPT WAS MADE TO REPAIR THE ARTERY. IT WAS NOTED THAT THE "SHEATH LAYERS WERE SEPARATED," WITH THE INNER LINER AT THE DISTAL TIP APPEARING VERY INVOLVED WITH THE DELIVERY SYSTEM. DURING THE REMOVAL OF THE SYSTEM, AN ATTEMPT WAS MADE TO SEPARATE THE DELIVERY SYSTEM AND ESHEATH BY SLIGHT ROTATION OF THE DELIVERY SYSTEM AND ESHEATH, AND CONFIRMING DEFLATION OF THE BALLOON. NO OTHER DAMAGE WAS NOTED TO ANY OTHER DEVICE AFTER THE WITHDRAWAL OF THE SYSTEM. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, AND INVESTIGATION CONCLUSIONS AND H11 HAVE BEEN UPDATED. THE DEVICE WAS NOT RETURNED AS IT'S IN POSSESSION OF THE HOSPITAL. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. IMAGING WAS PROVIDED AND REVEALED THAT TORTUOSITY WAS PRESENT IN THE ACCESS VESSEL ON THE RIGHT SIDE. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING DIFFICULTY/INABILITY WITHDRAWING CATHETER THROUGH PATIENT VASCULATURE OR THROUGH THE SHEATH HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, NON-COAXIAL WITHDRAWAL, ALTERED BALLOON PROFILE, DAMAGED SHEATH, AND/OR DAMAGED GUIDEWIRE. IN THIS CASE, THE WITHDRAWAL CATHETER THROUGH SHEATH WITH DIFFICULTY TO WITHDRAW SYSTEM WAS UNABLE TO BE CONFIRMED. THE AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (TORTUOSITY) AND/OR PROCEDURAL FACTORS (DAMAGE SHEATH) LIKELY CONTRIBUTED TO THE REPORTED EVENT. PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, SCAR TISSUE, OR SMALL VESSEL SIZE) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING WITHDRAWAL. ADDITIONALLY, IF THE SHEATH WAS PREVIOUSLY DAMAGED, IT CAN CAUSE FURTHER RESISTANCE AS THE DELIVERY SYSTEM IS WITHDRAWN INTO IT. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED BASED ON ADDITIONAL INFORMATION RECEIVED FROM THE EDWARDS TERRITORY MANAGER. AN UPDATE TO B5 WAS MADE TO REFLECT THE ADDITIONAL INFORMATION THAT WAS RECEIVED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21437755&pc=LNS | 3003300201-2025-00002 | 2024-09-07 04:00 | Death | UNITED SERVICE SOLUTIONS, LLC | 2025-02-21 05:00 | LNS | PIEZOLITH 3000 TRIPLE FOCUS LITHOTRIPTER | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Insufficient Information | K072538 |  | 1 | Event Description: ON (B)(6) 2025, WE WERE NOTIFIED THAT THE PATIENT PASSED AWAY AFTER HAVING AN EXTRACORPOREAL SHOCKWAVE LITHOTRIPSY TREATMENT USING THE PIEZOLITH 3000 TRIPLE FOCUS MACHINE. THE PATIENT WAS TREATED ON (B)(6) 2024 AND PASSED AWAY ON (B)(6) 2024. AN INVESTIGATION INTO THIS EVENT IS UNDERWAY. IT IS UNCLEAR WHETHER THE EVENT IS RELATED TO THE DEVICE. A SUPPLEMENTAL MDR WILL BE SUBMITTED AS NEEDED AS THE INVESTIGATION EVOLVES.Event Description: (B)(6) REPORT NOTES THAT THE PATIENT HAD A DIAGNOSIS OF RIGHT RENAL STONES. RIGHT EXTRACORPOREAL SHOCK WAVE LITHOTRIPSY WAS PERFORMED ON (B)(6) 2024. NO COMPLICATIONS DURING PROCEDURE NOTED. THE REPORTING HEALTHCARE PROFESSIONAL DID NOT REPORT ANY DEVICE MALFUNCTIONS OR COMPLICATIONS DURING THE PROCEDURE. THE DAY AFTER THE PROCEDURE INVOLVING THE DEVICE, THE PATIENT WAS DISCHARGED TO HOME WITH HIS DAUGHTER WITH WRITTEN AND VERBAL DISCHARGE INSTRUCTIONS. Manufacturer Narrative: AFTER NOTIFICATION OF THIS INCIDENT, THE MAINTENANCE RECORDS FOR THE MACHINE WERE REVIEWED. ON NOVEMBER 11, 2024, THE MACHINE RECEIVED ITS REGULARLY SCHEDULED PREVENTATIVE MAINTENANCE WHERE NO MECHANICAL ISSUES WERE NOTED. THE PATIENT WAS THE THIRD OF EIGHT TREATMENTS THAT DAY. FIVE TREATMENTS WERE PERFORMED AFTER THE PATIENT'S TREATMENT USING THE SAME MACHINE WITH NO KNOWN COMPLICATIONS. SINCE THE PATIENT'S TREATMENT, THERE HAVE BEEN 188 TREATMENTS PERFORMED ON THIS MACHINE WITH NO KNOWN COMPLICATIONS. AN INVESTIGATION INTO THIS EVENT IS UNDERWAY. IT IS UNCLEAR WHETHER THE EVENT IS RELATED TO THE DEVICE. A SUPPLEMENTAL MDR WILL BE SUBMITTED AS NEEDED AS THE INVESTIGATION EVOLVES. Manufacturer Narrative: THE HEALTHCARE PROFESSIONAL SUBMITTING THE REPORT PROVIDED ADDITIONAL INFORMATION AND COMMUNICATED THAT BASED ON THE USER FACILITY'S ((B)(6)) INVESTIGATION AND INQUIRY INTO THE PATIENT EVENT THERE WAS NO CONCERN WITH THE FUNCTION OR THE USE OF THE DEVICE, AND THERE WAS NO BELIEF THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21438423&pc=DSQ | 2916596-2025-00987 | 2025-02-06 05:00 | Death | THORATEC CORPORATION | 2025-02-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem | Cardiac Arrest; Hemorrhage/Bleeding; Tachycardia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER SIGNIFICANT BLEEDING FROM THE MOUTH AFTER A TOOTH EXTRACTION, THE PATIENT WAS TAKEN TO HOSPITAL AND SUBSEQUENTLY CARDIAC ARRESTED. SYSTEM CONTROLLER REVIEW OF ALARMS SHOWED MULTIPLE LOW FLOW ALARMS, THE LONGEST WAS 25MINS AT 5:16AM ON (B)(6) 2025 (PERI-ARREST TIME). THE REASON FOR THE CARDIAC ARREST WAS VENTRICULAR TACHYCARDIA ARREST. THE ARREST WAS THOUGHT TO BE THERAPY RELATED, AND THE DEVICE OPERATED AS INTENDED. THE PATIENT EXPIRED. THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE DEVICE PARAMETERS WERE WITHIN NORMAL LIMITS. Manufacturer Narrative: SECTION B2 DATE OF DEATH CORRECTION. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS OF BLEEDING AND VENTRICULAR TACHYCARDIA (VT) ARREST, AS WELL AS THE PATIENT'S OUTCOME, COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES COULD NOT CONFIRM THE REPORTED LOW FLOW ALARMS. ALTHOUGH A SPECIFIC CAUSE FOR THE REPORTED ALARMS COULD NOT BE CONCLUSIVELY DETERMINED, THE ACCOUNT INDICATED THAT THE LONGEST ALARM OCCURRED IN THE SETTING OF CARDIAC ARREST. THE SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 11:39:47 THROUGH 12:43:51 ON (B)(6) 2025. EVENTS CAPTURED PRIOR TO THIS TIMEFRAME WERE OVERWRITTEN BY NEWER INCOMING EVENTS, PER DESIGN. NO NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU, REV. D, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. CHAPTER 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING BLEEDING, CARDIAC ARRHYTHMIA, AND DEATH. THIS CHAPTER ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. CHAPTER 4 OF THE IFU, "HEARTMATE TOUCH¿ COMMUNICATION SYSTEM", DESCRIBES PUMP FLOW AND HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. CHAPTER 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THIS CHAPTER ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. CHAPTER 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS ALL SYSTEM ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21438644&pc=JWH | 1038671-2025-01139 |  | Death | EXACTECH, INC. | 2025-02-21 05:00 | JWH | SPECIFIC DEVICE NOT REPORTED | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT'S KNEE WAS REVISED. PATIENT SUFFERED FROM AN INFECTION, WHICH RESULTED IN THE REMOVAL OF HIS LEG IMPLANT AND ULTIMATELY RESULTED IN HIS PASSING. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: THIS COMPLAINT EVENT WAS RECEIVED VIA KROLL CLAIM (CLAIMS RECEIVED AS PART OF EXACTECH'S FINANCIAL RESTRUCTURING PROCESS, MANAGED BY KROLL, LLC). ALL AVAILABLE EVENT AND PRODUCT INFORMATION WAS COLLECTED AT THE TIME OF CLAIM INTAKE; NO FURTHER INFORMATION IS AVAILABLE OR EXPECTED. DUE TO THE LIMITED INFORMATION, THIS COMPLAINT EVENT CANNOT BE FURTHER INVESTIGATED OR CONFIRMED, AND A DEFINITIVE CAUSE CANNOT BE DETERMINED. POTENTIAL PRODUCT, PATIENT, AND/OR USER-RELATED CONTRIBUTING FACTORS CANNOT BE DEFINITIVELY DETERMINED. THERE IS NO INDICATION WITHIN THE AVAILABLE INFORMATION THAT A NEW OR UNEXPECTED FAILURE MODE OR HARM/ADVERSE EVENT TYPE HAS OCCURRED OR THAT THE BENEFIT-RISK ANALYSIS FOR EXACTECH PRODUCTS HAS BEEN ALTERED; THEREFORE, NO FURTHER RISK ANALYSIS WILL BE PERFORMED. HOWEVER, THIS EVENT WILL BE INCLUDED IN COMPLAINT TRENDING PER (B)(4) COMPLAINT TRENDING AND ACTIONS WILL BE TAKEN AS REQUIRED UPON DETECTION OF ANY TREND SIGNALS. NO FURTHER ACTION OR ESCALATION WILL BE TAKEN AT THIS TIME CONCERNING THIS REPORTED EVENT. SHOULD ADDITIONAL, RELEVANT INFORMATION BE RECEIVED CONCERNING THIS EVENT, THE COMPLAINT INVESTIGATION WILL BE RE-OPENED AND ACTIONS TAKEN AS APPROPRIATE. D1: CORRECTED. H6: CORRECTED HEALTH EFFECT AND INVESTIGATION CLINICAL CODES. Manufacturer Narrative: IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21418199&pc=MIH | 9680654-2025-00003 | 2025-01-26 05:00 | Death | WILLIAM A. COOK AUSTRALIA, PTY LTD | 2025-02-20 05:00 | MIH | ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT PROXIMAL BODY | Adverse Event Without Identified Device or Use Problem; Migration | Ruptured Aneurysm | P020018 |  | 1 | Event Description: THE PATIENT PRESENTED WITH UNKNOWN SYMPTOMS TO THE EMERGENCY DEPARTMENT ON LATE (B)(6) 2025. THE PATIENT'S SCAN CONFIRMED THE SEPARATION OF THE SEAL BETWEEN THE ZFEN-P-2-34-122-R AND ZFEN-D-12-28-76-C GRAFTS. WITHIN THE TIME BETWEEN THE SCAN RESULTS AND PLANNING INTERVENTION, THE ANEURYSM RE-PRESSURISED ENOUGH TO BURST. THE PATIENT EXPIRED. Manufacturer Narrative: DATE OF DEATH IS UNKNOWN. DATE OF AWARENESS WAS SELECTED AS AN APPROXIMATE DATE. THE GRAFT REMAINED IN SITU AND THEREFORE WAS NOT RETURNED FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED AND REVIEWED BY CLINICAL PERSONNEL, AND THE FOLLOWING WAS DETERMINED "THIS IS A CASE OF GRAFT COMPONENT SEPARATION CAUSED BY LATERAL BOWING OF THE GRAFT INSIDE THE ANEURYSM OVER TIME (ABOUT 18 MONTHS BETWEEN CT SCANS, FIRST ONE SHOWING OVERLAP OF 2 STENTS, SECOND ONE SHOWING BOWING AND SEPARATION.) THE DEVICE WAS IMPLANTED IN 2021, AND THE FIRST CT WE HAVE IS FROM (B)(6) 2023, AND IT SHOWS THE OVERLAP BETWEEN ZFEN-P AND ZFEN-D TO BE A GOOD 2-STENT OVERLAP. THE RECENT CT SCAN FROM (B)(6) 2025 SHOWS MARKED LATERAL BOWING OF THE DEVICE, WITH THE DISTAL ZFEN-D COMPONENT HAVING PULLED OUT OF THE PROXIMAL ZFEN-P COMPONENT. UNFORTUNATELY, THE PATIENT RUPTURED THE AAA AND DIED. THIS GRAFT SEPARATION IS TYPICAL OF THE ONES WE¿VE SEEN PREVIOUSLY INVOLVING ZFEN DEVICES, WHICH INVOLVES A ¿SLIP FIT¿ BETWEEN THE ZFEN-P AND THE ZFEN-D. OVER TIME, ESPECIALLY IF THERE IS ROOM IN THE AAA SAC, SOME GRAFTS MOVE LATERALLY, CAUSING A BOWING OF THE DEVICE. THE ZFEN-D THEN ¿SLIPS¿ OUT OF THE ZFEN-P AND IS SEEN VERY CLEARLY IN THIS CASE. SUMMARY: ZFEN-P AND ZFEN-D SEPARATION DUE TO BOWING OF THE DEVICE, AND A SLIP FIT". REVIEW OF THE DEVICE HISTORY RECORD FOR LOT NUMBER AC1082823 FOUND THE WORK ORDER APPEARED COMPLETE AND THE QUALITY CONTROL INSPECTION RECORD WAS VERIFIED TO ENSURE THAT THE DEVICE PASSED INSPECTION. THERE WERE NO NON-CONFORMANCES OR RELATED TEMPORARY DEVIATIONS IN PLACE AT THE TIME OF MANUFACTURING. UPON REVIEWING THE PHOTOS TAKEN OF THE COMPLAINT DEVICE DURING MANUFACTURE, IT APPEARS THAT THE DEVICE WAS MANUFACTURED TO SPECIFICATION. THE SAME REVIEW WAS CONDUCTED FOR THE ZFEN-D GRAFT WITH NO ISSUES DETECTED FOR THAT DEVICE EITHER. REVIEW OF THE INSTRUCTIONS FOR USE (IFU) SUPPLIED WITH THE DEVICE FOR GENERAL INFORMATION FOUND IT TO CONTAIN APPROPRIATE WARNINGS, PRECAUTIONS, AND INSTRUCTIONS TO THE USER, INCLUDING: 4. WARNINGS AND PRECAUTIONS. 4.1 GENERAL USE INFORMATION. THE LONG-TERM PERFORMANCE OF FENESTRATED ENDOVASCULAR GRAFTS, INCLUDING THE STENTS PLACED IN FENESTRATIONS/SCALLOPS, HAS NOT YET BEEN ESTABLISHED. ALL PATIENTS SHOULD BE ADVISED THAT ENDOVASCULAR TREATMENT REQUIRES LIFE-LONG, REGULAR FOLLOW-UP TO ASSESS THEIR HEALTH AND THE PERFORMANCE OF THEIR ENDOVASCULAR GRAFT. PATIENTS WITH SPECIFIC CLINICAL FINDINGS (E.G., ENDOLEAKS, ENLARGING ANEURYSMS, CHANGES IN THE STRUCTURE OR POSITION OF THE ENDOVASCULAR GRAFT, OR STENOSIS/OCCLUSION OF VESSELS ACCOMMODATED BY FENESTRATIONS) SHOULD RECEIVE ENHANCED FOLLOW-UP. AFTER ENDOVASCULAR GRAFT PLACEMENT, PATIENTS SHOULD BE REGULARLY MONITORED FOR PERIGRAFT FLOW, ANEURYSM GROWTH, PATENCY OF VESSELS ACCOMMODATED BY A FENESTRATION/SCALLOP, OR CHANGES IN THE STRUCTURE OR POSITION OF THE ENDOVASCULAR GRAFT. AT A MINIMUM, ANNUAL IMAGING IS RECOMMENDED, INCLUDING: 1) ABDOMINAL RADIOGRAPHS TO EXAMINE DEVICE INTEGRITY (SEPARATION BETWEEN COMPONENTS, STENT FRACTURE OR BARB SEPARATION) AND 2) CONTRAST AND NON-CONTRAST CT TO EXAMINE ANEURYSM CHANGES, PERIGRAFT FLOW, PATENCY, TORTUOSITY AND PROGRESSIVE DISEASE. IF RENAL COMPLICATIONS OR OTHER FACTORS PRECLUDE THE USE OF IMAGE CONTRAST MEDIA, ABDOMINAL RADIOGRAPHS AND DUPLEX ULTRASOUND MAY PROVIDE SIMILAR INFORMATION. 4.2 PATIENT SELECTION, TREATMENT AND FOLLOW-UP KEY ANATOMIC ELEMENTS THAT MAY AFFECT SUCCESSFUL EXCLUSION OF THE ANEURYSM INCLUDE SEVERE PROXIMAL NECK ANGULATION (> 45 DEGREES FOR INFRARENAL NECK TO AXIS OF AAA OR > 45 DEGREES FOR SUPRARENAL NECK RELATIVE TO THE IMMEDIATE INFRARENAL NECK); SHORT PROXIMAL AORTIC NECK (<4 MM); GREATER THAN 10% INCREASE IN DIAMETER OVER 15 MM OF PROXIMAL AORTIC NECK LENGTH; AND CIRCUMFERENTIAL THROMBUS AND/OR CALCIFICATION AT THE ARTERIAL IMPLANTATION SITES, SPECIFICALLY THE PROXIMAL AORTIC NECK AND DISTAL ILIAC ARTERY INTERFACE. IRREGULAR CALCIFICATION AND/OR PLAQUE MAY COMPROMISE THE FIXATION AND SEALING OF THE IMPLANTATION SITES. NECKS EXHIBITING THESE KEY ANATOMIC ELEMENTS MAY BE MORE CONDUCIVE TO GRAFT MIGRATION. 5. ADVERSE EVENTS POTENTIAL ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: ENDOPROSTHESIS: IMPROPER COMPONENT PLACEMENT; INCOMPLETE COMPONENT DEPLOYMENT; COMPONENT MIGRATION; SUTURE BREAK; OCCLUSION; INFECTION; STENT FRACTURE; GRAFT MATERIAL WEAR; DILATATION; EROSION; PUNCTURE; PERIGRAFT FLOW; BARB SEPARATION AND CORROSION. 11.4.8 DISTAL BIFURCATED BODY PLACEMENT REPEAT ANGIOGRAM TO VERIFY: - THE DEGREE OF OVERLAP WITH PROXIMAL BODY (NO LESS THAN 2 STENTS) - THE POSITION OF THE CONTRALATERAL LIMB - THE POSITION OF THE IPSILATERAL ILIAC LIMB WITH RESPECT TO THE COMMON ILIAC BIFURCATION." AND "REPOSITION DISTAL BIFURCATED BODY AS REQUIRED." A REVIEW OF THE MANUFACTURING RECORDS AND SPECIFICATIONS DID NOT REVEAL ANY DISCREPANCIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INVESTIGATION CONCLUDED THAT THE COMPLAINT DEVICE WAS MANUFACTURED TO SPECIFICATION. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED FROM THE INVESTIGATION. IT IS POSSIBLE THAT THE FOLLOWING MAY HAVE CONTRIBUTED TO THE OCCURRENCE OF THE REPORTED EVENT: INADEQUATE DIAMETER TOLERANCE TO ENSURE INTERFERENCE AT OVERLAP, INADEQUATE SEPARATION FORCE, PATIENT-RELATED FACTORS. AFTER CONSIDERING THIS EVENT THE RISK ASSOCIATED WITH THE USE OF THIS DEVICE IS STILL DEEMED ADEQUATE. THE APPROPRIATE PERSONNEL WILL CONTINUE TO MONITOR FOR SIMILAR COMPLAINTS. AN INTERNAL ACTION IS NOT DEEMED NECESSARY AT THIS TIME. TRENDING WILL MONITOR IF ANY FUTURE ACTIONS ARE REQUIRED. Manufacturer Narrative: DATE OF DEATH IS UNKNOWN. DATE OF AWARENESS WAS SELECTED AS AN APPROXIMATE DATE .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21418971&pc=OUT | 2032493-2025-00145 | 2025-01-18 05:00 | Death | MICROVENTION, INC. | 2025-02-20 05:00 | OUT | FRED 27 | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Ruptured Aneurysm; Thrombosis/Thrombus | P180027 |  | 1 | Event Description: AS REPORTED TO MICROVENTION: "CEREBRAL INFARCTION, DELAYED SAH AND DEATH: THE FRED WAS IMPLANTED. AFTER THE FRED WAS IMPLANTED, THE PATIENT HAD RECURRENT CEREBRAL INFARCTION, 4 TO 5 TIMES IN THE STENTED AREA OVER THE SIX MONTHS. FINALLY, THE PATIENT DIED OF SEVERE SUBARACHNOID HEMORRHAGE (SAH), WHICH COULD ONLY HAVE BEEN CAUSED BY THE ANEURYSM IN THE SAME AREA. THE PATIENT WAS TREATED FOR THE RECURRENT THROMBOSIS WITH PHARMACOLOGICAL THERAPY ON A CASE-BY-CASE BASIS." "IT IS UNKNOWN WHETHER OR NOT AUTOPSY WAS PERFORMED. THE PRIMARY DISEASE IS BASILAR TRUNK ANEURYSM. THE CAUSE OF DEATH WAS ASSUMED TO BE RUPTURE OF THE TREATED ANEURYSM. THE PATIENT DIED OF SEVERE HEMORRHAGE (SAH) WITH RUPTURE OF THE TREATED ANEURYSM." NO IMAGE AVAILABLE. MEDICAL HISTORY: BASILAR TRUNK ANEURYSM. STENT IMPLANTATION SITE ANEURYSM LOCATION: BA SHAPE: FUSIFORM LESION SITE: CENTER. ANTIPLATELET AND ANTICOAGULANT THERAPY: ADMINISTERED, PREOPERATIVE AND POSTOPERATIVE.Event Description: AS REPORTED TO MICROVENTION: "CEREBRAL INFARCTION, DELAYED SAH AND DEATH: (...) THE FRED WAS IMPLANTED. AFTER THE FRED WAS IMPLANTED, THE PATIENT HAD RECURRENT CEREBRAL INFARCTION (...), 4 TO 5 TIMES IN THE STENTED AREA OVER THE SIX MONTHS. FINALLY, THE PATIENT DIED OF SEVERE SUBARACHNOID HEMORRHAGE (SAH), WHICH COULD ONLY HAVE BEEN CAUSED BY THE ANEURYSM IN THE SAME AREA. THE PATIENT WAS TREATED FOR THE RECURRENT THROMBOSIS WITH PHARMACOLOGICAL THERAPY ON A CASE-BY-CASE BASIS." "IT IS UNKNOWN WHETHER OR NOT AUTOPSY WAS PERFORMED. THE PRIMARY DISEASE IS BASILAR TRUNK ANEURYSM. THE CAUSE OF DEATH WAS ASSUMED TO BE RUPTURE OF THE TREATED ANEURYSM. THE PATIENT DIED OF SEVERE HEMORRHAGE (SAH) WITH RUPTURE OF THE TREATED ANEURYSM." NO IMAGE AVAILABLE. MEDICAL HISTORY: BASILAR TRUNK ANEURYSM. STENT IMPLANTATION SITE. ANEURYSM LOCATION: BA. SHAPE: FUSIFORM. LESION SITE: CENTER. ANTIPLATELET AND ANTICOAGULANT THERAPY: ADMINISTERED, PREOPERATIVE AND POSTOPERATIVE ADDITIONAL INFORMATION AS REPORTED TO MICROVENTION: "THE CAUSE OF DEATH IS AN ANEURYSM, WHICH OCCURRED WITHIN HALF A YEAR AFTER IMPLANT PLACEMENT. PATIENT DIED (...)." ACCORDING TO INDEX PROCEDURE: "AFTER THE FEMORAL ARTERY WAS PUNCTURED, A GUIDING CATHETER AND A DAC [DISTAL ACCESS CATHETER] WERE DELIVERED INTO THE LEFT (VA) RADIAL ARTERY AND A [MICROCATHETER] WAS THEN DELIVERED INTO THE RIGHT PCA (POSTERIOR CEREBRAL ARTERY). THE DISTAL PORTION OF A FUSIFORM ANEURYSM WAS LOCATED NEAR THE DISTAL END OF THE BA (BASILAR ARTERY). THE FRED WAS DEPLOYED FROM THE RIGHT PCA. THE DEPLOYMENT ITSELF WAS SMOOTH, AND THE PROXIMAL END OF THE FRED WAS DEPLOYED INTO THE LEFT VA. AS BOTH THE BA AND VA WERE EXPANDED TO A FUSIFORM SHAPE IN TWO PLACES AND THE ARTERIES WERE UNEVEN DUE TO ATHEROSCLEROSIS, THE [FRED] WAS NOT WELL APPOSED TO THE VESSEL WALL IN SOME AREAS. HOWEVER, THE PHYSICIAN DETERMINED THAT THE APPOSITION OVER ALL WAS GOOD ENOUGH AND COMPLETED THE PROCEDURE." AS REPORTED APPROXIMATELY SIX AND A HALF MONTHS AFTER INDEX PROCEDURE: "[T]HE PATIENT HAD POOR ARTICULATION, NUMBNESS IN THE LEFT HAND, AND WEAKNESS IN THE LOWER EXTREMITIES. THE PATIENT WAS TAKEN TO THE (...) HOSPITAL BY AMBULANCE AND HOSPITALIZED." AS REPORTED SEVEN MONTHS AFTER INDEX PROCEDURE: "[T]HE PATIENT WAS FOUND UNCONSCIOUS WITH DILATED PUPILS. CT SCAN REVEALED SUBARACHNOID HEMORRHAGE. AS RESPIRATORY CONDITION WAS UNSTABLE, THE PATIENT WAS PLACED ON A VENTILATOR. IN THE EARLY MORNING [THE NEXT DAY], THE PATIENT PASSED AWAY. FROM THE FRED IMPLANTATION UNTIL [HOSPITAL] ADMISSION (...), THE PATIENT WAS HOSPITALIZED [6 TIMES FOR CEREBRAL INFARCTION]. DURING THIS PERIOD, ANGIOGRAPHY TAKEN (...) HAD SHOWN THAT THE ANEURYSM ITSELF HAD SHRUNK AND MOST OF THE ANEURYSM WAS OCCLUDED." ACCORDING TO INFORMATION PROVIDED TO MICROVENTION, PATIENT WAS ON DAPT BEFORE AND AFTER DEVICE IMPLANTATION. DAPT STARTED 2 WEEKS BEFORE THE INDEX PROCEDURE. AFTER THE PROCEDURE, DAPT WAS CONTINUED WITH SOME MEDICATION CHANGES. A MONTH AFTER THE INDEX PROCEDURE, MEDICATION WAS CHANGED TO TAPT. THE NEXT MONTH, THE MEDICATION WAS CHANGED DUE TO THE ONSET OF TIA. AFTER THAT, THE MEDICATION WAS STOPPED AND RESTARTED TWICE, AND THE PATIENT PASSED AWAY.Event Description: AS REPORTED TO MICROVENTION: "CEREBRAL INFARCTION, DELAYED SAH AND DEATH: (...) THE FRED WAS IMPLANTED. AFTER THE FRED WAS IMPLANTED, THE PATIENT HAD RECURRENT CEREBRAL INFARCTION (...), 4 TO 5 TIMES IN THE STENTED AREA OVER THE SIX MONTHS. FINALLY, THE PATIENT DIED OF SEVERE SUBARACHNOID HEMORRHAGE (SAH), WHICH COULD ONLY HAVE BEEN CAUSED BY THE ANEURYSM IN THE SAME AREA. THE PATIENT WAS TREATED FOR THE RECURRENT THROMBOSIS WITH PHARMACOLOGICAL THERAPY ON A CASE-BY-CASE BASIS." "IT IS UNKNOWN WHETHER OR NOT AUTOPSY WAS PERFORMED. THE PRIMARY DISEASE IS BASILAR TRUNK ANEURYSM. THE CAUSE OF DEATH WAS ASSUMED TO BE RUPTURE OF THE TREATED ANEURYSM. THE PATIENT DIED OF SEVERE HEMORRHAGE (SAH) WITH RUPTURE OF THE TREATED ANEURYSM." NO IMAGE AVAILABLE. MEDICAL HISTORY: BASILAR TRUNK ANEURYSM. STENT IMPLANTATION SITE. ANEURYSM LOCATION: BA. SHAPE: FUSIFORM. LESION SITE: CENTER. ANTIPLATELET AND ANTICOAGULANT THERAPY: ADMINISTERED, PREOPERATIVE AND POSTOPERATIVE. ADDITIONAL INFORMATION AS REPORTED TO MICROVENTION: "THE CAUSE OF DEATH IS AN ANEURYSM, WHICH OCCURRED WITHIN HALF A YEAR AFTER IMPLANT PLACEMENT. PATIENT DIED (...)." ACCORDING TO INDEX PROCEDURE: "AFTER THE FEMORAL ARTERY WAS PUNCTURED, A GUIDING CATHETER AND A DAC [DISTAL ACCESS CATHETER] WERE DELIVERED INTO THE LEFT (VA) RADIAL ARTERY AND A [MICROCATHETER] WAS THEN DELIVERED INTO THE RIGHT PCA (POSTERIOR CEREBRAL ARTERY). THE DISTAL PORTION OF A FUSIFORM ANEURYSM WAS LOCATED NEAR THE DISTAL END OF THE BA (BASILAR ARTERY). THE FRED WAS DEPLOYED FROM THE RIGHT PCA. THE DEPLOYMENT ITSELF WAS SMOOTH, AND THE PROXIMAL END OF THE FRED WAS DEPLOYED INTO THE LEFT VA. AS BOTH THE BA AND VA WERE EXPANDED TO A FUSIFORM SHAPE IN TWO PLACES AND THE ARTERIES WERE UNEVEN DUE TO ATHEROSCLEROSIS, THE [FRED] WAS NOT WELL APPOSED TO THE VESSEL WALL IN SOME AREAS. HOWEVER, THE PHYSICIAN DETERMINED THAT THE APPOSITION OVER ALL WAS GOOD ENOUGH AND COMPLETED THE PROCEDURE." AS REPORTED APPROXIMATELY SIX AND A HALF MONTHS AFTER INDEX PROCEDURE: "[T]HE PATIENT HAD POOR ARTICULATION, NUMBNESS IN THE LEFT HAND, AND WEAKNESS IN THE LOWER EXTREMITIES. THE PATIENT WAS TAKEN TO THE (...) HOSPITAL BY AMBULANCE AND HOSPITALIZED." AS REPORTED SEVEN MONTHS AFTER INDEX PROCEDURE: "[T]HE PATIENT WAS FOUND UNCONSCIOUS WITH DILATED PUPILS. CT SCAN REVEALED SUBARACHNOID HEMORRHAGE. AS RESPIRATORY CONDITION WAS UNSTABLE, THE PATIENT WAS PLACED ON A VENTILATOR. IN THE EARLY MORNING [THE NEXT DAY], THE PATIENT PASSED AWAY. FROM THE FRED IMPLANTATION UNTIL [HOSPITAL] ADMISSION (...), THE PATIENT WAS HOSPITALIZED [6 TIMES FOR CEREBRAL INFARCTION]. DURING THIS PERIOD, ANGIOGRAPHY TAKEN (...) HAD SHOWN THAT THE ANEURYSM ITSELF HAD SHRUNK AND MOST OF THE ANEURYSM WAS OCCLUDED." ACCORDING TO INFORMATION PROVIDED TO MICROVENTION, PATIENT WAS ON DAPT BEFORE AND AFTER DEVICE IMPLANTATION. DAPT STARTED 2 WEEKS BEFORE THE INDEX PROCEDURE. AFTER THE PROCEDURE, DAPT WAS CONTINUED WITH SOME MEDICATION CHANGES. A MONTH AFTER THE INDEX PROCEDURE, MEDICATION WAS CHANGED TO TAPT. THE NEXT MONTH, THE MEDICATION WAS CHANGED DUE TO THE ONSET OF TIA. AFTER THAT, THE MEDICATION WAS STOPPED AND RESTARTED TWICE, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: ITEMS RETURNED: N/A VISUAL ANALYSIS: A VISUAL INSPECTION OF THE DEVICE CAPTURED IN THIS FILE COULD NOT BE PERFORMED AS A PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION, NOR WERE ANY IMAGES OF THE DEVICE PROVIDED IN PLACE OF A DEVICE RETURN. PROCEDURE AND MEDICAL IMAGING WERE NOT PROVIDED FOR THIS INVESTIGATION. INVESTIGATION FINDINGS: WITHOUT THE RETURN AND PHYSICAL EVALUATION OF THE DEVICE, THE INVESTIGATION CANNOT FURTHER EXAMINE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. WITHOUT IMAGING, THE INVESTIGATION CANNOT FURTHER VERIFY IF THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EVALUATE THE CAUSE OF THE REPORTED EVENT. BASED ON A REVIEW OF THE DEVICE¿S RISK DOCUMENTATION, THE REPORTED EVENT DID NOT INDICATE THERE WERE THE PRESENCE OF ANY POTENTIAL OR NEW MANUFACTURING, DESIGN, QUALITY, OR OTHER SYSTEMIC ISSUES, OR NON-CONFORMANCES. THE COMPLAINT CODE IS MONITORED THROUGH THE TRENDING PROCESS; CORRECTIVE ACTION IS DETERMINED, AS NEEDED, THROUGH THIS PROCESS. INVESTIGATIONS OF HISTORIC COMPLAINT FILES WITH SIMILAR COMPLAINT CATEGORY CODING ARE RECORDED IN THE COMPLAINT HANDLING SYSTEM; WITHOUT THE ABILITY TO PERFORM AN ANALYSIS OF THE DEVICE, THIS INVESTIGATION CANNOT FURTHER EXAMINE ANY POTENTIAL ROOT CAUSES. BATCH REVIEW: A SEARCH FOR NON-CONFORMANCES ASSOCIATED WITH THIS PART/LOT NUMBER COMBINATION DID NOT REVEAL ANY PRODUCTION-RELATED ISSUES RELEVANT TO THE COMPLAINT THAT OCCURRED DURING MANUFACTURING OF THE DEVICE. COMPLAINT SYSTEM REVIEW: BASED ON A REVIEW OF THE LAST 2 YEARS OF COMPLAINT DATA, AND AT THE TIME OF THIS INVESTIGATION, NO SYSTEMIC ISSUES HAVE BEEN IDENTIFIED FOR THIS BATCH NUMBER THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. IFU REVIEW: "POTENTIAL COMPLICATIONS POSSIBLE COMPLICATIONS INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: BLEEDING OR HEMORRHAGE INCLUDING INTRACEREBRAL, RETROPERITONEAL OR OTHER LOCATIONS COMPLICATIONS OF ARTERIAL PUNCTURE INCLUDING PAIN, LOCAL BLEEDING (HEMATOMA) OR INJURY TO THE ARTERY OR ADJACENT NERVES. DEVICE MIGRATION, DISTAL EMBOLIZATION, HEADACHE, INCOMPLETE ANEURYSM OCCLUSION, NEUROLOGIC DEFICITS INCLUDING STROKE AND/OR DEATH, PERFORATION OR DISSECTION OF THE VESSEL(S), PSEUDOANEURYSM FORMATION, RUPTURE OR PERFORATION OF ANEURYSM, TRANSIENT ISCHEMIC ATTACK (TIA) OR ISCHEMIC STROKE, VASOSPASM, VESSEL OCCLUSION, VESSEL STENOSIS OR THROMBOSIS. WARNINGS: SHOULD UNUSUAL RESISTANCE BE FELT AT ANY TIME DURING ACCESS OR REMOVAL, THE INTRODUCER/GUIDE CATHETER/MICROCATHETER AND FRED SYSTEM SHOULD BE REMOVED AS A SINGLE UNIT. APPLYING EXCESSIVE FORCE DURING DELIVERY OR RETRIEVAL OF THE FRED SYSTEM CAN POTENTIALLY RESULT IN LOSS OR DAMAGE TO THE DEVICE AND DELIVERY COMPONENTS. IT IS IMPERATIVE TO USE THE FRED SYSTEM WITH A .027 INCH (0.69 MM) INNER DIAMETER HEADWAY® 27 MICROCATHETER. IF REPEATED FRICTION IS ENCOUNTERED DURING FRED SYSTEM DELIVERY, VERIFY MICROCATHETER IS NOT KINKED OR IN EXTREMELY TORTUOUS ANATOMY. CONFIRM THAT THE MICROCATHETER DOES NOT OVALIZE. CONFIRM THAT THERE IS ADEQUATE STERILE HEPARINIZED FLUSH SOLUTION. DO NOT REPOSITION THE FRED SYSTEM IN THE PARENT VESSEL WITHOUT FULLY RETRIEVING THE DEVICE. THE FRED SYSTEM MUST BE RETRIEVED/RESHEATH INTO THE MICROCATHETER AND RE-DEPLOYED AT THE DESIRED TARGET LOCATION OR REMOVED COMPLETELY FROM THE PATIENT. CAUTIONS: EXERCISE CAUTION WHEN CROSSING THE DEPLOYED/DETACHED FRED SYSTEM WITH ADJUNCTIVE DEVICES SUCH AS GUIDEWIRES, CATHETERS, MICROCATHETERS OR BALLOON CATHETERS TO AVOID DISRUPTING THE DEVICE GEOMETRY AND DEVICE PLACEMENT. DIRECTIONS FOR USE 12. ADVANCE THE DELIVERY WIRE TO TRANSFER THE FRED SYSTEM FROM WITHIN THE INTRODUCER INTO THE MICROCATHETER. WARNING: DO NOT TORQUE THE DELIVERY WIRE WHILE ADVANCING OR RETRACTING THE FRED SYSTEM. 13. CONTINUE ADVANCING THE DELIVERY WIRE INTO THE MICROCATHETER UNTIL THE PROXIMAL TIP OF THE DELIVERY WIRE ENTERS THE INTRODUCER. LOOSEN THE RHV LOCKING RING, REMOVE THE INTRODUCER, AND SET IT ASIDE. WARNING: DO NOT APPLY UNDUE FORCE. IF RESISTANCE IS ENCOUNTERED AT ANY POINT DURING DELIVERY OR MANIPULATION, WITHDRAW THE UNIT AND SELECT A NEW FRED SYSTEM. 15. POSITION THE FRED SYSTEM FOR DEPLOYMENT BY ALIGNING THE FRED SYSTEM IMPLANT DISTAL RADIOPAQUE END MARKERS APPROXIMATELY 7 MM PAST THE ANEURYSM NECK. [FIGURE 7] NOTE: A SLOW, PROPER PUSH/PULL TECHNIQUE, ENCOMPASSING SUFFICIENT DELIVERY WIRE PUSH FORCE, IN ADDITION TO AN OPPOSING MICROCATHETER WITHDRAWAL FORCE, TO REMOVE EXCESS MICROCATHETER SLACK WHILE MAINTAINING THE MICROCATHETER TIP WITHIN THE CENTER OF THE PARENT VESSEL, WILL FACILITATE PROPERLY DEPLOYING THE FRED SYSTEM AT THE PROPER LOCATION, TO ACHIEVE FULL EXPANSION AND GOOD VESSEL APPOSITION. CAUTION: USING A RAPID MICROCATHETER WITHDRAWAL TECHNIQUE TO DEPLOY THE FRED SYSTEM IS NOT RECOMMENDED AND MAY RESULT IN DEVICE ELONGATION OR IMPROPER DEPLOYMENT. BE AWARE OF DELIVERY WIRE TIP POSITION DURING DEPLOYMENT. 16. IF FRED SYSTEM POSITIONING IS NOT SATISFACTORY, THE FRED SYSTEM IMPLANT MAY BE RECAPTURED AND REPOSITIONED IF IT IS NOT FULLY DEPLOYED. THE IMPLANT MAY BE RECAPTURED UNTIL THE POINT WHERE THE DISTAL-MOST WIRE MARKER, COLLOCATED DISTAL TO THE IMPLANT PROXIMAL MARKERS, IS ALIGNED APPROXIMATELY 50% OF LENGTH PROXIMAL TO THE MICROCATHETER DISTAL MARKER BAND. [FIGURE 8] CAUTION: IF RESISTANCE IS FELT WHILE RECAPTURING THE DEVICE, DO NOT CONTINUE TO RECAPTURE. WITHDRAW THE MICROCATHETER SLIGHTLY TO UNSHEATH THE DEVICE (WITHOUT EXCEEDING THE RECAPTURE LIMIT), AND THEN ATTEMPT TO RECAPTURE AGAIN. CAUTION: THE FRED SYSTEM MUST NOT BE RE-DEPLOYED MORE THAN THREE TIMES. CAUTION: THE FRED SYSTEM DELIVERY WIRE SHOULD NOT BE UTILIZED AS A GUIDEWIRE. DO NOT TORQUE THE FRED SYSTEM. A TORQUE DEVICE SHOULD NOT BE USED. 17. IF FRED SYSTEM POSITIONING IS SATISFACTORY, CAREFULLY ADVANCE THE DELIVERY WIRE WHILE RETRACTING THE MICROCATHETER AS NEEDED TO MINIMIZE SLACK, MAINTAINING THE MICROCATHETER AROUND THE CENTER OF THE PARENT VESSEL, TO ALLOW THE IMPLANT TO DEPLOY ACROSS THE NECK OF THE ANEURYSM. ENSURE THE IMPLANT PROXIMAL RADIOPAQUE END MARKERS ARE APPROXIMATELY 7 MM PROXIMAL TO THE ANEURYSM NECK TO ENSURE AN ADEQUATE COVERAGE. NOTE: THE FRED SYSTEM WILL EXPAND AND MAY FORESHORTEN UP TO 60% FROM ITS UNDEPLOYED LENGTH. NOTE: VISUALIZE AND REFER TO IMPLANT RADIOPAQUE END MARKERS TO MAINTAIN ADEQUATE IMPLANT LENGTH OF APPROXIMATELY 7 MM ON EACH SIDE OF THE ANEURYSM NECK/TARGET LOCATION TO ENSURE APPROPRIATE COVERAGE. [FIGURE 7] WARNING: DO NOT DETACH THE FRED SYSTEM IF IT IS NOT PROPERLY POSITIONED IN THE PARENT VESSEL. WARNING: IF APPLICABLE, OBSERVE FRED SYSTEM MARKER POSITION DURING COILING PROCEDURE TO ENSURE THAT THE DEVICE DOES NOT MIGRATE. 19. CAREFULLY INSPECT THE DEPLOYED FRED IMPLANT UNDER FLUOROSCOPY TO CONFIRM THAT IT IS COMPLETELY APPOSED TO THE VESSEL WALL AND NOT KINKED. IF THE IMPLANT IS NOT FULLY APPOSED OR IS KINKED, CONSIDER UTILIZING A SUITABLE MICROGUIDEWIRE AND/OR OCCLUSION BALLOON CATHETER TO FULLY OPEN THE IMPLANT. 20. VERIFY THAT THE IMPLANT REMAINS PATENT AND PROPERLY POSITIONED. NOTE: THE JAILED MICROCATHETER SHOULD BE CAREFULLY REMOVED TO AVOID DISLODGING THE FRED SYSTEM IMPLANT. 21. CAUTION: CAREFULLY WATCH THE FRED SYSTEM IMPLANT DISTAL AND PROXIMAL MARKERS WHEN PASSING THROUGH THE IMPLANTED DEVICE WITH OTHER DEVICES TO AVOID DISPLACING THE IMPLANT." THE PHYSICAL DEVICE WAS NOT AVAILABLE FOR EVALUATION TO INVESTIGATE IF A CONDITION EXISTED THAT WOULD HAVE CONTRIBUTED TO EVENT. SUPPLEMENTAL IMAGING WAS ALSO UNAVAILABLE FOR REVIEW; WITHOUT IMAGING, THE INVESTIGATION CANNOT VERIFY THE EVENT OCCURRED AS DESCRIBED, NOR COULD THE INVESTIGATION FURTHER EXAMINE THE CAUSE OF THE REPORTED EVENT. THIS INFORMATION MAY BE UPDATED IF ADDITIONAL INFORMATION IS PROVIDED AT A LATER DATE. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED, SECTION B (HISTORY) ABOVE HAS BEEN UPDATED. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED, SECTION B ABOVE HAS BEEN UPDATED. MICROVENTION IS SUBMITTING THIS REPORT TO COMPLY WITH FDA REPORTING REGULATIONS UNDER 21 CFR PARTS 4 AND 803. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY MICROVENTION, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. MICROVENTION HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION AND HAS PROVIDED AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MICROVENTION, OR ITS EMPLOYEES THAT THE DEVICE, MICROVENTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. A GOOD FAITH EFFORT WILL BE MADE TO OBTAIN THE APPLICABLE INFORMATION RELEVANT TO THE REPORT. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21419175&pc=MIH | 9680654-2025-00004 | 2025-01-27 05:00 | Death | WILLIAM A. COOK AUSTRALIA, PTY LTD | 2025-02-20 05:00 | MIH | ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT PROXIMAL BODY | Adverse Event Without Identified Device or Use Problem; Material Split, Cut or Torn |  | P020018 |  | 1 | Event Description: PATIENT EXPIRED. A SCAN CONFIRMED A GRAFT SEAL SEPARATION OF THE ZFEN-P-2-34-137-R AND ZFEN-D-12-28-91-C.Event Description: THIS (B)(6) PATIENT EXPIRED. ORIGINALLY REPORTED AS: 'A SCAN CONFIRMED A GRAFT SEAL SEPARATION OF THE ZFEN-P-2-34-137-R AND ZFEN-D-12-28-91-C.' REVIEW OF IMAGING FOUND NO ZFEN-D-12-28-91 DEVICE IMPLANTED. UPDATED DESCRIPTION OF EVENT SINCE IMAGING REVIEW: 'PATIENT EXPIRED. TYPE 3B (TEARS) IN GRAFT MATERIAL.' SEE (B)(6): PATIENT EXPIRED. TYPE 3C ENDOLEAK WAS DISCOVERED ON IMAGING REVIEW FOR THE RELATED (B)(6). Manufacturer Narrative: THE GRAFT REMAINED IN SITU AND THEREFORE WAS NOT RETURNED FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED AND REVIEWED BY CLINICAL PERSONNEL, AND THE FOLLOWING WAS DETERMINED "THIS CASE SUPPOSEDLY INVOLVES A ZFEN (P AND D) THAT WAS IMPLANTED IN 2022, WITH WHAT APPEARS TO BE PREEXISTING ALPHA THORACIC GRAFTS IN THE THORACIC AORTA. THE IMAGING WE HAVE IS FROM 2023 AND 2024, BUT ON BOTH OF THOSE FOLLOW-UP CT SCANS, THERE IS NO SIGN OF A BIFURCATED COMPONENT LIKE A ZFEN-D. BUT THE SITE REPORTS SEPARATION OF THE ZFEN-P AND THE ZFEN-D. I CAN¿T SEE A ZFEN-D THERE AT ALL, AND THE LARGE ENDOLEAKS ARE HIGHER THAN THAT. THE PATIENT APPARENTLY HAD A SIGNIFICANT FALL IN (B)(6) 2024, REQUIRING HOSPITALISATION, AND SUBSEQUENTLY DIED WITH LARGE ENDOLEAKS FROM THE ENDOGRAFT. THE CT SCAN FROM (B)(6) 2024 SHOWS AT LEAST THREE LARGE TYPE 3 ENDOLEAKS: 1. LARGE ENDOLEAK FROM THE REGION OF THE TOP OF THE OVERLAP BETWEEN THE FENESTRATED GRAFT AND THE OUTER ALPHA THORACIC GRAFT. THIS MAY BE A FABRIC TEAR DUE TO THE BARBS OF THE FENESTRATED GRAFT HAVING TORN THE FABRIC OF THE ALPHA GRAFT ¿ JUST MY SPECULATION BUT WOULD EXPLAIN THIS LARGE ENDOLEAK. IT IS HARD TO SEE AND CONFIRM FROM THE CT SCAN THAT THERE WERE INDEED FIXATION BARBS ON THE FENESTRATED GRAFT. 2. SLIGHTLY LESS VOLUME ENDOLEAK FROM THE REGION OF THE 2ND STENT DOWN THE FENESTRATED GRAFT. 3. SIGNIFICANT ENDOLEAK FROM THE SMA FENESTRATION. ENDOLEAK 1 WOULD BE A TYPE 3B (FABRIC TEAR), ENDOLEAK 2 ALSO PROBABLY A TYPE 3B, ENDOLEAK 3 WOULD BE A TYPE 3C. GIVEN THAT I CAN¿T SEE ANY OBVIOUS GRAFT SEPARATION OF COMPONENTS APART FROM THE LOOSENING OF THE SMA STENT IN THE SMA FENESTRATION, I THINK THE MOST LIKELY EXPLANATION FOR THIS UNFORTUNATE CASE IS THAT THE PATIENT¿S FALL MAY HAVE CAUSED DOWNWARD (DISTAL) JOLTING MOVEMENT OF THE FENESTRATED GRAFT THAT IN TURN CAUSED FABRIC TEARS IN THE ALPHA THORACIC GRAFT FABRIC, AND THE TYPE 3B ENDOLEAK(S). SIMILARLY, THE SMA STENT MAY HAVE BEEN JOLTED LOOSE DUE TO THE MOVEMENT OF THE MAIN FENESTRATED GRAFT BODY, LEADING TO THE TYPE 3C ENDOLEAK FROM THE SMA FENESTRATION. I CAN¿T SEE ANY CONVINCING EVIDENCE OF GRAFT COMPONENT SEPARATION APART FROM THE LOOSE SMA STENT, AND I CAN¿T SEE ANY EVIDENCE OF A BIFURCATED GRAFT COMPONENT EG ZFEN-D. SUMMARY: LIKELY GRAFT FABRIC TEARS DUE TO PATIENT FALL, LEADING TO MULTIPLE LARGE ENDOLEAKS, AND PATIENT DEATH. NOT ANY INTRINSIC FAULT OR FAILURE OF ANY OF THE DEVICES". REVIEW OF THE DEVICE HISTORY RECORD FOUND THE WORK ORDER FOR LOT AC1100373 APPEARED COMPLETE AND THE QUALITY CONTROL INSPECTION WAS VERIFIED TO ENSURE THAT THE DEVICE PASSED INSPECTION. THERE WERE NO TEMPORARY DEVIATIONS OR NON-CONFORMANCES IN PLACE AT THE TIME OF MANUFACTURE. REVIEW OF SPECIFICATIONS FOUND THAT UPON REVIEWING THE PHOTOS TAKEN OF THE COMPLAINT DEVICE DURING MANUFACTURE IT APPEARS THAT THE DEVICE WAS MANUFACTURED AS PER SPECIFICATIONS. REVIEW OF INSTRUCTIONS FOR USE (IFU) AND LABELLING FOUND IT TO CONTAIN APPROPRIATE WARNINGS, PRECAUTIONS, AND INSTRUCTIONS TO THE USER. THERE IS NO EVIDENCE TO SUGGEST THAT THE USER DID NOT FOLLOW THE INSTRUCTIONS FOR USE. A REVIEW OF THE MANUFACTURING RECORDS AND SPECIFICATIONS DID NOT REVEAL ANY DISCREPANCIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INVESTIGATION CONCLUDED THAT THE COMPLAINT DEVICE WAS MANUFACTURED TO SPECIFICATION. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS NOT ABLE TO BE DETERMINED, BASED ON THE MEDICAL DIRECTOR'S ASSESSMENT, THE MOST LIKELY ROOT CAUSE OF THE GRAFT FABRIC TEARS (ENDOLEAKS 1 AND 2) IS DUE TO PATIENT FALL, WHICH LEAD TO MULTIPLE LARGE ENDOLEAKS, AND PATIENT DEATH. AFTER CONSIDERING THIS EVENT THE RISK ASSOCIATED WITH THE USE OF THIS DEVICE IS STILL DEEMED ADEQUATE. THE APPROPRIATE PERSONNEL WILL CONTINUE TO MONITOR FOR SIMILAR COMPLAINTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21419798&pc=DTC | 2182208-2025-00905 | 2025-01-31 05:00 | Death | MEDTRONIC, INC. | 2025-02-20 05:00 | DTC | CARELINK SMARTSYNC COMMON APP IOS | Display or Visual Feedback Problem | Cardiac Arrest | P890003 |  | 1 | Event Description: IT WAS REPORTED THAT THE MOBILE PROGRAMMER APPLICATION PROMPTED THAT AN UPDATE WAS REQUIRED WHEN ATTEMPTING TO INTERROGATE A PATIENT THAT WAS "CODING". THE UPDATE WAS INITIATED, HOWEVER THE UPDATE WAS TAKING LONGER THAN EXPECTED RESULTING IN THE NECESSITY TO SWAP OUT THE MOBILE PROGRAMMER APPLICATION. THE INTERROGATION WAS THEN SUCCESSFULLY COMPLETED WITH THE REPLACEMENT MOBILE PROGRAMMER APPLICATION, HOWEVER THE PATIENT SUBSEQUENTLY PASSED AWAY. IT WAS REPORTED THAT THE PATIENT DEATH WAS NOT DUE TO THE APPLICATION UPDATE. IT WAS NOTED THAT DISSATISFACTION WAS EXPRESSED. Manufacturer Narrative: CORRECTION FDC CODE H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS: PERFORMANCE DATA COLLECTED FROM THE MOBILE PROGRAMMER WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DATA/DATABASE FOUND THE CUSTOMER COMMENT THAT THE MOBILE PROGRAMMER APPLICATION PROMPTED THAT AN UPDATE WAS REQUIRED WAS CONFIRMED. ANALYSIS OF THE DATA/DATABASE FOUND THE CUSTOMER COMMENT THAT THE UPDATE TOOK LONGER THAN EXPECTED COULD NOT BE CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21419804&pc=PNJ | 9612164-2025-00912 | 2025-01-28 05:00 | Death | MEDTRONIC IRELAND | 2025-02-20 05:00 | PNJ | MICRA¿ VR2 | Positioning Problem; Appropriate Term/Code Not Available | Pulmonary Embolism | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), INITIAL DEPLOYMENT WAS UNSATISFACTORY, AND THE DEVICE DISLODGED FOLLOWING TUG TEST. THOUGH THE DEVICE WAS RECAPTURED SUCCESSFULLY, PATIENT¿S BLOOD PRESSURE DETERIORATED DURING OPERATION, WITH OXYGEN LEVELS DROPPING BELOW NORMAL LEVELS. THE LEADLESS IPG WAS ATTEMPTED, NOT USED. RESUSCITATIVE MEASURES AND INTUBATION WERE GIVEN. EVENTUALLY THE PATIENT PASSED AWAY. PULMONARY EMBOLISM WAS SUSPECTED TO BE THE CAUSE OF DEATH. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21420142&pc=PNJ | 9612164-2025-00921 | 2025-01-14 05:00 | Death | MEDTRONIC IRELAND | 2025-02-20 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY AFTER THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED CARDIAC PERFORATION AND PERICARDIAL EFFUSION. MEDICATION WAS ADMINISTERED. THE NEXT DAY FOLLOWING IMPLANT, THE PATIENT PASSED AWAY DURING A CARDIAC SURGERY. Manufacturer Narrative: CORRECTED: H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21420457&pc=MBB | 1818910-2025-02811 | 2025-01-24 05:00 | Death | DEPUY CMW - 9610921 | 2025-02-20 05:00 | MBB | DEPUY CMW 3G GENTAMICIN 40G | Adverse Event Without Identified Device or Use Problem | Loss of consciousness; Respiratory Failure; Asystole; Cardiovascular Insufficiency |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A SURGICAL PROCEDURE FOR IMPLANTATION OF HIP ENDOPROSTHESIS UTILIZING BONE CEMENT, THE PATIENT SUDDENLY BEGAN EXPERIENCING SYMPTOMS OF CARDIAC DISTRESS, ARREST AND ASYSTOLE. THE CONDITION AT THE TIME OF NOTIFICATION OF AN ADVERSE REACTION WAS SEVERE, WITHOUT DYNAMICS, UNSTABLE. MECHANICAL VENTILATION WAS CARRIED OUT. THE SEVERITY WAS DUE TO CARDIOVASCULAR, CEREBRAL, RESPIRATORY FAILURE. THE PATIENT DIED. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY ORTHOPAEDICS IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY ORTHOPAEDICS HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY ORTHOPAEDICS OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY ORTHOPAEDICS, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. H11 ADDITIONAL NARRATIVE: D4/G4: DEVICE IS NOT DISTRIBUTED IN THE UNITED STATES, BUT IS SIMILAR TO DEVICE MARKETED IN THE USA. THEREFORE, (01)GTIN IS NOT AVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). INVESTIGATION SUMMARY: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 3335040, LOT NUMBER: 4071992, AND NO NON-CONFORMANCES OR MANUFACTURING IRREGULARITIES WERE IDENTIFIED. ALL QC AND MICROBIOLOGY TESTING MET SPECIFICATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. DEVICE HISTORY LOT: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 3335040 LOT NUMBER: 4071992, AND NO NON-CONFORMANCES OR MANUFACTURING IRREGULARITIES WERE IDENTIFIED. DEVICE HISTORY REVIEW: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 3335040 LOT NUMBER: 4071992, AND NO NON-CONFORMANCES OR MANUFACTURING IRREGULARITIES WERE IDENTIFIED. H11 ADDITIONAL NARRATIVE: CORRECTED: D3, G1 (MANUFACTURING SITE NAME). Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). INVESTIGATION SUMMARY: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 3335040 LOT NUMBER: 4071992, AND NO NON-CONFORMANCES OR MANUFACTURING IRREGULARITIES WERE IDENTIFIED. ALL QC AND MICROBIOLOGY TESTING MET SPECIFICATION. THERE ARE KNOWN ADVERSE EVENTS ASSOCIATED WITH THE USE OF BONE CEMENTS INCLUDING SERIOUS ADVERSE EVENTS CONCERNING POTENTIAL BLOOD PRESSURE AND CARDIOVASCULAR ISSUES THAT COULD RESULT IN FATAL OUTCOMES DURING THE CEMENTATION AND IMPLANTATION PROCESS. OUR INSTRUCTION FOR USE (IFU) PROVIDES AWARENESS OF THESE ADVERSE EVENTS AND COMPLICATIONS. A COMPLAINT SEARCH FOR GENTAMICIN BONE CEMENT WAS PERFORMED AND NO PREVIOUS COMPLAINTS WERE FOUND RELATED TO A FATAL OUTCOME. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL MONITORING FOR ANY POTENTIAL SAFETY SIGNALS IS ALSO CONDUCTED THROUGH COMPLAINT TRENDING AND OTHER POST-MARKET SAFETY SURVEILLANCE ACTIVITIES. BASED ON THE INVESTIGATION, NO CORRECTIVE AND/OR PREVENTIVE ACTIONS WERE IDENTIFIED. DEVICE HISTORY LOT: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 3335040 LOT NUMBER: 4071992, AND NO NON-CONFORMANCES OR MANUFACTURING IRREGULARITIES WERE IDENTIFIED. DEVICE HISTORY BATCH: NULL DEVICE HISTORY REVIEW: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 3335040 LOT NUMBER: 4071992, AND NO NON-CONFORMANCES OR MANUFACTURING IRREGULARITIES WERE IDENTIFIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21421236&pc=BZD | 2518422-2025-102188 | 2024-11-08 05:00 | Death | RESPIRONICS, INC. | 2025-02-20 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A PHILIPS CPAP DEVICE. THE PATIENT'S WIFE REPORTED THAT HER HUSBAND HAD PASSED AWAY AND WAS ADVISED BY HER ATTORNEY TO OBTAIN THE SERIAL NUMBER FOR THE UNIT. SHE WAS INFORMED THAT THE SERIAL NUMBER IS LOCATED ON THE DEVICE ITSELF. HOWEVER, THE WIFE INDICATED THAT SHE NO LONGER POSSESSES THE UNIT AND WAS INFORMED THAT NO REGISTRATION RECORDS FOR HER HUSBAND COULD BE FOUND. NO DEVICE-RELATED INFORMATION WAS PROVIDED. FURTHERMORE, NO ADDITIONAL DETAILS WERE SHARED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21421237&pc=NPT | 2025587-2025-01344 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-20 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 14 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATENT WAS HOSPITALIZED FOR CONGESTIVE HEART FAILURE (CHF) AND MULTIPLE ORGAN FAILURE. 7 DAYS LATER, THE PATIENT TESTED POSITIVE FOR INFLUENZA. THE PATIENT'S GENERAL CONDITION WORSENED, AND THE PATIENT DIED. THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE AND HEART FAILURE.Event Description: IT WAS REPORTED THAT 14 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATENT WAS HOSPITALIZED FOR CONGESTIVE HEART FAILURE (CHF) AND MULTIPLE ORGAN FAILURE. 7 DAYS LATER, THE PATIENT TESTED POSITIVE FOR INFLUENZA. THE PATIENT'S GENERAL CONDITION WORSENED, AND THE PATIENT DIED. THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE AND HEART FAILURE. Manufacturer Narrative: UPDATED DATA: D. SUSPECT MEDICAL DEVICE EXPIRATION DATE, SERIAL #, UDI # H4. DEVICE MFG DATE MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21421530&pc=LPB | 2029046-2025-00485 | 2024-12-20 05:00 | Death | BIOSENSE WEBSTER INC | 2025-02-20 05:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation | P030031 |  | 1 | Event Description: DURING A CLINICAL TRIAL, IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND DEATH OCCURRED. PATIENT RECEIVED AN INDEX CARDIAC ABLATION FOR NONISCHEMIC VENTRICULAR TACHYCARDIA ON (B)(6) 2024. ON (B)(6) 2024, PATIENT EXPERIENCED VENTRICULAR FIBRILLATION CATEGORIZED AS SEVERE AND SERIOUS AS DEFINED BY INPATIENT/PROLONGED HOSPITALIZATION WITH AN ADMISSION DATE OF (B)(6) 2024 AND A DISCHARGE DATE OF (B)(6) 2024. RELATIONSHIP TO STUDY DEVICE IS NOT RELATED AND RELATIONSHIP TO THE INDEX STUDY PROCEDURE IS NOT RELATED. THE OUTCOME IS RECOVERED/RESOLVED WITH AN END DATE OF (B)(6) 2024. INTERVENTION WAS UNKNOWN. ON AN UNKNOWN DATE, PATIENT EXPERIENCED VENTRICULAR ARRHYTHMIAS CATEGORIZED AS SEVERE AND SERIOUS DEFINED BY DEATH. RELATIONSHIP TO STUDY DEVICE IS UNKNOWN AND RELATIONSHIP TO THE INDEX STUDY PROCEDURE IS UNKNOWN. THE OUTCOME IS FATAL. INTERVENTION WAS UNKNOWN. DATE OF DEATH WAS (B)(6) 2025. TYPE OF MORTALITY IS UNKNOWN. THE SUMMARY OF EVENTS LEADING TO SUBJECT'S DEATH WAS THAT THE PATIENT HAD 2X ABLATION PROCEDURE AND INSERTION OF ICD, RECURRENT ARRHYTHMIAS AND ICD SHOCKS. ON DAY OF DEATH, PATIENT HAD VA (VENTRICULAR ARRHYTHMIA), 5 ICD SHOCKS, LOSS OF CONSCIOUSNESS, "EMAS", CPR (CARDIOPULMONARY RESUSCITATION). THERE WAS NO REGAIN OF CONSCIOUSNESS WHICH LEAD TO DEATH. THIS CLINICAL STUDY EVENT WAS ORIGINALLY CONSIDERED NON-REPORTABLE, HOWEVER, BWI BECAME AWARE ON 31-JAN-2025 OF THE PATIENT¿S DEATH AND HAVE REASSESSED THE EVENT AS REPORTABLE. THE REPORTABLE AWARENESS DATE FOR THIS RECORD IS 31-JAN-2025. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 31301718L AND NO INTERNAL ACTIONS RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: INITIALLY IT WAS REPORTED IN THE 3500A INITIAL, ¿ON AN UNKNOWN DATE, PATIENT EXPERIENCED VENTRICULAR ARRHYTHMIAS CATEGORIZED AS SEVERE AND SERIOUS DEFINED BY DEATH. RELATIONSHIP TO STUDY DEVICE IS UNKNOWN AND RELATIONSHIP TO THE INDEX STUDY PROCEDURE IS UNKNOWN. THE OUTCOME IS FATAL. INTERVENTION WAS UNKNOWN.¿ ADDITIONAL INFORMATION WAS RECEIVED ON 25-FEB-2025, WHICH UPDATED THIS INFORMATION TO THE FOLLOWING: ¿ON (B)(6) 2025, PATIENT EXPERIENCED DEATH CATEGORIZED AS SEVERE AND SERIOUS DEFINED BY DEATH. RELATIONSHIP TO STUDY DEVICE IS NOT RELATED AND RELATIONSHIP TO THE INDEX STUDY PROCEDURE IS NOT RELATED. THE OUTCOME IS FATAL WITH AN END DATE ON (B)(6) 2025. INTERVENTION WAS NONE.¿ IN ADDITION, DURING AN INTERNAL REVIEW ON 18-MAR-2025, NOTED A CORRECTION TO THE 3500A INITIAL. FOR THE VENTRICULAR FIBRILLATION ADVERSE EVENT IT WAS REPORTED, ¿THE OUTCOME IS RECOVERED/RESOLVED WITH AN END DATE OF (B)(6) 2024.¿ IT SHOULD HAVE STATED ¿. THE OUTCOME IS RECOVERED/RESOLVED WITH SEQUELAE AND AN END DATE OF (B)(6) 2024. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: DURING INTERNAL REVIEW ON 18-MAR-2025, NOTED CORRECTION TO THE 3500A FOLLOW-UP #1 UNDER "G3. DATE RECEIVED BY MANUFACTURER" AS PROCESSED AS 23-FEB-2025 AND SHOULD HAVE BEEN PROCESSED AS 25-FEB-2025. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON 07-APR-2025 WHICH INDICATED THAT THE VENTRICULAR FIBRILLATION ADVERSE EVENT WAS INACTIVATED. INITIALLY IT WAS REPORTED UNDER THE 3500A FOLLOW-UP #1, "ON (B)(6) 2025, PATIENT EXPERIENCED DEATH CATEGORIZED AS SEVERE AND SERIOUS DEFINED BY DEATH." HOWEVER, ADDITIONAL INFORMATION WAS RECEIVED ON 15-APR-2025 WHICH UPDATED IT TO, "ON (B)(6) 2025, PATIENT EXPERIENCED CARDIAC ARREST CATEGORIZED AS SEVERE AND SERIOUS DEFINED BY DEATH." IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 17-JUL-2025 WHICH PROVIDED ADDITIONAL CONCOMITANT DEVICES OF AN ¿OCTARAY¿ CATHETER, CARTO 3 SYSTEM AND AN NGEN GENERATOR. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21421573&pc=NPT | 2025587-2025-01347 | 2024-09-08 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-20 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY DUE TO HEART FAILURE. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21421897&pc=PNJ | 9612164-2025-00929 | 2024-12-01 05:00 | Death | MEDTRONIC IRELAND | 2025-02-20 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Anemia; Hemorrhage/Bleeding; Cardiac Tamponade; Cardiogenic Shock; Pericardial Effusion; Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE IMPACT OF BODY MASS INDEX ON IN-HOSPITAL OUTCOMES IN PATIENTS RECEIVING LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS IN-HOSPITAL MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED VENOUS THROMBOEMBOLISMS, CARDIOGENIC SHOCK, BLEEDING COMPLICATIONS WHICH REQUIRED BLOOD TRANSFUSIONS, PERICARDIAL COMPLICATIONS, RENAL COMPLICATIONS, AND DEVICES WHICH REQUIRED REMOVALS OR REVISIONS FOR UNKNOWN REASONS. PERICARDIAL COMPLICATIONS INCLUDED NONINFECTIVE ACUTE PERICARDITIS, NONINFLAMMATORY PERICARDIAL EFFUSION, NONTRAUMATIC HEMOPERICARDIUM, CARDIAC TAMPONADE, AND UNSPECIFIED PERICARDIAL COMPLICATIONS. RENAL COMPLICATIONS INCLUDED ACUTE KIDNEY INJURY. BLEEDING COMPLICATIONS COVERED PERIPROCEDURAL BLEEDING, POSTPROCEDURAL ANEMIA, AND HEMOPERITONEUM/RETROPERITONEAL BLEEDING. THE STATUS OF THE LEADLESS IPGS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/79 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPACT OF BODY MASS INDEX ON IN-HOSPITAL OUTCOMES IN PATIENTS RECEIVING LEADLESS PACEMAKERS: A FIVE-CATEGORY ANALYSIS. HEART RHYTHM O2. 2024; 5:883¿889. DOI: 10.1016/J.HROO.2024.09.017 THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: MICRA PRODUCT ID: MDT-TPS-DELSYS SERIAL: UNKNOWN D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21422529&pc=LOX | 2124215-2025-10049 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-20 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21422587&pc=OZD | 1220648-2025-26226 | 2025-01-27 05:00 | Death | ABIOMED, INC. | 2025-02-20 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT'S LABS HEMOLYZED DURING IMPELLA CP SUPPORT. THE PATIENT WAS GIVEN BLOOD PRODUCTS TO TREAT THE CONDITION, BUT WAS LATER REPORTED TO HAVE EXPIRED. IT IS UNKNOWN IF THE REPORTED CONDITION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF HEMOLYSIS WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. H.1 REVISED TYPE OF REPORTABLE EVENT AS IT WAS PREVIOUSLY SELECTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26226. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423020&pc=NPT | 2025587-2025-01353 | 2025-02-13 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-20 05:00 | NPT | EVOLUT PRO PLUS VALVE | Malposition of Device; Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION.Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED. FOLLOWING IMPLANT OF THE VALVE, THE VALVE FRAME APPEARED UNDER EXPANDED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED. THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION.Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED. FOLLOWING IMPLANT OF THE VALVE, THE VALVE FRAME APPEARED UNDER EXPANDED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED. THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID PACED DURING POST-IMPLANT DILATION TO SLOW 120 BECAUSE OF NOT ENOUGH OUTPUT. A HIGH IMPLANT CONTRIBUTED TO THE DISLODGEMENT. PER THE PHYSICIAN, THE DISLODGEMENT OCCURRED DURING THE BALLOON DILATION.Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED. FOLLOWING IMPLANT OF THE VALVE, THE VALVE FRAME APPEARED UNDER EXPANDED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED. THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID PACED DURING POST-IMPLANT DILATION TO SLOW 120 BECAUSE OF NOT ENOUGH OUTPUT. A HIGH IMPLANT CONTRIBUTED TO THE DISLODGEMENT. PER THE PHYSICIAN, THE DISLODGEMENT OCCURRED DURING THE BALLOON DILATION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE VALVE WAS NOT INTENTIONALLY IMPLANTED IN A HIGH POSITION. RATHER, WHEN THE FIRST DELIVERY CATHETER SYSTEM WAS USED, THE USER INADVERTENTLY IMPLANTED THE VALVE AT THE HIGH IMPLANT DEPTH. THIS VALVE WAS ASSOCIATED WITH THE UNDER EXPANSION OF THE VALVE FRAME, DISLODGEMENT, AND OCCLUSION OF THE CORONARY ARTERIES. Manufacturer Narrative: UPDATED DATA: B5. A FIFTH PARAGRAPH WAS ADDED TO INDICATE THE VALVE NOTED IN SECTION D OF THIS REPORT WAS THE MAIN CONTRIBUTOR TO THE ADVERSE EVENTS. WHEN USING THE FIRST DCS, THE USER INADVERTENTLY DEPLOYED THE VALVE TO AN INACCURATE IMPLANT DEPTH WHICH LED TO THE VALVE DISLODGEMENT AND THE CASCADE OF SUBSEQUENT ADVERSE EVENTS. H6. CODES RELATED TO THE SYSTEM REPORTED DCS WERE ADDED. H10. THE REGULATORY REPORT RELATES TO THE SECOND VALVE IMPLANTED WHICH CANNOT BE EXCLUDED AS A POTENTIAL CONTRIBUTOR TO THE PATIENT'S DEATH. ADDITIONAL CODES. AN ANNEX G CODE WAS ADDED RELATED TO THE SYSTEM REPORTED DCS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE EXPIRATION DATE, SERIAL#, AND UDI# D10. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID X2: D-EVPROP34 (LOT: 0012514794); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID X2: L-EVPROP34 (LOT: 0012497770); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: EVPROPLUS-34 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. H4. DEVICE MFG DATE H6. DEVICE CODE WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.4, B.5, H.6 AND IMF CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: D-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: D-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: L-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: L-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: EVPROPLUS-34 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423054&pc=NPT | 2025587-2025-01354 | 2025-02-13 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-20 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION.Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED. FOLLOWING IMPLANT OF THE VALVE, THE VALVE FRAME APPEARED UNDER EXPANDED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED. THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION.Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED. FOLLOWING IMPLANT OF THE VALVE, THE VALVE FRAME APPEARED UNDER EXPANDED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED. THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID PACED DURING POST-IMPLANT DILATION TO SLOW 120 BECAUSE OF NOT ENOUGH OUTPUT. A HIGH IMPLANT CONTRIBUTED TO THE DISLODGEMENT. PER THE PHYSICIAN, THE DISLODGEMENT OCCURRED DURING THE BALLOON DILATION.Event Description: IT WAS REPORTED A PATIENT UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, FOR AN UNKNOWN REASON, IMPLANT WAS ATTEMPTED WITH TWO DIFFERENT 34MM MEDTRONIC VALVES (SERIAL NUMBERS NOT PROVIDED). HOWEVER, A CORONARY OCCLUSION OCCURRED AND THE PATIENT DIED DURING THE PROCEDURE. IT WAS REPORTED ONE OF THE VALVES OCCLUDED THE CORONARY ARTERIES. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE FIRST VALVE DISLODGED AND REQUIRED A SECOND VALVE TO BE IMPLANTED. THE DISLODGED VALVE CAUSED THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED. FOLLOWING IMPLANT OF THE VALVE, THE VALVE FRAME APPEARED UNDER EXPANDED AND A POST-IMPLANT BALLOON DILATION WAS PERFORMED. THE CAUSE OF DEATH WAS DUE TO THE CORONARY OCCLUSION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID PACED DURING POST-IMPLANT DILATION TO SLOW 120 BECAUSE OF NOT ENOUGH OUTPUT. A HIGH IMPLANT CONTRIBUTED TO THE DISLODGEMENT. PER THE PHYSICIAN, THE DISLODGEMENT OCCURRED DURING THE BALLOON DILATION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE VALVE WAS NOT INTENTIONALLY IMPLANTED IN A HIGH POSITION. RATHER, WHEN THE FIRST DELIVERY CATHETER SYSTEM WAS USED, THE USER INADVERTENTLY IMPLANTED THE VALVE AT THE HIGH IMPLANT DEPTH. THIS VALVE WAS ASSOCIATED WITH THE UNDER EXPANSION OF THE VALVE FRAME, DISLODGEMENT, AND OCCLUSION OF THE CORONARY ARTERIES. Manufacturer Narrative: UPDATED DATA: B5. A FIFTH PARAGRAPH WAS ADDED TO INDICATE THE VALVE NOTED IN SECTION D OF THIS REPORT WAS NOT THE MAIN CONTRIBUTOR TO THE ADVERSE EVENT. WHEN USING THE FIRST DCS, THE USER INADVERTENTLY DEPLOYED THE VALVE TO AN INACCURATE IMPLANT DEPTH WHICH LED TO THE VALVE DISLODGEMENT AND THE CASCADE OF SUBSEQUENT ADVERSE EVENTS. HOWEVER, THE VALVE NOTED IN SECTION D OF THIS REPORT CANNOT BE EXCLUDED AS A POTENTIAL CONTRIBUTOR TO THE PATIENT'S DEATH. H10. THE REGULATORY REPORT RELATES TO THE SYSTEM REPORTED DCS AND THE FIRST VALVE IMPLANTED AS THE MAIN CONTRIBUTORS TO THE ADVERSE EVENTS AND PATIENT'S DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. D10. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID X2: D-EVPROP34 (LOT: 0012514794); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID X2: L-EVPROP34 (LOT: 0012497770); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: EVPROPLUS-34 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6)2025 H4. DEVICE MFG DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: D-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: D-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: L-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: L-EVPROP34 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NOT IMPLANTABLE PRODUCT ID: EVPROPLUS-34 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423316&pc=GDW | 3005075853-2025-01370 | 2025-01-29 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-20 05:00 | GDW | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Exsanguination; Hemorrhage/Bleeding | K141952 |  | 1 | Event Description: PER USER FACILITY MEDWATCH FORM, #MW5165696, A PATIENT HAD SCHEDULED PROCEDURE FOR LEFT VIDEO ASSISTED THORACIC SURGERY WITH WEDGE RESECTION, POSSIBLE THORACOTOMY WITH LOBECTOMY, LEFT THORACIC NODE DISSECTION. SURGEON NOTED MALFUNCTION WITH STAPLER DEVICE, EXCESSIVE BLEEDING OCCURRED, AND PATIENT EXSANGUINATED IN THE OPERATING ROOM. DEVICE IS AVAILABLE FOR RETURN. Manufacturer Narrative: (B)(4). DATE SENT: 4/2/2025. H1: DEATH. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WERE THE INDICATIONS FOR SURGERY? PET POSITIVE NODULE, LEFT UPPER LOBE. LOBECTOMY. POST DIAGNOSIS WAS CANCER. WHAT WAS THE SURGICAL PROCEDURE? LEFT VIDEO ASSISTED THORACIC SURGERY WITH WEDGE RESECTION POSSIBLE THORACOTOMY WITH LOBECTOMY THORACIC NODE DISSECTION. WHAT ANATOMICAL STRUCTURE/VESSEL WAS THE DEVICE BEING USED ON WHEN THE BLEEDING OCCURRED? MOST SUPERIOR BRANCH OF THE PULMONARY ARTERY. WHAT WAS THE APPROXIMATE COMPRESSED WIDTH OF THE VESSEL THAT THE DEVICE WAS USED ON? ABOUT 1CM. CAN MORE DETAIL BE PROVIDED ON THE MALFUNCTION OF DEVICE? PUT THE STAPLER ON, PROXIMAL TO THE OTHER STAPLER, CLAMPED DOWN, FIRED, BUT WHEN RELEASED THE PULMONARY ARTERY STARTED TO MASSIVELY BLEED. WHAT WAS THE APPROXIMATE AMOUNT OF BLOOD LOSS? OVER 3 LITERS (FROM OP NOTE); TOTAL BLOOD LOSS. WHAT BLOOD PRODUCTS WERE ADMINISTERED? 5 UNITS OF PRBCS AND 1 UNIT OF PLATELETS HOW WAS THE BLEEDING ADDRESSED? APPLY DIRECT PRESSURE. PACKED WITH PACKING AND STILL CONTINUED TO BLEED AROUND IT. WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? USING DEVICE FOR OVER 10 YEARS. NEWER MODEL, BUT IS VERY WELL EXPERIENCED WITH EQUIPMENT. WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? ABOUT 1CM. DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? YES, SURGEON ACKNOWLEDGED THAT HE WAITED AT LEAST 15 SECONDS. DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? YES. DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? NO TISSUE DISTAL WERE ANY SURGICAL CLIPS USED IN THE AREA OF THE DEVICE FIRING? NO SURGICAL CLIPS IN THE AREA. WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? THE DEVICE HAD FIRED ON 3 PRIOR STAPLES LINES AS NORMAL. HOW WAS THE POST-OP BLEEDING IDENTIFIED? MASSIVE BLEEDING FROM MAIN PA. HOW MANY DAYS POSTOPERATIVE WAS THE BLEEDING IDENTIFIED? N/A. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? OVER 3 LITERS (FROM OP NOTE); TOTAL BLOOD LOSS. WAS A TRANSFUSION REQUIRED? YES. WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? N/A. WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? NO. WERE THERE ANY PRE-EXISTING PATIENT CONDITIONS? ELDERLY, AGE, FRAIL TISSUES. IS THERE A RECORDING AVAILABLE OF THE PROCEDURE? NO. WHAT IS THE PATIENT¿S CURRENT STATUS? EXPIRED. Manufacturer Narrative: (B)(4). H1: NOT REPORTABLE. UPON REVIEW OF THE ADDITIONAL INFORMATION PROVIDED, IT WAS CONCLUDED THAT THIS EVENT IS A DUPLICATE OF THE EVENT REPORTED UNDER (B)(4). Manufacturer Narrative: (B)(4). DATE SENT: 2/20/2025. D4 BATCH#: UNK. #MW5165696. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NO LOT OR BATCH NUMBER WAS PROVIDED THEREFORE A DEVICE HISTORY COULD NOT BE DONE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? WHAT WAS THE SURGICAL PROCEDURE? WHAT ANATOMICAL STRUCTURE/VESSEL WAS THE DEVICE BEING USED ON WHEN THE BLEEDING OCCURRED? WHAT WAS THE APPROXIMATE COMPRESSED WIDTH OF THE VESSEL THAT THE DEVICE WAS USED ON? CAN MORE DETAIL BE PROVIDED ON THE MALFUNCTION OF DEVICE? WHAT WAS THE APPROXIMATE AMOUNT OF BLOOD LOSS? WHAT BLOOD PRODUCTS WERE ADMINISTERED? HOW WAS THE BLEEDING ADDRESSED? WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? WERE ANY SURGICAL CLIPS USED IN THE AREA OF THE DEVICE FIRING? WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? HOW WAS THE POST-OP BLEEDING IDENTIFIED? HOW MANY DAYS POSTOPERATIVE WAS THE BLEEDING IDENTIFIED? WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? IS THERE A RECORDING AVAILABLE OF THE PROCEDURE? WHAT IS THE PATIENT¿S CURRENT STATUS? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423423&pc=MNQ | 3007666314-2025-00088 | 2025-01-20 05:00 | Death | INSPIRE MEDICAL SYSTEMS | 2025-02-20 05:00 | MNQ | INSPIRE MODEL 3028 IPG | Adverse Event Without Identified Device or Use Problem | Angioedema |  |  | 1 | Event Description: ON JANUARY 21, 2025, INSPIRE RECEIVED A REPORT OF A PATIENT DEATH WHICH OCCURRED ON (B)(6) 2025, TWO DAYS FOLLOWING THE IMPLANT OF THE INSPIRE SYSTEM WHICH CONSISTED OF THE FOLLOWING COMPONENTS: AN IMPLANTABLE PULSE GENERATOR, A STIMULATION LEAD, AND A RESPIRATORY SENSING LEAD. THE INFORMATION CONTAINED IN THIS MDR WAS REPORTED TO INSPIRE BY THE IMPLANTING PHYSICIAN, THE PATIENT¿S SPOUSE, AND THE INSPIRE REPRESENTATIVE PRESENT AT THE IMPLANT PROCEDURE. ON (B)(6) 2025, THE PATIENT WAS CLEARED FOR SURGERY AND THE IMPLANT PROCEDURE WAS COMPLETED WITHOUT ANY ADVERSE EVENTS OCCURRING. IT WAS REPORTED THAT, PRIOR TO THE PROCEDURE, THE PATIENT TOOK THEIR LISINOPRIL AS USUAL. FOLLOWING THE PROCEDURE, THE PATIENT AWOKE IN THE CARE UNIT AND COMPLAINED OF HAVING A SORE THROAT. SHORTLY THEREAFTER, THE PATIENT WAS CLEARED AND DISCHARGED. UPON ARRIVING HOME, THE PATIENT CONTINUED TO COMPLAIN OF HAVING A SORE THROAT. SHORTLY THEREAFTER, THE PATIENT COMPLAINED OF TONGUE SWELLING. THE SPOUSE CALLED THE IMPLANTING SURGEON¿S ACCESS LINE TO REPORT THE PATIENT¿S CONDITION. IN DESCRIBING THE PATIENT¿S CONDITION TO THE ACCESS LINE PHYSICIAN (A PHYSICIAN OTHER THAN THE IMPLANTING PHYSICIAN), THE PATIENT¿S SPOUSE INDICATED THAT THE PATIENT WAS ABLE TO BREATHE, BUT CONTINUING TO COMPLAIN OF HAVING A SORE THROAT. THE ACCESS LINE PHYSICIAN ADVISED THE PATIENT¿S SPOUSE TO SEEK EMERGENCY MEDICAL CARE FOR THE PATIENT IF THEIR SYMPTOMS WORSENED. APPROXIMATELY 30 MINUTES AFTER THE CALL WITH THE ACCESS LINE, THE PATIENT BEGAN TO COMPLAIN OF HAVING DIFFICULTY BREATHING, AND, SHORTLY THEREAFTER, LOST CONSCIOUSNESS. THE PATIENT¿S SPOUSE CALLED 911 AND INITIATED CHEST COMPRESSIONS ON THE PATIENT UNTIL EMERGENCY RESPONDERS ARRIVED. UPON THEIR ARRIVAL, THE EMERGENCY RESPONDERS PERFORMED A CRICOTHYROTOMY TO ESTABLISH AN EMERGENCY AIRWAY AFTER UNSUCCESSFULLY ATTEMPTING TO INTUBATE THE PATIENT. THE PATIENT WAS THEN TRANSPORTED TO THE HOSPITAL AND ADMITTED INTO THE ICU. THE ATTENDING ICU PHYSICIAN SUSPECTED THAT THE PATIENT WAS EXPERIENCING AN ANGIOEDEMA EVENT. NO IMAGING WAS PERFORMED OF THE PATIENT¿S CHEST AND NECK AREA. WHILE AT THE HOSPITAL, THE PATIENT WAS ON LIFE SUPPORT AND REMAINED IN THE ICU OVERNIGHT. ON (B)(6) 2025, THE IMPLANTING PHYSICIAN WAS NOTIFIED OF THE PATIENT¿S CONDITION AND ASKED TO CONSULT REGARDING THE POSSIBLE TRANSFER OF THE PATIENT FROM THE TREATING HOSPITAL TO THE IMPLANTING PHYSICIAN¿S CARE FACILITY. UPON FURTHER DISCUSSION BETWEEN THE IMPLANTING PHYSICIAN AND THE ICU PHYSICIAN, IMAGING WAS PERFORMED WHICH SHOWED NO SURGICAL SITE ABNORMALITIES, NO HEMATOMA, NO CERVICAL ABNORMALITIES, AND NO VASCULAR INJURIES. IT WAS DETERMINED THAT THE PATIENT WOULD REMAIN AT THE TREATING HOSPITAL. THE PATIENT REMAINED ON LIFE SUPPORT AND IN THE ICU OVERNIGHT. ON (B)(6) 2025, THE FAMILY AUTHORIZED THE WITHDRAWAL OF LIFE SUPPORT. THE PATIENT PASSED AWAY SHORTLY THEREAFTER. ACCORDING TO THE IMPLANTING PHYSICIAN, A REACTION FROM THE ANESTHESIA AND THE PATIENT¿S LISINOPRIL LED TO THE DEVELOPMENT OF THE ANGIOEDEMA WHICH IS THE SUSPECTED CAUSE OF DEATH. AN AUTOPSY WAS NOT PERFORMED. INSPIRE HAS NOT RECEIVED THE PATIENT¿S DEATH CERTIFICATE REPORTING AN OFFICIAL CAUSE OF DEATH. OTHER THAN ROUTINE TESTING REQUIRED DURING THE IMPLANT PROCEDURE, THE PATIENT¿S INSPIRE THERAPY WAS NOT ACTIVE AT ANY POINT FOLLOWING THE IMPLANT PROCEDURE AND IS NOT BELIEVED TO BE THE CAUSE OF, OR A CONTRIBUTING FACTOR TO, THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423441&pc=GDW | 3005075853-2025-01371 | 2025-02-12 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-20 05:00 | GDW | POWERED 60 ECHELON +, 340MM SHAFT | Patient Device Interaction Problem | Insufficient Information | K110385 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LEFT ATRIAL APPENDAGE CLOSURE, THE PATIENT DIED WHEN ECHELON WAS USED. THERE IS A POSSIBILITY THAT OTHER TISSUES WERE INVOLVED AND PROBABLY THERE WAS A PROBLEM WITH THE PROCEDURE NOT BECAUSE OF ECHELON. THE ISSUE HAS OCCURRED BUT THE DETAILS ARE UNKNOWN AT THIS TIME, SO WILL BE ADDED AS ADDITIONAL INFORMATION BECOMES AVAILABLE. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). B3: UNKNOWN: CAPTURED AS AWARENESS DATE. DATE SENT 2/20/2025. D4 BATCH#: UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. INVESTIGATION SUMMARY: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT COLOR CARTRIDGE WAS USED? WERE THERE ANY DIFFICULTIES WITH DEVICE DURING SURGERY CAN A TIMELINE OF EVENTS BE PROVIDED? WHAT WAS THE CAUSE OF THE PATIENT DEATH? WAS AN AUTOPSY PERFORMED, AND CAN THE REPORT BE PROVIDED? DOES THE SURGEON BELIEVE THE PATIENT¿S DEATH WAS RELATED IN ANY WAY TO AN ALLEGED DEFICIENCY WITH THE ETHICON DEVICE? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING PERSONNEL? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT: 4/1/2025. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT IS BELIEVED TO BE THE COD? NOT CLEAR. WHAT COLOR CARTRIDGE WAS USED? WERE THERE ANY DIFFICULTIES WITH DEVICE DURING SURGERY? CAN A TIMELINE OF EVENTS BE PROVIDED? WHAT WAS THE CAUSE OF THE PATIENT DEATH? WAS AN AUTOPSY PERFORMED, AND CAN THE REPORT BE PROVIDED? COULD NOT OBTAIN DETAILED INFORMATION. IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING PERSONNEL? NO.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423623&pc=DSQ | 2916596-2025-01215 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-02-20 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A STROKE ON (B)(6) 2025. THE HEAD COMPUTED TOMOGRAPHY FINDINGS WERE A LARGE LEFT POSTERIOR CEREBRAL ARTERY (LPCA) DISTRIBUTION STROKE. THERE WAS SMALL EXTRACRANIAL RIGHT VERTEBRAL ARTERY THROMBOSIS (OPPOSITE SIDE). THE INTERNATIONAL NORMALIZED RATIO (INR) WAS 1.25. THE PATIENT WAS STILL INTUBATED/SEDATED AND PER PROTOCOL NURSE TURNED DOWN SEDATION TO ASSESS NEURO FUNCTION AND FOUND PATIENT WAS FLACCID ON THE RIGHT SIDE. THE FAMILY THEN DECIDED TO WITHDRAW CARE AND THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PUMP WOULD NOT BE RETURNED.Event Description: ADDITIONAL INFORMATION FROM THE CT SCANS WERE THE PATIENT WAS CARDIOEMBOLIC IN ETIOLOGY IN THE SETTING OF RECENT CARDIAC SURGERY AND COMORBIDITIES. NO TREATMENT WAS INITIATED, NEUROLOGY WAS FOLLOWING AND NEUROINTERVENTIONAL RADIOLOGY CONSULTED AND DOCUMENTED THE FOLLOWING. THE PATIENT HAD KNOWN STAGE D HEART FAILURE WITH PLANNED LEFT VENTRICULAR ASSIST DEVICE (LVAD) AS WELL AS OTHER CARDIAC SURGERIES PLANNED PER INSTITUTE OF HEALTH (IOH). ON (B)(6) 2025 PATIENT DID UNDERGO SCHEDULED SURGERY WITH NOTED NEUROLOGICAL CHANGES POSTOPERATIVELY. STROKE IMAGING OBTAINED SHOWED QUESTION OF SUBACUTE/CHRONIC LEFT PCA INFARCT, POSSIBLE RIGHT CEREBELLAR HYPOATTENUATION, FILLING DEFECT OF THE RIGHT DISTAL VERTEBRAL ARTERY. IMAGING WAS REVIEWED WITH THE DOCTOR, LEFT VERTEBRAL ARTERY THROMBUS WAS CONFIRMED. HEPARIN INFUSION WAS ON HOLD DUE TO RISK OF HEMORRHAGIC CONVERSION. IT WAS RECOMMENDED DUAL ANTIPLATELET THERAPY IN THE INTERIM WITH ASPIRIN AND PLAVIX AS PER DISCUSSION WITH HEART FAILURE TEAM. CONTINUED NEURO CHECKS AS TOLERATED; PATIENT DID BECOME QUITE AGITATED WITH ANY STIMULATION. CONTINUED TO WEAN OFF SEDATION AS ABLE. RECOMMENDED REPEAT CT ANGIOGRAPHY OF NECK IN 1 WEEK TO REASSESS LEFT VERTEBRAL ARTERY. TIMING AND SAFETY OF ANTICOAGULATION PER NEUROLOGY. THERE WAS NO ACUTE ENDOVASCULAR TREATMENT AS THERE WAS NO LARGE VESSEL OCCLUSION. THE PATIENT WAS UNABLE TO HAVE MAGNETIC RESONANCE IMAGING (MRI) DUE TO LVAD. IT WAS NOTED THAT THE PATIENT'S SYMPTOMS WERE AS FOLLOWS: THEY WERE SEDATED AND IT WAS DECREASED FOR A NEURO CHECK. THE NURSE NOTES THE PATIENT WAS NOT MOVING ON THE RIGHT SIDE. IT WAS CONFIRMED THE CARDIAC SURGERY DONE ON (B)(6) 2025 WAS THE LVAD IMPLANT. Manufacturer Narrative: SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 4644 - MEDICATION REQUIRED ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21423742&pc=DSQ | 2916596-2025-01253 | 2023-06-08 04:00 | Death | THORATEC CORPORATION | 2025-02-20 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Renal Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS RENAL. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WOULD NOT BE EXPLANTED. WHETHER THIS WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED.Event Description: THE PATIENT HAD CHRONIC KIDNEY DISEASE PRIOR TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. THE PATIENT STARTED CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) IN THE POST OPERATION PERIOD (B)(6) 2023. THE PATIENT WAS TRANSITIONED TO SCHEDULED HEMODIALYSIS ON (B)(6) 2023. THE CAUSE OF THE RENAL FAILURE WAS ACUTE KIDNEY INJURY MAINLY DUE TO CARDIORENAL SYNDROME IN THE SETTING OF VOLUME OVERLOAD AND SECONDARY TO HYPERTENSIVE NEPHROSCLEROSIS. THE EVENT WAS TREATED WITH DIURETICS, CRRT AND HEMODIALYSIS. THE RENAL FAILURE WAS NOT DEVICE RELATED. THE DEVICE FUNCTIONED AS INTENDED. THE LVAD DID NOT WORSEN THE PATIENT'S RENAL DISEASE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THERE WERE NO DEVICE-RELATED ISSUES WITH (B)(6) REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. IT WAS DETERMINED THAT THE REPORTED INFORMATION DID NOT MEET THE REQUIREMENTS OF A COMPLAINT; HOWEVER, THIS RECORD WILL REMAIN A COMPLAINT AS AN INITIAL MDR (MEDICAL DEVICE REPORT) HAS ALREADY BEEN SUBMITTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE RELATED ISSUES WERE REPORTED BY THE ACCOUNT, THE IFU LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21424114&pc=FKX | 0008030665-2025-00421 | 2025-02-06 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-02-20 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY AT A NURSING HOME FACILITY WHILE CONNECTED TO THE PD CLINIC¿S LIBERTY SELECT CYCLER. THE DEATH WAS REPORTEDLY NOT RELATED TO PD THERAPY. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION, AND THUS FAR NO FURTHER DETAILS HAVE BEEN PROVIDED. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS (PD) UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. HOWEVER, THERE WAS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE CYCLER. ADDITIONALLY, THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THE EVENT. THE INFORMATION INITIALLY REPORTED STATED THAT ALTHOUGH THE PATIENT WAS CONNECTED TO THE CYCLER AT THE TIME OF DEATH, THE DEATH WAS NOT PD RELATED. ALTHOUGH MAINTENANCE DIALYSIS PREVENTS DEATH FROM UREMIA, MORTALITY AMONG PATIENTS WITH END-STAGE KIDNEY DISEASE (ESKD) REMAINS HIGH. AMONG PATIENTS WITH ESKD, CARDIOVASCULAR DISEASE ACCOUNTS FOR APPROXIMATELY 50 PERCENT OF DEATHS. BASED ON THE LIMITED INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21425351&pc=GAG | 2647580-2025-00579 | 2021-09-16 04:00 | Death | US SURGICAL PUERTO RICO | 2025-02-20 05:00 | GAG | TA | Mechanics Altered; Adverse Event Without Identified Device or Use Problem; Insufficient Information | Failure to Anastomose; Emotional Changes; Inflammation; Pain; Shock; Ascites; Unspecified Tissue Injury; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K231491 |  | 1 | Event Description: THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE AND THE ANASTOMOSIS DID NOT SEAL. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, LEUKOCYTOSIS, ASCITES, SEROSITIS, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, DIVERTING LOOP ILEOSTOMY, & CT SCAN. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFO: FURTHER REVIEW OF THIS EVENT FOUND THE PREVIOUSLY REPORTED DEVICE WAS NOT INVOLVED WITH THE REPORTED ISSUE. NO FURTHER REPORTS WILL BE SUBMITTED ON THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6 IME E2402: LEUKOCYTOSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21425379&pc=GAG | 2647580-2025-00580 | 2021-09-16 04:00 | Death | US SURGICAL PUERTO RICO | 2025-02-20 05:00 | GAG | GIA | Mechanics Altered; Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Emotional Changes; Inflammation; Pain; Shock; Ascites; Unspecified Tissue Injury; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K221013 |  | 1 | Event Description: THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE AND THE ANASTOMOSIS DID NOT SEAL. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, LEUKOCYTOSIS, ASCITES, SEROSITIS, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, DIVERTING LOOP ILEOSTOMY, & CT SCAN. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION.Event Description: ACCORDING TO THE REPORTER, DURING AN OPEN SMALL BOWEL ANASTOMOSIS, THE SURGEONS WERE REPAIRING THE FAILED ANASTOMOSIS AND THE FIRST FIRING WAS SATISFACTORY TO THE SURGEONS BUT THE SECOND FIRING HAD A HOLE AND DID NOT SEAL. A TEMPORARY STAPLE LINE WAS CREATED USING AN ULTRA UNIVERSAL STAPLER FOR THE INTRALUMENAL ANASTOMOSIS AND THEN CLOSED THE ANASTOMOSIS WITH A LINEAR STAPLER AND THUS THIS STAPLE LINE WAS CUT OFF AS A NORMAL PART OF CLOSING THE COMMON OPENING WITH THE LINEAR STAPLER. IT WAS NOTED THAT THERE WAS TISSUE DAMAGE. Manufacturer Narrative: CORRECTION: D4 (UNIQUE IDENTIFIER (UDI) #). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6 IME E2402: LEUKOCYTOSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFO: B5, G1, H6 (REMOVED AND ADDED PATIENT CODES, DEVICE CODES), CHANGED ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21425755&pc=OAE | 2029046-2025-00497 | 2025-01-21 05:00 | Death | BIOSENSE WEBSTER INC | 2025-02-20 05:00 | OAE | QDOT MICRO¿ CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Respiratory Failure | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION WITH A QDOT MICRO¿ CATHETER AND THE PATIENT EXPERIENCED CARDIAC ARREST THAT REQUIRED CPR AND MEDICATION. THE PATIENT DIED. AFTER THE PROCEDURE, THE PATIENT BECAME UNSTABLE. ALL CATHETERS WERE REMOVED PRIOR TO THE PATIENT BECOMING UNSTABLE. THE CODE RESPONSE TEAM WAS BROUGHT INTO THE LAB TO ATTEMPT TO RESUSCITATE THE PATIENT. CPR WAS PERFORMED ON THE PATIENT. MULTIPLE MEDICATIONS WERE GIVEN TO THE PATIENT. THE PATIENT WAS TRANSPORTED TO THE INTENSIVE CARE UNIT OF THE HOSPITAL. THE PATIENT PASSED AWAY. THE DATE OF DEATH WAS UNKNOWN. THE PHYSICIAN BELIEVES THE EVENT WAS RELATED TO A RESPIRATORY ISSUE. PRIOR TO CATHETERS BEING REMOVED, THE PHYSICIAN UTILIZED THE ULTRASOUND CATHETER TO CONFIRM THAT NO PERICARDIAL EFFUSION WAS PRESENT. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ON 24-FEB-2025, BWI RECEIVED ADDITIONAL INFORMATION INDICATING THAT THERE WERE NO ISSUES SEEN ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE). SECTION A ON THIS REPORT HAS BEEN UPDATED WITH THE PATIENT'S GENDER AND ETHNICITY (MALE, LATINO). ADDITIONALLY, THE PRODUCT INVESTIGATION WAS COMPLETED AS THE COMPLAINT DEVICE WAS NOT RETURNED FOR ANALYSIS. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21425777&pc=LWS | 9614453-2025-00725 | 2025-02-05 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-20 05:00 | LWS | EVERA MRI S DR SURESCAN | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Myocardial Infarction | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SPOUSE THAT THE PATIENT IS DECEASED AND SPECIFIED THAT THE CAUSE OF DEATH WAS LOW BLOOD PRESSURE AND A HEART ATTACK. THE PATIENT'S SPOUSE STATED THAT THEY DID NOT THINK THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS FAULTY, BUT THEY DID NOT THINK THE ICD WAS NOT RIGHT FOR THE PATIENT AND THE PATIENT DIED BECAUSE OF IT (THE ICD). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426380&pc=OZD | 1220648-2025-26234 | 2025-01-27 05:00 | Death | ABIOMED, INC. | 2025-02-20 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Renal Failure | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT HAD HEMOLYSIS AND ACUTE RENAL FAILURE. THE PATIENT SUFFERED SHOCK LIVER AND ACUTE RENAL FAILURE WITH INCREASING PRESSOR REQUIREMENTS AND PERSISTENT LACTIC ACIDOSIS. HEMOLYSIS WAS ALSO NOTED. DEVICE WAS REPOSITIONED AS IT WAS INITIALLY DEEP. HOWEVER, IT WAS UNCLEAR IF HEMOLYSIS WAS RESOLVING OR NOT GIVEN ANURIA, BUT LACTATE DEHYDROGENASE WAS TRENDING UP. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND RENAL FAILURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26234 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS: INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES: MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426508&pc=NIQ | 2124215-2025-06524 | 2025-01-28 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-20 05:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Chest Pain; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE 80% STENOSED TARGET LESION WAS LOCATED IN THE RIGHT CORONARY ARTERY. A 12 X 2.75 PROMUS ELITE DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO A CARDIAC CARE UNIT. AT NIGHT, THE PATIENT EXPERIENCED CHEST PAIN AND DEVELOPED HYPOTENSION; AND THE BLOOD PRESSURE DROPPED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED IMMEDIATELY BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST, AND AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426509&pc=DSQ | 2916596-2025-00853 | 2024-11-29 05:00 | Death | THORATEC CORPORATION | 2025-02-20 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Pneumonia; Renal Failure; Multiple Organ Failure; Hemorrhagic Stroke; Ischemia Stroke; Embolism/Embolus; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT LOG FILES WERE SUBMITTED FOR REVIEW AND CONTAINED LOW FLOW ESTIMATES AS WELL AS SUSTAINED LOW FLOW HAZARD EVENTS BETWEEN (B)(6) 2025 6:37AM AND (B)(6) 2025 7:04PM. THE LOW FLOW EVENTS APPEARED TO HAVE RESOLVED FOLLOWING THIS TIME FRAME. THE LOW FLOW READINGS DID NOT APPEAR TO BE THE RESULT OF ANY EXTERNAL EQUIPMENT MALFUNCTION. NO PATIENT SYMPTOMS WERE OBSERVED. THE PATIENT WAS ON VENTILATOR AND SEDATED. THE CAUSE OF LOW FLOWS WAS THOUGHT TO BE DUE TO PUMP THROMBOSIS. LOW FLOWS RESOLVED AFTER CARDIOPULMONARY RESUSCITATION (CPR) AND DRUG/FLUID RESUS. THE PATIENT PASSED AWAY ON (B)(6) 2025. IT WAS NOTED THAT THE PATIENT WAS READMITTED ON (B)(6) 2024 WITH SUPERIMPOSED BACTERIAL HOSPITAL-ACQUIRED PNEUMONIA (HAP). THIS RESULTED IN PROLONGED VENTILATION AND ESCALATION OF INOTROPIC SUPPORT. THE PATIENT ALSO DEVELOPED MULTIPLE STROKES THAT WERE LIKELY CARDIOEMBOLIC EVENTS. THE PATIENT EXPERIENCED GENERAL DECLINE SINCE (B)(6) 2025 PRESUMPTIVE ISCHEMIC BOWEL WITH PROGRESSIVE MULTIORGAN FAILURE, REQUIRING RENAL DIALYSIS. PATIENT HAD AN EPISODE OF CARDIOVASCULAR COLLAPSE ON (B)(6) 2025, REQUIRING CPR AND MAXIMAL INOTROPIC SUPPORT. PATIENT PROGRESSIVELY DECLINED AND DEMISED ON (B)(6) 2025 AS NON-CORONER'S CASE. CAUSE OF DEATH WAS ISCHEMIC BOWEL BLOOD GLUCOSE (B/G) AND HEART FAILURE.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD SOME LOW FLOW EVENTS WHILE HEMODYNAMICALLY STABLE ON (B)(6) 2025 AND AGAIN ON (B)(6) 2025. OMISSIONS OF ANTICOAGULANT INFRAVESICAL OBSTRUCTION (IVO) LED TO HEMORRHAGIC CONVERSION OF THE PATIENT'S PREVIOUS CEREBRAL INFARCT WHICH MAY HAVE CONTRIBUTED TO THE PUMP THROMBOSIS. CAUSE OF DEATH WAS ISCHEMIC BOWEL WITH BACKGROUND OF ADVANCED HEART FAILURE AND ILD. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE, AND MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS. A SPECIFIC CAUSE WAS NOT ABLE TO BE DETERMINED THROUGH THIS EVALUATION; HOWEVER, THE ACCOUNT COMMUNICATED THE LOW FLOW ALARMS WERE CAUSED BY PUMP THROMBOSIS. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED CONTROLLER EVENT LOG FILES COLLECTIVELY CONTAINED EVENTS FROM 11JAN2025 THROUGH 28JAN2025. LOW FLOW HAZARD ALARMS WERE CAPTURED ON 18JAN2025, 22JAN2025, 26JAN2025, AND 27JAN2025. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS, STROKE, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. ADDITIONALLY, IT ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION, UNDER ¿ANTICOAGULATION¿, PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING,¿ AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING,¿ OUTLINE SYSTEM CONTROLLER ALARMS, INCLUDING THE LOW FLOW HAZARD ALARM, AND PROVIDE INFORMATION REGARDING HOW TO RESPOND TO AND TROUBLESHOOT THE ALARMS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426524&pc=BTR | 3000219639-2025-00020 |  | Death | PARKER MEDICAL | 2025-02-20 05:00 | BTR | PARKER | Mechanical Problem; Gas/Air Leak | Cardiac Arrest; Low Oxygen Saturation; Diminished Pulse Pressure | K984528 |  | 1 | Event Description: PATIENT WAS SAT UP IN BED FOR A CHEST X-RAY. ET TUBE HAD BEEN IN PLACE SINCE (B)(6) 2025. SUDDEN CHANGE IN OXYGENATION AND VENTILATION NOTED. OBVIOUS CUFF LEAK. RN AND RT INJECTED 3 CC OF AIR WITH IMPROVEMENT AND THEN RAPID SOFTENING OF THE PILOT BALLOON. RECURRENCE OF OBVIOUS CUFF LEAK. UNABLE TO MAINTAIN CUFF PRESSURE. TUBE EMERGENTLY REPLACED. AFTER REINTUBATION, CARDIAC DYSFUNCTION WITH LOSS OF PULSE. CODE CALLED. PILOT LINE WAS CUT DURING EXTUBATION AND REINTUBATION. POST PROCEDURE EXAMINATION OF THE BALLOON SHOWED BALLOON PARTIALLY INFLATED. UTILIZING NEEDLE AND SYRINGE STAFF UNABLE TO INFLATE CUFF. PRESSURE TO CUFF ON THE OUTSIDE OF THE CUFF DOES NOT CREATE A LEAK. UNABLE TO TEST PILOT LINE OR INJECTION PORT.Event Description: PATIENT WAS SAT UP IN BED FOR A CHEST X-RAY. ET TUBE HAD BEEN IN PLACE SINCE (B)(6) 2025. SUDDEN CHANGE IN OXYGENATION AND VENTILATION NOTED. OBVIOUS CUFF LEAK. RN AND RT INJECTED 3 CC OF AIR WITH IMPROVEMENT AND THEN RAPID SOFTENING OF THE PILOT BALLOON. RECURRENCE OF OBVIOUS CUFF LEAK. UNABLE TO MAINTAIN CUFF PRESSURE. TUBE EMERGENTLY REPLACED. AFTER REINTUBATION, CARDIAC DYSFUNCTION WITH LOSS OF PULSE. CODE CALLED. PILOT LINE WAS CUT DURING EXTUBATION AND REINTUBATION. POST PROCEDURE EXAMINATION OF THE BALLOON SHOWED BALLOON PARTIALLY INFLATED. UTILIZING NEEDLE AND SYRINGE STAFF UNABLE TO INFLATE CUFF. PRESSURE TO CUFF ON THE OUTSIDE OF THE CUFF DOES NOT CREATE A LEAK. UNABLE TO TEST PILOT LINE OR INJECTION PORT. Manufacturer Narrative: H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): ENDOTRACEAL TUBE.THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. COMPLAINT WAS CONSIDERED CONFIRMED BASED ON CUSTOMER NARRATIVE OF EVENTS. A 24 MONTH REVIEW WAS CONDUCTED AND 49 ADDITIONAL COMPLAINTS WERE IDENTIFIED RELATED TO THIS ISSUE HOWEVER NO TREND WAS OBSERVED. THE PART AND LOT NUMBER WERE NOT PROVIDED, HOWEVER THE VENDOR WAS NOTIFIED OF THE INCIDENT. THIS INCIDENT WAS NOT IDENTIFIED AS AN INCREASING TREND. NO FURTHER ACTION TO BE TAKEN. Manufacturer Narrative: H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): ENDOTRACEAL TUBE THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426525&pc=BTM | 1314417-2025-00007 | 2025-01-19 05:00 | Death | SUNMED AP, LTD. | 2025-02-20 05:00 | BTM | CURAPLEX | Incomplete or Missing Packaging | Low Oxygen Saturation | N/A |  | 1 | Event Description: ROULETTE CHEMICAL ENG CO, HAS REPORTED "EMS SERVICES WAS AT A CALL, NEEDED A BVM. 2 BAGS THEY RECEIVED DID NOT HAVE THE FACE MASK IN THE BAGS, DELAYING PATIENT CARE. DAWN CONFIRMED THERE WAS A DEATH INVOLVED AFTER THE PATIENT WAS BROUGHT TO THE HOSPITAL.." AS THERE IS A EPORT OF TWO BAGS WITH SAME ISSUE WE WILL SUBMITTING 2 MDR REPORTS TOR THIS INCIDENT. THE FIRST MDR RELATED TO THIS INCIDENT IS 1314417-2025-00008.Event Description: (B)(6) HAS REPORTED "EMS SERVICES WAS AT A CALL, NEEDED A BVM. 2 BAGS THEY RECEIVED DID NOT HAVE THE FACE MASK IN THE BAGS, DELAYING PATIENT CARE. (B)(6) CONFIRMED THERE WAS A DEATH INVOLVED AFTER THE PATIENT WAS BROUGHT TO THE HOSPITAL." AS THERE IS A REPORT OF TWO BAGS WITH SAME ISSUE WE WILL SUBMITTING 2 MDR REPORTS TOR THIS INCIDENT. THE FIRST MDR RELATED TO THIS INCIDENT IS 1314417-2025-00008. Manufacturer Narrative: H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): RESUSCITATION BAG W/MASK THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. Manufacturer Narrative: H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): RESUSCITATION BAG W/MASK. THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. THE COMPLAINT OF "2 BAGS THEY RECEIVED DID NOT HAVE THE FACE MASK IN THE BAGS, DELAYING PATIENT CARE" REGARDING PART 2442-BVM1500 WAS NOT CONFIRMED. THE ROOT CAUSE WAS DETERMINED TO BE THAT THERE IS NO VISUAL AID OR WEIGHING SYSTEM TO DETERMINE IF ALL OF THE REQUIRED COMPONENTS ARE PRESENT IN THE FINISHED GOOD. A RISK ASSESSMENT WAS PERFORMED AND THE ULTIMATE RISK WAS DETERMINED TO BE HIGH WHICH DOES REQUIRE ESCALATION TO THE CAPA REVIEW BOARD. THERE HAVE BEEN 7 OTHER COMPLAINTS REGARDING THE SAME PART AND A SIMILAR ISSUE WITHIN THE 24 MONTHS PRECEDING THIS REPORTED EVENT. A RESOLUTION LETTER WAS SENT TO THE CUSTOMER. COMPLAINTS WILL CONTINUE TO BE MONITORED FOR POSSIBLE TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426527&pc=BTM | 1314417-2025-00006 | 2025-01-19 05:00 | Death | SUNMED AP, LTD. | 2025-02-20 05:00 | BTM | CURAPLEX | Incomplete or Missing Packaging | Low Oxygen Saturation | N/A |  | 1 | Event Description: ROULETTE CHEMICAL ENG CO, HAS REPORTED "EMS SERVICES WAS AT A CALL, NEEDED A BVM. 2 BAGS THEY RECEIVED DID NOT HAVE THE FACE MASK IN THE BAGS, DELAYING PATIENT CARE. DAWN CONFIRMED THERE WAS A DEATH INVOLVED AFTER THE PATIENT WAS BROUGHT TO THE HOSPITAL." AS THERE IS A REPORT OF TWO BAGS WITH SAME ISSUE, WE WILL SUBMITTING 2 MDR REPORTS TOR THIS INCIDENT. THE SECOND MDR RELATED TO THIS INCIDENT IS 1314417-2025-00007.Event Description: (B)(6), HAS REPORTED "EMS SERVICES WAS AT A CALL, NEEDED A BVM. 2 BAGS THEY RECEIVED DID NOT HAVE THE FACE MASK IN THE BAGS, DELAYING PATIENT CARE. (B)(6) CONFIRMED THERE WAS A DEATH INVOLVED AFTER THE PATIENT WAS BROUGHT TO THE HOSPITAL." AS THERE IS A EPORT OF TWO BAGS WITH SAME ISSUE WE WILL SUBMITTING 2 MDR REPORTS TOR THIS INCIDENT. THE SECOND MDR RELATED TO THIS INCIDENT IS 1314417-2025-00007. Manufacturer Narrative: H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): RESUSCITATION BAG W/MASK. THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT: (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. Manufacturer Narrative: H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): RESUSCITATION BAG W/MASK. THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. THE COMPLAINT OF "2 BAGS THEY RECEIVED DID NOT HAVE THE FACE MASK IN THE BAGS, DELAYING PATIENT CARE" REGARDING PART 2442-BVM1500 WAS NOT CONFIRMED. THE ROOT CAUSE WAS DETERMINED TO BE THAT THERE IS NO VISUAL AID OR WEIGHING SYSTEM TO DETERMINE IF ALL OF THE REQUIRED COMPONENTS ARE PRESENT IN THE FINISHED GOOD. A RISK ASSESSMENT WAS PERFORMED AND THE ULTIMATE RISK WAS DETERMINED TO BE HIGH WHICH DOES REQUIRE ESCALATION TO THE CAPA REVIEW BOARD. THERE HAVE BEEN 7 OTHER COMPLAINTS REGARDING THE SAME PART AND A SIMILAR ISSUE WITHIN THE 24 MONTHS PRECEDING THIS REPORTED EVENT. A RESOLUTION LETTER WAS SENT TO THE CUSTOMER. COMPLAINTS WILL CONTINUE TO BE MONITORED FOR POSSIBLE TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426540&pc=DSQ | 2916596-2025-00863 | 2024-08-05 04:00 | Death | THORATEC CORPORATION | 2025-02-20 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATE AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS UNKNOWN WHETHER THE PUMP WOULD BE EXPLANTED OR RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426698&pc=DSQ | 2916596-2025-01025 | 2025-01-15 05:00 | Death | THORATEC CORPORATION | 2025-02-20 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: SECTION A3 - CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE CUSTOMER WILL NOT PROVIDE ANY FURTHER INFORMATION REGARDING THE CAUSE OF DEATH. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: B5 ADDITIONAL INFORMATION NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426804&pc=ITX | 3002808148-2025-02422 | 2025-01-06 05:00 | Death | SHIRAKAWA OLYMPUS CO., LTD. | 2025-02-20 05:00 | ITX | ULTRASONIC PROBE | Peeled/Delaminated; Detachment of Device or Device Component | Insufficient Information | K982323 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ENDOBRONCHIAL ULTRASOUND-GUIDED TRANSBRONCHIAL NEEDLE ASPIRATION PROCEDURE FOR SUSPECTED PERIPHERAL LUNG CANCER, THE ULTRASONIC PROBE WAS USED IN COMBINATION WITH THE SINGLE-USE GUIDE SHEATH. UPON REMOVAL FROM THE PATIENT, IT WAS OBSERVED THAT APPROXIMATELY 5CM OF THE PROBE'S COATING HAD PEELED OFF. CT AND FLUOROSCOPY WERE IMMEDIATELY PERFORMED TO CHECK FOR RETAINED FRAGMENTS, AND NONE WERE DETECTED. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS DISCHARGED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT'S CONDITION SUDDENLY DETERIORATED, REQUIRING EMERGENCY INTERVENTION. THE PATIENT WAS TRANSFERRED TO THE HOSPITAL WHERE DEATH WAS CONFIRMED THE SAME DAY. THE CAUSE OF DEATH WAS DOCUMENTED AS UNKNOWN, AND NO ADDITIONAL INFORMATION REGARDING AUTOPSY FINDINGS OR CONTRIBUTING FACTORS WAS AVAILABLE. THERE WERE NO REPORTS OF PROCEDURAL COMPLICATIONS, IMMEDIATE POST-PROCEDURAL ADVERSE EVENTS, OR CONFIRMED DEVICE-RELATED INJURY. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THE DEVICE WAS RETURNED TO OLYMPUS FOR INSPECTION AND THE REPORTED FAILURE WAS CONFIRMED. IN ADDITION, THE FOLLOWING REPORTABLE MALFUNCTIONS WERE IDENTIFIED DURING THE DEVICE EVALUATION: THERE IS LEAKAGE OF ULTRASONIC MEDIA BASED ON THE RESULTS OF THE INVESTIGATION, IT IS LIKELY THE FOLLOWING LED TO THE MALFUNCTION: THE ELONGATION OF THE MATERIAL AT THE TIP AND THE FACT THAT THE FRACTURE CROSS-SECTION LOOKS LIKE IT HAS BEEN TORN OFF, RATHER THAN CUT WITH A BLADE, INDICATE THAT STRESS WAS APPLIED TO THE TIP SHEATH IN A PULLING MANNER. IN ADDITION, THERE IS NO TWISTING OR CRUSHING OF THE INSERTION SHEATH, INDICATING THAT THE STRESS WAS NOT APPLIED TO THE TIP SHEATH WHILE THE INTERNAL BLADE (FLEXIBLE SHAFT) WAS BEING DRIVEN. THEREFORE, IT IS ASSUMED THAT THE TIP SHEATH WAS DAMAGED AND FELL OFF DUE TO THE STRESS APPLIED TO THE TIP SHEATH WHILE THE INTERNAL BLADE (FLEXIBLE SHAFT) WAS NOT BEING DRIVEN. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21426812&pc=EOQ | 9614641-2025-00182 | 2025-01-06 05:00 | Death | AOMORI OLYMPUS CO., LTD. | 2025-02-20 05:00 | EOQ | SINGLE USE GUIDE SHEATH KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K060243 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ENDOBRONCHIAL ULTRASOUND-GUIDED TRANSBRONCHIAL NEEDLE ASPIRATION PROCEDURE FOR SUSPECTED PERIPHERAL LUNG CANCER, THE ULTRASONIC PROBE WAS USED IN COMBINATION WITH THE SINGLE-USE GUIDE SHEATH. UPON REMOVAL FROM THE PATIENT, IT WAS OBSERVED THAT APPROXIMATELY 5CM OF THE PROBE'S COATING HAD PEELED OFF. CT AND FLUOROSCOPY WERE IMMEDIATELY PERFORMED TO CHECK FOR RETAINED FRAGMENTS, AND NONE WERE DETECTED. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS DISCHARGED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT'S CONDITION SUDDENLY DETERIORATED, REQUIRING EMERGENCY INTERVENTION. THE PATIENT WAS TRANSFERRED TO THE HOSPITAL WHERE DEATH WAS CONFIRMED THE SAME DAY. THE CAUSE OF DEATH WAS DOCUMENTED AS UNKNOWN, AND NO ADDITIONAL INFORMATION REGARDING AUTOPSY FINDINGS OR CONTRIBUTING FACTORS WAS AVAILABLE. THERE WERE NO REPORTS OF PROCEDURAL COMPLICATIONS, IMMEDIATE POST-PROCEDURAL ADVERSE EVENTS, OR CONFIRMED DEVICE-RELATED INJURY. THERE WAS NO REPORTED MALFUNCTION OR ISSUE WITH THE SINGLE-USE GUIDE SHEATH. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THERE WAS NO ALLEGED DEVICE FAILURE. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21427698&pc=QZI | 2124215-2025-10225 | 2025-01-21 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-20 05:00 | QZI | FARAWAVE | Off-Label Use | Arrhythmia; Bradycardia |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WITH A FARAWAVE CATHETER WAS COMPLETED WITHOUT ISSUES. THE PHYSICIAN COMPLETED AN OFF-LABEL CTI ABLATION AND POSTERIOR WALL ABLATION. AFTER THE PROCEDURE WAS COMPLETED, ALL DEVICES WERE REMOVED FROM THE PATIENT AND THE PHYSICIAN TOLD THE ANESTHESIA TEAM TO WAKE THE PATIENT UP. IT WAS THEN REPORTED THAT THE PATIENT WAS HAVING BRADYCARDIA EPISODES, A TRANSVENOUS PACING WIRES WERE PLACED AND AN ARTERIAL LINE. AN ULTRASOUND WAS PERFORMED AND RULED OUT ANY PERICARDIAL EFFUSION. THE PATIENT HAD A LOW EJECTION FRACTION (EF) PRIOR TO ABLATION. A SWAN CATHETER WAS PLACED AND THE PATIENT WAS SENT TO INTENSIVE CARE UNIT (ICU) WHERE WENT INTO VENTRICULAR TACHYCARDIA AND DID NOT MAKE IT. NO EKG WAS OBSERVED DURING THE PROCEDURE. NO ABNORMALITIES ON THE WALL WERE OBSERVED. NO EFFUSION WAS FOUND. THE DEVICE IS NOT EXPECTED TO RETURN FOR ANALYSIS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21428428&pc=LOZ | 3003761017-2025-00090 | 2025-02-15 05:00 | Death | SYNCARDIA SYSTEMS, LLC | 2025-02-20 05:00 | LOZ | SYNCARDIA 70CC TAH | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P030011 |  | 1 | Event Description: REPORTED DEATH OF PATIENT DUE TO MULTI-SYSTEM ORGAN FAILURE. PATIENT'S FAMILY DECIDED TO WITHDRAW CARE DUE TO HEALTH STATUS. NO AUTOPSY WAS PERFORMED, THE DEVICE WAS NOT EXPLANTED, AND HOSPITAL STAFF REPORTED DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. Manufacturer Narrative: PATIENT DEATH DUE TO MULTI-SYSTEM ORGAN FAILURE. PATIENT'S FAMILY MADE THE DECISION TO WITHDRAW SUPPORT. DEVICE CONFIRMED BY HOSPITAL STAFF TO NOT HAVE CAUSED OR CONTRIBUTED TO PATIENT DEATH. DEVICE WAS NOT EXPLANTED AND NO AUTOPSY WAS PERFORMED. SYNCARDIA HAS COMPLETED ITS EVALUATION AND IS CLOSING THIS FILE. IF NEW OR ADDITIONAL INFORMATION IS RECEIVED IN THE FUTURE, SYNCARDIA WILL FILE A FOLLOW-UP MDR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21428605&pc=NPT | 2015691-2025-01185 | 2025-01-27 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-02-20 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Burst Container or Vessel; Material Fragmentation; Difficult to Remove | No Clinical Signs, Symptoms or Conditions | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST (FCS), DURING A TAVR PROCEDURE WITH A 20MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION, DURING INFLATION THE BALLOON RUPTURED ON CALCIUM IN THE STJ. THE DELIVERY SYSTEM WAS IMMEDIATELY ATTEMPTED TO BE REMOVED BUT BECAME STUCK PARTIALLY RETRIEVED INTO THE SHEATH. VASCULAR WAS CALLED; THEY WERE ABLE TO REMOVE THE SHEATH LEAVING THE DELIVERY SYSTEM BEHIND. WHEN ATTEMPTING TO REMOVE THE DELIVERY SYSTEM IT BECAME STUCK IN THE FEMORAL AND BROKE LEAVING BEHIND THE NOSECONE AND BALLOON. THE ARTERY WAS STABILIZED WITH MANUAL PRESSURE AND THE PATIENT REMAINED STABLE. THE PATIENT WAS INTUBATED AND THE ILLIAC WAS BALLOON OCCLUDED WHILE THEY DID A FEMORAL CUTDOWN TO REMOVE THE BROKEN END OF THE DELIVERY SYSTEM. THE BALLOON WAS REMOVED AND STENTS WERE PLACED IN THE FEMORAL ARTERY. THE VESSEL WAS SURGICALLY CLOSED. THE PATIENT REMAINED STABLE THROUGHOUT. UPON FOLLOW UP, THE FCS ADVISED THAT THE PATIENT WAS DISCHARGED HOME AFTER THE PROCEDURE. ON (B)(6) 2025 THEY WERE ADMITTED BACK TO THE HOSPITAL WITH ABDOMINAL PAIN. THEY WERE FOUND TO HAVE AN ISCHEMIC BOWEL. THE PATIENT EXPIRED ON (B)(6) 2025 FROM COMPLICATIONS FROM THEIR ISCHEMIC BOWEL.Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST (FCS), DURING A TAVR PROCEDURE WITH A 20MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION, DURING INFLATION THE BALLOON RUPTURED ON CALCIUM IN THE STJ. THE DELIVERY SYSTEM WAS IMMEDIATELY ATTEMPTED TO BE REMOVED BUT BECAME STUCK PARTIALLY RETRIEVED INTO THE SHEATH. VASCULAR WAS CALLED, THEY WERE ABLE TO REMOVE THE SHEATH LEAVING THE DELIVERY SYSTEM BEHIND. WHEN ATTEMPTING TO REMOVE THE DELIVERY SYSTEM IT BECAME STUCK IN THE FEMORAL AND BROKE LEAVING BEHIND THE NOSECONE AND BALLOON. THE ARTERY WAS STABILIZED WITH MANUAL PRESSURE AND THE PATIENT REMAINED STABLE. THE PATIENT WAS INTUBATED AND THE ILLIAC WAS BALLOON OCCLUDED WHILE THEY DID A FEMORAL CUTDOWN TO REMOVE THE BROKEN END OF THE DELIVERY SYSTEM. THE BALLOON WAS REMOVED AND STENTS WERE PLACED IN THE FEMORAL ARTERY. THE VESSEL WAS SURGICALLY CLOSED. THE PATIENT REMAINED STABLE THROUGHOUT. UPON FOLLOW UP, THE FCS ADVISED THAT THE PATIENT WAS DISCHARGED HOME AFTER THE PROCEDURE. THE PATIENT WAS ADMITTED TO THE HOSPITAL 10 DAYS POST TAVR WITH ABDOMINAL PAIN. THEY WERE FOUND TO HAVE AN ISCHEMIC BOWEL. THE PATIENT EXPIRED 2 DAYS AFTER ADMISSION FROM COMPLICATIONS FROM THEIR ISCHEMIC BOWEL. THE FCS WAS NOT AWARE OF A PREVIOUS BOWEL ISSUE WITH THIS PATIENT, BUT IN SPEAKING TO ONE OF THE HEART TEAM MEMBERS, THEY THOUGHT IT WAS LIKELY RELATED TO THE TAVR PROCEDURE. THEY THOUGHT IT COULD POSSIBLY HAVE BEEN CAUSED FROM AN AORTIC OCCLUSION BALLOON BEING INFLATE FOR TOO LONG DECREASING BLOOD FLOW TO THE ARTERIES THAT SUPPLY THE BOWEL OR COULD HAVE BEEN SOMETHING THAT EMBOLIZED DOWN THOSE ARTERIES THAT DECREASED OR STOPPED BLOOD FLOW. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS, SECTIONS G6, H2, H6: COMPONENT CODES, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAS BEEN ADDED. PER THE INSTRUCTIONS FOR USE (IFU), POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE INCLUDE THROMBUS FORMATION, PLAQUE DISLODGMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION, STROKE, DISTAL PERIPHERAL OCCLUSION, AND/OR DEATH. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS WIPING AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE THE CLOT. THE DEVICE TRAINING MANUALS AND IFU INSTRUCT THE OPERATOR TO ADMINISTER HEPARIN AND MAINTAIN THE ACT AT = 250 SEC. THERE ARE MULTIPLE ETIOLOGIES FOR BOWEL ISCHEMIA/INFARCTION INCLUDING EMBOLIZATION OCCURRING AFTER DEVICE MANIPULATION, BALLOON VALVULOPLASTY, OR VALVE DEPLOYMENT AND THESE EVENTS CAN RESULT IN VARYING DEGREES OF PERMANENT IMPAIRMENT. ADDITIONALLY, PROLONGED HYPOTENSION CAN CONTRIBUTE TO DECREASED BOWEL PERFUSION AND ISCHEMIA POTENTIALLY LEADING TO TISSUE NECROSIS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE (EXTENDED OCCLUSION BALLOON INFLATION OR EMBOLUS) CAUSED OR CONTRIBUTED TO THE BOWEL INFARCTION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, BALLOON BURST EVENTS DURING VALVE DEPLOYMENT HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO CALCIFICATION IN THE LANDING ZONE OR OVER-INFLATION OF THE BALLOON. THE COMPLAINT FOR BALLOON BURST WAS UNABLE TO CONFIRM AS DEVICE WAS NOT RETURNED AND NO APPLICABLE IMAGERY FOR EVALUATION. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THE BALLOON RUPTURED DURING INFLATION ON CALCIUM IN THE STJ AND 3MENSIO REVEALED THE PRESENCE OF CALCIFICATION ON STJ. THE PRESENCE OF CALCIFICATION CAN CREATE A CHALLENGING ANATOMY FOR BALLOON INFLATION. WHILE THE BALLOONS ARE SUFFICIENTLY DESIGNED AND TESTED TO ENSURE THE BURST PRESSURE IS AT OR ABOVE THE RATED BURST PRESSURE, CALCIFIED NODULES CAN COMPROMISE THE STRUCTURE OF THE BALLOON WALL VIA FOLLOWING MECHANISMS SUCH AS PUNCTURE, LOCAL OVERSTRETCHING, OPEN CELL IMPINGEMENT, OR STRESS CONCENTRATION. THE COMPLAINT FOR INABILITY TO WITHDRAW SYSTEM THROUGH SHEATH WAS UNABLE TO BE CONFIRMED AS DEVICE WAS NOT RETURNED AND NO APPLICABLE IMAGERY FOR EVALUATION. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (WITHDRAWAL OF BURST BALLOON) MAY HAVE CONTRIBUTED TO THE EVENT. AS THE BALLOON WAS BURST, THE ALTERED BALLOON PROFILE COULD HAVE MADE IT MORE SUSCEPTIBLE TO CATCH ON THE DISTAL END OF SHEATH TIP WHICH WOULD HAVE THEN LED TO THE EXPERIENCED RETRIEVAL DIFFICULTY. THE COMPLAINT FOR INABILITY TO WITHDRAW SYSTEM THROUGH VASCULATURE WAS UNABLE TO BE CONFIRMED AS DEVICE WAS NOT RETURNED AND NO APPLICABLE IMAGERY FOR EVALUATION. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION/TORTUOSITY/UNDERSIZED VESSEL) AND/OR PROCEDURAL FACTORS (IMPROPER WITHDRAWAL TECHNIQUES) MAY HAVE CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT 'VASCULAR WAS CALLED, THEY WERE ABLE TO REMOVE THE SHEATH LEAVING THE DELIVERY SYSTEM BEHIND. WHEN ATTEMPTING TO REMOVE THE DELIVERY SYSTEM IT BECAME STUCK IN THE FEMORAL AND BROKE LEAVING BEHIND THE NOSECONE AND BALLOON.' PER TRAINING MANUAL, USER IS INSTRUCTED TO UTILIZE SNARE TECHNIQUE OR WITHDRAW THE DELIVERY SYSTEM THROUGH THE TIP OF THE SHEATH. ATTEMPTING TO REMOVE JUST THE SHEATH CAN CAUSE THE DELIVERY SYSTEM TO GET STUCK INSIDE PATIENT'S BODY, RESULTING INTO WITHDRAWAL DIFFICULTIES. ADDITIONALLY, 3MENSIO REVEALED TORTUOSITY, CALCIFICATION AND UNDERSIZED VESSEL IN PATIENT'S ANATOMY AND THESE FACTORS MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING WITHDRAWAL. THE COMPLAINT FOR AND DISTAL TIP WAS UNABLE TO BE CONFIRMED AS DEVICE WAS NOT RETURNED AND NO APPLICABLE IMAGERY FOR EVALUATION. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (DEVICE MANIPULATION) MAY HAVE CONTRIBUTED TO THE EVENT. ADDITIONAL PULL FORCE/DEVICE MANIPULATION COULD HAVE BEEN APPLIED TO OVERCOME THE WITHDRAWAL DIFFICULTY WHICH THEN LED TO THE REPORTED SEPARATION. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS, SECTIONS G6, H2, B5 UPDATED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21428704&pc=DSQ | 2916596-2025-00979 | 2024-12-11 05:00 | Death | THORATEC CORPORATION | 2025-02-20 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Tricuspid Valve Insufficiency/ Regurgitation; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE AND END ORGAN FAILURE.Event Description: THE RIGHT VENTRICULAR FAILURE DID NOT EXIST BEFORE LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. HOWEVER, DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO IT. ON (B)(6) 2024, THE PATIENT PRESENTED FROM THE CLINIC WITH DYSPNEA, POOR APPETITE, AND SUSPECTED CONGESTIVE HEART FAILURE (CHF). THE PATIENT HAD PROGRESSIVE RIGHT VENTRICULAR DYSFUNCTION AND HEPATIC CONGESTION GIVEN SIGNIFICANT TRICUSPID REGURGITATION. THE PATIENT WAS AGGRESSIVELY DIURESED WITH LASIX (FUROSEMIDE) DRIP. RIGHT HEART CATHETERIZATION (RHC) WAS DONE ON (B)(6) 2024 WITH FEMORAL SWAN DUE TO DIFFICULT ACCESS SHOWED BILATERALLY ELEVATED FILLING PRESSURES WITH BORDERLINE CARDIAC INDEX (RA 26 V 34, PA 40/21/31, W 23 V TO 35, FCI 2.2, WT 67.4 KG). THE PATIENT'S DECOMPENSATION WAS MOSTLY RELATED TO RIGHT VENTRICULAR DYSFUNCTION AND SEVERE TRICUSPID REGURGITATION. BILATERAL ULTRASOUND OF UPPER EXTREMITY PERFORMED TO RULE OUT DEEP VEIN THROMBOSIS (DVT) OR ATYPICAL ANATOMY FOUND RIGHT INTERNAL JUGULAR (IJ) THROMBUS AND CHRONIC POST THROMBOTIC CHANGES IN LEFT INTERNAL JUGULAR (LIJ). THE RHC POOR HEMODYNAMICS WERE TREATED WITH ONGOING AGGRESSIVE DIURESIS AND MILRINONE TO HELP AUGMENT DIURESIS. HOWEVER, THE PATIENT DEVELOPED NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) WITH MILRINONE AND THE RENAL FUNCTION WORSENED WITH INOTROPE SUPPORT, ALSO WITH HYPONATREMIA STATUS POST TOLVAPTAN, AND BECAME HYPOTENSIVE. THE PATIENT REQUIRED MULTIPLE PRESSORS AND WAS TRANSFERRED TO CLINICAL TRIALS COORDINATION CENTER (CTCC) FOR FURTHER EVALUATION. BLOOD CULTURES WERE ALSO FOUND POSITIVE, AND THE PATIENT WAS TREATED WITH ANTIBIOTICS. THE PRESSOR REQUIREMENTS IMPROVED SLOWLY, AND EVENTUALLY WEANED OFF EPINEPHRINE ON (B)(6) 2024 AND REMAINED STABLE. THE PATIENT'S DECOMPENSATION WAS THOUGHT TO BE SYSTEMIC INFLAMMATORY RESPONSE SYNDROME (SIRS) FROM BACTEREMIA AND URINARY TRACT INFECTION. THE PATIENT REMAINED WEAK WITH POOR PROGNOSIS WITH RIGHT VENTRICULAR FAILURE. THE PATIENT HAD NOT BEEN A TRANSPLANT CANDIDATE DUE TO HIGH PANEL-REACTIVE ANTIBODIES (PRA) DESPITE DESENSITIZATION. TRANSFERRING TO ANOTHER INSTITUTION FOR DESENSITIZATION, BUT AFTER DISCUSSIONS WITH PALLIATIVE CARE, THE PATIENT WAS MORE INTERESTED IN HOSPICE PHILOSOPHY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SHOCK THERAPIES WERE TURNED OFF BEFORE DISCHARGE. THE PATIENT WAS ULTIMATELY DISCHARGED ON (B)(6) 2024 WITH PLANS TO SIGN ON WITH HOSPICE ON (B)(6) 2024. THE PATIENT UNDERSTOOD AND DID NOT WANT TO WAIT TO DISCHARGE ON (B)(6) 2024. 3 DAYS OF HOSPICE MEDICATIONS SENT PER HOSPICE NURSING INSTRUCTIONS. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: SECTION B3: EVENT DATE CORRECTED. SECTION H6: 4440 - THROMBUS. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, VENOUS THROMBOEMBOLISM, ARTERIAL NON-CENTRAL NERVOUS SYSTEM (CNS) THROMBOEMBOLISM, ARRHYTHMIA, INFECTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HM3 LVAS. SECTION 6 "PATIENT CARE AND MANAGEMENT" DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 ALSO LISTS INFECTION AND THROMBOEMBOLISM AS POTENTIAL LATE POST-IMPLANT COMPLICATIONS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. THE IFU ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21407233&pc=LWP | 2017865-2025-14451 | 2025-01-30 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-19 05:00 | LWP | ASSURITY MRI | Incorrect Measurement | Arrhythmia | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT A REVIEW OF A REMOTE TRANSMISSION SHOWING A HIGH VENTRICULAR RATE (HVR) ELECTROGRAM WITH HYSTERESIS INITIATING VENTRICULAR FIBRILLATION (VF). THE PATENT DIED. UPON FURTHER REVIEW IN THE CLINIC, THE RIGHT VENTRICULAR (RV) AUTOCAPTURE WAS TURNED ON AT 3-MONTH POST-OP CHECK. THE DEVICE TECHNICIAN DID NOT REALIZE HYSTERESIS AUTOMATICALLY TURNS ON WHEN AUTOCAPTURE IS TURNED ON IN A SINGLE-CHAMBER DEVICE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: THE PROVIDED SESSION RECORD WAS REVIEWED BY TECHNICAL SERVICES AND IT WAS DETERMINED BASED ON THE EGMS THAT THE DEVICE FUNCTIONED PER PROGRAMMED SETTINGS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21409783&pc=BZD | 2518422-2025-102137 | 2024-11-07 05:00 | Death | RESPIRONICS, INC. | 2025-02-19 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT¿S WIFE REPORTED RECEIVING A LETTER INDICATING THAT HER HUSBAND¿S CPAP DEVICE WOULD BE SHIPPED. SHE CLARIFIED THAT HER HUSBAND PASSED AWAY LAST YEAR AND WILL NO LONGER REQUIRE THE DEVICE. ACCORDING TO THE SYSTEM RECORDS, THE DEVICE WAS SHIPPED OUT A COUPLE OF YEARS AGO. HOWEVER, THE WIFE CONFIRMED THAT SHE DOES NOT HAVE THE DEVICE IN HER POSSESSION. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21409906&pc=OZD | 1220648-2025-26196 | 2025-01-26 05:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Increase in Pressure; Positioning Problem | Unspecified Infection; Tachycardia; Thrombocytopenia; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THERE WAS PURGE SYSTEM BLOCKED ALARMS AND PRESSURE WAS INCREASING. THE STAFF ADDED TISSUE PLASMA ACTIVATOR TO THE PURGE AND THE PATIENT REMAINS ON SUPPORT.Event Description: ADDITIONALLY, THE PATIENT WAS HEPARIN-INDUCED THROMBOCYTOPENIA (HIT) POSITIVE DURING SUPPORT. FURTHERMORE, THE PATIENT ALSO HAD VENTRICULAR TACHYCARDIA (VT) DURING SUPPORT. ALSO, THERE WERE SUCTION ALARMS THROUGHOUT SUPPORT. ADDITIONALLY, THE PUMP WAS FOUND DEEP. THE PUMP WAS EVENTUALLY REPOSITIONED. FURTHERMORE, THE PATIENT HAD AN INFECTION DURING SUPPORT. THE SOURCE OF THE INFECTION IS UNKNOWN. THE PATIENT EXPIRED ON SUPPORT. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: B1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B2 REVISED TO REFLECT DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B5 REVISED TO INCLUDE THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. D6B EXPLANT DATE ADDED. H1 TYPE OF REPORT REVISED TO REFLECT DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 REVISED TO INCLUDE THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & CP WITH SMARTASSIST. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS PER THE DATA LOG REVIEW, THE ROOT CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL BASED ON DATA LOG REVIEW AND THE CLINICAL DETAILS PROVIDED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26196 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21410081&pc=MKJ | 1220908-2025-00549 | 2025-01-25 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-02-19 05:00 | MKJ | X SERIES | Patient Device Interaction Problem | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE ASSOCIATED CPR STAT-PADZ WERE UNABLE TO ADHERE TO THE PATIENT. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER SET OF PADS TO CONTINUE TREATING THE PATIENT. HOWEVER, THIS SET WAS ALSO UNABLE TO ADHERE TO THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE ELECTRODE PADS INVOLVED WERE NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. INSTEAD, RETAINED SAMPLE TESTING OF THE SAME LOT NUMBER (3023E) WAS PERFORMED. THE RETAINED SAMPLE INVESTIGATION RESULTED IN NO FAULTS FOUND AND CONCLUDED THAT THE ELECTRODES WERE ASSEMBLED ACCORDING TO SPECIFICATION. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21410449&pc=HRS | 0008031020-2025-00331 | 2020-06-01 04:00 | Death | STRYKER GMBH | 2025-02-19 05:00 | HRS | UNKNOWN AXSOS 3 TITANIUM LOCKING PLATE | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  |  | 1 | Event Description: ¿THE MANUFACTURER RECEIVED STRYKER COLLABORATIVE STUDY (SCS) NAMED "A RETROSPECTIVE DATA COLLECTION OF THE INTERNAL FRACTURE FIXATION OF LONG BONES WITH THE AXSOS3 TITANIUM LOCKING PLATE SYSTEM" THAT CONTAINS COLLECTED DATA ON THE USAGE AND THE OUTCOMES OF AXSOS 3 TITANIUM LOCKING PLATE SYSTEM. THE REPORT DETAILS ANALYSIS PROVIDED FOR SURGICAL PROCEDURES PERFORMED BETWEEN JANUARY 1, 2007 THROUGH JUNE 1, 2020. DURING THE REVIEW OF THE REPORT, IT WAS NOT POSSIBLE TO ESTABLISH A SPECIFIC DEVICE DETAIL, PATIENT INFORMATION, AND CURRENTLY NO ADDITIONAL DEVICE INFORMATION IS AVAILABLE; HOWEVER, THE FOLLOWING ADVERSE EVENT WAS REPORTED: INFECTION, DEEP (EARLY) IN 4 PATIENTS.¿ THIS IS PATIENT 4 OF 4.Event Description: ¿THE MANUFACTURER RECEIVED STRYKER COLLABORATIVE STUDY (SCS) NAMED "A RETROSPECTIVE DATA COLLECTION OF THE INTERNAL FRACTURE FIXATION OF LONG BONES WITH THE AXSOS3 TITANIUM LOCKING PLATE SYSTEM" THAT CONTAINS COLLECTED DATA ON THE USAGE AND THE OUTCOMES OF AXSOS 3 TITANIUM LOCKING PLATE SYSTEM. THE REPORT DETAILS ANALYSIS PROVIDED FOR SURGICAL PROCEDURES PERFORMED BETWEEN JANUARY 1, 2007 THROUGH JUNE 1, 2020. DURING THE REVIEW OF THE REPORT, IT WAS NOT POSSIBLE TO ESTABLISH A SPECIFIC DEVICE DETAIL, PATIENT INFORMATION, AND CURRENTLY NO ADDITIONAL DEVICE INFORMATION IS AVAILABLE; HOWEVER, THE FOLLOWING ADVERSE EVENT WAS REPORTED: INFECTION, DEEP (EARLY) IN 4 PATIENTS.¿ THIS IS PATIENT 4 OF 4. IN 4TH PATIENT OUT OF 4, THIS PATIENT ORIGINALLY PRESENTED FOR A PERIPROSTHETIC DISTAL FEMUR FRACTURE AND REQUIRED THIGH FASCIOTOMIES AND INITIAL STABILIZATION WITH AN EXTERNAL FIXATOR PRIOR TO DEFINITIVE FIXATION. FOLLOWING DISCHARGE TO REHAB, THE PATIENT DEVELOPED SIGNS OF INFECTIONS AND WAS ADMITTED TO THE HOSPITAL FOR FURTHER CARE OF THE INFECTION. THEY UNDERWENT 2 IRRIGATION AND DEBRIDEMENT BUT DEVELOPED SEPTICEMIA AND RESPIRATORY FAILURE THAT ULTIMATELY CAUSED THEM TO EXPIRE. Manufacturer Narrative: THE REPORTED EVENT COULD NOT BE CONFIRMED, SINCE THE DEVICE(S) WERE NOT RETURNED FOR EVALUATION AND NO OTHER ADDITIONAL INFORMATION WAS RECEIVED FROM THE CLINICAL SITE. MORE DETAILED INFORMATION ABOUT THE PATIENT'S MEDICAL HISTORY, THE EVENT DETAILS AND THE INVOLVED DEVICE(S) MUST BE AVAILABLE TO DETERMINE THE ROOT CAUSE. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE, THE INVESTIGATION WILL BE REOPENED AND RE-EVALUATED ACCORDINGLY. Manufacturer Narrative: ADDITIONAL INFORMATION AND CORRECTED DATA: IN 4TH PATIENT OUT OF 4, THIS PATIENT ORIGINALLY PRESENTED FOR A PERIPROSTHETIC DISTAL FEMUR FRACTURE AND REQUIRED THIGH FASCIOTOMIES AND INITIAL STABILIZATION WITH AN EXTERNAL FIXATOR PRIOR TO DEFINITIVE FIXATION. FOLLOWING DISCHARGE TO REHAB, THE PATIENT DEVELOPED SIGNS OF INFECTIONS AND WAS ADMITTED TO THE HOSPITAL FOR FURTHER CARE OF THE INFECTION. THEY UNDERWENT 2 IRRIGATION AND DEBRIDEMENT BUT DEVELOPED SEPTICEMIA AND RESPIRATORY FAILURE THAT ULTIMATELY CAUSED THEM TO EXPIRE. THIS COMPLAINT HAS BEEN GENERATED BASED ON FINDINGS DISCOVERED DURING THE POST-MARKET SURVEILLANCE LITERATURE REVIEW. THE ALLEGED EVENT OF COULD NOT BE CONFIRMED, SINCE NO ADDITIONAL INFORMATION WAS RECEIVED FROM THE AUTHOR OR THE ARTICLE. MORE DETAILED INFORMATION ABOUT THE PATIENT MEDICAL HISTORY, THE EVENT CIRCUMSTANCES, AND MEDICAL REPORTS MUST BE AVAILABLE TO DETERMINE THE ROOT CAUSE. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE, THE INVESTIGATION WILL BE REOPENED AND RE-EVALUATED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411150&pc=LTI | 3012638928-2025-03273 |  | Death | SPATZ FGIA INC. | 2025-02-19 05:00 | LTI | SPATZ3 ADJUSTABLE BALLOON SYSTEM | Insufficient Information | Insufficient Information | P190012 |  | 1 | Event Description: WE HAVE BEEN NOTIFIED ABOUT A DEATH IN COLOMBIA, SOUTH AMERICA. WE HAVE VERY LIMITED INFORMATION. WE HAVE A CAT SCAN WHICH SHOWS A DISTENDED STOMACH WITH A BALLOON IN THE ANTRUM. WE WERE TOLD BY THE DISTRIBUTOR THAT THE BALLOON WAS REMOVED, AND THE PATIENT'S DEATH WAS DUE TO ASPIRATION AND A CARDIAC EVENT - BUT THAT HAS NOT BEEN CONFIRMED WITH MEDICAL REPORTS YET. Manufacturer Narrative: THE COMPANY HAS REQUESTED MORE INFORMATION AND THE AUTOPSY REPORT TO DETERMINE THE CAUSE OF DEATH AND ITS RELATION TO THE DEVICE/PROCEDURE AND WE WILL PROVIDE A FOLLOW UP REPORT WHEN NEW INFORMATION WILL BE REVEALED. ADDITIONAL INFORMATION: A DELAY IN THIS EVENT REPORTING WAS DUE TO INCOMPLETE INITIAL INFORMATION WHICH REQUIRES MORE TIME THAN ANTICIPATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411223&pc=OZD | 1220648-2025-26203 | 2025-01-27 05:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED PUMP MALPOSITIONING, LOW PUMP FLOW AND SUBSEQUENTLY EXPIRED. THE NEXT DAY AFTER START OF SUPPORT, THE STAFF WAS PRONING THE PATIENT WHEN THE IMPELLA CP DEVICE WAS PULLED BACK INTO THE AORTA DUE TO THE TUOHY-BORST BEING NOT COMPLETELY CLOSED. AN IMPELLA IN AORTA ALARM WAS TRIGGERED WITH IMPELLA FLOWS 2.8/2.6 (2.7) LITERS PER MINUTE SUDDENLY DROPPING TO FLOWS OF 0.1/0.0 (0.0) LITERS PER MINUTE. THERE WERE STILL NO SIGNS OF PULSATILE MOTOR CURRENT AND THE FLOWS WERE STILL 0.1/0.0 (0.0) LITERS PER MINUTE. AN ATTEMPT WAS MADE TO REPOSITION THE PUMP. HOWEVER, IT WOULD NOT MOVE. SEVERAL TIMES LOST FLOWS WERE POINTED OUT, AND THE PUMP COULD HAVE BEEN STUCK/CAUGHT SOMEWHERE AND/OR HAVE A KINK. THE POSITION NOR FLOWS COULD BE RESTORED. THE LEFT VENTRICLE WAS DISTENDED AND CONTRACTILITY WAS VERY BAD. A DISCUSSION ON HOW TO PROCEED WAS MADE AND STEPS TAKEN THEREAFTER ARE UNKNOWN. HOWEVER, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY THE NEXT DAY. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE. THE ROOT CAUSE OF THE POSITIONING ISSUE WAS TOUHY-BORST RELATED TO IMPROPER TECHNIQUE D6B WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26203. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411241&pc=MOZ | 3007093114-2025-00003 | 2025-01-01 05:00 | Death | ANIKA | 2025-02-19 05:00 | MOZ | MONOVISC | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THIS REPORT IS BEING SUBMITTED IN PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ANIKA OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, ANIKA, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A POTENTIAL PATIENT EVENT DOCUMENTED ON THIS REPORT. ON (B)(6) 2025 THE DISTRIBUTOR REPORTED TO ANIKA THAT A 92-YEAR-OLD FEMALE PATIENT WHO RECEIVED A MONOVISC INJECTION (COURSE UNKNOWN) ON AN UNSPECIFIED DATE IN (B)(6) 2024 EXPIRED ON AN UNSPECIFIED DATE, PRESUMABLY IN 2025. THE PATIENT REPORTEDLY RECEIVED A ROUTINE MONOVISC INJECTION EVERY SIX MONTHS. THERE WAS NO SPECIFIC ALLEGATION OF A TEMPORAL ASSOCIATION BETWEEN THE PATIENT'S DEATH AND THE MONOVISC INJECTIONS. THERE ARE NO SPECIFIC ALLEGATION OF A PRODUCT DEFICIENCY. THE LOT NUMBER OF THE MOST RECENT INJECTION WAS NOT PROVIDED. ADDITIONAL INFORMATION WAS SOLICITED.Event Description: THIS REPORT IS BEING SUBMITTED IN PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ANIKA OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, ANIKA, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A POTENTIAL PATIENT EVENT DOCUMENTED ON THIS REPORT. ON 27JAN2025 THE DISTRIBUTOR REPORTED TO ANIKA THAT A (B)(6)-YEAR-OLD FEMALE PATIENT WHO RECEIVED A MONOVISC INJECTION (COURSE UNKNOWN) ON AN UNSPECIFIED DATE IN (B)(6) 2024 EXPIRED ON AN UNSPECIFIED DATE, PRESUMABLY IN 2025. THE PATIENT REPORTEDLY RECEIVED A ROUTINE MONOVISC INJECTION EVERY SIX MONTHS. THERE WAS NO SPECIFIC ALLEGATION OF A TEMPORAL ASSOCIATION BETWEEN THE PATIENT'S DEATH AND THE MONOVISC INJECTIONS. THERE ARE NO SPECIFIC ALLEGATION OF A PRODUCT DEFICIENCY. THE LOT NUMBER OF THE MOST RECENT INJECTION WAS NOT PROVIDED. ADDITIONAL INFORMATION WAS SOLICITED. Manufacturer Narrative: THIS CASE IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON RECEIPT OF NEW AND RELEVANT INFORMATION OR UPON THE COMPLETION OF THE INVESTIGATION AT THE MANUFACTURING PLANT. Manufacturer Narrative: THIS CASE IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON RECEIPT OF NEW AND RELEVANT INFORMATION OR UPON THE COMPLETION OF THE INVESTIGATION AT THE MANUFACTURING PLANT. SUPPLEMENTAL REPORT. THE REPORTED EVENT IS NOT CONFIRMED. ADDITIONAL INFORMATION, PATIENT DEATH CERTIFICATE, MEDICAL RECORDS WAS NOT PROVIDED UPON REQUEST. THE LOT NUMBER WAS NOT PROVIDED UPON REQUEST. A REVIEW OF THE BATCH RECORD WAS NOT PERFORMED. A THREE-YEAR RETROSPECTIVE REVIEW OF THE NCRS THIS PRODUCT WAS PERFORMED. THERE WAS NO NONCONFORMANCES DOCUMENTED THAT WAS RELATED TO THE REPORTED EVENT. A THREE YEAR RETROSPECTIVE REVIEW OF THE PRODUCT RETENTION INSPECTION REPORTS WAS PERFORMED. THERE ARE NO NONCONFORMANCES RELATED TO THE REPORTED EVENT. A REVIEW OF THE RECENT STABILITY STUDY REPORT FOR MONOVISC WAS PERFORMED. THERE WERE NO NON-CONFORMANCES DOCUMENTED. THE CONCLUSION FOR THE PRODUCT STATED THAT THE PRODUCT HAS MET ALL REQUIRED SPECIFICATIONS AND TOLERANCES. THE CAUSE OF THE PATIENT DEATH COULD NOT BE ESTABLISHED BASED ON A CLINICAL ASSESSMENT OF THIS COMPLAINT. THERE IS INSUFFICIENT INFORMATION TO ESTABLISH A TEMPORAL ASSOCIATION BETWEEN THE USE OF THE DEVICE AND THE PATIENT'S DEATH. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON RECEIPT OF NEW AND RELEVANT INFORMATION. THIS CASE WILL BE MONITORED AND TRENDED FOR FUTURE ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411371&pc=DSQ | 2916596-2025-01058 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-19 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿TEMPORAL ANALYSIS IN OUTCOMES OF LONG-TERM MECHANICAL CIRCULATORY SUPPORT: RETROSPECTIVE STUDY¿ THAT HEARTMATE 3 (HM 3) MAY BE ASSOCIATED WITH DEVICE THROMBOSIS, BLEEDING, INFECTION, DRIVELINE INFECTIONS, NEUROLOGICAL COMPLICATIONS, RIGHT VENTRICULAR FAILURE, AND DEATH. THIS SINGLE-CENTER RETROSPECTIVE STUDY EVALUATED TWO TIME INTERVALS TO REFLECT CHANGES IN VENTRICULAR ASSIST DEVICE TECHNOLOGY (PERIOD 1: 2007-2015; PERIOD 2: 2016-2022). 181 PATIENTS WHO UNDERWENT LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION WERE EVALUATED. DEVICE UTILIZATION WAS THE FOLLOWING: HEARTMATE II=52 (76.4%) AND HEARTWARE=16 (23.6%) IN PERIOD 1 AND HEARTMATE II=2(1.8%), HEARTMATE 3=70 (61.9%), HEARTWARE=29 (25.7%), SYNCARDIA TAH=10 (8.8%) AND BERLINHEART EXCOR=2 (1.8%) IN PERIOD 2. THE MEAN AGE (IQR) OF ALL THE PATIENTS WAS 53 YEARS AND 11.6% OF ALL THE PATIENTS WERE WOMEN. RESULTS SHOWED THE INCIDENCE OF PUMP THROMBOSIS AND GASTROINTESTINAL BLEEDING WAS HIGHER IN THE FIRST TIME PERIOD, BUT NOT SIGNIFICANTLY DIFFERENT, BUT BLEEDING COMPLICATIONS WERE MORE PREVALENT IN THE FIRST PERIOD. THE INCIDENCE OF ANY INFECTIONS, DRIVELINE INFECTIONS, NEUROLOGICAL EVENTS, AND RIGHT VENTRICULAR FAILURE REQUIRING SHORT-TERM MECHANICAL SUPPORT WERE THE SAME IN BOTH TIME INTERVALS. HOWEVER, MORTALITY WAS SIGNIFICANTLY HIGHER IN THE FIRST PERIOD THAN IN THE SECOND ONE. OUT OF THE 113 PATIENTS EVALUATED IN PERIOD 2, THE FOLLOWING ADVERSE EVENTS WERE REPORTED: 5 PATIENTS EXPERIENCED DEVICE THROMBOSIS, 11 PATIENTS EXPERIENCED ANY BLEEDING EVENTS, 0 PATIENTS EXPERIENCED GASTROINTESTINAL BLEEDING, 45 PATIENTS EXPERIENCED ANY INFECTIONS, 27 PATIENTS EXPERIENCED DRIVELINE INFECTIONS, 19 PATIENTS EXPERIENCED NEUROLOGICAL COMPLICATIONS, 12 PATIENTS EXPERIENCED RIGHT VENTRICULAR ASSIST DEVICE (RVAD) IMPLANTATION, AND 32 PATIENTS DIED. RESULTS SHOWED SURVIVAL WAS SIGNIFICANTLY HIGHER DURING THE SECOND TIME PERIOD. DESPITE THE HIGHER RISK PROFILE OF THE PATIENTS AND PERSISTENT CHALLENGES DURING THE SECOND PERIOD, THERE WAS A SIGNIFICANT DECREASE IN MORTALITY AND MORBIDITY. THE USE OF THE HEARTMATE 3 DEVICE MAY HAVE CONTRIBUTED TO THIS RESULT. Manufacturer Narrative: SECTION B2: DATE OF DEATH ESTIMATED AS EVENT DATE. DATE OF EVENT HAS BEEN ENTERED AS 01JAN2022 AS DATA WAS COLLECTED BETWEEN 2007 AND 2022. SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. ONDRUSEK, M., ET AL. (2024). TEMPORAL ANALYSIS IN OUTCOMES OF LONG-TERM MECHANICAL CIRCULATORY SUPPORT: RETROSPECTIVE STUDY. THE THORACIC AND CARDIOVASCULAR SURGEON, 72(07), 521¿529. HTTPS://DOI.ORG/10.1055/S-0044-1782600. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411452&pc=OMP | MW5166398 | 2025-02-06 05:00 | Death | 3M / KCI USA, INC. | 2025-02-19 05:00 | OMP | 3M ACTIV.A.C. THERAPY SYSTEM |  | Hemorrhage/Bleeding; Thrombosis/Thrombus |  |  | 1 | Event Description: SUBJECT HAD A SARCOMA REMOVED FROM THE LEFT THIGH IN (B)(6) AND HAD AN ACTIV.A.C. PLACED OVER THE WOUND. SCENE PICTURES SHOW ABUNDANT BLOOD LOSS IN THE BATHROOM AND BEDROOM (SOAKING THE CARPET). SUBJECT WAS FOUND DEAD BY WIFE UPON RETURN FROM WORK IN THE AFTERNOON. AT AUTOPSY, BLOOD-SOAKED TOWEL WRAPPED AROUND THE LEG, ABUNDANT BLOOD CLOTS UNDER THE DRESSING, VERY LITTLE BLOOD (10 CC) RETRIEVABLE FROM THE HEART, AND ACTIV.AC. FULL OF BLOOD. HE BLEEDS FROM THE WOUND APPROXIMATELY 200 CC ON (B)(6) AND WAS MANAGED AT THE ER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411455&pc=DSQ | 2916596-2025-00981 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-19 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿END OF LIFE CARE IN PATIENTS WITH LONG TERM MECHANICAL CIRCULATORY SUPPORT IN WESTERN AUSTRALIA¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH STROKE, HEART FAILURE, DEATH. THE GOAL OF THIS STUDY WAS TO IMPROVE THE PROVISION AND UNDERSTANDING OF END-OF-LIFE CARE (EOLC) FOR LONG TERM MECHANICAL CIRCULATORY SUPPORT (LTMCS). A RETROSPECTIVE OBSERVATIONAL ANALYSIS OF ALL PATIENT DEATHS WITH HEARTMATE 2 (HM2), HEARTMATE 3 (HM3) AND HEARTWARE (HVAD) LMCS IN WESTERN AUSTRALIA FROM 2007 TO 2023 WAS PERFORMED. THE AVERAGE LENGTH OF LTMCS SUPPORT AT TIME OF DEATH WAS 1,115 DAYS (RANGE 3¿4100 DAYS). 65% OF DEATHS HAD .90 DAYS OF SUPPORT AND 60% OF THOSE DEATHS OCCURRED IN A HOSPITAL. 40% OF THE DEATHS WERE CAUSED BY STROKE, AND 20% OF THE DEATHS WERE CAUSED BY HEART FAILURE. 55% OF PATIENTS HAD DOCUMENTED GOALS OF CARE (GOC) AND 35% HAD PALLIATIVE REVIEW. PALLIATIVE CARE INVOLVEMENT WAS CORRELATED WITH LESS INVASIVE INTERVENTIONS. EOLC FOR LTMCS REQUIRES EARLY REFERRAL TO PALLIATIVE CARE. DEATHS WITH PALLIATIVE CARE INVOLVEMENT HAD FEWER MEDICAL INTERVENTIONS, WERE MORE LIKELY TO DIE AT HOME WITH SYMPTOM CONTROL AND MEDICATION TITRATION AND HAD GREATER FAMILY INVOLVEMENT. STRATEGIES TO INCREASE PALLIATIVE CARE SERVICES TO LTMCS PATIENTS MAY LEAD TO IMPROVED QUALITY OF LIFE. Manufacturer Narrative: SECTIONS A AND D4: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE. SMIT, R., ROBINSON, C., MCLEAN, J., DONOVAN, A., DELLA-BOSCA, F., HAYES, H., & LAM, K. (2024). END OF LIFE CARE IN PATIENTS WITH LONG TERM MECHANICAL CIRCULATORY SUPPORT IN WESTERN AUSTRALIA. HEART LUNG AND CIRCULATION, 33, S261. HTTPS://DOI.ORG/10.1016/J.HLC.2024.06.289. (B)(6) HOSPITAL, PERTH, WA, AUSTRALIA. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING STROKE AND DEATH. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411637&pc=OZD | 1220648-2025-26205 | 2025-01-14 05:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Cardiogenic Shock; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT UNDERGOING AN ABLATION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). IMPELLA WAS IMPLANTED USING BEST PRACTICE. AND THE ABLATION WAS PERFORMED. THE DECISION MADE TO KEEP THE IMPELLA IN PLACE. THE PEEL-AWAY REMOVED, REPOSITIONING SHEATH ADVANCED, ANGLE MATCHED, AND IMPELLA SECURED. THE PATIENT WAS SENT TO THE UNIT ON IMPELLA MCS. PLAN TO CONTINUE FLUID RESUSCITATION, MONITOR HEART FUNCTION AND HEMODYNAMICS. LATER IN THE DATE IT WAS REPORTED THE PATIENT HAD WORSENING HEMODYNAMICS. THE PATIENT WAS CANNULATED FOR VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO), AND BLOOD WAS TRANSFUSED WITH FLUIDS. APPROXIMATELY 13 LATER, THE PATIENT HAD A GASTROINTESTINAL BLEED AND BLOOD PRODUCTS: 1 UNIT OF PACKED RED BLOOD CELLS, PLATELETS 2 UNITS AND FRESH FROZEN PLASMA 1 UNIT WERE GIVEN. HEPARIN WAS USED IN THE IMPELLA PURGE SYSTEM, AND IT IS UNKNOWN IF THE USE OF THE HEPARIN IMPACTED OR EXACERBATED THE OUTCOME OF THE BLEED. THE PATIENT WOULD EXPIRE THREE DAYS LATER.Event Description: ADDITIONAL INFORMATION RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED PROGRESSION OF CARDIOGENIC SHOCK WITH A "DEFINITE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿DESPITE MAXIMAL EFFORTS, FAMILY DECIDED TO PROCEED WITH WITHDRAWAL OF CARE.¿. THE PATIENT OUTCOME WAS FATAL. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT ENDURED HEMORRHAGIC SHOCK WITH A "DEFINITE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿POST IMPELLA IMPLANT, PATIENT DEVELOPED A LARGE HEMATOMA IN THE LEFT GROIN. LACTIC BEGAN INCREASING, REQUIRED VA ECMO CANNULATION. MULTIPLE BLOOD PRODUCTS ADMINISTERED.¿. THE PATIENT IS NOT RECOVERED AND ISSUE WAS NOT RESOLVED. Manufacturer Narrative: B.5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H.6 UPDATED TO REFLECT THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H.10 UPDATED TO ADD INSTRUCTIONS FOR USE FOR THE H.6 CODES THAT WERE ADDED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ B.2 REVISED AS LIFE THREATENING WAS INCORRECTLY SELECTED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205. B.3 REVISED EVENT DATE DUE TO INFORMATION IN B.5 SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205 WAS SUBMITTED. D.3 E-MAIL ADDRESS SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205. D.4 REVISED LOT NUMBER AND UNIQUE IDENTIFIER (UDI) # AS THEY WERE PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205. G.1 REVISED MANUFACTURING SITE NAME AND ADDRESS AS THEY WERE PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205. G.2 ADDED STUDY AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205. H.4 REVISED DEVICE MANUFACTURE DATE AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26205. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411653&pc=OZD | 1220648-2025-26206 | 2025-01-26 05:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA 5.5 HAD HIGH PURGE PRESSURE, PURGE SYSTEM BLOCKED ALARM WHILE ON PATIENT SUPPORT. ANTICOAGULATION HAD NOT BEEN IN RANGE FOR THE LAST FEW HOURS. MOTOR CURRENT WAS IN RANGE. THE PHYSICIAN WAS ADVISED TO PERFORM TISSUE PLASMINOGEN ACTIVATOR-LYSIS. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON IMPELLA SUPPORT. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE PRODUCT WAS NOT RETURNED. THE PURGE PRESSURE WAS STABLE FOR 23 HOURS AND THEN THERE WERE SOME FLUCTUATIONS IN THE PURGE PRESSURE AND THE PURGE FLOW STARTED TO DECREASE. THERE WERE NO MOTOR CURRENT SPIKES OR LONG BAG/CASSETTE CHANGES. THE FLUCTUATIONS LASTED FOR THREE HOURS AND THEN THERE WAS A RAPID RISE OVER 38 MINUTES, WHICH TRIGGERED HIGH PURGE PRESSURE/BLOCKED ALARMS. IT REMAINED ELEVATED FOR 12 HOURS AND THEN IT STARTS TO SLOWLY DECREASE AND IS STABLE THROUGHOUT THE REST OF THE CASE. THIS IS ASSOCIATED WITH BIOMATERIAL BUILDUP IN THE PURGE GAP. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE IS MOST LIKELY DUE TO BIOMATERIAL BUILDUP BASED ON THE DATA LOG ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411678&pc=OZD | 1220648-2025-25747 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | IMPELLA CP | Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND APPROXIMATELY 16 HOURS AFTER START OF SUPPORT, IT WAS REPORTED THE PUMP'S POSITION WAS COMPLETELY IN THE AORTA AND REPOSITIONING WAS IMPOSSIBLE. PATIENT CRITICAL WAS TAKEN BACK TO THE CATHETERIZATION LAB, AND THE PUMP WAS EXPLANTED AND REPLACED WITH NEW IMPELLA CP DEVICE. THE PATIENT WAS RETURNED TO THE INTENSIVE CARE UNIT ON IMPELLA MCS. HOWEVER, WITHIN A COUPLE OF HOURS ON THE SECOND IMPELLA CP DEVICE, IT WAS NOTED TO BE DISLOCATED; THIS PUMP SLIPPED INTO THE AORTA AS WELL. ACCORDING TO THE PHYSICIAN, THE TUOHY BORST VALVE WAS FIXED. CONSULTATION WAS MADE WITH THE CHIEF PHYSICIAN, AND THE DECISION WAS MADE TO EXPLANT AND PLACE ANOTHER NEW (NOW THIRD IMPLANT) IMPELLA CP DEVICE AS IT WAS NOTED THE PATIENT WAS IN ¿VERY¿ CRITICAL CONDITION. THE THIRD REPLACEMENT PUMP WAS SUCCESSFULLY PLACED. HOWEVER, THE FOLLOWING DAY WITHIN SOME HOURS THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. PRODUCT WAS NOT RETURNED FOR REVIEW. THE ROOT CAUSE OF POSITIONING ISSUE WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25747 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: PATIENT-SPECIFIC INFORMATION IS UNAVAILABLE AT THE TIME OF THIS MDR WRITING; THEREFORE, RESPECTIVE FIELDS REFLECT "UNKNOWN" OR LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO MEDICAL DEVICE REPORTS RELATED TO THE PATIENT IMPLANT EXPERIENCE. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-25744.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21411928&pc=NPT | 2015691-2025-01123 | 2025-01-29 05:00 | Death | EDWARDS LIFESCIENCES | 2025-02-19 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P140031 |  | 1 | Event Description: AS REPORTED BY OUR UNITED KINGDOM AFFILIATES, DURING IMPLANT OF A 20MM SAPIEN 3 ULTRA VALVE IMPLANT WITHIN A NON-EDWARDS DEGENERATED 21MM AORTIC VALVE BIOPROSTHESIS, PRE-DILATATION WITH A 18MM NON-EDWARDS BALLOON WAS PERFORMED WITHOUT ISSUES AND THE SAPIEN 3 VALVE WAS DEPLOYED AT NOMINAL VOLUME SUCCESSFULLY. AT THIS POINT, ALL THE CORONARY ARTERIES WERE PERFUSED, BUT AFTER POST-DILATATION WITH A 20MM NON-EDWARDS BALLOON TO FRACTURE THE RING, THE PATIENT DEVELOPED CHEST PAIN AND ISCHEMIC ECG CHANGES. THE SAPIEN 3 ULTRA VALVE WAS STABLE WITH NO MOVEMENT AND DEPLOYED IN THE DESIRED POSITION. HOWEVER, THE AORTOGRAM SHOWED LEFT MAIN CORONARY ARTERY (LMA) OCCLUDED AND ACCESS TO THE LMA VESSEL USING THE GUIDEWIRE WAS OBTAINED BUT THE PATIENT ARRESTED (ECG ' VENTRICULAR FIBRILLATION (VF)), THERE WAS NOT ANY OUTPUT AND THE PATIENT NEEDED RESUSCITATION. A BALLOONING TO OPEN LMA WAS PERFORMED AND A 4.0X 20 STENT WAS IMPLANTED. THE PATIENT RECOVERED AND WAS STABLE INITIALLY BUT ARRESTED AGAIN SHORTLY AFTER. ANOTHER AORTOGRAM SHOWED THAT THE LMA WITH DEPLOYED STENT HAD BLOCKED OFF AGAIN. ACCESS WAS GAINED AGAIN, AND A FURTHER BALLOONING AND STENTING WERE PERFORMED, THIS TIME USING A 4.0 X 28 STENT. THE PATIENT REMAINED IN VF WITH NO OUTPUT AND DID NOT RESPOND TO RESUSCITATION THEREAFTER. CPR AND FULL ADVANCED LIFE SUPPORT (ALS) PROTOCOL WAS FOLLOWED BUT PATIENT PASSED AWAY AFTER 1.5 HOURS OF RESUSCITATION. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE AORTIC VALVE REPLACEMENT PROSTHESIS COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412521&pc=DSQ | 3007042319-2025-00530 | 2025-02-14 05:00 | Death | HEARTWARE, INC. | 2025-02-19 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Obstruction of Flow; Pumping Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF HEART FAILURE EXPERIENCED A LOW FLOW ALARM, INDICATING POTENTIAL COMPLICATIONS WITH THE VENTRICULAR ASSIST DEVICE (VAD). THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT, WHERE IT WAS SUSPECTED THAT THERE WAS AN INFLOW OCCLUSION WITH LITTLE TO NO FLOW THROUGH THE VAD AND LOW POWER. AN ECHOCARDIOGRAM AND COMPUTED TOMOGRAPHY SCAN WERE CONDUCTED, BOTH INDICATING NO FLOW THROUGH THE VAD. COMFORT MEASURES WERE ADMINISTERED, AND THE VAD WAS TURNED OFF. THE PATIENT SUBSEQUENTLY EXPIRED AN HOUR AFTER THE VAD WAS TURNED OFF. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED A SUDDEN DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS ON (B)(6) 2025, LEADING TO PARAMETERS BELOW NORMAL OPERATING RANGE, AND THIRTEEN (13) LOW FLOW ALARMS LOGGED SINCE (B)(6) 2025. AS A RESULT, THE REPORTED LOW FLOW EVENT WAS CONFIRMED. THE REPORTED OCCLUSION EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AN HOUR AFTER THE VAD WAS TURNED OFF. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412846&pc=DSQ | 2916596-2025-01052 | 2018-10-01 04:00 | Death | THORATEC CORPORATION | 2025-02-19 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "DYNAMIC RISK ESTIMATION OF ADVERSE EVENTS IN AMBULATORY LVAD PATIENTS A MOMENTUM 3 ANALYSIS" HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST DEVICE (LVAD) RECIPIENTS WHO SURVIVED HOSPITAL DISCHARGED NUMBERING 2,056 PATIENTS WITH A MEDIAN AGE OF 59.4 YEARS (20.4% WOMEN) WERE STUDIED FOR THE INCIDENCE OF GASTROINTESTINAL BLEEDING (GIB) (25.6%), STROKE (6.0%), AND DEATH (12.3%) AFTER 2 YEARS. AFTER 1 YEAR IT WAS 19.6%, 4.4% AND 7.1%. RECENT DATA FROM MOMENTUM 3 ESTIMATED THAT OVERALL PATIENT SURVIVAL AFTER 5 YEARS WAS 58.4%, 1 YEAR WAS 85.9% AND 2 YEARS WAS 78.8%. AFTER 1 YEAR IT WAS FOUND THAT INCIDENCE OF STROKE AND GIB REMAIN HIGH AT 11.9% AND 21.1% OF PATIENTS OVERALL. THE POOL OF PATIENT FOR THIS STUDY WERE PATIENTS ACROSS 69 U.S. CENTERS WERE ELIGIBLE IF THEY HAD ADVANCED HEART FAILURE REFRACTORY TO STANDARD MEDICAL THERAPY, HAD AN EJECTION FRACTION LESS THAN OR EQUAL TO 25%, AND WERE NOT EXPECTED TO REQUIRE BIVENTRICULAR SUPPORT OR HAVE IRREVERSIBLE END-ORGAN DAMAGE. PATIENTS IN THE IDE COHORT UNDERWENT 1:1 RANDOMIZATION TO HEARTMATE II OR HM3 AND WERE IMPLANTED BETWEEN SEPTEMBER 2014 AND AUGUST 2016 WITH FOLLOW-UP THROUGH AUGUST 2018. CAP PATIENTS RECEIVED THE HM3 FROM AUGUST 2016 TO OCTOBER 2018 WITH FOLLOW-UP THROUGH OCTOBER 2020. ONLY HM3 PATIENTS WERE INCLUDED IN THE PRESENT ANALYSIS. PATIENT¿S WERE NOTED TO HAVE HISTORY OF ISCHEMIC OR NONISHEMIC/OTHER CARDIOMYOPATHY TYPES. 41.8 % OF PATIENT HAD ATRIAL FIBRILLATION. PATIENT PRIOR HISTORY INCLUDED STROKE 8%, GIB 1.1%, VALVE REPLACEMENT OR REPAIR 6.9% AND CORONARY ARTERY BYPASS SURGERY 18.3 %. PATIENT HAD VALVULAR REGURGITATION OF ANY SEVERITY AORTIC 77.2%, MITRAL 68.0%, AND TRICUSPID 57.7%. GIB INCLUDED BOTH UPPER AND LOWER GI TRACT BLEEDING EVENTS. STROKE INCLUDED BOTH ISCHEMIC AND HEMORRHAGIC STROKES. PATIENT¿S WITH RESPIRATORY FAILURE, RIGHT HEART FAILURE, AND INFECTION WERE NOTED TO BE AT HIGHER RISK OF DEATH. RIGHT HEART FAILURE LED TO A 5-FOLD INCREASE IN THE ODDS OF DYING. Manufacturer Narrative: SECTIONS A-D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. B3: THE EVENT DATE HAS BEEN ENTERED AS 01OCT2018 AS PATIENTS WERE IMPLANTED BETWEEN SEP2014 AND OCT2018. INOVA SCHAR HEART AND VASCULAR, FALLS CHURCH, VIRGINIA, USA COLUMBIA UNIVERSITY IRVING MEDICAL CENTER AND NEW YORK-PRESBYTERIAN HOSPITAL, NEW YORK, NEW YORK, USA CARDIOVASCULAR INSTITUTE AT ALLEGHENY HEALTH NETWORK, PITTSBURGH, PENNSYLVANIA, USA HENRY FORD HOSPITAL, DETROIT, MICHIGAN, USA UNIVERSITY OF MICHIGAN, ANN ARBOR, MICHIGAN, USA BRIGHAM AND WOMEN'S HOSPITAL HEART AND VASCULAR CENTER AND HARVARD MEDICAL SCHOOL, BOSTON, MASSACHUSETTS, USA UNIVERSITY OF COLORADO SCHOOL OF MEDICINE, AURORA, COLORADO, USA THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412875&pc=KGC | 9611594-2025-00031 |  | Death | AVANOS MEDICAL INC. | 2025-02-19 05:00 | KGC | GASTROINTESTINAL ANCHOR SET, SAF-T-PEXY T-FASTENERS | Break | Peritonitis |  |  | 1 | Event Description: IT WAS REPORTED THE "SUTURES RUPTURED." THE PATIENT DIED FROM PERITONITIS. Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. THE LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. ALL INFORMATION REASONABLY KNOWN AS OF 18-FEB-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412960&pc=OZD | 1220648-2024-22743 | 2022-06-01 04:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. AUTOMATED IMPELLA CONTROLLER (AIC) MALFUNCTIONED IN WHICH THE AIC FAILED TO RECOGNIZE THE PUMP. THE INSTRUCTIONS FOR USE (IFU) WAS FOLLOWED AND A BACKUP CONTROLLER WAS USED. ADDITIONAL INFORMATION NOTED THAT THE CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE.Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE IMPELLA WAS PLACED MID PROCEDURE DUE TO MULTIPLE EPISODES OF SUSTAINED VENTRICULAR TACHYCARDIA. DEFIBRILLATION WAS REQUIRED AND THERAPY WAS DELIVERED OVER 30 TIMES IN THE PROCEDURAL AREA, INCLUDING SEVERAL ROUNDS DELIVERED FROM PATIENTS IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR. UPON PLUGGING IN THE AUTOMATED IMPELLA CONTROLLER (AIC), IT WAS REPORTED THAT THE AIC FAILED TO RECOGNIZE THE PUMP. THE INSTRUCTIONS FOR USE (IFU) WAS FOLLOWED AND A BACKUP CONTROLLER WAS USED. ADDITIONAL INFORMATION NOTED THAT THE CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED PUMP NOT DETECTED HAS BEEN COMPLETED. THE PRODUCT WAS RETURNED AND THE PUMP NOT DETECTED ISSUE WAS CONFIRMED. THE CONSOLE WAS THOROUGHLY TESTED ON AN ENGINEERING LAB BENCH. THE EPM CRYSTAL ISSUE WAS EASILY REPRODUCED BUT WAS EFFECTIVELY MITIGATED BY PMD RESET. THE PUMP DETECTION ISSUES THAT OCCURRED IN THE FIELD WERE MOST LIKELY DUE TO THE EPM CRYSTAL ISSUE. THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS. Manufacturer Narrative: ADDITIONAL INFORMATION WAS PROVIDED IN B.5 AND H.10. ADDITIONAL REVIEW OF THIS COMPLAINT EVENT WAS CONDUCTED VIA ABIOMED'S MEDICAL SAFETY OFFICE. WHILE ABIOMED IS UNABLE TO DISSOCIATE THE PATIENT'S OUTCOME TO THE IMPELLA DEVICE, OUR SAFETY OFFICE BELIEVES THE PRODUCT ISSUE IS UNLIKELY TO BE RELATED TO THE PATIENT OUTCOME. THE DEATH WAS LIKELY DUE TO PATIENTS CLINICAL CONDITION (PRE-PROCEDURAL CARDIOGENIC SHOCK AND SUSTAINED VT THAT REQUIRED CARDIOVERSION PRIOR TO INITIATION OF IMPELLA THERAPY). THE DELAY IN SUPPORT WAS TRANSIENT AND THERAPY WAS INITIATED ON THE BACK-UP CONSOLE. THERAPY CONTINUED FOR APPROXIMATELY 8.5 HOURS UNTIL THE PATIENT EXPIRED. D.2 REVISED COMMON DEVICE NAME IN ACCORDANCE WITH UPDATED PROCEDURES SINCE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2024-22743 WAS SUBMITTED. G.1 REVISED REPORTING CONTACT EMAIL IN ACCORDANCE WITH UPDATED PROCEDURES SINCE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2024-22743 WAS SUBMITTED. H.6 CODE 4629 WAS ADDED AS IT WAS INADVERENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2024-22743. H.10 CORRECTED INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412964&pc=OZD | 1220648-2024-22733 | 2022-09-26 04:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK ON AN AUTOMATED IMPELLA CONTROLLER (AIC) FOR MECHANICAL CIRCULATORY SUPPORT. THE COMPLAINANT REPORTED THAT THE AUTOMATED IMPELLA CONTROLLER (AIC) EXHIBITED A PURGE DISC NOT READING ALARM. THE STAFF TRIED TO TROUBLESHOOT; HOWEVER, IT WAS UNSUCCESSFUL. THE DEVICE WAS REPLACED WITH A NEW AIC. IT WAS REPORTED THAT THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH NO HARM TO PATIENT RELATED TO THIS MALFUNCTION. THE PROCEDURAL OUTCOME IS PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: B2: THE PATIENT'S DATE OF DEATH IS UNKNOWN. THIS DEVICE WAS RETURNED FOR EVALUATION. THE CONSOLES PURGE ASSEMBLY WAS INSPECTED WHERE IT WAS OBSERVED THAT THE PURGE DISC RETAINER HAD A LARGE FRACTURE IN IT BUT WAS STILL INTACT ON THE PURGE ASSEMBLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412965&pc=OZD | 1220648-2024-22736 | 2022-09-18 04:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS ON AN AUTOMATED IMPELLA CONTROLLER (AIC) FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED THAT THE CONSOLE WAS STUCK ON PRIMING DURING A CASSETTE CHANGE. THE DEVICE WAS REPLACED WITH ANOTHER AIC AND THE PROCEDURE WAS SUCCESSFUL. NO REPORT OF PATIENT HARM OR INJURY RELATED TO THIS MALFUNCTION. ADDITIONAL INFORMATION NOTED THE PATIENT WAS MADE A DO NOT RESUSCITATE, CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: B2: THE DATE OF PATIENT DEATH IS UNKNOWN. THE DEVICE WAS RETURNED FOR EVALUATION. CONSOLE WAS INSPECTED AND TESTED ON AN ENGINEERING LAB BENCH. GLUCOSE BUILDUP WAS OBSERVED IN THE PURGE INSERT TRAY BUT DID NOT IMPACT THE ABILITY TO RUN A KNOWN-GOOD (KG) PUMP AND PURGE (NO ISSUES OBSERVED). THE ROOT CAUSE OF THE PURGE FLUID NOT DETECTED ERROR WAS MOST LIKELY DUE TO IMPROPER CASSETTE PLACEMENT (USE ISSUE). THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412966&pc=OZD | 1220648-2024-22747 | 2022-05-08 04:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT NOTED WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE AUTOMATED IMPELLA CONTROLLER (AIC) EXPERIENCED CONTROLLER FAILURE (RED ALARM). ISSUE WAS RESOLVED BY SWITCHING TO BACKUP AIC. ADDITIONAL INFORMATION NOTED THAT THE CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED AIC - CONTROLLER FAILURE (RED ALARM) HAS BEEN COMPLETED. THE PRODUCT WAS RETURNED AND THE AIC - CONTROLLER FAILURE (RED ALARM) WAS CONFIRMED. THE CAUSE OF THE AIC ISSUE WAS MOST LIKELY DUE TO A DEFECTIVE PURGE ASSEMBLY, HOWEVER THE OVERALL ROOT CAUSE COULD NOT BE DETERMINED BECAUSE OF THE INABILITY TO REPRODUCE. THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21412969&pc=OZD | 1220648-2024-22744 | 2022-09-28 04:00 | Death | ABIOMED, INC. | 2025-02-19 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK ON AN AUTOMATED IMPELLA CONTROLLER (AIC) FOR MECHANICAL CIRCULATORY SUPPORT. THE COMPLAINANT REPORTED THAT THE AUTOMATED IMPELLA CONTROLLER (AIC) EXHIBITED A CONTROLLER ERROR ALARM. THE DEVICE WAS REPLACED WITH ANOTHER AIC, THE PROCEDURE WAS SUCCESSFUL WITH NO REPORTED HARM TO PATIENT REGARDING THIS MALFUNCTION. ADDITIONAL INFORMATION NOTED THE PATIENT WAS MADE A DO NOT RESUSCITATE, CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: B2: THE DATE OF THE PATIENT DEATH IS UNKNOWN. THE CAUSE OF THE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413024&pc=MIH | 3007284313-2025-03779 | 2025-01-22 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-19 05:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Insufficient Information | Insufficient Information | P020004 |  | 1 | Event Description: THE LONG-TERM EVAR ASSESSMENT AND FOLLOW-UP (LEAF) RETROSPECTIVE STUDY PROVIDED AN INITIAL UPDATE REGARDING THE PROCEDURAL OUTCOMES OF CURRENTLY ENROLLED PARTICIPANTS FOLLOWING IMPLANT OF GORE® EXCLUDER® AAA ENDOPROSTHESIS. THIS EVENT COVERS THE DATA REPORTED FROM THE VASCULAR QUALITY INITIATIVE (VQI) SITES AND LINKED TO MEDICARE FEE-FOR-SERVICE CLAIMS DATA. THE INITIAL REPORT COVERED 21,297 PATIENTS WHO ENROLLED IN THE REGISTRY BETWEEN JANUARY 1, 2014 AND DECEMBER 31, 2022. 393 PATIENTS HAD DEATHS REPORTED WHILE STILL IN THE HOSPITAL POST-PROCEDURE. THERE IS NO INFORMATION RELATED TO THE CAUSE OF DEATH. A KAPLAN-MEIER GRAPH WAS INCLUDED IN THE REPORT SHOWING FREEDOM FROM ANEURYSM-RELATED MORTALITY. AT THE 5 YEAR FOLLOW-UP THE KAPLAN MEIER ESTIMATE WAS THAT APPROXIMATELY 5% OF SUBJECTS HAD ANEURYSM-RELATED MORTALITY. Manufacturer Narrative: THE PATIENTS REFERENCED IN THIS EVENT FILE ARE ENROLLED IN THE COLLABORATIVE LONG-TERM EVAR ASSESSMENT AND FOLLOW-UP (LEAF) POST-MARKET SURVEILLANCE PROJECT THROUGH THE SOCIETY FOR VASCULAR SURGERY VASCULAR QUALITY INITIATIVE (SVS VQI). THE VQI IS A COLLABORATIVE OF REGIONAL QUALITY GROUPS COLLECTING AND ANALYZING DATA IN AN EFFORT TO IMPROVE PATIENT CARE. THE VQI COLLECTS PERIOPERATIVE AND FOLLOW-UP DATA. THE VQI IS GOVERNED BY THE SVS PATIENT SAFETY ORGANIZATION (SVS PSO), WHICH PROVIDES OVERSIGHT OF DATA SHARING ARRANGEMENTS, KEY OUTCOME AND QUALITY MEASURE ANALYSES, AND DISSEMINATION OF INFORMATION TO PARTICIPATING PROVIDERS. ALL INFORMATION CONTAINED IN THIS EVENT FILE WAS RECEIVED FROM THE VQI DATA BASE. VQI IS NOT MANAGED, MAINTAINED OR SUPERVISED BY GORE OR ANY REPRESENTATIVE OF GORE. THE INITIAL INFORMATION OBTAINED FROM VQI IS THE ONLY INFORMATION BEING PROVIDED. FURTHER INVESTIGATION IS NOT POSSIBLE AS IDENTIFIABLE DATA AS TO SITE(S), PHYSICIAN(S), PATIENT(S), DEVICE(S) AND/OR SPECIFIC DATE(S) ARE NOT BEING PROVIDED. THE COMPLAINT FILE WILL REFLECT LEAF AS THE STUDY NAME. IT IS LIKELY THESE INDIVIDUAL EVENTS HAVE BEEN PREVIOUSLY CAPTURED AND REPORTED BY GORE, BUT DUE TO THE NATURE OF THIS STUDY AND ANONYMITY OF THE PATIENT AND DEVICE DATA, GORE IS UNABLE TO CONFIRM THIS. FOR THIS REASON, GORE WILL SUBMIT ONE REPORT FOR MULTIPLE REPORTABLE EVENTS WHERE THE INDIVIDUALLY IDENTIFIABLE INFORMATION OF THE DEVICE IS NOT AVAILABLE AND NO ADDITIONAL INFORMATION WILL BE PROVIDED BECAUSE WE CANNOT OBTAIN ENOUGH INFORMATION ABOUT EACH IDENTIFIED PATIENT AND/OR DEVICE MENTIONED IN ORDER TO PROVIDE A COMPLETE REPORT FOR EACH REPORTABLE EVENT. THE REGISTRIES ARE LIKELY TO CONTAIN DEVICES FOR THE GORE® EXCLUDER® AAA ENDOPROSTHESIS, THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS, AND POTENTIALLY THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS. THE DATA GIVEN DID NOT DIFFERENTIATE WHICH OF THE SPECIFIC DEVICES CONTRIBUTED TO THE SERIOUS INJURIES. THEREFORE, WE REPORT ALL EVENTS FOR ALL DEVICES ON THE SAME REPORT. DEVICE BRAND NAME: GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS PMA NUMBER: P200030. DEVICE BRAND NAME: GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS PMA NUMBER: P020004. A.2. AGE AT THE TIME OF EVENT: THE MEAN AGE FOR THE SUBJECTS IN THIS DATASET IS 75.7 YEARS. THEREFORE, 76 YEARS HAS BEEN USED AS THE ESTIMATED PATIENT AGE. A.3. GENDER: THE MAJORITY OF SUBJECTS IN THIS DATASET ARE MALE. THEREFORE, MALE HAS BEEN USED AS THE ESTIMATED PATIENT GENDER. A.4. PATIENT WEIGHT: THE AVERAGE WEIGHT OF THE PATIENTS WAS NOT GIVEN IN THE REPORTS. A.5. / A.6. PATIENT ETHNICITY / RACE: THE MAJORITY OF SUBJECTS IN THIS DATASET HAVE AN ETHNICITY LISTED AS "NOT HISPANIC OR LATINO" AND AN IDENTIFIED RACE OF "WHITE". THEREFORE, THESE HAVE BEEN USED AS THE ESTIMATED PATIENT ETHNICITY AND RACE. B.3. DATE OF EVENT: AS THE INDEX PROCEDURE DATE(S) AND DATE OF EVENT(S) REMAIN UNAVAILABLE, THE DATE OF EVENT HAS BEEN ESTIMATED AS THE DATE GORE AWARE OF THE INITIAL DATA EXPORT: 22-JAN-2025. D.4. DEVICE INFORMATION: THE DEVICE LOT/SERIAL NUMBER(S) REMAINS UNAVAILABLE. THEREFORE, DEVICE INFORMATION REMAINS UNKNOWN. D.10. CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: THIS INFORMATION REMAINS UNAVAILABLE TO GORE. H.4. DEVICE MANUFACTURE DATE: AS THE DEVICE LOT/SERIAL NUMBERS ARE UNAVAILABLE, THE DEVICE MANUFACTURE DATES REMAINS UNKNOWN. H.6. TYPE OF INVESTIGATION: CODE B22 - AS THE DEVICE LOT/SERIAL NUMBERS ARE UNAVAILABLE, NO DEVICE HISTORY RECORD REVIEW WAS ABLE TO BE COMPLETED. H.6. INVESTIGATION CONCLUSIONS CODE D12: IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE, COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE® EXCLUDER® AAA ENDOPROSTHESIS THAT MAY OCCUR INCLUDE, BUT ARE NOT LIMITED TO: DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413667&pc=CBK | 2518422-2025-102159 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-02-19 05:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A PATIENT HAD AN EVENT. ADDITIONAL INFORMATION REQUESTED AND REPORTER PROVIDED INFORMATION ON 28-JAN-2025, VIA EMAIL, THAT PATIENT WAS DECEASED. THE REPORTED EVENT PROVIDED IN GOOD FAITH EFFORTS STATING THE PATIENT WAS DECEASED MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USER ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413718&pc=MVK | 3008642652-2025-01792 | 2025-01-22 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-19 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 06:21:57 ON (B)(6) 2025. AT 09:30:08, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 09:30:44, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT/INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 09:31:11, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 09:31:38, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 09:32:06, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 09:32:54 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413727&pc=JDI | 1038671-2025-01095 | 2015-09-01 04:00 | Death | EXACTECH, INC. | 2025-02-19 05:00 | JDI | UNKNOWN HIP COMPONENTS | Insufficient Information | Unspecified Infection |  |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT'S HIP BECAME INFECTED. PATIENT PASSED AWAY. IT WAS ALSO NOTED BY THE PATIENT THAT THEY RECEIVED A HIP REPLACEMENT AS SURGICAL TREATMENT. IT IS UNKNOWN IF THEY INTENDED THE INITIAL PROCEDURE OR IF THEY UNDERWENT A REVISION PROCEDURE. NO FURTHER INFORMATION AVAILABLE. Manufacturer Narrative: THE REPORTED EVENT WAS UNABLE TO BE CONFIRMED DUE TO LIMITED INFORMATION RECEIVED FROM THE CUSTOMER. NO DEVICE WAS RETURNED FOR EVALUATION; FURTHER, PHOTOGRAPHS AND/OR RADIOGRAPH IMAGES WERE NOT PROVIDED FOR REVIEW. OPERATIVE NOTES AND/OR MEDICAL RECORDS WERE NOT PROVIDED FOR REVIEW OF USAGE/ TECHNIQUE. A REVIEW OF MANUFACTURING DATA WAS UNABLE TO BE PERFORMED AS THE LOT INFORMATION OF THE PRODUCT INVOLVED IN THE EVENT WAS NOT AVAILABLE. A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AS THE NECESSARY INFORMATION TO ADEQUATELY INVESTIGATE THE REPORTED EVENT WAS NOT PROVIDED. IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY CORRECTED INFORMATION. IMPACT CODE AND COMPONENT CODE. THE FOLLOWING SECTIONS HAVE ADDITIONAL INFORMATION: E1 (EMAIL), H6 TYPE OF INVESTIGATION IF ANY FURTHER INFORMATION IS OBTAINED THAT WOULD CHANGE OR ALTER ANY INFORMATION PROVIDED, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413731&pc=MVK | 3008642652-2025-01795 | 2025-01-18 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-19 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE NON-LIFEVEST DEFIBRILLATIONS. THE DEVICE WAS STARTED UP AT 07:47:13 ON (B)(6) 2025. AT 22:46:55 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SVT @ 180 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF TRANSITIONING TO VT @ 210 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VT. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. BETWEEN 22:48:09 AND 22:56:14, THE PATIENT RECEIVED FIVE NON-LIFEVEST DEFIBRILLATIONS. THE PATIENT WAS LAST SEEN IN VF WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF AT THE TIME OF THE ELECTRODE BELT DISCONNECTION. THE ELECTRODE BELT WAS DISCONNECTED AT 23:01:13 ON (B)(6) 2025. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413751&pc=MVK | 3008642652-2025-01794 | 2025-01-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-19 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 08:39:29 ON (B)(6) 2025. AT 22:26:28, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT. AT 22:27:43, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT FOR 28 SECONDS TRANSITIONING TO AF @ 90 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413854&pc=NPT | 2025587-2025-01325 | 2024-02-13 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-19 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Anemia; Stroke/CVA; Ambulation Difficulties | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WOKE UP APPROXIMATELY 30 MINUTES AFTER THEIR USUAL TIME AND WERE FOUND BY FAMILY TO BE EXPERIENCING STANDING DIFFICULTIES. THE PATIENT WAS TRANSPORTED BY EMERGENCY SERVICES. DIFFUSION WEIGHTED IMAGING REVEALED HIGH SIGNAL IN BOTH CEREBELLAR HEMISPHERES, LOW SIGNAL IN THE SAME AREA ACCORDING TO THE APPARENT DIFFUSION COEFFICIENT, AND NO MICROBLEEDS ON TRANSVERSE RELAXATION WEIGHTED IMAGING. OF NOTE, THE PATIENT WAS TAKING 30 MG OF ANTICOAGULANT MEDICATION AT THE TIME OF ONSET AND THROMBOLYTIC THERAPY WAS NOT APPLICABLE. PROGRESSION OF ANEMIA WAS NOTED ALONG WITH BLACK STOOL, AND INCREASED BUN AND CREATININE. GASTROINTESTINAL ENDOSCOPY SHOWED NO FINDINGS OF GASTROINTESTINAL BLEEDING, HOWEVER. THE PATIENT'S CONDITION WAS STABILIZED AND THE PATIENT WAS TRANSFERRED FOR REHABILITATION. APPROXIMATELY 11 MONTHS LATER, AFTER HOSPITALIZATION AT MULTIPLE FACILITIES, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21413997&pc=PBN | 2435151-2025-00001 | 2022-06-05 04:00 | Death | KANEKA CORPORATION | 2025-02-19 05:00 | PBN | LIPOSORBER LA-15 SYSTEM | Adverse Event Without Identified Device or Use Problem | Pulmonary Edema |  |  | 1 | Event Description: DEATH OF AN 8-YEAR-OLD FSGS PATIENT WAS REPORTED. (MFR#3002808904-2025-00005). THIS WAS AN 8-YEAR-OLD GIRL WHO WAS DIAGNOSED WITH NEPHROTIC SYNDROME WITH ACUTE KIDNEY INJURY AND HYPERTENSION ON (B)(6) 2022. A RENAL BIOPSY ON (B)(6) 2022 SHOWED COLLAPSING GLOMERULOPATHY. SHE REQUIRED INITIATION OF HEMODIALYSIS ON (B)(6) 2022 FOR UREMIA AND FLUID OVERLOAD. SHE REMAINED DIALYSIS DEPENDENT. HEMODIALYSIS WAS PERFORMED AT DAVITA 3X / WK ON MONDAYS, WEDNESDAYS, AND FRIDAYS UNDER THE DIRECTION OF CEDARS PEDIATRIC NEPHROLOGISTS. SHE WAS RESISTANT TO STEROIDS AND RITUXIMAB. CALCINEURIN INHIBITORS WERE NOT GIVEN DUE TO POOR RENAL FUNCTION. LDL APHERESIS WAS STARTED WITHOUT PROBLEMS FOR REFRACTORY FSGS ON (B)(6) 2022. ON (B)(6) 2022, SHE UNDERWENT THE 10TH TREATMENT OUT OF 12 TREATMENTS WITHOUT PROBLEMS. SHE WAS GIVEN SOLUMEDROL PER PROTOCOL. SHE RECEIVED HEMODIALYSIS TREATMENTS AT DAVITA ON (B)(6) 2022. THEY WERE UNEVENTFUL. THE NEXT HEMODIALYSIS WAS SCHEDULED ON (B)(6) 2022. ON (B)(6) 2022 AT 1:25 AM, THE PATIENT PRESENTED AT KP LAMC ED FOR SHORTNESS OF BREATH. THE PARENT REPORTED THAT HE WAS GETTING THE PATIENT READY FOR BED AND NOTED THE INCREASED WORK OF BREATHING SO THE PATIENT WAS BROUGHT TO THE ED. HER O2 SATURATION WAS IN THE 40S ON ROOM AIR AND IMPROVED TO AS HIGH AS 79% ON A NONREBREATHER MASK THUS THE DECISION WAS MADE TO INTUBATE THE PATIENT. VISUALIZATION WAS DIFFICULT GIVEN THE GENEROUS OROPHARYNGEAL SOFT TISSUE AND COPIOUS BLOOD-TINGED THIN LIQUID POOLING IN THE OROPHARYNX. THIS WAS ABORTED AND SHE WAS AGAIN BAGGED WITH A RETURN TO ONLY ~60% SPO2. SHE BECAME BRADYCARDIC, SO CHEST COMPRESSIONS WERE INITIATED, AND THE FIRST DOSE OF EPINEPHRINE WAS GIVEN SOMETIME AROUND 1:45 AM. THE SECOND ATTEMPT WITH MAC 3 LARYNGOSCOPE BLADE WAS SUCCESSFUL AT FULL VISUALIZATION OF THE GLOTTIS AND CORDS AFTER SUCTIONING COPIOUS BLOOD-TINGED THIN LIQUID FROM THE OROPHARYNX, WITH CONTINUAL RECURRENT FILLING OF THE OP WITH FLUID FROM THE TRACHEA. HOWEVER, A 5.5 CUFFED AND STYLETED ET TUBE WAS UNABLE TO PASS THROUGH THE CORDS. THE PATIENT WAS BAGGED AGAIN AND ON 3RD ATTEMPT, AGAIN WITH MAC 3 BLADE ON C-MAC WITH DIRECT VIDEO VISUALIZATION, A 5.0 UNCUFFED ETT WAS OBSERVED PASSING THROUGH THE CORDS. BREATH SOUNDS WERE APPRECIATED BILATERALLY ON AUSCULTATION BY AN ED PHYSICIAN, HOWEVER END-TIDAL MONITORING DID NOT SHOW A WAVEFORM. A SUCTION CATHETER WAS PASSED DOWN THE ETT AND COPIOUS BLOOD-TINGED THIN LIQUID RETURNED AND POOLED IN THE SUCTION CANISTER. BRIEFLY FOLLOWING ETT SUCTION, O2 SATS IMPROVED TO THE 30S TO 40S AND END TIDAL MONITORING REGISTERED A NUMBER BUT NO WAVEFORM, BUT SATS AND END TIDAL QUICKLY FELL AGAIN TO EVENTUALLY UNDETECTABLE DESPITE SUCTIONING COPIOUS FLUID FROM THE ETT PERIODICALLY THROUGHOUT THE CODE EVENT. FLUID ALSO POURED OUT OF THE PATIENT'S NOSE AND MOUTH THROUGHOUT THE CODE EVENT. CPR WITH CHEST COMPRESSIONS AND 7 ROUNDS OF EPINEPHRINE WERE GIVEN. A DOSE OF CALCIUM CHLORIDE WAS ALSO GIVEN DURING THE 30-MINUTE CODE, 2 DOSES OF SODIUM BICARBONATE, AND INSULIN AND A GLUCOSE BOLUS WERE ALSO GIVEN, ALTHOUGH WAS FSBG OF WELL OVER 100. AT AROUND 25 MINUTES INTO THE EVENT, FRANK BLOOD RETURNED FROM THE ETT AND BEGAN TO POUR FROM THE PATIENT'S MOUTH AND NOSE. AT NO TIME DURING REGULAR PULSE CHECKS WAS THERE A PALPABLE EFFECTIVE PULSE DETECTED. AT 2:15 AM, DURING THE LAST PULSE CHECK, THE MONITOR SHOWED PEA, THERE WAS NO PALPABLE PULSE, O2 SATS WERE UNDETECTABLE, AND THE PATIENT WAS CYANOTIC CENTRALLY. DEATH WAS PRONOUNCED AT 2:15AM. SHE WAS THOUGHT TO HAVE CARDIORESPIRATORY FAILURE DUE TO ACUTE ONSET PULMONARY EDEMA, CAUSE UNKNOWN. AUTOPSY WAS DECLINED. MEDICAL DIRECTOR REPORTED THAT THE PT HAD NO COMPLICATIONS DURING THE LIPOSORBER TREATMENT AND DOES NOT BELIEVE PT'S PASSING WAS DUE TO LIPOSORBER TREATMENT. Manufacturer Narrative: BELONGS, PASSED ALL IN-PROCESS INSPECTIONS TEST. NO NONCONFORMITY OR ABNORMALITY IN THE MANUFACTURING PROCESSES OF THE DEVICE CONCERNED WAS FOUND. THE ACTUAL DEVICE WAS NOT RETURNED AND COULD NOT BE INVESTIGATED. TEN TREATMENTS WITH OUR PRODUCT WERE PERFORMED AND COMPLETED WITHOUT PROBLEMS AND WITHOUT PRODUCT FAILURE. AFTER THAT, TWO HEMODIALYSIS TREATMENTS USING ANOTHER COMPANY'S PRODUCT WERE ALSO COMPLETED WITHOUT ANY PROBLEMS. ALSO, WE RECEIVED A COMMENT FROM THE DOCTOR IN CHARGE THAT THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND OUR PRODUCT. AS A RESULT OF THE ABOVE, WE THINK THAT THE ADVERSE EVENT WAS CAUSED BY THE PATIENT'S CONDITION AND THAT THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND OUR PRODUCT. REASON FOR SUBMISSION OF THIS MDR: WE PERFORMED A DETAILED INVESTIGATION OF THE PATIENT'S MEDICAL HISTORY AND HOSPITAL TREATMENT PROCESS AND COLLECTED SUFFICIENT ADDITIONAL INFORMATION TO CONFIRM THAT THERE WAS NO DEVICE FAILURE, NO PROBLEM WITH THE DHR, AND THAT A CAUSAL RELATIONSHIP WAS RULED OUT AS A PHYSICIAN'S COMMENT IN CHARGE. AS THE MANUFACTURER, WE HAD DETERMINED ON 15 FEB., 2022, IN ACCORDANCE WITH OUR PROCEDURES, THAT THERE WAS NO CAUSAL RELATIONSHIP WITH OUR PRODUCT AND THAT THE MDR WAS NOT APPLICABLE. HOWEVER, AT THE FDA AUDIT RECEIVED ON 6 FEB., 2025, THE AUDITOR IN CHARGE ADVISED US TO SUBMIT AN MDR WITH THE STATEMENT "AS A MANUFACTURER, WE HAVE DETERMINED THAT THERE IS NO CAUSAL RELATIONSHIP" AND WE ARE SUBMITTING THIS MDR. AS THE IMPORTER, KANEKA MEDICAL AMERICA WILL SUBMIT DUE TO MANUFACTURER SUBMISSION PER 21CFR803.40(A). THIS SUBMISSION IS NOT AN ADMISSION OF CAUSE OR CONTRIBUTION BUT SENT OUT AN ABUNDANCE OF CAUTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414053&pc=LWR | 2135147-2025-00843 | 2024-12-04 05:00 | Death | ABBOTT MEDICAL | 2025-02-19 05:00 | LWR | TRIFECTA¿ | Adverse Event Without Identified Device or Use Problem | Angina; Cardiac Arrest; Low Blood Pressure/ Hypotension | P100029 |  | 1 | Event Description: IT WAS REPORTED THAT ON AN UNKNOWN DATE, A 25MM TRIFECTA VALVE WAS IMPLANTED. ON AN UNKNOWN DATE, A 23MM NON-ABBOTT VALVE WAS IMPLANTED WITHIN THE TRIFECTA VALVE VIA VALVE-IN-VALVE. POST PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE PACU. "A LITTLE TIME PASSED," THEN THE PATIENT REPORTED CHEST PAIN AND WAS BROUGHT BACK TO THE CATH LAB WHERE ANGIOGRAM WAS PERFORMED. "THE RIGHT CORONARY ARTERY (TCA) WAS THE CULPRIT" AND A STENT WAS PLACED IN THE RCA WITH REDUCTION OF CHEST PAIN. LATER IN THE EVENING, THE PATIENT AGAIN REPORTED CHEST PAIN AND SUBSEQUENT CABG GRAFT TO THE RCA WAS PERFORMED. THE NEXT DAY, THE PATIENT WAS REPORTED TO BE DOING OK. HOWEVER, "AT SOME POINT WHILE SITTING" IN A CHAIR, THE PATIENT EXPERIENCED LOSS OF PRESSURE AND PULSELESS ACTIVITY AND PASSED AWAY. THERE ARE NO ALLEGATIONS THAT THE IMPLANTED VALVE WAS DEFICIENT OR INADEQUATE. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF PATIENT DEATH AFTER IMPLANT WAS REPORTED. INFORMATION FROM FIELD INDICATED THAT THE PATIENT HAD CHEST PAIN. THE PATIENT WAS HOSPITALIZED AND PASSED AWAY PRIOR TO SURGICAL INTERVENTION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS NO OTHER INFORMATION WAS PROVIDED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE ARE NO ALLEGATIONS THAT THE IMPLANTED VALVE WAS DEFICIENT OR INADEQUATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414269&pc=NRY | 2029214-2025-00457 | 2024-12-28 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-19 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Ischemia Stroke |  |  | 1 | Event Description: LITERATURE REVIEWED: STUDY TITLE: "DEMOGRAPHIC CHARACTERISTICS AND TREATMENT OUTCOMES OF INTRACRANIAL ATHEROSCLEROSIS STENTING: A RETROSPECTIVE CASE-SERIES OF 216 CONSECUTIVE PATIENTS." OBJECTIVE: TO ANALYZE DEMOGRAPHIC CHARACTERISTICS, TREATMENT OUTCOMES, AND PROCEDURAL CHALLENGES AS SOCIATED WITH INTRACRANIAL ATHEROSCLEROSIS (ICAS) IN PATIENTS TREATED AT A SINGLE INSTITUTION. STUDY TIME FRAME: FROM JANUARY 2016 TO DECEMBER 2023. MANUFACTURERS: MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION, INCLUDING PRODUCTS FROM TERUMO CORPORATION, MEDTRONIC INC., BOSTON SCIENTIFIC, MERIL LIFE SCIENCES, AND OTHERS. MEDTRONIC DEVICES USED: RESOLUTE INTEGRITY, ONYX (19 CASES) AND SOLITAIRE (2 CASES). DEATHS: DEATHS OCCURRED IN THE STUDY POPULATION. CAUSES OF DEATH: TWO DEATHS: ONE DUE TO ISCHEMIC COMPLICATIONS AND THE OTHER FROM INTRAVENTRICULAR TAMPONADE, A HEMORRHAGIC COMPLICATION. DURING THE 72-H PERIPROCEDURAL PERIOD FOLLOWING THE PROCEDURE, TWO DEATHS FROM STROKES WERE REPORTED, RESULTING IN A 0.7% PERIPROCEDURAL COMPLICATION RATE. THESE FATALITIES OCCURRED IN PATIENTS WITH BASELINE STENOSIS OF 90% OR GREATER IN THE LEFT VERTEBRAL ARTERY. ONE DEATH WAS ATTRIBUTED TO ISCHEMIC COMPLICATIONS, WHILE THE OTHER RESULTED FROM INTRAVENTRICULAR TAMPONADE, CLASSIFIED AS A HEMORRHAGIC COMPLICATION. ADVERSE EVENTS: THE PERIPROCEDURAL COMPLICATION RATE WAS 0.7%. ADVERSE EVENTS INCLUDED STENT MIGRATION, FAILURE TO DEPLOY DUE TO VESSEL TORTUOSITY, IMMEDIATE THROMBOSIS, AND WORSENING MRS SCORES. NO FURTHER INFORMATION PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED NO MEDTRONIC PRODUCTS WERE USED IN THE CASES OF THE TWO PATEINT'S THAT EXPIRED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-ONYX (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID UNK-NV-ONYX (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: SARSHAYEV, M.; TURDALIYEVA, B.; TANBAYEVA, G.; MAKHANBETKHAN, S.; MUSSABEKOV, M.; DAVLETOV, D.; MAIDAN, A.; B ERDIKHOJAYEV, M.. DOI.ORG/ 10.3390/JCM14010125. J. CLIN. MED. 14, 125 2025. DOI: 10.3390/JCM14010125. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414394&pc=LOX | 2124215-2025-09978 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, CARDIOGENIC SHOCK; DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON. Manufacturer Narrative: DUE TO SAFE HARBOR DE IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414423&pc=LOX | 2124215-2025-09985 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, CARDIOGENIC SHOCK; DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414474&pc=LOX | 2124215-2025-09981 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, CARDIOGENIC SHOCK; DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414477&pc=LOX | 2124215-2025-09820 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE; CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414478&pc=NWX | 2124215-2025-10004 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414483&pc=NWX | 2124215-2025-10006 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST AND DEATH AT DISCHARGE. CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414487&pc=LOX | 2124215-2025-09827 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414494&pc=LOX | 2124215-2025-10019 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC TAMPONADE, CARDIOGENIC SHOCK, OTHER BLEED; DEATH DURING PROCEDURE, CAUSE: CARDIOVASCULAR HEMORRHAGE. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414502&pc=OOB | 2124215-2025-10065 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414504&pc=LOX | 2124215-2025-10039 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC TAMPONADE, CARDIOGENIC SHOCK, OTHER BLEED; DEATH DURING PROCEDURE, CAUSE: CARDIOVASCULAR HEMORRHAGE. Manufacturer Narrative: DUE TO SAFE HARBOR DE IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414515&pc=OOB | 2124215-2025-10073 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE; CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414557&pc=NWX | 2124215-2025-10044 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414572&pc=NKM | 2135147-2025-00849 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-19 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRICUSPID REGURGITATION AND OUTCOMES IN MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PROGNOSTIC VALUE OF TR BEFORE AND AFTER MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER). DEVICES MENTIONED INCLUDE MITRACLIIP. THE ARTICLE CONCLUDED THAT *CONCLUSION*. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS YOHEI OHNO AT TOKAI UNIVERSITY SCHOOL OF MEDICINE, 143 SHIMOKASUYA, ISEHARA, JAPAN; WITH CORRESPONDING EMAIL: YOHNO@TOKAI.AC.JP. THE TIME FRAME OF THE STUDY WAS BLANK TO BLANK. SAY NOT CONFIRMED OR REPORTED IF APPLICABLE. A TOTAL OF 3666 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDE OBESITY, HEART FAILURE, PRIOR HEART FAILURE HOSPITALIZATION, DIABETES, HYPERTENSION, MYOCARDIA INFARCTION/ANGINA, ATRIAL FBIRILLATION, VENTRICULAR TACHYCARDIA/FIBRILLATION, COPD, STROKE, DIALYSIS, LIVER CIRRHOSIS, MITRAL REGURGITATION, TRICUSPID REGURGITATION, ANTICOAGULANT MEDICATION, HEART FAILURE MEDICATION PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED CARDIOVASCULAR DEATH, ALL CAUSE DEATH, HEART FAILURE, HOSPITALIZATION. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "TRICUSPID REGURGITATION AND OUTCOMES IN MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR." SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING OBESITY, HEART FAILURE, PRIOR HEART FAILURE HOSPITALIZATION, DIABETES, HYPERTENSION, MYOCARDIAL INFARCTION/ANGINA, ATRIAL FIBRILLATION, VENTRICULAR TACHYCARDIA/FIBRILLATION, COPD, STROKE, DIALYSIS, LIVER CIRRHOSIS, MITRAL REGURGITATION, TRICUSPID REGURGITATION, ANTICOAGULANT MEDICATION, HEART FAILURE MEDICATION. COMPLICATIONS REPORTED INCLUDED CARDIOVASCULAR DEATH, ALL CAUSE DEATH, HEART FAILURE, HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414732&pc=NIK | 2017865-2025-14521 | 2025-01-28 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-19 05:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Respiratory Failure | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14519 AND 2017865-2025-14520. IT WAS REPORTED FOLLOWING IMPLANT PROCEDURE THAT THE PATIENT WENT INTO RESPIRATORY DISTRESS. PATIENT BECAME ACIDOTIC, HEART RATE SLOWED, AND EVENTUALLY THE PATIENT ARRYTHMIA WENT INTO TORSADES DE POINTES. ANESTHESIOLOGIST WAS UNABLE TO RESTORE NORMAL RESPIRATORY FUNCTION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTION: COMPONENT CODE UPDATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414754&pc=NKE | 2017865-2025-14519 | 2025-01-28 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-19 05:00 | NKE | QUADRA ALLURE MP CRT-P | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Respiratory Failure | P030035 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14520, 2017865-2025-14521. IT WAS REPORTED FOLLOWING IMPLANT PROCEDURE THAT THE PATIENT WENT INTO RESPIRATORY DISTRESS. PATIENT BECAME ACIDOTIC, HEART RATE SLOWED, AND EVENTUALLY THE PATIENT ARRYTHMIA WENT INTO TORSADES DE POINTES. ANESTHESIOLOGIST WAS UNABLE TO RESTORE NORMAL RESPIRATORY FUNCTION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTION: COMPONENT CODE UPDATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414779&pc=DSQ | 2916596-2025-00940 | 2024-12-11 05:00 | Death | THORATEC CORPORATION | 2025-02-19 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN EXTENSIVE DRIVELINE INFECTION THAT LED TO SEPSIS. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE PATIENT DIED FROM DISSEMINATED METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS (MSSA) (BLOOD, URINE, DRIVELINE SITE, PACEMAKER, AND TOE) AFTER A FAILED TOE AMPUTATION. THE PATIENT HAD NO DRIVELINE SYMPTOMS PRIOR TO THIS. IT WAS ASSUMED THAT THE DISSEMINATED MSSA CONTAMINATED THE DRIVELINE. CULTURES WERE OBTAINED AFTER PRESENTATION WITH SEPSIS ON (B)(6) 2024 AND IT WAS MSSA. THIS WAS TREATED WITH MULTIPLE INTRAVENOUS (IV) AND ORAL ANTIBIOTICS: RIFAMPIN, CEFAZOLIN, ZOSYN (PIPERACILLIN/TAZOBACTAM), VANCOMYCIN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21414995&pc=MFE | 2029214-2025-00458 | 2024-12-28 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-02-19 05:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Ischemia Stroke |  |  | 1 | Event Description: HERE'S A DETAILED ANALYSIS OF THE STUDY FROM THE PDF YOU PROVIDED: LITERATURE REVIEWED: STUDY TITLE: "DEMOGRAPHIC CHARACTERISTICS AND TREATMENT OUTCOMES OF INTRACRANIAL ATHEROSCLEROSIS STENTING: A RETROSPECTIVE CASE-SERIES OF 216 CONSECUTIVE PATIENTS." OBJECTIVE: TO ANALYZE DEMOGRAPHIC CHARACTERISTICS, TREATMENT OUTCOMES, AND PROCEDURAL CHALLENGES ASSOCIATED WITH INTRACRANIAL ATHEROSCLEROSIS (ICAS) IN PATIENTS TREATED AT A SINGLE INSTITUTION. (PAGE 1) STUDY TIME FRAME: FROM JANUARY 2016 TO DECEMBER 2023. (PAGE 3) MANUFACTURERS: MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION, INCLUDING PRODUCTS FROM TERUMO CORPORATION, MEDTRONIC INC., BOSTON SCIENTIFIC, MERIL LIFE SCIENCES, AND OTHERS. (PAGE 3) MEDTRONIC DEVICES USED: RESOLUTE INTEGRITY, ONYX (19 CASES) AND SOLITAIRE (2 CASES). (PAGE 3) DEATHS: DEATHS OCCURRED IN THE STUDY POPULATION. (PAGE 2) CAUSES OF DEATH: TWO DEATHS: ONE DUE TO ISCHEMIC COMPLICATIONS AND THE OTHER FROM INTRAVENTRICULAR TAMPONADE, A HEMORRHAGIC COMPLICATION. (PAGE 3) DURING THE 72-H PERIPROCEDURAL PERIOD FOLLOWING THE PROCEDURE, TWO DEATHS FROM STROKES WERE REPORTED, RESULTING IN A 0.7% PERIPROCEDURAL COMPLICATION RATE. THESE FATALITIES OCCURRED IN PATIENTS WITH BASELINE STENOSIS OF 90% OR GREATER IN THE LEFT VERTEBRAL ARTERY. ONE DEATH WAS ATTRIBUTED TO ISCHEMIC COMPLICATIONS, WHILE THE OTHER RESULTED FROM INTRAVENTRICULAR TAMPONADE, CLASSIFIED AS A HEMORRHAGIC COMPLICATION. ADVERSE EVENTS: THE PERIPROCEDURAL COMPLICATION RATE WAS 0.7%. ADVERSE EVENTS INCLUDED STENT MIGRATION, FAILURE TO DEPLOY DUE TO VESSEL TORTUOSITY, IMMEDIATE THROMBOSIS, AND WORSENING MRS SCORES. (PAGE 2, PAGE 3) NO FURTHER INFORMATION PERTAINING TO MEDTRONIC PRODUCTS:Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED NO MEDTRONIC PRODUCTS WERE USED IN THE CASES OF THE TWO PATEINT'S THAT EXPIRED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID UNK-NV-ONYX (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: SARSHAYEV, M.; TURDALIYEVA, B.; TANBAYEVA, G.; MAKHANBETKHAN, S.; MUSSABEKOV, M.; DAVLETOV, D.; MAIDAN, A.; B ERDIKHOJAYEV, M.. DOI.ORG/ 10.3390/JCM14010125. J. CLIN. MED. 14, 125 2025. DOI: 10.3390/JCM14010125 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415206&pc=OOB | 2124215-2025-10081 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, CARDIOGENIC SHOCK; DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415214&pc=LOX | 2124215-2025-10054 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC TAMPONADE, CARDIOGENIC SHOCK, OTHER BLEED; DEATH DURING PROCEDURE, CAUSE: CARDIOVASCULAR HEMORRHAGE. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415217&pc=OOB | 2124215-2025-10045 | 2024-09-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation of Vessels |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415222&pc=LOX | 2124215-2025-10068 | 2024-08-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415224&pc=MCX | 2124215-2025-10107 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | MCX | ROTALINK PLUS | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, CARDIOGENIC SHOCK; DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415225&pc=LOX | 2124215-2025-10071 | 2024-08-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415226&pc=NWX | 2124215-2025-10082 | 2024-09-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415227&pc=NIQ | 2124215-2025-10078 | 2024-08-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415232&pc=LOX | 2124215-2025-10055 | 2024-08-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415243&pc=NIQ | 2124215-2025-09833 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE. CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415250&pc=LOX | 2124215-2025-10079 | 2024-09-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415254&pc=NWX | 2124215-2025-10075 | 2024-08-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415257&pc=LOX | 2124215-2025-10063 | 2024-08-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415349&pc=LOX | 2124215-2025-10048 | 2024-09-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-19 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation of Vessels |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: OTHER CARDIOVASCULAR REASON. Manufacturer Narrative: THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415420&pc=NVN | 2017865-2025-14520 | 2025-01-28 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-19 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Respiratory Failure | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-14519 AND 2017865-2025-14521. IT WAS REPORTED FOLLOWING IMPLANT PROCEDURE THAT THE PATIENT WENT INTO RESPIRATORY DISTRESS. PATIENT BECAME ACIDOTIC, HEART RATE SLOWED, AND EVENTUALLY THE PATIENT ARRYTHMIA WENT INTO TORSADES DE POINTES. ANESTHESIOLOGIST WAS UNABLE TO RESTORE NORMAL RESPIRATORY FUNCTION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTION: COMPONENT CODE UPDATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21415921&pc=MVK | 3008642652-2025-01811 | 2025-02-04 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-02-19 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 08:06:44 ON (B)(6) 2025. AT 19:29:45, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 19:30:35, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 19:38:06 ON 2/4/2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21417590&pc=QQK | 3011560054-2025-00021 | 2025-01-08 05:00 | Death | ORGANOX LIMITED | 2025-02-19 05:00 | QQK | ORGANOX METRA | Restricted Flow rate | Liver Damage/Dysfunction; Sepsis | P200035 |  | 1 | Event Description: IT WAS REPORTED THAT DURING PERFUSION THERE WAS DIMINISHED ARTERIAL FLOW AND IMPACTS TO VENOUS OUTFLOW IN THE RIGHT LOBE OF THE LIVER. THE DEVICE USER (DU) VERIFIED THAT THE FLOW WAS LOW BY USING AN EXTERNAL FLOW PROBE TO DIRECTLY MEASURE ARTERIAL FLOW. THE DU NOTED THAT THE NEW CANNULA IN THE DISPOSABLE SET WAS MUCH LESS FORGIVING IN TERMS OF ITS PLACEMENT IN THE HEPATIC VEINS. AFTER THE LIVER WAS TRANSPLANTED, THE PATIENT EXPERIENCED BAD GRAFT DYSFUNCTION AND SIGNIFICANT LIVER NECROSIS.Event Description: IT WAS REPORTED THAT DURING PERFUSION THERE WAS DIMINISHED ARTERIAL FLOW AND IMPACTS TO VENOUS OUTFLOW IN THE RIGHT LOBE OF THE LIVER. THE DEVICE USER (DU) VERIFIED THAT THE FLOW WAS LOW BY USING AN EXTERNAL FLOW PROBE TO DIRECTLY MEASURE ARTERIAL FLOW. THE DU NOTED THAT THE NEW CANNULA IN THE DISPOSABLE SET WAS MUCH LESS FORGIVING IN TERMS OF ITS PLACEMENT IN THE HEPATIC VEINS. AFTER THE LIVER WAS TRANSPLANTED, THE PATIENT EXPERIENCED BAD GRAFT DYSFUNCTION AND SIGNIFICANT LIVER NECROSIS. SUBSEQUENTLY, THE PATIENT EXPERIENCED SEPSIS WHICH ULTIMATELY LED TO PATIENT DEATH.Event Description: IT WAS REPORTED THAT DURING PERFUSION THERE WAS DIMINISHED ARTERIAL FLOW AND IMPACTS TO VENOUS OUTFLOW IN THE RIGHT LOBE OF THE LIVER. THE DEVICE USER (DU) VERIFIED THAT THE FLOW WAS LOW BY USING AN EXTERNAL FLOW PROBE TO DIRECTLY MEASURE ARTERIAL FLOW. THE DU NOTED THAT THE NEW CANNULA IN THE DISPOSABLE SET WAS MUCH LESS FORGIVING IN TERMS OF ITS PLACEMENT IN THE HEPATIC VEINS. AFTER THE LIVER WAS TRANSPLANTED, THE PATIENT EXPERIENCED BAD GRAFT DYSFUNCTION AND SIGNIFICANT LIVER NECROSIS. SUBSEQUENTLY, THE PATIENT EXPERIENCED SEPSIS WHICH ULTIMATELY LED TO PATIENT DEATH. Manufacturer Narrative: THE ROOT CAUSE INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS AND ADDITIONAL DETAILS ARE BEING SOUGHT FROM THE DEVICE USER. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: B2, B5, H1, AND H6 WERE UPDATED BECAUSE THE RECIPIENT OF THE DONOR LIVER DIED DUE TO SEPSIS. ADDITIONAL INFORMATION SHALL BE REQUESTED TO HELP DETERMINE IF THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: NO DEVICE AND/OR ADDITIONAL INFORMATION IS AVAILABLE FOR EVALUATION. INTERNAL REVIEW OF THE REPORTED EVENT ISOLATED THE REPORTED ISSUE TO COMPLAINTS INITIATED FROM ONE SAME SITE. ADDITIONAL TRAINING WAS COMPLETED BY THE SITE. TO DATE, THE CENTER HAS HAD NO SUBSEQUENT ISSUES. BASED ON AVAILABLE INFORMATION, CAUSE OF THE LOW FLOW COULD NOT BE DEFINITIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21392081&pc=LWR | MW5166215 | 2023-03-17 04:00 | Death | ABBOTT / ST. JUDE MEDICAL | 2025-02-18 05:00 | LWR | TRIFECTA GT TISSUE HEART VALVE |  | Insufficient Information |  |  | 1 | Event Description: ABBOTT TRIFECTA HEART VALVE MALFUNCTIONED AND WAS THE CAUSE OF MY MOM HAVING TO HAVE A SECOND OPERATION THAT LED TO HER DEATH. WHEN SHE WENT IN FOR A DIVERTICULITIS SCAN IT WAS THAT DOCTOR THAT SEEN IT WASN'T WORKING. THE HEART DOCTOR NEVER INFORMED HER OF ANYTHING. THE DEVICE HAD BEEN ON SOME TYPE OF LIST WHERE DR'S SHOULD HAVE WARNED THEIR PATIENTS. UNLESS THE FDA FAILED TO GIVE THIS INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21392112&pc=LYJ | 1644487-2025-00174 | 2024-11-25 05:00 | Death | CYBERONICS - HOUSTON | 2025-02-18 05:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: THE PHYSICIAN HAD REPORTED THAT ONE OF THEIR PREVIOUS PATIENTS HAD PASSED AWAY. NO REASON OR CAUSE WAS PROVIDED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21392230&pc=NWX | 2124215-2025-09025 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE; CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21392786&pc=DQY | 2124215-2025-07686 | 2025-02-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | DQY | WATCHMAN TRUSTEER ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Hypovolemic Shock; Perforation of Vessels; Retroperitoneal Hemorrhage | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC), A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND A 35MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED USING A WATCHMAN TRUSTEER SHEATH (WAS). THERE WERE NO RECAPTURES OF THE CLOSURE DEVICE NEEDED AND WAS IMPLANTED AFTER IT MET RELEASE CRITERIA. IT WAS NOTED THERE WERE NO ISSUES OR ABNORMALITIES USING THE WAS AND IT REQUIRED MINIMAL STEERING. THE PROCEDURE WAS COMPLETED WITH NO KNOWN ISSUES. AFTER THE PROCEDURE, WHEN THE PATIENT ARRIVED TO POST ANESTHESIA CARE UNIT (PACU), THE PATIENT CODED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS DISCOVERED THAT THE PATIENT HAD A RETROPERITONEAL BLEED CAUSED BY TWO HOLES AT THE INFRARENAL INFERIOR VENA CAVA (IVC). ONE OF THE HOLES WAS SUTURED AND THE SECOND HOLE (WHICH WAS APPROXIMATELY A CENTIMETER LACERATION) WAS UNABLE TO BE REPAIRED. THE PATIENT HAD SEVERAL UNKNOWN COMORBIDITIES THAT CONTRIBUTED TO HIS CONDITION. THE PHYSICIAN EXPERIENCED NO DIFFICULTIES WITH ACCESS, REMOVAL, OR CLOSURE WITH A NON-BOSTON SCIENTIFIC PRE-CLOSE SUTURE DEVICE. IN THE PHYSICIAN OPINION, THE LACERATIONS ARE FROM AN UNKNOWN CAUSE. THE PATIENT EXPERIENCED NO HEMODYNAMIC CHANGES DURING THE PROCEDURE AND ONLY EXHIBITED ISSUES AFTER LEAVING THE PROCEDURE ROOM. THE OFFICIAL CAUSE OF DEATH DETERMINED WAS HEMORRHAGIC SHOCK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21393139&pc=DSQ | 2916596-2025-00723 | 2025-01-20 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Renal Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. AFTER IMPLANT, THE PATIENT REQUIRED RIGHT VENTRICULAR ASSIST DEVICE (RVAD) AND CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) SUPPORT. IN SPITE OF THE SUPPORT, THE PATIENT DID NOT MAINTAIN HEMODYNAMICS AND WAS DECLARED. THE DEATH WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THAT THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RENAL DYSFUNCTION AND RIGHT HEART FAILURE) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21393148&pc=NPT | 2135147-2025-00787 | 2019-11-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-18 05:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "SEX-SPECIFIC ANATOMIC DIFFERENCES IN PATIENTS UNDERGOING TRANSCATHETER AORTIC VALVE IMPLANTATION: INSIGHTS FROM THE ST-TAVI REGISTRY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO DETERMINE SEX-SPECIFIC ANATOMIC CHARACTERISTICS OF PATIENTS UNDERGOING EVALUATION FOR TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, TO DETERMINE SEX-SPECIFIC PROPORTION OF ILIO-FEMORAL UNFEASIBILITY FOR TRANSFEMORAL TAVI PROCEDURE BASED ON THE MANUFACTURER¿S RECOMMENDATIONS, TO COMPARE THE SEX-SPECIFIC DISTRIBUTION OF DIFFERENT VALVE SIZES BASED ON THE MANUFACTURER¿S RECOMMENDATIONS, AS WELL AS TO DETERMINE SEX-BASED DIFFERENCES IN CLINICAL OUTCOMES POST-TAVI AND ITS ASSOCIATION WITH ANATOMIC MEASURES. DEVICES INCLUDED IN THE STUDY WERE SAPIEN 3, EVOLUT PRO+, ACURATE NEO2, AND NAVITOR VALVE. THE ARTICLE CONCLUDED FEMALE PATIENTS SHOWED SMALLER ABSOLUTE DIMENSIONS OF LVOT, AORTIC ROOT AND ILIO-FEMORAL ARTERIES COMPARED TO MALE PATIENTS. THERE WERE NO DIFFERENCES IN THE PREVALENCE OF ILIO-FEMORAL UNFEASIBILITY FOR TRANSFEMORAL TAVI PROCEDURE, BUT THERE WERE SEX-SPECIFIC DIFFERENCES IN THE AVAILABILITY OF CONVENTIONAL VALVE SIZES ACROSS DIFFERENT PLATFORMS. FEMALE PATIENTS EXHIBITED HIGHER PERIPROCEDURAL MORTALITY WITH NO DIFFERENCE IN OTHER CLINICAL OUTCOMES. [THE PRIMARY AND CORRESPONDING AUTHOR WAS FRANE RUNJIC, HEAD OF THE DEPARTMENT OF INTERVENTIONAL CARDIOLOGY, DEPARTMENT OF CARDIOLOGY, UNIVERSITY HOSPITAL OF SPLIT, SPLIT, CROATIA, WITH CORRESPONDING EMAIL: FRANE.RUNJIC@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM NOVEMBER 2019 TO JULY 2023. A TOTAL OF 320 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT SPECIFIED HOW MANY RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 81.3 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ARTERIAL HYPERTENSION, DIABETES MELLITUS, PRIOR ACUTE MYOCARDIAL INFARCTION, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, CEREBROVASCULAR INCIDENT, SMOKING HISTORY, DYSLIPIDEMIA, CANCER, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, ASTHMA, HYPOTHYROIDISM, AND ARTHRITIS. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH WAS ESTIMATED. B3: DATE OF EVENT WAS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLED " SEX-SPECIFIC ANATOMIC DIFFERENCES IN PATIENTS UNDERGOING TRANSCATHETER AORTIC VALVE IMPLANTATION_ INSIGHTS FROM THE ST-TAVI REGISTRY." Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF NAVITOR TRANSCATHETER VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ARTERIAL HYPERTENSION, DIABETES MELLITUS, PRIOR ACUTE MYOCARDIAL INFARCTION, PERCUTANEOUS CORONARY INTERVENTION, CORONARY ARTERY BYPASS GRAFT, CEREBROVASCULAR INCIDENT, SMOKING HISTORY, DYSLIPIDEMIA, CANCER, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, ASTHMA, HYPOTHYROIDISM, AND ARTHRITIS. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), UNEXPECTED MEDICAL INTERVENTION (ECMO), STROKE, CARDIAC TAMPONADE, ANNULAR RUPTURE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH WAS ESTIMATED. B3: DATE OF EVENT WAS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLED " SEX-SPECIFIC ANATOMIC DIFFERENCES IN PATIENTS UNDERGOING TRANSCATHETER AORTIC VALVE IMPLANTATION_ INSIGHTS FROM THE ST-TAVI REGISTRY. "
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21393600&pc=OOB | 2124215-2025-09013 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21393738&pc=DSQ | 2916596-2025-00855 | 2025-01-31 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Mechanical Problem; Decreased Pump Speed; Pumping Stopped; Infusion or Flow Problem | Cardiac Arrest; Aortic Valve Insufficiency/ Regurgitation | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT AN ALARM OCCURRED AROUND 04:00 PER THE PATIENT. UPON INVESTIGATION, THE CUSTOMER NOTED THAT THE PATIENT APPEARED TO HAVE BEEN HAVING INTERMITTENT LOW FLOW "ALARMS" AND PUMP STOPS. IT WAS NOTED THAT THE PATIENT'S DRIVELINE WAS PREVIOUSLY DAMAGED STARTING AROUND 10 INCHES FROM THE INSERTION SITE AND HAD BEEN REPAIRED WITH RESCUE TAPE, WITH THE LAST "OFFICIAL REPAIR" BEING MADE IN 2012. THE PATIENT WAS ADMITTED TO THE HOSPITAL AND WAS MONITORED ON TELEMETRY WITH THEIR DRIVELINE SECURED. THERE WAS NO REPORTED CHANGE IN THE PATIENT'S LACTATE DEHYDROGENASE (LDH) MARKERS, AND NO REPORTED CONCERN FOR THROMBUS OR HEMOLYSIS. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED 3 PUMP STOPS AND 6 LOW SPEED ALARMS ON (B)(6) 2025 OCCURRING WHILE ON BATTERY POWER, INDICATING A PHASE TO PHASE SHORT CAUSED BY DRIVELINE DAMAGE, AS WELL AS POWER AND FLOW ELEVATIONS POTENTIALLY INDICATING THAT SOMETHING OTHER THAN BLOOD WAS PASSING THROUGH THE PUMP. WHILE INPATIENT, THE PATIENT EXPERIENCED A PUMP STOP ON AN UNGROUNDED CABLE, WITH SUBSEQUENT CARDIAC ARREST. THE PUMP STOP WAS REPORTED BY THE CUSTOMER TO HAVE BEEN CAUSED BY THE PREVIOUSLY INDICATED PHASE TO PHASE SHORT DUE TO THE DRIVELINE DAMAGE. THE PATIENT RECEIVED FULL RESUSCITATIVE EFFORTS INCLUDING CARDIO-PULMONARY RESUSCITATION (CPR), INTUBATION, AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CANNULATION, HOWEVER THESE EFFORTS WERE DEEMED INSUFFICIENT DUE TO THE PATIENT'S AORTIC CLOSURE. THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST ON (B)(6) 2025. RELATED MFR'S COVERING DRIVELINE DAMAGE- 2916596-2022-14426; 2916596-2021-07028. Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED. SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED LOW SPEED / PUMP STOP EVENTS CONSISTENT WITH SUSPECTED DRIVELINE WIRE DAMAGE; HOWEVER, A SPECIFIC CAUSE FOR THE EVENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION AS THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. THE SUBMITTED DRIVELINE IMAGES SHOWED THAT TAPE SURROUNDED MOST OF THE EXTERNAL DRIVELINE, INCLUDING ON THE PROXIMAL AND DISTAL SIDES OF THE PREVIOUS DRIVELINE REPAIR, PREVIOUSLY ADDRESSED UNDER 2916596-2012-01234. A REVIEW OF THE SUBMITTED LOG FILE CONFIRMED LOW SPEED / PUMP STOP EVENTS ON (B)(6) 2025 WHILE THE DEVICE WAS CONNECTED TO 14-VOLT BATTERY(S). POWER ELEVATIONS WERE NOTED DURING THE PUMP STOP EVENTS, ALONG WITH ACTIVE LOW SPEED ADVISORY, LOW SPEED HAZARD, LOW FLOW HAZARD, AND MOTOR STOP ALARM FLAGS. BASED ON HISTORICAL EXPERIENCE WITH THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND SIMILAR REPORTED EVENTS, THESE EVENTS COULD BE INDICATIVE OF AN ISSUE WITH THE DRIVELINE. THERE WERE NO OTHER NOTABLE ALARMS ACTIVE IN THE LOG FILE. IT WAS REPORTED THAT HEARTMATE II LVAS WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS, INCLUDING DEATH. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ DISCUSSES DAMAGE DUE TO WEAR AND FATIGUE OF THE DRIVELINE. THIS SECTION ALSO CONTAINS INFORMATION ON ¿CARING FOR THE DRIVELINE¿ (UNDER ¿ONGOING SYSTEM ASSESSMENT AND CARE¿) AND PROVIDES POSSIBLE INDICATIONS OF DRIVELINE DAMAGE AS WELL AS HOW TO RESPOND TO SUCH EVENTS. SECTION 7 "ALARMS AND TROUBLESHOOTING" OUTLINES SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. SECTION 8 ¿EQUIPMENT STORAGE AND CARE¿ ALSO CONTAINS INFORMATION ON ¿CARE OF THE DRIVELINE,¿ AND PROVIDES POSSIBLE INDICATIONS OF DRIVELINE DAMAGE. THE HEARTMATE II LVAS PATIENT HANDBOOK IS ALSO CURRENTLY AVAILABLE. SECTION 4 ¿LIVING WITH THE HEARTMATE II¿ CONTAINS INFORMATION ON ¿CARING FOR THE DRIVELINE¿; HOWEVER, ALL HEARTMATE II LVAD DRIVELINES HAVE THE POTENTIAL FOR WIRE/SHIELD BREAKDOWN TO OCCUR DEPENDENT UPON LENGTH OF USE AND PATIENT HANDLING.¿ SECTION 5 "ALARMS AND TROUBLESHOOTING" OUTLINES SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D4: THE DEVICE EXPIRATION DATE WAS UNABLE TO BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21393822&pc=NIQ | 2124215-2025-09033 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21394327&pc=LOX | 2124215-2025-09029 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21394594&pc=NWX | 2124215-2025-09012 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21395680&pc=NPT | 2025587-2025-01285 | 2025-02-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-18 05:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Device Dislodged or Dislocated; Patient Device Interaction Problem | Cardiac Arrest; Renal Failure; Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, A PATIENT EXPERIENCED A CARDIAC ARREST DURING INITIAL VALVE DEPLOYMENT TO THE THIRD NODE REQUIRING INTUBATION AND 5 MINUTES OF CARDIAC MASSAGE. THE EVOLUT VALVE HAD TO BE RECAPTURED AND MOVED TO THE ASCENDING AORTA. AFTER STABILIZATION, A SECOND VALVE DEPLOYMENT WAS ATTEMPTED TO 80%, BUT THE VALVE WAS AGAIN RECAPTURED DUE TO DEPTH ISSUES. THE VALVE WAS AT A DEPTH OF 10 MM AND, PER THE PHYSICIAN, THE REASON FOR THE DEPTH ISSUE WAS PANNUS ON THE PREVIOUSLY-IMPLANTED SURGICAL VALVE. A THIRD DEPLOYMENT POSITIONED THE VALVE AT A 2-MILLIMETER DEPTH, BUT UPON REMOVAL OF THE DELIVERY CATHETER SYSTEM, THE VALVE WAS PULLED TOWARDS THE ASCENDING AORTA BY THE NOSE CONE. THE MEDICAL TEAM DETERMINED THAT THE VALVE'S POSITION DID NOT PRESENT ACUTE COMPLICATIONS, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT FOR CLOSE MONITORING. THE PATIENT WAS SCHEDULED FOR A TAVI IMPLANTATION AND HAD EXTENDED HOSPITALIZATION. THE PATIENT ULTIMATELY PASSED AWAY DUE TO ACUTE KIDNEY FAILURE AND A NEW CARDIAC ARREST.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DEATH WAS NOT DUE TO THE VALVE OR DELIVERY CATHETER SYSTEM (DCS), HOWEVER WAS DUE TO THE PATIENT'S CLINICAL CONDITION.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, A PATIENT EXPERIENCED A CARDIAC ARREST DURING INITIAL VALVE DEPLOYMENT TO THE THIRD NODE REQUIRING INTUBATION AND 5 MINUTES OF CARDIAC MASSAGE. THE EVOLUT VALVE HAD TO BE RECAPTURED AND MOVED TO THE ASCENDING AORTA. AFTER STABILIZATION, A SECOND VALVE DEPLOYMENT WAS ATTEMPTED TO 80%, BUT THE VALVE WAS AGAIN RECAPTURED DUE TO DEPTH ISSUES. THE VALVE WAS AT A DEPTH OF 10 MM AND, PER THE PHYSICIAN, THE REASON FOR THE DEPTH ISSUE WAS PANNUS ON THE PREVIOUSLY-IMPLANTED SURGICAL VALVE. A THIRD DEPLOYMENT POSITIONED THE VALVE AT A 2-MILLIMETER DEPTH, BUT UPON REMOVAL OF THE DELIVERY CATHETER SYSTEM, THE VALVE WAS PULLED TOWARDS THE ASCENDING AORTA BY THE NOSE CONE. THE MEDICAL TEAM DETERMINED THAT THE VALVE'S POSITION DID NOT PRESENT ACUTE COMPLICATIONS, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT FOR CLOSE MONITORING. THE PATIENT WAS SCHEDULED FOR A TAVI IMPLANTATION AND HAD EXTENDED HOSPITALIZATION. THE PATIENT ULTIMATELY PASSED AWAY DUE TO ACUTE KIDNEY FAILURE AND A NEW CARDIAC ARREST. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DEATH WAS NOT DUE TO THE VALVE OR DELIVERY CATHETER SYSTEM (DCS), HOWEVER WAS DUE TO THE PATIENT'S CLINICAL CONDITION. ADDITIONAL INFORMATION WAS RECEIVED THAT A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. SLOW DEPLOYMENT OF THE VALVE WAS PERFORMED, AND THE NOSE CONE WAS CENTERED WITHIN THE FRAME INFLOW BY SLIGHTLY RETRACTING THE GUIDEWIRE BEFORE SLOWLY REMOVING THE DCS. PER THE PHYSICIAN, THE DISLODGEMENT WAS CAUSED BY THE AMOUNT OF PANNUS PRESENT IN THE NON-MEDTRONIC (DOKIMUS) SURGICAL VALVE AND THE POSSIBLE INCORRECT VISUALIZATION OF THE SURGICAL VALVE. Manufacturer Narrative: IMAGE REVIEW: THERE WERE 18 IMAGES PROVIDED. THERE WAS A COMPUTED TOMOGRAPHY (CT) PROVIDED, BUT NO EXECUTIVE SUMMARY PROVIDING SURGICAL VALVE INFORMATION. CT IMAGE DOES SHOW SURGICAL VALVE IN PLACE WITH POSSIBLE THICKENING OR THROMBUS AROUND THE LEAFLETS. A VALVE FLUOROSCOPY CHECK WAS PROVIDED AND THE VALVE WAS NOT IDENTIFIED AS A MISLOAD. IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, A PATIENT EXPERIENCED A CARDIAC ARREST DURING INITIAL VALVE DEPLOYMENT TO THE THIRD NODE REQUIRING INTUBATION AND 5 MINUTES OF CARDIAC MASSAGE. THE IMAGE OF THE VALVE DEPLOYMENT TO THE THIRD NODE WAS PROVIDED AND APPEARS 10MM BELOW THE PIGTAIL CATHETER. MEDTRONIC DEPLOYMENT BEST PRACTICES STATE TO BEGIN DEPLOYMENT WITH MARKER BAND NEAR THE BOTTOM OF THE PIGTAIL IN THE CUSP OVERLAP PROJECTION. THERE IS EVIDENCE OF MULTIPLE VIEW BEING USED AT DIFFERENT TIMES OF THE PROCEDURE DURING DEPLOYMENT AND RECAPTURE. THEREFORE, IT IS UNABLE TO CONFIRM PROPER DEPTH DURING DEPLOYMENT, RECAPTURE, AND REDEPLOYMENT. EVIDENCE SHOWS THAT THE LAST IMAGE PRIOR TO RELEASE THE VALVE APPEARS TO BE AT A 0MM DEPTH. AS STATED IN THE INSTRUCTIONS FOR USE (IFU); CONSIDER RECAPTURING IF THE IMPLANT DEPTH IS 1 OR >5MM. 1 MM DEPTH MAY CONTRIBUTE TO AN INCREASED RISK OF PROSTHETIC VALVE DISLODGEMENT. IT WAS REPORT THAT UPON REMOVAL OF THE DELIVERY CATHETER SYSTEM, THE VALVE WAS PULLED TOWARDS THE ASCENDING AORTA BY THE NOSE CONE. THE MEDICAL TEAM DETERMINED THAT THE VALVE'S POSITION DID NOT PRESENT ACUTE COMPLICATIONS, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT FOR CLOSE MONITORING. THE PATIENT WAS SCHEDULED FOR A TAVI IMPLANTATION AND HAD EXTENDED HOSPITALIZATION. THE PATIENT ULTIMATELY PASSED AWAY DUE TO ACUTE KIDNEY FAILURE AND A NEW CARDIAC ARREST. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID GWBC30 (LOT: UNKNOWN); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A UPDATED: B5, H11 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VOLUTFX-23, SERIAL/LOT #: (B)(6), UBD: 25-JUL-2026, UDI#: (B)(4) SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21395796&pc=PNJ | 9612164-2025-00836 | 2024-08-14 04:00 | Death | MEDTRONIC IRELAND | 2025-02-18 05:00 | PNJ | MICRA | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Tachycardia; Cardiac Perforation; Pericardial Effusion; Tricuspid Valve Insufficiency/ Regurgitation; Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING OUTCOMES IN LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) PATIENTS ACCORDING TO AGE AND COMORBIDITIES. CHARLSON-COMORBIDITY-INDEX (CCI) WAS DETERMINED, AND PATIENTS WERE DIVIDED INTO A LOW- (CCI = 5) AND HIGH-COMORBIDITY (CCI > 5) GROUP. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS IN-HOSPITAL DEATHS OR ALL-CAUSE. NONE OF THE IN-HOSPITAL DEATHS WERE DEEMED PERI-PROCEDURAL COMPLICATIONS. OVERALL IN-HOSPITAL DEATHS AND ALL-CAUSE MORTALITY WAS SIGNIFICANTLY HIGHER IN THE HIGH COMORBIDITY GROUP. THERE WERE CASES OF ABORTED IMPLANTS: HIGH THRESHOLDS AFTER SEVERAL ATTEMPTS IN TWO CASES, AND ONE PERFORATION WHICH REQUIRED ACUTE STERNOTOMY. THERE WERE PATIENTS WHO EXPERIENCED PERICARDIAL EFFUSION WHICH REQUIRED PERICARDIOCENTESIS, UNKNOWN VASCULAR COMPLICATIONS WHICH REQUIRED INTERVENTIONAL OR SURGICAL TREATMENT, RUPTURE OF TRICUSPID TENDON WHICH RESULTED IN SIGNIFICANT TRICUSPID REGURGITATION, CARDIAC DECOMPENSATION DURING INTERVENTION WHICH NECESSITATED NON-INVASIVE VENTILATION AND INTRAVENOUS DIURETICS, VENTRICULAR TACHYCARDIA (VT) WHICH REQUIRED CARDIOVERSION, DEEP VEIN THROMBOSIS, AND LEADLESS IPG DISLOCATION WHICH NECESSITATED CAPTURING AND REPOSITIONING. IN THE FOLLOW-UP PERIOD, THERE WERE DEVICES WHICH EXHIBITED PREMATURE BATTERY DEPLETION DUE TO HIGH THRESHOLDS WITH SOME REQUIRING A DEVICE REVISION, AND PACING INDUCED LEFT VENTRICULAR (LV) DYSFUNCTION WHICH REQUIRED A DEVICE UPGRADE. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED.Event Description: IT WAS FURTHER REPORTED THAT SOME OF THE CAUSES OF DEATH WERE DUE TO CARDIAC CAUSES, INFECTION, AND CEREBROVASCULAR CAUSES. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/81 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPACT OF COMORBIDITY ON COMPLICATION RATES AND LIFE EXPECTANCY IN PATIENTS WITH A LEADLESS PACEMAKER. INTERNATIONAL JOURNAL OF CARDIOLOGY. 2024. 415: 132453 DOI: 10.1016/J.IJCARD.2024.132453. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: MICRA PRODUCT ID: MDT-TPS-DELSYS SERIAL: UNKNOWN, D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21395896&pc=MHY | 3006630150-2025-00830 | 2025-01-20 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-02-18 05:00 | MHY | VERCISE GEVIA? | Unexpected Therapeutic Results | Cardiac Arrest; High Blood Pressure/ Hypertension; Movement Disorder; Insufficient Information | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT A DEEP BRAIN STIMULATION (DBS) PATIENT SUFFERED A CARDIAC EVENT FOLLOWING A LOSS OF DBS THERAPY. HOWEVER, THIS COULD NOT BE CONFIRMED, AS THE PATIENT DID NOT HAVE THEIR REMOTE CONTROL TO VERIFY THE STATUS OF THEIR DBS SYSTEM. THE PHYSICIAN THEORIZED THAT THE DBS SHUTDOWN MIGHT HAVE TRIGGERED PARKINSONISM-HYPERPYREXIA SYNDROME AND RHABDOMYOLYSIS, WHICH LIKELY LED TO THE HEART ATTACK. FURTHERMORE, THE PATIENT HAD BEEN TAKING A NEW MEDICATION KNOWN TO HAVE RHABDOMYOLYSIS AS A POTENTIAL SIDE EFFECT. THE PATIENT WAS HOSPITALIZED BUT, UNFORTUNATELY, PASSED AWAY.Event Description: IT WAS REPORTED THAT A DEEP BRAIN STIMULATION (DBS) PATIENT SUFFERED A CARDIAC EVENT FOLLOWING A LOSS OF DBS THERAPY. HOWEVER, THIS COULD NOT BE CONFIRMED, AS THE PATIENT DID NOT HAVE THEIR REMOTE CONTROL TO VERIFY THE STATUS OF THEIR DBS SYSTEM. THE PHYSICIAN THEORIZED THAT THE DBS SHUTDOWN MIGHT HAVE TRIGGERED PARKINSONISM-HYPERPYREXIA SYNDROME AND RHABDOMYOLYSIS, WHICH LIKELY LED TO THE HEART ATTACK. FURTHERMORE, THE PATIENT HAD BEEN TAKING A NEW MEDICATION KNOWN TO HAVE RHABDOMYOLYSIS AS A POTENTIAL SIDE EFFECT. THE PATIENT WAS HOSPITALIZED BUT, UNFORTUNATELY, PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED THAT REVEALED THAT NEITHER THE FAMILY MEMBERS OF THE PATIENT NOR THE INTENSIVE CARE UNIT PHYSICIAN BELIEVED THE PATIENT'S PASSING WAS RELATED TO THEIR DBS SYSTEM. DURING THIS EVENT, THE PATIENT SHOWED SYMPTOMS OF DARK URINE AND ELEVATED BLOOD LEVELS. THE PATIENT WAS INTUBATED, AND CARE WAS SUBSEQUENTLY WITHDRAWN. UPON RECEIVING THE IMPLANTABLE PULSE GENERATOR (IPG), THE BOSTON SCIENTIFIC REPRESENTATIVE NOTED THAT IT WAS STILL FUNCTIONING.Event Description: IT WAS REPORTED THAT A DEEP BRAIN STIMULATION (DBS) PATIENT SUFFERED A CARDIAC EVENT FOLLOWING A LOSS OF DBS THERAPY. HOWEVER, THIS COULD NOT BE CONFIRMED, AS THE PATIENT DID NOT HAVE THEIR REMOTE CONTROL TO VERIFY THE STATUS OF THEIR DBS SYSTEM. THE PHYSICIAN THEORIZED THAT THE DBS SHUTDOWN MIGHT HAVE TRIGGERED PARKINSONISM-HYPERPYREXIA SYNDROME AND RHABDOMYOLYSIS, WHICH LIKELY LED TO THE HEART ATTACK. FURTHERMORE, THE PATIENT HAD BEEN TAKING A NEW MEDICATION KNOWN TO HAVE RHABDOMYOLYSIS AS A POTENTIAL SIDE EFFECT. THE PATIENT WAS HOSPITALIZED BUT, UNFORTUNATELY, PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED THAT REVEALED THAT NEITHER THE FAMILY MEMBERS OF THE PATIENT NOR THE INTENSIVE CARE UNIT PHYSICIAN BELIEVED THE PATIENT'S PASSING WAS RELATED TO THEIR DBS SYSTEM. DURING THIS EVENT, THE PATIENT SHOWED SYMPTOMS OF DARK URINE AND ELEVATED BLOOD LEVELS. THE PATIENT WAS INTUBATED, AND CARE WAS SUBSEQUENTLY WITHDRAWN. UPON RECEIVING THE IMPLANTABLE PULSE GENERATOR (IPG), THE BOSTON SCIENTIFIC REPRESENTATIVE NOTED THAT IT WAS STILL FUNCTIONING. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BOSTON SCIENTIFIC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: THE RETURNED IMPLANTABLE PULSE GENERATOR (IPG) WAS ANALYZED AND PASSED VISUAL INSPECTION AND THE FUNCTIONAL TEST WITHOUT ANY ANOMALIES. HOWEVER, THE IPG DATA LOGS WERE UNABLE TO BE CAPTURED. A LABELLING REVIEW FOUND THAT THE REPORTED EVENT IS A KNOWN RISK WITH THE USE OF THE DEEP BRAIN STIMULATION (DBS) SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21396008&pc=BZD | 2518422-2025-102107 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-02-18 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A DREAMSTATION AUTO CPAP DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21396049&pc=LYJ | 1644487-2025-00176 | 2025-01-23 05:00 | Death | LIVANOVA USA, INC. | 2025-02-18 05:00 | LYJ | PULSE GEN MODEL UNKNOWN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: THE PATIENT'S MOTHER REPORTED THAT THE PATIENT PASSED AWAY WITH THE VNS THAT WAS SUGGESTED BY HER PHYSICIAN. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397075&pc=DSQ | 2916596-2025-01035 | 2021-09-01 04:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "OUTCOMES OF INTRACORPOREAL CONTINUOUS AND PARACORPOREAL PULSATILE VENTRICULAR ASSIST DEVICES IN PEDIATRIC PATIENTS 10¿30 KG" THAT HEARTMATE 3 MAY BE ASSOCIATED WITH DEATH, STROKE, HEMOLYSIS, BLEEDING, INFECTION, NEUROLOGICAL DYSFUNCTION, RENAL DYSFUNCTION, AND RIGHT HEART FAILURE. THIS RETROSPECTIVE STUDY EVALUATED 95 PATIENTS WHO WEIGHED BETWEEN 10 KILOGRAMS (KG) AND 30 KG AND RECEIVED AN INTRACORPOREAL CONTINUOUS (IC) OR PARACORPOREAL PULSATILE (PP) VENTRICULAR ASSIST DEVICE (VAD) BETWEEN JUN2018 AND SEP2021. 10 OF THE PATIENTS WERE ON HEARTMATE 3 AND 31 PATIENTS HAD HEARTWARE, WITH A TOTAL OF 41 PATIENTS ON IC VADS. THE REMAINING 54 PATIENTS HAD PP VADS, 54 WITH EXCOR, 1 WITH PEDIMAG, AND 2 WITH ROTAFLOW. EACH DEVICE WAS USED AS A LEFT VAD. THE PATIENTS WITH IC VADS WERE OLDER (MEDIAN [IQR] = 7.7 [5.6¿10.1] YEARS, P < 0.001) AND HAD A HIGHER WEIGHT (MEDIAN [IQR] = 21.9 [17.5¿27] KG, P = 0.001) COMPARED TO THE PP GROUP (MEDIAN AGE [IQR] = 3.7 [2.9¿6.1] YEARS, MEDIAN WEIGHT [IQR] = 14.3 [11.9¿17.8] KG). THERE WAS NO STATISTICAL DIFFERENCE IN DIAGNOSIS, OR INOTROPE USE AT THE TIME OF IMPLANTATION, HOWEVER THE PARACORPOREAL GROUP HAD MORE THAN TWICE AS MANY PATIENTS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AT THE TIME OF VAD IMPLEMENTATION, COMPARED TO THE INTRACORPOREAL GROUP (16 [29.6%] VS. 6 [14.6%], P = 0.086). IN THE PATIENTS WITH IC VADS, 11 PATIENTS REMAINED ALIVE ON THE DEVICE, 6 PATIENTS PASSED AWAY, 23 PATIENTS WERE TRANSPLANTED, AND 1 PATIENT HAD VENTRICULAR RECOVERY AND WAS WEANED. OF THE IC VAD IMPLANTED PATIENTS, 1 HAD HEMOLYSIS, 14 HAD MAJOR BLEEDING, 1 HAD AN EXTRA-AXIAL HEMORRHAGE, 1 HAD A HEMORRHAGIC STROKE, 3 HAD AN ISCHEMIC STROKE, 3 HAD RENAL DYSFUNCTION, 4 HAD RIGHT HEART FAILURE, AND 20 HAD MAJOR INFECTIONS (INCLUDING MEDIASTINITIS, SEPSIS, LOCALIZED NON-DEVICE INFECTIONS, MAJOR INFECTIONS ASSOCIATED WITH THE INTERNAL PUMP, INFLOW OR OUTFLOW TRACTS, AND PERCUTANEOUS SITE OR POCKET INFECTIONS.) STROKE WAS NUMERICALLY MORE COMMON IN THE PP COHORT WITH 4.6 ADVERSE EVENTS PER 100 PATIENT MONTHS COMPARED TO 1.45 ADVERSE EVENTS PER 100 PATIENT MONTHS IN THE IC GROUP, BUT DID NOT REACH STATISTICAL SIGNIFICANCE (P = 0.07). HEMOLYSIS WAS GREATER IN THE PP GROUP WITH 4.2 ADVERSE EVENTS PER 100 PATIENT MONTHS COMPARED TO 0.36 ADVERSE EVENTS PER 100 PATIENT MONTHS (P = 0.004). THERE WERE NO STATISTICALLY SIGNIFICANT DIFFERENCES IN MAJOR BLEEDING OR MAJOR INFECTION. DEVICE MALFUNCTION WAS HIGHER IN THE PP COHORT WITH 6.04 ADVERSE EVENTS PER 100 PATIENT MONTHS COMPARED TO NONE IN THE IC COHORT (P = 0.001). THERE WAS ALSO NO DIFFERENCE IN CLINICAL OUTCOMES, INCLUDING PATIENTS ALIVE ON DEVICE (P = 0.343) AND PATIENTS ALIVE AND TRANSPLANTED (P = 0.42). ONE PATIENT (2%) IN THE IC COHORT HAD VENTRICULAR RECOVERY AND WEANED FROM DEVICE. THERE WAS NO DIFFERENCE IN NUMBER OF DEVICE DAYS (IC MEDIAN [IQR] = 124 [47.5¿257] DAYS, PP MEDIAN [IQR] = 107.5 [51.2¿161] DAYS, P = 0.170) OR TOTAL OF ICU DAYS POSTIMPLANT (IC MEDIAN [IQR] = 34.5 [21¿66.75] DAYS, PP MEDIAN [IQR] = 60 [20¿114] DAYS, P = 0.109). FOUR PATIENTS (9.8%) IN THE IC GROUP WERE DISCHARGED ON DEVICE. IN THIS STUDY OF PEDIATRIC PATIENTS WEIGHING 10¿30 KG, WE FOUND NO STATISTICALLY SIGNIFICANT DIFFERENCE IN MORTALITY BETWEEN THE IC AND PP COHORTS. SIGNIFICANT DIFFERENCES WERE FOUND IN HIGHER RATES OF HEMOLYSIS AND DEVICE MALFUNCTION IN THE PP GROUP, WITH NUMERICALLY HIGHER BUT NOT STATISTICALLY DIFFERENT RATES OF STROKE. FURTHER STUDIES TO EVALUATE THE POTENTIAL BENEFITS OF THESE DEVICES IN SMALLER PEDIATRIC PATIENTS REMAIN IMPORTANT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. HERMA HEART INSTITUTE, CHILDREN¿S WISCONSIN, AND DEPARTMENT OF PEDIATRICS, SECTION OF CRITICAL CARE, MEDICAL COLLEGE OF WISCONSIN, MILWAUKEE, WISCONSIN; AND ¿HEART INSTITUTE, CINCINNATI CHILDREN¿S HOSPITAL MEDICAL CENTER, CINCINNATI, OHIO. LEE, J., SHEZAD, M. F., ZAFAR, F., & NIEBLER, R. A. (2024). OUTCOMES OF INTRACORPOREAL CONTINUOUS AND PARACORPOREAL PULSATILE VENTRICULAR ASSIST DEVICES IN PEDIATRIC PATIENTS 10-30 KG. ASAIO JOURNAL, 70(7), 616-620. DOI:HTTP://DX.DOI.ORG/10.1097/MAT.000000000000216. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397080&pc=DSP | 3010532612-2025-00122 | 2025-01-23 05:00 | Death | ARROW INTERNATIONAL LLC | 2025-02-18 05:00 | DSP | ARROW AC3 OPTIMUS IABP NA/EMEA | Defective Device | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K162820 |  | 1 | Event Description: IT WAS REPORTED "THE INCORRECT TRIGGER HAD BEEN SELECTED AND THE PATIENT HAD DIED, POSSIBLY AS A RESULT. IABP SET INCORRECTLY - OPERATOR MODE WITH IABP ON INTERNAL MODE AT RATE OF 85BPM - NOT AUGMENTING IN PHASE WITH PRESSURE OWN BLOOD PRESSURE - NOT IN PHASE WITH CARDIAC RHYTHM AND PATIENT OWN AUGMENTATION OF BLOOD PRESSURE. NO ECG READING ON IABP MACHINE (ARTEFACTS +++)". ADDITIONAL INFORMAITON STATES " THE PATIENT WAS VERY SICK WHEN THEY ARRIVED WITH US AND PRIOR AND DURING IABP INSERTION EXPERIENCED MULTIPLE RUNS OF VENTRICULAR F IBRILLATION AND VENTRICULAR TACHYCARDIA. ON LEAVING THE CATH LABS AND MOVING TO THE ICU THE PATIENT WAS ON MULTIPLE INOTROPES AS WELL AS THE IABP. AT THIS TIME WE ARE STILL UNSURE AS TO WHETHER THE IABP CONTRIBUTED TO THE DEATH OF THIS PATIENT, HOWEVER, ALONG WITH THIS INFORMATION AND ACCOUNTS AND INVESTIGATIONS WITH OUR CLINICAL TEAM WE HOPE TO BE ABLE TO PROVIDE FURTHER INFORMATION".Event Description: IT WAS REPORTED "THE INCORRECT TRIGGER HAD BEEN SELECTED AND THE PATIENT HAD DIED, POSSIBLY AS A RESULT. IABP SET INCORRECTLY - OPERATOR MODE WITH IABP ON INTERNAL MODE AT RATE OF 85BPM - NOT AUGMENTING IN PHASE WITH PRESSURE OWN BLOOD PRESSURE - NOT IN PHASE WITH CARDIAC RHYTHM AND PATIENT OWN AUGMENTATION OF BLOOD PRESSURE. NO ECG READING ON IABP MACHINE (ARTEFACTS +++)". ADDITIONAL INFORMATION STATES " THE PATIENT WAS VERY SICK WHEN THEY ARRIVED WITH US AND PRIOR AND DURING IABP INSERTION EXPERIENCED MULTIPLE RUNS OF VENTRICULAR FIBRILLATION AND VENTRICULAR TACHYCARDIA. ON LEAVING THE CATH LABS AND MOVING TO THE ICU THE PATIENT WAS ON MULTIPLE INOTROPES AS WELL AS THE IABP. AT THIS TIME WE ARE STILL UNSURE AS TO WHETHER THE IABP CONTRIBUTED TO THE DEATH OF THIS PATIENT, HOWEVER, ALONG WITH THIS INFORMATION AND ACCOUNTS AND INVESTIGATIONS WITH OUR CLINICAL TEAM WE HOPE TO BE ABLE TO PROVIDE FURTHER INFORMATION". Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). NO IABP PART WAS RETURNED TO TELEFLEX CHELMSFORD FOR INVESTIGATION. THE ALARM HISTORY LOG FILE WAS REVIEWED. THE LOG FILE SHOWS 6 "POSSIBLE ECG TRIGGER DETECTED" ALARMS, WHICH CAN ONLY BE ISSUED WHILE IN INTERNAL TRIGGER. ADDITIONALLY, IT WAS NOTED THAT THE USER SWITCHED FROM AUTOPILOT TO OPERATOR MODE AT LEAST TWICE BASED ON THE ALARMS THAT WERE ISSUED INCLUDING ALARMS WHICH CAN ONLY BE ISSUED WHILE IN OPERATOR MODE ("AP TRIGGER LOSS" AND "ECG TRIGGER LOSS") AND AUTOPILOT MODE ("TRIGGER LOSS"). DUE TO THE FINDINGS OF INTERNAL TRIGGER BEING USED WHILE THE PATIENT POTENTIALLY HAD INTRINSIC CARDIAC ACTIVITY, AN IN-SERVICE HAS BEEN REQUESTED TO REVIEW THE INSTRUCTIONS FOR USE (IFU) WITH THE CUSTOMER. THE INSTRUCTIONS FOR USE (IFU) STATES: "WARNING: INTERNAL TRIGGER SHOULD NOT BE USED WHEN THE PATIENT HAS INTRINSIC CARDIAC ACTIVITY. THIS CAN CAUSE INCORRECT TIMING WHICH MAY IMPAIR THE PATIENT HEMODYNAMICS. NOTE: INTERNAL TRIGGER MODE SHOULD ONLY BE USED IF THE PATIENT HAS NO MYOCARDIAL ACTIVITY AND/OR VENTRICULAR EJECTION. YOU MUST SELECT OPERATOR MODE, THEN TRIGGER MODE, INTERNAL AND CONFIRM TO SELECT THIS TRIGGER MODE. NOTE: UPON DETECTION OF AN ECG SIGNAL, THE IABP WILL ISSUE THE "POSSIBLE ECG SIGNAL DETECTED" ALARM. CHECK THE PATIENT AND ECG SIGNAL AND CONFIRM IF THE ECG HAS RETURNED. IF SO, SELECT A PATIENT-BASED TRIGGER SIGNAL -- EITHER ECG OR AP. FAILURE TO DO THIS MAY RESULT IN ASYNCHRONOUS PUMPING." A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE LOT NUMBER/SERIAL NUMBER WITH NO RELEVANT FINDINGS. THE DEVICE PASSED ALL MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE REPORTED COMPLAINT OF "INCORRECT TRIGGER" IS CONFIRMED BASED THE ALARM LOG PROVIDED. THE LOG FILE SHOWS MULTIPLE "POSSIBLE ECG TRIGGER DETECTED" ALARMS, WHICH INDICATE THAT THE USER WAS USING THE INTERNAL TRIGGER MODE WHILE THE PATIENT POTENTIALLY HAD INTRINSIC CARDIAC ACTIVITY. NO PART WAS RETURNED TO TELEFLEX CHELMSFORD FOR INVESTIGATION. BASED ON THE REVIEW OF THE LOG FILE, AN IN-SERVICE HAS BEEN REQUESTED TO REVIEW THE IFU WITH THE CUSTOMER. BASED ON A REVIEW OF THE DEVICE HISTORY RECORD (DHR), THE PRODUCT MET SPECIFICATION UPON RELEASE. THE MOST PROBABLE ROOT CAUSE OF THE COMPLAINT IS USER RELATED. THIS WILL BE MONITORED FOR ANY DEVELOPING TRENDS. THIS WILL BE MONITORED FOR ANY DEVELOPING TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397185&pc=NWX | 2124215-2025-09034 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397274&pc=OZD | 1220648-2025-26158 | 2025-01-25 05:00 | Death | ABIOMED, INC. | 2025-02-18 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THREE DAYS INTO IMPELLA CP SUPPORT, THE PATIENT BEGAN TO EXPERIENCE PERSISTENT BLEEDING FROM THEIR ACCESS SITE. THE REPOSITIONING SHEATH WAS ADJUSTED AND FEMOSTOP WAS APPLIED AT THE SITE, BUT THE BLEED CONTINUED. IN ADDITION, IT WAS DOCUMENTED THAT THE PUMP WAS UNABLE TO PROVIDE ADEQUATE SUPPORT DUE TO SUCTION BEING ENCOUNTERED AT LEVELS ABOVE P5. AS A RESULT OF THE PATIENT'S WORSENING CONDITION AND THE RESTRICTED SUPPORT LEVEL, DISCUSSIONS WERE INITIATED ON THE DISCONTINUATION OF SUPPORT. THE PATIENT EXPIRED AT THE END OF SUPPORT. IT IS UNKNOWN IF THE COMPLICATIONS DIRECTLY CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND LOW PUMP FLOW HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE PATIENT WAS BLEEDING FROM MULTIPLE SITE AND THE AT THE MOMENT ACTIVATED CLOTTING TIME'S WERE NOT AVAILABLE. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED. DATA LOGS WERE NOT RETURNED AS WELL. WITHOUT THE PUMP OR THE DATA LOGS THE FAILURE COULD NOT BE FURTHER INVESTIGATED. THEREFORE, THE ROOT CAUSE OF THE LOW FLOW ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND DATA LOGS WERE NOT RETURNED AS WELL. B.2 REQUIRED INTERVENTION WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26158. G.1 ADDED ADDRESS LINE 1, CITY, STATE, ZIP CODE AND COUNTRY WERE ADDED AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26158. H.6 A HEALTH EFFECT - IMPACT CODE WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26158. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397729&pc=DQY | 2124215-2025-09027 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | DQY | TRAPPER EXCHANGE DEVICE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397744&pc=PBN | 3002808904-2025-00005 | 2022-06-05 04:00 | Death | KANEKA CORPORATION | 2025-02-18 05:00 | PBN | LIPOSORBER LA-15 SYSTEM | Adverse Event Without Identified Device or Use Problem | Pulmonary Edema | P910018 |  | 1 | Event Description: DEATH OF AN 8-YEAR-OLD FSGS PATIENT WAS REPORTED. THIS WAS AN 8-YEAR-OLD GIRL WHO WAS DIAGNOSED WITH NEPHROTIC SYNDROME WITH ACUTE KIDNEY INJURY AND HYPERTENSION ON (B)(6) 2022. A RENAL BIOPSY ON (B)(6) 2022 SHOWED COLLAPSING GLOMERULOPATHY. SHE REQUIRED INITIATION OF HEMODIALYSIS ON (B)(6) 2022 FOR UREMIA AND FLUID OVERLOAD. SHE REMAINED DIALYSIS DEPENDENT. HEMODIALYSIS WAS PERFORMED AT DAVITA 3X / WK ON MONDAYS, WEDNESDAYS, AND FRIDAYS UNDER THE DIRECTION OF CEDARS PEDIATRIC NEPHROLOGISTS. SHE WAS RESISTANT TO STEROIDS AND RITUXIMAB. CALCINEURIN INHIBITORS WERE NOT GIVEN DUE TO POOR RENAL FUNCTION. LDL APHERESIS WAS STARTED WITHOUT PROBLEMS FOR REFRACTORY FSGS ON (B)(6) 2022. ON (B)(6) 2022, SHE UNDERWENT THE 10TH TREATMENT OUT OF 12 TREATMENTS WITHOUT PROBLEMS. SHE WAS GIVEN SOLUMEDROL PER PROTOCOL. SHE RECEIVED HEMODIALYSIS TREATMENTS AT DAVITA ON (B)(6) 2022. THEY WERE UNEVENTFUL. THE NEXT HEMODIALYSIS WAS SCHEDULED ON (B)(6) 2022. ON (B)(6) 2022 AT 1:25 AM, THE PATIENT PRESENTED AT KP LAMC ED FOR SHORTNESS OF BREATH. THE PARENT REPORTED THAT HE WAS GETTING THE PATIENT READY FOR BED AND NOTED THE INCREASED WORK OF BREATHING SO THE PATIENT WAS BROUGHT TO THE ED. HER O2 SATURATION WAS IN THE 40S ON ROOM AIR AND IMPROVED TO AS HIGH AS 79% ON A NON-REBREATHER MASK THUS THE DECISION WAS MADE TO INTUBATE THE PATIENT. VISUALIZATION WAS DIFFICULT GIVEN THE GENEROUS OROPHARYNGEAL SOFT TISSUE AND COPIOUS BLOOD-TINGED THIN LIQUID POOLING IN THE OROPHARYNX. THIS WAS ABORTED AND SHE WAS AGAIN BAGGED WITH A RETURN TO ONLY ~60% SPO2. SHE BECAME BRADYCARDIC, SO CHEST COMPRESSIONS WERE INITIATED, AND THE FIRST DOSE OF EPINEPHRINE WAS GIVEN SOMETIME AROUND 1:45 AM. THE SECOND ATTEMPT WITH MAC 3 LARYNGOSCOPE BLADE WAS SUCCESSFUL AT FULL VISUALIZATION OF THE GLOTTIS AND CORDS AFTER SUCTIONING COPIOUS BLOOD-TINGED THIN LIQUID FROM THE OROPHARYNX, WITH CONTINUAL RECURRENT FILLING OF THE OP WITH FLUID FROM THE TRACHEA. HOWEVER, A 5.5 CUFFED AND STYLETED ET TUBE WAS UNABLE TO PASS THROUGH THE CORDS. THE PATIENT WAS BAGGED AGAIN AND ON 3RD ATTEMPT, AGAIN WITH MAC 3 BLADE ON C-MAC WITH DIRECT VIDEO VISUALIZATION, A 5.0 UNCUFFED ETT WAS OBSERVED PASSING THROUGH THE CORDS. BREATH SOUNDS WERE APPRECIATED BILATERALLY ON AUSCULTATION BY AN ED PHYSICIAN, HOWEVER END-TIDAL MONITORING DID NOT SHOW A WAVEFORM. A SUCTION CATHETER WAS PASSED DOWN THE ETT AND COPIOUS BLOOD-TINGED THIN LIQUID RETURNED AND POOLED IN THE SUCTION CANISTER. BRIEFLY FOLLOWING ETT SUCTION, O2 SATS IMPROVED TO THE 30S TO 40S AND END TIDAL MONITORING REGISTERED A NUMBER BUT NO WAVEFORM, BUT SATS AND END TIDAL QUICKLY FELL AGAIN TO EVENTUALLY UNDETECTABLE DESPITE SUCTIONING COPIOUS FLUID FROM THE ETT PERIODICALLY THROUGHOUT THE CODE EVENT. FLUID ALSO POURED OUT OF THE PATIENT'S NOSE AND MOUTH THROUGHOUT THE CODE EVENT. CPR WITH CHEST COMPRESSIONS AND 7 ROUNDS OF EPINEPHRINE WERE GIVEN. A DOSE OF CALCIUM CHLORIDE WAS ALSO GIVEN DURING THE 30-MINUTE CODE, 2 DOSES OF SODIUM BICARBONATE, AND INSULIN AND A GLUCOSE BOLUS WERE ALSO GIVEN, ALTHOUGH WAS FSBG OF WELL OVER 100. AT AROUND 25 MINUTES INTO THE EVENT, FRANK BLOOD RETURNED FROM THE ETT AND BEGAN TO POUR FROM THE PATIENT'S MOUTH AND NOSE. AT NO TIME DURING REGULAR PULSE CHECKS WAS THERE A PALPABLE EFFECTIVE PULSE DETECTED. AT 2:15 AM, DURING THE LAST PULSE CHECK, THE MONITOR SHOWED PEA, THERE WAS NO PALPABLE PULSE, O2 SATS WERE UNDETECTABLE, AND THE PATIENT WAS CYANOTIC CENTRALLY. DEATH WAS PRONOUNCED AT 2:15AM. SHE WAS THOUGHT TO HAVE CARDIORESPIRATORY FAILURE DUE TO ACUTE ONSET PULMONARY EDEMA, CAUSE UNKNOWN. AUTOPSY WAS DECLINED. MEDICAL DIRECTOR REPORTED THAT THE PT HAD NO COMPLICATIONS DURING THE LIPOSORBER TREATMENT AND DOES NOT BELIEVE PT'S PASSING WAS DUE TO LIPOSORBER TREATMENT. Manufacturer Narrative: THE DEVICE HISTORY RECORDS (DHR) OF THE DEVICE CONCERNED WAS REVIEWED.THE PRODUCTION LOT, TO WHICH THE DEVICE CONCERNED BELONGS, PASSED ALL IN-PROCESS INSPECTIONS TEST. NO NONCONFORMITY OR ABNORMALITY IN THE MANUFACTURING PROCESSES OF THE DEVICE CONCERNED WAS FOUND. THE ACTUAL DEVICE WAS NOT RETURNED AND COULD NOT BE INVESTIGATED. TEN TREATMENTS WITH OUR PRODUCT WERE PERFORMED AND COMPLETED WITHOUT PROBLEMS AND WITHOUT PRODUCT FAILURE. AFTER THAT, TWO HEMODIALYSIS TREATMENTS USING ANOTHER COMPANY'S PRODUCT WERE ALSO COMPLETED WITHOUT ANY PROBLEMS. ALSO, WE RECEIVED A COMMENT FROM THE DOCTOR IN CHARGE THAT THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND OUR PRODUCT. AS A RESULT OF THE ABOVE, WE THINK THAT THE ADVERSE EVENT WAS CAUSED BY THE PATIENT'S CONDITION AND THAT THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND OUR PRODUCT. REASON FOR SUBMISSION OF THIS MDR: WE PERFORMED A DETAILED INVESTIGATION OF THE PATIENT'S MEDICAL HISTORY AND HOSPITAL TREATMENT PROCESS AND COLLECTED SUFFICIENT ADDITIONAL INFORMATION TO CONFIRM THAT THERE WAS NO DEVICE FAILURE, NO PROBLEM WITH THE DHR, AND THAT A CAUSAL RELATIONSHIP WAS RULED OUT AS A PHYSICIAN'S COMMENT IN CHARGE. AS THE MANUFACTURER, WE HAD DETERMINED ON 15 FEB., 2022, IN ACCORDANCE WITH OUR PROCEDURES, THAT THERE WAS NO CAUSAL RELATIONSHIP WITH OUR PRODUCT AND THAT THE MDR WAS NOT APPLICABLE. HOWEVER, AT THE FDA AUDIT RECEIVED ON 6 FEB., 2025, THE AUDITOR IN CHARGE ADVISED US TO SUBMIT AN MDR WITH THE STATEMENT "AS A MANUFACTURER, WE HAVE DETERMINED THAT THERE IS NO CAUSAL RELATIONSHIP" AND WE ARE SUBMITTING THIS MDR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21397802&pc=NIQ | 2124215-2025-09017 | 2025-01-25 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY. A 2.50 X 48 MM SYNERGY XD WAS DEPLOYED AND THE PERCUTANEOUS CORONARY INTERVENTION WAS COMPLETED WITH THROMBOLYSIS IN MYOCARDIAL INFARCTION (TIMI) 3 FLOW. HALF AN HOUR LATER THE PATIENT WENT INTO HEART FAILURE. ADRENALINE WAS PROVIDED AND AN INTRA-AORTIC BALLOON PUMP (IABP) WAS PLACED. DEFIBRILLATION AND CARDIOPULMONARY RESUSCITATION WERE PERFORMED FOR 30-40 MINUTES UNSUCCESSFULLY AND THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS VENTRICULAR FIBRILLATION. Manufacturer Narrative: E1 INITIAL REPORTER ADDRESS: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21398263&pc=LOX | 2124215-2025-09023 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21398742&pc=OZD | 1220648-2025-26170 | 2025-01-26 05:00 | Death | ABIOMED, INC. | 2025-02-18 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE) IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND SIX DAYS FOLLOWING THE IMPLANT, LOW PUMP FLOW WAS ENCOUNTERED. THE DEVICE WAS EXPLANTED WITH AN UNSUCCESSFUL ATTEMPT TO PLACE AND IMPELLA 5.5 HOWEVER ANOTHER IMPELLA CP WAS IMPLANTED. THREE DAYS FOLLOWING, THE REPLACEMENT IMPELLA CP DEVICE WOULD HAVE SUCTION ALARMS, PATIENT EXPERIENCE VENTRICULAR FIBRILLATION AND EXPIRES. THE PATIENT PRESENTED TO THE HOSPITAL WITH CHEST PAIN AND VOMITING AND FOUND TO HAVE ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WITH EJECTION FRACTION 10-15%. THE PATIENT HAD A STENT PLACED FIVE DAYS PRIOR. THE PATIENT DECOMPENSATED IN THE EMERGENCY DEPARTMENT AND WAS TAKEN BACK TO THE CATHETERIZATION LAB FOR AN IMPELLA CP DEVICE PLACEMENT, WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT WAS TRANSFERRED TO THE UNIT ON SUPPORT. APPROXIMATELY SIX DAYS AFTER START OF THE IMPELLA MCS, CONTINUOUS SUCTION ALARM WAS REPORTED. THERE WERE LOW FLOWS AT HIGH P LEVELS, BUT NO KNOWN HEMOLYSIS. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR A HIGHER LEVEL OF CARE AND ADVANCED THERAPIES EVALUATION. THE DECISION WAS MADE TO ESCALATE TO AN IMPELLA 5.5 DEVICE. PATIENT BROUGHT TO OPERATING ROOM, PREPPED/DRAPED FOR LEFT AXILLARY IMPELLA 5.5 DEVICE IMPLANT. LEFT AXILLARY CUT DOWN, GRAFT SEWN DIRECTLY ONTO ARTERY AND TUNNEL INFERIOR. LEFT VENTRICLE WAS THEN ACCESSED VIA WIRE/CATHETER TECHNIQUE UNDER FLUOROSCOPY/TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THE .018 GUIDEWIRE WAS LOADED AND PLACED INTO LV AND CATHETER REMOVED. THE IMPELLA 5.5 WAS THEN BACKLOADED ONTO WIRE; HOWEVER, THERE WERE FAILED ATTEMPTS TO CROSS THE ANASTOMOSIS. IMPELLA 5.5 THEN REMOVED AND AN IMPELLA CP WAS OPENED AND PLACED. PLACEMENT CONFIRMED WITH TEE, MEASURING AT 3.7 CM. IMPELLA CP RAMPED TO P-8 WITH FLOWS 2.3 L/MIN. CHIMNEY GRAFT/REPO SHEATH SUTURED TO THE LEFT CHEST. STERILE DRESSING APPLIED TO LEFT AXILLARY SITE PER THEIR PROTOCOL. THE PATIENT WAS TRANSFERRED TO THE BED FOR TRANSPORT BACK TO THE UNIT, BUT UPON LEAVING THE ROOM, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. THE PATIENT WAS SHOCKED TWO TIMES AND RETURNED TO NORMAL SINUS RHYTHM. TEE PROBE WAS RE-INSERTED AND PLACEMENT CONFIRMED AND THE PATIENT WAS TAKEN TO THE UNIT. SAME DAY, LATER THAT NIGHT, ON THE UNIT, THE AUTOMATED IMPELLA CONTROLLER DISPLAYED/SOUNDED CONTINUOUS SUCTION ALARM. REPLACING PA CATHETER AND PLAN TO OBTAIN AN ECHO AS SOON AS POSSIBLE TO CONFIRM PLACEMENT. THE NEXT DAY, AXILLARY IMPELLA CP CONTINUED ON P-8 AND CONTINUOUS SUCTION ALARMS CONTINUE. AUDIBLE ALARMS TURNED OFF. FLOWS OF 2.5. URINE OUTPUT IS CLEAR/YELLOW. FOLLOW UP THE NEXT DAY NOTED IMPELLA MCS CONTINUED WITH PATIENT'S PROGNOSIS AS GUARDED. ECPELLA (EXTRACORPOREAL MEMBRANE OXYGENATION AND IMPELLA MCS) WAS CONSIDERED BUT THE PATIENT WAS NOT A CANDIDATE FOR ECMO. THE PLAN WAS TO WEAN THE PATIENT AND ASSESS THE PATIENT'S RESPONSE. HOWEVER, THE PATIENT WOULD EXPIRE; CARE WAS WITHDRAWN. MEDICAL SAFETY REVIEW OF THE EVENT DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE REPLACEMENT IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF THREE DEVICES RELATED TO THE PATIENT'S IMPLANT EXPERIENCE. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-26083 FOR THE IMPELLA 5.5 AND REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE INITIALLY IMPLANTED IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT ON 23-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). Manufacturer Narrative: THE INVESTIGATION FOR THE VENTRICULAR FIBRILLATION (VF) HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE VF COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21398808&pc=NIQ | 2124215-2025-09030 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21399016&pc=LOX | 2124215-2025-09000 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21399046&pc=OZD | 1220648-2025-26173 | 2025-01-20 05:00 | Death | ABIOMED, INC. | 2025-02-18 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED A HEMORRHAGIC STROKE AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT WAS ADMITTED IN FOR AN EMERGENCY PERCUTANEOUS CORONARY INTERVENTION TO THE RIGHT CORONARY ARTERY APPROXIMATELY 15 DAYS PRIOR TO THE IMPELLA 5.5 DEVICE IMPLANT PROCEDURE. THE NEXT DAY FOLLOWING THE PCI, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION AND UNDERWENT PCI TO THE LEFT ANTERIOR DESCENDING ARTERY. INTRAOPERATIVE BLOOD CLOTS AROSE AND LEFT MAIN ARTERY INFARCTION OCCURRED; EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND INTRA-AORTIC BALLOON PUMP (IABP) WERE PLACED. SOMETIME THEREAFTER, THE PATIENT WAS TAKEN OFF ECMO AND IABP AND CRASHED IN A FEW HOURS. THE PATIENT WAS PLACED ON ECMO AND IABP AGAIN. EVENTUALLY, THE PATIENT WAS PLACED ON THE IMPELLA 5.5 DEVICE FOR MCS. NINE DAYS LATER AFTER START OF THE IMPELLA 5.5 MCS, HOWEVER, THE PATIENT WOULD EXPERIENCE A HEMORRHAGIC STROKE. THE TEAM WAS UNABLE TO MAINTAIN THE PATIENT'S HEMODYNAMICS, AND THE PATIENT WOULD EXPIRE THIS DAY. ACTIONS TAKEN WERE USE OF SODIUM BICARBONATE IN THE PURGE, AND BLOOD COAGULATION MANAGEMENT. IT WAS FURTHER NOTED ACTIVATED CLOTTING TIME WAS NOT EXCESSIVELY PROLONGED, AND THE HEMORRHAGE WAS AN UNPREDICTABLE ACCIDENT. THE RELATIONSHIP BETWEEN THE IMPELLA DEVICE AND STROKE IS NOTED TO BE UNKNOWN, AND THE PATIENT EXPIRED FROM MULTI-ORGAN FAILURE DUE TO THE CEREBRAL HEMORRHAGE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR. PER THE REPORT, THE RELATIONSHIP BETWEEN THE DEVICE AND HEMORRHAGIC STROKE IS UNKNOWN AND THEREFORE THE DEMISED OUTCOME CANNOT BE DISSOCIATED. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 THE SELECTION LIFE-THREATENING WAS INADVERTENTLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26173 H3 YES WAS INADVERTENTLY SELECTED FOR EVALUATION SUMMARY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26173, THE INVESTIGATION RESULTS WERE NOT INCLUDED IN THE INITIAL REPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21399505&pc=LYJ | 1644487-2025-00175 | 2000-01-01 05:00 | Death | LIVANOVA USA, INC. | 2025-02-18 05:00 | LYJ | PULSE GEN MODEL UNKNOWN | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD MORE INTENSE SEIZURES AND ULTIMATELY PASSED AWAY FROM A SEIZURE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21399577&pc=LOX | 2124215-2025-09022 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE; CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21399979&pc=NIQ | 2124215-2025-09019 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21400097&pc=LOX | 2124215-2025-08995 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Perforation of Vessels; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, NEW REQUIREMENT FOR DIALYSIS, OTHER VASCULAR COMPLICATIONS REQUIRING TREATMENT, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21400144&pc=DSQ | 2916596-2025-00805 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY.Event Description: THE CAUSE OF DEATH WAS UNKNOWN. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21400270&pc=LOX | 2124215-2025-09021 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE; CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21400531&pc=FSA | 2183887-2025-00001 | 2025-01-11 05:00 | Death | EZ WAY INC | 2025-02-18 05:00 | FSA | 500 LB SMART LIFT WITH SCALE | Use of Device Problem; Separation Problem | Insufficient Information |  |  | 1 | Event Description: RESIDENT HAD PRIOR CARE PROVIDED BY A NURSE. WHILE STILL IN BED, THE CNA AND CMA PUT THE SLING UNDER THE RESIDENT. THE CNA AND CMA HOOKED UP THE SLING TO THE LIFT AND BACKED THE LIFT AWAY FROM THE BED. THE LEFT LEG STRAP CAME LOOSE, THE LEFT LEG WAS RELEASED AND THE RESIDENT FELL OUT OF THE SLING AND HIT HER HEAD ON THE GROUND. THE RIGHT LEG STAYED IN SLING. Manufacturer Narrative: MET WITH (B)(6). WE WENT INTO THEIR SIMULATOR ROOM THAT HAD A DUMMIE IN THE BED. THEY HAD THE DUMMIE HOOKED UP TO THE MACHINE WITH ALL 4 STRAPS DOUBLE LOOPED. THEY SAID THAT THEY DO NOT DOUBLE LOOP AND DID NOT KNOW WHY IT WAS THAT WAY. I STRESSED THE DIFFERENCES OF THE STRAPS HAD TO DO WITH POSITIONING. IN THE ROOM THEY HAD THE SLING AND LIFT THAT WAS USED IN THE EVENT. LIFT AND SLING INSPECTED - TESTED OK. *L500PS-03 - 140966, THE MACHINE DID NOT HAVE ANY VISIBLE ISSUES. 50578 - 162245, MEDIUM MESH SLING WITH HEAD SUPPORT, DID NOT SEE ANY ISSUES WITH THE SLING- ALL LOOPS WERE INTACT. OPERATION MANUAL, PG 6. STEP 4 - NUMBER 2.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21400779&pc=DSQ | 2916596-2025-00808 | 2025-01-04 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Anemia; Hematoma; Renal Failure; Multiple Organ Failure; Gastrointestinal Hemorrhage; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT BECAUSE OF NERVE PAIN, THE PATIENT'S PAIN MEDICATION WAS CHANGED/INTENSIFIED FROM THE BEGINNING OF AUTUMN. THEIR INTERNATIONAL NORMALIZED RATIO (INR) INCREASED DESPITE THE ANTICOAGULATION DOSE REDUCTION. ON (B)(6) 2025 THE PATIENT PRESENTED TO THE HOSPITAL DUE TO MELENA, EXTENSIVE RETROPERITONEAL HEMATOMAS, AND CONTINUOUS ANEMIA. THE NEED FOR RED BLOOD CELL TRANSFUSIONS WAS FOUND. THE PATIENT HAD LIVER FAILURE, KIDNEY FAILURE, AND MULTIPLE ORGAN FAILURE (MOF). THE PATIENT'S LEVEL OF CONSCIOUSNESS STARTED TO FADE. THE PATIENT WAS PLACED ON PALLIATIVE SEDATION AFTER WHICH THE PUMP WAS TURNED OFF. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WORKED WELL THE WHOLE TIME. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MOF. THE OUTCOME WAS DEVICE OR THERAPY RELATED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE BLEEDING WAS IDENTIFIED ON (B)(6) 2025. THE PATIENT UNDERWENT A COMPUTED TOMOGRAPHY (CT) SCAN. THE SOURCE OF BLEEDING WAS A RETROPERITONEAL HEMATOMA. THE PATIENT WAS TREATED WITH BLOOD INFUSION AND IV KONAKION (VITAMIN K). THE BLEEDING DID NOT RESOLVE BEFORE THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RENAL AND HEPATIC DYSFUNCTION), BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, UNDER "ANTICOAGULATION", ALSO OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4 - PATIENT WEIGHT NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B5 NARRATIVE INFORMATION. HEALTH EFFECT - IMPACT CODE IMAGING REQUIRED F2203. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21400872&pc=PNJ | 9612164-2025-00825 | 2025-01-22 05:00 | Death | MEDTRONIC IRELAND | 2025-02-18 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS DIFFICULTY TO POSITIONING / MANEUVERING. THE DEVICE WAS DEPLOYED AND TESTED AND THE ELECTRICAL PARAMETERS WERE NOT ACCEPTABLE. THE DEVICE WAS SUCCESSFULLY RECAPTURED AND WITHIN MINUTES OF FINDING A NEW LOCATION, THE PATIENT EXPERIENCED A CARDIAC PERFORATION. THE PATIENT CODED DURING WHICH AN ARTERIAL LINE WAS ACHIEVED, AN ECHOCARDIOGRAM WAS PERFORMED WITH SUSPECTED PERICARDIAL EFFUSION AND A PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT STABILIZED UNTIL THEIR BLOOD PRESSURE DROPPED AND THE PATIENT CODED AGAIN. RESCUE WAS INEFFECTIVE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE APPROVED FINAL INVESTIGATION. A REVIEW OF THE DEVICE HISTORY RECORD FOUND NO DEVIATIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE MALFUNCTION WAS NOT CONFIRMED DURING EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE. Manufacturer Narrative: CORRECTION TO H6 FOR PLI 20. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401029&pc=OOB | 2124215-2025-09015 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC TAMPONADE, CARDIOGENIC SHOCK, OTHER BLEED; DEATH DURING PROCEDURE, CAUSE: CARDIOVASCULAR HEMORRHAGE. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401052&pc=OOB | 2124215-2025-09014 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401079&pc=LOX | 2124215-2025-09565 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401163&pc=DSQ | 2916596-2025-00869 | 2025-01-26 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem | Low Blood Pressure/ Hypotension; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation; Tricuspid Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. THE ETIOLOGY COULD BE MULTIFACTORIAL AND POTENTIALLY BE PROGRESSIVE RIGHT VENTRICLE (RV) FAILURE, SEVERE AORTIC REGURGITATION (AR) AND TRICUSPID REGURGITATION (TR), AND/OR LEFT VENTRICULAR OUTFLOW GRAFT OBSTRUCTION (LVOGO). THERE WAS MULTIPLE ORGAN FAILURE (MOF) (HEART, KIDNEY, LIVER) WITH PROFOUND HYPOTENSION AND ANURIA. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF THE PUMP WOULD BE EXPLANTED. THE PUMP WOULD NOT BE RETURNED.Event Description: LOG FILES WERE NOT AVAILABLE, AND AN OUTFLOW GRAFT OBSTRUCTION WAS NOT CONFIRMED. IT WAS UNKNOWN WHETHER THE PATIENT HAD A HISTORY OF RV FAILURE PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED OUTFLOW GRAFT OBSTRUCTION COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, AND HEPATIC DYSFUNCTION), DEVICE THROMBOSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 5, ¿SURGICAL PROCEDURES¿, WARNS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. IF NOT ADDRESSED, THE LVAD WILL NOT BE ABLE TO PROVIDE THE INTENDED FLOW. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5 OF THE IFU, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5 OF THE IFU, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT. THIS SECTION OUTLINES INDICATIONS OF RIGHT HEART FAILURE AND PROVIDES THE RECOMMENDED TREATMENT OPTIONS FOR PATIENTS WITH RIGHT HEART FAILURE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A. PATIENT INFORMATION UPDATED. B5 NARRATIVE INFORMATION. D4 - ADDITIONAL DEVICE INFORMATION UPDATED INCLUDING H4 - DEVICE MANUFACTURE DATE UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION A: PATIENT DATE OF BIRTH AND WEIGHT WAS NOT PROVIDED. D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401166&pc=OZD | 1220648-2025-26177 | 2025-01-25 05:00 | Death | ABIOMED, INC. | 2025-02-18 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE VIA THE LEFT FEMORAL ARTERY FOR MECHANICAL CIRCULATORY SUPPORT (MCS). TWO DAYS AFTER START OF SUPPORT IT WAS NOTED CIRCULATION COULD NOT BE MAINTAINED, AND THE PATIENT EXPIRED. THERAPY STARTED PRIOR TO IMPELLA MCS WAS NOTED AS VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). PROBLEMS WITH THE IMPELLA DEVICE OPERATION WAS NOTED AS "NONE." NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME. AS THE IMPELLA DEVICE WAS PRESENT, PROVIDING MECHANICAL CIRCULATORY SUPPORT TO THE PATIENT, WHEN THE EVENT OCCURRED OR OTHERWISE ASSOCIATED WITH THE EVENT AND NO ALTERNATIVE CAUSE FOR THE EVENT IS IDENTIFIED AS A LIKELY CAUSE OF SUCH EVENT, THE DEMISED OUTCOME CANNOT BE DISSOCIATED FROM THE IMPELLA DEVICE. PER THE REPORTER, AN AUTOPSY IS BEING COMPLETED, AND THE CAUSE IS CURRENTLY BEING INVESTIGATED. UPON RECEIPT OF ADDITIONAL INFORMATION, A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. Manufacturer Narrative: THE INVESTIGATION FOR THE UNSPECIFIED DEATH HAS BEEN COMPLETED. NEITHER THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THE ROOT CAUSE OF DEATH COULD NOT BE DETERMINED AS THE DEVICE WAS NOT RETURNED NOR SUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401357&pc=DSQ | 2916596-2025-00883 | 2025-01-31 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; Low Blood Pressure/ Hypotension; Tachycardia; Ventricular Fibrillation; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED ACUTE SHORTNESS OF BREATH, CHEST PAIN, AND HYPOTENSION. INTRAVENOUS FLUIDS AND NOREPINEPHRINE (LEVOPHED) WERE STARTED. AN ECHOCARDIOGRAM SHOWED THROMBUS AT THE AORTIC VALVE. THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION REQUIRING MULTIPLE SHOCKS. AMIODARONE, LIDOCAINE, AND MAGNESIUM WERE ADMINISTERED. A REPEAT ECHOCARDIOGRAM SHOWED EJECTION FRACTION OF LESS THAN 5%, A DILATED RIGHT VENTRICLE, AND SEVERE DYSFUNCTION. THE PATIENT WAS INTUBATED, SEDATED AND PLACED ON AN VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). INTERNATIONAL NORMALIZED RATIO (INR) ON DAY OF EVENT WAS 1.75. THE PATIENT HAD A POOR PROGNOSIS AND THE FAMILY OPTED TO WITHDRAW CARE. THE DEVICE OPERATED AS EXPECTED.Event Description: IT WAS STATED THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND THE PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE OPERATED AS EXPECTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, THROMBOEMBOLISM, ARRHYTHMIA, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 ALSO LISTS THROMBOEMBOLISM AND ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401383&pc=OOB | 2124215-2025-09580 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401580&pc=LOX | 2124215-2025-09579 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401644&pc=PYX | 1220648-2025-26163 | 2025-01-25 05:00 | Death | ABIOMED, INC. | 2025-02-18 05:00 | PYX | IMPELLA RP FLEX | Difficult to Remove; Use of Device Problem |  | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA RP FLEX AND IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT (MCS). INITIALLY, THE IMPELLA CP WAS IMPLANTED AND DURING THE PERCUTANEOUS CORONARY INTERVENTION, THE PATIENT BEGAN TO HAVE FREQUENT SUCTION ALARMS. RIGHT HEART CATHETERIZATION WAS PERFORMED AND A PULMONARY ARTERY PULSATILITY INDEX WAS CALCULATED OF 0.5. THE DECISION WAS THEN MADE TO PLACE AN IMPELLA RP FLEX FOR BIPELLA SUPPORT. WITHIN APPROXIMATELY THREE TO FOUR HOURS AFTER RETURNING TO UNIT, THE PATIENT BEGAN TO LOSE PULSATILITY ON THE IMPELLA CP. MEAN ARTERIAL PRESSURES WERE SUSTAINED IN THE 40¿S AND 50¿S DESPITE CURRENT SUPPORT AND VASOPRESSOR MEDICATIONS. THE IMPELLA CP WAS REPOSITIONED AT BEDSIDE WHICH APPEARED TO BE STUCK IN THE PAPILLARY MUSCLE. FLOWS WERE DIMINISHED AT 2L ON PERFORMANCE LEVEL P-8. IMPELLA RP FLEX FLOWS REDUCED TO AVOID PULMONARY EDEMA. DECISION WAS MADE TO TRANSPORT PATIENT TO CATHETERIZATION LAB TO CANNULATE FOR VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). AFTER ARRIVING, THE PATIENT WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS INITIATED. THE IMPELLA CP WAS REMOVED AND LEFT FEMORAL ARTERY CANNULATED FOR ECMO. THE LEFT FEMORAL VEIN WAS ALSO CANNULATED. THE IMPELLA RP FLEX WAS BEING EXPLANTED AT TIME OF ECMO CANNULATION AND INITIATION. AT TIME OF IMPELLA RP FLEX EXPLANT, AIR WAS INDUCED INTO ECMO VENOUS LINE, WHICH AIR LOCKED AND STOPPED ECMO CIRCUIT. NO PULSE WAS ASSESSED. DECISION WAS MADE TO CEASE CARDIOPULMONARY AND RESUSCITATION EFFORTS AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THE PATIENT IMPLANT EXPERIENCE. REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT (B)(6) 2025 AND PATIENT IDENTIFIER ENDING IN (B)(6) . Manufacturer Narrative: THE INVESTIGATION FOR THE REMOVAL ISSUES HAS BEEN COMPLETED. CLINICAL DETAILS STATE THAT THE PATIENT WENT INTO CARDIAC ARREST AFTER CONTINUED DECLINE ON BOTH RP AND CP SUPPORT. THE CP WAS EXPLANTED AND ECMO WAS CANNULATED ON THE LEFT FEMORALS AND INITIATED. AFTER EXPLANT OF THE RP FLEX, AIR WAS INTRODUCED INTO THE ECMO VENOUS LINE. THIS WAS CAUGHT BY THE AIR LOCK BUT THE PATIENT ULTIMATELY EXPIRED. NO PRODUCT WAS RETURNED AND LOGS WERE NOT AVAILABLE FOR THE CASE. THE CAUSE OF THE UNSPECIFIED (ALLEGATION) - DEATH AS IT RELATES TO THE REMOVAL OF THE DEVICE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND LACK OF CLINICAL DETAILS. THE INSTRUCTIONS FOR USE FOR THE IMPELLA RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER STATED THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ CORRECTIONS TO THE INITIAL MFR REPORT: G1-REPORTING CONTACT LAST NAME CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401744&pc=DSQ | 2916596-2025-01026 | 2025-01-30 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Renal Failure; Respiratory Failure; Ischemia Stroke; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT, WHILE PLACING THE BYPASS CANNULA FEMORALLY, THE GUIDE WIRE PERFORATED THE VESSEL, AND GENERAL AND VASCULAR SURGERY SCRUBBED IN. THE PATIENT REQUIRED A RIGHT VENTRICULAR ASSIST DEVICE OF UNKNOWN MANUFACTURER (RVAD) TO BE IMPLANTED. THE PATIENT HAD A COMPLICATED COURSE WITH INTRAOPERATIVE AND POSTOPERATIVE BLEEDING, AS WELL AS HEART, RENAL, AND RESPIRATORY FAILURE, AS WELL AS AN ISCHEMIC STROKE, REQUIRING BLOOD TRANSFUSIONS, AND THE FAMILY DECIDED TO WITHDRAW LVAD SUPPORT ON (B)(6)2025. IT IS UNKNOWN WHAT OTHER SYMPTOMS THE PATIENT EXPERIENCED. THE PUMP WAS DECOMMISSIONED AND HAD BEEN NOTED TO HAVE FUNCTIONED AS INTENDED. THERE WERE NO RECENT CHANGES TO THE PATIENT CONDITION OR ANTICOAGULATION STATUS THAT WOULD HAVE CONTRIBUTED TO STROKE. THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE (MSOF) ON (B)(6)2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OR THERAPY WAS NOT THOUGHT TO HAVE CONTRIBUTED TO THE MSOF. THE DEVICE WAS NOT EXPLANTED, AND NO AUTOPSY WAS PERFORMED.Event Description: LABS AND COMPUTED TOMOGRAPHY (CT) SCANS WERE PERFORMED. THE PATIENT WAS INCAPACITATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS BLEEDING, MULTIPLE TYPES OF ORGAN FAILURE, STROKE, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO VALUES, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401795&pc=DSQ | 2916596-2025-00897 | 2025-01-05 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE INFORMATION REGARDING THE CAUSE AND DETAILS LEADING UP TO DEATH WAS NOT PROVIDED BY THE HOSPITAL DUE TO DATA PRIVACY REASON. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND IT REPORTEDLY OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401873&pc=LOX | 2124215-2025-09564 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401927&pc=CAW | 2518422-2025-102123 | 2025-01-10 05:00 | Death | RESPIRONICS, INC. | 2025-02-18 05:00 | CAW | EVERFLO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K061261 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING AN EVERLO OXYGEN CONCENTRATOR. THE CAREGIVER WAS REQUESTING TO RETURN THE DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401938&pc=DSQ | 2916596-2025-00852 | 2025-01-29 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY (B)(6) 2025 DUE TO VENTRICULAR FIBRILLATION (V-FIB). WHETHER THE DEATH WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED WAS NOT PROVIDED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED OR WOULD BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21401964&pc=LOX | 2124215-2025-09559 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402314&pc=LWS | 2649622-2025-04729 | 2024-12-29 05:00 | Death | MPRI | 2025-02-18 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Muscle Weakness; Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED AND A CAUSE OF DEATH WAS PROVIDED AS SUDDEN CARDIAC DEATH. IT WAS NOTED THAT THE PATIENT PRESENTED TO DIALYSIS FEELING WEAK AND COLLAPSED OUTSIDE THE BUILDING. CARDIOPULMONARY RESUSCITATION WAS STARTED WITH MULTIPLE EXTERNAL DEFIBRILLATIONS AND ROUNDS OF EPINEPHRINE. AT THE EMERGENCY DEPARTMENT, THE PATIENT REMAINED IN VENTRICULAR FIBRILLATION (VF) AND TWO MORE ROUNDS WERE PERFORMED. POINT OF CARE ULTRASOUND WAS PERFORMED WHICH REVEALED COMPLETE CARDIAC STANDSTILL AND CODE WAS CALLED AT THAT POINT. THE DISPOSITION OF THE DEVICE SYSTEM REMAINS UNKNOWN. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID DDMB2D1 ICD IMPLANTED: (B)(6) 2019. EXPLANTED: (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402338&pc=GDW | 3005075853-2025-01280 | 2022-02-04 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-02-18 05:00 | GDW | ECHELON CONTOUR W/ GREEN RELOAD | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Anemia | K200420 |  | 1 | Event Description: IT WAS REPORTED VIA CLINICAL TRIAL PATIENT (B)(6): 11001-089 EXPERIENCE, ANEMIA. RELATIONSHIP TO STUDY DEVICE: POSSIBLY RELATED. Manufacturer Narrative: (B)(4). DATE SENT 2/18/2025. D4 BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. INVESTIGATION SUMMARY: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. ADDITIONAL INFORMATION RECEIVED: STUDY ID (B)(6). SITE ID 11001. SUBJECT ID (B)(6). AE SEQUENCE NUMBER 1. ADVERSE EVENT TERM ANEMIA. START DATE (B)(6) 2022. SEVERITY MILD. SERIOUS AE NO. SERIOUS: DEATH 0. DATE OF DEATH 0. SERIOUS: LIFE THREATENING INJURY 0. SERIOUS: PERMANENT IMPAIRMENT 0. SERIOUS: IN-PATIENT/PROLONGED HOSPITALIZATION. HOSPITALIZATION ADMISSION DATE (B)(6) 2022. DISCHARGE DATE (B)(6) 2022. SERIOUS: MEDICAL OR SURGICAL INTERVENTION 0. SERIOUS: FETAL DISTRESS OR DEATH 0. RELATIONSHIP TO STUDY DEVICE POSSIBLY RELATED. PATIENT FINAL DATA. STUDY ID (B)(6). SITE ID 11001. SUBJECT ID 89. AGE IN YEARS 85. PATIENTS AGE 90 OR OLDER 0. GENDER OF THE SUBJECTS 1 - MALE. BMI 22. ASA SCORE AT PRE-OPERATION II - A PATIENT WITH MILD SYSTEMIC DISEASE. ETHNICITY OF THE PATIENT 2 - NOT HISPANIC OR LATINO. SMOKING HISTORY 3 - NEVER. TNM CLASSIFICATION FOR T T3. TNM CLASSIFICATION FOR N N2. TNM CLASSIFICATION FOR M M0. RESIDUAL TUMOR CLASSIFICATION R0 - NO RESIDUAL TUMOR. NUMBER OF TIMES HAD ABDOMINAL SURGERY IN THE PAST 0 - NONE. PREOPERATIVE CHEMOTHERAPY IN THE PAST 6 MONTHS NO. DURATION OF PREOPERATIVE CHEMOTHERAPY (MONTHS) 0. PREOPERATIVE RADIATION IN THE PAST 6 MONTHS NO. DURATION OF PREOPERATIVE RADIATION (WEEKS) 0. DOES THE PATIENT HAS ANY COMORBIDITY YES. ICD-10-CM CODE FOR COMORBIDITY I10 HYPERTENSION. ICD-10-CM CODE FOR COMORBIDITY N28.9 KIDNEY DISEASE. ICD-10-CM CODE FOR COMORBIDITY 0. ICD-10-CM CODE FOR COMORBIDITY 0. ECHELON CONTOUR CURVED CUTTER PRODUCT CODE - 1ST GCS40G. ECHELON CONTOUR CURVED CUTTER PRODUCT CODE - 2ND 0. WAS ANASTOMOSIS PERFORMED? NO. WAS A CIRCULAR STAPLER USED TO COMPLETE THE ANASTOMOSIS? NO. PRODUCT CODE FOR CIRCULAR STAPLER USED N/A. TOTAL NUMBER OF RELOADS FOR ECHELON1 0. TOTAL NUMBER OF RELOADS FOR ECHELON2 0. ESTIMATED BLOOD LOSS IN VOLUME IN ML 100-200 ML. ANY INTRAOPERATIVE BLOOD TRANSFUSION YES. ANY POSTOPERATIVE BLOOD TRANSFUSION YES. DATE OF OPERATION PERFORMED (B)(6) 2022. DATE OF HOSPITAL DISCHARGE (B)(6) 2022. OPERATION PERFORMED 0. OPERATION PERFORMED AS OTHER, PLEASE SPECIFY 0. PRIMARY INDICATION FOR CURRENT SURGERY 4 -COLORECTAL DISEASE. IF THE PRIMARY INDICATION FOR CURRENT SURGERY IS OTHER, PLEASE SPECIFY 0. WAS THE PROCEDURE PLANNED OR EMERGENCY? PLANNED. CURRENT PROCEDURAL TERMINOLOGY (CPT) CODE 44143 - COLECTOMY, PARTIAL; WITH END. COLOSTOMY AND CLOSURE OF DISTAL SEGMENT (HARTMANN TYPE PROCEDURE). CURRENT PROCEDURAL TERMINOLOGY (CPT) CODE 0. CURRENT PROCEDURAL TERMINOLOGY (CPT) CODE 0. STAPLE LINE LEAKAGE NO. DATE OF STAPLE LINE LEAKAGE 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE FOR K63.2 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE K65.0 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE K65.1 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE N73.0 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE OW9JOZ 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE Y83.2 0. ICD-10-CM CODE FOR STAPLE LINE LEAKAGE K91.89 0. OTHER ICD-10-CM FOR STAPLE LINE LEAKAGE 0. REOPERATION OR REVISION SURGERY WITHIN 30-DAYS NO. DATE OF REOPERATION OR REVISION SURGERY 0. STAPLE LINE BLEEDING OR HEMORRHAGE 0. DATE OF STAPLE LINE BLEEDING OR HEMORRHAGE NO. ICD-10-CM CODE FOR POSTPROCEDURAL HEMORRHAGE 0. ICD-10-CM CODE FOR HEMATOMA 0. OTHER ICD-10-CM CODE FOR HEMORRHAGE OR HEMATOMA, PLEASE SPECIFY 0. STRICTURE AND STENOSIS NO. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: (B)(4). DATE SENT 3/11/2025. CORRECTED DATA: B2, H1. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WERE THE INDICATIONS FOR SURGERY? COLORECTAL DISEASE DID THE PATIENT RECEIVE ANY PREOPERATIVE CHEMOTHERAPY OR RADIATION? NO WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? NOT REPORTED WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. NOT REPORTED WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? 100 - 200 ML. WAS A TRANSFUSION REQUIRED? YES, BLOOD TRANSFUSION PERFORMED. HOW WAS THE BLEEDING ADDRESSED? BLOOD TRANSFUSION. WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? PREOP ANTIC 5000 UNITS SQ HEPARIN, BEFORE SURGERY. POST OP LOVENOX QD. WAS THE BLEEDING INTRALUMINAL OR EXTRALUMINAL? INTRALUMINAL. WHAT IS THE CURRENT PATIENT STATUS? DECEASED. WHEN WAS THE STAPLE RETAINING CAP REMOVED? IN SURGERY. WHAT PROVISIONS DID THE SURGEON TAKE TO ENSURE THE TISSUE WAS LYING EVENLY WITHIN THE JAWS? CHECKED TWICE. WAS THE DEVICE DIFFICULT TO CLOSE? NO. WAS THE DEVICE CLOSED AND REPOSITIONED MULTIPLE TIMES PRIOR TO FIRING? NO. WAS THE DEVICE DIFFICULT TO FIRE? NO. WAS THE DISTAL TRANSECTION COMPLETED IN ONE OR MULTIPLE FIRINGS? ONE. WHAT IS THE SCRUB'S EXPERIENCE WITH RELOADING THE DEVICE? IT IS OUR ROUTINE THAT THE SURGEON RELOADS THE DEVICE. WERE THERE ANY DIFFICULTIES LOADING THE DEVICE? NO. WHAT DID THEY DO TO ENSURE THAT THE CARTRIDGE WAS SNAPPED IN PLEASE PRIOR TO CLOSING THE DEVICE? SURGEON DOUBLE CHECKS. WAS THE RETAINING PIN PLACED MANUALLY OR AUTOMATICALLY? MANUALLY. IN THE PHYSICIAN'S OPINION, WHY IS THIS EVENT PROBABLY RELATED TO THE DEVICE AND PROCEDURE? STAPLERS ARE NOT HEMOSTATIC AND BLEEDING AFTER ANASTOMOSIS IS A COMMON AE AFTER SURGERY. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHY DID THE PATIENT EXPIRE? DOES THE SURGEON BELIEVE THAT THE ALLEGED ISSUES WITH THE DEVICE CAUSED THE DEATH OF THE PATIENT? IF YES, WHY? WOULD THE SURGEON LIKE A CALL WITH ETHICON? Manufacturer Narrative: (B)(4). DATE SENT; 4/2/2025. CORRECTED DATA: B1, B2, H1. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHY DID THE PATIENT EXPIRE? SEPSIS, B CELL LYMPHOMA - DOES THE SURGEON BELIEVE THAT THE ALLEGED ISSUES WITH THE DEVICE CAUSED THE DEATH OF THE PATIENT? NO. IF YES, WHY? - WOULD THE SURGEON LIKE A CALL WITH ETHICON? NO. UPON REVIEW OF THE INFORMATION PROVIDED, IT WAS CONCLUDED THAT THIS EVENT DOES NOT MEET THE FDA DEFINED CRITERIA FOR A REPORTABLE EVENT AND IS BEING CONSIDERED A NON-PRODUCT ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402469&pc=DSQ | 2916596-2025-01203 | 2024-09-25 04:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON HOSPICE, AND THEY PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ALTHOUGH THE CAUSE OF DEATH WAS NOT REPORTED, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402590&pc=NVN | 2182208-2025-00866 | 2024-06-24 04:00 | Death | MEDTRONIC, INC. | 2025-02-18 05:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ATRIOVENTRICULAR NODE ABLATION AFTER LEFT BUNDLE BRANCH AREA PACING (LBBAP). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE IS NO ALLEGATION OF LEAD-DEATH RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND LEAD-RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS, SEVERE TRICUSPID REGURGITATION, AND PALPITATIONS. THERE WERE LEADS WHICH EXHIBITED DISLODGEMENTS AND LBBAP IMPLANT FAILURES FOR UNKNOWN REASONS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS FEMALE/77 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SINGLE-CENTER EXPERIENCE OF EFFICACY AND SAFETY OF ATRIOVENTRICULAR NODE ABLATION AFTER LEFT BUNDLE BRANCH AREA PACING FOR THE MANAGEMENT OF ATRIAL FIBRILLATION. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY. 2024. 67:1865¿1876. DOI: 10.1007/S10840-024-01847-2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402596&pc=LWS | 2649622-2025-04733 | 2025-01-31 05:00 | Death | MPRI | 2025-02-18 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD APPEARED TO BE OVERSENSING AND UNDERSENSING WHICH MAY HAVE DELAYED DETECTION OF A TREATED VENTRICULAR FIBRILLATION (VF) EPISODE. THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT APPROXIMATELY NINE DAYS LATER THE PATIENT DIED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT HAS NOT BEEN RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 429888 LEAD, IMPLANTED (B)(6) 2019. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402615&pc=CAW | 2518422-2025-017843 | 2024-02-27 05:00 | Death | RESPIRONICS, INC. | 2025-02-18 05:00 | CAW | EVERFLO | Insufficient Information | Insufficient Information | K061261 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402698&pc=DSQ | 2916596-2025-00929 | 2024-07-10 04:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE OUTCOME WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE WORKED AS INTENDED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. SECTION E1: REPORTER ADDRESS LINE 2: (B)(6) . Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. CAUSE OF DEATH WAS NOT PROVIDED FOR DATA PROTECTION REASONS. AN AUTOPSY WAS NOT PERFORMED, AND IT WAS REPORTED THAT THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR THE DRIVELINE WERE ALSO REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402701&pc=LOX | 2124215-2025-09545 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402855&pc=DSQ | 3007042319-2025-00521 | 2025-02-11 05:00 | Death | HEARTWARE, INC. | 2025-02-18 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Pumping Problem | Tachycardia; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT WITH A HISTORY OF CARDIOMYOPATHY AND REFRACTORY VENTRICULAR TACHYCARDIA (VT) DESPITE MAXIMAL MEDICAL MANAGEMENT, EXPERIENCED HIGH WATT ALARMS WITH WATTS INCREASING FROM 7 TO 10, AND A SUSPECTED IMPELLER THROMBUS. THE PATIENT HAD SUBTHERAPEUTIC ANTICOAGULATION WITH AN INTERNATIONAL NORMALISED RATIO (INR) OF 1.8, BELOW THE GOAL OF 2.5-3.5, DUE TO VARIABLE ORAL INTAKE WHILE ON HOSPICE CARE FOR THE PAST FOUR MONTHS. THE PATIENT AND FAMILY DECLINED FURTHER HOSPITAL EVALUATION AND TREATMENT. THE FAMILY DECIDED TO WITHDRAW CARE AND TURN OFF THE VAD AT HOME, RESULTING IN THE PATIENT'S DEATH. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PER FORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE CONTROLLER LOG FILES WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED HIGH-POWER EVENT COULD NOT BE CONFIRMED. OF NOTE, THE SITE INDICATED THAT CARE WAS WITHDRAWN AND THE VAD TURNED OFF, RESULTING IN THE PATIENT¿S DEATH. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE REPORTED HIGH-POWER EVENT INCLUDING BUT NOT LIMITED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION, PATIENT RELATED FACTORS, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, CARDIAC ARRHYTHMIA, THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON A REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THE PATIENT HAD A HISTORY OF CARDIAC ARRHYTHMIA. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21402861&pc=LOX | 2124215-2025-09571 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED HEART FAILURE, CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403091&pc=NGV | 2135147-2025-00814 | 2024-12-28 05:00 | Death | ABBOTT MEDICAL | 2025-02-18 05:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A 25MM AMPLATZER AMULET OCCLUDER WAS SUCCESSFULLY IMPLANTED WITHOUT ISSUE. THE FOLLOWING DAY, THE PATIENT SUFFERED CARDIAC ARREST. CPR WAS PERFORMED WITHOUT SUCCESS AND THE PATIENT PASSED AWAY. IT'S UNKNOWN WHAT THE CAUSE OF CARDIAC ARREST WAS.Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A 25MM AMPLATZER AMULET OCCLUDER WAS SUCCESSFULLY IMPLANTED WITHOUT ISSUE. THE FOLLOWING DAY, THE PATIENT SUFFERED CARDIAC ARREST. CPR WAS PERFORMED WITHOUT SUCCESS AND THE PATIENT PASSED AWAY. IT'S UNKNOWN WHAT THE CAUSE OF CARDIAC ARREST WAS. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF CARDIAC ARREST AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. THE PROVIDED IMAGING WAS REVIEWED BY MEDICAL AFFAIRS. BASED ON THE INFORMATION RECEIVED, A CAUSE FOR THE REPORTED CARDIAC ARREST AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403157&pc=OOB | 2124215-2025-09537 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403327&pc=OOB | 2124215-2025-09493 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT MYOCARDIAL INFARCTION, CARDIAC ARREST AND DEATH OCCURRED. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403378&pc=LYJ | 1644487-2025-00186 | 2005-05-01 04:00 | Death | LIVANOVA USA, INC. | 2025-02-18 05:00 | LYJ | PULSE GEN MODEL UNK | Adverse Event Without Identified Device or Use Problem | Paralysis; Convulsion, Clonic; Post Operative Wound Infection | P970003 |  | 1 | Event Description: SCIENTIFIC ARTICLE WAS RECEIVED AND REVIEWED WHERE IT WAS REPORTED THAT 3.8% OF THE PATIENTS IN THIS STUDY EXHIBITED INCREASED SEIZURE FREQUENCY. INFECTIONS WERE SEEN IN 2.6% OF THE PATIENTS AND 1.1% OF THE PATIENTS HAD VOCAL CORD PARALYSIS. ERYTHEMA IN THE INSERTION SITE, KELOID, DIZZINESS AND MALIGNANT HYPERTHERMIA-LIKE EPISODE IN ONE PATIENT EACH. THIRTY-FOUR PATIENTS (18.0 %) DISCONTINUED VNS. THE VNS DEVICE WAS REMOVED IN NINE (4.8 %) PATIENTS, WITH FIVE CASES ATTRIBUTED TO INFECTION AND THE REMAINING FOUR DUE TO LACK OF RESPONSE. FOURTEEN PATIENTS DIED DURING THE OVERALL FOLLOW-UP. BASED ON THE DEATH STATISTICS, THE CAUSE OF DEATH WAS EPILEPSY OR STATUS EPILEPTICUS IN SIX PATIENTS. ONE SUDDEN UNEXPECTED DEATH (SUDEP) WAS IDENTIFIED. NO PATIENT IDENTIFYING INFORMATION WAS PROVIDED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY HAVE NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS¿ OR MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE. HEALTH EFFECT - CLINICAL CODE: E010904, HEALTH EFFECT - CLINICAL CODE: E010901, HEALTH EFFECT - IMPACT CODE: F1202, HEALTH EFFECT - IMPACT CODE: F02.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403415&pc=DSQ | 2916596-2025-01194 | 2025-01-29 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Paralysis; Coma; Hemorrhagic Stroke; Ischemia Stroke; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT, IMPLANTED ON (B)(6) 2025 WITH THE HEARTMATE 3 (HM3) FOLLOWED SIMPLE COMMANDS AND MOVED ALL EXTREMITIES UNTIL 1630 ON (B)(6) 2025. A STROKE ALERT WAS CALLED FOR NOT FOLLOWING COMMANDS AND BEING UNRESPONSIVE TO PAIN. UPON INITIAL EXAMINATION, THE PATIENT WAS UNRESPONSIVE, NOT FOLLOWING COMMANDS. THE PATIENT WAS SUBSEQUENTLY OBSERVED SPONTANEOUSLY MOVING THEIR RIGHT SIDE AND FOLLOWING COMMANDS ON THE RIGHT (SQUEEZING HAND) HOWEVER NOT ON THE LEFT. STAT COMPUTED TOMOGRAPHY (CT) OF THE HEAD WAS PERFORMED, REVEALING MIDDLE CEREBRAL ARTERY (MCA) OCCLUSION. THE PATIENT WAS URGENTLY TAKEN TO INTERVENTIONAL RADIOLOGY (IR) FOR A MECHANICAL THROMBECTOMY. POST-THROMBECTOMY, THE PATIENT HAD A WORSENING OF MENTAL STATUS AND FOLLOW UP CT REVEALED A VERY LARGE HEMISPHERIC INFARCT WITH SIGNIFICANT EDEMA AND MIDLINE SHIFT. SIGNIFICANT DISCUSSIONS OCCURRED REGARDING EXPECTED OUTCOMES FROM THE STROKE AND CONCERN FOR WORSENING EDEMA. THE DECISION WAS MADE TO PURSUE COMFORT MEASURES ONLY. LOG FILES WERE SUBMITTED FOR REVIEW FOR ANY ABNORMITIES. THE EVENT LOG DISPLAYED (1) TRANSIENT LOWER FLOW ON (B)(6) 2025 AT 2:40 AM WHICH SEEMS TO BE PHYSIOLOGICAL IN NATURE WHERE THE LOW FLOW ESTIMATOR DROPPED BELOW 2.5 LITERS PER MINUTE (LPM). THE LOWER FLOW WAS NOT SUSTAINED LONG ENOUGH (AT LEAST 10 SECONDS) TO ACTIVATE THE AUDIBLE ALARM. THE PUMP WAS SEEING A REDUCTION IN BLOOD VOLUME. 101 PULSATILITY INDEX (PI) EVENTS WERE CAPTURED IN THE EVENT LOG. THERE WERE NO OTHER UNUSUAL EVENTS SEEN WITHIN THE LOG FILES. THE PATIENT PASSED AWAY FOLLOWING WITHDRAW OF LIFE SUPPORT AND HM3 ON (B)(6) 2025. THE DEVICE OPERATED AS EXPECTED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT NOT BEING STARTED ON ANTICOAGULATION MEDICINE POST INITIAL IMPLANT ON (B)(6) 2025 MAY HAVE CONTRIBUTED TO THEIR DEATH. IT WAS ALSO REPORTED THAT THE PATIENT WAS FOUND TO HAVE RIGHT MIDDLE CEREBRAL ARTERY (R-MCA) OCCLUSION. ADDITIONAL NOTES WERE PROVIDED ABOUT THE PATIENT'S THROMBECTOMY THAT REVEALED THE PATIENT HAD RIGHT ICA, MCA, PCA AND ACA CLOTS, MULTIPLE MEDIUM-SIZED CLOTS REMOVED ON THE STENT RETRIEVER PASSES. AROUND 05:00, THE PATIENT HAD SIGNIFICANT NEUROLOGICAL CHANGES AND PROCEEDED TO WITHDRAW TO PAIN ON THE RIGHT UPPER AND LOWER EXTREMITIES, WITH NO MOVEMENT ON THE LEFT SIDE, LEFT SIDE GAZE DEVIATION, ALONGSIDE NO GAG OR COUGH REFLEX. THE EVENT WAS NOT CONSIDERED TO BE DEVICE RELATED. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403492&pc=NPT | 2025587-2025-01302 | 2025-01-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-18 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THE VLV EVPROPLUS-23, THE PATIENT EXPIRED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403522&pc=LWR | 2025587-2025-01301 | 2024-11-15 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-18 05:00 | LWR | FREESTYLE AORTIC ROOT BIOPROSTHESIS | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure; Insufficient Information | P970031 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT POST IMPLANT OF THIS AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT DIED AT HOME AND THE CAUSE OF DEATH WAS REPORTED AS CHRONIC SYSTOLIC HEART FAILURE. IT WAS STATED THAT THE PHYSICIAN DOES NOT BELIEVE THE MEDTRONIC PRODUCT CAUSED OR CONTRIBUTED TO DEATH AND NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CHANGED THE EVENT DESCRIPTION AND DEVICE RELATEDNESS OF THIS PREVIOUSLY REPORTED EVENT. THERE IS NO EVIDENCE TO SUGGEST THE DEVICE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY. B5 AND H6 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED A2 UPDATED D1 UPDATED D4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403620&pc=DSQ | 2916596-2025-00878 | 2025-01-27 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403681&pc=LOX | 2124215-2025-09546 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403707&pc=OOB | 2124215-2025-09553 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21403789&pc=LOX | 2124215-2025-09534 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT MYOCARDIAL INFARCTION, CARDIAC ARREST AND DEATH OCCURRED. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21404047&pc=DSQ | 2916596-2025-00882 | 2025-01-31 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED AND WOULD BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21404147&pc=LWS | 9614453-2025-00687 | 2024-12-29 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-18 05:00 | LWS | EVERA MRI XT DR SURESCAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Muscle Weakness; Ventricular Fibrillation | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED AND A CAUSE OF DEATH WAS PROVIDED AS SUDDEN CARDIAC DEATH. IT WAS NOTED THAT THE PATIENT PRESENTED TO DIALYSIS FEELING WEAK AND COLLAPSED OUTSIDE THE BUILDING. CARDIOPULMONARY RESUSCITATION WAS STARTED WITH MULTIPLE EXTERNAL DEFIBRILLATIONS AND ROUNDS OF EPINEPHRINE. AT THE EMERGENCY DEPARTMENT, THE PATIENT REMAINED IN VENTRICULAR FIBRILLATION (VF) AND TWO MORE ROUNDS WERE PERFORMED. POINT OF CARE ULTRASOUND WAS PERFORMED WHICH REVEALED COMPLETE CARDIAC STANDSTILL AND CODE WAS CALLED AT THAT POINT. THE DISPOSITION OF THE DEVICE SYSTEM REMAINS UNKNOWN. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 6935 (B)(6); PRODUCT TYPE: 0264-LEAD-HV; IMPLANT DATE (B)(6) 2012; EXPLANT DATE (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21404354&pc=DSQ | 2916596-2025-00984 | 2023-08-12 04:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE INFORMATION REGARDING THE CAUSE AND DETAILS LEADING UP TO DEATH WAS NOT PROVIDED BY THE HOSPITAL DUE TO DATA PRIVACY REASON. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND IT REPORTEDLY OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. SECTION E1: REPORTER ADDRESS LINE 2: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21404655&pc=LOX | 2124215-2025-09560 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIAC ARREST, DEATH AT DISCHARGE, CAUSE: ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21404662&pc=NPT | 2135147-2025-00815 | 2025-01-24 05:00 | Death | ABBOTT MEDICAL | 2025-02-18 05:00 | NPT | NAVITOR VISION | Difficult to Fold, Unfold or Collapse | Cardiac Arrest | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). THE ANATOMY OF THE PATIENT WAS: PERIMETER OF 80MM, MINIMUM DIAMETER OF 22.6, MAX DIAMETER OF 29.1, AND AN AVERAGE OF 25.8. DURING THE PROCEDURE, A PRE-BAV WAS PERFORMED USING A 24MM NON-ABBOTT BALLOON. THE PHYSICIAN DIDN'T THINK THE BALLOON WAS FULLY EXPANDED, SO THE BAV WAS PERFORMED AGAIN WITH BETTER EXPANSION. THE DS WAS THEN POSITIONED IN THE ANNULUS, AND WHEN THE VALVE WAS STARTING TO BE DEPLOYED, IT WAS NOTICED THAT IT WAS BIASING THE INNER CURVE AND WAS ALSO CONSTRAINED. THE SYSTEM WAS FULLY RECAPTURED AND TAKEN BACK INTO THE DESCENDING AORTA. THE GOAL WAS TO TRY TO GET THE NON-ABBOTT GUIDEWIRE TO SIT BETTER ACROSS THE VALVE. AGAIN THE SYSTEM WAS BROUGHT TO THE ANNULUS AND DEPLOYED, BUT AT 80% IT WAS STILL NOTICED TO BE SIGNIFICANTLY CONSTRAINED DUE TO CALCIFICATION. AS RECAPTURING STARTED, THE PATIENT STARTED TO DECOMPENSATE. THE VALVE WAS FULLY RECAPTURED AND PLACED IN THE DESCENDING AORTA AND THEN CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. A SECOND VALVE AND DELIVERY SYSTEM WAS PREPPED IN CASE THEY WANTED TO PROCEED WITH THE TAVR BUT THIS NEVER OCCURRED. THE PATIENT WAS LATER TRANSITIONED TO ECMO, AND SHORTLY AFTER THAT, THE PATIENT PASSED AWAY. THE PHYSICIAN ALLEGED THE CAUSE OF DEATH BEING RELATED TO THE PROCEDURE AND THE PATIENTS' COMORBIDITIES INCLUDING SEVERE AORTIC STENOSIS, KIDNEY DISEASE WITH DIALYSIS AND DISCUSSION ON HOSPICE CARE PRIOR TO THE IMPLANT PROCEDURE. IT WAS NOTED THAT AORTIC INSUFFICIENCY WAS NOT IMMEDIATELY NOTICED AFTER THE BAV HOWEVER, THERE WAS SEVERE LEAK AROUND THE VALVE AT 80% DEPLOYMENT. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF VALVE UNDEREXPANSION, CARDIAC ARREST, AND DEATH WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE CASE AND PROVIDED IMAGES WERE REVIEWED BY AN ABBOTT SENIOR DIRECTOR OF MEDICAL AFFAIRS. BASED ON ALL AVAILABLE INFORMATION, THE REPORTED VALVE UNDEREXPANSION APPEARS TO BE RELATED TO CHALLENGING PATIENT ANATOMY (CALCIFICATION). THE REPORTED CARDIAC ARREST AND DEATH APPEAR TO BE RELATED TO A COMBINATION OF PATIENT CONDITION (SIGNIFICANT COMORBIDITIES) AND PROCEDURAL CONDITIONS (DECOMPENSATION AFTER PRE-BAV). THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS THE RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21404963&pc=BZD | 2518422-2025-102121 | 2024-12-04 05:00 | Death | RESPIRONICS, INC. | 2025-02-18 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405076&pc=DSQ | 2916596-2025-00785 | 2025-01-27 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THE PATIENT'S GENERAL CONDITION HAD DETERIORATED. THIS WAS NOT CONSIDERED DEVICE RELATED AND IT FUNCTIONED AS INTENDED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE CAUSE OF DEATH WAS JUST THE PATIENT'S GENERAL CONDITION DETERIORATING. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405203&pc=LOX | 2124215-2025-09572 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405298&pc=NIQ | 2124215-2025-08705 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE MIDDLE RIGHT CORONARY ARTERY (RCA) WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.7 MM. THE TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 3.00 MM X 24 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THREE DAYS AFTER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND TICAGRELOR. ON AN UNKNOWN DATE IN 2024, THE SUBJECT DIED. AT THE TIME OF THE EVENT, THE SUBJECT WAS ON ASPIRIN AND TICAGRELOR. NO ACTION WAS TAKEN TO TREAT THE EVENT, AND THE CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405376&pc=LOX | 2124215-2025-09533 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT MYOCARDIAL INFARCTION, CARDIAC ARREST AND DEATH OCCURRED. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405507&pc=DSQ | 2916596-2025-00905 | 2020-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Malposition of Device; Adverse Event Without Identified Device or Use Problem | Aneurysm; Fatigue; Hematoma; Intracranial Hemorrhage; Pain; Fungal Infection; Weight Changes; Thrombosis/Thrombus; Intra-Abdominal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿SUCCESSFUL MANAGEMENT OF MYCOTIC VISCERAL ANEURYSMS IN A LVAD PATIENT: NAVIGATING POSTOPERATIVE CHALLENGES¿ THAT HEARTMATE 3 MAY BE ASSOCIATED WITH PUMP THROMBOSIS, INFECTION, AND BLEEDING. THE CASE REPORT PRESENTED A 64-YEAR-OLD MALE WITH MEDICAL HISTORY OF STAGE III CHRONIC KIDNEY DISEASE, PRE-DIABETES, HYPERTENSION, CORONARY ARTERY DISEASE, SEVERE ISCHEMIC CARDIOMYOPATHY REQUIRING A HEARTMATE (HM) 2 LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT IN 2014, FOLLOWED BY A HM2-TO-HM2 EXCHANGE IN 2017, AND A HM2-TO-HM3 EXCHANGE IN 2020 DUE TO RECURRENT PUMP THROMBOSIS. THE PATIENT PRESENTED TO THE CLINIC WITH UNINTENTIONAL WEIGHT LOSS, FATIGUE, LEFT UPPER QUADRANT PAIN, LEUKOCYTOSIS, AND MULTIPLE PULSATILITY INDEX (PI) EVENTS ON LVAD INTERROGATION. A MONTH PRIOR, THEY HAD BEEN ADMITTED WITH ODONTOGENIC MAXILLARY SINUSITIS, ENDOPHTHALMITIS, AND FUNGEMIA DUE TO CANDIDA ALBICANS WHICH REQUIRED INTRAVENOUS AND CHRONIC SUPPRESSIVE ORAL ANTIFUNGAL THERAPY. CONCERNS REGARDING SYSTEMIC FUNGAL INFECTION AND POTENTIAL LVAD INVOLVEMENT PROMPTED ADMISSION. ABDOMINAL IMAGING REVEALED A LARGE COMPLEX SUPERIOR MESENTERIC ARTERY (SMA) ANEURYSM. SELECTIVE EMBOLIZATION OF THE SPLENIC ARTERY WAS PERFORMED BY INTERVENTIONAL RADIOLOGY (IR). HOWEVER, THE MYCOTIC NATURE OF THE SMA ANEURYSM AND ITS RELATION TO THE MAJOR SMA BRANCHES RENDERED ENDOVASCULAR INTERVENTION UNFEASIBLE. VASCULAR SURGERY CONSULT WAS OBTAINED, AND THE PATIENT WAS TAKEN TO THE OPERATING ROOM WHERE THE ANEURYSM AND CONCURRENT INFECTION WERE EXCISED WITH RECONSTRUCTION OF THE SMA WITH AN ENDOGENOUS FEMORAL VEIN GRAFT. CONCOMITANT SPLENECTOMY WAS ALSO PERFORMED. SURGICAL PATHOLOGY CONFIRMED INVASIVE FUNGAL PRESENCE IN BOTH TISSUES. WHILE THESE ANEURYSMS WERE SUCCESSFULLY MANAGED, THE PATIENT¿S POSTOPERATIVE CARE WAS COMPLICATED BY POST-OPERATIVE ABDOMINAL BLEEDING AND SPONTANEOUS INTRACRANIAL BLEEDING. THE PATIENT REFUSED FURTHER INTERVENTIONS AND TRANSITIONED TO HOSPICE.Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿SUCCESSFUL MANAGEMENT OF MYCOTIC VISCERAL ANEURYSMS IN A LVAD PATIENT AND NAVIGATING LVAD POSTOPERATIVE CHALLENGES?¿ IDENTIFYING THAT HM3 MAY BE RELATED TO WEIGHT LOSS, FATIGUE, THROMBUS, INFECTION, BLEEDING, HEMORRHAGIC STROKE, PUMP MALPOSITION, LOW POWER ALARMS, AND DEATH. THIS IS A CASE STUDY OF A 64-YEAR-OLD MALE WHO UNDERWENT PUMP EXCHANGE TO HM3 IN 2020. A MONTH PRIOR, THE PATIENT WAS ADMITTED DUE TO ODONTOGENIC MAXILLARY FUNGAL SINUSITIS, COMPLICATED BY ENDOPHTHALMITIS AND FUNGEMIA. INTRAVENOUS MEDICATION WAS ADMINISTERED. CONCERNS REGARDING SYSTEMIC FUNGAL INFECTION AND POTENTIAL LEFT VENTRICULAR ASSIST DEVICE (LVAD) RELATED COMPLICATIONS PROMPTED IMMEDIATE ADMISSION. COMPUTED TOMOGRAPHY IMAGING UNVEILED A SPLENIC ANEURYSM WITH ACTIVE ARTERIAL EXTRAVASATION AND A MID-SUPERIOR MESENTERIC ARTERY (SMA) ANEURYSM. WHEN DISSECTING THE ANEURYSM, AN INFECTED THROMBUS WAS NOTED. INTRAOPERATIVE ULTRASOUND WAS UTILIZED TO IDENTIFY THE OUTFLOW BRANCHES, AND THEY WERE INDIVIDUALLY DISSECTED AND SLINGED. THE PATIENT HAD MANY MEDICAL COMPLICATIONS SECONDARY TO THEIR LVAD, SPLENECTOMY, AND POOR NUTRITION. NUTRITION WAS A MAJOR ISSUE AS THEY DID NOT HAVE AN APPETITE POSSIBLY SECONDARY TO THEIR ANTIFUNGAL MEDICATION AND OR COMPRESSION BY THEIR LVAD AS THEY PREVIOUSLY HAD AN HM2, WHICH WAS NOTABLY LARGER THAN THEIR HM3; THIS COULD HAVE LED TO MALPOSITIONING OF THE DEVICE, IMPACTING STOMACH CAPACITY. THE PATIENT¿S NEED FOR ANTICOAGULATION LED TO MULTIPLE COMPLICATIONS INCLUDING A POST SPLENECTOMY SITE HEMATOMA, WHICH REQUIRED RETURNING TO THE OPERATING ROOM (OR) FOR WASHOUT AND DRAIN PLACEMENT. UPON RESUMING HEPARIN, THE PATIENT DEVELOPED A BRAIN BLEED IDENTIFIED ON POST-OPERATIVE DAY (POD) 18 WHEN THE PATIENT BEGAN HAVING LEFT ARM PARESTHESIA AND WEAKNESS. THE PATIENT AND FAMILY REQUESTED DO NOT RESUSCITATE, AND A DECISION WAS MADE TO DISCHARGE THEM HOME TO HOSPICE ON POD 21. THEY PASSED 6 DAYS LATER AT HOSPICE. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT CORRECTED DATE OF DEATH WAS NOT PROVIDED. ARTICLE TITLE: SUCCESSFUL MANAGEMENT OF MYCOTIC VISCERAL ANEURYSMS IN A LVAD PATIENT AND NAVIGATING LVAD POSTOPERATIVE CHALLENGES? KENDALL WERMINE ET AL. J VASC SURG CASES INNOV TECH 2025;11:101715. HTTPS://DOI.ORG/10.1016/J.JVSCIT.2024.101715; DEPARTMENT OF SURGERY, BAYLOR SCOTT & WHITE ALL SAINTS MEDICAL CENTER FORT WORTH, FORT WORTH, TEXAS. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), DOCUMENT REV D, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV A, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, BLEEDING, INFECTION, THROMBOEMBOLISM, AND DEATH SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 ALSO LISTS THROMBOEMBOLISM AND INFECTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", EXPLAINS HOW TO INSERT THE PUMP IN THE VENTRICLE AND INSTRUCTS THE USER TO TAKE CARE TO AVOID ORIENTING THE INLET TOWARDS THE INTERVENTRICULAR SEPTUM AS PUMP FUNCTION WILL BE COMPROMISED IN THE PRESENCE OF INLET OBSTRUCTION. THE STEPS TO ADJUST THE ORIENTATION OF THE PUMP ARE ALSO OUTLINED IN THIS SECTION. THE SERIAL NUMBER OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM IN USE WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2020 AS WHEN THE HEARTMATE 3 WAS IMPLANTED. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. WERMINE, K., GOHAR, S., MILHOAN, R., YOUREE, B., RAMARATHNAM, V., & ALQAIM, M. (2024). SUCCESSFUL MANAGEMENT OF MYCOTIC VISCERAL ANEURYSMS IN A LVAD PATIENT: NAVIGATING POSTOPERATIVE CHALLENGES. THE JOURNAL OF HEART AND LUNG. BAYLOR SCOTT AND WHITE HEALTH, FORT WORTH, TX TRANSPLANTATION, 43(4), S598. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.909. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: SECTION B2: LIFE-THREATENING CHECKED. SECTION E1: REPORTER ESTABLISHMENT NAME CORRECTED. SECTION E2: HEALTH PROFESSIONAL CORRECTED. SECTION E3: OCCUPATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED TO INCLUDE COMPANY REPRESENTATIVE AND HEALTH PROFESSIONAL. SECTION H6: DUE TO SYSTEM LIMITATIONS, THE FOLLOWING CODING WAS NOT INCLUDED IN THE INITIAL REPORT: HEALTH EFFECT - CLINICAL CODE: 4594 HIGH WHITE BLOOD CELL COUNT, HEALTH EFFECT - CLINICAL CODE: 4844 SINUSITIS, HEALTH EFFECT - CLINICAL CODE: 1835 ENDOPTHALMITIS, HEALTH EFFECT - CLINICAL CODE: 1708 ANEURYSM. SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, BLEEDING, AND INFECTION. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE SERIAL NUMBER OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM IN USE WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405592&pc=DSQ | 2916596-2025-00939 | 2025-01-10 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: G2 - REPORT SOURCE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. ADDITIONALLY, DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT PROVIDE ANY FURTHER PATIENT OUTCOME INFORMATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405615&pc=OZD | 1220648-2025-26178 | 2025-01-26 05:00 | Death | ABIOMED, INC. | 2025-02-18 05:00 | OZD | IMPELLA CP | Restricted Flow rate |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED DEVICE MALPOSITIONING AND LOW PUMP FLOW. THE PATIENT WOULD LATER EXPIRE. THE PATIENT WAS STATUS POST CORONARY ARTERY BYPASS GRAFT (CABG) SURGERY X 3 VESSELS ONE DAY PRIOR TO THE IMPELLA CP IMPLANT. CABG, THERE WAS A LEFT INTERNAL MAMMARY ARTERY AND OBTUSE MARGINAL ARTERY LIMA AND OM GRAFT FAILURE, POOR TARGETS, AND EJECTION FRACTION WAS REDUCED FROM 55-60% PRE-OPERATIVE TO 20-25% POST OPERATIVE. THE DECISION WAS MADE TO DO HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION WITH THE IMPELLA CP MCS FOR RECOVERY. THE PATIENT HAD SEVERE DISEASE IN THE LEFT MAIN, LEFT ANTERIOR DESCENDING AND LEFT CIRCUMFLEX ARTERIES. THE IMPELLA CP WAS SUCCESSFULLY IMPLANTED, AND PERCUTANEOUS CORONARY INTERVENTION WAS COMPLETED TO THE THREE VESSELS, THE PATIENT REMAINED STABLE THROUGHOUT THE PROCEDURE AND WAS TRANSPORTED TO THE UNIT ON IMPELLA MCS IN CRITICAL CONDITION. THE NEXT DAY VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WITH IMPELLA CP MCS WAS PLANNED AND PLACED WITH PERFUSION SHEATHS TO THE LOWER LEGS. FOUR DAYS LATER, IT WAS REPORTED THERE WAS A BRIEF SUCTION EVENT IN THE MORNING OF SAID DAY AFTER REPOSITIONING THE PATIENT. THE PATIENT REMAINED ON ECMO WITH FLOWS OF 4.5-5.5 LITERS PER MINUTE. THERE WERE OCCASIONAL PLACEMENT IN AORTA ALARMS AS THE MORNING CONTINUED. MOTOR CURRENT REMAINED PULSATILE. WITH SIMULTANEOUS RAMP TRIAL, BEDSIDE TRANSTHORACIC ECHOCARDIOGRAM SHOWED DEEP POSITION WITH PIGTAIL IN CAPILLARY MUSCLE ALTHOUGH IMAGES SUBOPTIMAL DUE TO BODY HABITUS. THERE WAS DECOUPLING OF WAVEFORMS WHEN PERFORMANCE LEVEL WAS INCREASED WITH ECMO FLOW REDUCED. THE DEVICE WAS REPOSITIONED PULLED BACK IN 0.5-1.0 INCREMENTS TO A DEPTH OF 4.02 CENTIMETERS ACROSS THE AORTIC VALVE AND APPROXIMATELY 10 MINUTES AFTER REPOSITIONING, IMPELLA POSITION IN AORTA ALARM PRESENTED DESPITE NO MOVEMENT FROM THE PATIENT. WAVEFORMS AND IMAGES WERE REVIEWED, DISCUSSED AND IT WAS NOTED THE IMPELLA CP PUMP POSITION APPEARED TO BE SHALLOW. THE DECISION WAS MADE NOT TO REPOSITION AGAIN BUT RATHER DISABLE THE ALARM. THE RISK OF DISABLING THE ALARM WAS DISCUSSED WITH THE PHYSICIAN AND IT WAS SUGGESTED TO CLOSELY MONITOR PULMONARY ARTERY CATHETER WAVEFORMS, URINE OUTPUT, MOTOR CURRENT, AND VASOPRESSOR REQUIREMENTS. ADDITIONALLY, IT WAS FURTHER REPORTED ON THIS SAME DAY, THE PATIENT STARTED HAVING A GASTROINTESTINAL (GI) BLEED. PLATELETS WERE LOW AND HEPARIN-INDUCED THROMBOCYTOPENIA PANEL WAS SENT WITH NEGATIVE RESULTS. ALL ANTI-COAGULATION WAS STOPPED. AN ENDOSCOPY WAS COMPLETED AND REVEALED AN UPPER GI BLEED FROM AN ACTIVE ULCER. CAUTERY WAS PERFORMED BY THE GASTROENTEROLOGIST AND SEVERAL BLOOD PRODUCTS HAD BEEN GIVEN SINCE THE START OF THE BLEED. IT WAS NOTED THE PHYSICIAN DID NOT ATTRIBUTE ANY OF THE PATIENT¿S BLEEDING TO THE IMPELLA CP DEVICE, AND OF NOTE, THERE WAS NO HEPARIN IN THE IMPELLA PURGE SYSTEM. ONE DAY LATER, ECMO WAS WEANED DOWN WITH IMPELLA CP INCREASED WITH DECANNULATION AND ESCALATION TO AN IMPELLA 5.5 PLANNED FOR THE NEXT DAY. THE NEXT DAY, THE PATIENT LOST ALL PULSATILITY AND CONTINUED TO OOZE FROM NOSE, MOUTH, AND TRIPLE LUMEN. DUE TO NO PULSATILITY, THE DECISION WAS MADE TO GO BACK UP ON ECMO AND DECREASE IMPELLA. A DISCUSSION WAS MADE WITH THE FAMILY AND THE PATIENT CONTINUED ECMO AND IMPELLA MCS WITH THE DECISION TO WITHDRAW CARE INSTEAD OF ESCALATING TO AN IMPELLA 5.5. THE PATIENT EXPIRED THE FOLLOWING DAY. THE CAUSE OF DEATH IS UNKNOWN, NOT PROVIDED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR TO THE OVERALL EVENT OUTCOME ALTHOUGH [IT SHOULD BE NOTED] ESCALATION WAS REJECTED, AND CARE WAS WITHDRAWN. Manufacturer Narrative: THE INVESTIGATION FOR THE LOW PUMP FLOW HAS BEEN COMPLETED. CLINICAL DETAILS STATES THAT SUCTION ALARMS OCCURRED WHEN ECMO WEANING. NO PRODUCT AND DATA WERE RETURNED. THE ROOT CAUSE OF LOW FLOW WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21405817&pc=BZD | 2518422-2025-101884 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-02-18 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A DREAM STATION DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21406031&pc=CAT | 3014810929-2025-00004 | 2025-01-04 05:00 | Death | Unknown Manufacturer | 2025-02-18 05:00 | CAT | UKNOWN | Use of Device Problem | Partial thickness (Second Degree) Burn | K090007 |  | 1 | Event Description: PATIENT SUFFERED FLASH BURN TO FACE WHILE SMOKING AND USING OXYGEN CONCENTRATOR. PATIENT SUSTAINED 2ND DEGREE BURNS TO FACE. PATIENT WAS TRANSFERRED TO TRAUMA CENTER AND WAS PLACED ON A VENTILATOR. PATIENT WAS EXTUBATED SAME DAY AND SUBSEQUENTLY DIED. THERE IS NO DATA PROVIDED TO SUPPORT THAT THE DEVICE(S) IN QUESTION MALFUNCTIONED. Manufacturer Narrative: ADVERSE EVENT CAUSED BY USER SMOKING. NO EVIDENCE CANNULA MADE BY JUMAO. CANNULA WAS CONNECTED TO TRUAIR-5 OXYGEN CONCENTRATOR (MADE BY JUMAO) MODEL # 1-02C5L (GENERATOR, OXYGEN, PORTABLE, PROCODE: CAW) WITH SERIAL # (B)(6). THE OXYGEN CONCENTRATOR AT ISSUE FALLS WITHIN A RANGE OF SERIALS NUMBERS SUBJECT TO RECALL. RECALL NUMBER: Z-0795-2025. RECALL INITIATION DATE: 12/09/2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21406061&pc=OOB | 2124215-2025-09552 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-18 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS, OTHER BLEED, DEATH AT DISCHARGE, CAUSE: NEUROLOGICAL. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE FIELD IS LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. NO FURTHER INFORMATION IS AVAILABLE FOR THESE EVENTS, AS THE USER FACILITY IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21406646&pc=NAY | 2955842-2025-01603 | 2024-08-12 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-02-18 05:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiogenic Shock; Respiratory Failure | K131861 |  | 1 | Event Description: A LITERATURE ARTICLE THAT DESCRIBED A RETROSPECTIVE REVIEW OF 1714 ROBOTIC ENDOSCOPIC CARDIAC SURGERIES, PERFORMED AT ONE SINGLE INSTITUTION BETWEEN SEPTEMBER 2013 AND FEBRUARY 2024, STUDIED OUTCOMES OF 550 CONSECUTIVE PATIENTS THAT UNDERWENT ROBOTIC MITRAL VALVE (MV) SURGERY. MV REPAIR WAS PERFORMED IN 98% OF PATIENTS WITH DEGENERATIVE MITRAL REGURGITATION (MR). CONCOMITANT PROCEDURES INCLUDED COX-MAZE CRYOABLATION IN 127 (23%) PATIENTS, TRICUSPID VALVE REPAIR IN 54 (9.8%), SEPTAL MYECTOMY IN 15 (2.7%), TOTALLY ENDOSCOPIC CORONARY BYPASS IN 6 (1.1%), AND AORTIC VALVE REPLACEMENT IN 3 (0.5%). ENDOAORTIC BALLOON OCCLUSION WAS USED IN 392 PATIENTS (71%), VENTRICULAR FIBRILLATORY ARREST IN 114 (21%), AND TRANSTHORACIC AORTIC CLAMP IN 44 (8%). THIRTY-DAY MORTALITY EVENTS OCCURRED IN 7 PATIENTS (1.3%); THE CAUSES OF DEATH PROVIDED INCLUDED CARDIOGENIC SHOCK REQUIRING VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION IN 4 PATIENTS, 1 PATIENT EXPERIENCED AORTIC DISSECTION WITH RESPIRATORY FAILURE, 1 PATIENT EXPERIENCED PULMONARY EMBOLISM, AND 1 PATIENT DIED FROM AN UNKNOWN CAUSE. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS REPORTED, NOR DID THE AUTHORS ALLEGED THAT ANY INTUITIVE PRODUCTS CAUSED OR CONTRIBUTED TO THE DEATHS. THE DESIGNATED AUTHOR CONTACT HAS NOT RESPONDED TO A REQUEST FOR ADDITIONAL INFORMATION.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: AN INVESTIGATION COULD NOT BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATES OR DA VINCI SYSTEM IDENTIFIERS. THERE WERE NO DEVICE ISSUES DOCUMENTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THIS REPORT IS FOR A LITERATURE REVIEW REGARDING A LARGE SERIES OF ROBOTIC CARDIAC PROCEDURES. THE ARTICLE LISTS THE COMPLICATIONS THAT OCCURRED FOLLOWING THESE PROCEDURES. THERE ARE NO ALLEGATIONS MADE OR ISSUES DESCRIBED REGARDING ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. THE COMPLICATIONS LISTED ARE WELL KNOWN TO CARDIAC SURGERY AND NOT SPECIFIC TO ANY DEVICE OR INSTRUMENTATION. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THESE EVENTS. CITATION: KITAHARA H, NISIVACO S, BHASIN R, HAMZAT I, GRADY K, BALKHY HH. 550 ROBOTIC TOTALLY ENDOSCOPIC MITRAL VALVE SURGERIES WITHIN A COMPREHENSIVE ROBOTIC CARDIAC PROGRAM. ANN THORAC SURG. 2024 AUG 24:S0003-4975(24)00685-4. DOI: 10.1016/J.ATHORACSUR.2024.08.005. EPUB AHEAD OF PRINT. PMID: 39186997. SECTION A2 AGE: THE MEAN AGE OF ALL PATIENTS WAS 63 YEARS (25-89 YRS). SECTION A3 GENDER: MALE - 61%; FEMALE - 39%. Manufacturer Narrative: THE DESIGNATED AUTHOR WAS CONTACTED AND CONFIRMED THAT THERE WERE NO DEVICE MALFUNCTIONS NOR COMPLICATIONS THAT WERE RELATED TO THE DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21406833&pc=MIH | 3007284313-2025-03775 | 2025-01-22 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-02-18 05:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Insufficient Information | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE LONG-TERM EVAR ASSESSMENT AND FOLLOW-UP (LEAF) RETROSPECTIVE STUDY PROVIDED AN INITIAL UPDATE REGARDING THE PROCEDURAL OUTCOMES OF CURRENTLY ENROLLED PARTICIPANTS FOLLOWING IMPLANT OF GORE® EXCLUDER® AAA ENDOPROSTHESIS. THIS EVENT COVERS THE DATA REPORTED FROM THE KAISER PERMANENTE EVAR REGISTRY SITES AND LINKED TO MEDICARE FEE FOR SERVICE DATA. IN THE INITIAL REPORT, 1987 SUBJECTS WERE IMPLANTED WITH THE GORE EXCLUDER C3, (B)(6)2014 TO (B)(6) 2022. IN THAT PERIOD, THE FOLLOWING NEW ADVERSE EVENTS WERE REPORTED: 21 DEATHS DURING HOSPITALIZATION (AVERAGE HOSPITAL STAY IS 2 DAYS), 25 RE-TREATMENT DURING HOSPITALIZATION. A KAPLAN-MEIER GRAPH WAS INCLUDED IN THE REPORT SHOWING FREEDOM FROM ANEURYSM RELATED MORTALITY. AT THE 5 YEAR FOLLOW-UP THE KAPLAN MEIER ESTIMATE WAS THAT APPROXIMATELY 1.5% OF SUBJECTS HAD ANEURYSM-RELATED MORTALITY. POST-OP COMPLICATIONS -GRAFT RELATED (REVISION) (2). RETREATMENT DURING HOSPITALIZATIONS INDICATIONS INCLUDES: ABDOMINAL COMPARTMENT SYNDROME (2) ISCHEMIC BOWEL (1) DISSECTION (2) ENDOLEAK (3) OCCLUSION (11) RUPTURE-2 STENOSIS-2. RE-INTERVENTIONS AT 1 YEAR FOLLOW-UP, THE FOLLOWING WAS OBSERVED: DISSECTION (3), ENDOLEAK (31), FEMORAL-FEMORAL BYPASS GRAFT (1), GRAFT INFECTION (5), MIGRATION (4), OCCLUSION-17, PSEDO ANEURYSM-2, RUPTURE-2, SAC GROWTH-1, STENOSIS-8, TREATMENT OF INTERNAL ILIAC-2. Manufacturer Narrative: UPDATED SECTION G3/G4, AND H10. THE PATIENTS REFERENCED IN THIS EVENT FILE ARE ENROLLED IN THE COLLABORATIVE LONG-TERM EVAR ASSESSMENT AND FOLLOW-UP (LEAF) POST-MARKET SURVEILLANCE PROJECT THROUGH (B)(6). ALL INFORMATION CONTAINED IN THIS EVENT FILE WAS RECEIVED FROM THE (B)(6) DATA BASE. (B)(6) IS NOT MANAGED, MAINTAINED OR SUPERVISED BY GORE OR ANY REPRESENTATIVE OF GORE. THE INITIAL INFORMATION OBTAINED FROM (B)(6) IS THE ONLY INFORMATION BEING PROVIDED. FURTHER INVESTIGATION IS NOT POSSIBLE AS IDENTIFIABLE DATA AS TO SITE(S), PHYSICIAN(S), PATIENT(S), DEVICE(S) AND/OR SPECIFIC DATE(S) ARE NOT BEING PROVIDED. THE COMPLAINT FILE WILL REFLECT LEAF AS THE STUDY NAME. IT IS LIKELY THESE INDIVIDUAL EVENTS HAVE BEEN PREVIOUSLY CAPTURED AND REPORTED BY GORE, BUT DUE TO THE NATURE OF THIS STUDY AND ANONYMITY OF THE PATIENT AND DEVICE DATA, GORE IS UNABLE TO CONFIRM THIS. FOR THIS REASON, GORE WILL SUBMIT ONE REPORT FOR MULTIPLE REPORTABLE EVENTS WHERE THE INDIVIDUALLY IDENTIFIABLE INFORMATION OF THE DEVICE IS NOT AVAILABLE AND NO ADDITIONAL INFORMATION WILL BE PROVIDED BECAUSE WE CANNOT OBTAIN ENOUGH INFORMATION ABOUT EACH IDENTIFIED PATIENT AND/OR DEVICE MENTIONED IN ORDER TO PROVIDE A COMPLETE REPORT FOR EACH REPORTABLE EVENT. THE REGISTRIES ARE LIKELY TO CONTAIN DEVICES FOR THE GORE® EXCLUDER® AAA ENDOPROSTHESIS, THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS, AND POTENTIALLY THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS. THE DATA GIVEN DID NOT DIFFERENTIATE WHICH OF THE SPECIFIC DEVICES CONTRIBUTED TO THE SERIOUS INJURIES. THEREFORE, WE REPORT ALL EVENTS FOR ALL DEVICES ON THE SAME REPORT. DEVICE BRAND NAME: GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS. PMA NUMBER: P200030. DEVICE BRAND NAME: GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS. PMA NUMBER: P020004. Manufacturer Narrative: THE PATIENT REFERENCED IN THIS EVENT FILE IS ENROLLED IN THE LONG-TERM EVAR ASSESSMENT AND FOLLOW-UP (LEAF) RETROSPECTIVE STUDY PROVIDED AN INITIAL UPDATE REGARDING THE PROCEDURAL OUTCOMES OF CURRENTLY ENROLLED PARTICIPANTS FOLLOWING IMPLANT OF GORE® EXCLUDER® AAA ENDOPROSTHESIS. ALL INFORMATION CONTAINED IN THIS EVENT FILE WAS RECEIVED FROM KAISER PERMANENTE EVAR REGISTRY SITES AND LINKED TO MEDICARE FEE FOR SERVICE DATA. THE INITIAL INFORMATION OBTAINED FROM KAISER PERMANENTE IS THE ONLY INFORMATION BEING PROVIDED. FURTHER INVESTIGATION IS NOT POSSIBLE AS IDENTIFIABLE DATA AS TO SITE(S), PHYSICIAN(S), PATIENT(S), DEVICE(S) AND/OR SPECIFIC DATE(S) ARE NOT BEING PROVIDED. A.2. AGE AT THE TIME OF EVENT: THE MEAN AGE FOR THE SUBJECTS IN THIS DATASET IS 76 YEARS. THEREFORE, 76 YEARS HAS BEEN USED AS THE ESTIMATED PATIENT AGE. A.3. GENDER: THE MAJORITY OF SUBJECTS IN THIS DATASET ARE MALE. THEREFORE, MALE HAS BEEN USED AS THE ESTIMATED PATIENT GENDER. A.4. PATIENT WEIGHT: THE AVERAGE WEIGHT OF THE PATIENTS WAS NOT GIVEN IN THE REPORTS. B.3. DATE OF EVENT: AS THE INDEX PROCEDURE DATE(S) AND DATE OF EVENT(S) REMAIN UNAVAILABLE, THE DATE OF EVENT HAS BEEN ESTIMATED AS THE DATE GORE AWARE OF THE INITIAL DATA EXPORT: 22-JAN-2025. D.4. DEVICE INFORMATION: THE DEVICE LOT/SERIAL NUMBER(S) REMAINS UNAVAILABLE. THEREFORE, DEVICE INFORMATION REMAINS UNKNOWN. D.10. CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: THIS INFORMATION REMAINS UNAVAILABLE TO GORE. H.4. DEVICE MANUFACTURE DATE: AS THE DEVICE LOT/SERIAL NUMBERS ARE UNAVAILABLE, THE DEVICE MANUFACTURE DATES REMAINS UNKNOWN. H.6. TYPE OF INVESTIGATION: CODE B22 - AS THE DEVICE LOT/SERIAL NUMBERS ARE UNAVAILABLE, NO DEVICE HISTORY RECORD REVIEW WAS ABLE TO BE COMPLETED. IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21406921&pc=DSQ | 2916596-2025-00862 | 2024-12-04 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. IT WAS UNKNOWN IF THE PUMP WOULD BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21407058&pc=DSQ | 2916596-2025-00906 | 2023-01-24 05:00 | Death | THORATEC CORPORATION | 2025-02-18 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS ALSO COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21488878&pc=PNJ | 9612164-2025-01063 | 2025-01-31 05:00 | Death | MEDTRONIC IRELAND | 2025-02-28 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ONE WEEK POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21489077&pc=OZP | 2032227-2025-130112 | 2025-01-15 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-28 05:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hypoglycemia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 DUE TO DROWNED IN THE SWIMMING POOL WHILE SWIMMING IN A RESORT POOL. THE CUSTOMER HAD LOW BLOOD GLUCOSE/HYPOGLYCEMIA WHILE SWIMMING IN THE POOL. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNK_RESERVOIR, UNOMEDICAL, MMT-1780KL. INFORMATION WAS OBTAINED FROM THE REPORTING PERSON AS PER DECEASED REPORTING TROUBLESHOOTING GUIDELINES. THE PUMP WAS REMOVED JUST BEFORE GOING TO SWIMMING POOL AND LAST WORN DATE FOR PUMP WAS (B)(6) 2025. UNK_RESERVOIR WAS REQUESTED AND RESPONSE WAS THE DEVICE WILL BE RETURNED. UNOMEDICAL WAS REQUESTED AND RESPONSE WAS THE DEVICE WILL BE RETURNED. MMT-1780KL WAS REQUESTED AND RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21489354&pc=OYC | 2032227-2025-130176 | 2022-09-11 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-28 05:00 | OYC | PARADIGM REAL-TIME REVEL INSULIN INFUSION PUMP MMT-723LNAS | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis | P150019 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2022, AND THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. IT WAS ALSO REPORTED THE CUSTOMER HAD HYPERGLYCEMIA. TROUBLESHOOTING WAS PARTIALLY PERFORMED, AND THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE MMT-723LNAS, MMT-397T, MMT-332AT, MMT-7703NA, AND MMT-7008A. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. NO PRODUCT RETURN IS REQUIRED FOR MMT-723LNAS. NO PRODUCT RETURN IS REQUIRED FOR MMT-397T. NO PRODUCT RETURN IS REQUIRED FOR MMT-332AT. NO PRODUCT RETURN IS REQUIRED FOR MMT-7703NA. NO PRODUCT RETURN IS REQUIRED FOR MMT-7008A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21489749&pc=OZP | 2032227-2025-130248 | 2025-01-30 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-02-28 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 DUE TO COMPLICATIONS OF DIABETES. CUSTOMER DIED IN THE TUB WHILE TAKING A BATH AT HOME. THE CUSTOMER'S BLOOD GLUCOSE WAS UNKNOWN. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884. INFORMATION WAS OBTAINED FROM THE REPORTING PERSON AS PER DECEASED REPORTING TROUBLESHOOTING GUIDELINES. THE INSULIN PUMP WAS NOT WORN AT THE TIME OF PASSING SINCE IT WAS REMOVED SAME DAY BEFORE CUSTOMER WAS GOING FOR BATH. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490123&pc=PNJ | 9612164-2025-01079 | 2025-02-08 05:00 | Death | MEDTRONIC IRELAND | 2025-02-28 05:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED TWO DAYS AFTER THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490283&pc=NPT | 2025587-2025-01538 | 2025-01-27 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-28 05:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE INCIDENCE AND IMPACT OF CORONARY EVENTS AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). THE STUDY POPULATION CONSISTED OF 64,660 PATIENTS WHO UNDERWENT TAVR WITH EITHER A NON-MEDTRONIC BALLOON-EXPANDABLE VALVE (BEV = 41,266) OR A MEDTRONIC SELF-EXPANDING VALVE (SEV = 23,394). IN THE SEV GROUP, THE AUTHORS OBSERVED DEATHS BEFORE DISCHARGE AFTER TAVR AND OBSERVED ADDITIONAL DEATHS FOLLOWING A FIRST CORONARY EVENT AFTER TAVR DISCHARGE (DEFINED AS REHOSPITALIZATION FOR STABLE ANGINA OR ACUTE CORONARY SYNDROME). NONE OF THE DEATHS WERE RECOUNTED IN DETAIL. OTHER ADVERSE EVENTS THAT OCCURRED IN THE SEV GROUP: MORE THAN ONE VALVE IMPLANTED; AORTIC REGURGITATION (MODERATE TO SEVERE); CORONARY OCCLUSION; ACUTE CORONARY SYNDROME; CORONARY EVENT NECESSITATING CORONARY ANGIOGRAPHY, PERCUTANEOUS CORONARY INTERVENTION, OR CORONARY ARTERY BYPASS GRAFTING; VALVE EMBOLIZATION; ANNULUS RUPTURE; AORTIC DISSECTION; TAMPONADE; NEED FOR PERMANENT PACEMAKER IMPLANT; URGENT SURGICAL AORTIC VALVE REPLACEMENT; CARDIAC SURGERY; STROKE; NEED FOR TRANSFUSION; ENDOCARDITIS; ACUTE KIDNEY INJURY/KIDNEY FAILURE; AND REHOSPITALIZATION FOR HEART FAILURE.Event Description: LITERATURE WAS REVIEWED REGARDING THE INCIDENCE AND IMPACT OF CORONARY EVENTS AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). THE STUDY POPULATION CONSISTED OF 64,660 PATIENTS WHO UNDERWENT TAVR WITH EITHER A NON-MEDTRONIC BALLOON-EXPANDABLE VALVE (BEV = 41,266) OR A MEDTRONIC SELF-EXPANDING VALVE (SEV = 23,394). IN THE SEV GROUP, THE AUTHORS OBSERVED DEATHS BEFORE DISCHARGE AFTER TAVR AND OBSERVED ADDITIONAL DEATHS FOLLOWING A FIRST CORONARY EVENT AFTER TAVR DISCHARGE (DEFINED AS REHOSPITALIZATION FOR STABLE ANGINA OR ACUTE CORONARY SYNDROME). NONE OF THE DEATHS WERE RECOUNTED IN DETAIL. OTHER ADVERSE EVENTS THAT OCCURRED IN THE SEV GROUP: MORE THAN ONE VALVE IMPLANTED; AORTIC REGURGITATION (MODERATE TO SEVERE); CORONARY OCCLUSION; ACUTE CORONARY SYNDROME; CORONARY EVENT NECESSITATING CORONARY ANGIOGRAPHY, PERCUTANEOUS CORONARY INTERVENTION, OR CORONARY ARTERY BYPASS GRAFTING; VALVE EMBOLIZATION; ANNULUS RUPTURE; AORTIC DISSECTION; TAMPONADE; NEED FOR PERMANENT PACEMAKER IMPLANT; URGENT SURGICAL AORTIC VALVE REPLACEMENT; CARDIAC SURGERY; STROKE; NEED FOR TRANSFUSION; ENDOCARDITIS; ACUTE KIDNEY INJURY/KIDNEY FAILURE; AND REHOSPITALIZATION FOR HEART FAILURE. ADDITIONAL INFORMATION RECEIVED FROM THE CORRESPONDING AUTHOR STATED THAT MEDTRONIC PRODUCT DID NOT CAUSE OR CONTRIBUTE TO ANY OF THE ADVERSE EVENTS (DEATHS AND NON-DEATH OUTCOMES) OBSERVED DURING THE STUDY. HOWEVER, NO EVIDENCE WAS PROVIDED TO CORROBORATE THIS STATEMENT. Manufacturer Narrative: CITATION: ZENDJEBIL S, AKODAD M, IUNG B, ET AL. CORONARY EVENTS AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT: INSIGHTS FROM THE FRANCE TAVI REGISTRY. JACC CARDIOVASC INTERV. 2025;18(2):229-243. DOI: 10.1016/J.JCIN.2024.09.005 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. MEDTRONIC TRANSCATHETER VALVE BRANDS USED IN THE STUDY: COREVALVE (PRODUCT CODE NPT, PMA#: P130021), EVOLUT R (PRODUCT CODE NPT, PMA#: P130021), AND EVOLUT PRO (PRODUCT CODE NPT, PMA#: P130021). EARLIEST APPROVED PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: SECTION B5 - SECOND PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490697&pc=JDI | 0002249697-2025-00249 | 2025-01-30 05:00 | Death | STRYKER ORTHOPAEDICS-MAHWAH | 2025-02-28 05:00 | JDI | EXETER V40(TM). FEM. STEM 44MM | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest; Embolism/Embolus | K011623 |  | 1 | Event Description: REVISION THA WAS PERFORMED ON (B)(6) 2025 TO REPLACE THE IMPLANT DEVICES (COMPETITOR PRODUCTS) WITH NEW ONES (STRYKER'S PRODUCTS) BECAUSE OF THE NECK FRACTURE OF THE COMPETITOR STEM. THE CUP WAS REPLACED SUCCESSFULLY, AND THE SURGEON BEGAN TO REPLACE THE STEM AS A NEXT STEP. HOWEVER, DURING STEM REPLACEMENT, THE PATIENT'S ELECTROCARDIOGRAM SUDDENLY SHOWED ABNORMAL WAVEFORMS AND HER CONDITION WORSENED. ALTHOUGH IT WAS NOT CONFIRMED PROPERLY, THE ANESTHESIOLOGIST WHO PARTICIPATED IN THE OPERATION STATED, "IT'S A PULMONARY EMBOLISM" DURING A LIFESAVING PROCEDURE OF THE SUDDEN CHANGE IN THE PATIENT'S CONDITION, SO THERE IS A POSSIBILITY THAT THE PATIENT WENT INTO CARDIAC ARREST DUE TO A PULMONARY EMBOLISM. AS THE PATIENT WENT INTO CARDIAC ARREST, THE SURGERY DISCONTINUED, AND THE PATIENT MANAGED TO REGAIN CONSCIOUSNESS ON THE SAME DAY THROUGH 30 MINUTES OF CARDIAC MASSAGE AND ECMO. THE PATIENT'S FINAL OUTCOME AFTER REGAINING CONSCIOUSNESS IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT HAD NO ABNORMALITIES SUCH AS PRIMARY DISEASE AT THE TIME OF SURGERY OTHER THAN STEM NECK FRACTURE.Event Description: REVISION THA WAS PERFORMED ON (B)(6) 2025 TO REPLACE THE IMPLANT DEVICES (COMPETITOR PRODUCTS) WITH NEW ONES (STRYKER'S PRODUCTS) BECAUSE OF THE NECK FRACTURE OF THE COMPETITOR STEM. THE CUP WAS REPLACED SUCCESSFULLY, AND THE SURGEON BEGAN TO REPLACE THE STEM AS A NEXT STEP. HOWEVER, DURING STEM REPLACEMENT, THE PATIENT'S ELECTROCARDIOGRAM SUDDENLY SHOWED ABNORMAL WAVEFORMS AND HER CONDITION WORSENED. ALTHOUGH IT WAS NOT CONFIRMED PROPERLY, THE ANESTHESIOLOGIST WHO PARTICIPATED IN THE OPERATION STATED "IT'S A PULMONARY EMBOLISM" DURING A LIFESAVING PROCEDURE OF THE SUDDEN CHANGE IN THE PATIENT'S CONDITION, SO THERE IS A POSSIBILITY THAT THE PATIENT WENT INTO CARDIAC ARREST DUE TO A PULMONARY EMBOLISM. AS THE PATIENT WENT INTO CARDIAC ARREST, THE SURGERY DISCONTINUED, AND THE PATIENT MANAGED TO REGAIN CONSCIOUSNESS ON THE SAME DAY THROUGH 30 MINUTES OF CARDIAC MASSAGE AND ECMO. THE PATIENT'S FINAL OUTCOME AFTER REGAINING CONSCIOUSNESS IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT HAD NO ABNORMALITIES SUCH AS PRIMARY DISEASE AT THE TIME OF SURGERY OTHER THAN STEM NECK FRACTURE. UPDATE: PATIENT DIED ON (B)(6) 2025 DUE TO A PULMONARY EMBOLUS. Manufacturer Narrative: CORRECTED DATA: DEVICE DETAILS (CAT#, GTIN, PRODUCT CODE, COMMON DEVICE NAME). AN EVENT REGARDING PATIENT FACTORS (PULMONARY EMBOLISM) INVOLVING A EXETER STEM WAS REPORTED. THE EVENT WAS NOT CONFIRMED. METHOD & RESULTS: PRODUCT EVALUATION AND RESULTS: MATERIAL ANALYSIS, VISUAL, FUNCTIONAL AND DIMENSIONAL INSPECTIONS COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED. CLINICIAN REVIEW: NO MEDICAL RECORDS WERE RECEIVED FOR REVIEW WITH CLINICAL CONSULTANT. PRODUCT HISTORY REVIEW: REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW: THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED. CONCLUSIONS: IT WAS REPORTED THAT DURING THE SURGERY, AFTER THE CUP PLACEMENT WAS COMPLETED AND THE STEM ON THE FEMORAL SIDE WAS BEING RE-REPLACED (THE SALES COMMENT WAS, "THE SURGICAL FIELD DID NOT LOOK CLEAR, SO THIS IS JUST SPECULATION, BUT WE THINK IT WAS WHEN WE WERE REMOVING THE SOFT TISSUE"), THE PATIENT'S ELECTROCARDIOGRAM SUDDENLY SHOWED ABNORMAL WAVEFORMS AND HER CONDITION SUDDENLY WORSENED. THE EXACT CAUSE COULD NOT BE DETERMINED AS INSUFFICIENT INFORMATION WAS PROVIDED. NO FURTHER INVESTIGATION FOR THIS EVENT IS POSSIBLE AT THIS TIME. IF DEVICES AND / OR ADDITIONAL INFORMATION BECOME AVAILABLE TO INDICATE FURTHER EVALUATION IS WARRANTED, THIS RECORD WILL BE REOPENED. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490825&pc=OZD | 1220648-2025-26415 | 2025-01-30 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT, CRITICAL, WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). OF NOTE, PRIOR TO IMPLANT OF THE IMPELLA, THE PATIENT WAS INITIALLY ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) FOR SUSPECTED BIVENTRICULAR DISFUNCTION, THEN REJECTED AND DECISION WAS MADE TO PLACE THE IMPELLA CP EVEN WITH THE FIRST SYMPTOMS OF MULTI-ORGAN FAILURE APPEARING. LATER, SAME DAY OF IMPELLA IMPLANT, AN OPTICAL SENSOR SYSTEM ERROR OCCURRED WITH THE AUTOMATED IMPELLA CONTROLLER (AIC). IT WAS FURTHER EXPLAINED THE LIGHT WAS VISIBLE; HOWEVER, THE CONSOLE WAS UNABLE TO RECOGNIZE THE OPTICAL SENSOR. THERE WAS AN AIC-TO-AIC EXCHANGE, THE ISSUE RESOLVED, AND THE PATIENT CONTINUED ON THE SAME PUMP WITH IMPELLA MCS. FOLLOWING, THREE DAYS LATER, THE PATIENT WAS REPORTED TO HAVE DEVELOPED HEMOLYSIS AND WAS IN SYSTEMIC MULTI-ORGAN FAILURE. TREATMENT FOR THE HEMOLYSIS IS UNKNOWN; HOWEVER, IT WAS SHARED THROUGH ADDITIONAL INFORMATION, THE HEMOLYSIS EVENT WAS UNRESOLVED. THE NEXT THREE DAYS WOULD PASS, AND IT WAS REPORTED THE PATIENT EXPIRED FROM MULTI-ORGAN FAILURE AFTER A TOTAL OF 6 DAYS ON IMPELLA MCS. AN AIC OPTICAL SENSOR FAILURE OCCURRED; HOWEVER, SUPPORT CONTINUED WITH THE SAME PUMP THAT CONTINUES TO OPERATE WITHOUT PLACEMENT MONITORING. A STANDARD AIC-TO-AIC EXCHANGE APPEARS TO HAVE OCCURRED; NO LOSS OR INTERRUPTION IN SUPPORT. THE PATIENT DEVELOPED HEMOLYSIS THAT WAS NOT RESOLVED AND THIS, HOWEVER, MAY HAVE ADDED TO AN ALREADY COMPLEX SITUATION LEADING TO A POSSIBLE ASSOCIATION TO THE OUTCOME. MEDICAL SAFETY REVIEW EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA CP PUMP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME ALTHOUGH THE LIKELY CAUSE OF DEATH APPEARS TO BE, AS REPORTED, MULTI-ORGAN FAILURE ALREADY PRESENT BEFORE THE IMPELLA CP IMPLANT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ THIS IS ONE OF TWO DEVICE MANUFACTURING REPORTS ASSOCIATED WITH THE EVENT. REFER TO THE FOLLOWING: MEDICAL DEVICE REPORT FOR AUTOMATED IMPELLA CONTROLLER SERIAL NUMBER (B)(6) WITH DATE OF EVENT 27-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 30-JAN-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). (THIS REPORT). Manufacturer Narrative: THE INVESTIGATION FOR THE HEMOLYSIS HAS BEEN COMPLETED. CLINICAL DETAILS STATED THAT PLACEMENT SIGNAL NOT RELIABLE ALARMS TRIGGERED DURING THE CASE UNTIL CONSOLE CHANGED. INSUFFICIENT DATA LOG PROVIDED. NO POSITIONING ALARMS OBSERVED. NO SUCTIONS OCCURRED. UNCLEAR IF THE PUMP REPOSITIONED OR RESOLVED AT ANY POINT DURING IMPELLA SUPPORT. THE CAUSE OF THE HEMOLYSIS ISSUE CANNOT BE DETERMINED SINCE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490861&pc=KDI | 3003464075-2025-00046 | 2025-01-31 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-02-28 05:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | K050525 |  | 1 | Event Description: A REPORT WAS RECEIVED ON 03 FEB 2025 FROM THE NURSE OF A MALE PATIENT WITH A MEDICAL HISTORY INCLUDING CHRONIC HYPOTENSION AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED 05-07 FEB 2025 FROM THE HOME THERAPY MANAGER (HTM) WHO STATED THE PATIENT WAS FOUND BY THE CAREGIVER SURROUNDED BY AN UNSPECIFIED AMOUNT OF BLOOD (SOURCE OF BLOOD LOSS NOT SPECIFIED). THE CAREGIVER PERFORMED CARDIOPULMONARY RESUSCITATION (CPR) AND CALLED EMERGENCY MEDICAL SERVICES (EMS). UPON EMS ARRIVAL, RESUSCITATION EFFORTS WERE CONTINUED BUT UNSUCCESSFUL, AND THE PATIENT EXPIRED, TIME NOT PROVIDED. Manufacturer Narrative: THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490904&pc=FPA | 21490904 | 2024-12-15 05:00 | Death | BD SWITZERLAND SARL | 2025-02-28 05:00 | FPA | ALARIS, SMARTSITE |  | Cardiac Arrest; Pallor; Respiratory Arrest |  |  | 1 | Event Description: AN INFANT ADMITTED TO NICU AND UMBILICAL ARTERIAL (UAC) AND UMBILICAL VENOUS CATHETERS (UVC) WERE TO BE PLACED. PRIOR TO PLACEMENT, THE UAC AND UVC LINES WERE SPIKED AND HANGING IN THE IV PUMPS. THE PHYSICIAN INSERTED THE UAC AND UVC WITH ASSISTANCE FROM THE RN. FOLLOWING THE DRAWING OF 2 SETS OF LABS AND A X-RAY, THE LINES WERE CONNECTED TO UAC AND UVC. SHORTLY AFTER, THE RN NOTICED THE BEDSIDE MONITOR WENT BLANK AND THE RN BEGAN TO TROUBLESHOOT. THE RN NOTICED THE INFANT WAS VERY PALE WITH SHALLOW CHEST RISE. IMMEDIATELY, THE RN CALLED FOR ASSISTANCE AND INFANT ARRESTED. THE PHYSICIAN NOTICED AIR IN THE UAC TUBING. PROMPTLY, THE RN STOPPED THE LINE AND REMOVED THE AIR. THE PHYSICIAN WAS ABLE TO VISUALIZE AIR ON THE X-RAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21490962&pc=OZD | 1220648-2025-26416 | 2025-02-03 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA FOR MECHANICAL CIRCULATORY SUPPORT AND POST IMPLANT, SAME DAY, ACCESS SITE BLEEDING OCCURRED AND WAS MANAGED (STEPS TO MANAGE THE BLEED ARE UNKNOWN). HEPARIN WAS PAUSED DUE TO HIGH PARTIAL THROMBOPLASTIN TIME (PTT) LESS THAN 120. THE FOLLOWING DAY THE PATIENT WAS NOTED TO BE UNSTABLE AND STILL BLEEDING FROM THE GROIN SITE. IT WAS NOTED HEPARIN WAS REDUCED, PTT 71. DISCUSSION AROUND ESCALATION TO EXTRACORPOREAL MEMBRANE OXYGENATION WAS NOTED. HOWEVER, TWO DAYS LATER IT WAS INDICATED THAT THE PATIENT EXPIRED ON SUPPORT. NO FURTHER DETAILS WERE MADE KNOWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491097&pc=LWP | MW5167043 | 2024-10-19 04:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | LWP | CARDIAC PACEMAKER |  | Abdominal Pain; Chest Pain; Nausea; Ventricular Fibrillation; Vomiting; Loss of consciousness |  |  | 1 | Event Description: (B)(6) WAS IMPLANTED WITH ABBOTT CARDIAC PACEMAKER PM2272, SERIAL NUMBER (B)(6) ON (B)(6) 2021. THE PACEMAKER WAS CHECKED AT THE PHYSICIAN OFFICE EVERY 6 MONTHS. THE NEXT CHECK WAS DUE AT THE END OF (B)(6) 2024. ON (B)(6) 2024 (B)(6) EXPERIENCED SEVERE UPPER CHEST PAIN/ABDOMINAL PAIN WHILE ON VACATION. HE WAS NAUSEATED AND THREW UP HIS BREAKFAST AND SAID HE WANTED TO GO TO THE HOSPITAL. HE THEN BENT OVER AND WHEN I TRIED TO RAISE HIS HEAD, HE WAS UNRESPONSIVE. THE EMTS WERE CALLED, AND HE WAS TRANSPORTED TO (B)(6) HOSPITAL VIA AMBULANCE. I WAS AT THE EMERGENCY DEPARTMENT ROOM AS THE CODE PROGRESSED. I HEARD THE NURSE QUESTION THE PHYSICIAN AS TO HOW THE EKG COULD BE CORRECT AS THE PATIENT HAD A PACEMAKER. THE PHYSICIAN STATED THAT THE PACEMAKER HAD FAILED. (B)(6) WAS PRONOUNCED DEAD, AND THE CAUSE OF DEATH WAS. VENTRICULAR FIBRILLATION. I CALLED ABBOTT ON (B)(6) 2024 TO REPORT THIS ISSUE AND TALKED WITH (B)(4) AT (B)(4). I CALLED AGAIN ON (B)(6) 2024 AND REPORTED TO HER THE SYMPTOMS AND THAT I HAD HEARD A SOFT DOUBLE BEEP 2 WEEKS PRIOR TO THIS EVENT BUT JUST ASSUMED THAT IT WAS HIS SONICARE TOOTHBRUSH THAT HE SOMETIMES DID NOT PROPERLY PUT ON THE CHARGER. I HAD ALSO HEARD THIS SOFT BEEP WHILE AT HILTON HEAD BUT ALSO ASSUMED IT WAS THE TOOTHBRUSH BATTERY. LATER WHEN I TALKED WITH THE ART TEACHER, SHE MENTIONED THAT SHE THOUGHT SHE HAD ALSO HEARD A FAINT DOUBLE BEEP BUT COULD NOT LOCATE THE ORIGIN AS IT DID NOT REPEAT. (B)(6) HAD VISITED THE CARDIOLOGISTS DR. (B)(6) ON (B)(6) 2024, AND NO PROBLEMS DETECTED. (B)(6) STATED THAT SHE WOULD CALL BACK AFTER THE MIDDLE OF (B)(6) BUT NOW MY CALLS TO HER WILL NOT GO THROUGH. SHE TOLD ME THAT SHE DID NOT KNOW OF ANY "DOUBLE BEEP'' WARNING EMITTED FROM THIS DEVICE. REF REPORTS: MW5167044, MW5167045.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491213&pc=LWS | MW5167044 | 2024-10-19 04:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | LWS | CARDIAC PACEMAKER | Low Audible Alarm; Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Chest Pain; Nausea; Ventricular Fibrillation; Vomiting; Loss of consciousness |  |  | 1 | Event Description: (B)(6) WAS IMPLANTED WITH ABBOTT CARDIAC PACEMAKER PM2272, SERIAL NUMBER (B)(6) ON (B)(6) 2021. THE PACEMAKER WAS CHECKED AT THE PHYSICIAN OFFICE EVERY 6 MONTHS. THE NEXT CHECK WAS DUE AT THE END OF (B)(6) 2024. ON (B)(6) 2024 (B)(6) EXPERIENCED SEVERE UPPER CHEST PAIN/ABDOMINAL PAIN WHILE ON VACATION. HE WAS NAUSEATED AND THREW UP HIS BREAKFAST AND SAID HE WANTED TO GO TO THE HOSPITAL. HE THEN BENT OVER AND WHEN I TRIED TO RAISE HIS HEAD, HE WAS UNRESPONSIVE. THE EMTS WERE CALLED, AND HE WAS TRANSPORTED TO (B)(6) HOSPITAL VIA AMBULANCE. I WAS AT THE EMERGENCY DEPARTMENT ROOM AS THE CODE PROGRESSED. I HEARD THE NURSE QUESTION THE PHYSICIAN AS TO HOW THE EKG COULD BE CORRECT AS THE PATIENT HAD A PACEMAKER. THE PHYSICIAN STATED THAT THE PACEMAKER HAD FAILED. (B)(6) WAS PRONOUNCED DEAD, AND THE CAUSE OF DEATH WAS VENTRICULAR FIBRILLATION. I CALLED ABBOTT ON (B)(6) 2024 TO REPORT THIS ISSUE AND TALKED WITH (B)(6) AT (B)(6). I CALLED AGAIN ON (B)(6) 2024 AND REPORTED TO HER THE SYMPTOMS AND THAT I HAD HEARD A SOFT DOUBLE BEEP 2 WEEKS PRIOR TO THIS EVENT, BUT JUST ASSUMED THAT IT WAS HIS SONICARE TOOTHBRUSH THAT HE SOMETIMES DID NOT PROPERLY PUT ON THE CHARGER. I HAD ALSO HEARD THIS SOFT BEEP WHILE AT HILTON HEAD BUT ALSO ASSUMED IT WAS THE TOOTHBRUSH BATTERY. LATER WHEN I TALKED WITH THE ART TEACHER, SHE MENTIONED THAT SHE THOUGHT SHE HAD ALSO HEARD A FAINT DOUBLE BEEP BUT COULD NOT LOCATE THE ORIGIN AS IT DID NOT REPEAT. (B)(6) HAD VISITED THE CARDIOLOGISTS DR. (B)(6) ON (B)(6) 2024, AND NO PROBLEMS DETECTED. MS. (B)(6) STATED THAT SHE WOULD CALL BACK AFTER THE MIDDLE OF (B)(6) BUT NOW MY CALLS TO HER WILL NOT GO THROUGH. SHE TOLD ME THAT SHE DID NOT KNOW OF ANY "DOUBLE BEEP'' WARNING EMITTED FROM THIS DEVICE. REF REPORTS: MW5167043, MW5167045. Event Description: ADDITIONAL INFORMATION RECEIVED ON 3/3/2025 FOR REPORT MW5167044 TO UPDATE PROCODE TO LWS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491241&pc=NVN | MW5167045 | 2024-10-19 04:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | NVN | CARDIAC PACEMAKER |  | Abdominal Pain; Chest Pain; Ventricular Fibrillation; Vomiting; Dizziness; Loss of consciousness |  |  | 1 | Event Description: (B)(6) WAS IMPLANTED WITH ABBOTT CARDIAC PACEMAKER PM2272, SERIAL NUMBER (B)(6) ON (B)(6) 2021. THE PACEMAKER WAS CHECKED AT THE PHYSICIAN OFFICE EVERY 6 MONTHS. THE NEXT CHECK WAS DUE AT THE END OF (B)(6) 2024. ON (B)(6) 2024 (B)(6) EXPERIENCED SEVERE UPPER CHEST PAIN/ABDOMINAL PAIN WHILE ON VACATION. HE WAS NAUSEATED AND THREW UP HIS BREAKFAST AND SAID HE WANTED TO GO TO THE HOSPITAL. HE THEN BENT OVER AND WHEN I TRIED TO RAISE HIS HEAD, HE WAS UNRESPONSIVE. THE EMTS WERE CALLED, AND HE WAS TRANSPORTED TO (B)(6) HOSPITAL VIA AMBULANCE. I WAS AT THE EMERGENCY DEPARTMENT ROOM AS THE CODE PROGRESSED. I HEARD THE NURSE QUESTION THE PHYSICIAN AS TO HOW THE EKG COULD BE CORRECT AS THE PATIENT HAD A PACEMAKER. THE PHYSICIAN STATED THAT THE PACEMAKER HAD FAILED. (B)(6) WAS PRONOUNCED DEAD, AND THE CAUSE OF DEATH WAS VENTRICULAR FIBRILLATION. I CALLED ABBOTT ON (B)(6) 2024 TO REPORT THIS ISSUE AND TALKED WITH (B)(6) AT (B)(6). I CALLED AGAIN ON (B)(6) 2024 AND REPORTED TO HER THE SYMPTOMS AND THAT I HAD HEARD A SOFT DOUBLE BEEP 2 WEEKS PRIOR TO THIS EVENT, BUT JUST ASSUMED THAT IT WAS HIS SONICARE TOOTHBRUSH THAT HE SOMETIMES DID NOT PROPERLY PUT ON THE CHARGER. I HAD ALSO HEARD THIS SOFT BEEP WHILE AT HILTON HEAD BUT ALSO ASSUMED IT WAS THE TOOTHBRUSH BATTERY. LATER WHEN I TALKED WITH THE ART TEACHER, SHE MENTIONED THAT SHE THOUGHT SHE HAD ALSO HEARD A FAINT DOUBLE BEEP BUT COULD NOT LOCATE THE ORIGIN AS IT DID NOT REPEAT. (B)(6) HAD VISITED THE CARDIOLOGISTS DR. (B)(6) ON (B)(6) 2024, AND NO PROBLEMS DETECTED. MS. (B)(6) STATED THAT SHE WOULD CALL BACK AFTER THE MIDDLE OF (B)(6) BUT NOW MY CALLS TO HER WILL NOT GO THROUGH. SHE TOLD ME THAT SHE DID NOT KNOW OF ANY "DOUBLE BEEP'' WARNING EMITTED FROM THIS DEVICE. REF REPORTS: MW5167043, MW5167044.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491500&pc=QLG | MW5167049 |  | Death | UNK | 2025-02-28 05:00 | QLG | FREESTYLE LIBRE 2 |  | Insufficient Information |  |  | 1 | Event Description: PATIENT PASSED AWAY (B)(6) 2022 AND WAS A DIABETIC. HE WAS USING THE DEVICE ON THE PA(PHYSICIAN ASSISTANT) REQUEST FOR HIS DIABETES DIAGNOSIS. HCP (HEALTHCARE PROFESSIONAL) DID CONSENT TO FOLLOW UP (B)(6) NO FURTHER INFORMATION PROVIDED OR AVAILABLE REGARDING THIS REPORT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491533&pc=PXU | MW5167050 | 2025-02-23 05:00 | Death | MERIT MEDICAL SYSTEMS | 2025-02-28 05:00 | PXU | MERIT MEDICAL PERICARDIOCENTESIS KIT - 8.3 F (2.77 MM) HIGH FLOW PIGTAIL CATHETE |  | Pericardial Effusion |  |  | 1 | Event Description: IN THE CATH LAB WHILE PERFORMING THE PERICARDIAL FLUID DRAINAGE PATIENT INITIALLY BECAME HEMODYNAMICALLY UNSTABLE AND THEN CODE BLUE WAS CALLED. CODE BLUE WAS TERMINATED AS PER FAMILY REQUEST AND PATIENT PASSED. POST-PROCEDURE, GUIDEWIRE FROM PERICARDIOCENTESIS KIT WAS FOUND TO BE SHEARED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491654&pc=DSI | 1218950-2025-000129 | 2025-02-03 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-02-28 05:00 | DSI | INTELLIVUE MX40 802.11A/B/G | Incorrect Measurement | Insufficient Information | K113125 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE MONITORING THE PATIENT THE READINGS OF THE SP02 WERE INCONSISTENT. THE SOURCE NOTES INDICATED THAT THE PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. ADDITIONAL INFORMATION IS REQUESTED FOR FURTHER CLARIFICATION AND ASSESSMENT OF THE CUSTOMER¿S ALLEGED HARM. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE TO EVALUATE THE DEVICE IN QUESTION. THE FSE INDICATED DURING AN EVALUATION THAT THE SPO2 CABLES REVEALED THE VALUES OBTAINED VIA THE PATIENT SIMULATOR WERE NUMERICALLY IDENTICAL REGARDLESS OF WHICH CABLE WAS TESTED. THE INTERMITTENT SIGNAL LOSS STOPPED AFTER SWITCHING TO ANOTHER SPO2 CABLE. THEN, WHEN SWITCHING BACK TO THE ORIGINAL CABLE, THE SIGNAL LOSS DID NOT RECUR AND COULD NOT BE DUPLICATED. BY SWITCHING THE CABLES, IT APPEARS THE ORIGINAL CABLE WAS RESEATED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS NOT CONFIRMED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491702&pc=DQA | MW5167051 | 2024-12-15 05:00 | Death | COVIDIEN | 2025-02-28 05:00 | DQA | NELLCOR BESIDE SPO2 MONITOR |  | Respiratory Distress Syndrome of Newborns |  |  | 1 | Event Description: PATIENT WAS CO-SLEEPING WITH PARENT(S) AND WAS USING A NELLCOR BEDSIDE SPO2 MONITORING DEVICE. PATIENT PASSED AWAY WHILE SLEEPING OVERNIGHT ON THE DEVICE. FAMILY STATES THE DEVICE DID NOT ALARM AT THE TIME OF PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491863&pc=HHS | MW5167052 | 2011-06-01 04:00 | Death | BAYER HEALTHCARE, LLC. | 2025-02-28 05:00 | HHS | ESSURE |  | Failure of Implant; Unspecified Reproductive System or Breast Problem |  |  | 1 | Event Description: ON (B)(6) 2011 DR. (B)(6) IN (B)(6) AND SHE WORKED FOR (B)(6) HOSPITAL IN (B)(6). SHE AIN'T PLAYING ME WITH THE BIRTH CONTROL DEVICE CALLED ESSURE BIRTH CONTROL THEIR TWO RODS THEY'RE STUCK IN MY FALLOPIAN TUBES ON (B)(6) AT 2011 SHE KNEW THEY WOULD BE RECALLED IN 2008, BUT SHE STILL PUT THEM INSIDE ME IN 2017 I HAD A YOUNG DAUGHTER ON (B)(6) 2017 I HAD A DAUGHTER BORN WITHOUT REPRODUCTIVE PARTS. I WAS NOT SUPPOSED TO HAVE A BABY (B)(6) OF 2018 I GAVE BIRTH TO A STILLBORN BABY WHO WAS BORN WITHOUT REPRODUCTIVE ORGANS, AND SHE DIED 6 HOURS LATER AFTER GIVING ME HER LAST BREATH I AM CURRENTLY LIVING IN (B)(6) WITHOUT MY FALLOPIAN TUBES AND MY OVARIES. DR (B)(6) OF (B)(6) HOSPITAL IN (B)(6) REMOVED THE DETECTED AND CORRODED OVARIES AND MY FALLOPIAN TUBES (B)(6) OF 2019 AND EVER SINCE THEN I'VE BEEN HAVING PROBLEMS IN MY LOWER REGION OF MY PUBLIC AREA AND MY UTERUS AREA. I NEED HELP AND I NEED A LAWYER TO GO AFTER DR (B)(6). REF REPORT: MW5167053.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491865&pc=OZD | 1220648-2025-26418 | 2025-02-07 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING THE IMPELLA 5.5 PROCEDURE, THE INTRODUCER WAS LEAKING AT THE PEEL AWAY POINT. THE IMPELLA 5.5 WAS THEN ADVANCED INTO THE VENTRICLE AND SUPPORT WAS INITIATED. THE INTRODUCER WAS THEN REMOVED. IT WAS FURTHER REPORTED THAT CARE WAS WITHDREW AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA 5.5 AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491873&pc=NPT | 2025587-2025-01542 | 2024-10-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-28 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Pulmonary Edema; Renal Failure; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE PATIENT ACCUMULATED WATER IN THE LUNGS, HEART FAILURE, AND RENAL FAILURE. SUBSEQUENTLY, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, IT WAS UNKNOWN IF THE VALVE AND VALVE IMPLANT PROCEDURE CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21491874&pc=HHS | MW5167053 | 2011-06-01 04:00 | Death | BAYER HEALTHCARE, LLC. | 2025-02-28 05:00 | HHS | ESSURE |  | Unspecified Reproductive System or Breast Problem; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ON (B)(6) 2011 DR. (B)(6) IN (B)(6) AND SHE WORKED FOR (B)(6) HOSPITAL IN (B)(6). SHE AIN'T PLAYING ME WITH THE BIRTH CONTROL DEVICE CALLED ESSURE BIRTH CONTROL THEIR TWO RODS THEY'RE STUCK IN MY FALLOPIAN TUBES ON (B)(6) AT 2011 SHE KNEW THEY WOULD BE RECALLED IN 2008, BUT SHE STILL PUT THEM INSIDE ME IN 2017 I HAD A YOUNG DAUGHTER ON (B)(6) AND 2017 I HAD A DAUGHTER BORN WITHOUT REPRODUCTIVE PARTS. I WAS NOT SUPPOSED TO HAVE A BABY (B)(6) OF 2018 I GAVE BIRTH TO A STILLBORN BABY WHO WAS BORN WITHOUT REPRODUCTIVE ORGANS, AND SHE DIED 6 HOURS LATER AFTER GIVING ME HER LAST BREATH I AM CURRENTLY LIVING IN (B)(6) WITHOUT MY FALLOPIAN TUBES AND MY OVARIES. DR (B)(6) OF (B)(6) HOSPITAL IN (B)(6) REMOVED THE DETECTED AND CORRODED OVARIES AND MY FALLOPIAN TUBES (B)(6) OF 2019 AND EVER SINCE THEN I'VE BEEN HAVING PROBLEMS IN MY LOWER REGION OF MY PUBLIC AREA AND MY UTERUS AREA. I NEED HELP AND I NEED A LAWYER TO GO AFTER DR (B)(6). REFERENCE REPORT: MW5167052.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492132&pc=OZD | 1220648-2025-26423 | 2025-02-02 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT EXPERIENCED CONTINUOUS OOZING AT THE ACCESS SITE. THE DEVICE CONTINUED TO MIGRATE AWAY FROM THE PATIENT¿S BODY SO THE DOCTOR MADE THE DECISION TO SUTURE DOWN THE SUTURE PAD USING ALL FOUR EYELETS. THE PATIENT CONTINUED ON SUPPORT. ADDITIONAL INFORMATION NOTED THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL INFORMATION WAS SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED MD MADE THE DECISION TO SUTURE DOWN THE SUTURE PAD USING ALL FOUR EYELETS TO PREVENT THE BLEEDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492171&pc=OZD | 1220648-2025-26426 | 2025-01-12 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT WHO EXPERIENCED VENTRICULAR TACHYCARDIA DURING IMPELLA CP SUPPORT. THE PATEINT REQUIRED ANTITACHYCARDIA PACING AND MEXITIL WAS RESUMED. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B3 DATE OF EVENT REVISED ON MANUFACTURER DEVICE REPORT 1220648-2025-26426 TO REFLECT CORRECT EVENT DATE. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26426 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G2 REPORT SOURCE; STUDY WAS MISSING FROM MANUFACTURER DEVICE REPORT 1220648-2025-26426.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492186&pc=NPT | 2025587-2025-01545 | 2024-07-15 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT SUDDENLY DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE VALVE AND THE VALVE IMPLANT PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. Manufacturer Narrative: UPDATED D.4 AND H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492577&pc=MHX | 9610105-2025-00001 | 2025-02-04 05:00 | Death | GE HEALTHCARE FINLAND OY | 2025-02-28 05:00 | MHX | CARESCAPE B450 | No Audible Alarm | Respiratory Failure | K213363 |  | 1 | Event Description: IT WAS REPORTED THAT NO AUDIBLE ALARMS WERE PROVIDED WHEN THE PATIENT EXPIRED. Manufacturer Narrative: LEGAL MANUFACTURER: (B)(4). A1-2, 4-6: NOT PROVIDED BY THE CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492838&pc=OZD | 1220648-2025-26431 | 2025-01-30 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND SUBSEQUENTLY EXPERIENCED INTESTINAL ISCHEMIA. THE REPORT NOTED IT IS POSSIBLE A CLOT DEVELOPED DURING PERCUTANEOUS CORONARY INTERVENTION OR DUE TO HEPARIN-INDUCED THROMBOCYTOPENIA (HIT). HOWEVER, THE CAUSE IS UNCLEAR. THUS, THE RELATIONSHIP TO THE IMPELLA DEVICE IS UNKNOWN. THE PATIENT WOULD EXPIRE AS A RESULT OF THE EVENT. ADDITIONAL COMMENTS NOTED, THE CAUSE OF DEATH AS INTESTINAL ISCHEMIA, DEATH RELATIONSHIP TO THE IMPELLA IS UNKNOWN (HOWEVER THE PUMP WAS FUNCTIONING NORMALLY) AND MEYLON WAS ADDED TO THE PURGING SOLUTION DUE TO THE SUSPECTED HIT. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. AN ATTEMPT TO CLARIFY THE DATE OF DEATH, DATE OF EVENT AND AWARENESS DATE HAS BEEN MADE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492917&pc=LWS | 2124215-2025-12495 | 2025-02-20 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-28 05:00 | LWS | EMBLEM S-ICD | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | No Clinical Signs, Symptoms or Conditions | P110042/S043 |  | 1 | Event Description: IT WAS REPORTED THAT UPON INTERROGATION, IT WAS NOTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) EXHIBITED A PREMATURE BATTERY DEPLETION ERROR MESSAGE. DEVICE REPLACEMENT WAS SCHEDULED. AT THIS TIME, THE S-ICD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT UPON INTERROGATION, IT WAS NOTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) EXHIBITED A PREMATURE BATTERY DEPLETION ERROR MESSAGE. DEVICE REPLACEMENT WAS SCHEDULED. AT THIS TIME, THE S-ICD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT, THIS S-ICD WAS EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT UPON INTERROGATION, IT WAS NOTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) EXHIBITED A PREMATURE BATTERY DEPLETION ERROR MESSAGE. DEVICE REPLACEMENT WAS SCHEDULED. AT THIS TIME, THE S-ICD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT, THIS S-ICD WAS EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT UPON INTERROGATION, IT WAS NOTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) EXHIBITED A PREMATURE BATTERY DEPLETION ERROR MESSAGE. DEVICE REPLACEMENT WAS SCHEDULED. AT THIS TIME, THE S-ICD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT, THIS S-ICD WAS EXPLANTED. UNFORTUNATELY, THE PATIENT PASSED AWAY SEVERAL DAYS AFTER THIS PROCEDURE. THIS S-ICD WAS RECEIVED FOR ANALYSIS. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS S-ICD WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF A COMPROMISED CAPACITOR. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITOR WAS COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF EMBLEM DEVICES THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE EMBLEM S-ICD ACCELERATED DEPLETION ADVISORY POPULATION. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS S-ICD WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF A COMPROMISED CAPACITOR. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITOR WAS COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF EMBLEM DEVICES THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE EMBLEM S-ICD ACCELERATED DEPLETION ADVISORY POPULATION. THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS B2 (OUTCOMES ATTRIB TO ADV EVENT), B5 (EVENT DESCRIPTION), D6B (EXPLANT DATE), H1 (TYPE OF REPORTABLE EVENT), AND H6 (IMPACT CODES).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492932&pc=DRA | 2124215-2025-12254 | 2024-11-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-28 05:00 | DRA | FARADRIVE STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A FARADRIVE SHEATH WAS INVOLVED IN A PULSE FIELD ABLATION PROCEDURE TO TREAT ATRIAL FIBRILLATION THAT RESULTED IN THE DEATH OF A PATIENT. Manufacturer Narrative: FURTHER INFORMATION REGARDING THE SPECIFIC PRODUCT INFORMATION, THE CAUSE OF DEATH, OR THE NATURE OF THE EVENT ARE UNABLE TO BE PROVIDED. WE WERE INFORMED OF THIS EVENT VIA MEDWATCH NOTIFICATION AND WERE UNABLE TO CORRELATE IT TO ANY EVENT BOSTON SCIENTIFIC WAS ALREADY AWARE OF. DUE TO THE ANONYMOUS NATURE OF THE INITIAL REPORTER WE ARE UNABLE TO CONDUCT ANY GOOD FAITH EFFORTS TO OBTAIN MORE INFORMATION. NO FURTHER DEVELOPMENTS ARE EXPECTED BUT ADDITIONAL MEDWATCH REPORTS WILL BE SUBMITTED SHOULD MORE INFORMATION BECOME AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21492938&pc=QZI | 2124215-2025-12253 | 2024-11-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-28 05:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A FARAWAVE PFA CATHETER WAS INVOLVED IN A PULSE FIELD ABLATION PROCEDURE TO TREAT ATRIAL FIBRILLATION THAT RESULTED IN THE DEATH OF A PATIENT. Manufacturer Narrative: FURTHER INFORMATION REGARDING THE SPECIFIC PRODUCT INFORMATION, THE CAUSE OF DEATH, OR THE NATURE OF THE EVENT ARE UNABLE TO BE PROVIDED. WE WERE INFORMED OF THIS EVENT VIA MEDWATCH NOTIFICATION AND WERE UNABLE TO CORRELATE IT TO ANY EVENT BOSTON SCIENTIFIC WAS ALREADY AWARE OF. DUE TO THE ANONYMOUS NATURE OF THE INITIAL REPORTER WE ARE UNABLE TO CONDUCT ANY GOOD FAITH EFFORTS TO OBTAIN MORE INFORMATION. NO FURTHER DEVELOPMENTS ARE EXPECTED BUT ADDITIONAL MEDWATCH REPORTS WILL BE SUBMITTED SHOULD MORE INFORMATION BECOME AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493045&pc=NKM | 2135147-2025-01024 | 2012-11-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "DIRECT ORAL ANTICOAGULANTS VERSUS VITAMIN K ANTAGONISTS AFTER MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR IN PATIENTS WITH ATRIAL FIBRILLATION: A SINGLE-CENTER OBSERVATIONAL STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE DIRECT ORAL ANTICOAGULANTS VERSUS VITAMIN K ANTAGONISTS AFTER MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR IN PATIENTS WITH ATRIAL FIBRILLATION. DEVICES MENTIONED INCLUDE MITRACLIP AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT DIRECT ORAL ANTICOAGULANTS MAY OFFER COMPARABLE EFFICACY AND SAFETY TO VKAS IN PREVENTING THROMBOEMBOLIC EVENTS FOLLOWING M-TEER IN PATIENTS WITH ATRIAL FIBRILLATION. FURTHER RANDOMIZED TRIALS ARE NEEDED TO CONFIRM THESE RESULTS AND ESTABLISH OPTIMAL ANTICOAGULATION STRATEGIES IN THIS PATIENT POPULATION. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS JAN-HENDRIK SCHIPPER AT UNIVERSITY HOSPITAL COLOGNE INSTITUTION WITH CORRESPONDING EMAIL JAN-HENDRIK.SCHIPPER@UK-KOELN.DE]. THE TIME FRAME OF THE STUDY WAS NOVEMBER 2012 TO APRIL 2019. A TOTAL OF 206 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT OF PATIENTS RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, PRIOR HEART SURGERY, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDE STROKE, BLEEDING, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, PRIOR HEART SURGERY, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED STROKE, BLEEDING, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493230&pc=NPT | 2025587-2025-01555 | 2025-01-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-28 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S CONDITION DETERIORATED AND THE PATIENT SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493254&pc=DSQ | 2916596-2025-01158 | 2025-02-07 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: A1-A4: PATIENT INFORMATION WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493457&pc=NPS | 2135147-2025-01036 | 2020-06-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "EFFICACY OF TRANSCATHETER EDGE-TO-EDGE REPAIR FOR CARDIAC IMPLANTABLE ELECTRONIC DEVICE-ASSOCIATED TRICUSPID REGURGITATION: INSIGHTS FROM THE TRI-SPA REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE EFFECTIVENESS AND CLINICAL OUTCOMES OF TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) IN PATIENTS WITH CARDIAC IMPLANTABLE ELECTRONIC DEVICES (CIEDS). DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT IN A REAL-WORLD SETTING, T-TEER SEEMS TO BE AN EFFECTIVE THERAPEUTIC OPTION FOR SELECTED PATIENTS WITH MORE THAN MODERATE TR AND CIEDS WITH NO OR MILD LEAD-LEAFLET INTERACTION, OFFERING COMPARABLE CARDIOVASCULAR OUTCOMES AND CLINICAL IMPROVEMENT TO THOSE WITHOUT LEADS. HOWEVER, THE PRESENCE OF CIEDS MAY REPRESENT AN INDEPENDENT RISK FACTOR FOR TR RECURRENCE. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS XAVIER FREIXA AT HOSPITAL CLÍNIC DE BARCELONA INSTITUTION WITH CORRESPONDING EMAIL XAVIERFREIXA@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS JUNE 2020 TO MAY 2023. A TOTAL OF 310 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 247 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, MAJOR BLEEDING, MYOCARDIAL INFARCTION, ENDOCARDITIS, RECURRENT TR, ADDITIONAL TRICLIP PROCEDURE, TISSUE DAMAGE, SLDA, MIGRATION, DEATH. (B)(6), UNK TRICLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, MAJOR BLEEDING, MYOCARDIAL INFARCTION, ENDOCARDITIS, RECURRENT TR, ADDITIONAL TRICLIP PROCEDURE, TISSUE DAMAGE, SLDA, MIGRATION, DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLES ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "EFFICACY OF TRANSCATHETER EDGE-TO-EDGE REPAIR FOR CARDIAC IMPLANTABLE ELECTRONIC DEVICE-ASSOCIATED TRICUSPID REGURGITATION: INSIGHTS FROM THE TRI-SPA REGISTRY." INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, MAJOR BLEEDING, MYOCARDIAL INFARCTION, ENDOCARDITIS, RECURRENT TR, ADDITIONAL TRICLIP PROCEDURE, TISSUE DAMAGE, SLDA, MIGRATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493673&pc=NPT | 2025587-2025-01558 | 2024-09-19 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-02-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Stroke/CVA; Dysphagia/ Odynophagia; Paralysis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED FOR MYELODYSPLASTIC SYNDROME. DURING HOSPITALIZATION, THE PATIENT WAS DIAGNOSED WITH SUSPECTED PARKINSON'S DISEASE AND MULTIPLE SYSTEM ATROPHY. APPROXIMATELY 1 WEEK LATER, THE PATIENT SUFFERED A STROKE THAT RESULTED IN HEMIPLEGIA. THE PATIENT BECAME BEDRIDDEN AND EXPERIENCED DIFFICULTY WITH FOOD INTAKE. THE PATIENT PASSED AWAY DUE TO SEQUELAE OF CEREBRAL INFARCTION APPROXIMATELY 3 MONTHS AND 3 WEEKS LATER.Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED FOR MYELODYSPLASTIC SYNDROME. DURING HOSPITALIZATION, THE PATIENT WAS DIAGNOSED WITH SUSPECTED PARKINSON'S DISEASE AND MULTIPLE SYSTEM ATROPHY. APPROXIMATELY 1 WEEK LATER, THE PATIENT SUFFERED A STROKE THAT RESULTED IN HEMIPLEGIA. THE PATIENT BECAME BEDRIDDEN AND EXPERIENCED DIFFICULTY WITH FOOD INTAKE. THE PATIENT PASSED AWAY DUE TO SEQUELAE OF CEREBRAL INFARCTION APPROXIMATELY 3 MONTHS AND 3 WEEKS LATER. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE STROKE WAS CONFIRMED VIA COMPUTED TOMOGRAPHY OF THE HEAD. THE CAUSE OF THE STROKE IS UNKNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. D. SUSPECT MEDICAL DEVICE: EXPIRATION DATE, SERIAL #, AND FULL UDI# H4. DEVICE MFG DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493758&pc=DSQ | 2916596-2025-00964 | 2024-12-14 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS UNDER HOSPICE CARE AND PASSED AWAY AT HOME ON
(B)(6) 2024. AS OF (B)(6) 2025, THE CAUSE OR DETAILS LEADING UP TO THE DEATH, AS WELL AS WHETHER THE PATIENT OUTCOME HAD BEEN RELATED TO THE DEVICE OR IF THE DEVICE HAD OPERATED AS EXPECTED, WERE UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493782&pc=LWS | 2649622-2025-05847 | 2022-10-11 04:00 | Death | MPRI | 2025-02-28 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOUR DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: 429878 LEAD; IMPLANTED (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493832&pc=NKM | 2135147-2025-01037 | 2020-06-01 04:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "EFFICACY OF TRANSCATHETER EDGE-TO-EDGE REPAIR FOR CARDIAC IMPLANTABLE ELECTRONIC DEVICE-ASSOCIATED TRICUSPID REGURGITATION: INSIGHTS FROM THE TRI-SPA REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE EFFECTIVENESS AND CLINICAL OUTCOMES OF TRICUSPID TRANSCATHETER EDGE-TO-EDGE REPAIR (T-TEER) IN PATIENTS WITH CARDIAC IMPLANTABLE ELECTRONIC DEVICES (CIEDS). DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT IN A REAL-WORLD SETTING, T-TEER SEEMS TO BE AN EFFECTIVE THERAPEUTIC OPTION FOR SELECTED PATIENTS WITH MORE THAN MODERATE TR AND CIEDS WITH NO OR MILD LEAD-LEAFLET INTERACTION, OFFERING COMPARABLE CARDIOVASCULAR OUTCOMES AND CLINICAL IMPROVEMENT TO THOSE WITHOUT LEADS. HOWEVER, THE PRESENCE OF CIEDS MAY REPRESENT AN INDEPENDENT RISK FACTOR FOR TR RECURRENCE. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS XAVIER FREIXA AT HOSPITAL CLÍNIC DE BARCELONA INSTITUTION WITH CORRESPONDING EMAIL XAVIERFREIXA@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS JUNE 2020 TO MAY 2023. A TOTAL OF 310 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 247 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, MAJOR BLEEDING, MYOCARDIAL INFARCTION, ENDOCARDITIS, RECURRENT TR, ADDITIONAL TRICLIP PROCEDURE, TISSUE DAMAGE, SLDA, MIGRATION, DEATH (B)(6), UNK TRICLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, MAJOR BLEEDING, MYOCARDIAL INFARCTION, ENDOCARDITIS, RECURRENT TR, ADDITIONAL TRICLIP PROCEDURE, TISSUE DAMAGE, SLDA, MIGRATION, DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLES ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "EFFICACY OF TRANSCATHETER EDGE-TO-EDGE REPAIR FOR CARDIAC IMPLANTABLE ELECTRONIC DEVICE-ASSOCIATED TRICUSPID REGURGITATION: INSIGHTS FROM THE TRI-SPA REGISTRY." INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, MAJOR BLEEDING, MYOCARDIAL INFARCTION, ENDOCARDITIS, RECURRENT TR, ADDITIONAL TRICLIP PROCEDURE, TISSUE DAMAGE, SLDA, MIGRATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE OF THE DEVICE WAS ASSOCIATED WITH THE MITRACLIP DEVICE BEING USED ON THE TRICUSPID VALVE TO TREAT TRICUSPID REGURGITATION (TR). A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493873&pc=NKM | 2135147-2025-01048 | 2016-01-11 05:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRENDS AND OUTCOMES OF READMISSIONS FOLLOWING POST-PROCEDURAL STROKE IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE NATIONAL READMISSION DATABASE (2016- 2020)" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE IMPACT OF POST-PROCEDURAL STROKE ON OUTCOMES AND READMISSIONS IN PATIENTS WHO UNDERWENT THE TEER PROCEDURE. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT PATIENTS IN THE ACUTE CVA GROUP HAD HIGHER RATES OF READMISSIONS, MEAN LENGTH OF STAY IN THE HOSPITAL, AND HIGHER HEALTHCARE BURDEN. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NIRAV PATEL AND CHADI ALRAIES AT DETROIT MEDICAL CENTER INSTITUTION WITH CORRESPONDING EMAIL ALRAIES@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS JANUARY 2016 TO DECEMBER 2020. A TOTAL OF 28,997 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 16,719 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, CORONARY ARTERY DISEASE. (B)(6), UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED STROKE, ACUTE KIDNEY INJURY, HEART FAILURE, MYOCARDIAL INFARCTION, BLEEDING, TAMPONADE, CARDIOGENIC SHOCK, CARDIAC ARREST, DEATH.Event Description: THE ARTICLE "TRENDS AND OUTCOMES OF READMISSIONS FOLLOWING POST-PROCEDURAL STROKE IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE NATIONAL READMISSION DATABASE (2016- 2020)" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE IMPACT OF POST-PROCEDURAL STROKE ON OUTCOMES AND READMISSIONS IN PATIENTS WHO UNDERWENT THE TEER PROCEDURE. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT PATIENTS IN THE ACUTE CVA GROUP HAD HIGHER RATES OF READMISSIONS, MEAN LENGTH OF STAY IN THE HOSPITAL, AND HIGHER HEALTHCARE BURDEN. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NIRAV PATEL AND CHADI ALRAIES AT DETROIT MEDICAL CENTER INSTITUTION WITH CORRESPONDING EMAIL ALRAIES@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS (B)(6) 2016 TO (B)(6) 2020. A TOTAL OF 28,997 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 16,719 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, CORONARY ARTERY DISEASE. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED STROKE, ACUTE KIDNEY INJURY, HEART FAILURE, MYOCARDIAL INFARCTION, BLEEDING, TAMPONADE, CARDIOGENIC SHOCK, CARDIAC ARREST, DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "TRENDS AND OUTCOMES OF READMISSIONS FOLLOWING POST-PROCEDURAL STROKE IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE NATIONAL READMISSION DATABASE (2016- 2020)". INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, CORONARY ARTERY DISEASE. COMPLICATIONS REPORTED INCLUDED STROKE, ACUTE KIDNEY INJURY, HEART FAILURE, MYOCARDIAL INFARCTION, BLEEDING, TAMPONADE, CARDIOGENIC SHOCK, CARDIAC ARREST, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. CODES REMOVED: HEALTH EFFECT-CLINICAL CODE: 4582. B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "TRENDS AND OUTCOMES OF READMISSIONS FOLLOWING POST-PROCEDURAL STROKE IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE NATIONAL READMISSION DATABASE (2016- 2020)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493928&pc=LLI | 1417592-2025-00113 | 2025-02-01 05:00 | Death | MEDLINE INDUSTRIES LP | 2025-02-28 05:00 | LLI | BED, BASIC, FULL ELECTRIC, 450 LB | Electrical /Electronic Property Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: RESIDENT APPEARED TO HAVE ROLLED OVER ON HISSIDE IN SUCH A MANNER THAT HIS CHEST TRIGGERED THE REMOTE CONTROL FUNCTIONS CAUSING BOTH THE HEAD AND FOOTSECTIONS OF THE BED TO EXTEND TO THEIR MAXIMUM VERTICAL POSITIONS. Manufacturer Narrative: IT WAS REPORTED "RESIDENT WAS OBSERVED LYING PARTIALLY ON BED, PARTIALLY OFF THE BED, WITH UPPER PART OF BODY PRESSED BETWEEN HIS HOSPITAL BED, BED REMOTE AND HALO DEVICE. THE RESIDENT HAD EXPIRED. RESIDENT WAS RELEASED TO MEDICAL EXAMINER'S OFFICE FOR FOLLOW-UP. ON SCENE INVESTIGATORS NOTED THAT THE HEAD AND FOOT OF THE ELECTRIC HOSPITAL BED WERE IN AT THEIR MAX EXTENSIONS. INVESTIGATORS NOTED THAT IT APPEARED THAT EVEN AT MAX EXTENSION THE MOTOR REMAINED ENGAGED AND CONTINUED TO TURN TO THE POINT THE PLASTIC CASING HAD POPPED APART". THE BED WAS BEING USED WITH A DOUBLE-SIDED HALO SAFETY RING. IT WAS REPORTED THAT THE USER KEPT THE BED PENDANT BESIDE HIM FOR EASE OF USE. IT HAS BEEN DETERMINED THAT THE REPORTED EVENT CAUSED OR CONTRIBUTED TO SERIOUS INJURY, THEREFORE, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21493929&pc=NIK | 9614453-2025-00839 | 2022-10-11 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-02-28 05:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOUR DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND RIGHT VENTRICULAR (RV) LEAD REPLACEMENT PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: 429878 LEAD, IMPLANTED (B)(6) 2020. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21494003&pc=OZD | 1220648-2025-26435 | 2025-02-05 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT EXPERIENCED HEMATOMA WITH A HEMOGLOBIN DECREASE. IMPELLA WAS REMOVED AND THE PUNCTURE SITE WAS CLOSED WITH A PROGLIDE AND AN ANGIOSEAL AND A PRESSURE BANDAGE WAS APPLIED. BLOOD UNITS WERE ORDERED. PATIENT WAS STABLE. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY PATIENT MOVEMENT SINCE THE HEMATOMA OCCURRED AFTER THE PATIENT WAS TRANSPORTED FROM THE CATHETERIZATION LABORATORY TO THE INTENSIVE CARE UNIT. B.2 ADDED DEATH AND DATE OF DEATH AS THEY WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435. B.7 ADDED INFORMATION AS IT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435. E.1 REVISED ADDRESS LINE 1 AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. F.6 AND F.8 DATES WERE REPORTED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435 AND SHOULD NOT HAVE BEEN. G.1 ADDED REPORTING CONTACT ADDRESS LINE 1, US CITY, STATE, COUNTRY AND US ZIP CODE AS THEY WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435. H.6 HEALTH EFFECT IMPACT CLINICAL CODE 1888 WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26435 AND A NEW CODE WAS ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21494236&pc=DSQ | 2916596-2025-01230 | 2025-02-13 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Respiratory Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PULMONARY INFECTION DURING THEIR POST-OPERATIVE INTENSIVE CARE UNIT (ICU) STAY. DESPITE REINTUBATION, NON-INVASIVE VENTILATION, AND ANTIBIOTIC TREATMENT, THE PATIENT WAS NOT RECOVERING. THE PATIENT WAS UNDER CLOSE MONITORING IN THE ICU. THE SOURCE OF THE INFECTION WAS BACTEREMIA, AND THE PATIENT EXHIBITED RESPIRATORY INSUFFICIENCY. THE PATIENT EXPIRED ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21494944&pc=NKM | 2135147-2025-01055 | 2021-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-02-28 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "SAFETY AND EFFICACY IN MITRAL REGURGITATION MANAGEMENT WITH THE MITRACLIP® G4 SYSTEM: INSIGHTS FROM A SINGLE-CENTER STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE PROCEDURAL, ECHOCARDIOGRAPHIC, FUNCTIONAL, AND QUALITY OF LIFE (QOL) OUTCOMES IN PATIENTS WHO UNDERWENT TEER WITH THE MITRACLIP® G4 SYSTEM, ALONG WITH POSSIBLE PREDICTORS OF NEW YORK HEART ASSOCIATION (NYHA) CLASS I AT 30 DAYS AND AT 1 YEAR. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THE MITRACLIP® G4 SYSTEM PROVIDES SIGNIFICANT IMPROVEMENTS IN MR SEVERITY, FUNCTIONAL CLASS, AND QOL. LOWER NT-PROBNP AND EUROSCORE II WERE STRONG PREDICTORS OF ACHIEVING OPTIMAL FUNCTIONAL STATUS (NYHA CLASS I). [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS GEORGIOS PAPADOPOULOS AT INTERBALKAN MEDICAL CENTER INSTITUTION WITH CORRESPONDING EMAIL GEORGIOS.E.PAPADOPOULOS@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS JANUARY 2021 TO DECEMBER 2023. A TOTAL OF 83 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 83 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, ATRIAL FIBRILLATION (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, RECURRENT MR, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, RECURRENT MR, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: B2: DATE OF DEATH ESTIMATED B3: DATE OF EVENT ESTIMATED. IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21495345&pc=OZD | 1220648-2025-26437 | 2025-02-08 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED AN IMPELLA CP PUMP VIA THE NON-ABIOMED PROVIDED SHEATH. THE SHEATH WAS THE SENTRANT 14 FRENCH AND THEN THE 18 FRENCH SENTRANT WHEN THE 14 FRENCH FAILED TO ALLOW THE IMPELLA CP TO DELIVER THROUGH. THE LIMB BECAME ISCHEMIC. THE DISCUSSION WAS MADE FOR A PERFUSION CATHETER BUT THE FAMILY MADE THE CHOICE NOT TO ESCALATED CARE AND INSTEAD WITHDREW CARE. THEREFORE, THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21495422&pc=DSQ | 2916596-2025-01153 | 2025-02-12 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD SEPSIS AND LIVER FAILURE, AND THE PATIENT PASSED AWAY. THE SEPSIS WAS FROM A LUNG INFECTION AND THE LIVER FAILURE WAS NOT PRE-EXISTING. AN AUTOPSY WOULD NOT BE PERFORMED, AND THE PUMP WOULD NOT BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. G, ARE CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING INFECTION, HEPATIC DYSFUNCTION, AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21495943&pc=OEZ | 3015910259-2025-00013 | 2024-08-27 04:00 | Death | MEDLINE INDUSTRIES, LP | 2025-02-28 05:00 | OEZ | LHK, CA, LEFT HEART 360I | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE FACILITY ON (B)(6) 2024 DURING A LEFT HEART CATHETERIZATION, AN AIR INJECTION WAS RECORDED WITH SUBSEQUENT DEATH OF THE PATIENT. THE CUSTOMER COULD NOT ATTRIBUTE THE AIR INJECTION TO ANY DEFINABLE MALFUNCTION OF THE MANIFOLD SYSTEM. Manufacturer Narrative: ACCORDING TO THE FACILITY ON (B)(6) 2024 DURING A LEFT HEART CATHETERIZATION, AN AIR INJECTION WAS RECORDED WITH SUBSEQUENT DEATH OF THE PATIENT. THE CUSTOMER COULD NOT ATTRIBUTE THE AIR INJECTION TO ANY DEFINABLE MALFUNCTION OF THE MANIFOLD SYSTEM. THERE IS NO FURTHER INFORMATION AT THIS TIME. IT HAS BEEN DETERMINED THAT THE REPORTED EVENT CAUSED OR CONTRIBUTED TO SERIOUS INJURY REQUIRING MEDICAL INTERVENTION, THEREFORE, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496171&pc=DSQ | 2916596-2025-01295 | 2025-01-21 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Infusion or Flow Problem | Bacterial Infection; Fatigue; Hemorrhage/Bleeding; Muscle Weakness; Respiratory Failure; Lactate Dehydrogenase Increased | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT LOG FILES WERE SUBMITTED FOR REVIEW CONCERNING A PATIENT WITH AN INCREASE IN LACTATE DEHYDROGENASE (LDH) AND RECURRENT LOW FLOW ALARMS. THE LOG FILE CAPTURED LOW FLOW EVENTS ON 21JAN2025 AND 30JAN2025. THERE ARE ALSO SEVERAL LOW FLOW FLAGS WHICH DID NOT PERSIST LONG ENOUGH TO TRIGGER A LOW FLOW ALARM. THESE OCCURRED AT TIMES WHEN PULSATILITY INDEX WAS ELEVATED. THERE DID NOT APPEAR TO BE ANY TYPE OF MECHANICAL CIRCULATORY SYSTEM EQUIPMENT ISSUES CAUSING THESE EVENTS. IT WAS LATER REPORTED THAT THE CAUSE OF THE LOW FLOW ALARMS WAS DUE TO HYPOVOLEMIA DUE TO BLEEDING AND HYPERTENSION. THE PATIENT EXPERIENCED SYMPTOMS OF WEAKNESS AND FATIGUE AT THE TIME OF THE LOW FLOWS. THIS WAS TREATED WITH BLOOD TRANSFUSIONS AND ADJUSTED ANTIHYPERTENSIVES. THE LOW FLOWS WERE SAID TO HAVE RESOLVED DUE TO THE TREATMENT. PATIENT WAS ALSO GIVEN HEMODIALYSIS DUE TO A MISSED SESSION. PATIENT LDH REMAINED ELEVATED AND IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PER THE VAD TEAM, THE PATIENT HAD BEEN DETERIORATING RECENTLY, HAD MISSED MULTIPLE DIALYSIS SESSIONS WHICH ULTIMATELY LED TO RESPIRATORY FAILURE. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. IT WAS ALSO SAID THAT THE PATIENT HAD MULTIPLE CO-MORBIDITIES, INCLUDING END-STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) FOR SEVERAL YEARS, BREAST CANCER IN WHICH THEY WERE RECEIVING CHEMO TREATMENT, DRIVELINE INFECTION ON INTRAVENOUS ANTIBIOTICS, AND PULMONARY/CARDIAC SARCOIDOSIS ON IMMUNOSUPPRESSION. THE PATIENT HAD BEEN LOSING WEIGHT DUE TO NOT EATING BECAUSE OF ALL THE PREVIOUSLY LISTED CO-MORBIDITIES AND MEDICATIONS. Manufacturer Narrative: ADDITIONAL H6 CODES: E130501 - RENAL FAILURE, E1208 - WEIGHT LOSS, E2343 - HEMODYNAMIC INSTABILITY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS, WHICH THE ACCOUNT ATTRIBUTED TO BLEEDING AND SUBSEQUENT HYPOVOLEMIA. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2024 ¿ (B)(6) 2025. INTERMITTENT LOW FLOW FAULT FLAGS WERE CAPTURED FROM THE BEGINNING OF THE FILE THROUGH 22:13:09 ON (B)(6) 2025, WITH SEVERAL OF THESE FAULTS RESULTING IN TRANSIENT LOW FLOW HAZARD ALARMS. THE SUBMITTED CONTROLLER PERIODIC LOG FILE CAPTURED AN ADDITIONAL LOW FLOW ALARM ON (B)(6) 2025. OF NOTE, PULSATILITY INDEX (PI) VALUES APPEARED TO BE ELEVATED DURING THE LOW FLOW EVENTS. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED AT THE STORED PATIENT SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, RENAL FAILURE, DRIVELINE INFECTION, BLEEDING, HEMOLYSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. THIS SECTION ALSO NOTES THAT, IN GENERAL, THE MAGNITUDE OF THE PULSATILITY INDEX (PI) VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E., THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. FOR PATIENTS WITH LOW FLOW SOFTWARE SYSTEM CONTROLLERS, THE HEARTMATE 3 LVAS POCKET GUIDES TO ALARMS FOR PATIENTS, REV. A, AND CLINICIANS, REV. A EXPLAIN THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.0 LPM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496358&pc=DSQ | 2916596-2025-01127 | 2019-01-01 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE REVIEW ARTICLE, "CARDIAC REVERSE REMODELING MEDIATED BY HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE: COMPARISON TO OLDER GENERATION DEVICES" THAT PATIENTS WHO RECEIVED A HEARTWARE (HVAD) N=115, HEARTMATE II (HM2) N = 104 AND HEARTMATE 3 (HM3) N = 186 FROM JANUARY 2016 TO JANUARY 2019 WERE ELIGIBLE. PATIENT WITH ACUTE MYOCARDITIS, DEATH WITHIN 3 MONTHS OF IMPLANT, BI-VENTRICULAR ASSIST DEVICES, PATIENTS NOT IMPLANTED WITH THE STUDY DEVICES, AND PATIENTS WITHOUT A BASELINE OR FOLLOW-UP ECHOCARDIOGRAPHIC DATA WERE EXCLUDED. PATIENTS WERE CENSORED AT DEATH, TRANSPLANT, DEVICE EXPLANT OR END OF STUDY FOLLOW UP WHICH WAS 1 YEAR AFTER IMPLANT. BASELINE CLINICAL DATA, INCLUDING PATIENT DEMOGRAPHICS, MEDICAL COMORBIDITIES, LABORATORY RESULTS, ECHOCARDIOGRAMS, AND INVASIVE HEMODYNAMICS WERE OBTAINED BEFORE DEVICE IMPLANTATION. SERIAL ECHOCARDIOGRAMS WERE OBTAINED AT 1, 3, 6, AND 12 MONTHS AFTER DEVICE IMPLANTATION, AND REPEAT INVASIVE HEMODYNAMICS WERE OBTAINED BETWEEN 3 AND 6 MONTHS AFTER LVAD IMPLANTATION. GUIDELINE-DIRECTED MEDICAL THERAPY AT 1, 3, 6, AND 12 MONTHS WAS ALSO RECORDED. THE RATE OF ACUTE POST-IMPLANT RV FAILURE (DEFINED AS CVP>18 MM HG WITH CI <2.0 WHILE IN THE ABSENCE OF ELEVATED LEFTSIDED FILLING PRESSURE, OR PROLONGED NITRIC OXIDE OR INOTROPIC SUPPORT) WAS 40.5%. PATIENTS IMPLANTED WITH HEARTMATE 3 HAD SIGNIFICANTLY LOWER RATE OF ACUTE RV FAILURE (P = 0.03). HOWEVER, THERE WERE NO SIGNIFICANT DIFFERENCES IN THE RATE OF SEVERE RV FAILURE BETWEEN THE GROUPS AT 3, 6, AND 12 MONTHS FOLLOWING LVAD IMPLANTATION. THERE WERE NO DIFFERENCES IN THE SEVERITY OF TRICUSPID REGURGITATION (NONE, MILD, MODERATE, SEVERE) BETWEEN THE THREE DEVICES AT ANY TIME POINTS. DURING THE 12 MONTH POST-LVAD FOLLOW-UP PERIOD, 85 PATIENTS UNDERWENT TRANSPLANTATION, 74 PATIENTS DIED, 229 PATIENTS ARE ALIVE WITH ONGOING LVAD SUPPORT, AND 14 PATIENTS WERE LOST TO FOLLOW-UP. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 DEVICES AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ARTICLE CONCLUDED THAT THE MECHANICAL UNLOADING AND CIRCULATORY SUPPORT PROVIDED BY THE HEARTMATE 3 RESULTED IN COMPARABLE CARDIAC REVERSE REMODELING AS HEARTMATE II AND HEARTWARE. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A AND D - ALL KNOWN INFORMATION WAS PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. YIN, M. Y., MANETA, E., KYRIAKOPOULOS, C. P., MICHAELS, A. T., GENOVESE, L. D., INDARAM, M. B., . . . DRAKOS, S. G. (2024). CARDIAC REVERSE REMODELING MEDIATED BY HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE: COMPARISON TO OLDER GENERATION DEVICES. ASAIO JOURNAL, 70(12), 1060-1066. DOI: HTTP://DX.DOI.ORG/10.1097/MAT.0000000000002245.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496379&pc=NPT | 2015691-2025-01436 | 2024-12-27 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-02-28 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, A 26 MM SAPIEN 3 ULTRA RESILIA VALVE WAS DEPLOYED IN THE NATIVE AORTIC POSITION VIA TRANSFEMORAL APPROACH. ON POSTOPERATIVE DAY (POD) 1, CEREBRAL INFARCTION WAS OBSERVED. THE PATIENT'S ACTIVITIES OF DAILY LIVING WERE DECREASING AND THE PATIENT BECAME DIFFICULT TO EAT. THE PATIENT'S CONDITION WAS GRADUALLY WORSENING. ON POD 10, THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURAL EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE EXACT CAUSE OF THE CEREBRAL INFARCTION IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496522&pc=BTR | 3012307300-2025-02353 | 2023-11-01 04:00 | Death | SMITH¿S MEDICAL ASD, INC. | 2025-02-28 05:00 | BTR | PORTEX NASAL SOFT SEAL CUFFED TRACHEAL TUBE | Imprecision | Cardiac Arrest | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THAT A PERSON WAS FOUND FLOATING IN THE SEA AND RESCUED IN A STATE OF CARDIOPULMONARY ARREST. THE PARAMEDICS PERFORMED INTUBATION TO DELIVER OXYGEN. AFTER ARRIVING AT THE HOSPITAL, IT WAS FOUND THAT THE TUBE HAD ENTERED THE ESOPHAGUS INSTEAD OF THE TRACHEA AND THE PATIENT WAS PRONOUNCED DEAD AT THE HOSPITAL. THE REPORTER STATED, DURING THE INTUBATION, THE PARAMEDIC ONLY CHECKED THE PATIENT WITH A STETHOSCOPE AFTER INTUBATION BUT DID NOT CHECK WITH ANY OTHER INSTRUMENTS. IT WAS ADDITIONALLY REPORTED THAT THE DOCTOR STATED THAT THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND THE TUBE NOT BEING IN THE TRACHEA. MEDICAL REVIEW FOLLOWS: AN 84-YEAR-OLD MAN WHO WAS FLOATING IN THE SEA WAS RESCUED IN A STATE OF CARDIOPULMONARY ARREST, AND PARAMEDICS PERFORMED TRACHEAL INTUBATION TO DELIVER OXYGEN. LATER, IT WAS FOUND THAT IT HAD ENTERED THE ESOPHAGUS INSTEAD OF THE TRACHEA AT THE HOSPITAL WHERE HE WAS TRANSPORTED. THE MAN WAS PRONOUNCED DEAD AT THE HOSPITAL. AT THAT TIME, THE PARAMEDIC ONLY CHECKED THE MAN WITH A STETHOSCOPE AFTER INSERTING THE TUBE, BUT DID NOT CHECK WITH THE PRESCRIBED MULTIPLE INSTRUMENTS. THE DOCTOR STATED THAT THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE MAN'S DEATH AND THE TUBE NOT BEING IN THE TRACHEA. Manufacturer Narrative: THE EXACT DATE IS UNKNOWN, AN APPROXIMATE DATE WAS USED. IF MORE INFORMATION ON THE DATE BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SENT. NO PRODUCT OR PHOTOS WERE RETURNED THEREFORE NO DEVICE ANALYSIS COULD BE COMPLETED. A DEVICE HISTORY RECORD (DHR) REVIEW COULD NOT BE PERFORMED AS THE LOT NUMBER WAS UNKNOWN. IF THE PRODUCT IS RETURNED THE MANUFACTURER WILL RE-OPEN THE COMPLAINT FOR FURTHER DEVICE ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496619&pc=DRA | 2124215-2025-12529 | 2025-02-20 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-28 05:00 | DRA | FARADRIVE STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Air Embolism; Arrhythmia; Non specific EKG/ECG Changes; Intracranial Hemorrhage |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED A ST SEGMENT ELEVATION, BRADYCARDIA AND CEREBRAL AIR EMBOLISM. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE A FARADRIVE WAS SELECTED FOR USE. THE PHYSICIAN SECURED TRANSEPTAL ACCESS USING NON-BOSTON SCIENTIFIC DEVICES. THE FIRST ATTEMPTED FAILED, RESULTING IN A PUNCTURE OF THE PERICARD WITH NO FURTHER ISSUES. AFTER GETTING TRANSEPTAL DONE THE FARADRIVE SHEATH WAS INSERTED, DURING THIS TIME THE PATIENT MOVED A LOT AS SEDATION WAS NOT SUFFICIENTLY AT FIRST. UPON SHEATH ADVANCEMENT INTO THE LEFT ATRIUM, A BRADYCARDIA AND ST SEGMENT ELEVATION WAS NOTICED ON THE SURFACE ECG. A CORONARY ANGIOGRAPHY WAS PERFORMED TO RULE OUT ANY AIR INGRESS. NO FURTHER ISSUES WERE OBSERVED DURING THE PROCEDURE AND THE PULMONARY VEINS WERE SUCCESSFULLY TREATED. AT THE END OF THE PROCEDURE, NEUROLOGICAL ABNORMALITIES OCCURRED LEADING TO PATIENT INTUBATION AND A COMPUTED TOMOGRAPHY SCAN WHICH REVEALED A CEREBRAL AIR EMBOLISM THAT REQUIRED FURTHER TREATMENT. NO FURTHER INFORMATION WAS PROVIDED. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS. 28FEB2025 - PER GFE: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT RECEIVED AN MRI SHOWING SIGNS OF BRAIN EDEMA. AFTER IT, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT AND RECEIVED HIGH-CONCENTRATED OXYGEN THERAPY AND WAS INTUBATED. ON (B)(6) 2025 THE PATIENT MANIFESTED ANISOCORIA AND A SECOND CT SCAN WAS PERFORMED SHOWING A SWELLING ON THE RIGHT SIDE, THE PATIENT WAS TRANSFERRED TO THE NEUROSURGICAL OPERATION ROOM TO HAVE A HEMI-CRANIECTOMY. ON (B)(6) 2025 ANOTHER IMAGING WAS PERFORMED AFTER THE PATIENT CONDITION DETERIORATED WITH SIGNS OF MASSIVE BLEEDING ON THE RIGHT SIDE AND ENTRAPMENT OF THE LEFT SIDE RESULTING IN A COMPLETE CRANIECTOMY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496738&pc=DSQ | 2916596-2025-01122 | 2024-12-30 05:00 | Death | THORATEC CORPORATION | 2025-02-28 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Sepsis; Diaphoresis; Confusion/ Disorientation; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS IN THE HOSPITAL BEING TREATED FOR SEPSIS. INOTROPES WERE INITIATED. THE PATIENT PASSED AWAY. THEIR MAIN CAUSE OF DEATH WAS SEPTICEMIA.Event Description: IT WAS REPORTED THAT THE SOURCE OF THE INFECTION WAS THE DRIVELINE. THE PATIENT PRESENTED WITH ELEVATED PROCALCITONIN (PCT) AND C-REACTIVE PROTEIN (CRP) MARKERS, AN IRREGULAR HEART BEAT, CONFUSION, AND CLAMMINESS. THE PATIENT WAS IN ATRIAL FIBRILLATION (AFIB), WHICH DID NOT RESULT IN CLINICAL COMPROMISE. IT WAS UNKNOWN IF THE PATIENT HAD A HISTORY OF ARRHYTHMIA PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE ARRHYTHMIA WAS NOT THOUGHT TO HAVE BEEN DEVICE OR THERAPY RELATED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. A, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), SEPSIS, CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496870&pc=NIO | 2124215-2025-12081 | 2024-12-26 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-28 05:00 | NIO | EPIC? VASCULAR | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Retroperitoneal Hemorrhage | P110035 |  | 1 | Event Description: IT WAS REPORTED THAT A HEMORRHAGE OCCURRED, AND THE PATIENT DIED. THE 100% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED EXTERNAL ILIAC ARTERY. A 10 X 60 X 75 EPIC VASCULAR STENT WAS IMPLANTED. DURING THE FOLLOW-UP PERIOD, EXTRAVASCULAR BLEEDING WAS NOTED FROM THE END OF THE STENT. A SURGICAL PROCEDURE WAS PERFORMED TO RESOLVE THE BLEEDING, AND THE PATIENT WAS UNDER OBSERVATION. FOUR DAYS LATER, THE PATIENT'S CONDITION SUDDENLY CHANGED, AND THEY WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, BUT THE PATIENT DID NOT REGAIN CONSCIOUSNESS AND PASSED AWAY. A CT SCAN REVEALED BLEEDING IN THE PERIPHERY OF THE STENT. THE CAUSE OF DEATH WAS RETROPERITONEAL HEMATOMA.Event Description: IT WAS REPORTED THAT A HEMORRHAGE OCCURRED, AND THE PATIENT DIED. THE 100% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED EXTERNAL ILIAC ARTERY. A 10X60X75 EPIC VASCULAR STENT WAS IMPLANTED. DURING THE FOLLOW-UP PERIOD, EXTRAVASCULAR BLEEDING WAS NOTED FROM THE END OF THE STENT. A SURGICAL PROCEDURE WAS PERFORMED TO RESOLVE THE BLEEDING, AND THE PATIENT WAS UNDER OBSERVATION. FOUR DAYS LATER, THE PATIENT'S CONDITION SUDDENLY CHANGED, AND THEY WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, BUT THE PATIENT DID NOT REGAIN CONSCIOUSNESS AND PASSED AWAY. A CT SCAN REVEALED BLEEDING IN THE PERIPHERY OF THE STENT. THE CAUSE OF DEATH WAS RETROPERITONEAL HEMATOMA. IT WAS FURTHER REPORTED THAT THE PATIENT HAD GASTRIC CANCER. THE PATIENT VOMITED BLOOD DUE TO GASTRIC CANCER BETWEEN THE INDEX PROCEDURE AND THE CARDIAC ARREST. Manufacturer Narrative: B5. DESCRIBE EVENT OR PROBLEM: ADDED INFORMATION, BASED ON ADDITIONAL INFORMATION UPDATED B7: OTHER RELEVANT HISTORY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496885&pc=LWP | 2124215-2025-12582 | 2003-02-03 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-02-28 05:00 | LWP | INSIGNIA I | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N970003/S016 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS SUCCESSFULLY IMPLANTED WITHOUT COMPLICATIONS. SUBSEQUENTLY, THE PATIENT DIED ON THE SAME DAY AS IMPLANT. IT WAS NOTED THAT THE PATIENT WAS BEING TREATED FOR SINUS TACHYCARDIA. NO FURTHER INFORMATION COULD BE OBTAINED AS THE EVENT OCCURRED IN 2003. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21496969&pc=OZD | 1220648-2025-26439 | 2025-02-01 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Failure to Advance |  | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE IMPLANT OF IMPELLA CP DEVICE TO A PATIENT (UNKNOWN RACE AND ETHNICITY) IN CARDIOGENIC SHOCK (CGS) DUE TO AN ACUTE MYOCARDIAL INFARCTION WAS UNSUCCESSFUL. THE PATIENT WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION HAD SUCCESSFUL DEPLOYMENT OF PERCUTANEOUS CORONARY INTERVENTION TO THE LEFT ANTERIOR DESCENDING ARTERY. WHILE COMPLETING CASE FOR TRANSFER TO INTENSIVE CARE UNIT, THE PATIENT ARRESTED AND DECISION TO PLACE IMPELLA CP FOR CGS. THE INITIAL PUMP ATTEMPT TO PLACE WITH A 14FR X 13CM SHEATH WAS UNSUCCESSFUL. THE PHYSICIAN REMOVED AND DISCARDED THE IMPELLA CP PUMP AND THE 14FR X 13CM SHEATH WAS EXCHANGED FOR 14FR X 25CM SHEATH OVER STIFFER WIRE. THE REPLACEMENT IMPELLA CP WAS PLACED OVER ABIOMED 0.018¿ WIRE AND STARTED ON AUTO. SHOCK ALERT CALLED WITH POTENTIAL PLANS TO TRANSFER. HOWEVER, WITHIN TWO HOURS, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WHEN INTO CARDIAC ARREST AND EXPIRED IN THE PROCEDURAL AREA. THE FAMILY STOPPED CARE. BASED ON THE INFORMATION AT HAND, THE PATIENT WAS VERY SICK. THE PHYSICIAN FOR SOME REASON WHICH IS UNKNOWN (HOWEVER COULD HAVE BEEN DUE TO VESSEL CHARACTERISTICS ETC.; DETAILS ARE UNKNOWN) WAS UNABLE TO DELIVER THE IMPELLA CP DEVICE. THE PATIENT NEEDED SUPPORT, AND THERE WAS A CONSEQUENTIAL DELAY WITH PROVIDING IMPELLA MECHANICAL CIRCULATORY SUPPORT. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21497059&pc=OZD | 1220648-2025-26441 | 2025-02-02 05:00 | Death | ABIOMED, INC. | 2025-02-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED SEPSIS AND PASSED AWAY DURING IMPELLA CP SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-26441 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21497113&pc=NVN | 2017865-2025-15475 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-28 05:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-15473. RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-15474. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DECEASED. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21497132&pc=NVZ | 2017865-2025-15473 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-28 05:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-15474. RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-15475. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. Manufacturer Narrative: DEVICE RETURNED FOR ANALYSIS DUE TO PATIENT DEATH. DEVICE WAS PROGRAMMED TO HIGH FIELD SETTINGS WITH AUTOCAPTURE AND RATE RESPONSIVE FEATURE ENABLED. ANALYSIS PERFORMED, INCLUDING ELECTRICAL TESTS INDICATED THAT THE DEVICE WAS AT END-OF-LIFE LEVEL. LONGEVITY ASSESSMENT WAS PERFORMED; DEVICE MET THE PROJECTED LONGEVITY AND IS CONSIDERED NORMAL BATTERY DEPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21497143&pc=NVN | 2017865-2025-15474 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-02-28 05:00 | NVN | ISOFLEX S LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15473. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-15475. IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21497607&pc=NPT | 2015691-2025-01440 | 2024-08-16 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-02-28 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem |  |  |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, A 26MM SAPIEN 3 VALVE WAS TRANSFEMORALLY DEPLOYED IN THE AORTIC ANNULUS. ON POSTOPERATIVE DAY 15, AORTIC DISSECTION OCCURRED AND THE PATIENT EXPIRED AT DIFFERENT HOSPITAL. IT WAS REPORTED FROM THE PATIENT FAMILY AND NO INFORMATION INCLUDING INTERVENTION AND THE DETAILS OF DEATH WAS OBTAINED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE AORTIC DISSECTION IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21497636&pc=LPB | 2029046-2025-00614 | 2025-01-31 05:00 | Death | BIOSENSE WEBSTER INC | 2025-02-28 05:00 | LPB | THERMOCOOL® SMART TOUCH® SF BI-DIRECTIONAL NAVIGATION CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation; Ischemic Heart Disease | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA ABLATION WITH A THERMOCOOL® SMART TOUCH® SF BI-DIRECTIONAL NAVIGATION CATHETER AND THE PATIENT EXPERIENCED ST ELEVATION, VENTRICULAR FIBRILLATION THAT RESULTED IN DEATH. DURING A PVC (PREMATURE VENTRICULAR CONTRACTION) ABLATION, AN STSF FJ CATHETER WAS USED TO MAP AND LOCALIZE THE PVC IN LVOT (LEFT VENTRICULAR OUTFLOW TRACT). CARTO 3D MAPPING WAS USED; HOWEVER, ICE (INTRACARDIAC ECHO) WAS NOT USED DURING THIS PROCEDURE. THE EARLIEST ACTIVATION WAS LOCALIZED IN THE LVOT JUST BELOW THE AORTIC CUSPS. ABLATION WAS PERFORMED ON THE EARLIEST SITE. AFTER ABLATION, THE PATIENT HAD ST ELEVATION, BLOOD PRESSURE DROPPED SEVERELY, AND PATIENT EVENTUALLY WENT INTO VF (VENTRICULAR FIBRILLATION). ECHO WAS DONE TO CHECK FOR EFFUSION/TAMPONADE. NEITHER WAS OBSERVED. A MECHANICAL PUMP WAS USED FOR CHEST COMPRESSIONS. AFTER NUMEROUS ATTEMPTS TO STABILIZE THE RHYTHM, THE PHYSICIAN EVENTUALLY CALLED IT AND THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO COMPLICATIONS FROM ABLATION. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21498573&pc=DSQ | 3004582654-2025-00014 | 2025-02-22 05:00 | Death | BERLIN HEART GMBH | 2025-02-28 05:00 | DSQ | EXCOR® | Detachment of Device or Device Component | Air Embolism | P160035 |  | 1 | Event Description: ON (B)(6) 2025, THE SURGEON CONTACTED BERLIN HEART TO REPORT AN ISSUE WITH A SINGLE VENTRICLE PATIENT SUPPORTED WITH A EXCOR PEDIATRIC VAD WITH AN ATRIAL INFLOW CANNULA AND AN AORTIC OUTFLOW WHO HAD BEEN TAKEN TO THE OPERATING ROOM FOR A CARDIAC TRANSPLANT. IN THE OPERATING ROOM, THE PATIENT WAS TRANSITIONED FROM THE EXCOR VAD TO CARDIOPULMONARY BYPASS WITH THE BERLIN HEART CANNULAS. WHILE DISSECTING THE HEART, THE SURGEON NOTICED AIR IN THE VENOUS LINE OF THE BYPASS CIRCUIT THAT RESOLVED WITHOUT INTERVENTION AND WAS THEN NOTICED AGAIN. WHILE INVESTIGATING THE SOURCE OF THE AIR, THE SURGEON DISSECTED THE SITE OF THE ATRIAL CANNULA CONNECTION AND OBSERVED THAT THE HEAD OF THE ATRIAL CANNULA HAD SEPARATED FROM THE ATRIAL CANNULA SEWING RING WHICH REMAINED ATTACHED TO THE HEART ALLOWING AIR TO ENTER THE PATIENT. THE PATIENT ACUTELY DECOMPENSATED AND WAS FOUND TO HAVE A MASSIVE AIR EMBOLISM. EFFORTS TO RESUSCITATE THE PATIENT WERE NOT EFFECTIVE, THE PROCEDURE WAS ABORTED, AND THE PATIENT WAS DECLARED DEAD IN THE OPERATING ROOM. BERLIN HEART INC REQUESTED THAT THE SITE SEND ANY PICTURES AVAILABLE AND RETURN THE PRODUCT FOR INVESTIGATION AFTER THE AUTOPSY HAD BEEN PERFORMED. AT THE TIME OF THE EVENT THAT PATIENT WAS ON CARDIOPULMONARY BYPASS. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED IN RESPONSE TO THE FDA LETTER DATED 4/17/2025. THE INITIAL MDR REPORT 3004582654-2025-00014 HAS BEEN SUBMITTED ON 02/28/2025. Manufacturer Narrative: THE EXCOR ATRIAL CANNULA FOR CHILDREN; Ø 12/9 MM; L 32 CM; HEAD 23 MM (LOT NO. 00132882) WAS IN USE ON THE PATIENT FROM ON (B)(6) 2023 UNTIL THE TIME OF THE EVENT ON 2025-02-22 (501 DAYS). THIS CANNULA IS NO LONGER ON THE PATIENT. WE HAVE REVIEWED THE PRODUCTION RECORDS OF THE CANNULA LOT NO. 00132882. THIS PRODUCT WAS PRODUCED ACCORDING TO OUR SPECIFICATION. ACCORDING TO THE PRODUCTION RECORDS, A TOTAL OF FOUR CANNULAS WITH THIS LOT NO. WERE PRODUCED. OUT OF THE FOUR CANNULAS, THREE WERE CANNULAS DISTRIBUTED IN THE US WERE USED ON PATIENTS. THE FOURTH CANNULA WHICH WAS DISTRIBUTED TO CANADA WAS NOT USED. THERE ARE NO OTHER COMPLAINTS ASSOCIATED WITH THIS LOT NUMBER. A DETAILED INVESTIGATION REPORT WILL BE PROVIDED AS SOON AS IT IS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21498789&pc=NPT | 2015691-2025-01447 | 2025-02-02 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-02-28 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, FOUR (4) DAYS AFTER TAVR PROCEDURE OF 23MM SAPIEN 3 ULTRA RESILIA VALVE, AORTIC DISSECTION AND CEREBRAL INFARCTION WERE OBSERVED. THE PATIENT WAS UNDER OBSERVATION. ADDITIONAL INFORMATION WAS RECEIVED BY THE SALES REP 7 DAYS POST TAVR THAT THE PATIENT PASSED AWAY. THE PATIENT WAS (B)(6) YEARS OLD FEMALE. AS REGARDING THE CAUSE OF AORTIC DISSECTION, THE DEVICE PUSHER MAY HAVE TOUCHED. THERE WAS NO RESISTANCE FELT UPON DEVICE INSERTION/WITHDRAWAL. AS FOR CEREBRAL INFARCTION, IT WAS UNDER OBSERVATION. BOTH THE CAUSE OF DEATH AND THE DATE OF DEATH WERE UNKNOWN. NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE THV PROCEDURE AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE STROKE MAY BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.