## Jan'25
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121544&pc=BZD | 2518422-2025-001457 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121646&pc=BZD | 2518422-2025-001468 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121684&pc=BZD | 2518422-2025-001437 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Headache; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS 32 ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121789&pc=BZD | 2518422-2025-001487 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Inflammation; Dizziness; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121953&pc=BZD | 2518422-2025-001498 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121969&pc=BZD | 2518422-2025-001534 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD); Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, LIVER DISEASE TOXICITY AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122018&pc=BZD | 2518422-2025-001543 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Nausea; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA/VOMITING, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122120&pc=NVZ | 3004209178-2025-00530 |  | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-10 05:00 | NVZ | ADAPTA | Appropriate Term/Code Not Available | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PULSE GENERATOR (IPG) WAS EXPLANTED FOR AN UNKNOWN REASON. FOLLOW-UP WAS ATTEMPTED TO OBTAIN THE RELEVANT INFORMATION; HOWEVER, NO ADDITIONAL DETAILS ARE AVAILABLE AT THIS TIME. THIS EVENT IS BEING CONSERVATIVELY REPORTED IN AN ABUNDANCE OF CAUTION IN RESPONSE TO THE ADVISORY DESCRIBING THE POTENTIAL FOR DEVICE CIRCUIT ERROR TO OCCUR WHILE THE DEVICE IS PROCESSING AN ATRIAL-SENSED EPISODE.Event Description: IT WAS NOTED THAT THE IPG WAS NOT RELATED TO DEVICE CIRCUIT ERROR AND THAT THE PATIENT HAS PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. ANALYSIS OF THE DEVICE REVEALED NORMAL BATTERY DEPLETION. THE BATTERY INDICATOR SIGNIFYING THAT IT IS TIME FOR DEVICE REPLACEMENT OCCURRED AFTER EXPLANT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-58 LEAD; THERAPY DATE: (B)(6) 2008. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122145&pc=BZD | 2518422-2025-001492 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122261&pc=BZD | 2518422-2025-001593 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Dizziness; Respiratory Failure; Cancer; Heart Failure/Congestive Heart Failure; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER, DEATH RESPIRATORY FAILURE, HEART FAILURE PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122391&pc=BZD | 2518422-2025-001599 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING) CANCER, DEATH, LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122568&pc=BZD | 2518422-2025-001645 | 2023-08-24 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Unspecified Heart Problem; Unspecified Respiratory Problem | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG AND HEART PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN BOX H. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H: TYPE OF REPORTED COMPLAINT UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122615&pc=BZD | 2518422-2025-001611 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD); Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE) OTHER (UNSPECIFIED INFORMATION). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122628&pc=BZD | 2518422-2025-001587 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION CANCER; DEATH, BRAIN CANCER. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122699&pc=BZD | 2518422-2025-001636 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21122929&pc=BZD | 2518422-2025-001671 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE, CANCER, NOSE IRRITATION AND SKIN IRRITATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123394&pc=BZD | 2518422-2025-001710 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS AND/OR HEADACHE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123396&pc=LDD | 3003832357-2025-000025 | 2025-01-06 05:00 | Death | SCHILLER AG | 2025-01-10 05:00 | LDD | TEMPUS LS | Failure to Deliver Shock/Stimulation; Defibrillation/Stimulation Problem | Unspecified Heart Problem | K200849 |  | 1 | Event Description: IT HAS BEEN REPORTED TO PHILIPS THE TEMPUS LS DEFIBRILLATOR DISPLAYED, AFTER TURNING THE DEVICE ON AND APPLYING PADS TO THE PATIENT, "HARDWARE FAILURE".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123447&pc=ILS | 3007231105-2025-00002 | 2024-11-02 04:00 | Death | INVACARE GMBH | 2025-01-10 05:00 | ILS | OCEAN VIP | Device Tipped Over | Intracranial Hemorrhage |  |  | 1 | Event Description: IT WAS REPORTED A FALL OCCURRED FROM AN OCEAN VIP SHOWER CHAIR IN THE BEDROOM WHEN THE SLING WAS BEING REMOVED BY AN ASSISTANT. THE SHOWER CHAIR TIPPED FORWARD, AND THE PATIENT FELL TOWARDS THE CLOSET DOOR AND HIT HIS HEAD. AFTER THE INCIDENT, THE PATIENT DID NOT WANT TO GO TO THE HOSPITAL ACCORDING TO THE ASSISTANT. THE PATIENT HAS REGULAR CONTACT WITH THE HEALTH CARE PROVIDER AS HE RECEIVES DIALYSIS THREE TIMES A WEEK. AFTER 5 DAYS, THE PATIENT'S CONDITION DETERIORATED, AND LEVEL OF CONSCIOUSNESS DECREASED. WHEN THE PRESCRIBER FOLLOWED UP ON THE INCIDENT THE FOLLOWING WEEK, IT WAS DISCOVERED THAT THE PATIENT WAS HOSPITALIZED WITH A CEREBRAL HEMORRHAGE AND TREATMENT HAD BEEN TERMINATED. THE PATIENT DIED ON (B)(6) 2024. Manufacturer Narrative: INVACARE WAS MADE AWARE OF AN EVENT THAT OCCURRED IN SWEDEN INVOLVING AN OCEAN VIP MOBILE SHOWER CHAIR. INVACARE IS FILING THIS MEDWATCH IN AN ABUNDANCE OF CAUTION DUE TO THE OCEAN VIP SHOWER CHAIR IS SOLD IN THE U.S. THE SUBJECT DEVICE WAS MANUFACTURED BY INVACARE GERMANY, AQUATEC OPERATIONS. THE SHOWER CHAIR HAS BEEN DISPOSED OF AND NO PHOTOS ARE AVAILABLE. THE REPORTED ADVISED THE ANTI-TIPPERS AND OPTIONAL PELVIC AND CHEST BELTS WERE NOT INSTALLED ON THE SHOWER CHAIR. A MALFUNCTION OF THE DEVICE WAS NOT ALLEGED, AND THE EVENT MAY BE THE RESULT OF THE USE ERROR. IT WAS ALSO NOTED IT COULD NOT BE DETERMINED IF THE CEREBRAL HEMORRHAGE WAS DUE TO THE PATIENT¿S MEDICAL HISTORY OR THE RESULT OF THE FALL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123852&pc=FRN | 9610825-2024-01020 | 2024-12-10 05:00 | Death | B BRAUN MELSUNGEN AG | 2025-01-10 05:00 | FRN | PERFUSOR® | Insufficient Flow or Under Infusion | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE CUSTOMER: THERE WAS AN UNDERDOSE - PATIENT DEATH. THE CUSTOMER HAS STATED THEY DO NOT BELIEVE THE PUMP IS AT FAULT, THEY WANT THE DATA HISTORY OF THE PUMP. Manufacturer Narrative: THIS REPORT HAS BEEN IDENTIFIED AS B. BRAUN INTERNAL REPORT NUMBER (B)(4). NOTE: THIS REPORT IS BEING FILED FOR AN ITEM NUMBER THAT IS NOT SOLD IN THE UNITED STATES, HOWEVER SIMILAR ITEMS ARE SOLD IN THE UNITED STATES BY B. BRAUN MEDICAL, INC. UNDERDOSE GENERAL INFORMATION: COMPLAINT: (B)(4). INFORMATION TO THE SAMPLE: MODEL: PERFUSOR SPACE ARTICLE NUMBER: 8713030 SERIAL NUMBER/BATCH: (B)(6) SOFTWARE VERSION: 688N030005 HOURS OF OPERATION: 2579 FURTHER INFORMATION: N/A INVESTIGATION RESULTS: HISTORY INSPECTION: THE DEVICE HISTORY FILES WERE READ OUT AND ANALYZED. THE CUSTOMER SPECIFIED 2024-12-10 AS THE DATE OF THE INCIDENT. THE LAST THERAPY ON THIS DAY WAS ANALYZED. 2024-12-10 AT 10:45 A BD PLASTIPAK 50ML WAS SELECTED. THE SYRINGE WAS FILLED TO APPROX. 49ML. THE DRUG SHORT NAME MORPHMT WAS SELECTED. THE THERAPY WAS STARTED WITH A FLOW RATE OF 0 ,19ML/H AT 10:45. THE THERAPY WAS TERMINATED BY THE USER AFTER 3 HOURS AND 37 MINUTES. 0.73ML WERE INFUSED IN TOTAL. NO ABNORMALITIES CAN BE FOUND IN THE DEVICE HISTORY FILES. VISUAL INSPECTION: A VISUAL INSPECTION WAS PERFORMED. THE COVER CAPS ON THE SCREW PILLARS, AND THE PRODUCTION SEAL ON THE LOWER HOUSING WERE INTACT AND UNDAMAGED. NO VISIBLE DAMAGED PARTS ARE TO LOCATE. FUNCTIONAL INSPECTION: A FUNCTIONAL TEST WAS PERFORMED. THE DEVICE PASSES THE SELF-TEST. A BD 50 ML SYRINGE WAS INSERTED, AND THE PUMP IDENTIFIED THE SYRINGE, AND IT COULD BE SELECTED FROM THE MENU. IT WAS POSSIBLE TO PUT THE PUMP IN OPERATION. INDIVIDUAL INSPECTION: A DELIVERY ACCURACY MEASUREMENT ACCORDING TO IEC 60601-2-24 WAS ARRANGED. HERE A NOMINAL FLOW RATE OF 5 ML/H WAS CHOSEN. THE ASSESSED MEAN DEVIATION "A" OF THE SECOND OPERATING HOUR WAS MEASURED AND RESULTED IN A VALUE OF + 0.29%. ACCURACY OF SET DELIVERY RATE SHOULD BE ± 2 % ACCORDING TO IEC/EN 60601-2-24. DISASSEMBLING: DURING THE INVESTIGATION NO FAULTS COULD BE DETECTED, TO INVESTIGATE THE INSIDE OF THE DEVICE, ONLY THE UPPER HOUSING WAS REMOVED. NO DAMAGE OR SOILING COULD BE FOUND. JUDGMENT: THE COMPLAINT COULD NOT BE CONFIRMED. SUMMING UP ALL TESTS, THE PERFUSOR SPACE OPERATED WITHIN OUR SPECIFICATION. NO PRODUCT DEVIATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124036&pc=GDW | 2647580-2025-00161 | 2024-08-04 04:00 | Death | US SURGICAL PUERTO RICO | 2025-01-10 05:00 | GDW | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 28 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED FOUR (2.3%) 30-DAY MORTALITY RATE AND 8 (4.7%) 90-DAY MORTALITY FOR PATIENTS WITH NO ANASTOMOTIC STRICTURE. THE CAUSE OF DEATHS WAS NOT REPORTED. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN) RISK FACTORS FOR AND TREATMENT OF ANASTOMOTIC STRICTURES AFTER IVOR LEWIS ESOPHAGECTOMY ¦ SOPHIE L. F. DORAN, MARIA G. DIGBY, SOPHIE V. GREEN, CLIVE J. KELTY, ANAND P. TAMHANKAR, ¦ SURGICAL ENDOSCOPY (2024) 38:6771¿6777 ¦ HTTPS://DOI.ORG/10.1007/S00464-024-11150-W ¦ PUBLISHED ONLINE: 19 AUGUST 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124063&pc=GDW | 2647580-2025-00162 | 2024-08-04 04:00 | Death | US SURGICAL PUERTO RICO | 2025-01-10 05:00 | GDW | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 31 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED FOUR (2.3%) 30-DAY MORTALITY RATE AND 8 (4.7%) 90-DAY MORTALITY FOR PATIENTS WITH NO ANASTOMOTIC STRICTURE. THE CAUSE OF DEATHS WAS NOT REPORTED. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN) RISK FACTORS FOR AND TREATMENT OF ANASTOMOTIC STRICTURES AFTER IVOR LEWIS ESOPHAGECTOMY ¦ SOPHIE L. F. DORAN, MARIA G. DIGBY, SOPHIE V. GREEN, CLIVE J. KELTY, ANAND P. TAMHANKAR, ¦ SURGICAL ENDOSCOPY (2024) 38:6771¿6777 ¦ HTTPS://DOI.ORG/10.1007/S00464-024-11150-W ¦ PUBLISHED ONLINE: 19 AUGUST 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124096&pc=QZI | 2124215-2024-83627 | 2024-12-11 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-10 05:00 | QZI | FARAWAVE PULSED FIELD ABLATION CATHETER, 31MM | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Liver Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY A WEEK AFTER USE OF THE FARAWAVE PULSED FIELD ABLATION CATHETER, THE PATIENT EXPERIENCED LIVER FAILURE. THE DEVICE IS NOT EXPECTED TO BE RETURN AS IT WAS DISPOSED. NO FURTHER INFORMATION WAS PROVIDED.Event Description: IT WAS REPORTED THAT APPROXIMATELY A WEEK AFTER USE OF THE FARAWAVE PULSED FIELD ABLATION CATHETER, THE PATIENT EXPERIENCED LIVER FAILURE. THE DEVICE IS NOT EXPECTED TO BE RETURN AS IT WAS DISPOSED. NO FURTHER INFORMATION WAS PROVIDED. IT WAS FURTHER REPORTED THAT THE EVENT OCCURRED ABOUT A WEEK AFTER THE PERSISTENT ATRIAL FIBRILLATION (OFF-LABEL USE) PULSE FIELD ABLATION (PFA) PROCEDURE. THE EVENT WAS NOT SPECIFIC TO THE FARAWAVE CATHETER, BUT IT WAS RELATED TO THE PROCEDURE IN GENERAL. THE PATIENT PASSED AWAY AT THE END OF (B)(6) 2024 DUE TO LIVER FAILURE. THE PHYSICIAN SHARED SHE DID NOT THINK IT WAS DUE TO THE PFA PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO DESCRIBE EVENT OR PROBLEM, IMPACT CODES, AND TYPE OF REPORTABLE EVENT. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION; THEREFORE, A FAILURE ANALYSIS OF THE COMPLAINT DEVICE COULD NOT BE COMPLETED. THERE WERE NO ALLEGATIONS OF A QUALITY OR MANUFACTURING DEFICIENCY LEADING TO THE ADVERSE EVENT AS REPORTED TO BSC, AND THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS AT THIS TIME. THE EXACT CAUSE OF THE PATIENT DEATH WAS UNKNOWN. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124261&pc=NPV | 2025587-2025-00266 | 2024-12-18 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-10 05:00 | NPV | MELODY TRANSCATHETER PULMONARY VALVE | Patient Device Interaction Problem | Endocarditis; Septic Shock | P140017 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE EFFECTIVENESS OF EMERGENCY PERCUTANEOUS TREATMENTS IN 16 PATIENTS WITH SEVERE RIGHT VENTRICULAR OUTFLOW TRACT (RVOT) OBSTRUCTION DUE TO INFECTIVE ENDOCARDITIS (IE). ALL BUT ONE OF THESE PATIENTS HAD A PREVIOUSLY IMPLANTED MEDTRONIC HEART VALVE DEVICE (CONTEGRA CONDUIT, MELODY TRANSCATHETER PULMONARY VALVE, MELODY IN CONTEGRA, OR MELODY IN HANCOCK PROSTHESIS). AS DESCRIBED IN THE ARTICLE, EVERY PATIENT IN THE STUDY WAS DIAGNOSED WITH IE AFTER A MEDIAN DURATION OF 66.5 (28.3-87.3) MONTHS POST-DEVICE IMPLANT. ADVERSE EVENTS NOTED IN THE ARTICLE: MELODY IMPLANTED VALVE-IN-VALVE INSIDE CONTEGRA PRIOR TO INCLUSION IN THIS STUDY (3 CASES); MELODY IMPLANTED VALVE-IN-VALVE INSIDE HANCOCK PRIOR TO INCLUSION IN THIS STUDY (2 CASES); AND IE PRESENTING WITH RVOT OBSTRUCTION, VEGETATIONS ON VALVE, EMBOLI, HIGH PRESSURE GRADIENTS, REDUCED RV FUNCTION, OBSTRUCTIVE CARDIAC SHOCK, SEPTIC SHOCK, FEVER, AND BACTEREMIA. INTERVENTIONS PERFORMED TO ADDRESS IE INCLUDED: INOTROPIC SUPPORT, BALLOON DILATION, COVERED OR UNCOVERED STENT PLACEMENT, PERCUTANEOUS VALVE IMPLANT, SURGICAL VALVE REPLACEMENT, AND HEART TRANSPLANTATION. DURING FOLLOW-UP, THE AUTHORS RECOUNTED THAT FOUR CASES OF PERSISTENT BACTERIEMIA AND ONE CASE OF UNCONTROLLED SEPSIS OCCURRED BUT DID NOT IDENTIFY THE DEVICES IMPLANTED IN THESE PATIENTS. ADDITIONALLY, ONE OF THE DEATHS THAT OCCURRED DURING FOLLOW-UP INVOLVED A MELODY PATIENT WHO DIED DUE TO EARLY IE RELAPSE AND SEPTIC SHOCK (PATIENT 16 IN SUPPLEMENTARY TABLE 1). THE OTHER DEATH OCCURRED IN A PATIENT WITH A NON-MEDTRONIC DEVICE (PATIENT 7 IN SUPPLEMENTARY TABLE 1). Manufacturer Narrative: CITATION: CALLEGARI A, ALBERTINI M, REVERDITO G, BONNET D, MALEKZADEH-MILANI S. INTERVENTIONAL TREATMENT OF ACUTE RIGHT VENTRICULAR OUTFLOW TRACT INFECTIOUS ENDOCARDITIS: A BRIDGE TO SURGICAL OR PERCUTANEOUS PULMONARY VALVE REPLACEMENT. CATHETER CARDIOVASC INTERV. PUBLISHED ONLINE DECEMBER 18, 2024. DOI:10.1002/CCD.31348 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124284&pc=BZD | 2518422-2025-001607 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124346&pc=GDW | 2647580-2025-00160 | 2024-08-04 04:00 | Death | US SURGICAL PUERTO RICO | 2025-01-10 05:00 | GDW | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY AIMED TO ASSESS THE RATE OF AND RISK FACTORS FOR ANASTOMOTIC STRICTURE IN 207 PATIENTS WHO UNDERWENT IVOR LEWIS ESOPHAGECTOMY FOR ESOPHAGEAL CANCER BETWEEN 2013 AND 2021. A STANDARD CIRCULAR STAPLED END-TO-SIDE ESOPHAGO-GASTRIC ANASTOMOSIS WAS PERFORMED USING A 25 MILLIMETER (MM) CIRCULAR STAPLER. POST OPERATIVE COMPLICATIONS INCLUDED FOUR (2.3%) 30-DAY MORTALITY RATE AND 8 (4.7%) 90-DAY MORTALITY FOR PATIENTS WITH NO ANASTOMOTIC STRICTURE. THE CAUSE OF DEATHS WAS NOT REPORTED. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN); UNKNOWN EEA, UNKNOWN EEA (LOT#UNKNOWN). RISK FACTORS FOR AND TREATMENT OF ANASTOMOTIC STRICTURES AFTER IVOR LEWIS ESOPHAGECTOMY ¦ SOPHIE L. F. DORAN, MARIA G. DIGBY, SOPHIE V. GREEN, CLIVE J. KELTY, ANAND P. TAMHANKAR, ¦ SURGICAL ENDOSCOPY (2024) 38:6771¿6777 ¦ HTTPS://DOI.ORG/10.1007/S00464-024-11150-W ¦ PUBLISHED ONLINE: 19 AUGUST 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124469&pc=OZD | 1220648-2025-25478 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemolysis; Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, THE PATIENT BECAME HYPERTENSIVE WITH SUCTION ALARMS AND BLOOD PRODUCTS GIVEN. STERILE SLEEVE ADVANCED AND TOUHY BOOST VALVE TIGHTENED AND T LOCK SECURE. PLACEMENT VERIFIED WITH TRANS-ESOPHAGEAL ECHOCARDIOGRAPHY (TEE). AFTER BLOOD PRODUCTS WERE GIVEN, THE IMPELLA INCREASED BACK TO P9. ADDITIONALLY, URINE WAS NOTED TO BE PINK. PLACEMENT VERIFIED PER ECHOCARDIOGRAPHY. THE URINE WAS RED. NO PLASMA FREE HGB DRAWN. NO HEMOLYZED LABS. SUSPECTED HEMOLYSIS. PATIENT WAS MADE DO NOT RESUSCITATE, AND DUE TO FAILURE TO THRIVE DECISION WAS MADE BY FAMILY TO WITHDRAW CARE. THE PATIENT EXPIRED. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUES HAS BEEN COMPLETED. DEVICE HISTORY REVIEW CONFIRMED PUMP PASSED ALL POST STERILE INSPECTION CHECKS. DATA LOGS CONFIRMED THE FAILURE MODE AND CLINICAL NARRATIVE. LOGS SHOWED LOW FLOW OCCURRED DURING THE CASE THAT TRIGGERED AUTO SUCTION RESPONSE & MULTIPLE SUCTION ALARMS. CLINICAL INFORMATION STATES PUMP WAS IN PROPER POSITION DURING THE INCIDENTS. NO MOTOR CURRENT SPIKES OUTSIDE P LEVEL CHANGES WERE OBSERVED. CLINICAL INFORMATION ALSO STATES THAT AFTER GIVING BLOOD ABLE TO TURN PUMP BACK TO P-9. NO OTHER ISSUES WERE FOUND ON THE LOGS. THE DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON CLINICAL DESCRIPTION AND DATA REVIEW, THE ROOT CAUSE OF THE REPORTED ISSUE OF LOW FLOW WAS MOST LIKELY PATIENT CONDITION. THE ROOT CAUSE OF HEMOLYSIS WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124470&pc=DYB | 2015691-2025-00215 | 2024-12-17 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-10 05:00 | DYB | EDWARDS ESHEATH PLUS INTRODUCER SET | Material Puncture/Hole; Physical Resistance/Sticking | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Perforation of Vessels; Vascular Dissection; Asystole | K200258 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, REGARDING IMPLANTATION OF A 29MM SAPIEN 3 ULTRA RESILIA CLAVE IN THE AORTIC POSITION. A 29MM SAPIEN 3 ULTRA RESILIA VALVE WAS INSERTED INTO THE 16F ESHEATH. THE TEAM EXPERIENCED QUITE A BIT OF RESISTANCE ABOUT A 1/4 OF THE WAY THROUGH THE BLUE PORTION OF THE SHEATH. THE IMPLANTERS NOTICED THEY WERE MAKING SOME PROGRESS WITH FORCE THROUGH THE SHEATH SO THEY OPTED NOT TO FLUORO THE VALVE. WHEN RESISTANCE WAS NO LONGER FELT THE TEAM PANNED DOWN TO THE VALVE AND NOTICED THE DELIVERY SYSTEM AND VALVE WERE PROLAPSED THROUGH THE SHEATH AND HAD PERFORATED THE RIGHT COMMON ILIAC. THE VALVE DID NOT MAKE IT OUT OF THE TIP OF THE SHEATH. THE TEAM QUICKLY TOOK AN AORTOGRAM AND CONFIRMED THE PERFORATION BUT ALSO NOTICED MASSIVE DISSECTIONS RUNNING UP TO THE DISTAL AORTA. THEY INSERTED AN AORTIC BALLOON AND INFLATED IN THE DISTAL AORTA TO TAMPONADE, HOWEVER THE PATIENT'S PRESSURE QUICKLY DROPPED AND BECAME ASYSTOLIC. GIVEN THE EXTREME HIGH RISK TO SURGERY AND THE LACK OF PERCUTANEOUS OPTIONS, THE TEAM MADE A DECISION THAT THERE WASN'T A FEASIBLE OPTION TO SALVAGE. THE TEAM CALLED IT AND THE PATIENT PASSED AWAY. THE TEAM PULLED THE COMMANDER DELIVERY SYSTEM THROUGH THE SHEATH WHILE THE VALVE WAS PULLED OFF THE BALLOON AND STUCK IN SHEATH. THE ESHEATH WAS STILL IN PLACE WHEN THEY PICKED UP PATIENT FOR MORGUE. TEAM QUICKLY DISCARDED THE REMAINING DEVICES. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS: G3, G6, H2, H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PROVIDED IMAGING WAS REVIEWED. 3MENSIO IMAGERY WAS PROVIDED AND THE FOLLOWING WAS OBSERVED: TORTUOSITY AND CALCIFICATION PRESENT IN PATIENTS RIGHT ACCESS VESSEL. TORTUOSITY AND CALCIFICATION PRESENT IN PATIENTS RIGHT ILIAC VESSEL. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE REPORTED RESISTANCE WAS UNABLE TO BE CONFIRMED AS NO COMPLAINT UNIT WAS RETURNED FOR EVALUATION AND NO PROCEDURAL IMAGERY WAS PROVIDED. 3MENSIO SHOWS THE ACCESS VESSEL SIZE WAS APPROPRIATE FOR USE (MINIMUM LUMINAL DIAMETER GREATER OR EQUAL TO 5.5MM FOR USE WITH THE 14F ESHEATH PLUS) AND A STEEP INSERTION ANGLE WAS NOT USED. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (TORTUOSITY, CALCIFICATION) POTENTIALLY CONTRIBUTED TO THE EVENT AS THE 3MENSIO REVIEW SHOWED TORTUOSITY AND CALCIFICATION. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. THE REPORTED LINER PUNCTURE WAS UNABLE TO BE CONFIRMED AS NO COMPLAINT UNIT WAS RETURNED FOR EVALUATION AND NO PROCEDURAL IMAGERY WAS PROVIDED. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THE VALVE WERE PROLAPSED THROUGH THE SHEATH AND HAD PERFORATED THE RIGHT COMMON ILIAC. DURING DELIVERY SYSTEM ADVANCEMENT THROUGH A CHALLENGING PATHWAY, IT IS POSSIBLE THAT THE DEVICES MAY NOT BE COAXIALLY ALIGNED. THIS CAN LEAD TO THE CRIMPED VALVE TO CATCH ONTO THE SHEATH LINER AND LEAD TO DAMAGE. DEVICE MANIPULATION APPLIED TO OVERCOME THE RESISTANCE CAN FURTHER DAMAGE THE SHEATH THROUGH CONTINUED INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. VESSEL CALCIFICATION AND/OR TORTUOSITY IF PRESENT CAN EXASPERATE THE INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. SHARP CALCIFIED NODULES CAN DIRECTLY WEAKEN THE SHEATH SHAFT MAKING IT MORE SUSCEPTIBLE TO DAMAGE AS THE DELIVERY SYSTEM WITH CRIMPED VALVE IS ADVANCED THROUGH. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS UNDERWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124569&pc=OZD | 1220648-2025-25480 | 2024-12-21 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT HAD HEMOLYSIS, WHICH WAS CONFIRMED VIA LABS. HOWEVER, THE PATIENT PASSED AWAY THE SAME DAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. H5 WAS INADVERTENTLY OMITTED FROM MANUFACTURER DEVICE REPORT 1220648-2025-25480.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124598&pc=BZD | 2518422-2025-001751 | 2023-07-25 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124599&pc=BZD | 2518422-2025-001759 | 2022-02-08 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE THAT A USER OF THE PHILIPS CPAP DEVICE HAD PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124935&pc=DRA | 2135147-2025-00249 | 2024-12-18 05:00 | Death | ABBOTT MEDICAL | 2025-01-10 05:00 | DRA | MITRACLIP | Failure to Advance | Perforation; Pulmonary Edema | K190167 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024 A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MITRAL REGURGITATION (MR). PATIENT HAD SIGNIFICANT RA DILATATION, MINOR SCOLIOSIS OF THE LOWER LUMBAR REGION ALONG WITH A HYPERELASTIC SEPTUM. TRANSSEPTAL ACCESS WAS DIFFICULT. ACCESS WAS GAINED WITH THE BAYLIS SYSTEM ONCE, THEN AS WE WERE ADVANCING THE STEERABLE GUIDE CATHETER (SGC) ACROSS THE SEPTUM, LOST ACCESS TO THE LEFT SIDE WITH THE WIRE. ENDED UP HAVING TO DO ANOTHER TRANSSEPTAL USING AN AGILIS STEERABLE SHEETH. ACCESS WAS GAINED BUT STRUGGLED TO GET THE SGC TO ACTUALLY TRAVERSE THE SEPTUM. AFTER SEVERAL ATTEMPTS OF TRYING TO CROSS A SEPTOPLASTY WAS PERFORMED WITH A 14MM BALLOON THAT THE PHYSICIAN DECIDED TO USE. EVEN AFTER DILATING THE SEPTUM WITH THE BALLOON, THE SGC STILL WOULD NOT CROSS. ALSO ATTEMPTED A BUDDY WIRE TECHNIQUE AFTER THE SEPTOPLASTY TO TRY AND PROVIDE ADDITIONAL SUPPORT FOR THE GUIDE. THIS TOO DID NOT WORK AND AFTER SEVERAL HOURS, THE PHYSICIAN DECIDED TO ABORT THE CASE. NO CLIP WAS INTRODUCED. POST- PROCEDURE, THE PATIENT IN RECOVERY HAD FLASH PULMONARY EDEMA. ON AN UNKNOWN DATE, THE PATIENT DIED. IT WAS THOUGHT THE DEATH WAS RELATED TO POOR STATE OF HEALTH ALONG WITH THE PROGRESSIVE SEVERE MR THAT WENT UNSOLVED DUE TO THE ABORTED MITRACLIP PROCEDURE.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: THE PATIENT PRESENTED WITH FUNCTIONAL MITRAL REGURGITATION (MR) GRADE 4. DURING THE PROCEDURE, A NON-ABBOTT TRANSSEPTAL SYSTEM WAS INITIALLY USED, AND AFTER THE FIRST TRANSSEPTAL PUNCTURE, GUIDEWIRE ACCESS TO THE LEFT ATRIUM WAS LOST. ANOTHER TRANSSEPTAL PUNCTURE WAS PERFORMED USING A NON-ABBOTT STEERABLE SHEATH, BUT THE STEERABLE GUIDE CATHETER (SGC) STILL FAILED TO CROSS THE SEPTUM. DESPITE PERFORMING A SEPTOPLASTY WITH A 14MM BALLOON AND ATTEMPTING A BUDDY WIRE TECHNIQUE, THE SGC COULD NOT TRAVERSE THE SEPTUM. THE PROCEDURE WAS ULTIMATELY ABORTED AFTER SEVERAL HOURS, AND THE MR REMAINED UNCHANGED AT GRADE 4. AN ATRIAL SEPTAL DEFECT (ASD) WAS OBSERVED HOWEVER, NO TREATMENT WAS PERFORMED. POST-PROCEDURE WHILE IN RECOVERY, THE PATIENT EXPERIENCED FLASH PULMONARY EDEMA. THE CAUSE OF THE PULMONARY EDEMA WAS UNKNOWN. THEN, ON AN UNKNOWN DATE THE PATIENT DIED. THE POSSIBLE CAUSE OF THE DEATH WAS RELATED TO THE PATIENT¿S POOR OVERALL HEALTH ALONG WITH PROGRESSIVE SEVERE MR THAT WAS UNRESOLVED AFTER THE ABORTED MITRACLIP PROCEDURE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 DEATH DATE IS ESTIMATED AS IT WAS NOT PROVIDED. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS FROM THE LOT. ALL AVAILABLE INFORMATION WAS INVESTIGATED AND THE REPORTED FAILURE TO ADVANCE THE SGC APPEARS TO BE RELATED TO PATIENT CONDITIONS (HYPER ELASTIC SEPTUM). PERFORATION (ATRIAL SEPTAL DEFECT/ASD) APPEARS TO BE RELATED TO PROCEDURAL CONDITIONS ASSOCIATED WITH THE ATRIAL SEPTAL PUNCTURE. A CAUSE FOR THE REPORTED PULMONARY EDEMA CANNOT BE DETERMINED. ADDITIONALLY, BASED ON THE INFORMATION PROVIDED BY THE ACCOUNT AND THE MEDICAL REVIEW, THE REPORTED DEATH APPEARS TO BE DUE TO PATIENT'S COMORBIDITIES IN COMBINATION WITH SEVERE MR. PULMONARY EDEMA, PERFORATION AND DEATH ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125029&pc=DTB | 2649622-2025-01036 | 2024-12-24 05:00 | Death | MPRI | 2025-01-10 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING THAT RESULTED IN FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) DETECTION AND INAPPROPRIATE ATRIAL PACING. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Event Description: IT WAS FURTHER REPORTED THROUGH FOLLOW-UP THAT THE PATIENT IS DECEASED. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125103&pc=OZD | 1220648-2025-25483 | 2024-12-28 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND WENT INTO VENTRICULAR TACHYCARDIA AND SUBSEQUENTLY EXPIRED. THE IMPELLA CP DEVICE WAS SUCCESSFULLY IMPLANTED AND FOLLOWING THE IMPLANT THE PUMP "DOVE" TOO FAR INTO THE LEFT VENTRICLE CAUSING THE PATIENT TO GO INTO VENTRICULAR TACHYCARDIA. THE PATIENT WAS SHOCKED ONE TIME TO REGAIN A STABILE RHYTHM. PERCUTANEOUS CORONARY INTERVENTION WITH FOUR DRUG-ELUTING STENTS TO THE LEFT ANTERIOR DESCENDING AND LEFT CIRCUMFLEX ARTERIES WAS COMPLETED. SUPPORT WAS CONTINUED FOR ONE MORE DAY. THE PATIENT REPORTEDLY EXPIRED THE FOLLOWING DAY. IT IS UNKNOWN IF THE VENTRICULAR TACHYCARDIA EVENT CONTRIBUTED TO THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125313&pc=DQA | 2936999-2025-00007 | 2024-11-19 05:00 | Death | MEDIANA CO. LTD. | 2025-01-10 05:00 | DQA | NELLCOR | Device Alarm System | Cardiac Arrest | K142865 |  | 1 | Event Description: IT WAS REPORTED THAT THE PULSE OXIMETER WAS INVOLVED IN AN INCIDENT WHERE THE ALARM ALLEGEDLY DID NOT SOUND. THE EVENT TOOK PLACE AT HOME, AND THE PARAMEDIC REPORTED ATTENDING THE PATIENT AT 04:12AM, PRONOUNCING RECOGNITION OF LIFE ENDED (ROLE) AT 04:13AM. THE ONSET OF DETERIORATION APPEARED TO BE AROUND 02:41AM ACCORDING TO THE DEVICE'S TABULAR AND GRAPHICAL TRENDS. UPON INVESTIGATION, THE DEVICE WAS FOUND IN NORMAL MODE AND SET TO ADULT MODE, RATHER THAN SLEEP STUDY MODE. THE DEVICE HAD BEEN QUARANTINED SINCE THE INCIDENT, AND DATA WAS DOWNLOADED FROM THE DEVICE USING THE NELLCOR ANALYTICS TOOL BY NON-MEDTRONIC STAFF. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G3, H3, H6 H3 EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS AVAILABLE FOR EVALUATION. THE EVALUATION FOUND NO POTENTIALLY CONTRIBUTING FACTORS, AND THE SAMPLE MET ALL RELATED SPECIFICATIONS. IT WAS REPORTED THAT THE PULSE OXIMETER WAS INVOLVED IN AN INCIDENT WHERE THE ALARM ALLEGEDLY DID NOT SOUND. THE REPORTED ISSUE COULD NOT BE CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE IDENTIFIED BECAUSE NO RELATED PROBLEM WAS DETECTED WITH THE DEVICE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125427&pc=MKJ | 1220908-2025-00088 | 2024-12-30 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-10 05:00 | MKJ | X SERIES | Incorrect Interpretation of Signal | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 71-YEAR-OLD MALE PATIENT, THE DEVICE ISSUED A "SHOCK ADVISED" PROMPT FOR A HEART RHYTHM THEY BELIEVED WAS NON-SHOCKABLE. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE CUSTOMER WAS CONTACTED FOR RETURN OF THE SUSPECT PRODUCT AND CLINICAL LOG. THE PRODUCT AND CLINICAL LOG FILES HAVE NOT BEEN RETURNED TO ZOLL FOR EVALUATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125937&pc=OZD | 1220648-2025-25493 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Obstruction/Occlusion | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT COMPLICATIONS WERE ENCOUNTERED DURING ATTEMPTED DELIVERY OF AN IMPELLA CP PUMP FOR MECHANICAL CIRCULATORY SUPPORT. FOLLOWING INSERTION, IT WAS NOTED THAT THE PATIENT HAD A CALCIFIED ANATOMY THAT PREVENTED SUCCESSFUL ADVANCEMENT OF THE DEVICE. THE IMPLANT WAS ABORTED. THE PATIENT PASSED AWAY IN THE PROCEDURAL AREA SUBSEQUENTLY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL DETAILS PROVIDED WERE SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, IMPELLA UNABLE TO PASS THROUGH THE VASCULATURE AS PER THE CLINICAL DESCRIPTION PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126177&pc=DSQ | 2916596-2025-00211 | 2022-09-01 04:00 | Death | THORATEC CORPORATION | 2025-01-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿ASPIRIN AND HEMOCOMPATIBILITY EVENTS WITH A LEFT VENTRICULAR ASSIST DEVICE IN ADVANCED HEART FAILURE¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH ISCHEMIC STROKE, HEMORRHAGIC STROKE, MAJOR BLEEDING, GASTROINTESTINAL BLEEDING, AND DEATH. UNDER THE ARIES-HM3 TRIAL, 628 PATIENTS WITH HEARTMATE 3 LVADS WERE RANDOMIZED TO RECEIVE VITAMIN K AND EITHER ASPIRIN (100 MG/D) OR A PLACEBO FROM JUL2020 TO SEP2022. 39 PATIENTS WERE NOT INCLUDED IN THE PRIMARY END POINT ANALYSIS, DUE TO SURGICAL EVENTS THAT RESULTED IN CESSATION OF THE TREATMENT GROUP MEDICATION (EG, BLEEDING EVENTS) OR SURGICAL OUTCOMES, SPECIFICALLY DEATH OR WITHDRAWAL (3 PATIENTS PASSED AWAY AFTER RANDOMIZATION). THUS, THE PRIMARY ANALYSIS POPULATION INCLUDED 296 PATIENTS ASSIGNED TO THE PLACEBO GROUP AND 293 PATIENTS ASSIGNED TO THE ASPIRIN GROUP. MORE PATIENTS REACHED 12-MONTH SURVIVAL FREE OF STROKE, PUMP THROMBOSIS, MAJOR BLEEDING, OR ARTERIAL PERIPHERAL THROMBOEMBOLISM IN THE PLACEBO GROUP THAN IN THE ASPIRIN GROUP [68.1% VS. 74.2%], LARGELY DRIVEN BY REDUCTION IN BLEEDING EVENTS IN THE PLACEBO GROUP (22.5% VS 28.2%). 1.5% OF PATIENTS IN BOTH GROUPS PASSED AWAY, AND 1.8% OF THE PLACEBO GROUP EXPERIENCED STROKE VS. 2.2% IN THE ASPIRIN GROUP. THERE WERE FEWER PATIENTS WHO DIED OR EXPERIENCED A NONSURGICAL MAJOR HEMOCOMPATIBILITY-RELATED ADVERSE EVENT IN THE PLACEBO GROUP THAN THE ASPIRIN GROUP THROUGH 24 MONTHS (36.9% VS 45.9%; P= .03), FEWER FIRST NONSURGICAL BLEEDING EVENTS IN THE PLACEBO THAN ASPIRIN GROUP AT 24 MONTHS (30.0 VS. 42.4, P = 0.1), LOWER RATES OF NONSURGICAL BLEEDING HOSPITALIZATIONS (13.6 VS. 23.9 EVENTS PER 100 PATIENT YEARS, P = 0.002), AND 293 (47%) FEWER DAYS SPENT IN THE HOSPITAL FOR BLEEDING. THERE WERE NO DIFFERENCES IN ISCHEMIC AND HEMORRHAGIC STROKE EVENTS, AND THERE WERE NO INSTANCES OF PUMP THROMBOSIS OR ARTERIAL PERIPHERAL THROMBOEMBOLISM. Manufacturer Narrative: SECTION A AND D - SPECIFIC PATIENT AND DEVICE INFORMATION ARE UNKNOWN. MEHRA MR, NETUKA I, URIEL N, ET AL. ASPIRIN AND HEMOCOMPATIBILITY EVENTS WITH A LEFT VENTRICULAR ASSIST DEVICE IN ADVANCED HEART FAILURE: THE ARIES-HM3 RANDOMIZED CLINICAL TRIAL. JAMA. 2023;330(22):2171¿2181. DOI:10.1001/JAMA.2023.23204. (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126693&pc=DSQ | 2916596-2025-00032 | 2024-12-05 05:00 | Death | THORATEC CORPORATION | 2025-01-10 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pain; Depression; Ascites; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Renal Impairment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS IN A COMPLEX STATE. THEY WERE ORIGINALLY ADMITTED WITH RIGHT HEART FAILURE (RHF), RENAL INSUFFICIENCY, PAIN SYNDROME, DEPRESSION, AND PHLEGMONA ON THE LOWER LIMB. COMPLEX THERAPEUTIC INTERVENTIONS WERE PERFORMED THAT INCLUDED INOTROPES, DIALYSIS, AND NUTRITIONAL SUPPORT. EVACUATION OF ASCITES WAS PERFORMED WHICH WAS COMPLICATED BY HEMOPERITONEUM AND SURGICAL INTERVENTION WAS REQUIRED. DESPITE COMPLEX THERAPY THE PATIENT SUFFERED FROM MULTIPLE ORGAN FAILURE (MOF) AND SUBSEQUENTLY DIED.Event Description: THE RIGHT HEART FAILURE DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT AND THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE RIGHT HEART FAILURE. THE RIGHT HEART FAILURE WAS TREATED WITH INOTROPES. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING BLEEDING, PSYCHIATRIC EPISODE, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS NEUROLOGIC DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6 ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126876&pc=NPU | 2015691-2025-00238 | 2024-12-06 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-10 05:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Difficult or Delayed Positioning; Difficult to Remove; Device Slipped; Adverse Event Without Identified Device or Use Problem | Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN CANADA, THIS WAS A VALVE IN MAC CASE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE, IN OPEN MITRAL POSITION. DURING THE PROCEDURE, THE ANTERIOR LEAFLET WAS RESECTED, AND THE VALVE WAS INTRODUCED IN THE MITRAL ORIENTATION. THE VALVE WAS CRIMPED DIRECTLY ON THE BALLOON. IT WAS DIFFICULT TO ADVANCE THROUGH THE MITRAL VALVE AND THE VALVE MOVED ON THE BALLOON. DURING INITIAL INFLATION, NO TRANSLATION ON VALVE EXPANSION WAS NOTED AND IT WAS RECOGNIZED THAT THE VALVE SEEMED TO BE OFF BALLOON. THE BALLOON CAME OFF THE VALVE WHILE ATTEMPTING TO ADVANCE THE DELIVERY SYSTEM IN THE TIGHT ANATOMICAL MITRAL VALVE. NO LEAK WAS NOTICED ON THE DELIVERY SYSTEM. THEREFORE, THE BALLOON THAT WAS STUCK IN THE VENTRICLE WAS RETRACTED BY PERFORMING SMALL BALLOON INFLATIONS AND RETRACTIONS TO ALIGN THE VALVE ON THE BALLOON. ONCE THE VALVE WAS ON THE BALLOON, THE VALVE WAS SUCCESSFULLY DEPLOYED AT A NOMINAL VOLUME. AFTER VALVE DEPLOYMENT, A VENTRICULAR SEPTAL DEFECT (VSD) WAS NOTICED. A PATCH REPAIR WAS SUCCESSFULLY DONE. THE HEART WAS CLOSED AND NO VENTRICULAR SEPTAL DEFECT OR PARAVALVULAR LEAK WAS NOTED ON ECHO. NO DAMAGE WAS OBSERVED ON THE BALLOON. THE PATIENT PASSED AWAY. THE PERCEIVED CAUSE OF DEATH WAS VSD COMPLICATIONS.Event Description: AS REPORTED BY OUR AFFILIATES IN CANADA, THIS WAS A VALVE IN MAC CASE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE, IN OPEN MITRAL POSITION. DURING THE PROCEDURE, THE VALVE WAS CRIMPED DIRECTLY ON THE BALLOON. THE FLEX TIP REMAINED IN THE CRIMPING POSITION AND NOT AGAINST THE VALVE FOR THE ENTIRETY OF THE CASE. THE ANTERIOR LEAFLET WAS RESECTED, AND THE VALVE WAS INTRODUCED IN THE MITRAL ORIENTATION. THE VALVE WAS CRIMPED DIRECTLY ON THE BALLOON. IT WAS DIFFICULT TO ADVANCE THROUGH THE MITRAL VALVE AND THE VALVE MOVED ON THE BALLOON. DURING INITIAL INFLATION, NO TRANSLATION ON VALVE EXPANSION WAS NOTED AND IT WAS RECOGNIZED THAT THE VALVE SEEMED TO BE OFF BALLOON. THE BALLOON CAME OFF THE VALVE WHILE ATTEMPTING TO ADVANCE THE DELIVERY SYSTEM IN THE TIGHT ANATOMICAL MITRAL VALVE. NO LEAK WAS NOTICED ON THE DELIVERY SYSTEM. THEREFORE, THE BALLOON THAT WAS STUCK IN THE VENTRICLE WAS RETRACTED BY PERFORMING SMALL BALLOON INFLATIONS AND RETRACTIONS TO ALIGN THE VALVE ON THE BALLOON. THE VALVE ALIGNMENT FUNCTION WAS NOT UTILIZED SINCE THE FLEX TIP WAS NOT AGAINST THE VALVE, MAKING THIS FUNCTION NOT POSSIBLE. ONCE THE VALVE WAS ON THE BALLOON, THE VALVE WAS SUCCESSFULLY DEPLOYED AT A NOMINAL VOLUME. AFTER VALVE DEPLOYMENT, A VENTRICULAR SEPTAL DEFECT (VSD) WAS NOTICED. A PATCH REPAIR WAS SUCCESSFULLY DONE. THE HEART WAS CLOSED AND NO VENTRICULAR SEPTAL DEFECT OR PARAVALVULAR LEAK WAS NOTED ON ECHO. NO DAMAGE WAS OBSERVED ON THE BALLOON. THE PATIENT PASSED AWAY. THE VSD COULD NOT BE ISOLATED TO A SINGLE EVENT, AS IT WAS ONLY NOTED AFTER A SERIES OF ALL THE STEPS WERE TAKEN. THE PERCEIVED CAUSE OF DEATH WAS VSD COMPLICATIONS. Manufacturer Narrative: UPDATED SECTION B5 AND H6- DEVICE CODE. INVESTIGATION IS ONGOING. Manufacturer Narrative: UPDATED SECTION H6- TYPE OF INVESTIGATION, FINDINGS, AND CONCLUSIONS. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE COMPLAINTS FOR DIFFICULTY TO CROSS NATIVE ANNULUS, VALVE MOVEMENT ON BALLOON, AND DIFFICULTY TO WITHDRAW SYSTEM WITH VALVE THROUGH VASCULATURE WERE UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF DEVICE OR IMAGERY. IN THIS CASE, THERE WAS NO ALLEGATION A DEVICE MALFUNCTION CONTRIBUTED TO THE REPORTED EVENTS. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. AS REPORTED, "IT WAS DIFFICULT TO ADVANCE THROUGH THE TIGHT ANATOMICAL MITRAL VALVE AND THE VALVE MOVED ON THE BALLOON." PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, SMALL VESSEL SIZE, PRE-EXISTING STENTS, ETC.) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. IN THIS CASE, THE OPEN-HEART PROCEDURE WITH A TIGHT ANATOMICAL MITRAL VALVE WAS REPORTED TO HAVE CAUSED THE CROSSING DIFFICULTY. AS SUCH, AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (PATIENT ANATOMY) AND/OR OFF LABEL USE (OPEN HEART THV PROCEDURE) MAY HAVE CONTRIBUTED TO DIFFICULTY TO CROSS NATIVE ANNULUS. HOWEVER, A DEFINITE ROOT CAUSE WAS UNABLE TO BE DETERMINED. AS REPORTED, "DURING PROCEDURE, THE VALVE WAS CRIMPED DIRECTLY ON THE BALLOON. THE FLEX TIP REMAINED IN THE CRIMPING POSITION AND NOT AGAINST THE VALVE FOR THE ENTIRETY OF THE CASE. ANTERIOR LEAFLET RESECTED AND THE VALVE WAS INTRODUCED IN MITRAL ORIENTATION. IT WAS DIFFICULT TO ADVANCE THROUGH THE TIGHT ANATOMICAL MITRAL VALVE AND THE VALVE MOVED ON THE BALLOON". PER TRAINING MANUAL, IT STATES "ENSURE THE FLEX CATHETER TIP IS FLUSH WITH THE THV FOR SUPPORT DURING CROSSING". IN THIS CASE IT WAS REPORTED THAT THE FLEX TIP REMAINED IN DEFAULT POSITION DURING THE ENTIRE PROCEDURE AND THEREFORE THE VALVE WAS NOT SUPPORTED BY THE FLEX TIP WHILE CROSSING THE MITRAL VALVE. ADDITIONALLY, IT WAS REPORTED THAT DIFFICULTIES WERE ENCOUNTERED TO CROSS THE MITRAL VALVE. IT IS POSSIBLE THAT THE CRIMPED VALVE INTERACTED OR CAUGHT ON THE MITRAL VALVE WHILE ATTEMPTING TO CROSS THE MITRAL VALVE, AND IN COMBINATION WITH ADDITIONAL DEVICE MANIPULATION TO OVERCOME THE RESISTANCE, MAY HAVE CAUSED THE THV TO MOVE PROXIMALLY ON THE BALLOON. AS SUCH, AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (NOT FOLLOWING INSTRUCTIONS, DEVICE MANIPULATION) MAY HAVE CONTRIBUTED TO VALVE MOVEMENT ON BALLOON. HOWEVER, A DEFINITE ROOT CAUSE WAS UNABLE TO BE DETERMINED. AS REPORTED, "THE BALLOON WAS STUCK IN THE VENTRICLE BECAUSE THE S3UR VALVE WAS TOO TIGHT IN THE MITRAL VALVE (RECALL THAT THE SAPIEN VALVE WAS SO TIGHT IN THE NATIVE ANATOMY THAT THE BALLOON CAME OFF THE VALVE WHEN ATTEMPTING TO ADVANCE THE SYSTEM IN THE TIGHT ANATOMICAL MITRAL VALVE). IT WAS DIFFICULT TO RETRACT THE BALLOON TO GET IT ON THE VALVE BECAUSE IT WAS INCREDIBLY TIGHT. A SERIES OF MICRO RETRACTIONS ON THE BALLOON WITH MICRO BALLOON INFLATIONS WERE ATTEMPTED TO MARGINALLY INFLATE THE VALVE ENOUGH TO GET THE BALLOON ON THE VALVE." PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT CAUSES DIFFICULTY DURING WITHDRAWAL. IN THIS CASE, IT WAS REPORTED THAT THE BALLOON WAS STUCK IN THE VENTRICLE BECAUSE THE S3UR VALVE WAS TOO TIGHT IN THE MITRAL VALVE AND IT WAS DIFFICULT TO RETRACT THE BALLOON TO GET IT ON THE VALVE BECAUSE IT WAS INCREDIBLY TIGHT. AS SUCH, AVAILABLE INFORMATION SUGGEST THAT PATIENT FACTORS (PATIENT ANATOMY) MAY HAVE CONTRIBUTED TO DIFFICULTY TO WITHDRAW SYSTEM WITH VALVE THROUGH VASCULATURE. HOWEVER, A DEFINITE ROOT CAUSE WAS UNABLE TO BE DETERMINED. A VENTRICULAR DEFECT (VSD) IS A HOLE IN THE SEPTUM THAT SEPARATES THE TWO LOWER CHAMBERS (VENTRICLE). A VSD ALLOWS AN ABNORMAL SHUNT, MEANING AN ABNORMAL PASSAGEWAY FOR BLOOD THAT SHOULD NOT BE OCCURRING. A MINOR SEPTAL DEFECT OR LEAK MAY GO UNNOTICED AND NOT CAUSE ANY SYMPTOMS. WHILE A LARGER LEAK MAY CAUSE SYMPTOMS SUCH AS DIFFICULT BREATHING, PALPATIONS, FAINTING, TIA, STROKE, AND/OR HEMODYNAMIC INSTABILITY. THESE LARGER SEPTAL DEFECTS MAY REQUIRE A CLOSURE DEVICE (SEPTAL OCCLUDER) TO CLOSE THE SEPTAL DEFECT. EXTREME CASES MAY REQUIRE SURGICAL INTERVENTION TO TREAT. TMVR PROCEDURES SOMETIMES ARE DONE REQUIRING TRANSSEPTAL ACCESS FOR PERCUTANEOUS MITRAL VALVE REPLACEMENT. IT IS THE MDS DECISION TO EITHER CLOSE THE TRANSSEPTAL OPENING WITH A SEPTAL OCCLUDER OR LEAVE THE OPENING TO CLOSE ON ITS OWN. CLOSING THE SEPTAL ACCESS DURING THE INITIAL TMVR WITH A SEPTAL OCCLUDER IS CONSIDERED PROCEDURAL AND NOT A SERIOUS INJURY (LIKE PERCLOSE IN THE TRANSFEMORAL APPROACH). THIS WAS A CASE IN WHICH THE THV WAS IMPLANTED IN AN OPEN-HEART SURGERY, SO NO NEED TO IN CROSS THE SEPTUM FOR THV IMPLANTATION. IN THIS CASE, AFTER VALVE DEPLOYMENT, VENTRICULAR SEPTAL DEFECT WAS NOTICED. A PATCH REPAIR WAS SUCCESSFULLY DONE. THE HEART WAS CLOSED AND NO VENTRICULAR SEPTAL DEFECT OR PARAVALVULAR LEAK WAS NOTED ON ECHO. HOWEVER, THE PATIENT PASSED AWAY. THE VSD COULD NOT BE ISOLATED TO A SINGLE EVENT, AS IT WAS ONLY NOTED AFTER A SERIES OF ALL THE STEPS WERE TAKEN. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126937&pc=OZD | 1220648-2025-25500 | 2024-12-24 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ATTEMPTED TO BE IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. UPON ATTEMPTING TO PASS THE IMPELLA CP THROUGH ILIACS THROUGH THE AORTA, IT WAS NOT POSSIBLE TO ADVANCE THE IMPELLA CP DUE TO PATIENT ANATOMY AND THE PROCEDURE WAS ABORTED. IT WAS REPORTED THAT THE PATIENT DECOMPENSATED FURTHER AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE THE INSERTION WAS ABORTED DUE TO PATIENT ANATOMY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21127115&pc=BZD | 2518422-2025-001782 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP | Degraded | Respiratory Failure | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND RESPIRATORY FAILURE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21127168&pc=NPT | 2015691-2025-00242 | 2024-07-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-01-10 05:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, AFTER A TRANS-AORTIC TAVR PROCEDURE WITH A 23MM SAPIEN 3 VALVE, BLEEDING OCCURRED FROM THE STERNOTOMY SITE THE SAME DAY AS THE PROCEDURE. RETHORACOTOMY WAS PERFORMED TO STOP THE BLEEDING, AND THE PATIENT RECOVERED SLOWLY. ON POSTOPERATIVE DAY 14, THE PATIENT'S CONDITION WAS NOTED TO DETERIORATE, AND A DETAILED EXAMINATION REVEALED WIDESPREAD INTESTINAL ISCHEMIA. CONSERVATIVE TREATMENT WAS CONTINUED, BUT THE PATIENT PASSED AWAY THE FOLLOWING DAY. THE CAUSE OF DEATH WAS REPORTED AS INTESTINAL ISCHEMIA. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. A PROCEDURAL OR POST PROCEDURAL BLEED/ HEMORRHAGE WITHOUT AN OBVIOUS SOURCE OF BLEEDING IS A RARE BUT POTENTIALLY LIFE-THREATENING COMPLICATION OF CORONARY/CARDIAC INTERVENTIONAL PROCEDURES AND TENDS TO OCCUR MORE FREQUENTLY IN THE PRESENCE OF MORE AGGRESSIVE ANTICOAGULATION REGIMENS. POSSIBLE SOURCES INCLUDE PUNCTURE OF THE FEMORAL ARTERY ABOVE THE INGUINAL LIGAMENT AND ABOVE THE INFERIOR EPIGASTRIC ARTERY, ALLOWING THE RESULTANT BLEEDING TO EXTEND INTO THE RETROPERITONEAL SPACE, OR INADVERTENT TRAUMA OR PERFORATION OF THE ANNULAR STRUCTURE OR VENTRICLES DURING THE PROCEDURE. THIS TYPE OF BLEEDING MAY NOT BE RECOGNIZED UNTIL THE POST PROCEDURAL PERIOD. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL PROCEDURE INCLUDE THROMBUS FORMATION, PLAQUE DISLODGMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION, STROKE, DISTAL PERIPHERAL OCCLUSION, AND/OR DEATH. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS WIPING AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE THE CLOT. THE THV TRAINING MANUALS AND IFU INSTRUCT THE OPERATOR TO ADMINISTER HEPARIN AND MAINTAIN THE ACT AT = 250 SEC. THERE ARE MULTIPLE ETIOLOGIES FOR BOWEL ISCHEMIA/INFARCTION INCLUDING EMBOLIZATION OCCURRING AFTER DEVICE MANIPULATION, BAV, OR VALVE DEPLOYMENT AND THESE EVENTS CAN RESULT IN VARYING DEGREES OF PERMANENT IMPAIRMENT. ADDITIONALLY, PROLONGED HYPOTENSION CAN CONTRIBUTE TO DECREASED BOWEL PERFUSION AND ISCHEMIA POTENTIALLY LEADING TO TISSUE NECROSIS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE HEMORRHAGE/BLEEDING AND ISCHEMIA CANNOT BE DETERMINED, IT IS POSSIBLE THE MECHANISMS OF THE TAVR PROCEDURE DESCRIBED ABOVE OR PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21127169&pc=OZD | 1220648-2025-25501 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A HEMATOMA WAS NOTED AT THE GROIN SITE SHORTLY AFTER ARRIVAL TO THE INTENSIVE CARE UNIT FROM IMPELLA CP PLACEMENT. MANUAL PRESSURE WAS HELD AND HEPARIN DRIP WAS STOPPED AND ON HOLD. THE NEXT DAY, THE PATIENT HAD SLIGHT OOZING AT THE SITE WITHOUT HEMATOMA. IT WAS FURTHER REPORTED THAT ON DAY SIX OF IMPELLA SUPPORT, THE PATIENT HAD A PERCUTANEOUS CORONARY INTERVENTION PROCEDURE AND WAS COMPLAINING OF SEVERE ABDOMINAL PAIN AND BECAME LESS RESPONSIVE. THE FAMILY DECIDED TO MAKE THE PATIENT DO-NOT-RESUSCITATE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL INFORMATION WAS SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING IS DETERMINED TO BE PATIENT MOVEMENT SINCE THE HEMATOMA APPEARED UPON TRANSFERRING TO THE INTENSIVE CARE UNIT. B.2 REVISED AS REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25501. H.6 CODE 1888 WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25501.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21127181&pc=PYX | 1220648-2025-25502 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | PYX | IMPELLA RP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT EXPERIENCED OOZING AT THE INSERTION SITE POST PEEL AWAY EXCHANGE WITH SHEATH DURING IMPELLA RP PLACEMENT. OOZING RESOLVED WITH EXTRA MATTRESS SUTURE ADDED BY WRAPPING AROUND SHEATH. THE PATIENT¿S FAMILY DECIDED ON COMFORT CARE UPON ARRIVAL TO THE INTENSIVE CARE UNIT AND PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL INFORMATION WAS SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE OOZING WAS RESOLVED BY ADDING EXTRA MATTRESS SUTURE. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA RP WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS AND CAUTIONS. ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND. ¿BE SURE THAT THE STOPCOCK ON THE REPOSITIONING SHEATH IS ALWAYS KEPT IN THE CLOSED POSITION. SIGNIFICANT BLEED BACK CAN RESULT IF THE STOPCOCK IS OPEN.¿ ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE FEMORAL VEIN. CLOSE AND DRESS THE WOUND.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21127568&pc=MGB | 2024168-2025-00437 | 2016-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-01-10 05:00 | MGB | PERCLOSE PROGLIDE¿ | Insufficient Information | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THIS IS A RETROSPECTIVE STUDY OF PATIENTS WHO UNDERWENT TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURES IN COMMON FEMORAL ARTERIES. PROGLIDE DEPLOYMENT WAS THE CLOSURE METHOD WITH VASCULAR COMPLICATIONS OF STROKE, BLEEDING, PLACEMENT OF PACEMAKER, DISSECTION, PSEUDOANEURYSM, HEMATOMA, ISCHEMIA, STENOSIS, ARTERIOVENOUS FISTULA, AND DEATH WITH INTERVENTION OF MANUAL COMPRESSION, BALLOON INFLATION, STENTING, COVERED STENT, PACEMAKER, AND SURGERY INCLUDING SOME CASES OF PROLONGED HOSPITALIZATION. DATA SHOWED A SHORTER PROCEDURE DURATION, USE OF THE RIGHT FEMORAL ACCESS SITE, AND USE OF A SMALLER INTRODUCER SHEATH SIZE (14F) LED TO FEW COMPLICATIONS. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B6 - DATE OF EVENT ESTIMATED. D4: PRIMARY UDI NUMBER ¿ THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND/OR MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: TITLE "PREVALENCE OF POSTTRANSCATHETER AORTIC VALVE IMPLANTATION VASCULAR COMPLICATIONS IN REAL LIFE." Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF THE PRODUCTION RECORD AND COMPLAINT HISTORY OF THE REPORTED LOT COULD NOT BE CONDUCTED BECAUSE THE PART AND LOT NUMBER WERE NOT PROVIDED. THE PATIENT EFFECT OF DEATH IS LISTED IN THE ELECTRONIC PROGLIDE INSTRUCTIONS FOR USE (EIFU), AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF THE DEVICE. BASED ON THE INFORMATION REPORTED THROUGH THE RESEARCH ARTICLE, A CONCLUSIVE CAUSE FOR THE UNSPECIFIED DEVICE ISSUE COULD NOT BE DETERMINED. ADDITIONALLY, A DEFINITIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21127642&pc=BZD | 2518422-2025-001785 | 2024-03-27 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DS2ADV AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP THAT THE PATIENT ALLEGES DEATH. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER¿S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 22 MAY 2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP THAT THE PATIENT ALLEGES DEATH. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G AND H UPDATED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21128008&pc=DSQ | 2916596-2025-00125 | 2024-08-13 04:00 | Death | THORATEC CORPORATION | 2025-01-10 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Infection; Sepsis; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO INTRACEREBRAL HEMORRHAGE. THE OUTCOME WAS DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED AND AN AUTOPSY WAS NOT PERFORMED.Event Description: LEADING UP TO PATIENT'S DEATH, THEY HAD SEPSIS SECONDARY TO A DRIVELINE INFECTION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED. SECTION H3: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21128353&pc=BZD | 2518422-2025-001778 | 2022-09-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129057&pc=OZD | 1220648-2025-25505 | 2024-12-13 05:00 | Death | ABIOMED, INC. | 2025-01-10 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD ESCALATED CARE FROM AN IMPELLA CP TO AN IMPELLA 5.5 PUMP. THE ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK PATIENT SUFFERED VENTRICULAR TACHYCARDIA AND ATRIAL FIBRILLATION WHILE ON IMPELLA 5.5 SUPPORT. THE PATIENT WAS ABLE TO BE CONVERTED TO NORMAL SINUS RHYTHM AND REMAINED ON SUPPORT FOR A TOTAL OF 12 DAYS UNTIL CARE WAS WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED VENTRICULAR TACHYCARDIA (VT) HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE VT COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129257&pc=BZD | 2518422-2025-100406 | 2025-01-02 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | REP DREAMSTATION AUTO CPAP, JP | Gas/Air Leak; Adverse Event Without Identified Device or Use Problem | Insufficient Information | UNKNOWN |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY WHILE USING THE REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT WAS FOUND UNRESPONSIVE AND HAD STOPPED BREATHING WHILE USING THE DEVICE. THE TREATMENT DATA WAS REVIEWED, AND A LARGE LEAK WAS FOUND DURING THE FIRST HOUR OF THE EVENT AND THE REST OF THE TIME WAS LEFT IN AN UNMASKED STATE. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129276&pc=DYB | 2015691-2025-00247 | 2024-12-13 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-10 05:00 | DYB | EDWARDS ESHEATH+ INTRODUCER SET | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES JAPAN AFFILIATE, THE PATIENT UNDERWENT A TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 20MM SAPIEN 3 ULTRA RESILIA VALVE IN THE NATIVE AORTIC POSITION. WITHIN THE FOLLOWING 4 DAYS POST PROCEDURE, THE PATIENT DEVELOPED VASCULAR OCCLUSION OF THE LOWER LIMBS AND PATIENT PASSED AWAY DUE TO A REPERFUSION INJURY. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS ARE INCONCLUSIVE DUE TO THE LIMITED INFORMATION PROVIDED. HOWEVER, PATIENT FACTORS MIGHT HAVE CONTRIBUTED TO THE EVENT, INCLUDING THE SEVERE CALCIFICATION OF THE NATIVE VALVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129523&pc=CAW | 2518422-2025-001791 | 2023-11-16 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | CAW | SIMPLYGO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K111885 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A SIMPLYGO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DEATH. THERE WAS REPORT OF PATIENT DEATH.
THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE
MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A SIMPLYGO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DEATH. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129528&pc=BZD | 2518422-2025-001790 | 2024-03-19 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973 AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129628&pc=BZD | 2518422-2025-001796 | 2023-08-07 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129829&pc=BZD | 2518422-2025-001795 | 2024-06-26 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129848&pc=FRN | 1314492-2024-04099 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-10 05:00 | FRN | SPECTRUM IQ | Filling Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K222048 |  | 1 | Event Description: IT WAS REPORTED BY THE NURSE, THAT DURING AN INFUSION OF HEPARIN WITH THE SPECTRUM PUMP, ALLEGEDLY THE PUMP DID NOT INFUSE PROPERLY, AND SUBSEQUENTLY THE PATIENT DIED. THE SPECTRUM PUMP WAS SET TO DELIVER THE HEPARIN (25000 UNITS/250ML) WITH A DOSE OF 15.2 UNITS/KG/HOUR (THOUGH IT WAS REPORTED BY THE NURSE THE PUMP "SHOULD HAVE BEEN SET TO THIS" WHICH WAS NOT FURTHER SPECIFIED) WITH A VOLUME TO BE INFUSED OF 14.5ML OVER 50 MINUTES WITH A RATE OF 17.6ML/HOUR. HOWEVER, ACCORDING TO THE NURSE THE HEPARIN BAG REMAINED FULL, AND IT APPEARS THE SOLUTION ¿DID NOT INFUSE AT ALL¿. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: THE DEVICE HAS BEEN RECEIVED AND THE EVALUATION IS IN PROGRESS. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION H2, H3, H6 AND H11: THE DEVICE WAS RECEIVED FOR EVALUATION. FUNCTIONAL TESTING WAS PERFORMED AND DID NOT IDENTIFY ANY ISSUES RELATED TO THE CUSTOMER REPORTED CONDITION. THE REPORTED WAS NOT REPRODUCED DURING EVALUATION. EVALUATION SENT DEVICE FOR FLOW TESTING AT 40 ML/HR FOR 30 MINS AT 20 VTBI WITH THE RESULTS OF 19.956 AND PASSING WITHIN DEVICE SPECIFICATION AT ERROR PERCENTAGE OF -0.22%. THE DEVICE FUNCTIONED AS INTENDED. A SERVICE HISTORY REVIEW WAS PERFORMED AND REVEALED THAT THE DEVICE HAS NO PREVIOUS SERVICE EVENTS; THEREFORE, SERVICING DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. A DEVICE HISTORY REVIEW REVEALED NO ISSUES THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THE EVENT HISTORY LOG REVIEW WAS PERFORMED. THE USER FACILITY DID NOT PROVIDE AN EVENT OCCURRENCE DATE, HOWEVER REVIEW OF THE LOG REVEALED AN INFUSION OF HEPARIN PROGRAMMED AND STARTED A FEW DAYS PRIOR TO THE EVENT BEING REPORTED. THE INFUSION CONTINUED INTO THE FOLLOWING DAY. AT TIMESTAMP 04:53 THE USER ACTIVATED STANDBY MODE. AT 06:00, THE USER RESUMED THE INFUSION AT A RATE OF 8.4 ML/HR AND EXPERIENCED AN "UPSTREAM OCCLUSION!" ALARM IN THE SAME TIMESTAMP. THE USER CLEARED THE ALARM AND RESTARTED THE PUMP IN THE SAME TIMESTAMP. THE PUMP CONTINUED FOR 3 HOURS AND 14 MINUTES WITHOUT INTERRUPTION. AT THAT TIME, THE USER STOPPED THE PUMP AND ACTIVATED STANDBY. THE USER RESTARTED THE PUMP 13 MINUTES LATER. ONE HOUR AND 3 MINUTES LATER, THE USER UPDATED THE RATE TO 13 ML/HR. AFTER 6 HOURS AND 42 MINUTES, THE USER UPDATED THE RATE TO 17.6 ML/HR. THE PUMP RAN FOR 2 HOURS AND 56 MINUTES WITHOUT INTERRUPTION BEFORE THE USER STOPPED THE PUMP AND DID NOT RESUME THE INFUSION. DETAILS REGARDING PUMP SETUP WERE NOT PROVIDED BY THE USER FACILITY AND COULD IMPACT FLOW ACCURACY. THE REPORTED CONDITION WAS NOT VERIFIED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129869&pc=DSQ | 2916596-2024-52864 | 2023-09-23 04:00 | Death | THORATEC CORPORATION | 2025-01-10 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE IS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129998&pc=BZD | 2518422-2025-001800 | 2024-01-09 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Dyspnea; Cough; Wheezing; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. AFTER THE THIRD ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED THAT THE DEVICE WILL BE AVAILABLE FOR EVALUATION, BUT IT HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. AFTER DEVICE RETURN AND COMPLETION OF EVALUATION, A FOLLOW-UP REPORT WILL BE FILED. DATE OF DEATH, PATIENT OUTCOME CODE HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129999&pc=BZD | 2518422-2025-001801 | 2024-05-01 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES .THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: PATIENT OUTCOME CODE GRID HAS BEEN UPDATED. IN BOX E:REPORTER PHONE HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130001&pc=BZD | 2518422-2025-001806 | 2024-05-13 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130002&pc=BZD | 2518422-2025-001798 | 2024-03-08 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130100&pc=PNJ | 9612164-2025-00209 | 2024-03-08 05:00 | Death | MEDTRONIC IRELAND | 2025-01-10 05:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Insufficient Information | P150033 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THIS LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). THE AUTHORS DESCRIBED ONE PATIENT DEATH WHICH OCCURRED IN THE EARLY PERI-PROCEDURAL TIMEFRAME; HOWEVER, THE CAUSE OF DEATH WAS UNKNOWN. THERE WERE TWO PATIENTS WHO EXPERIENCED PERICARDIAL EFFUSION; ONE REQUIRED PERCUTANEOUS DRAINAGE AND THE OTHER RESOLVED WITH OBSERVATION. THERE WERE TWO FAILED IMPLANTS DUE TO INADEQUATE MEASUREMENTS AFTER SEVERAL REPOSITIONING ATTEMPTS AND OTHER DEVICES WHICH EXHIBITED HIGH THRESHOLDS AT IMPLANT AND AT THREE MONTHS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE AGE CHARACTERISTICS IS 81 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: AVEIR VR REAL-WORLD PERFORMANCE AND CHRONIC PACING THRESHOLD PREDICTION USING MAPPING AND FIXATION ELECTRICAL DATA. EUROPACE. 2024. 26, 1¿8. DOI: 10.1093/EUROPACE/EUAE051 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130178&pc=DSQ | 2916596-2025-00029 | 2024-12-07 05:00 | Death | THORATEC CORPORATION | 2025-01-10 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PNEUMONIA, URINARY TRACT INFECTION (UTI), RIGHT VENTRICULAR DILATATION, RENAL CONGESTION, AND MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT RETURN.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE CAUSE OF THE RIGHT VENTRICULAR DILATION AND RENAL DYSFUNCTION WHERE DUE TO CHRONIC RIGHT VENTRICULAR FAILURE AND CONGESTION IN THE PATIENT. IT WAS NOTED THAT A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE, AND THE EVENT WAS NOT THOUGHT TO BE DEVICE RELATED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD MILD RIGHT VENTRICULAR DYSFUNCTION PRIOR TO IMPLANT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THIS IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HM3 LVAS. SECTION 6 "PATIENT CARE AND MANAGEMENT" DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 ALSO STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP.". NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130336&pc=BZD | 2518422-2025-001814 | 2024-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130348&pc=BZD | 2518422-2025-001789 | 2024-06-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION AUTO BIPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130398&pc=BZD | 2518422-2025-001784 | 2023-12-14 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION FROM PATIENT'S WIFE WHO STATED THAT SHE HAS BEEN GETTING PRODUCT NOTICE. SHE SAID HER HUSBAND PASSED AWAY LAST MARCH, AND SHE HAS GOTTEN RID OF THE DEVICE. HE HAD USED THE REPLACEMENT DEVICE. HE DID NOT PASS AWAY DUE TO USE OF THE DEVICE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130413&pc=BZD | 2518422-2025-001822 | 2024-06-14 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD DECEASED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130547&pc=BZD | 2518422-2025-001812 | 2023-05-02 04:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DS2ADV AUTO CPAP | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K200480 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130936&pc=NPT | 2015691-2025-00257 | 2024-12-12 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-10 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS AFFILIATES IN CANADA, DURING A TAVR PROCEDURE USING A 23MM SAPIEN 3 ULTRA VIA TRANSFEMORAL APPROACH, A PRE-DILATION WITH A 16MM EDWARDS BALLOON WAS PERFORMED BUT DUE TO THE AMOUNT OF CALCIUM, IT WAS NOT ABLE TO CROSS. A SECOND ATTEMPT WITH A 12MM NON-EDWARDS BALLOON WAS SUCCESSFUL. THE VALVE WAS SUCCESSFULLY DEPLOYED AT NOMINAL VOLUME AT 80/20 POSITION. THE DECISION WAS MADE TO DO A POST-DILATION WITH A 22MM NON-EDWARDS BALLOON. AFTER POST-DILATION, A LARGE ANNULAR INJURY WAS OBSERVED, AND THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE. ECMO WAS STARTED BUT THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS ARE INCONCLUSIVE AS RELEVANT PATIENT AND PROCEDURAL FACTORS WERE NOT PROVIDED. HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: CORRECTION TO H6; INVESTIGATION CONCLUSION. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PATIENT FACTORS (CALCIFICATION) AND ANATOMY OF THE SINUS, CAUSED OR CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21131344&pc=BZD | 2518422-2025-001825 | 2023-12-06 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21131659&pc=BZD | 2518422-2025-001815 | 2023-11-08 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer; Pancreatitis; Unspecified Gastrointestinal Problem |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, PANCREATIC, ESOPHAGEAL. THERE WAS REPORT OF PATIENT DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973, AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21131793&pc=NVY | 2017865-2025-10703 | 2024-12-17 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-10 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Signal Artifact/Noise; Failure to Capture | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED REMOTELY VIA MERLIN.NET. REVIEW OF THE TRANSMISSION REVEALED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED EPISODES OF FAILURE TO CAPTURE AND NOISE. THE PATIENT PASSED AWAY. CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21132014&pc=BZD | 2518422-2025-100427 | 2024-12-20 05:00 | Death | RESPIRONICS, INC. | 2025-01-10 05:00 | BZD | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY AND THE DEVICE WAS IN USE AT THE TIME THE PATIENT PASSED. IT IS REPORTED THE DEVICE DID NOT CAUSE ANY HARM TO THE PATIENT AND WAS NOT THE CAUSE OF THE PATIENT PASSING. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21132192&pc=NVN | 2017865-2025-10714 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-10 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10712. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10715. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21132193&pc=LWP | 2017865-2025-10712 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-10 05:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10714. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10715. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. ELECTRICAL, LONGEVITY, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21132194&pc=NVN | 2017865-2025-10715 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-10 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10712. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10714. IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: FOR D9 NOT BEING SELECTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138441&pc=NAY | MW5164582 |  | Death | INTUITIVE SURGICAL, INC. | 2025-01-10 05:00 | NAY | DA VINCI SURGICAL SYSTEM |  | Insufficient Information |  |  | 1 | Event Description: DRS. (B)(6), PLEASE NOTE, MS. (B)(6), WHO DIED WITHIN ONE YEAR FOLLOWING ROBOTIC MASTECTOMY FOR STAGE 2 BREAST CANCER: PLEASE NOTE - MS. (B)(6) CASE WAS ADVERTISED AS THE FIRST ROBOTIC MASTECTOMY OPERATION PERFORMED IN THE U.S. FOR TREATMENT OF BREAST CANCER: PLEASE NOTE, 5-YR SURVIVAL RATES FOR STAGE 2 BREAST CANCER USING THE CONVENTIONAL APPROACH TO MASTECTOMY IS OVER 80%. MS. (B)(6) DIED WITHIN ONE YEAR OF HER ROBOTIC MASTECTOMY SURGERY. PLEASE ENSURE THAT INTUITIVE SURGICAL HAS FORMALLY REPORTED THIS PATIENT MORTALITY FOLLOWING HER ROBOTIC MASTECTOMY TO THE FDA. ON JAN 5, 2025, AT 11:37 PM, (B)(6) WROTE: DRS. (B)(6), I AM WRITING THIS LETTER OF SAFETY AND ETHICS CONCERN/WARNING TO YOU BASED ON THE TWO RECENT HUMAN SUBJECT RESEARCH PUBLICATIONS FROM SOUTH KOREA, WHICH WERE SUPPORTED BY INTUITIVE SURGICAL: HTTPS://JAMANETWORK.COM/JOURNALS/JAMASURGERY/FULLARTICLE/2822315, HTTPS://WWW.NATURE.COM/ARTICLES/S41598-024-84807-0. THESE STUDIES INVOLVE THE USE OF THE DA VINCI ROBOTIC SYSTEM TO PERFORM MASTECTOMIES ON PATIENTS WITH BREAST CANCER, WHICH AS YOU KNOW WELL, IS CONSIDERED A NON-STANDARD AND OFF-LABEL APPROACH IN THE US AND WESTERN EUROPE. ADDITIONALLY, THESE EXPERIMENTAL SURGERIES WERE PERFORMED OUTSIDE OF A PROPERLY DESIGNED PROSPECTIVE CLINICAL TRIAL DESIGNED TO ESTABLISH ONCOLOGICAL NONINFERIORITY OF THIS NEW APPROACH TO MASTECTOMY FOR BREAST CANCER TREATMENT. IN BOTH THESE STUDIES INTUITIVE'S S. KOREAN PROXIES PERFORMED THIS EXPERIMENTAL ROBOTIC SURGICAL PROCEDURE ON WOMEN WITH BREAST CANCER - AND IN DIRECT BREACH OF THE WELL-ESTABLISHED STANDARD OF CARE FOR SURGICAL EXCISION OF BREAST CANCER. BOTH STUDIES EXPLICITLY STATE THEIR ANALYSES ARE "RETROSPECTIVE" - DESPITE THE FACT THAT PERFORMING THE QUITE SUBSTANTIAL NUMBER OF ROBOTIC MASTECTOMIES, DESCRIBED IN THESE PAPERS, REQUIRES EXTENSIVE "PROSPECTIVE" PLANNING. THUS, TO HAVE PERFORMED AS MANY ROBOTIC MASTECTOMIES AS THESE S. KOREAN SURGEONS DID PERFORM, OUTSIDE OF A PROPER PROSPECTIVE CLINICAL TRIAL SETTING, RAISES A MAJOR ETHICAL RED FLAG IN ANY REASONABLE HUMAN SUBJECT EXPERIMENTAL SETTING. IT IS ALMOST CERTAIN THAT THIS TRANSPARENT USE OF THE "RETROSPECTIVE" RUBRIC WAS INTENTIONALLY AIMED AT SKIRTING A ETHICALLY PROPER, WELL-DESIGNED AND ADEQUATELY POWERED PROSPECTIVE TRIAL TO ESTABLISH ONCOLOGICAL NON-INFERIORITY. WITH THIS LETTER, I AM INFORMING YOU THAT YOUR S. KOREAN COLLABORATORS' USE OF A "RETROSPECTIVE STUDY" RUBRIC IN DESCRIBING THESE PUBLICATIONS IS LIKELY TO BE DECEPTIVE AND ETHICALLY PROBLEMATIC - BY EXTENSION, INTUITIVE SURGICAL'S FINANCIAL SUPPORT OF THESE STUDIES OUGHT TO BE CAUSE FOR SIGNIFICANT CONCERN BY YOUR COMPANY AND ITS EXECUTIVES AND INVESTORS. SPECIFICALLY, BECAUSE IN NEITHER OF THE ABOVE REFERENCED STUDIES IS THERE ANY INDICATION AS TO WHETHER THE ROBOTIC MASTECTOMY OPERATIONS PERFORMED IN S. KOREA WERE DONE WITH PROPER INFORMED CONSENT TO THE PATIENTS ABOUT THE FACTS THAT: 1) THIS NEW APPROACH TO BREAST CANCER TREATMENT IS 100% EXPERIMENTAL, 2) ITS ONCOLOGICAL SAFETY OR EFFICACY HAS NOT BEEN ESTABLISHED, OR 3) THE ROBOTIC APPROACH TO MASTECTOMY FOR TREATMENT OF BREAST CANCER MAY BE ONCOLOGICALLY INFERIOR TO THE CONVENTIONAL APPROACH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108798&pc=BZD | 2518422-2025-000866 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21109229&pc=BZD | 2518422-2025-000927 | 2023-08-10 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: PATIENT OUTCOME CODE GRID, EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED. BOX B: ADVERSE EVENT/PRODUCT PROBLEM HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21109607&pc=LWR | 2025587-2025-00213 | 2024-10-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-09 05:00 | LWR | FREESTYLE AORTIC ROOT BIOPROSTHESIS | Insufficient Information; Patient Device Interaction Problem | Endocarditis; Insufficient Information | P970031 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION FROM THE PATIENTS SPOUSE THAT ON THE SAME DAY POST IMPLANT OF THIS 27MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. IT WAS REPORTED BY THE PATIENTS SPOUSE THAT THE PATIENT DIED DUE TO COMPLICATIONS OF THE VALVE SURGERY DUE TO A DAMAGED VALVE. THEREFORE, RELATEDNESS OF THE DEATH TO THE BIOPROSTHETIC VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED WHICH STATED THAT THE CAUSE OF DEATH WAS ENDOCARDITIS. Manufacturer Narrative: B5 AND H6 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110074&pc=BZD | 2518422-2025-001097 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110085&pc=BZD | 2518422-2025-000998 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110169&pc=BZD | 2518422-2025-001058 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110222&pc=BZD | 2518422-2025-001082 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110253&pc=BZD | 2518422-2025-001115 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110258&pc=BZD | 2518422-2025-001131 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LIVER DISEASE/TOXICITY, AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110376&pc=BZD | 2518422-2025-000921 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer; Lymphoma |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, LYMPHOMA/LUNG CANCER AND DEATH OF THE PATIENT. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973 AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110487&pc=LWR | 9612164-2025-00156 | 2024-12-24 05:00 | Death | MEDTRONIC MEXICO | 2025-01-09 05:00 | LWR | AVALUS AORTIC TISSUE VALVE | Insufficient Information | Insufficient Information | P170006 |  | 1 | Event Description: IT WAS REPORTED THAT POST IMPLANT OF THIS 27MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE PRODUCT COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: MEDTRONIC RECEIVED ADDITIONAL INFORMATION WHICH STATED THAT THE SITE ASSESSED THE EVENT AS NOT RELATED TO THE IMPLANT PROCEDURE AND UNKNOWN IF RELATED TO THE VALVE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110493&pc=BZD | 2518422-2025-100280 | 2024-02-19 05:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION GO AUTO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION GO AUTO DEVICE HAS PASSED AWAY WHILE USING THE DEVICE. THE PATIENT¿S SPOUSE REPORTED THAT THE PATIENT UNEXPECTEDLY PASSED AWAY WHILE USING THE CPAP. ACCORDING TO THE SPOUSE, THE PATIENT APPLIED THE MASK AT APPROXIMATELY 9:30 PM, AND SHE WENT TO BED AROUND 10:15 PM. AT 10:45 PM, SHE NOTICED THE PATIENT SNORING LOUDLY AND NUDGED HIM. HE CEASED SNORING, AND SHE BELIEVED HE HAD ROLLED OVER AND FALLEN BACK ASLEEP, AS WAS HIS USUAL PATTERN. THE SPOUSE AWOKE AT 5:30 AM TO SHOWER, AND UPON EXITING THE SHOWER, NOTICED SOMETHING WAS NOT QUITE RIGHT. SHE TURNED ON THE LIGHT AND OBSERVED THE PATIENT, WITH HIS HAND UNDER HIS HEAD, APPEARING UNRESPONSIVE. THE SPOUSE IMMEDIATELY CALLED 911 AT APPROXIMATELY 5:45-5:50 AM. EMERGENCY RESPONDERS ARRIVED WITHIN 5-10 MINUTES AND PRONOUNCED THE PATIENT DECEASED. THE OFFICIAL CAUSE OF DEATH WAS LISTED AS NATURAL CAUSES. THE PATIENT¿S OFFICIAL DATE OF DEATH IS RECORDED AS (B)(6) 2024; HOWEVER, THE SPOUSE BELIEVES THE PATIENT MAY HAVE PASSED THE NIGHT PRIOR. IT WAS REPORTED THAT THE PATIENT CLEANED THE DEVICE BY PUTTING THE MASK, TUBING AND WATER CHAMBER IN A BLUE BAG AND IT WOULD CLEAN IT SOMEHOW. THIS WAS COMPLETED ONCE A WEEK. DISTILLED WATER WAS USED FOR THE WATER RESERVOIR. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110501&pc=BZD | 2518422-2025-001143 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATENT HAS EXPIRED DUE TO EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER AND OTHER PANCREATIC ISSUE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21111338&pc=OZD | 1220648-2025-25441 | 2024-12-21 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA CP | Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: COMPLAINANT IN JAPAN REPORTED THE PATIENT HAD AN ATTEMPTED IMPELLA CP IMPLANT AND IT WAS ABORTED. THERE WAS AN ISSUE DURING THE INSERTION OF THE 16FR SHEATH. THE INSERTION WAS ABORTED, AND AN IMPELLA 5.5 WAS PLACED. THE PATIENT EVENTUALLY EXPIRED (REFER TO 24-40053, CI-39969).Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON SUPPORT. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA & RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED DELIVERY ISSUE HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE DELIVERY ISSUE COULD NOT BE DETERMINED. Manufacturer Narrative: CORRECTIONS TO THE INITIAL MFR REPORT: B1-SELECTED ADVERSE EVENT. B2-SELECTED DEATH AND ADDED DATE OF DEATH. B5-MSO REVIEW STATEMENT. H1-CHANGED TO DEATH. H6-ADDED IMPACT CODE 1802.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21111558&pc=BZD | 2518422-2025-001208 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21111601&pc=BZD | 2518422-2025-001184 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Liver Damage/Dysfunction; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21111806&pc=DSQ | 2916596-2024-52878 | 2024-12-18 05:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemothorax; Pneumothorax; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD RECURRENT BLEEDING FROM PULMONARY VASCULAR BED WHICH WAS NOT LEFT VENTRICULAR ASSIST DEVICE (LVAD) RELATED. THE LOG FILES CAPTURED SEVERAL LOW FLOW ALARMS BETWEEN 8 PM AND 9 PM ON (B)(6) 2024. THERE DID NOT APPEAR TO BE ANY TYPE OF MCS EQUIPMENT ISSUES WHICH CAUSED THESE EVENTS. THE LOW FLOW EVENTS SEEMED TO BE A PATIENT RELATED ISSUE AND NOT AN EQUIPMENT RELATED ISSUE.Event Description: LOG FILES FROM (B)(6)2025 CONTAINED LOW FLOW ESTIMATES AS WELL AS SUSTAINED LOW FLOW HAZARD EVENTS. THE LOW FLOW WAS DUE TO HEMOTHORAX, AND PNEUMOTHORAX AND A CHEST TUBE WAS PLACED. THE PATIENT HAS REMAINED IN THE HOSPITAL SINCE THE FAILED TRANSCATHETER AORTIC VALVE REPLACEMENT. THE HEMODYNAMIC STABILITY WAS FOLLOWED WITH FLUIDS AND BLOOD PRODUCTS. THE HEMOTHORAX AND PNEUMOTHORAX WAS NOT COMPLETELY RESOLVED. THE PATIENT REMAINED CRITICALLY ILL WITH DISCUSSION OF WITHDRAWAL OF CARE AND COMFORT MEASURES. THE PATIENT HAD POOR LUNG PARENCHYMA, RECURRENT BLEEDING, PNEUMOTHORAXES AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) DEPENDENCE. THE FLOW READINGS DID NOT APPEAR TO BE THE RESULT OF ANY EXTERNAL EQUIPMENT MALFUNCTION. THERE WERE NO UNUSUAL ROTOR FAULTS, PUMP TEMPERATURE, OR ANY ABNORMAL PUMP FAULTS SEEN IN THE LOG FILES. THE EQUIPMENT WAS OPERATING AS INTENDED. THE PATIENT WAS AWAITING HEART AND LIVER TRANSPLANT.Event Description: THE AORTIC VALVE ISSUES EXISTED PRE--IMPLANT AND A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO THE AORTIC VALVE ISSUES OR FAILURE OF THE AORTIC VALVE REPLACEMENT. THE PATIENT ULTIMATELY PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - IMPACT CODE AND MEDICAL DEVICE PROBLEM CODE CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS, WHICH THE ACCOUNT ATTRIBUTED TO HEMOTHORAX AND PNEUMOTHORAX. THE SUBMITTED CONTROLLER EVENT LOG FILE CAPTURED SEVERAL LOW-SPEED ADVISORIES, INTENTIONAL PUMP STOPS, AND BACKUP BATTERY NOT INSTALLED ALARMS ON 18DEC2024. ACCORDING TO THE ACCOUNT, THE DECREASE IN SPEED, INTENTIONAL PUMP STOPS, AND REMOVAL OF THE BACKUP BATTERY WERE PERFORMED BY THE CLINICIAN DURING TAVR DEVICE REMOVAL SURGERY. THESE EVENTS WERE ALSO ASSOCIATED WITH LOW FLOW HAZARD ALARMS, WHICH APPEARED TO RESOLVE ONCE THE STORED PATIENT SPEED WAS INCREASED FOLLOWING THE PROCEDURE. THE ADDITIONAL SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 19JAN2025 ¿ 21JAN2025. SEVERAL LOW FLOW HAZARD ALARMS WERE CAPTURED ON 19JAN2025 AND 20JAN2025. REVIEW OF THE SUBMITTED CONTROLLER PERIODIC LOG FILE REVEALED ADDITIONAL LOW FLOW ALARMS CAPTURED ON 18JAN2025. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED IN THE SUBMITTED LOG FILES, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED AT THE STORED PATIENT SPEED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, DOCUMENT (B)(4), REV. D, ARE CURRENTLY AVAILABLE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112015&pc=MIH | 9612164-2025-00168 | 2024-12-22 05:00 | Death | MEDTRONIC IRELAND | 2025-01-09 05:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Leak/Splash | Rupture | P100021 |  | 1 | Event Description: AN ENDURANT STENT GRAFT WAS IMPLANTED DURING THE INITIAL ENDOVASCULAR TREATMENT OF AN 52MM ABDOMINAL AORTIC ANEURYSM ON (B)(6) 2019. IT WAS NOTED THAT THE PATIENT HAD A SHORT ANGULATED AORTIC NECK. IT WAS REPORTED APPROX. 5 YEARS POST INDEX PROCEDURE THE PATIENT PRESENTED EMERGENTLY DUE TO AN ANEURYSM RUPTURE DUE TO A TYPE IA ENDOLEAK. DESPITE MEDICAL EFFORTS, THE PATIENT'S BLOOD PRESSURE DID NOT STABILIZE, AND TREATMENT WAS DISCONTINUED. PER THE PHYSICIAN THE CAUSE OF THE TYPE IA ENDOLEAK WAS DUE TO THE PATIENTS ANATOMY DUE TO CHANGES IN THE SHAPE OF THE BLOOD VESSELS OVER TIME ULTIMATELY CAUSING THE ANEURYSM TO RUPTURE. THE CAUSE OF DEATH WAS DUE TO ABDOMINAL AORTIC ANEURYSM RUPTURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112368&pc=LYJ | 1644487-2025-00021 | 2024-12-08 05:00 | Death | LIVANOVA USA, INC. | 2025-01-09 05:00 | LYJ | PULSE GEN MODEL 104 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S MOTHER THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. AT THE TIME OF THE REPORT, THE CAUSE OF DEATH HAD NOT BEEN CONFIRMED. THE PATIENT'S VNS WAS EXPLANTED HOWEVER IT HAS NOT BEEN RECEIVED TO DATE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: THE SUSPECTED DEVICE WAS RECEIVED INTO PRODUCT ANALYSIS. THE ANALYSIS OF THE DEVICE HAS NOT BEEN COMPLETED TO DATE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: ANALYSIS OF THE SUSPECT DEVICE WAS COMPLETED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112449&pc=DSQ | 2916596-2025-00141 | 2024-12-19 05:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO A TRAUMATIC FALL AND INTRACRANIAL HEMORRHAGE. PATIENT WAS WALKING INTO THEIR APPOINTMENT WHERE THEY LOST FOOTING AND FELL FORWARD, HITTING FACE/HEAD AND RESULTING IN TRAUMATIC BRAIN BLEEDING. IT WAS NOT DEVICE OR THERAPY RELATED AND WAS DUE TO THE PATIENT FALL. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS EXPLANTED AND WOULD NOT BE RETURNED. THERE WERE NO ISSUES WITH THE OPERATION OF THE DEVICE AT ANY TIME. THERE WERE NO CHANGES TO PATIENT CONDITION OF ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED TO THE ACCIDENT. THERE WERE NO RECENT CHANGES TO PUMP SETTINGS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) LISTS ADVERSE EVENTS, INCLUDING BLEEDING, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS. THE IFU ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112464&pc=DSQ | 2916596-2024-52867 | 2024-02-07 05:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. ADDITIONALLY, DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT PROVIDE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112527&pc=DSQ | 2916596-2024-52963 | 2024-07-02 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112606&pc=BZD | 2518422-2025-001269 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112620&pc=BZD | 2518422-2025-001259 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H6: PATIENT OUTCOME CODE GRID WAS CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112722&pc=DSQ | 2916596-2024-52868 | 2023-08-11 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY CUSTOMER. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. DEVICE HISTORY RECORDS FOR THE POWER MODULE WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QA SPECIFICATIONS. THE POWER MODULE WAS SHIPPED TO THE CUSTOMER ON 17MAY2023. THE INSPECTION DATA FOR THE SEALED LEAD ACID BATTERY (SERIAL NUMBER: (B)(6) WERE REVIEWED AND REVEALED THAT THE BATTERY PASSED. THE SEALED LEAD ACID BATTERY WAS INSPECTED ON 09AUG2022 THROUGH MATERIAL. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112734&pc=PNJ | 9612164-2025-00174 | 2023-01-01 05:00 | Death | MEDTRONIC IRELAND | 2025-01-09 05:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Insufficient Information | P150033 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). THE AUTHORS DESCRIBED ONE PATIENT DEATH IN THE THIRTY DAY POST INTERVENTIONAL PERIOD; HOWEVER, THE CAUSE OF DEATH WAS UNKNOWN. THIS DEATH OCCURRED IN THE GROUP WITHOUT TAVR AND WAS DEEMED UNRELATED TO THE LEADLESS IPG IMPLANTATION. THERE WERE PATIENTS WHO EXPERIENCED PERICARDIAL EFFUSIONS, FEMORAL OR INTRA-ABDOMINAL BLEEDING WITHIN THE THIRTY-DAY POST IMPLANTATION PERIOD. THREE PATIENTS REQUIRED DRAINAGE, BUT NONE REQUIRED CARDIAC SURGERY. ONE OF THE PERICARDIAL EFFUSIONS OCCURRED IN THE TAVR GROUP AND REQUIRED PERICARDIOCENTESIS. THREE PATIENTS HAD FEMORAL BLEEDING; ONE HAD ARTERIAL BLEEDING WITH SUBSEQUENT COVERED STENTING OF THE FEMORAL ARTERY; TWO HAD VENOUS BLEEDINGS THAT WERE MANAGED CONSERVATIVELY. BLEEDING EVENTS WERE TREATED WITH TRANSFUSIONS. ONE OCCURRENCE OF INTRA-ABDOMINAL BLEEDING OF UNKNOWN ORIGIN WAS NOTED IN A PATIENT WITH PERICARDIAL EFFUSION, AND BOTH ISSUES WERE MANAGED CONSERVATIVELY. THERE WERE TWO DEVICES WHICH WERE EXTRACTED DUE TO ELEVATED CAPTURE THRESHOLDS WITH PREMATURE BATTERY DEPLETION AND LOSS OF CAPTURE. BOTH WERE REPLACED WITH A CONVENTIONAL TRANSVENOUS IPG SYSTEM. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/81 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LEADLESS PACEMAKER IMPLANTATION, FOCUSING ON PATIENTS WITH CONDUCTION SYSTEM DISORDERS POST TRANSCATHETER AORTIC VALVE REPLACEMENT: A RETROSPECTIVE ANALYSIS. CJC OPEN. 2024. 6:96-103. DOI: 10.1016/J.CJCO.2023.10.009. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21112758&pc=BZD | 2518422-2025-001324 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113069&pc=NVY | 2017865-2025-10568 | 2024-12-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-09 05:00 | NVY | RIATA ST ACTIVE FIXATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10567. IT WAS REPORT THAT THE PATIENT WAS FOUND ASYSTOLE IN A HOUSE FIRE. THE CAUSE AND MANNER OF DEATH ARE PENDING FURTHER STUDIES. Manufacturer Narrative: A PARTIAL LEAD (ONLY THE CONNECTOR PORTION) WAS RETURNED IN THREE PIECES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113100&pc=BZD | 2518422-2025-001280 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION AND OTHER (UNSPECIFIED INFORMATION). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113106&pc=LWS | 2017865-2025-10567 | 2024-12-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-09 05:00 | LWS | FORTIFY ASSURA VR ICD, US | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10568. IT WAS REPORT THAT THE PATIENT WAS FOUND ASYSTOLE IN A HOUSE FIRE. THE CAUSE AND MANNER OF DEATH ARE PENDING FURTHER STUDIES. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113139&pc=OZD | 1220648-2025-25454 | 2024-12-23 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA 5.5 | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT HAD AN EMERGENT IMPELLA 5.5 IMPLANT DUE TO HAVING A PERICARDIAL EFFUSION. THE IMPELLA WOULD NOT PASS THE BEND INTO THE AORTA. MULTIPLE WIRES WERE ATTEMPTED. THE PATIENT BEGAN DECOMPENSATING AND THE EFFUSION WAS WORSENING. THE PATIENT ARRESTED MOST LIKELY FROM THE TAMPONADE. THE DECISION WAS MADE TO STOP CARDIOPULMONARY RESUSCITATION AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA & RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE PUMP WAS NOT ABLE TO PASS BEND INTO AORTA DUE TO PATIENT'S ANATOMY. H6 CODE 2993 WAS REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT 1220648-2025-25454. NEW CODE WAS ADDED TO MEDICAL DEVICE PROBLEM CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113168&pc=BZD | 2518422-2025-001186 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113226&pc=PNJ | 2017865-2025-10619 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-09 05:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM AN UNKNOWN CAUSE. THERE WAS NO INDICATED THAT THE DEVICE OR DEVICE RELATED PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113464&pc=OZD | 1220648-2025-25460 | 2024-12-24 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE IMPELLA CP WAS INSERTED TO SUPPORT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION / CARDIOGENIC SHOCK. THE PATIENT WAS BLEEDING AROUND THE SHEATH OF THE IMPELLA. THE BLEEDING WAS NOT ABLE TO BE CONTROLLED, AND THE CP WAS EXPLANTED. ADDITIONALLY, CARE WAS WITHDRAWN AND PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING COULD NOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. B.7 ADDED INFORMATION AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25460. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25460 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113471&pc=DSQ | 2916596-2025-00175 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE, "COMPARISON OF CONTINUOUS FLOW CENTRIFUGAL LEFT VENTRICULAR ASSIST DEVICES AS A BRIDGE TO TRANSPLANT STRATEGY IN A LOW ORGAN DONATION ENVIRONMENT: SINGLE CENTER EXPERIENCE", THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH DEATH, INFECTION, RIGHT VENTRICULAR FAILURE (RVF), ARRHYTHMIAS, AND OTHER VARIOUS COMPLICATIONS. THIS RETROSPECTIVE STUDY EVALUATED 68 PATIENTS WITH ADVANCED HEART FAILURE, WHO UNDERWENT CONTINUOUS FLOW LVAD IMPLANT AS BRIDGE TO TRANSPLANT BETWEEN 2009 AND 2023. BASELINE CHARACTERISTICS INCLUDED BLOOD TESTS, DEMOGRAPHICS, ECHOCARDIOGRAPHIC, AND RIGHT HEART CATHETERIZATION DATA BEFORE IMPLANTATION. OUT OF THE 68 PATIENTS, 41 PATIENTS RECEIVED A HEARTWARE (HVAD) PUMP AND 27 PATIENTS RECEIVED A HEARTMATE 3 (HM3) DEVICE. THE HM3 PATIENTS HAD SIGNIFICANTLY BETTER SURVIVAL RATES (P=0.010) AND LOWER COMPLICATION RATES (P=0.025) THAN THE HVAD PATIENTS. BOTH GROUPS WERE DIVIDED INTO ADDITIONAL SUBGROUPS BASED ON WHETHER THEY RECEIVED A HEART TRANSPLANT OR NOT. THE GROUP OF PATIENTS THAT DID NOT RECEIVE A HEART TRANSPLANT (NOHTX) HAD SIGNIFICANTLY HIGHER COMPLICATION RATES COMPARED WITH THE GROUP OF PATIENTS THAT RECEIVED A HEART TRANSPLANT (HTX) (P=0.00041). OF THE 27 HM3 PATIENTS, THERE WERE 2 RECORDED DEATHS AND 14 RECORDED TRANSPLANTS, WHICH WAS COMPARED WITH THE 17 DEATHS AND 21 TRANSPLANTS FROM THE GROUP OF HVAD PATIENTS. MORE PATIENTS IN THE NO HEART TRANSPLANT GROUP WERE ON ECMO THAN THE NUMBER OF PATIENTS IN THE GROUP THAT RECEIVED HEART TRANSPLANTS. WHILE SUPPORTED BY THE LVAD IMPLANT, THERE WERE 10 HM3 PATIENTS AND 21 HVAD PATIENTS ON TEMPORARY INTRA-AORTIC BALLOON PUMP (IABP) SUPPORT, 1 HM3 PATIENT AND 3 HVAD PATIENTS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT, AND 0 HM3 PATIENTS AND 1 HVAD PATIENT ON SUPPORT FROM AN IMPELLA DEVICE. THERE WERE 5 HM3 PATIENTS, COMPARED TO 23 HVAD PATIENTS, THAT EXPERIENCED ADVERSE EVENT COMPLICATIONS WHILE ON LVAD SUPPORT (P=0.0021). THERE WERE 2 HM3 PATIENTS, COMPARED TO 14 HVAD PATIENTS, THAT EXPERIENCED INFECTION WHILE ON LVAD SUPPORT (P=0.011), 3 HM3 PATIENTS EXPERIENCED RVF COMPARED TO 2 HVAD PATIENTS WHILE ON LVAD SUPPORT (P=0.3353), AND 1 HM3 PATIENT EXPERIENCED ARRHYTHMIAS COMPARED TO 0 HVAD PATIENTS WHILE ON LVAD SUPPORT (P=0.2144). THERE WERE NO HM3 PATIENTS THAT EXPERIENCED THE FOLLOWING ADVERSE EFFECTS, WHICH WERE EXPERIENCED BY HVAD PATIENTS: STROKE (P=0.009), GASTROINTESTINAL BLEEDING (P = 0.0594), PUMP THROMBOSIS (P=0.0594), AND DRIVELINE TRANSECTION (P=0.4136). THE STUDY CONCLUDED THAT PATIENTS WITH ADVANCED HEART FAILURE GAINED SUBSTANTIAL BENEFIT FROM LVADS AWAITING HEART TRANSPLANTATION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. INVOLVED HOSPITAL: HEART FAILURE AND TRANSPLANT UNIT, ONASSIS CARDIAC SURGERY CENTRE, 356 SYNGROU AVENUE, KALLITHEA, ATTICA 176 74, GREECE. BONIOS, M., MILIOPOULOS, D., GKOUZIOUTA, A., KOGERAKIS, N., FRAGKOULIS, S., ARMENIS, I., ZARKALIS, D., IEROMONACHOS, K., KOLIOPOULOU, A., LEONTIADIS, E., GEORGIADOU, P., VARTELA, V., TSIAPRAS, D., SFIRAKIS, P., KAPELIOS, C., DIMOPOULOS, S., KAKLAMANIS, L., NTEGIANNIS, D., ANTONIOU, T., ADAMOPOULOS, S. (2024). COMPARISON OF CONTINUOUS FLOW CENTRIFUGAL LEFT VENTRICULAR ASSIST DEVICES AS A BRIDGE TO TRANSPLANT STRATEGY IN A LOW ORGAN DONATION ENVIRONMENT: SINGLE CENTER EXPERIENCE. HELLENIC JOURNAL OF CARDIOLOGY. HTTPS://DOI.ORG/10.1016/J.HJC.2024.10.004 MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113837&pc=OMP | 8043484-2025-00005 | 2024-12-17 05:00 | Death | SMITH & NEPHEW MEDICAL LTD. | 2025-01-09 05:00 | OMP | UNKNOWN RENASYS | Adverse Event Without Identified Device or Use Problem; Device Handling Problem | Sepsis; Post Operative Wound Infection |  |  | 1 | Event Description: IT WAS REPORTED THAT, AFTER A WOUND DEBRIDEMENT PROCEDURE TOOK PLACE ON (B)(6)2024, ON A DIABETIC WOUND UNDERNEATH THE FOOT, ALL WENT WELL WITH THE PROCEDURE AND PATIENT WAS WELL ON THEIR WAY TO RECOVERY. NEVERTHELESS, AROUND THE END OF OCTOBER, BEGINNING NOVEMBER IT WAS SUSPECTED THEIR WOUND GOT INFECTED. ON (B)(6)2024 THEIR LEG WAS WARM AND THERE WAS A BAD SMELL COMING FROM THE FOOT. BY (B)(6)2024 THE PATIENT WAS TRANSPORTED BY AMBULANCE TO THE HOSPITAL DUE TO HIGH FEVER AND IMMENSE PAIN. IT WAS LATER FOUND THEIR FOOT WOUND HAD A PROFUSE GROWTH OF STAPHYLOCOCCUS AUREUS AND CORYNEBACTERIUM SP, THESE TWO CAUSED SEPTICEMIAS. ADDITIONALLY, A VACUUM MACHINE PLACED ON THEIR FOOT FELL OVER IN THE THEATRE AND WOULD NOT WORK ANYMORE AND WAS REPORTED TO BE REPLACED FOUR DAYS LATER. BY THE TIME THE BACTERIA GROWTH WAS UNDER CONTROL, THE PATIENT WAS PUT ON DIALYSIS TWO CONSECUTIVE DAYS BUT DECEASED AT THE AGE OF 57 ON (B)(6)2024.Event Description: IT WAS REPORTED THAT, AFTER A WOUND DEBRIDEMENT PROCEDURE TOOK PLACE ON (B)(6)2024, WITH A NON-S+N PRODUCT, ON A DIABETIC WOUND UNDERNEATH THE FOOT, ALL WENT WELL WITH THE PROCEDURE AND PATIENT WAS WELL ON THEIR WAY TO RECOVERY. NEVERTHELESS, AROUND THE END OF OCTOBER, BEGINNING (B)(6) IT WAS SUSPECTED THEIR WOUND GOT INFECTED. ON (B)(6)2024 THEIR LEG WAS WARM AND THERE WAS A BAD SMELL COMING FROM THE FOOT. BY (B)(6)2024 THEY WERE TRANSPORTED BY AMBULANCE TO THE HOSPITAL DUE TO HIGH FEVER AND IMMENSE PAIN. IT WAS LATER FOUND THEIR FOOT WOUND HAD A PROFUSE GROWTH OF STAPHYLOCOCCUS AUREUS AND CORYNEBACTERIUM SP, THESE TWO CAUSED SEPTICEMIAS. A RENASYS VACUUM MACHINE PLACED ON THEIR FOOT FELL OVER IN THE THEATRE ON (B)(6)2024 AND WOULD NOT WORK ANYMORE AND WAS REPORTED TO BE FIXED FOUR DAYS LATER ON (B)(6)2024. BY THE TIME THE BACTERIA GROWTH WAS UNDER CONTROL, THEY WERE PUT ON DIALYSIS TWO CONSECUTIVE DAYS, BUT THEY DECEASED AT THE AGE OF 57 ON (B)(6)2024. Manufacturer Narrative: CORRECTED DATA: B5, D1, D2A, D2B, H6 (HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE, MEDICAL DEVICE PROBLEM CODE). Manufacturer Narrative: H3,H6: THE DEVICE WAS NOT RETURNED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE PERFORMED. THE CLINICAL/MEDICAL INVESTIGATION CONCLUDED THAT, BASED ON THE INFORMATION PROVIDED, A POST PROCEDURAL COMPLICATION AND/OR THE PROGRESSIVE WORSENING OF THE PATIENT¿S SYSTEMIC CONDITIONS SUCH AS DIABETES, WOUND INFECTION, AND SEPSIS OVER THREE MONTHS MOST LIKELY LED TO THE DIALYSIS TREATMENTS, AND THE PATIENT¿S SUBSEQUENT DEATH. ACCORDING TO THE PROVIDED DEATH CERTIFICATE, THE PATIENT'S CAUSE OF DEATH WAS DUE TO NATURAL CAUSES. ADDITIONALLY, A NON-ADHERENCE TO THE INSTRUCTIONS FOR USE (IFU) COULD NOT BE RULED OUT. AS IT WAS REPORTED THE NEGATIVE PRESSURE WOUND THERAPY (NPWT) WAS INTERRUPTED FOR 4 DAYS AND THE IFU FOR THE RENASYS DEVICE DOES WARN THAT THE ¿NPWT SHOULD REMAIN FOR THE DURATION OF THE TREATMENT.¿ THEREFORE, A CAUSAL RELATIONSHIP BETWEEN THE SMITH AND NEPHEW (S+N) DEVICE, THE REPORTED SEPSIS AND SUBSEQUENT DEATH CANNOT BE ESTABLISHED; THEREFORE, A DEVICE MALFUNCTION IS NOT SUPPORTED. FURTHER IMPACT TO THE PATIENT BEYOND THAT WHICH HAS BEEN REPORTED COULD NOT BE CONCLUDED DUE TO THE PATIENT¿S DEATH UNRELATED TO THE USE OF THE S+N DEVICE. DEVICE SPECIFIC IDENTIFIERS WERE NOT PROVIDED. THEREFORE, AN EVALUATION OF THE MANUFACTURING RECORDS, INSTRUCTIONS FOR USE, SPECIFIC RISK MANAGEMENT FILE, PRIOR ACTIONS AND STERILIZATION RECORDS REVIEW COULD NOT BE PERFORMED. A REVIEW OF COMPLAINT HISTORY FOR THE PREVIOUS 13 MONTHS, DID NOT REVEAL SIMILAR EVENTS FOR "DECEASE" AND DID REVEALED SIMILAR EVENTS FOR "INFECTION" AND "DEVICE FALL", FOR THE DEVICE LISTED, THIS FAILURE MODE AND ADVERSE EVENT WILL BE MONITORED FOR FUTURE COMPLAINTS FOR ANY NECESSARY CORRECTIVE ACTIONS. FACTORS THAT COULD CONTRIBUTE TO THE REPORTED EVENT INCLUDE THE PATIENT'S MEDICAL CONDITION OR A POST PROCEDURAL COMPLICATION, AS IT WAS PREVIOUSLY MENTIONED. AT THIS TIME, WE HAVE NO EVIDENCE TO SUSPECT THAT THE PRODUCT FAILED TO MEET ANY PRODUCT SPECIFICATIONS AT THE TIME OF MANUFACTURE. BASED ON THIS INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. Manufacturer Narrative: INTERNAL COMPLAINT REFERENCE: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113853&pc=OZD | 1220648-2025-25461 | 2024-12-10 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SMART ASSIST SUPPORT FOR 732.63 HOURS BEFORE EXPIRING. THE CAUSE OF DEATH WAS HEMORRHAGIC SHOCK. THE LOCATION OF BLEEDING WAS UNKNOWN, AND THE CAUSE OF THE HEMORRHAGIC SHOCK WAS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE BLEEDING LOCATION IS UNKNOWN AND THERE IS NO OTHER INFORMATION AVAILABLE AS TO HOW THE BLEEDING OCCURRED. THEREFORE, THE ROOT CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113931&pc=DSQ | 2916596-2025-00173 | 2021-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT DISTRESSED COMMUNITY INDEX (DCI) IS A MODEL OF SOCIAL VULNERABILITY THAT ASSIGNS A SCORE FROM 0 (NOT DISTRESSED) TO 100 (SEVERELY DISTRESSED) TO A ZIP CODE BASED ON SEVERAL SOCIOECONOMIC STATUS (SES) DETERMINANTS. DCI WAS USED TO EVALUATE THE IMPACT OF SES ON PREOPERATIVE CHARACTERISTICS AND ONE YEAR POSTOPERATIVE OUTCOMES OF LEFT VENTRICULAR ASSIST DEVICE (LVAD) RECIPIENTS. THIS IS A RETROSPECTIVE COHORT STUDY. PATIENTS WERE STRATIFIED BY THEIR ZIP CODE AS BEING DISTRESSED (DS), DCI > 60, OR NOT DISTRESSED (ND), DCI <60. A TOTAL OF 125 PTS MET CRITERIA FOR INCLUSION- 45 (36%) ND AND 80 (64%) DS LVAD RECIPIENTS. BASELINE DEMOGRAPHICS WERE LARGELY SIMILAR, AND BOTH GROUPS WERE PREDOMINANTLY MALE (87% ND, 78% DS). A GREATER PORTION OF THE DS COHORT IDENTIFIED AS HISPANIC (38% VS 11%, P = 0.007) AND WERE INDICATED AS DESTINATION THERAPY (84% VS 71%, P = 0.03). PREIMPLANTATION, THE DS COHORT HAD A LOWER CARDIAC INDEX (1.7 VS 2 L/ MINUTE/M 2 , P = 0.02). THERE WAS NOT A SIGNIFICANT DIFFERENCE IN THE IMMEDIATE POSTOPERATIVE COURSE- HOURS INTUBATED (22 ND, 24 DS), DAYS IN THE ICU (8 ND, 8 DS), AND LENGTH OF STAY (21 ND, 22 DS). THIRTY DAY (87% ND, 94% DS) AND 1 YEAR (80% ND, 84% DS) SURVIVAL WERE SIMILAR BETWEEN GROUPS, WITH MOST PATIENTS BEING DISCHARGED DIRECTLY HOME (47% ND, 58% DS). THE DS COHORT HAD A HIGHER RATE OF REHOSPITALIZATION IN THE FIRST YEAR (61% VS 43%, P=0.04). THERE WAS NO DIFFERENCE IN BLEEDING EVENTS (33% ND, 41% DS) OR CEREBROVASCULAR ACCIDENTS (11% ND, 12% DS), THOUGH THE DS COHORT WAS MORE LIKELY TO DEVELOP A MAJOR INFECTION (66% VS 47%, P=0.03). WHILE DCI WAS ASSOCIATED WITH SOME DIFFERENCES IN BASELINE CLINICAL PROFILE, IT WAS NOT ASSOCIATED WITH DIFFERENCES IN SURVIVAL IN PATIENTS UNDERGOING LVAD IMPLANTATION. DCI, AS A MARKER OF SES, WAS INDICATIVE OF REHOSPITALIZATIONS AND MAJOR INFECTIONS. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE SAME AS 01JAN2024 AS DATA WAS COLLECTED BETWEEN 2017 AND 2021. AUTHOR CITATION: B.E. FERRELL, D. GOLDSTEIN. J.A. OLIVERA, R. ZHU, I.O. FAITH, S. PATEL. (2024). UTILIZATION OF DISTRESSED COMMUNITIES INDEX TO EXAMINE THE IMPACT OF SOCIOECONOMIC STATUS ON LEFT VENTRICULAR ASSIST DEVICE OUTCOMES. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION 43(4)254. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21113941&pc=OZD | 1220648-2025-25459 | 2024-12-17 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCE A CEREBRAL HEMORRHAGE WITH A DEMISED OUTCOME. THE PATIENT WAS ON BOTH EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND IMPELLA MCS, AND DURING SUPPORT IT WAS NOTED THE PATIENT HAD EXPERIENCED A LARGE CEREBRAL HEMORRHAGE THAT CAUSED A DEVIATION. IT WAS STATED THE PATIENT WAS MAINLY MANAGED WITH ECMO, AND THE ACTIVATED CLOTTING TIME WAS APPROXIMATELY 220. THE PATIENT'S HEART STARTED TO RECOVER; HOWEVER, THE PATIENT EXPIRED APPROXIMATELY THREE DAYS AFTER START OF SUPPORT FROM THE CEREBRAL HEMORRHAGE. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. NO DEVICE WAS RECEIVED FOR EVALUATION. DEVICE HISTORY RECORD REVIEW CONFIRMED THAT THE PUMP PASSED ALL POST INSPECTION CHECKS. AS INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED, THE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114105&pc=BZD | 2518422-2025-001350 | 2024-06-10 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114242&pc=BZD | 2518422-2025-100312 | 2024-12-17 05:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY DUE TO CHRONIC LUNG DISEASE. THE PATIENT'S WIFE DOES NOT BELIEVE THE DEATH WAS RELATED TO THE USE OF THE CPAP DEVICE AND HAS REQUESTED A RETURN LABEL TO RETURN THE REPLACEMENT DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAD PASSED AWAY DUE TO CHRONIC LUNG DISEASE. THE PATIENT'S WIFE DID NOT BELIEVE THE DEATH WAS RELATED TO THE USE OF THE CPAP DEVICE AND HAD REQUESTED A RETURN LABEL TO RETURN THE REPLACEMENT DEVICE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE, AND THE EVALUATION NOTES STATED, "DEVICE DOES NOT DUPLICATE CUSTOMER COMPLAINT." AT THE TIME OF INSPECTION, THE DEVICE RECORDED 2,408.2 TOTAL MACHINE HOURS, WITH ZERO BLOWER HOURS. ADDITIONALLY, IT WAS NOTED THAT THE PORT COVER (ACCESSORY DOOR) WAS MISSING. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H CODING IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114463&pc=DSQ | 2916596-2024-52900 | 2023-04-27 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. IT WAS NOT PROVIDED WHETHER THIS WAS DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WOULD NOT BE RETURNED. PATIENT PRIVACY LAWS WOULD NOT ALLOW THE COMMUNICATION OF PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT'S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF PATIENT DEATH. Manufacturer Narrative: SECTION A1-A4: PATIENT INFORMATION WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114471&pc=OZD | 1220648-2025-25463 | 2024-12-20 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD A BRAIN BLEED. CAROTID ARTERY WAS USED FOR IMPELLA ACCESS, SO IT IS POSSIBLE THE ACCESS SITE WAS CONNECTED TO THE BRAIN BLEED, BUT IT IS DIFFICULT TO KNOW FOR SURE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED STROKE HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE STROKE COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114562&pc=BZD | 2518422-2025-001278 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Asthma; Headache; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Vomiting; Dizziness; Cancer; Unspecified Heart Problem; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA / VOMITING, ASTHMA, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114582&pc=DSQ | 2916596-2025-00005 | 2024-12-17 05:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED AND WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS UNKNOWN. WHETHER THE DEVICE OPERATED AS EXPECTED WAS UNKNOWN. NO DEVICES WILL RETURN. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114684&pc=NPT | 2015691-2025-00186 | 2024-12-24 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-09 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. AFTER VALVE DEPLOYMENT, A POST-DILATION WAS PERFORMED USING SAME VOLUME WITH DELIVERY SYSTEM. THE PATIENT BECAME IMMEDIATELY HEMODYNAMICALLY UNSTABLE AND RESUSCITATION WAS PERFORMED. PERICARDIAL EFFUSION WAS NOTED ON TTE AND ASSUMED TO BE DUE TO AORTIC/ANNULAR RUPTURE. PERICARDIAL DRAIN WAS DONE. RESUSCITATION WAS HALTED AFTER THIRTY MINUTES GIVEN PATIENT AGE AND FRAILTY. THE PATIENT WAS NOT A SURGICAL CANDIDATE. THE PATIENT EXPIRED. THE PERCEIVED ROOT CAUSE OF THIS EVENT WAS HEAVILY CALCIFIED LEAFLETS AND IT WAS ASSUMED THAT THE INJURY OCCURRED WITH FIRST BALLOON INFLATION DURING DEPLOYMENT. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PERICARDIAL EFFUSION WAS NOTED ON TTE AND ASSUMED TO BE DUE TO AORTIC/ANNULAR RUPTURE. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THIS EVENT WAS HEAVILY CALCIFIED LEAFLETS AND IT WAS ASSUMED THAT THE INJURY OCCURRED WITH FIRST BALLOON INFLATION DURING DEPLOYMENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114834&pc=MKJ | 1220908-2024-04813 | 2024-12-19 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-09 05:00 | MKJ | R SERIES | Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO PACE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. COMPLAINANT INDICATED THAT THE CLINICIAN CHANGED PADS TWICE IN AN ATTEMPT TO RESOLVE NO ECG SIGNAL WITHOUT SUCCESS. CLINICIAN THEN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI, THIS DEVICE WAS MANUFACTURED BEFORE UDI REQUIREMENTS. ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCTS AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING WITHOUT DUPLICATING THE REPORT. AN INTERNAL INSPECTION RESULTED IN NO FINDINGS. THE ACCESSORIES USED DURING THE EVENT WERE NOT RETURNED FOR EVALUATION. THERE WAS NO FAULT FOUND WITH THE DEVICE. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. THE DEVICE ACTIVITY LOGS WERE NOT AVAILABLE FOR REVIEW. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21115047&pc=BZD | 2518422-2025-100310 | 2024-12-17 05:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21115425&pc=BZD | 2518422-2025-001288 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21115433&pc=KDI | 3003464075-2025-00024 | 2024-12-15 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-01-09 05:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Air Embolism; Cardiac Arrest | K050525 |  | 1 | Event Description: A REPORT WAS RECEIVED ON (B)(6) 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 62-YEAR-OLD MALE PATIENT APPROVED FOR PERFORMING SOLO HOME HEMODIALYSIS THERAPY, WITH A MEDICAL HISTORY INCLUDING DIABETES, CARDIOVASCULAR ISSUES (NOS) AND END STAGE RENAL DISEASE WHO STATED THE PATIENT EXPIRED AT AN UNSPECIFIED TIME DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 02 JAN 2025 FROM THE HTN WHO STATED THE PATIENT WENT TO USE THE BATHROOM AT AN UNSPECIFIED TIME DURING TREATMENT AND WAS LATER WAS FOUND ON THE FLOOR WITH HIS ARTERIAL LINE DISCONNECTED AND CLAMP TO BOTH THE ARTERIAL LINE AND CATHETER PORT OPEN. PER THE HTN, THE CAUSE OF DEATH WAS AIR EMBOLISM, UNRELATED TO NXSTAGE PRODUCT AND THERAPY. Manufacturer Narrative: THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21115965&pc=PNJ | 9612164-2025-00182 | 2024-01-18 05:00 | Death | MEDTRONIC IRELAND | 2025-01-09 05:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Sepsis | P150033 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE LIFE CYCLE AND POST-MORTEM TISSUE INGROWTH PATTERNS OF A LEADLESS PACEMAKER SYSTEM. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE CLASSIFIED AS CARDIAC OR NON-CARDIAC AND INCLUDED NON-DEVICE RELATED INFECTIONS OR SEPSIS AND OTHER UNKNOWN CAUSES. THERE WERE DEVICES WHICH EXHIBITED INCREASED PACING THRESHOLDS AND THE DEVICES WERE EXTRACTED. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS [OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/79 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LIFE CYCLE AND POST-MORTEM INGROWTH PATTERNS OF A LEADLESS PACEMAKER SYSTEM. EUROPACE. 2024. 26, 1¿4. DOI: 10.1093/EUROPACE/EUAE013 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21116421&pc=NEZ | 9710107-2025-00007 |  | Death | GIVEN IMAGING LTD., YOQNEAM | 2025-01-09 05:00 | NEZ | PILLCAM | Entrapment of Device | Foreign Body In Patient | K211684 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER IT WAS REPORTED THAT THE CAPSULE WAS RETAINED. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118330&pc=BZD | 2518422-2025-100325 | 2024-10-28 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | NO EQUIPMENT NUMBER | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION FROM WIFE OF THE PATIENT, STATING THAT THE PATIENT IS DECEASED, AND CALLER WOULD LIKE TO STOP RECEIVING ALL NOTICES. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118333&pc=BZD | 2518422-2025-100326 | 2024-10-28 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S WIFE, TO SEE WHAT THEY COULD DO FOR THE ELS BECAUSE THEY THREW AWAY THE DEVICE. PATIENT PASSED AWAY AND PATIENT'S WIFE WANTED TO KNOW IF THEY ARE STILL ELIGIBLE. THE CUSTOMER REPRESENTATIVE EXPLAINED SINCE THE DEVICE WAS DISCARDED, IT WOULD BE MARKED AS DISCARDED IN PHILIPS SYSTEM AND THEY ARE NOT ELIGIBLE FOR DEVICE RETURN AWARD. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118352&pc=GDW | 3005075853-2025-00293 | 2024-01-01 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-01-09 05:00 | GDW | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding; Unspecified Tissue Injury | K141952 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN UNKNOWN LIVER PROCEDURE THAT ON THE SECOND FIRING NEAR THE PORTAL VEIN THE DEVICE MISFIRED. WHEN THE DEVICE WAS REMOVED THE TISSUE STARTED BLEEDING PROFUSELY AND PATIENT DIED 2HR LATER. DURING THE PROCEDURE THE PATIENT WAS GIVEN 4 ROUNDS OF MTP. Manufacturer Narrative: (B)(4). DATE SENT: 1/24/2025. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT WERE THE INDICATIONS FOR SURGERY? RIGHT HEPATECTOMY-TUMOR-RIGHT HEPATIC VEIN. WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? A LOT. WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? UNKNOWN. DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? UNKNOWN. DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? NO. DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? NO. WAS THERE ANY TISSUE IN THE JAW DISTAL TO THE CUT LINE? UNKN. WHAT WAS THE EXACT ANATOMICAL STRUCTURE THAT WAS WITHIN THE JAWS OF THE DEVICE AT THE TIME OF FIRING? RIGHT HEPATIC VEIN. WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? UNKNOWN. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? UNKNOWN BUT THEY HAD TO TRANSFUSE THE PATIENT A LOT-TOLD REP THEY USED ALL THE BLOOD IN THE HOSPITAL. WAS A TRANSFUSION REQUIRED? YES. WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? UNKNOWN. WAS THE BLEEDING INTRALUMINAL OR EXTRALUMINAL? UNKNOWN. WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? UNKNOWN. WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? UNKNOWN. WOULD THE SURGEON BE INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING TEAM? ADDITIONAL INFORMATION CAN BE REQUESTED FROM THE SURGEON AT (B)(6) BUT SURGEON IS NOT REQUESTING A MEETING. WHEN SURGEON FINALLY GOT THE LIVER MOBILIZED WHEN LOOKING IVC THERE WAS A 2CM HOLE IN THE HEPATIC VEIN. Manufacturer Narrative: (B)(4). DATE SENT 1/9/2025, D4 BATCH # UNK, B3: ONLY EVENT YEAR KNOWN: 2024. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? WAS THERE ANY TISSUE IN THE JAW DISTAL TO THE CUT LINE? WHAT WAS THE EXACT ANATOMICAL STRUCTURE THAT WAS WITHIN THE JAWS OF THE DEVICE AT THE TIME OF FIRING? WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WAS THE BLEEDING INTRALUMINAL OR EXTRALUMINAL? WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? WOULD THE SURGEON BE INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING TEAM? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118511&pc=LWQ | 3008592544-2025-00005 | 2024-08-21 04:00 | Death | MEDTRONIC, INC. | 2025-01-09 05:00 | LWQ | OPEN PIVOT STANDARD AORTIC HEART VALVE | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Stroke/CVA; Endocarditis; Hemorrhage/Bleeding; Myocardial Infarction; Sepsis; Transient Ischemic Attack; Pseudoaneurysm; Vascular Dissection; Multiple Organ Failure; Pericardial Effusion; Thrombosis/Thrombus; Insufficient Information | P990046 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING DECELLULARIZED AORTIC HOMOGRAFTS VERSUS MECHANICAL FOR AORTIC ROOT REPLACEMENT. THE STUDY POPULATION INCLUDED 289 PATIENTS WITH A MEAN AGE OF 48 YEARS WHO WERE PREDOMINANTLY MALE. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION; 216 PATIENTS WERE IMPLANTED WITH A MEDTRONIC OPEN PIVOT AORTIC MECHANICAL VALVE. PATIENT DEATHS OCCURRED IN THE STUDY POPULATION; THE CAUSES OF DEATH WERE FULMINANT SEPSIS, INTRACRANIAL HEMORRHAGE, PERICARDIAL EFFUSION AND MULTIORGAN FAILURE WITH MYOCARDIAL INFARCTION OF THE ANTEROLATERAL PAPILLARY MUSCLE. AMONG OPEN PIVOT AORTIC MECHANICAL VALVE PATIENTS, ADVERSE EVENTS INCLUDED: BLEEDING (CEREBRAL AND PERIPHERAL), VALVE THROMBOSIS, CEREBRAL EMBOLISM, STROKE, TRANSIENT ISCHEMIC ATTACK (TIA), ENDOCARDITIS, PACEMAKER IMPLANT, AORTIC ROOT DISSECTION, PSEUDOANEURYSM AND RE-INTERVENTION OR REPLACEMENT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: AUTHORS: ANDREEVA ET AL.. DECELLULARIZED AORTIC HOMOGRAFTS VERSUS MECHANICAL COMPOSITE GRAFTS FOR AORTIC ROOT REPLACEMENT. EUROPEAN JOURNAL OF CARDIO-THORACIC SURGERY 66(2) 2024. 10.1093/EJCTS/EZAE314 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. MEDTRONIC PRODUCTS REFERENCED: OPEN PIVOT MECHANICAL HEART VALVE (PMA# P990046, PRODUCT CODE LWQ). EARLIEST APPROVED PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118518&pc=DSQ | 2916596-2024-52932 | 2024-10-06 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH AND WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WERE NOT PROVIDED BY THE CUSTOMER. THE DEVICE WAS NOT EXPLANTED.Event Description: DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118530&pc=DSQ | 2916596-2025-00093 | 2024-10-04 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE TO HEMORRHAGIC STROKE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY CUSTOMER. AN AUTOPSY WOULD NOT BE PERFORMED AND THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118678&pc=NPT | 2025587-2025-00249 | 2024-12-30 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-09 05:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem | Vascular Dissection; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE INVOLVING THE EVFXPLUS-26 VALVE, A PRE-IMPLANT BALLOON VALVULOPLASTY (BAV) WAS PERFORMED AS STANDARD PRACTICE FOR THE IMPLANTING PHYSICIAN. SUBSEQUENTLY, THE VALVE WAS IMPLANTED IN THE PATIENT. FOLLOWING VALVE PLACEMENT, PARAVALVULAR LEAK WAS OBSERVED. A POST-IMPLANT BAV WAS PERFORMED USING A 22 MM NON-MEDTRONIC BALLOON. THE VALVE DISLODGED FROM A DEPTH OF 4 MM ON THE LEFT CORONARY CUSP (LCC) AND 4 MM ON THE NON-CORONARY CUSP (NCC) TO A NEW DEPTH OF APPROXIMATELY 6 MM ABOVE THE ANNULUS ON BOTH THE LCC AND NCC DURING THE POST-IMPLANT BAV. PER THE PHYSICIAN, THE PACING CAPTURE WAS LOST AS THE BALLOON WAS INFLATED CAUSING THE RESULTING CONTRACTIONS CAUSED THE FULLY EXPANDED BALLOON TO COME INTO CONTACT WITH THE VALVE FRAME AND CAUSE THE DISLODGEMENT. THE VALVE WAS SNARED INTO THE ASCENDING AORTA PROXIMAL TO THE INNOMINATE ARTERY. A SECOND VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM (DCS) AND INSERTED INTO THE PATIENT OVER SMALL CURVE NON-MEDTRONIC GUIDEWIRE. DURING ADVANCEMENT, THE DCS WAS UNABLE TO SUCCESSFULLY PASS THROUGH THE FIRST DISLODGED VALVE. THE DECISION WAS MADE TO WITHDRAW THE SECOND VALVE AND DELIVERY SYSTEM FROM THE BODY. A 23 MM NON-MEDTRONIC VALVE WAS IMPLANTED IN THE PATIENT INSTEAD. THE PROCEDURE WAS DELAYED BY APPROXIMATELY 45-MINUTES. FOLLOWING THE PROCEDURE, AN INJURY TO THE ASCENDING AORTA WAS SUSPECTED. COMPUTED TOMOGRAPHY CONFIRMED THE INJURY, AND THE PATIENT WAS TAKEN TO SURGERY IN AN ATTEMPT TO REPAIR THE DAMAGED ASCENDING AORTA. THE PATIENT WAS REPORTED TO HAVE PASSED AWAY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER. PER THE PHYSICIAN, THE AORTIC DISSECTION CAUSED THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE SEVERITY OF THE PVL FOLLOWING VALVE DEPLOYMENT WAS MILD. DURING THE SURGERY, THE VALVE WAS EXPLANTED. Manufacturer Narrative: UPDATED DATA: B5. D6B. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118684&pc=NPT | 2025587-2025-00250 | 2024-12-30 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-09 05:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE INVOLVING THE EVFXPLUS-26 VALVE, A PRE-IMPLANT BALLOON VALVULOPLASTY (BAV) WAS PERFORMED AS STANDARD PRACTICE FOR THE IMPLANTING PHYSICIAN. SUBSEQUENTLY, THE VALVE WAS IMPLANTED IN THE PATIENT. FOLLOWING VALVE PLACEMENT, PARAVALVULAR LEAK WAS OBSERVED. A POST-IMPLANT BAV WAS PERFORMED USING A 22 MM NON-MEDTRONIC BALLOON. THE VALVE DISLODGED FROM A DEPTH OF 4 MM ON THE LEFT CORONARY CUSP (LCC) AND 4 MM ON THE NON-CORONARY CUSP (NCC) TO A NEW DEPTH OF APPROXIMATELY 6 MM ABOVE THE ANNULUS ON BOTH THE LCC AND NCC DURING THE POST-IMPLANT BAV. PER THE PHYSICIAN, THE PACING CAPTURE WAS LOST AS THE BALLOON WAS INFLATED CAUSING THE RESULTING CONTRACTIONS CAUSED THE FULLY EXPANDED BALLOON TO COME INTO CONTACT WITH THE VALVE FRAME AND CAUSE THE DISLODGEMENT. THE VALVE WAS SNARED INTO THE ASCENDING AORTA PROXIMAL TO THE INNOMINATE ARTERY. A SECOND VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM (DCS) AND INSERTED INTO THE PATIENT OVER SMALL CURVE NON-MEDTRONIC GUIDEWIRE. DURING ADVANCEMENT, THE DCS WAS UNABLE TO SUCCESSFULLY PASS THROUGH THE FIRST DISLODGED VALVE. THE DECISION WAS MADE TO WITHDRAW THE SECOND VALVE AND DELIVERY SYSTEM FROM THE BODY. A 23 MM NON-MEDTRONIC VALVE WAS IMPLANTED IN THE PATIENT INSTEAD. THE PROCEDURE WAS DELAYED BY APPROXIMATELY 45-MINUTES. FOLLOWING THE PROCEDURE, AN INJURY TO THE ASCENDING AORTA WAS SUSPECTED. COMPUTED TOMOGRAPHY CONFIRMED THE INJURY, AND THE PATIENT WAS TAKEN TO SURGERY IN AN ATTEMPT TO REPAIR THE DAMAGED ASCENDING AORTA. THE PATIENT WAS REPORTED TO HAVE PASSED AWAY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER. PER THE PHYSICIAN, THE AORTIC DISSECTION CAUSED THE DEATH. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: EVFXPLUS-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: EVFXPLUS-26 (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118704&pc=DSQ | 2916596-2024-52915 | 2024-05-11 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS UNKNOWN. IT WAS NOT PROVIDED WHETHER THIS WAS DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: THE PATIENT PRIVACY LAWS OF THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118838&pc=BZD | 2518422-2025-000809 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF CANCER AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21119065&pc=BZD | 2518422-2025-001421 | 2024-02-08 05:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION AUTO BIPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21119112&pc=BZD | 2518422-2025-001420 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-09 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21119141&pc=DSQ | 2916596-2024-52916 | 2024-06-06 04:00 | Death | THORATEC CORPORATION | 2025-01-09 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE DEATH WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: THE PATIENT PRIVACY LAWS OF THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: IT WAS REPORTED THAT THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: A4 PATIENT WEIGHT WAS NOT PROVIDED NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21120098&pc=FDT | 2429304-2025-00024 |  | Death | AIZU OLYMPUS CO., LTD. | 2025-01-09 05:00 | FDT | EVIS EXERA III DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Drug Resistant Bacterial Infection |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING AN ERCP (ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY) PROCEDURE IN (B)(6) 2023, THE PATIENT EXPERIENCED A VRE (VANCOMYCIN RESISTANT ENTEROCOCCI) INFECTION AND LATER PASSED AWAY. NO ADDITIONAL INFORMATION WAS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21120145&pc=FDT | 9610595-2025-00229 |  | Death | AIZU OLYMPUS CO., LTD. | 2025-01-09 05:00 | FDT | EVIS EXERA III DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Drug Resistant Bacterial Infection | K220587 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING AN ERCP (ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY) PROCEDURE IN (B)(6) 2023, THE PATIENT EXPERIENCED A VRE (VANCOMYCIN RESISTANT ENTEROCOCCI) INFECTION AND LATER PASSED AWAY. NO ADDITIONAL INFORMATION WAS AVAILABLE.Event Description: NO NEW INFORMATION RECEIVED FROM THE CUSTOMER. Manufacturer Narrative: THIS REPORT HAS BEEN SUBMITTED BY THE IMPORTER UNDER THIS MDR REPORT NUMBER (B)(4). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE RESULTS OF THE LEGAL MANUFACTURER'S FINAL INVESTIGATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION; THEREFORE, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21120415&pc=NPT | 2025587-2025-00252 | 2024-12-01 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-09 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A PRE-EXISTING SURGICAL AORTIC VA LVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: B3: DATE IS APPROXIMATE. YEAR IS CONFIRMED VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21120728&pc=OZD | 1220648-2025-25465 | 2024-12-16 05:00 | Death | ABIOMED, INC. | 2025-01-09 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT SUFFERED A CEREBRAL HEMORRHAGE. THE PATIENT ULTIMATELY EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED STROKE HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE STROKE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121055&pc=PJF | 3003925919-2024-00002 | 2024-11-27 05:00 | Death | CERUS CORPORATION | 2025-01-09 05:00 | PJF | INTERCEPT BLOOD SYSTEM FOR PLATELETS | Contamination /Decontamination Problem | Sepsis | BP140143 |  | 1 | Event Description: PROBABLE TRANSFUSION TRANSMITTED BACTERIAL INFECTION (TTBI). INITIAL INFORMATION RECEIVED: 10-DEC-2024, 11-DEC-2024, 12-DEC-2024, 17-DEC-2024. INITIAL ADVERSE EVENT REPORT: ON 10-DEC-2024, CERUS RECEIVED A FATAL ADVERSE EVENT REPORT OF PROBABLE TRANSFUSION TRANSMITTED BACTERIAL INFECTION (TTBI) WITH STAPHYLOCOCCUS UREOLYTICUS. THIS ADVERSE EVENT WAS REPORTED TO CERUS BY DR. (B)(6), MEDICAL DIRECTOR FROM (B)(6). MEDICAL HISTORY AND MEDICATIONS: THIS REPORT INVOLVED A 76-YEAR-OLD, MALE OF AN UNREPORTED RACE. THE PATIENT'S MEDICAL HISTORY INCLUDED ACUTE MYELOID LEUKEMIA AND MYELODYSPLASTIC SYNDROME. THE PATIENT'S PAST MEDICAL PROCEDURES WERE NOT REPORTED. CONCOMITANT MEDICATIONS WERE NOT REPORTED. ADVERSE EVENT DESCRIPTION: ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED AT (B)(6) HOSPITAL, FOR TRANSFUSION SUPPORT. UPON ARRIVAL, THE PATIENT WAS AFEBRILE, WITH NO CLINICAL SIGNS AND NO IMPLANTABLE PORT. ON THE SAME DAY, IN THE MORNING, THE PATIENT WAS TRANSFUSED WITH TWO UNITS OF RED BLOOD CELLS WITHOUT INCIDENT. ON THE SAME DAY, FROM 15:05H TO 15:40H, THE PATIENT WAS TRANSFUSED WITH ONE UNIT OF DAY 4 PSORALEN TREATED PLATELETS (PTP) (BLOOD COMPONENT DETAILS NOT REPORTED) THROUGH A PERIPHERAL VEIN CONNECTED TO A 3-WAY TAP/STOP COCK. AFTER TRANSFUSION WAS COMPLETED, THE IMPLICATED PTP WAS DISCONNECTED BY AN ACCREDITED NURSE, CLOSING THE TUBING WITH A KNOT AND CLAMP PLACED ON THE DISTAL END BEFORE UNHOOKING THE POCKET FROM THE STEM. PER REPORTER, THE "THE POCKET TUBING OF THE IMPLICATED PTP NEVER CAME INTO CONTACT WITH THE PATIENT." CLARIFICATION REGARDING THE PURPOSE OF THE "POCKET TUBING" IS PENDING. THE PATIENT'S PRE-TRANSFUSION / DURING TRANSFUSION / POST-TRANSFUSION VITAL SIGNS WERE: BODY TEMPERATURE: NOT REPORTED / NOT REPORTED / 38.7 °C (ONE HOUR AFTER TRANSFUSION) AND 39.4 °C (TWO HOURS AFTER TRANSFUSION), RESPIRATORY RATE: NOT REPORTED, BLOOD PRESSURE: NOT REPORTED, HEART RATE: NOT REPORTED. AROUND 20 MINUTES AFTER TRANSFUSION, THE PATIENT DEVELOPED CHILLS AND SLIGHT BRADYCARDIA. ONE HOUR AFTER TRANSFUSION, THE PATIENT BECAME FEBRILE WITH A TEMPERATURE OF 38.7°C, WHICH ROSE TO 39.4°C ONE HOUR LATER. AN INFECTION ASSESSMENT WAS CONDUCTED, CLAFORAN ANTIBIOTIC STARTED EMPIRICALLY, AND THE PATIENT WAS HOSPITALIZED FOR MONITORING. UNFORTUNATELY, HE DETERIORATED, BECAME HYPOTENSIVE, AND DIED BETWEEN THE NIGHT OF (B)(6) 2024 AND THE MORNING OF (B)(6) 2024. PRE-TRANSFUSION SAMPLE WAS NOT TESTED. ON (B)(6) 2024 AT 17:00H AND AT 21:00H, POST-TRANSFUSION BLOOD SAMPLES WERE COLLECTED FROM THE PATIENT AND SENT FOR GRAM STAIN AND BLOOD CULTURE. GRAM POSITIVE COCCI IDENTIFIED ON GRAM STAIN AND COAGULASE-NEGATIVE STAPHYLOCOCCI DETECTED IN BLOOD CULTURE. STAPHYLOCOCCUS UREILYTICUS WAS IDENTIFIED IN 8 VIALS USING MALDI TOF. IMPLICATED PTP GRAM STAIN SHOWED GRAM POSITIVE COCCI AND CULTURED POSITIVE FOR STAPHYLOCOCCUS UREILYTICUS. NO ANTIBIOTIC RESISTANCE IDENTIFIED. THE PATIENT EXPERIENCED A FATAL ADVERSE EVENT OF PROBABLE TTBI WITH STAPHYLOCOCCUS UREILYTICUS. BLOOD DONOR(S) INFORMATION: THE DONOR IS A REPEAT FEMALE DONOR AND A REVIEW OF HER PRIOR DONATIONS DURING 2024 DID NOT INDICATE ANY OTHER TRANSFUSION RECIPIENT ADVERSE EVENTS. BETADINE AND ALCOHOL SCRUB X4 WAS USED TO DISINFECT THE VENIPUNCTURE AREA. ATTEMPT TO SCHEDULE THE DONOR FOR SWAB COLLECTION AND CULTURE ON (B)(6) 2024 HAS BEEN DONE BUT THE DONOR HAS NOT CONFIRM THIS APPOINTMENT. THE BIOLOGICAL QUALIFICATIONS OF DONATION TUBES (I.E., PRE-DONATION TESTING TUBES) COULD NOT BE CULTURED BECAUSE THEY HAVE BEEN OPEN. IMPLICATED PATHOGEN REDUCED BLOOD COMPONENT(S) COLLECTION AND MANUFACTURING INFORMATION: (B)(6) 2024, (B)(6) CENTER COLLECTED A TRIMA APHERESIS PLATELET SUSPENDED IN PLATELET ADDITIVE SOLUTION (PAS) WITH A CO-COMPONENT PLASMA. ON (B)(6) 2024, THE IMPLICATED PTP WAS TREATED WITH INTERCEPT PROCESSING SET FOR PLATELETS (LOT # CE23I01A01) AND WAS ILLUMINATED WITH INTERCEPT ILLUMINATOR (SERIAL # (B)(6) AT THE EFS RUNGIS BLOOD CENTER. THE TREATMENT REPORT SHOWED SUCCESSFUL ILLUMINATION WITH NO INDICATION OF ILLUMINATOR MALFUNCTION OR IRREGULARITIES. RESIDUAL AMOSTOSALEN WAS 0.25 M AND PH WAS 7.07. THE APHERESIS COLLECTION KIT HAS BEEN DISCARDED. ALL BUT THE FINAL INTERCEPT PROCESSING BAGS HAVE BEEN DISCARDED. THE SPINNING INDEX (I.E. SWIRLING) WAS ACCEPTABLE AND NO AGGREGATE WERE OBSERVED AT THE TIME OF ISSUANCE BY THE EFS BLOOD CENTER ON THE DAY OF TRANSFUSION. THE IMPLICATED PTP WAS TRANSPORTED FROM EFS TO THE CLINICAL SERVICE BY PNEUMATIC TRANSPORT SYSTEM. THE CARTRIDGE THAT WAS USED FOR TRANSPORT WAS CULTURED AND THE RESULT IS NEGATIVE FOR ORGANISMS. CO-COMPONENT(S) INFORMATION: ON (B)(6) 2024, THE CO-COMPONENT PLASMA WAS CULTURED WITH NEGATIVE RESULTS AT DAY 5. ADVERSE EVENT INVESTIGATIONS: (B)(6) HOSPITAL: GRAM STAIN OF THE IMPLICATED PTP POCKET TUBING FOUND GRAM POSITIVE COCCI. GENOTYPE ANALYSIS OF POSITIVE CULTURE SAMPLES FROM THE PATIENT AND IMPLICATED PTP IS IN PROGRESS. ACCORDING TO CLINICAL MICROBIOLOGY, THIS ORGANISM HAS NEVER BEEN IDENTIFIED IN CLINICAL SAMPLES AT THIS HOSPITAL IN THE PAST. STAPHYLOCOCCUS UREILYTICUS IS AN ENVIRONMENTAL STRAIN/BIOFILM STAPHYLOCOCCUS INVOLVED IN NOSOCOMIAL INFECTIONS. RETROGRADE CONTAMINATION WAS RULED OUT DUE TO POSITIVE DIRECT EXAMINATION OF THE IMPLICATED PTP BAG, ARGUING IN FAVOR OF CONTAMINATION OF THE IMPLICATED PTP. NEGATIVE CULTURE RESULT FROM THE CO-COMPONENT PLASMA IS NON-INFORMATIVE BECAUSE HUMAN PLASMA MAY CONTAIN COMPOUNDS THAT INHIBIT THE MULTIPLICATION OF BACTERIA, LEADING TO LOSS OF VIABILITY. SINCE THIS STRAIN OF STAPHYLOCOCCUS UREILYTICUS DOES NOT HAVE ANY ANTIBIOTIC RESISTANCE TRAIT, ADDITIONAL STUDIES REGARDING ANTIBIOTIC SENSITIVITY WILL BE CHALLENGING. (B)(6) CENTER INVESTIGATION : VISUAL INSPECTION DID NOT IDENTIFY ANY OBVIOUS DAMAGE. THE LEAK TEST WAS DONE BY IMMERSING THE IMPLICATED PTP BAG IN WATER, COMPRESSING IT WITH AN AIR COMPRESSOR AT A PRESSURE SETTING OF 1.8 BAR. THERE WERE NO BUBBLES OBSERVED, AND THE OUTCOME OF THIS TEST WAS NEGATIVE. EFS STAFF WHO WAS INVOLVED IN THE COLLECTION OF THE IMPLICATED PTP HAVE BEEN SWABBED FOR CULTURE, RESULTS ARE PENDING. CERUS INVESTIGATION: A BATCH RECORD REVIEW FOR INTERCEPT PROCESSING SET FOR PLATELETS (PRODUCT CODE #INT2204B/LOT #CE23I01A01); EXPIRATION DATE: 30-NOV-2025) WAS PERFORMED. THE REVIEW OF THE BATCH RECORDS FOR THE FINISHED GOODS DID NOT INDICATE ANY OUT-OF-LIMITS RESULTS DURING TESTING AND CONTROLS. THE PRODUCT MET ALL QUALITY REQUIREMENTS AND SPECIFICATIONS AT THE TIME OF RELEASE. THIS BATCH WAS RELEASED ON 21-DEC-2023. NO OTHER COMPLAINT HAS BEEN RECEIVED FOR LOT #CE23I01A01. THE TOTAL BATCH SIZE WAS 6096 PROCESSING SETS. CERUS WILL WORK WITH (B)(6) TO SHIP BACTERIAL ISOLATES FROM THE PATIENT AND IMPLICATED PTP TO AN ACCREDITED LABORATORY FOR ADDITIONAL INVESTIGATIONS. FRESENIUS KABI LA CHÂTRE : CERUS WILL WORK WITH EFS TO SHIP IMPLICATED PTP BAG TO FRESENIUS KABI LA CHÂTRE FOR ADDITIONAL INVESTIGATIONS. ADVERSE EVENT ASSESSMENT: 1.~PROBABLE TTBI WITH STAPHYLOCOCCUS UREILYTICUS [PT: TRANSMISSION OF AN INFECTIOUS AGENT VIA PRODUCT] THE REPORTER ASSESSED THE EVENT AS DEATH IN SEVERITY, SERIOUS DUE TO BEING DEATH, PROBABLE TO CERTAIN IN RELATION TO PSORALEN TREATED PLATELETS AND PROBABLE TO CERTAIN IN RELATION TO INTERCEPT PROCESSING SET FOR PLATELETS. THE CERUS MEDICAL REVIEWER CONSIDERED THE EVENT AS SERIOUS DUE TO DEATH, PROBABLE IN RELATION TO PSORALEN TREATED PLATELETS AND INTERCEPT PROCESSING SET FOR PLATELETS. STAPHYLOCOCCUS UREILYTICUS WAS IDENTIFIED IN BOTH THE PATIENT AND IMPLICATED PTP AND GIVEN THE PATIENT'S CLINICAL PRESENTATION IT IS CONSISTENT WITH A SEPTIC TRANSFUSION REACTION. THAT SAID, IT IS UNCLEAR AT THIS TIME IF THE CONTAMINATION OCCURRED PRE- OR POST-INTERCEPT PATHOGEN INACTIVATION IN LIGHT OF THE FACT THAT A 3-WAY STOP COCKED WAS USED FOR THE TRANSFUSION WHICH COULD HAVE INTRODUCED CONTAMINATION INTO THE IMPLICATED PTP. ADDITIONAL INVESTIGATION INTO THE COLLECTION STAFF, HENRI MONDOR'S PLATELET TRANSFUSION PROCEDURES, PATIENT AND IMPLICATED PTP BACTERIAL GENETIC ANALYSIS ARE ONGOING. MOREOVER, INVESTIGATION INTO THE IMPLICATED PTP BAG INTEGRITY MUST BE COMPLETED TO RULE OUT ANY LEAKS OR BREACHES. FINALLY, CERUS WILL WORK TO OBTAIN THE PATIENT AND IMPLICATED PTP CULTURE ISOLATES TO HAVE PATHOGEN INACTIVATION STUDIES PERFORMED SINCE STAPHYLOCOCCUS UREILYTICUS HAS NOT BEEN EVALUATED IN ANY IN VITRO PATHOGEN INACTIVATION STUDIES IN THE PAST. THE CERUS MEDICAL REVIEWER CONSIDERED THE LISTEDNESS OF THIS EVENT IN RELATION TO THE PSORALEN TREATED PLATELETS AS LISTED BASED ON APPLICABLE CERUS REFERENCE SAFETY INFORMATION. THE CERUS MEDICAL REVIEWER CONSIDERED THE LISTEDNESS OF THIS EVENT IN RELATION TO INTERCEPT PROCESSING SET FOR PLATELETS AS LISTED BASED ON APPLICABLE CERUS REFERENCE SAFETY INFORMATION.Event Description: CERUS RECEIVED ADDITIONAL FOLLOW UP INFORMATION ON THIS EVENT ON 20-DEC-2024, 25-DEC-2024, 31-DEC-2024, 02-JAN-2025, 07-JAN-2025, 10-JAN-2025, 17-JAN-2025, 26-JAN-2025, AND 28-MAY-2025. ADVERSE EVENT DESCRIPTION: (REVISED INFORMATION): ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED AT (B)(6) HOSPITAL, (B)(6) FOR TRANSFUSION SUPPORT. UPON ARRIVAL, THE PATIENT WAS AFEBRILE, WITH NO CLINICAL SIGNS AND NO IMPLANTABLE PORT. ON THE SAME DAY, IN THE MORNING, THE PATIENT WAS TRANSFUSED WITH TWO UNITS OF RED BLOOD CELLS WITHOUT INCIDENT. ON THE SAME DAY, FROM 15:05H TO 15:40H, THE PATIENT WAS TRANSFUSED WITH ONE UNIT OF DAY 4 PSORALEN TREATED PLATELETS (PTP) (DONATION IDENTIFICATION NUMBER (DIN): (B)(4); EXPIRATION DATE: 30-NOV-2024) THROUGH A PERIPHERAL VEIN CONNECTED TO A 3-WAY TAP/STOP COCK, AS PART OF ROUTINE TRANSFUSION PRACTICE. AFTER TRANSFUSION WAS COMPLETED, AN ACCREDITED NURSE CLOSED THE TUBING OF THE IMPLICATED PTP WITH A KNOT AND CLAMP PLACED ON THE DISTAL END. BLOOD DONOR(S) INFORMATION: (NEW INFORMATION) SWAB SAMPLES TAKEN FROM THE DONOR'S NOSE, ARMPIT AND GROIN WERE CULTURED AND THE RESULTS WERE NEGATIVE. IMPLICATED PATHOGEN REDUCED BLOOD COMPONENT(S) COLLECTION AND MANUFACTURING INFORMATION: (NEW INFORMATION) ALL PREVENTIVE MAINTENANCE (PM) ON THE ILLUMINATOR WERE COMPLETED SUCCESSFULLY WITHIN THE LAST SEVERAL YEARS AND NO ISSUES WERE IDENTIFIED DURING THESE MAINTENANCE VISITS. THE LAST PM WAS SUCCESSFULLY COMPLETED ON (B)(6) 2024. (NEW AND UPDATED INFORMATION): ON (B)(6) 2024 AT AROUND 14:00H, THE IMPLICATED PTP'S SPINNING INDEX (I.E. SWIRLING) WAS ACCEPTABLE AND NO AGGREGATES WERE OBSERVED. THE IMPLICATED PTP WAS SENT TO THE HOSPITAL AROUND 14:44H VIA THE PNEUMATIC TRANSPORT SYSTEM. TRANSFUSION NURSE RECEIVED THE IMPLICATED PTP AROUND 15:00H. TRANSFUSION STARTED AT 15:05H AND ENDED AT 15:40H. THE PATIENT STARTED EXHIBITING SYMPTOMS AT 16:00H. PATIENT BLOOD SAMPLES WERE DRAWN AT 17:00H. IMPLICATED PTP WAS SENT BACK TO EFS AT 18:00H FOR INVESTIGATION. THE IMPLICATED PTP WAS STORED IN THE REFRIGERATOR UNTIL IT WAS GIVEN TO THE BACTERIOLOGY LAB ON (B)(6) 2024 AROUND 11:30H. THE CARTRIDGE THAT WAS USED FOR PNEUMATIC TRANSPORT WAS CULTURED AND THE RESULT WAS NEGATIVE FOR STAPHYLOCOCCUS UREILYTICUS. ADVERSE EVENT INVESTIGATIONS: (B)(6) HOSPITAL: (UPDATED INFORMATION) TWO SAMPLES WERE TAKEN FROM THE IMPLICATED PTP, ONE THROUGH A PORT AND ANOTHER THROUGH THE POCKET TUBING (I.E., SEGMENT), BOTH WERE IN DIRECT CONNECTION WITH THE INTERIOR OF THE BAG. BOTH SAMPLES SHOWED GRAM POSITIVE COCCI AND CULTURED POSITIVE FOR STAPHYLOCOCCUS UREILYTICUS. (NEW INFORMATION) GENOTYPE ANALYSIS OF STAPHYLOCOCCUS UREILYTICUS CULTURE ISOLATES FROM THE PATIENT AND IMPLICATED PTP SHOWED STRONG GENETICALLY LINKAGE. EFS BLOOD CENTER INVESTIGATION: (UPDATED INFORMATION) SWAB SAMPLES TAKEN FROM THE NOSE, ARMPIT AND GROIN OF THE EFS STAFF WHO COLLECTED THE IMPLICATED PTP WERE CULTURED, AND THE RESULTS WERE NEGATIVE. CERUS INVESTIGATION: PATHOGEN INACTIVATION INVESTIGATION: (NEW INFORMATION) BACTERIAL ISOLATES FROM THE PATIENT AND IMPLICATED PTP HAVE BEEN SHIPPED TO THE CERUS US HEADQUARTERS FOR PATHOGEN INACTIVATION STUDIES WHICH WERE COMPLETED ON (B)(6) 2025. ALL STUDIES WERE PERFORMED WITH 4 REPLICATES USING APHERESIS PLATELET UNITS ADJUSTED TO ~319 ML WITH 35% PLASMA/65% SSP+ TO MIMIC THE CHARACTERISTICS OF THE IMPLICATED PTP. THE GROWTH CHARACTERISTICS OF STAPHYLOCOCCUS UREILYTICUS IN PLATELETS WERE ASSESSED BY INOCULATING EACH PLATELET UNIT WITH A TARGET 10-25 CFU OF STAPHYLOCOCCUS UREILYTICUS AND INCUBATING UNDER STANDARD STORAGE CONDITIONS (22°C WITH AGITATION). BACTERIAL TITER WAS ASSESSED ON DAY 1 (~24 HOURS POST-INOCULATION), DAY 2 (~48 HOURS POST-INOCULATION), DAY 4, AND DAY 7. STAPHYLOCOCCUS UREILYTICUS GREW TO 0.8, 5.1, 8.4, AND 8.6 LOG CFU/ML, RESPECTIVELY. STAPHYLOCOCCUS. UREILYTICUS WAS GROWN TO A HIGH TITER AND SPIKED INTO EACH PLATELET COMPONENT AT ~7.7 LOG CFU/ML AND TREATED WITH THE INTERCEPT PLATELET LARGE VOLUME PROCESSING SET WITH APPROXIMATELY 150 ¿M AMOTOSALEN AND 3.9 J/CM2 UVA LIGHT WITH THE INT100 ILLUMINATOR. AT POST-ILLUMINATION, GREATER THAN OR EQUAL TO 7.7 LOG CFU/ML OF STAPHYLOCOCCUS. UREILYTICUS WAS INACTIVATED. A SECOND ASSESSMENT WAS PERFORMED WITH APPROXIMATELY 6.7 LOG CFU/ML OF STAPHYLOCOCCUS UREILYTICUS SPIKED INTO EACH PLATELET COMPONENT AND TREATED WITH THE INTERCEPT PLATELET LARGE VOLUME PROCESSING SET WITH 150 ¿M AMOTOSALEN AND 3.9 J/CM2 UVA LIGHT WITH THE INT100 ILLUMINATOR. AT POST-ILLUMINATION, >6.7 LOG CFU/ML OF STAPHYLOCOCCUS UREILYTICUS WAS INACTIVATED WITH NO RESIDUAL VIABLE BACTERIA OBSERVED UP TO DAY 7 POST-TREATMENT. IN CONCLUSION, THE INTERCEPT BLOOD SYSTEM FOR PLATELETS CAN ROBUSTLY INACTIVATE STAPHYLOCOCCUS UREILYTICUS AT 6.7 LOG CFU/ML WITH NO VIABLE BACTERIA OBSERVED AFTER STORAGE ON DAY 7 POST-TREATMENT. SINCE STAPHYLOCOCCUS UREILYTICUS GREW ONLY TO 5.1 LOG CFU/ML BY DAY 2 (~48 HOURS POST-INOCULATION) WHICH IS LOWER THAN THE CAPACITY OF THE SYSTEM (6.7 LOG CFU/ML), IT IS HIGHLY UNLIKELY THAT THE INTERCEPT BLOOD SYSTEM FOR PLATELETS COULD NOT INACTIVATE STAPHYLOCOCCUS UREILYTICUS, EVEN IF THE IMPLICATED PTP WAS TREATED AT THE END OF THE ACCEPTABLE TREATMENT TIME (END OF DAY 1). TAKEN TOGETHER, THE EVIDENCE SUGGESTS THAT CONTAMINATION LIKELY OCCURRED POST-ILLUMINATION. FRESENIUS KABI LA CHÂTRE INVESTIGATION: (NEW INFORMATION) THE IMPLICATED PTP BAG WAS SHIPPED TO FRESENIUS KABI LA CHÂTRE FOR ADDITIONAL INVESTIGATION AND THE FINDINGS WERE: ·~AN EMPTY STORAGE CONTAINER OF THE INTERCEPT PROCESSING SET FOR PLATELET WITH TWO PERFORATORS ATTACHED TO BOTH OUTLET TUBING WERE RECEIVED FOR INVESTIGATION. ·~VISUAL OBSERVATION OF THE SAMPLE SHOWED NO ANOMALY, FRICTION MARKS, OR DEFECTS. ·~THE CONTAINER UNDERWENT A PRESSURE TEST BETWEEN PLATES AT 0.8 BAR FOR 5 MINUTES, WHICH CONFIRMED THE ABSENCE OF LEAKAGE AT THE UPPER WELDS. ·~THE CONTAINER WAS THEN PRESSURE TESTED UNDERWATER AT 0.6 BAR AND NO LEAKS WERE DETECTED ON THE CONTAINER SHEETING AND WELDS. ·~THE INLET AND OUTLET TUBINGS WERE TESTED USING THE FOUR-WAY BEND, REVEALING NO ADHESIVE DEFECTS. ADVERSE EVENT ASSESSMENT: (UPDATED INFORMATION) THE REPORTER ASSESSED THE EVENT AS DEATH IN SEVERITY, SERIOUS DUE TO BEING DEATH, POSSIBLE IN RELATION TO PSORALEN TREATED PLATELETS AND INTERCEPT PROCESSING SET FOR PLATELETS. (NEW AND UPDATED INFORMATION) THE CERUS MEDICAL REVIEWER CONSIDERED THE EVENT AS SERIOUS DUE TO DEATH, POSSIBLE IN RELATION TO PSORALEN TREATED PLATELETS AND UNLIKELY IN RELATION TO THE INTERCEPT PROCESSING SET FOR PLATELETS. STAPHYLOCOCCUS UREILYTICUS WAS IDENTIFIED IN BOTH THE PATIENT AND IMPLICATED PTP AND GIVEN THE PATIENT'S CLINICAL PRESENTATION IT IS CONSISTENT WITH A SEPTIC TRANSFUSION REACTION. GENETIC SEQUENCING OF THE PATIENT AND IMPLICATED PTP'S CULTURE ISOLATES SHOWS STRONG GENETIC LINKAGE. STAPHYLOCOCCAL SPECIES ARE GENERALLY EFFECTIVELY INACTIVATED BY THE INTERCEPT SYSTEM AND PATHOGEN INACTIVATION STUDY COMPLETED BY THE CERUS MICROBIOLOGY DEPARTMENT CONFIRMED ROBUST INACTIVATION OF STAPHYLOCOCCUS UREILYTICUS AT 6.7 LOG CFU/ML WITH NO VIABLE BACTERIA OBSERVED AFTER STORAGE ON DAY 7 POST-TREATMENT. THESE DATA SUGGEST THAT CONTAMINATION OCCURRED POST-ILLUMINATION. SINCE THE CAD BAG IS NOT AVAILABLE FOR INVESTIGATION, IT IS POSSIBLE THAT THE CONTAMINATION OCCURRED THROUGH DEFECTS IN THE CAD BAG AFTER ILLUMINATION (I.E., PATHOGEN INACTIVATION) AND PRIOR TO STORAGE. ALTHOUGH REMOTE, IT IS NOT IMPLAUSIBLE THAT THE 3-WAY STOP COCK USED FOR THE TRANSFUSION COULD HAVE INTRODUCED CONTAMINATION INTO THE IMPLICATED PTP AFTER SUCCESSFUL PATHOGEN INACTIVATION. THE DONOR AND EFS'S STAFF MEMBER WHO PERFORMED THE APHERESIS PLATELET COLLECTION WERE CULTURED WITH NEGATIVE RESULTS. THE CO-COMPONENT PLASMA WAS ALSO CULTURE NEGATIVE. FRESENIUS KABI LA CHÂTRE CONDUCTED THOROUGH ANALYSIS OF THE IMPLICATE PTP BAG INTEGRITY. THEY DID NOT IDENTIFY ANY DEFECTS, MARKS, ANOMALIES AND THE UNDERWATER PRESSURE TESTING CONFIRMED THERE IS NO MICROSCOPIC LEAKS. CERUS COMPLETED BATCH RECORD REVIEW OF THE IMPLICATED PTP PROCESSING SET AND DID NOT IDENTIFY ANY OUT OF LIMIT RESULTS. ADDITIONALLY, CERUS HAS COMPLETED PATHOGEN INACTIVATION INVESTIGATION OF THE IMPLICATED PTP CULTURE AND CONFIRMED ROBUST INACTIVATION OF STAPHYLOCOCCUS UREILYTICUS. THE CERUS MEDICAL REVIEWER CONSIDERED THE LISTEDNESS OF THIS EVENT IN RELATION TO THE PSORALEN TREATED PLATELETS AS LISTED BASED ON APPLICABLE CERUS REFERENCE SAFETY INFORMATION. LISTEDNESS ASSESSMENT IS NOT APPLICABLE BECAUSE THIS EVENT IS NOT RELATED TO THE INTERCEPT PROCESSING SET FOR PLATELETS DEVICE. Manufacturer Narrative: UNIQUE DEVICE IDENTIFIER # IS NOT REPORTED IN SECTION D4 BECAUSE THE SUSPECT MEDICAL DEVICE WAS MANUFACTURED FOR MARKET OUTSIDE THE UNITED STATES (FRANCE).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21121220&pc=NAY | 2955842-2025-00301 | 2022-05-16 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-09 05:00 | NAY | DAVINCI SI | Insufficient Information | Hemorrhage/Bleeding; Unspecified Tissue Injury | K081137 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPIRED THE DAY AFTER A DA VINCI-ASSISTED RADICAL PROSTATECTOMY WITH LYMPHADENECTOMY PROCEDURE. DURING THE PROCEDURE, AN UNSPECIFIED FORCEPS INSTRUMENT WAS USED TO GRASP A BLOOD VESSEL, AND ¿DUE TO ITS STRONG GRIPPING FORCE¿ DAMAGE WAS CAUSED TO THE VESSEL RESULTING IN BLEEDING THAT WAS CONTROLLED, AND THE PROCEDURE WAS COMPLETED. POST-OPERATIVELY, UNSPECIFIED SIGNS OF SUSPECTED BLEEDING WERE OBSERVED ¿SUCH AS A DROP IN BLOOD PRESSURE¿. FOLLOW-UP WAS INITIATED AND NO ADDITIONAL INFORMATION WAS RECEIVED TO DATE. Manufacturer Narrative: A REVIEW OF THE DA VINCI SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS, MONOPOLAR CURVED SCISSORS, TWO LARGE NEEDLE DRIVER INSTRUMENTS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A PROSTATECTOMY PROCEDURE IN WHICH AN UNSPECIFIED FORCEPS INSTRUMENT WAS USED TO GRASP A VESSEL WHICH BECAME DAMAGED. THE EXACT VESSEL, HOW THE INSTRUMENT MAY HAVE DAMAGED THE VESSEL, AND THE DETAILS AROUND HOW THIS INJURY HAPPENED WERE NOT PROVIDED. THE PATIENT LATER DIED. HOW THEY DIED AND THE EVENTS LEADING TO THEIR DEATH WERE NOT PROVIDED. HOW THE VESSEL INJURY MAY OR MAY HAVE NOT CONTRIBUTED TO THIS DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124419&pc=MLV | MW5164524 | 2024-10-21 04:00 | Death | ABBOTT MEDICAL | 2025-01-09 05:00 | MLV | AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER |  | Stroke/CVA; Fever; Low Blood Pressure/ Hypotension; Failure of Implant; Septic Shock; Loss of Vision; Ulcer; Pressure Sores; Obstruction/Occlusion; Low Oxygen Saturation; Respiratory Failure; Cognitive Changes |  |  | 1 | Event Description: ON THE MORNING OF MONDAY (B)(6) 2024 AT (B)(6) GA MY MOTHER, (B)(6) HAD AN ELECTIVE OUTPATIENT PROCEDURE TO INSTALL A DEVICE CALLED THE ABBOTT AMPLATZER AMULET LAAO (LEFT ATRIAL APPENDAGE OCCLUDER) WHICH WAS MEANT TO PREVENT BLOOD CLOTS DUE TO AFIB(ATRIAL FIBRILLATION) WHICH COULD LEAD TO STROKE (SHE HAD BEEN TAKING A BLOOD THINNER CALLED ELIQUIS WHICH WAS WORKING FINE FOR HER). THE DEVICE MOVED (EMBOLIZED IS THE TERM HER DOCTORS USED) TO HER MITRAL VALVE IN POST-OP WHICH BLOCKED BLOOD FLOW TO HER HEART. THAT, IN ADDITION TO AN HOUR OF CPR (CARDIOPULMONARY RESUSCITATION) WHILE THEY RETRIEVED THE DEVICE LED TO MULTIPLE STROKES IN BOTH HEMISPHERES OR HER BRAIN, AMONG OTHER THINGS. SHE WAS LEFT BLIND, UNABLE TO MOVE, AND WITH A VERY LOW LEVEL OF AWARENESS. SHE HAD JUST TURNED 79 AND EVEN THOUGH SHE LIVED WITH ME, SHE WAS INDEPENDENT. SHE DID HER OWN SHOPPING, MADE HER OWN MEALS, DROVE A CAR, GARDENED AND CARED FOR THE BIRDS AND SQUIRRELS THAT LIVED IN THE BACKYARD. SHE LOVED READING AND TOOK CONTINUING EDUCATION CLASSES AND WAS TRYING TO LIVE HER BEST LIFE. MY MOTHER AND FAMILY WERE THEN MISTREATED BY THE MEDICAL SYSTEM THROUGHOUT THE ORDEAL THAT FOLLOWED. ABOUT 2 WEEKS LATER SHE WAS DISCHARGED FROM THE HOSPITAL WITH ONLY A FEW HOURS NOTICE AND WITHOUT BEING TOLD OF HER RIGHT TO APPEAL. WE WERE THREATENED WITH THOUSANDS OF DOLLARS A DAY IN CHARGES IF WE DIDN'T HAVE HER MOVED AND THE BED PREVIOUSLY OFFERED TO HER AT THE (B)(6) FOR "REHAB" WAS RESCINDED AND WE WERE LEFT WITH NO OTHER CHOICE BUT TO MOVE HER FURTHER AWAY FROM HOME TO THE (B)(6) IN (B)(6). SHE STAYED THERE ALMOST A MONTH BEFORE, DESPITE OUR APPEALS, SHE WAS DEEMED TO NO LONGER NEED THE LEVEL OF CARE SHE WAS GETTING THERE WHICH WAS A TOTAL FARCE. SO, IN THE LAST WEEK OF HER LIFE WE HAD HER MOVED TO THE (B)(6) IN (B)(6)FOR SUBACUTE CARE ON (B)(6) 2024. SHE HAD DEVELOPED A PRESSURE WOUND IN THE HOSPITAL THAT PERFORMED THE PROCEDURE AND IT WAS NOW STAGE 4. BY FRIDAY (B)(6) 2024 SHE WAS RUNNING A TEMPERATURE OF 104F, HER BREATHING WAS LABORED, HER PULSE OXYGEN WAS VERY LOW, AND HER BLOOD PRESSURE PLUMMETING. WE HAD HER MOVED TO THE(B)(6) EMERGENCY ROOM WHERE THEY STABILIZED HER. SHE WAS SOON ADMITTED TO THE ICU. OUR CHOICE, AT THAT POINT WAS TO CONTINUE TREATING HER WHICH WOULD LEAD TO ADDITIONAL PAIN AND SUFFERING OR CHANGE OUR COURSE TO JUST PROVIDING HER COMFORT. WE CHOSE THE LATTER, AND SHE PASSED PEACEFULLY 12 HOURS LATER IN THE ICU ON (B)(6) 2024 AT 7:21PM. THE ICU STAFF DID AN AMAZING JOB FOR MY MOTHER AND FAMILY (ONE OF THE FEW INSTANCES OF REAL KINDNESS AND CONCERN FROM THE MEDICAL SYSTEM WE SAW DURING THIS NIGHTMARE). SHE DIED OF SEPTIC SHOCK COMPLICATED BY HYPOXEMIC RESPIRATORY FAILURE COMPLICATED BY STAGE 4 SACRAL DECUBITUS ULCER COMPLICATED BY MALNUTRITION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124440&pc=NVN | MW5164525 | 2024-09-05 04:00 | Death | MEDTRONIC, INC. | 2025-01-09 05:00 | NVN | MEDTRONIC SELECTSECURE MRI SURESCAN HBP LEAD |  | Respiratory Failure; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSES OF DEATH WERE AN ACUTE ONSET CHRONIC DECOMPENSATED HEART FAILURE AND HYPOXIC RESPIRATORY FAILURE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125976&pc=LWR | MW5164559 | 2024-12-04 05:00 | Death | ST. JUDE MEDICAL/ABBOTT LABORATORIES | 2025-01-09 05:00 | LWR | TRIFECTA VALE - AORTIC VALVE 25MM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: AS REPORTED THROUGH THE THV HOTLINE, THE PATIENT RECEIVED A 23 MM SAPIEN 3 ULTRA RESILIA VALVE INSIDE A 25 TRIFECTA IN A VALVE-IN-VALVE PROCEDURE. AFTER THE PATIENT LEFT THE CATH LAB AND WENT TO PACU, A LITTLE TIME PASSED, AND THE PATIENT COMPLAINED OF CHEST PAIN. THE PATIENT WAS BROUGHT BACK TO THE CATH LAB, AND AN ANGIOGRAM WAS PERFORMED OF HIS LEFT AND RIGHT CORONARY. THE RIGHT CORONARY ARTERY WAS THE CULPRIT, SO A STENT WAS PLACED IN THE RIGHT CORONARY ARTERY. THE PATIENT HAD REDUCED CHEST PAIN, AND THE PROCEDURE WAS COMPLETED. LATER THAT EVENING THE PATIENT WAS HAVING MORE CHEST PAIN AND WAS BROUGHT BACK TO THE CATH LAB. A STENT WAS PLACED IN THE RIGHT CORONARY ARTERY. AFTER THIS PROCEDURE WAS COMPLETED, THE PATIENT WAS STILL COMPLAINING OF CHEST PAIN. THE CARDIAC SURGEON WAS CONTACTED TO HAVE HIM EXAMINE THE PATIENT AND TAKE HIM TO THE OPERATING ROOM FOR A CABG GRAFT TO THE RCA. THE NEXT DAY, THE PATIENT SEEMED TO BE DOING OKAY. AT SOME POINT WHILE SITTING IN HIS CHAIR HE HAD A LOSS OF PRESSURE AND PULSELESS ELECTRICAL ACTIVITY AND ENDED UP PASSING AWAY. THE INITIAL REPORTER DID NOT ALLEGE THAT ANY EW DEVICE WAS DEFICIENT OR INADEQUATE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21099044&pc=BZD | 2518422-2025-000561 | 2023-08-07 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Asthma; Headache; Inflammation; Pulmonary Dysfunction; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21099105&pc=BZD | 2518422-2025-000568 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Sore Throat; Cancer; Cough; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF KIDNEY DISEASE/TOXICITY, CANCER, DEATH, SORE THROAT AND COUGH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN BOX B. AFTER REVIEW, IT WAS DETERMINED THAT BOX B SHOULD BE CORRECTED. OUTCOMES ATTRIBUTED TO AE HAS BEEN CORRECTED AS AND REPORTED AS BELOW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21099540&pc=BZD | 2518422-2025-000573 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | REMSTAR PRO C-FLEX+ | Material Integrity Problem | Headache; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS HEADACHE, NAUSEA, VOMITING, KIDNEY DISEASE REDUCED CARDIOPULMONARY RESERVE AND DEATH PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21099858&pc=BZD | 2518422-2025-000455 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Headache; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE (UNSPECIFIED), DEATH AND OTHER (UNSPECIFIED EVENT). THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21099952&pc=FRN | 2016493-2024-46658 | 2024-12-06 05:00 | Death | CAREFUSION SD | 2025-01-08 05:00 | FRN | ALARIS SYSTEM | Break; Contamination; Use of Device Problem; Inaccurate Delivery | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION.Event Description: IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON (B)(6) 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915.Event Description: IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON 06 DECEMBER 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915. Manufacturer Narrative: OMIT: B21 - TYPE OF INVESTIGATION NOT YET DETERMINED, C21 - RESULTS PENDING COMPLETION OF INVESTIGATION, D16 - CONCLUSION NOT YET AVAILABLE. CORRECTION: DATE OF EVENT, DATE OF DEATH, DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: DEVICE EVAL BY MANUFACTURER? IMDRF ANNEX A, B, C, D, G CODE AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. EXECUTIVE SUMMARY: THE REPORTED ISSUE OF AN UNINTENDED RATE CHANGE FENTANYL WAS NOT DEFINITIVELY CONFIRMED BUT IS ATTRIBUTED TO BE THE RESULT OF USER PROGRAMMING BASED ON THE LOG ANALYSIS FINDINGS. LABORATORY TEST RESULTS PERFORMED ON THE REPORTED SUSPECT DEVICE CONFIRMED THE PUMP MODULE TO BE WITHIN SPECIFICATION FOR RATE ACCURACY AND WAS OPERATING AS INTENDED, WITH NO SIGNS OF UNREGULATED FLOW OBSERVED. DESPITE THE IDENTIFICATION OF A THIRD-PARTY PART, IT WAS NOT FOUND TO BE A CONTRIBUTING FACTOR TO THE REPORTED EVENT SINCE THE DEVICE WAS OPERATING AS INTENDED DURING TESTING. A PREVENTIVE MAINTENANCE WAS PERFORMED ON PUMP MODULE SN: (B)(6) AND PCU SN: (B)(6) THROUGH ASM V12.3 (COMPATIBLE VERSION), ALL TESTS PASSED. AN EXTERNAL INSPECTION WAS PERFORMED ON THE SUSPECT PUMP MODULE AND PCU FOUND NO IRREGULARITIES RELATED TO THE REPORTED EVENT. AN INTERNAL INSPECTION WAS PERFORMED ON THE SUSPECT PUMP MODULE FOUND NO IRREGULARITIES RELATED TO THE REPORTED EVENT. ON THE PCU, WAS FOUND A THIRD-PARTY PART BATTERY WHICH CAUSED INCIDENTAL ERRORS. THE PROFILE USE WAS IDENTIFIED AS UTSW 5.11, SUSPECT DRUG FENTANYL WAS CONFIRMED THROUGH A PCU EVENT LOG. FROM THE DOWNLOADED PCU AND PUMP MODULE DEVICE LOGS, A KEYPRESS LOG SUMMARY WAS PERFORMED AND IS DETAILED IN THE BELOW TABLE (COVERING THE PERIOD FROM WHEN MODULE WAS LAST POWERED ON (B)(6)2024 TO THE TIME THE MODULE WAS POWERED OFF (B)(6)2024) AT 9:27:08 PM (B)(6)2024 THE USER PROGRAMMED AN INFUSION THROUGH A REMOTE IV, BCM_ALIAS_NDC: (FENTAC); DRUG AMOUNT=2500MCG/250ML; RATE=25.6ML/HR; VTBI=250ML; DOSE=256 MICRO G/H; PRIMARY VOLUME INFUSED (PVI)=0ML. AT 9:27:10 PM, A GUARDRAILS CLINICAL ADVISORY WAS OBSERVED: HIGH RISK 2ND RN TO VERIFY. THE USER PRESSED CONFIRM. AT 10:20:31 PM (B)(6)2024 THE USER PRESSED THE CHANNEL KEY AND CHANGE TO A DOSE=25 MICRO G/H WHICH CHANGE THE RATE TO RATE=2.5ML/HR. BETWEEN 10:22:44 PM - 10:37:47 PM (B)(6)2024 THE USER PRESSED CHANNEL KEY SEVEN (7) TIMES WITHOUT PROGRAMMING AND SELECTED THE START KEY. AT 12:50:34 AM (B)(6)2024 THE USER PROGRAMMED A NINE (9) SECOND BOLUS INFUSION WITH A RATE= 999ML/HR; VTBI=2.5ML. BETWEEN 12:51:06 AM - 12:51:27 AM (B)(6)2024 TWO OCCLUDED PATIENT SIDE ALARM OCCURRED WHICH THE USER RESTARTED THE INFUSION. AT 12:51:38 AM (B)(6)2024 THE BOLUS INFUSION OF RATE= 999ML/HR; VTBI=2.5ML COMPLETED AND THE INFUSION CONTINUED WITH THE PRIMARY INFUSION AT 2.5ML/H. AT 8:19:31 AM (B)(6)2024 THE USER PRESSED KEY CHANNEL OFF, PVI=49.989ML. AT 8:32:15 AM (B)(6)2024 THE USER POWERED OFF THE SYSTEM. ACCORDING TO THE FACILITY'S REPORT, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION." HOWEVER, LOG ANALYSIS REVEALED THAT ON THE (B)(6)2024 AT 10:20:31 PM THE DOSE HAD BEEN MANUALLY MODIFIED DECREASING THE RATE FROM 25.6ML/HR. TO 2.5 ML/HR. THE ALLEGED RATE INCREASE FROM 2.5ML/HR. TO 25.6ML/HR. WAS NOT FOUND WITHIN THE LOGS. AFTER THE FIRST TEST TWO ERROR CODES APPEARED 133.6092 AND 120.4410 WHICH PREVENT TO CONTINUE WITH THE PREVENTIVE MAINTENANCE, THE PCU STOPS RECOGNIZING THE PUMP MODULE, THEREFORE, THE PCB POWER SUPPLY AND THE BATTERY WERE TEMPORARILY REPLACED FOR TESTING PURPOSES ONLY WITH A LABORATORY KNOWN GOOD BOARD WHICH FIXED THE FOUND ERRORS. THESE ERROR EVENTS WERE DEEMED INCIDENTAL ISSUES FOUND OCCURRING DURING THE INVESTIGATION; HOWEVER, THEY WERE NOT ASSOCIATED WITH THE REPORTED INCIDENT AND WERE NOT REPORTED BY THE FACILITY TO HAVE OCCURRED. A PREVENTIVE MAINTENANCE WAS PERFORMED ON PUMP MODULE SN: (B)(6) AND PCU SN: (B)(6) IN THE THROUGH ASM V12.3 (COMPATIBLE VERSION), PREVENTIVE MAINTENANCE TEST TASK WAS PERFORMED. ALL TEST PASSED. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED UNINTENDED RATE CHANGE OF FENTANYL WAS NOT ABLE TO BE CONFIRMED AS A DEVICE ERROR DURING THE INVESTIGATION; HOWEVER, IT IS BEING ATTRIBUTED TO BE A USER ERROR. THE LOG REVIEW SHOWED THE USER STARTED THE FENTANYL INFUSION WITH THE DOSE AT 256 MICRO G/H THEN DECREASED IT TO 25 MICRO G/H. A BOLUS INFUSION OF 2.5ML WAS ALSO FOUND PROGRAMMED AND INFUSED TO COMPLETION AFTER THE DOSE WAS INFUSING AT 25 MICRO G/H. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A0401, A1801, G04067, G02017, C07, C0601, D15 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE. Manufacturer Narrative: OMIT : G07001 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. ADDITIONAL INFORMATION: IMDRF ANNEX G CODE. Manufacturer Narrative: THE ACTUAL DATE OF EVENT IS UNKNOWN. THE ACTUAL DATE OF DEATH IS UNKNOWN. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE FAILURE INVESTIGATION HAS BEEN COMPLETED. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21099954&pc=FRN | 2016493-2024-46659 | 2024-12-06 05:00 | Death | CAREFUSION SD | 2025-01-08 05:00 | FRN | ALARIS SYSTEM | Contamination; Corroded; Use of Device Problem; Inaccurate Delivery | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION.Event Description: IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON (B)(6) 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915.Event Description: IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON 06 DECEMBER 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915. Manufacturer Narrative: OMIT: G07001 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. ADDITIONAL INFORMATION: IMDRF ANNEX G CODE. Manufacturer Narrative: THE ACTUAL DATE OF EVENT IS UNKNOWN. THE ACTUAL DATE OF DEATH IS UNKNOWN. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE FAILURE INVESTIGATION HAS BEEN COMPLETED. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT : B21 - TYPE OF INVESTIGATION NOT YET DETERMINED, C21 - RESULTS PENDING COMPLETION OF INVESTIGATION, D16 - CONCLUSION NOT YET AVAILABLE. CORRECTION : DATE OF EVENT, DATE OF DEATH, DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, B, C, D, G CODES, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. EXECUTIVE SUMMARY: THE REPORTED ISSUE OF AN UNINTENDED RATE CHANGE FENTANYL WAS NOT DEFINITIVELY CONFIRMED BUT IS ATTRIBUTED TO BE THE RESULT OF USER PROGRAMMING BASED ON THE LOG ANALYSIS FINDINGS. LABORATORY TEST RESULTS PERFORMED ON THE REPORTED SUSPECT DEVICE CONFIRMED THE PUMP MODULE TO BE WITHIN SPECIFICATION FOR RATE ACCURACY AND WAS OPERATING AS INTENDED, WITH NO SIGNS OF UNREGULATED FLOW OBSERVED. DESPITE THE IDENTIFICATION OF A THIRD-PARTY PART, IT WAS NOT FOUND TO BE A CONTRIBUTING FACTOR TO THE REPORTED EVENT SINCE THE DEVICE WAS OPERATING AS INTENDED DURING TESTING. A PREVENTIVE MAINTENANCE WAS PERFORMED ON PUMP MODULE SN: (B)(6) AND PCU SN: (B)(6) THROUGH ASM V12.3 (COMPATIBLE VERSION), ALL TESTS PASSED. AN EXTERNAL INSPECTION WAS PERFORMED ON THE SUSPECT PUMP MODULE AND PCU FOUND NO IRREGULARITIES RELATED TO THE REPORTED EVENT. AN INTERNAL INSPECTION WAS PERFORMED ON THE SUSPECT PUMP MODULE FOUND NO IRREGULARITIES RELATED TO THE REPORTED EVENT. ON THE PCU, WAS FOUND A THIRD-PARTY PART BATTERY WHICH CAUSED INCIDENTAL ERRORS. THE PROFILE USE WAS IDENTIFIED AS UTSW 5.11, SUSPECT DRUG FENTANYL WAS CONFIRMED THROUGH A PCU EVENT LOG. FROM THE DOWNLOADED PCU AND PUMP MODULE DEVICE LOGS, A KEYPRESS LOG SUMMARY WAS PERFORMED AND IS DETAILED IN THE BELOW TABLE (COVERING THE PERIOD FROM WHEN MODULE WAS LAST POWERED ON 05DEC2024 TO THE TIME THE MODULE WAS POWERED OFF 06DEC2024). AT 9:27:08 PM 05DEC2024 THE USER PROGRAMMED AN INFUSION THROUGH A REMOTE IV, BCM_ALIAS_NDC: (FENTAC); DRUG AMOUNT=2500MCG/250ML; RATE=25.6ML/HR; VTBI=250ML; DOSE=256 MICRO G/H; PRIMARY VOLUME INFUSED (PVI)=0ML. AT 9:27:10 PM, A GUARDRAILS CLINICAL ADVISORY WAS OBSERVED: HIGH RISK 2ND RN TO VERIFY. THE USER PRESSED CONFIRM. AT 10:20:31 PM 05DEC2024 THE USER PRESSED THE CHANNEL KEY AND CHANGE TO A DOSE=25 MICRO G/H WHICH CHANGE THE RATE TO RATE=2.5ML/HR. BETWEEN 10:22:44 PM - 10:37:47 PM 05DEC2024 THE USER PRESSED CHANNEL KEY SEVEN (7) TIMES WITHOUT PROGRAMMING AND SELECTED THE START KEY. AT 12:50:34 AM 06DEC2024 THE USER PROGRAMMED A NINE (9) SECOND BOLUS INFUSION WITH A RATE= 999ML/HR; VTBI=2.5ML. BETWEEN 12:51:06 AM - 12:51:27 AM (B)(6) 2024 TWO OCCLUDED PATIENT SIDE ALARM OCCURRED WHICH THE USER RESTARTED THE INFUSION. AT 12:51:38 AM (B)(6) 2024 THE BOLUS INFUSION OF RATE= 999ML/HR; VTBI=2.5ML COMPLETED AND THE INFUSION CONTINUED WITH THE PRIMARY INFUSION AT 2.5ML/H. AT 8:19:31 AM 06DEC2024 THE USER PRESSED KEY CHANNEL OFF, PVI=49.989ML. AT 8:32:15 AM 06DEC2024 THE USER POWERED OFF THE SYSTEM. ACCORDING TO THE FACILITY'S REPORT, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION." HOWEVER, LOG ANALYSIS REVEALED THAT ON THE 05DEC2024 AT 10:20:31 PM THE DOSE HAD BEEN MANUALLY MODIFIED DECREASING THE RATE FROM 25.6ML/HR. TO 2.5 ML/HR. THE ALLEGED RATE INCREASE FROM 2.5ML/HR. TO 25.6ML/HR. WAS NOT FOUND WITHIN THE LOGS. AFTER THE FIRST TEST TWO ERROR CODES APPEARED 133.6092 AND 120.4410 WHICH PREVENT TO CONTINUE WITH THE PREVENTIVE MAINTENANCE, THE PCU STOPS RECOGNIZING THE PUMP MODULE, THEREFORE, THE PCB POWER SUPPLY AND THE BATTERY WERE TEMPORARILY REPLACED FOR TESTING PURPOSES ONLY WITH A LABORATORY KNOWN GOOD BOARD WHICH FIXED THE FOUND ERRORS. THESE ERROR EVENTS WERE DEEMED INCIDENTAL ISSUES FOUND OCCURRING DURING THE INVESTIGATION; HOWEVER, THEY WERE NOT ASSOCIATED WITH THE REPORTED INCIDENT AND WERE NOT REPORTED BY THE FACILITY TO HAVE OCCURRED. A PREVENTIVE MAINTENANCE WAS PERFORMED ON PUMP MODULE SN: (B)(6) AND PCU SN:(B)(6) IN THE THROUGH ASM V12.3 (COMPATIBLE VERSION), PREVENTIVE MAINTENANCE TEST TASK WAS PERFORMED. ALL TEST PASSED. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED UNINTENDED RATE CHANGE OF FENTANYL WAS NOT ABLE TO BE CONFIRMED AS A DEVICE ERROR DURING THE INVESTIGATION; HOWEVER, IT IS BEING ATTRIBUTED TO BE A USER ERROR. THE LOG REVIEW SHOWED THE USER STARTED THE FENTANYL INFUSION WITH THE DOSE AT 256 MICRO G/H THEN DECREASED IT TO 25 MICRO G/H. A BOLUS INFUSION OF 2.5ML WAS ALSO FOUND PROGRAMMED AND INFUSED TO COMPLETION AFTER THE DOSE WAS INFUSING AT 25 MICRO G/H. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A040502, A1801, GG04037, G02017, CC0601, D15, D01 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21100008&pc=OZP | 2032227-2025-104640 | 2024-11-13 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-08 05:00 | OZP | PUMP MMT-1884L MM780G BLE 3.0 MG/ML | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hyperglycemia; Unspecified Infection; Cancer; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON AN UNKNOWN DATE OF DEATH. CAUSE OF DEATH WAS HIGH BLOOD SUGAR; THE PUMP HAS TO BE REMOVED. THE PREEXISTING MEDICAL CONDITIONS MAY HAVE SOME UNDERLYING INFECTIONS, BRAIN CANCER. THE CUSTOMER REPORTED THE BLOOD GLUCOSE VALUE OF 659 MG/DL. THE PUMP WAS NOT WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1884L, AND MMT-396A. TROUBLESHOOTING WAS PERFORMED, AND THE ISSUE WAS NOT RESOLVED. IT WAS UNKNOWN WHETHER THE CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED HIGH BLOOD GLUCOSE EVENT OR NOT AND WHETHER THE CUSTOMER WAS USED THE AUTO MODE FEATURE OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. PRODUCT RETURN WAS REQUESTED FOR MMT-1884L. THE CUSTOMER WILL DISCONTINUE USING THE DEVICE, AND THE CUSTOMER RESPONSE WAS THAT THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-396A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21100448&pc=MVK | 3008642652-2025-00199 | 2024-10-22 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-08 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SEVENTEEN APPROPRIATE TREATMENTS AND TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 11:22:22 ON (B)(6) 2024. AT 08:07:12 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 200 BPM. AT 08:07:46, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 190 BPM. POST SHOCK RHYTHM WAS VT @ 200 BPM WITH MOTION ARTIFACT. AT 08:08:12, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 08:09:08, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM. POST SHOCK RHYTHM WAS VT @ 210 BPM. AT 08:09:41, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM. POST SHOCK RHYTHM WAS VT @ 180 BPM. AT 08:10:07, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 08:10:39, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 30 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM. AT 08:13:06, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 240 BPM. AT 08:14:04, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 230 BPM. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM GRADING TO VT @ 240 BPM. AT 08:14:37, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM DEGRADING TO VT @ 230 BPM. AT 08:15:03, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 230 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM WITH PVC¿S. THE RHYTHM THEN DEGRADED TO VT @ 260 BPM. AT 08:15:27, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS VT @ 240 BPM. AT 08:15:53, THE PATIENT RECEIVED THE TENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 230 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH HB. AT 08:18:45, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 220 BPM. AT 08:19:16, THE PATIENT RECEIVED THE ELEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 220 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 90 BPM WITH PVC¿S. THE RHYTHM THEN DEGRADED TO VT @ 250 BPM. AT 08:20:22, THE PATIENT RECEIVED THE TWELFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 230 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH PAC¿S, PVC¿S, AND HB. AT 08:20:56, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 120 BPM WITH PVC¿S. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 130 BPM. AT 08:21:30, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 230 BPM. AT 08:23:06, THE PATIENT RECEIVED THE THIRTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM. POST SHOCK RHYTHM WAS VT @ 215 BPM. AT 08:23:33, THE PATIENT RECEIVED THE FOURTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 08:24:04, THE PATIENT RECEIVED THE FIFTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 220 BPM. POST SHOCK RHYTHM WAS ONE SINUS BEAT TRANSITIONING TO VT @ 230 BPM. AT 08:24:37, THE PATIENT RECEIVED THE SIXTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 220 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM DEGRADING TO VT @ 230 BPM. AT 08:25:03, THE PATIENT RECEIVED THE SEVENTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 230 BPM. POST SHOCK RHYTHM WAS VT @ 220 BPM. THE DEVICE WAS SHUT DOWN AT 09:03:36 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21100918&pc=DTB | 2124215-2025-00180 | 2024-12-11 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | DTB | FINELINE II EZ STEROX | Signal Artifact/Noise | Perforation of Vessels | P960004 |  | 1 | Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD HAD EXHIBITED NOISE. THE PATIENT WAS THOUGHT TO BE PACEMAKER DEPENDENT, SO A NEW RV LEAD WAS IMPLANTED FIRST. WHEN THE OLD LEAD WAS EXTRACTED, IT WAS NOTED TO BE BOUND TO THE RIGHT ATRIAL LEAD, WHICH BECAME DISLODGED. THE RV LEAD TORE THE SUPRA VENA CAVA (SVC) AND CAUSED MASSIVE BLEEDING. CARDIAC SURGERY WAS IMMEDIATELY CALLED BUT THE PATIENT EXPIRED. NO PRODUCTS WILL BE RETURNED FOR ANALYSIS. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21100921&pc=BZD | 2518422-2025-100209 | 2024-11-12 05:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE CALLER INDICATED THAT THE DEVICE IS NO LONGER IN THEIR POSSESSION. NO ADDITIONAL DETAILS WERE PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101092&pc=NVN | MW5164373 | 2024-10-14 04:00 | Death | MEDTRONIC, INC. | 2025-01-08 05:00 | NVN | CAPSUREFIX NOVUS MRI SURESCAN |  | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE REPORTED CAUSE OF DEATH ARE SYSTOLIC HEART FAILURE AND ETIOLOGY UNKNOWN. THE PATIENT HAD AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM INCLUDING MEDTRONIC LEADS. THERE IS NO KNOWN INDICATION THAT THE PATIENT¿S DEATH IS RELATED TO THE LISTED DEVICE(S). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5164374.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101100&pc=NVY | MW5164374 | 2024-10-14 04:00 | Death | MEDTRONIC, INC. | 2025-01-08 05:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN LEAD |  | Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE REPORTED CAUSE OF DEATH ARE SYSTOLIC HEART FAILURE AND ETIOLOGY UNKNOWN. THE PATIENT HAD AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM INCLUDING MEDTRONIC LEADS. THERE IS NO KNOWN INDICATION THAT THE PATIENT¿S DEATH IS RELATED TO THE LISTED DEVICE(S). THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5164373.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101267&pc=BZD | 2518422-2025-100212 | 2024-10-03 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | DREAMSTATION | Degraded | Sleep Dysfunction | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION CPAP PRO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S WIFE STATES ON BEHALF OF HER LATE HUSBAND THAT HER HUSBAND HAD SLEEP APNEA, AND HE PASSED AWAY TWO YEARS AGO. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101417&pc=BZD | 2518422-2025-000716 | 2021-10-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: IN THIS REPORT, SECTION BOX H: DEVICE MANUFACTURERS (EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID) HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101816&pc=BZD | 2518422-2025-000718 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101841&pc=OZD | 1220648-2025-25410 | 2024-12-13 05:00 | Death | ABIOMED, INC. | 2025-01-08 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD GASTROINTESTINAL (GI) BLEED. HEPARIN WAS ON HOLD DUE TO GI BLEED. BLEEDING WAS DUE TO AN OPEN CHEST, NOT IMPELLA RELATED. THE PATIENT WAS GIVEN TWO UNITS OF RED BLOOD CELLS (RBCS) THE PATIENT ENDURED ARRHYTHMIAS. SUPRAVENTRICULAR TACHYCARDIA (SVT) CAUSING SUCTION EVENTS AND VERY LOW FLOWS. CARDIOVERSION WAS PLANNED. THE PATIENT WAS GIVEN ONE UNIT OF RBCS. THE PATIENT REMAINS ON IMPELLA SUPPORT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: B.2 ADDED DEATH AND DATE OF DEATH DUE TO ADDITIONAL INFORMATION RECEIVED. B.5 ADDITIONAL INFORMATION WAS RECEIVED FROM THE CLINICAL SITE REGARDING THE PATIENT OUTCOME OF DEATH. D.6B EXPLANT DATE WAS ADDED. H.1 TYPE OF REPORTABLE EVENT WAS REVISED DUE TO NEW ADDITIONAL INFORMATION RECEIVED. H.6 HEALTH EFFECT - IMPACT CODE WAS ADDED DUE TO ADDITIONAL INFORMATION RECEIVED. E.1 ADDED FAX NUMBER AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25410. H.6 CODE 1721 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25410.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101966&pc=BZD | 2518422-2025-000725 | 2024-03-01 05:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102241&pc=OOB | 2124215-2024-83068 | 2024-08-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Angina; Cardiac Arrest; Chest Pain; Non specific EKG/ECG Changes |  |  | 1 | Event Description: JAPAN ALLIANCE REGISTRY IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. ON (B)(6) 2024, THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH RIGHT ANTERIOR CHEST PAINT AND WAS DIAGNOSED WITH NSTEMI. A PERCUTANEOUS CORONARY INTERVENTION WAS PERFORMED ON THE SAME DATE. THE TARGET LESION, WHICH WAS LOCATED IN THE HIGH LATERAL BRANCH AND WAS 75% STENOSED, WAS TREATED USING AN AGENT DCB MR 2.00 X 20MM. FOLLOWING A MITRAL VALVE REPLACEMENT ON (B)(6) 2024, THE PATIENT WAS EXPERIENCING CHEST PAINS AGAIN, AND AN ELECTROCARDIOGRAM REVEALED ST SEGMENT CHANGES. ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED WITH SUSPECTED CORONARYSPASTIC ANGINA. THEY WERE PRESCRIBED NIFEDIPINE AND WERE DISCHARGED ON (B)(6) 2024. ON (B)(6) 2024, THE PATIENT EXPERIENCED CARDIAC ARREST AND WAS NOT ABLE TO BE RESUSCITATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102538&pc=DSQ | 2916596-2024-52861 | 2024-11-13 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemothorax; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE PUMP WAS IMPLANTED, THE FLOW DROPPED SUDDENLY TO 0LPM. THERE WAS NO CHANGE IN POWER. THE PATIENT HAD LOW FLOW ALARMS. THE PUMP WAS EXCHANGED. THE LOW FLOW ALARMS RESOLVED WITH PUMP EXCHANGE AND VOLUME RESUSCITATION. WHEN THE LOW FLOWS STOPPED THERE WERE NO SYMPTOMS. HOWEVER, WHEN THE FLOWS STOPPED, THE PATIENT'S BLOOD PRESSURE (BP) AND MEAN ARTERIAL PRESSURE (MAP DECREASED ALONG WITH PERIPHERAL ARTERIAL OXYGEN SATURATION (SP02). THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS RIGHT VENTRICULAR (RV) FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE WAS OPERATING AS EXPECTED.Event Description: ON (B)(6) 2024, THE PATIENT WAS DECANNULATED FROM THE PERIPHERAL VA ECMO. AFTER THE IMPELLA RP WAS IN PLACE, A CHEST TUBE WAS PLACED IN THE OPERATING ROOM. THE PATIENT RETURNED TO THE UNIT AND REQUIRED MULTIPLE VASOPRESSORS AND INOTROPE SUPPORT. SHORTLY AFTER THE PATIENT ARRIVED BACK TO THE UNIT, A POST OPERATIONAL CHEST X-RAY WAS PERFORMED. THE CHEST X-RAY REVEALED NEAR COMPLETE WHITEOUT OF LEFT LUNG WITH MEDIASTINAL LEFT SHIFT. A BEDSIDE BRONCHOSCOPY WAS PERFORMED. AT 18:20 ON (B)(6) 2024, SHORTLY AFTER THE BRONCHOSCOPY, THE NURSING STAFF REPORTED A DROP IN HEMOGLOBIN (HGB) TO 8.5 AND RECEIVED ORDERS TO TRANSFUSE ONE-UNIT PACKED RED BLOOD CELLS (PRBC). THERE WERE LARGE AMOUNTS OF CHEST TUBE OUTPUT AT THIS TIME. 15 MINUTES LATER, THE LVAD FLOWS WERE DOWN TO 1 LPM AND DECREASED IMPELLA FLOWS WERE DOWN TO 2.5 LPM. THE MEAN ARTERIAL PRESSURE (MAP) AT THIS TIME WAS 20S-30S. MULTIPLE BLOOD PRODUCTS WERE ORDERED, THE LVAD FLOWS RETURNED TO NEAR NORMAL ONCE VOLUME WAS ADMINISTERED TO THE PATIENT. THE PATIENT REQUIRED MULTIPLE BLOOD PRODUCT TRANSFUSIONS OVERNIGHT ON (B)(6) 2024 INTO THE MORNING OF (B)(6) 2024. THE PRESSOR SUPPORT AND INOTROPE SUPPORT WERE CONTINUED. ON (B)(6) 2024 AT ABOUT 03:00, 2 CHEST TUBES WERE PLACED AT BEDSIDE DUE TO RIGHT SIDED HEMOTHORAX. THE PATIENT WAS BLEEDING AND HAD POOR SUPPORT FROM RP FLEX. THE RP FLEX WAS REMOVED AT BEDSIDE AND A PERCUTANEOUS RVAD WITH A SPECTRUM CANNULA AND A CENTRIMAG WITH OXYGENATOR WAS PLACED. ADEQUATE FLOWS WERE INITIATED WITH RVAD, BUT THE PATIENT WAS STILL BLEEDING FROM CHEST TUBES. NEAR CONSTANT BLOOD PRODUCT ADMINISTRATION WAS GIVEN IN ATTEMPT TO MAINTAIN FLOW AND MAP. THE CARE TEAM DISCUSSED WITH FAMILY REGARDING POOR PROGNOSIS AND THE PATIENT'S FAMILY REQUESTED TO PLACE PATIENT ON COMFORT MEASURES. ON (B)(6) 2024, THE PATIENT PASSED AWAY AFTER THE DEVICES WERE STOPPED. IT WAS REPORTED THAT MILD RIGHT VENTRICULAR (RV) FAILURE EXISTED PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT BUT THE PATIENT EXPERIENCED WORSENING RV FAILURE AT THE TIME OF THE EVENT. A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO RIGHT HEART FAILURE.Event Description: THERE WERE NO LOG FILES FROM (B)(6) 2024. THERE WERE NO PHOTOS OF THE PUMP EXCHANGE THAT WERE TAKEN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3, LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES. A SPECIFIC CAUSE FOR THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 8.0¿ FROM THE PUMP HEADER. THE DISTAL PORTION OF THE PUMP CABLE AND THE MODULAR CABLE WERE NOT RETURNED. THE SEALED OUTFLOW GRAFT, OUTFLOW GRAFT BEND RELIEF, AND GRAFT HARDWARE WERE NOT RETURNED. THE APICAL CUFF WAS RETURNED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE PUMP BODY, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED DEPOSITIONS OR THROMBUS FORMATIONS. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LITERS PER MINUTE (LPM) AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. ADDITIONALLY, SECTION 6 OF THE IFU (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102556&pc=NPT | 2015691-2025-00143 | 2024-12-17 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-01-08 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, DURING TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) FOR THE NATIVE AORTIC POSITION, HEMODYNAMIC INSTABILITY OCCURRED. AFTER CROSSING THE VALVE INTO THE AORTA, BLOOD PRESSURE WAS DECLINED. POST VALVE DEPLOYMENT OF A 23 MM SAPIEN 3 ULTRA RESILIA VALVE WITH NOMINAL VOLUME, VENTRICULAR FIBRILLATION WAS OBSERVED AND PROLONGED HYPOTENSION. EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED AND THE PATIENT LEFT OPERATION ROOM. ADDITIONAL INFORMATION WAS OBTAINED FROM THE CS, EVENTUALLY THE PATIENT PASSED AWAY. THE DETAILS ARE UNKNOWN AT THIS STAGE.Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, DURING TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) FOR THE NATIVE AORTIC POSITION, HEMODYNAMIC INSTABILITY OCCURRED. AFTER CROSSING THE VALVE INTO THE AORTA, BLOOD PRESSURE WAS DECLINED. POST VALVE DEPLOYMENT OF A 23 MM SAPIEN 3 ULTRA RESILIA VALVE WITH NOMINAL VOLUME, VENTRICULAR FIBRILLATION WAS OBSERVED AND PROLONGED HYPOTENSION. EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED AND THE PATIENT LEFT OPERATION ROOM. ADDITIONAL INFORMATION WAS OBTAINED FROM THE CS, EVENTUALLY THE PATIENT PASSED AWAY. THE DETAILS ARE UNKNOWN AT THIS STAGE. ADDITIONAL INFORMATION WAS OBTAINED THROUGH THE JAPANESE TAVI REGISTRY AS BELOW. AFTER TAVR PROCEDURE, CIRCULATORY FAILURE GRADUALLY PROGRESSED, LEADING TO LOWER LIMB ISCHEMIA, ACIDOSIS, AND HYPERKALEMIA. AT AROUND 4:00 ON (B)(6) 2024, THE PATIENT SUFFERED CARDIAC ARREST DUE TO VENTRICULAR TACHYCARDIA. NO FURTHER LIFE-SUSTAINING TREATMENT WAS DESIRED, AND THE PATIENT WAS DECLARED DEAD ON 9:04 WITHIN THE SAME DAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), HYPOTENSION, HEMODYNAMIC INSTABILITY &/OR CARDIOGENIC SHOCK IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE OVERALL TAVR PROCEDURE, INCLUDING VASCULAR AND APICAL ACCESS, USE OF ANGIOGRAPHY, BALLOON VALVULOPLASTY, USE OF CONSCIOUS SEDATION AND/OR GENERAL ANESTHESIA, AND THE BIO-PROSTHESIS IMPLANTATION. PATIENTS UNDERGOING THE TAVR PROCEDURE CAN BE NON-OPERATIVE OR HIGH RISK, HAVE COMPLEX MEDICAL HISTORIES AND MULTIPLE CO-MORBIDITIES. PATIENT RISK FACTORS FOR HYPOTENSION INCLUDE LOW EF, CAD, CHF, ARTERIOSCLEROSIS, HYPOVOLEMIA, AND ANEMIA. ADDITIONALLY, THESE PATIENTS ARE ROUTINELY ADMINISTERED MULTIPLE VASOACTIVE DRUGS DURING THE PROCEDURE AND ARE INTENTIONALLY MADE HYPOTENSIVE, UTILIZING RAPID VENTRICULAR PACING, TO FACILITATE ACCURATE VALVE DEPLOYMENT. AS A RESULT OF THESE FACTORS, INTRA-OPERATIVE HYPOTENSION IS NOT UNCOMMON AND IS TREATED WITH STANDARD THERAPIES, INCLUDING VASOACTIVE DRUGS. IT IS ALSO NOT UNCOMMON TO INITIATE BRIEF CHEST COMPRESSIONS OR CARDIAC MASSAGE TO FACILITATE DISTRIBUTION OF THESE VASOACTIVE DRUGS. IN SOME CASES, THESE STANDARD MANEUVERS ARE NOT ADEQUATE, AND INITIATION OF CARDIOPULMONARY BYPASS (CPB), INSERTION OF IABP, AND/OR CONVERSION TO OPEN SURGERY IS REQUIRED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE'SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. AS A PRECAUTION, THE THV TRAINING MANUALS INSTRUCT THE OPERATOR TO 1) MINIMIZE THE NUMBER AND DURATION OF RAPID BURST PACING EPISODES, AND 2) ALLOW SUFFICIENT HEMODYNAMIC RECOVERY BEFORE INITIATING ANOTHER EPISODE OF RAPID PACING. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT UNKNOWN PATIENT COMORBIDITIES MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADDITIONAL INFORMATION, SECTIONS G6, H2, B5, D1, D6, H6: COMPONENT CODE AND CLINICAL CODE HAVE BEEN UPDATED/CORRECTED. PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS'S DEVICE-RELATED BLEED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE (ADD RISK/CONTRIBUTING FACTORS) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. NO CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102762&pc=DYB | 3006260740-2024-08201 | 2024-12-10 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-01-08 05:00 | DYB | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH 4.5F | Material Frayed | Heart Failure/Congestive Heart Failure | K915167 |  | 1 | Event Description: PATIENTS NEED TO ENTER THE CHEMOTHERAPY DRUGS, SO THE PLACEMENT NURSE ULTRASOUND UNDER THE PLACEMENT OF GROSHONG, THE USE OF PUNCTURE AND THE USE OF THE GUIDE WIRE INTO THE BLOOD VESSEL ARE VERY SMOOTH, BEFORE THE DELIVERY OF THE GUIDE WIRE INTACT, WITHDRAWN AFTER THE DISCOVERY OF THE WINDING WIRE SCATTERED, THE LENGTH OF THE LEADING WIRE MEASURED 34.7CM. THE PATIENT WENT INTO CRITICAL CARE WITH ACUTE HEART FAILURE THE NIGHT AFTER PLACEMENT, WHETHER IT WAS RELATED TO THE GUIDEWIRE IS CURRENTLY UNKNOWN. ADDITIONAL INFORMATION RECEIVED ON 12/24/2024: AT PRESENT, IT IS NOT POSSIBLE TO CONFIRM WHETHER THERE ARE FRAGMENTS IN THE PATIENT'S BODY, THE PATIENT HAS PASSED AWAY. Event Description: PATIENTS NEED TO ENTER THE CHEMOTHERAPY DRUGS, SO THE PLACEMENT NURSE ULTRASOUND UNDER THE PLACEMENT OF GROSHONG, THE USE OF PUNCTURE AND THE USE OF THE GUIDE WIRE INTO THE BLOOD VESSEL ARE VERY SMOOTH, BEFORE THE DELIVERY OF THE GUIDE WIRE INTACT, WITHDRAWN AFTER THE DISCOVERY OF THE WINDING WIRE SCATTERED, THE LENGTH OF THE LEADING WIRE MEASURED 34.7CM. THE PATIENT WENT INTO CRITICAL CARE WITH ACUTE HEART FAILURE THE NIGHT AFTER PLACEMENT, WHETHER IT WAS RELATED TO THE GUIDEWIRE IS CURRENTLY UNKNOWN. ADDITIONAL INFORMATION RECEIVED 12/24/2024: AT PRESENT, IT IS NOT POSSIBLE TO CONFIRM WHETHER THERE ARE FRAGMENTS IN THE PATIENT'S BODY, THE PATIENT HAS PASSED AWAY. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE CUSTOMER REPORTED THE PATIENT PASSED AWAY BUT DID NOT PROVIDE THE PATIENT'S DATE OF DEATH. THE DATE IN THIS REPORT IS THE DATE BD BECAME AWARE OF THE PATIENT'S DEATH. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND IT WILL BE EVALUATED. A SUPPLEMENTAL WILL BE SUBMITTED WITH EVALUATION RESULTS. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A BROKEN GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF A GUIDEWIRE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED THE CORE WIRE OF THE GUIDEWIRE WAS BROKEN, CURVED SLIGHTLY LEADING UP TO THE BREAK, AND PROTRUDING FROM THE COILED WIRE. THE ENTIRE LENGTH OF THE COILED WIRE OF THE GUIDEWIRE APPEARED TO BE TIGHTLY COILED, EXCEPT AT THE LOCATION OF THE PROTRUDING CORE WIRE. NO DETAILS OF THIS LOCATION IN THE COILED WIRE AND NO DETAILS OF THE BREAK IN THE CORE WIRE COULD BE DISCERNED FROM THE RETURNED PHOTOGRAPH. THE DISTAL END OF THE COILED WIRE WAS OBSERVED TO BE MORE LUSTROUS COMPARED TO THE REST OF THE WIRE, WHICH SUGGESTS THAT WELD TIP OF THE GUIDEWIRE WAS INTACT; HOWEVER, THIS COULD NOT BE CONFIRMED FROM INSPECTION OF THE RETURNED PHOTOGRAPH. ADDITIONALLY, THE OTHER END OF THE GUIDEWIRE WAS NOT VISIBLE IN THE RETURNED PHOTOGRAPH; THEREFORE, IT COULD NOT BE CONFIRMED IF BOTH WELD TIPS WERE PRESENT FROM INSPECTION OF THE RETURNED PHOTOGRAPH. WHILE THE CORE WIRE OF THE GUIDEWIRE WAS OBSERVED TO BE BROKEN, IT WAS NOT POSSIBLE TO DETERMINE THE CAUSE OF THE BREAK FROM INSPECTION OF THE RETURNED PHOTOGRAPH, AND IT COULD NOT BE DETERMINED IF PIECES OF THE GUIDEWIRE WERE MISSING AS ONLY A PORTION OF THE GUIDEWIRE COULD BE SEEN IN THE RETURNED PHOTOGRAPH. FURTHERMORE, IT COULD NOT BE DETERMINED IF THE OBSERVED BREAK IN THE GUIDEWIRE WAS RELATED TO THE REPORT OF THE PATIENT¿S DEATH, AND NO CLINICAL INFORMATION OR DATA WAS PROVIDED TO SUGGEST THE GUIDEWIRE BREAK CONTRIBUTED TO THE PATIENT¿S DEATH. NO BATCH INFORMATION WAS RECEIVED FROM THE CUSTOMER AND, CONSEQUENTLY, IT WAS NOT POSSIBLE TO PERFORM A HISTORICAL REVIEW OF SIMILAR COMPLAINTS NOR WAS IT POSSIBLE TO PERFORM A REVIEW OF MANUFACTURING RECORDS AND PROCESSES FOR THE BATCH INVOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT, AND THIS EVALUATION WILL BE UPDATED UPON RECEIPT OF A PHYSICAL SAMPLE FOR EVALUATION AND/OR ANY ADDITIONAL INFORMATION PROVIDED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102848&pc=NIK | 2124215-2024-83084 | 2024-11-15 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | NIK | MOMENTUM X4 CRT-D | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Septic Shock | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM HAD BEEN HOSPITALIZED IN THE INTENSIVE CARE UNIT (ICU) DUE TO SEPSIS AND INFECTION. A REVISION PROCEDURE WAS PLANNED, WHERE THE BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) DEFIBRILLATION LEAD AND THE COMPETITOR'S RIGHT ATRIAL (RA), RIGHT VENTRICULAR (RV), AND LEFT VENTRICULAR (LV) PACING LEADS WERE GOING TO BE EXTRACTED. HOWEVER, ONLY THE CRT-D WAS REMOVED FROM THE PATIENT, THEN THE PATIENT DIED DURING THE PROCEDURE. THE OFFICIAL CAUSE OF DEATH WAS REPORTED TO BE SEPTIC SHOCK. THE EXPLANTED DEVICE WAS NOT EXPECTED TO BE RETURNED. ALL LEADS REMAIN ABANDONED IN THE PATIENT. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102886&pc=LWS | 2124215-2024-83096 | 2024-11-15 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | LWS | ENDOTAK RELIANCE SG | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Septic Shock | P910073 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM HAD BEEN HOSPITALIZED IN THE INTENSIVE CARE UNIT (ICU) DUE TO SEPSIS AND INFECTION. A REVISION PROCEDURE WAS PLANNED, WHERE THE BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) DEFIBRILLATION LEAD AND THE COMPETITOR'S RIGHT ATRIAL (RA), RIGHT VENTRICULAR (RV), AND LEFT VENTRICULAR (LV) PACING LEADS WERE GOING TO BE EXTRACTED. HOWEVER, ONLY THE CRT-D WAS REMOVED FROM THE PATIENT, THEN THE PATIENT DIED DURING THE PROCEDURE. THE OFFICIAL CAUSE OF DEATH WAS REPORTED TO BE SEPTIC SHOCK. THE EXPLANTED DEVICE WAS NOT EXPECTED TO BE RETURNED. ALL LEADS REMAIN ABANDONED IN THE PATIENT. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102898&pc=OZD | 1220648-2025-25417 |  | Death | ABIOMED, INC. | 2025-01-08 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: A PUBLICATION WAS RECEIVED TITLED CLINICAL AND PATHOLOGICAL CHARACTERISTICS OF IMMUNE CHECKPOINT INHIBITOR ASSOCIATED FULMINANT MYOCARDITIS," AND INDICATED A PATIENT (UNKNOWN RACE AND ETHNICITY) UNDERGOING TREATMENT FOR LUNG CANCER WAS ADMITTED DUE TO FEVER AND CHEST PAIN. ECHOCARDIOGRAPHY REVEALED DIFFUSE SEVERE LEFT VENTRICLE HYPOKINESIS WITH LEFT VENTRICLE EJECTION FRACTION OF 10%. CORONARY ANGIOGRAPHY REVEALED NO STENO-OCCLUSIVE LESIONS. DUE TO HEMODYNAMIC INSTABILITY, VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA-ECMO) AND INTRA-AORTIC BALLOON PUMP (IABP) WERE INTRODUCED. THE PATIENT WAS TRANSFERRED TO AN OUTSIDE HOSPITAL ON THE THIRD DAY AND AN IMPELLA CP DEVICE WAS INSERTED TO REPLACE THE IABP. THE NEXT DAY, THE PATIENT'S PUPILS WERE OBSERVED, AND A HEAD COMPUTED TOMOGRAPHY (CT) SCAN CONFIRMED "EXTENSIVE" CEREBRAL HEMORRHAGE IN THE LEFT FRONTOPARIETAL LOBES ASSOCIATED WITH VENTRICULAR PERFORATION AND UNCAL HERNIATION. THE PATIENT'S COMA WAS DEEMED IRREVERSIBLE, SO SURGERY WAS NOT AN OPTION. ALTHOUGH CARDIAC FUNCTION GRADUALLY IMPROVED, THE PATIENT EXPIRED OF INTRACEREBRAL HEMORRHAGE WITH UNCINATE HERNIATION ON THE 13TH DAY OF SUPPORT. Manufacturer Narrative: AT THE TIME OF THIS MDR WRITING SOME MANDATORY INFORMATION IS UNKNOWN: DATE OF ONSET AND DATE OF DEATH ARE UNKNOWN; RESPECTIVE FIELDS LEFT BLANK. DEVICE-SPECIFIC INFORMATION OTHER THAN THE BRAND NAME IS UNAVAILABLE, AND THEREFORE RESPECTIVE SECTIONS (D4, D6 AND H4) ARE REFLECTED AS "UNKNOWN" OR FIELD LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. RYO IZUMI, ET AL. (2024). CLINICAL AND PATHOLOGICAL CHARACTERISTICS OF IMMUNE CHECKPOINT INHIBITOR ASSOCIATED FULMINANT MYOCARDITIS. CARDIO-ONCOLOGY (2024) 10:82. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103128&pc=MNS | 2518422-2025-100236 | 2024-12-31 05:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS LOCATED IN THE LOGS, AND NO FAULTS WERE FOUND. THE DEVICE'S FILTERS, TUBING, AND MANUAL WERE REPLACED AND THE DEVICE SENT TO BE PUT BACK INTO LOAN STOCK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103151&pc=NVY | 2649622-2025-00664 | 2024-10-26 04:00 | Death | MPRI | 2025-01-08 05:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Signal Artifact/Noise; Over-Sensing; Inappropriate/Inadequate Shock/Stimulation | Hypovolemic Shock; Pneumonia; Shock from Patient Lead(s) | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH-RATE NON-SUSTAINED EPISODES AND SHORT VENTRICULAR INTERVENTRALS DURING A SURGICAL PROCEDURE AND HIGH-VOLTAGE THERAPIES WERE DELIVERED. ADDITIONALLY, IT WAS NOTED THAT THE RV LEAD EXHIBITED NOISE AND T-WAVE OVERSENSING (TWOS). A RIGHT ATRIAL (RA) LEAD POSITION CHECK HAD FAILED. BOTH LEAD REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN OFFICE NOTED THAT THE THERAPY WAS INAPPROPRIATE DUE TO RIGHT VENTRICULAR (RV) LEAD NOISE OVERSENSING LIKELY DURING SURGERY.Event Description: ADDITIONAL INFORMATION RECEIVED FROM THE PHYSICIAN OFFICE NOTED THAT THE PATIENT IS DECEASED. IT WAS NOTED THAT THE PATIENT DIED A DAY AFTER THE SURGICAL PROCEDURE. IT WAS ALSO NOTED THAT A MAGNET WAS NOT IN PLACE AND DETECTIONS WERE NOT TURNED OFF AT THE TIME OF THE SURGERY. THE PATIENT ULTIMATELY DIED FROM HYPOVOLEMIC SHOCK AS A CONSEQUENCE OF PNEUMONIA. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 407652 LEAD, IMPLANTED: (B)(6) 2019. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103388&pc=NVN | 2017865-2025-10474 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-08 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-10473, RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-10475. IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL EXAMINATION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103492&pc=OZD | 1220648-2025-25421 | 2024-12-21 05:00 | Death | ABIOMED, INC. | 2025-01-08 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, THE STAFF NOTICED A LARGE HEMATOMA FORMING AROUND THE IMPELLA SITE. DESPITE APPLYING MANUAL PRESSURE, THE HEMATOMA CONTINUED TO GROW. CRITICAL CARE ORDERED TO START MASSIVE TRANSFUSION PROTOCOL. NOREPINEPHRINE HAD TO BE STARTED, AND THE CONSOLE BEGAN ALARMING SUCTION. THE DOCTOR CAME TO THE BEDSIDE AND THE PLAN WAS TO RETURN TO THE CATHETERIZATION LAB TO CLOSE THE RIGHT GROIN AND GET ACCESS IN THE LEFT GROIN AND PLACE A NEW DEVICE. AFTER THE PLAN WAS FORMULATED, THE FAMILY CAME TO THE BEDSIDE AND DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL INFORMATION WAS SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY PATIENT MOVEMENT, PATIENT WAS RATHER DISORIENTED AND BEGAN THRASHING ABOUT IN BED AND ATTEMPTING MULTIPLE TIMES TO SIT UP AS PER THE CLINICAL DESCRIPTION PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103654&pc=DSQ | 2916596-2024-52927 | 2024-12-08 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Material Twisted/Bent; Adverse Event Without Identified Device or Use Problem | Dyspnea; Low Blood Pressure/ Hypotension; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT HAD BEEN ADMITTED ON (B)(6) 2024 WITH LOW FLOW ALARMS IN THE SETTING OF ACUTE DECOMPENSATED HEART FAILURE (ADHF). A COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) SCAN WAS DONE ON (B)(6) 2024 AFTER THEIR ACUTE KIDNEY INJURY RESOLVED. THE CTA SHOWED NON-OCCLUSIVE THROMBUS OUTFLOW TRACT RESULTING IN MODERATE TO SEVERE NARROWING CAUSED BY A SIGNIFICANT TWIST IN THE OUTFLOW GRAFT. THE RISKS AND BENEFITS OF STENTING WERE DISCUSSED, AND THE PATIENT OPTED OUT OF INTERVENTION. ON (B)(6) 2024 THE PATIENT WAS MADE COMFORT CARE AFTER CONTINUING TO HAVE FREQUENT LOW FLOW EVENTS, INCREASED SHORTNESS OF BREATH (SOB), HYPOXEMIA AND HYPOTENSION. THE PATIENT PASSED AWAY DUE TO PUMP MALFUNCTION. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED OUTFLOW GRAFT TWIST COULD NOT BE CONFIRMED THROUGH THE EVALUATION OF THE SUBMITTED IMAGES. ADDITIONALLY, THE REPORTED LOW FLOW ALARMS WERE UNABLE TO BE CONFIRMED THROUGH THIS EVALUATION, AS NO LOG FILES WERE SUBMITTED FOR ANALYSIS. A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RENAL DYSFUNCTION AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT SUBMITTED TWO IMAGES FOR REVIEW THAT SHOWED A COMPUTED TOMOGRAPHY (CT) SCAN OF THE OUTFLOW GRAFT AND BEND RELIEF. THE SECOND IMAGE SHOWS ARROWS POINTING TO A DARK AREA UNDERNEATH THE BEND RELIEF; HOWEVER, THE ORIENTATION OF THE GRAFT COULD NOT BE CONCLUSIVELY DETERMINED BASED ON THIS IMAGE. THEREFORE, THE REPORTED OUTFLOW GRAFT TWIST IS UNABLE TO BE CONCLUSIVELY CONFIRMED. THE HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS SYSTEM ALARM CONDITIONS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", CONTAINS INFORMATION ON "PREPARING THE SEALED OUTFLOW GRAFT" AND EXPLAINS THAT PRIOR TO IMPLANTATION, THE BEND RELIEF SHOULD BE DISENGAGED FROM THE GRAFT FOR THE DE-AIRING PROCEDURE. SECTION 5 ALSO CONTAINS A SUB-SECTION ON "ATTACHING THE SEALED OUTFLOW GRAFT TO THE AORTA", WHICH INSTRUCTS THE USER TO STRETCH THE GRAFT COMPLETELY AND THEN MEASURE AND CUT THE SEALED OUTFLOW GRAFT TO THE APPROPRIATE LENGTH. THIS SECTION ALSO INSTRUCTS THE USER TO VERIFY THAT THE OUTFLOW GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. ADDITIONALLY, THIS SECTION CAUTIONS THE USER: "DO NOT ROTATE/TWIST THE SEALED GRAFT. CHECK THE ALIGNMENT OF THE BLACK LINE ON THE GRAFT TO VERIFY THAT THE SEALED GRAFT IS NOT TWISTED OR KINKED." THIS SECTION ALSO EXPLAINS HOW TO ATTACH THE BEND RELIEF ONCE THE VENT NEEDLE HAS BEEN REMOVED FROM THE SEALED OUTFLOW GRAFT AND LEAKS HAVE BEEN RULED OUT. ALTHOUGH THE DEVICE WAS REPORTEDLY IMPLANTED PRIOR TO THE IMPLEMENTATION OF THE OUTFLOW GRAFT CLIP, THE HEARTMATE 3 LVAS IFU INCLUDES INSTRUCTIONS FOR INSTALLING THE OUTFLOW GRAFT CLIP, AND STATES TO ATTACH THE OUTFLOW GRAFT CLIP TO PREVENT POST-OPERATIVE OUTFLOW GRAFT TWISTING. THIS SECTION FURTHER WARNS THAT FAILURE TO INSTALL THE OUTFLOW GRAFT CLIP SO THAT IT IS FLUSH WITH THE BEND RELIEF CAN ALLOW GRAFT TWISTING OR ABRASION WHICH MAY LEAD TO SERIOUS ADVERSE EVENTS SUCH AS BLEEDING, GRAFT OCCLUSION, THROMBOSIS, AND/OR DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103746&pc=DSQ | 3007042319-2025-00037 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-01-08 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE WAS FOUND DECEASED AT HOME. THE PATIENT HAD EXPERIENCED SEVERAL HOSPITALIZATIONS FOR ACUTE DECOMPENSATED HEART FAILURE (ADHF) DURING THE YEAR AND HAD A DECLINING FUNCTIONAL STATUS PRIOR TO DEATH. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT HAD EXPERIENCED SEVERAL HOSPITALIZATIONS FOR ACUTE DECOMPENSATED HEART FAILURE (ADHF) DURING THE YEAR AND HAD A DECLINING FUNCTIONAL STATUS; HOWEVER, THE CAUSE OF DEATH WAS UNKNOWN. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103752&pc=DSQ | 2916596-2024-52966 | 2024-11-04 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONAL INFORMATION WAS RECEIVED THAT DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103860&pc=DSQ | 2916596-2024-52967 | 2024-05-15 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: SECTION E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONAL INFORMATION WAS RECEIVED THAT DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103938&pc=DSQ | 2916596-2024-52961 | 2024-09-04 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE HOSPITAL DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21104155&pc=DSQ | 2916596-2025-00096 | 2021-09-01 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Unspecified Infection; Brain Injury; Respiratory Failure; Cardiovascular Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿VENTRICULAR ASSIST DEVICE USING A THORACOTOMY-BASED IMPLANT TECHNIQUE: MULTI-CENTER IMPLANTATION OF THE HEARTMATE 3 IN SUBJECTS WITH HEART FAILURE USING SURGICAL TECHNIQUES OTHER THAN FULL MEDIAN STERNOTOMY (HM3 SWIFT)¿ THAT THERE WAS A PROSPECTIVE, MULTICENTER, SINGLE-ARM STUDY WAS CONDUCTED IN PATIENTS ELIGIBLE FOR HEARTMATE 3 IMPLANTATION WITH THORACOTOMY-BASED SURGICAL TECHNIQUE (BILATERAL THORACOTOMY OR PARTIAL UPPER STERNOTOMY WITH LEFT THORACOTOMY). THE STUDY EXAMINED HOW THE IMPLANTATION OF THE HEARTMATE 3 IN SUBJECTS WITH HEART FAILURE USING SURGICAL TECHNIQUES OTHER THAN FULL MEDIAN STERNOTOMY STUDY TO COMPARE CLINICAL OUTCOMES OF THE THORACOTOMY-BASED TECHNIQUE FOR IMPLANTATION OF THE HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) WITH A STANDARD MEDIAN STERNOTOMY IN PROPENSITY SCORE¿MATCHED PATIENTS FROM THE MOMENTUM 3 TRIAL PORTFOLIO. THE COMPOSITE PRIMARY END POINT WAS SURVIVAL FREE OF DISABLING STROKE, OR REOPERATION TO REMOVE OR REPLACE A MALFUNCTIONING DEVICE, OR CONVERSION TO MEDIAN STERNOTOMY AT 6-MONTHS POSTIMPLANT (ELECTIVE TRANSPLANTS WERE TREATED AS A SUCCESS). THE STUDY ENROLLED 102 PATIENTS BETWEEN DEC2020 AND JUL2022 IN THE THORACOTOMY-BASED ARM IN 23 CENTERS IN NORTH AMERICA. 204-STERNOTOMY-BASED PATIENTS FROM THE MOMENTUM 3 CAP STUDY WERE USED IN COMPARISON. PATIENTS WERE MOSTLY MEN (81.4% THORACOTOMY-BASED AND 83.8% STERNOTOMY BASED), WHITE (68.6% THORACOTOMY-BASED AND 72.1% STERNOTOMY-BASED). IN HM3 SWIFT, SURGEONS USED THE THORACOTOMY-BASED TECHNIQUE INVOLVING EITHER A BILATERAL THORACOTOMY OR A PARTIAL UPPER STERNOTOMY FOR ACCESS TO THE ASCENDING AORTA COUPLED WITH A LEFT ANTEROLATERAL THORACOTOMY INCISION TO GAIN ACCESS TO THE APEX OF THE LEFT VENTRICLE. CARDIOPULMONARY BYPASS COULD BE INITIATED BY DIRECTLY CANNULATING THE ASCENDING AORTA OR PERIPHERALLY BY CANNULATING THE COMMON FEMORAL ARTERY AND VEIN. THE PRECISE LOCATION OF THE THORACOTOMY INCISION WAS DETERMINED VIA AN INTRAOPERATIVE TRANSTHORACIC ECHOCARDIOGRAM BEFORE INITIATING THE THORACOTOMY-BASED SURGERY. IN THE ANALYSIS OF THE PRIMARY END POINT, 85.2% OF THORACOTOMY-BASED PATIENTS MET THE CRITERIA COMPARED WITH 86.1% OF STERNOTOMY-BASED PATIENTS. WHEN EVALUATING ALL EVENTS, THERE WERE NO DIFFERENCES BETWEEN THE 2 GROUPS IN OVERALL SURVIVAL (P = .85), CAUSE-SPECIFIC MORTALITY RATES, OR FREEDOM FROM DISABLING STROKE (P = .97). THERE WERE NO DIFFERENCES IN ICU LENGTH OF STAY BETWEEN THE 2 GROUPS (P= .84), HOWEVER, IN PATIENTS DISCHARGED, THE LENGTH OF STAY WITH THORACOTOMY-BASED IMPLANT WAS LONGER (P = .03). THE TOTAL IMPLANT AND CARDIOPULMONARY BYPASS TIME WAS LONGER IN THE THORACOTOMY-BASED IMPLANT COMPARED WITH THE STERNOTOMY-BASED IMPLANT (BOTH P<.0001). NO DIFFERENCES WERE OBSERVED FOR BLOOD PRODUCT UTILIZATION (P=.0009), ADVERSE EVENTS (INCLUDING RIGHT HEART FAILURE, P = .26), FUNCTIONAL STATUS, AND QUALITY OF LIFE BETWEEN COHORTS. HOWEVER, RIGHT VENTRICULAR ASSIST DEVICE USE WAS HIGHER IN THORACOTOMY-BASED COHORT (0.33 EPPY THORACOTOMY-BASED VS 0.12 STERNOTOMY-BASED; P = .02). THERE WAS NO DIFFERENCE IN STROKE OF ANY SUBTYPE AND SEVERITY (0.19 EPPY THORACOTOMY-BASED VS 0.18 STERNOTOMY-BASED; P ¼ .92). FOR THE PRESPECIFIED SUBGROUPS OF AGE, SEX, RACE, OR ETHNIC GROUP, INTENDED GOAL OF PUMP SUPPORT (BRIDGE TO TRANSPLANTATION OR DESTINATION THERAPY), OR INTERMACS PROFILE, NO INTERACTION BETWEEN THE INDIVIDUAL PRESPECIFIED SUBGROUPS AND IMPLANT METHOD WAS OBSERVED REGARDING THE PRIMARY END POINT. WITHIN THE THORACOTOMY-BASED TECHNIQUE, AORTIC ACCESS WAS OBTAINED BY RIGHT THORACOTOMY IN 41 OUT OF 102 (40%) OF PATIENTS, THE REMAINING PATIENTS HAD PARTIAL STERNOTOMY THAT INCLUDED BOTH TYPE-J IN 47 OUT OF 102 (46%) AND TYPE T IN 13 OUT OF 102 (13%). IN 1 SUBJECT WITH A PARTIAL STERNOTOMY, THE TYPE OF HEMI-STERNOTOMY WAS NOT RECORDED. AORTIC ACCESS PERFORMED VIA SMALL RIGHT THORACOTOMY OR PARTIAL UPPER STERNOTOMY IN THE THORACOTOMY-BASED IMPLANT SURGICAL TECHNIQUE DID NOT RESULT IN DIFFERENCES IN PRIMARY END POINT SUCCESS, SURVIVAL, OR FREEDOM FROM DISABLING STROKE. ONLY 1 PATIENT REQUIRED CONVERSION TO A MEDIAN STERNOTOMY. THE PRINCIPAL FINDING OF THE HM3SWIFT STUDY WAS THAT THE THORACOTOMY-BASED SURGICAL TECHNIQUE WAS NONINFERIOR WHEN COMPARED WITH A TRADITIONAL MEDIAN STERNOTOMY-BASED IMPLANT OF THE HM3 LVAD. NO MAJOR BENEFITS OR RISKS FOR THE THORACOTOMY-BASED IMPLANT TECHNIQUE IN ADVERSE EVENTS, HEALTH CARE RESOURCE USE PARAMETERS (LENGTH OF HOSPITAL STAY OR USE OF BLOOD PRODUCTS), OR IN FUNCTIONAL CAPACITY, QUALITY OF LIFE, OR SURVIVAL AMONG THE SURGICAL IMPLANT TECHNIQUES. THERE WERE 2 BRAIN RELATED MORTALITIES (2/102) IN THE THORACOTOMY-BASED PATIENTS AND 6 BRAIN RELATED MORTALITIES (6/204) IN STERNOTOMY-BASED PATIENTS. THERE WAS 1 STROKE MORTALITY (1/102) IN THE THORACOTOMY-BASED PATIENTS AND 5 STROKE MORTALITIES (5/204) IN STERNOTOMY-BASED PATIENTS. THERE WERE 6 CARDIOVASCULAR RELATED MORTALITIES (6/102) IN THE THORACOTOMY-BASED PATIENTS AND 16 CARDIOVASCULAR RELATED MORTALITIES (16/204) IN STERNOTOMY-BASED PATIENTS. THERE WAS 1 INFECTION RELATED MORTALITY (1/102) IN THE THORACOTOMY-BASED PATIENTS AND 2 INFECTION RELATED MORTALITIES (2/204) IN STERNOTOMY-BASED PATIENTS. THERE WAS 1 RESPIRATORY RELATED MORTALITY (1/102) IN THE THORACOTOMY-BASED PATIENTS AND 1 RESPIRATORY RELATED MORTALITY (1/204) IN STERNOTOMY-BASED PATIENTS. THERE WERE 2 OTHER MORTALITIES (2/102) IN THE THORACOTOMY-BASED PATIENTS AND 1 OTHER MORTALITY (1/204) IN STERNOTOMY-BASED PATIENTS. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01SEP2021 AS DATA WAS COLLECTED BETWEEN 01DEC2020 AND 01JUL2022. AUTHOR INFORMATION: GOSEV, I., PHAM, D. T., UM, J. Y., ANYANWU, A. C., ITOH, A., KOTKAR, K., TAKEDA, K., NAKA, Y., PELTZ, M., SILVESTRY, S. C., COUPER, G., LEACCHE, M., RAO, V., SUN, B., TEDFORD, R. J., MOKADAM, N., MCNUTT, R., CRANDALL, D., MEHRA, M. R., & SALERNO, C. T. (2024). VENTRICULAR ASSIST DEVICE USING A THORACOTOMY-BASED IMPLANT TECHNIQUE: MULTI-CENTER IMPLANTATION OF THE HEARTMATE 3 IN SUBJECTS WITH HEART FAILURE USING SURGICAL TECHNIQUES OTHER THAN FULL MEDIAN STERNOTOMY (HM3 SWIFT). THE JOURNAL OF THORACIC AND CARDIOVASCULAR SURGERY, 168(5). HTTPS://DOI.ORG/10.1016/J.JTCVS.2024.02.013. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS CANNOT BE CONCLUSIVELY DETERMINED. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LVAS IFU, REVISION B IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21104228&pc=DSQ | 2916596-2024-52876 | 2024-12-18 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Ascites; Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT DURING HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT IN THE OPERATING ROOM (OR), THE PATIENT BECAME VASOPLEGIC, REQUIRING MULTIPLE BLOOD PRODUCTS, FLUIDS, AND PRESSORS. THE SURGEON NOTED AN INCREASE IN THE PATIENT'S ABDOMINAL SIZE AND FIRMNESS. THE PATIENT WAS UNABLE TO BE WEANED OFF CARDIOPULMONARY BYPASS (CPB) AND THERE WAS A DROP IN THE PATIENT'S PARTIAL PRESSURE OF OXYGEN (PO2). AN EXPLORATORY LAPAROTOMY WAS PERFORMED, AT THE PATIENT'S BEDSIDE AND A CENTRIMAG RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS PLACED. THE RIGHT HEART FAILURE (RHF) WAS MULTIFACTORIAL AS THE PATIENT SPENT A PROLONGED TIME ON CARDIOPULMONARY BYPASS (CPB) DUE TO AORTIC VALVE REGURGITATION AND BLEEDING, NEEDING RESUSCITATION. THE SOURCE OF THE BLEEDING WAS CONFIRMED TO BE FRIABLE EPICARDIAL TISSUE AND WAS TREATED WITH BLOOD PRODUCTS AND FLUIDS. IT WAS UNKNOWN IF RIGHT HEART FAILURE EXISTED PRE-LVAD IMPLANT, BUT IT WAS DETERMINED NOT TO BE CAUSED BY A DEVICE-RELATED ISSUE. THE PATIENT BECAME HEMODYNAMICALLY STABLE AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). BOTH LVAD AND RVAD HAD ADEQUATE PUMP FLOW WITH NO ALARMS AND WERE WORKING AS INTENDED. THERE WAS A PLANNED WASH OUT IN THE OPERATING ROOM WITH A POTENTIAL CLOSING OF THE CHEST ON (B)(6) 2024, BUT WAS UNABLE TO LEAVE THE OPERATING ROOM DUE TO THE BLEEDING AND HEMODYNAMIC INSTABILITY. THE PATIENT PASSED AWAY IN THE OPERATING ROOM. THE DEVICE OPERATED AS EXPECTED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR PROVIDED AND THE PUMP WAS NOT EXPLANTED.Event Description: THE EXPLORATORY LAPAROTOMY FOUND NO OBVIOUS BLEEDING, AND THE INCREASE IN ABDOMINAL SIZE WAS ATTRIBUTED MOSTLY TO EDEMATOUS FLUID FILLING TISSUE. Manufacturer Narrative: H6: ADDITIONAL HEALTH EFFECT - IMPACT CODE - 4643 BLOOD TRANSFUSION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND BLEEDING, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, EXPLAINS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. THIS SECTION ALSO STATES THAT IF THE AORTIC INSUFFICIENCY IS NOT ADDRESSED, THE DEVICE WILL NOT BE ABLE TO PROVIDE THE INTENDED FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21104857&pc=MKJ | 3030677-2025-000106 | 2024-12-10 05:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-01-08 05:00 | MKJ | FRX (W B) DEFIB, US ENGLISH, EXCHG | Failure to Deliver Shock/Stimulation | Cardiac Arrest; Unspecified Heart Problem | P180028 |  | 1 | Event Description: THE USER IS QUESTIONING WHY THE DEVICE DID NOT ADVISE A SHOCK ON A SHOCKABLE RHYTHM. THE PATIENT SURVIVED.Event Description: UPDATE TO CONFIRM PATIENT DID NOT SURVIVE. Manufacturer Narrative: UPDATING THE CASE TO A DEATH BASED ON NEW INFORMATION RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21104904&pc=OZD | 1220648-2025-25428 | 2024-12-20 05:00 | Death | ABIOMED, INC. | 2025-01-08 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, BLEEDING WAS NOTED TO IMPELLA SITE. MANUAL PRESSURE WAS HELD. FLUID AND 2 UNITS OF PACKED RED BLOOD CELLS WERE ADMINISTERED FOR VOLUME. THE PATIENT WAS NOTED TO BE IN A-FIB, SYNCHRONIZED CARDIOVERSION ADMINISTERED TWICE. IMPELLA CP WAS REPOSITIONED UNDER FLUORO. THE PATIENT WAS MOVED TO THE INTENSIVE CARE UNIT AND FEMSTOP WAS PLACED. DISTAL PULSES WERE NOT DOPPLERABLE. THE PHYSICIAN WAS AWARE BUT DEEMED THE PATIENT TO BE DEPENDENT ON IMPELLA SUPPORT AND DISTAL PULSES WERE MONITORED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. REPORTABLE DUE TO PATIENT HARM AND POTENTIAL RELATIONSHIP BETWEEN THE CAUSE OF DEATH AND THE IMPELLA. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUES HAS BEEN COMPLETED. NO DEVICE WAS RECEIVED FOR EVALUATION. DEVICE HISTORY RECORD REVIEW CONFIRMED THE PUMP PASSED ALL POST STERILE INSPECTION CHECKS. TO DETERMINE THE TRUE ROOT CAUSE OF THE REPORTED ISSUES, SUFFICIENT CLINICAL INFORMATION WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS ALL NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE REPORTED ISSUES WAS UNABLE TO BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, AN INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105350&pc=KDI | 2521402-2024-00140 | 2024-11-16 05:00 | Death | B. BRAUN AVITUM AG | 2025-01-08 05:00 | KDI | DIALOG+® | No Apparent Adverse Event | Cardiac Arrest |  |  | 1 | Event Description: B. BRAUN WAS CONTACTED BY A TECHNICIAN ASKING FOR SUPPORT ON TROUBLESHOOTING. THE TECHNICIAN WAS DOING ELECTROLYTE TESTING ON A MACHINE ON WHICH A PATIENT HAD DIED DURING DIALYSIS THERAPY. THE TESTING WAS DONE AS PART OF THE CUSTOMER'S PROCEDURE WHEN A PATIENT DIED. THE CUSTOMER DID NOT SUGGEST THAT THE DEVICE CAUSED THE PATIENT TO CODE AND PASS. THE ELECTROLYTE SAMPLE FOR CA++ WAS 0.1 MEQ/L OUT OF DAVITA´S TOLERANCE FOR CALCIUM, WHICH 2.5 ± 0.1, THE SAMPLE RESULTED IN A CALCIUM VALUE OF 2.3 MEQ/L. THE CUSTOMER TOOK ANOTHER DIALYSATE SAMPLE, AND THIS SAMPLE WAS GOOD. THE CONVERSION FACTORS FOR THE CONCENTRATES BEING USED (BOTH BICARBONATE AND ACID CONCENTRATE) WERE FOUND TO BE CORRECT. NO SAMPLE, NO MACHINE DATA RECORD OF THE THERAPY AND NO DETAILS REGARDING THE TECHNICAL SAFETY INSPECTION WERE MADE AVAILABLE FOR INVESTIGATION. THUS, INVESTIGATION OF THE CASE HAS BEEN PERFORMED BASED ON THE REPORTED INFORMATION AND DATA. FURTHER INFORMATION THAT WAS RECEIVED STATES THAT TEMPERATURE AND CONDUCTIVITY WERE TESTED WITH AN EXTERNAL METER. THE COMPARISON MEASUREMENT WAS FOUND TO BE OK. AS THE CONCENTRATES USED AND THE CONVERSION FACTORS CONFIGURED ON THE MACHINE ARE UNKNOWN, IT IS NOT POSSIBLE TO ASSESS CORRECTNESS OF THOSE VALUES. HOWEVER, THE CUSTOMER STATED THAT THE CONVERSION FACTORS WERE CORRECT. THE CUSTOMER FURTHER INFORMED THAT THE MACHINE WAS RECALIBRATED AND PUT BACK TO SERVICE. IT IS NOT CLEAR WHICH VALUES WERE CALIBRATED AND WHY, SINCE PREVIOUS INFORMATION STATED THAT THE TEMPERATURE AND CONDUCTIVITY VALUES WERE OK, THAT IS, WITHIN THE SPECIFIED VALUES. BASED ON THE AVAILABLE INFORMATION, THE INVESTIGATION DOES NOT SHOW ANY EVIDENCE OR INDICATION OF A MALFUNCTION OF THE USED DIALOG+ DIALYSIS MACHINE. THERE IS NO PRODUCT DEVIATION. IN GENERAL, ALL SAFETY RELEVANT FUNCTIONS ARE TESTED DURING THE SELF-TESTS IN PREPARATION. WITHOUT HAVING PASSED THE SELF-TESTS A THERAPY START IS NOT POSSIBLE. IF DURING THERAPY A SAFETY RELEVANT MALFUNCTION OCCURS, THE MACHINE TRIGGERS AN ALARM AND SWITCHES TO PATIENT-SAFE MODE. THE BICARBONATE AND FINAL CONDUCTIVITY OF THE DIALYSIS FLUID ARE CHECKED DURING PREPARATION AND CONTINUOUSLY MONITORED THROUGHOUT THE WHOLE THERAPY. IF AN ERROR REGARDING CONDUCTIVITY HAPPENED IN PREPARATION AN ALARMED WOULD BE TRIGGERED AND A THERAPY START WOULD NOT BE POSSIBLE. IF AN ERROR REGARDING CONDUCTIVITY HAPPENED IN THERAPY, THE DEVICE WOULD TRIGGER AN ALARM AND WOULD DEFAULT TO PATIENT-SAFE MODE. DURING THE PREPARATION PHASE AND DURING THERAPY INCORRECTLY WORKING GEAR PUMPS ARE DETECTED BY THE MACHINE AND WILL BE ALARMED. IF A FAILURE OCCURRED IN PREPARATION PHASE THE MACHINE WOULD NOT PASS THE SELF-TEST AND THEREFORE A THERAPY START WOULD NOT BE POSSIBLE. IF THE FAILURE OCCURRED DURING THERAPY THE MACHINE WOULD TRIGGER AN ALARM AND WOULD DEFAULT TO PATIENT-SAFE MODE. IN THE IFU CHAPTER "CONNECTING CONCENTRATE" IS DESCRIBED: "RISK TO THE PATIENT DUE TO INCORRECT COMPOSITION OF DIALYSIS FLUID! ENSURE THAT THE CORRECT CONCENTRATES ARE PROVIDED FOR THE INTENDED THERAPY. ONLY USE CONCENTRATES WHOSE PRINTED USE-BY DATE HAS NOT EXPIRED. ONLY USE ORIGINALLY CLOSED AND INTACT CONCENTRATE CONTAINERS. OBSERVE STORAGE INFORMATION ON CONCENTRATE CONTAINERS. IT IS RECOMMENDED TO USE CONCENTRATES PRODUCED BY B. BRAUN AVITUM AG. WHEN CONCENTRATES ARE USED THAT ARE NOT PRODUCED BY B. BRAUN AVITUM AG THE CORRECT MIXING RATIO AND COMPOSITION HAS TO BE CHECKED ON THE CONCENTRATE LABEL." IN CONCLUSION, BASED ON THE ANALYSIS AND EVALUATION OF THE PROVIDED INFORMATION THERE WAS NOT MALFUNCTION OF THE DEVICE AND THE DEVICE HAS NOT LED TO THE PATIENT'S DEATH. THIS IS IN LINE WITH THE CUSTOMER'S STATEMENT. THIS CASE IS CLASSIFIED AS NOT REPORTABLE TO THE COMPETENT AUTHORITY BY THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105404&pc=DSQ | 2916596-2025-00146 | 2024-08-01 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: IT WAS LATER REPORTED THAT THE CAUSE OF THE PATIENT'S DEATH WAS DUE TO THEIR FAILURE TO THRIVE AND THEY WITHDREW FROM CARE. DEATH WAS SAID TO NOT HAVE BEEN DEVICE RELATED AND THE DEVICE OPERATE AS EXPECTED. Manufacturer Narrative: SECTION A4- PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE, REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105423&pc=DSQ | 2916596-2024-52951 | 2024-12-13 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM NEUROLOGICAL DYSFUNCTION, BRAIN DEATH. THE CAUSE OF DEATH WAS A LARGE MULTIFOCAL INTRACEREBRAL HEMORRHAGE OF THE RIGHT PARIETAL LOBE AND SUBARACHNOID HEMORRHAGE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS OPERATING AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND IT WAS UNKNOWN IF IT WOULD BE PROVIDED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THE PATIENT HAD A CATASTROPHIC HEMORRHAGIC STROKE THAT WAS INCOMPATIBLE WITH LIFE. THE FAMILY WAS UNWILLING TO WITHDRAW CARE. THE PATIENT HAD LOST ALL OF THEIR BRAIN STEM REFLEXES EXCEPT FOR SPONTANEOUS BREATHING IN THE VENT SO THEY COULD NOT BE DECLARED BRAIN DEAD. THE VENT CORRELATED PERFECTLY TO THE INTRINSIC PULSATILITY IN THE HEARTMATE 3. A NUCLEAR FLOW SCAN WAS OBTAINED, AND THE PATIENT HAD NO FLOW TO THEIR BRAIN, SO THEY WERE DECLARED BRAIN DEAD DESPITE SPONTANEOUS RESPIRATIONS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105490&pc=NPT | 2015691-2025-00159 | 2017-10-01 04:00 | Death | EDWARDS LIFESCIENCES | 2025-01-08 05:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEATH VALVE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140031 |  | 3 | Event Description: EDWARDS REVIEWED THE ARTICLE, ''COMPARABLE EFFICACY AND SAFETY FOR BICUSPID AORTIC VALVE STENOSIS PATIENTS UNDERGOING TRANSCATHETER AORTIC VALVE REPLACEMENT WITH BALLOON-EXPANDABLE OR SELF-EXPANDING VALVES USING WEI'S SIZING METHOD'', CORRESPONDING AUTHOR/S DR. WEI-HSIAN YINA ET AL. A TOTAL OF 43 PATIENTS WERE TREATED IN CHENG HSIN GENERAL HOSPITAL BETWEEN OCTOBER 2017 AND JUNE 2021 AND THE PRESENT STUDY AIMED TO INVESTIGATE WHETHER PROPER SIZING CAN IMPROVE THE PROCEDURAL AND CLINICAL OUTCOMES OF TAVR. THE FOLLOWING EVENTS WERE IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE: THREE PATIENTS WITH A SAPIEN 3 VALVE SUFFERED A POST-PROCEDURAL VALVE THROMBOSIS. IT WAS NEEDED AN ANTICOAGULANT THERAPY. THE THREE PATIENTS PASSED AWAY. Manufacturer Narrative: THE DATES OF THE EVENTS ARE UNKNOWN. HOWEVER, THE ARTICLE PROVIDED THE FOLLOWING DATE RANGE, ''FROM 10/2017 UNTIL 06/2021''. THEREFORE, 01/10/2017 USED AS THE OCCURRENCE DATE. PER THE INSTRUCTIONS FOR USE (IFU), VALVE THROMBOSIS IS A POTENTIAL RISK ASSOCIATED WITH THE USE OF THE TRANSCATHETER HEART VALVES (THV). USE OF THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS) IS CONTRAINDICATED IN PATIENTS WHO CANNOT TOLERATE AN ANTICOAGULATION/ANTIPLATELET REGIMEN. IN ADDITION, IT WARNS THAT VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THROMBOSIS IS A RARE AND WELL-RECOGNIZED COMPLICATION ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. POTENTIAL PATIENT RISK FACTORS SUCH AS ATRIAL FIBRILLATION, SYSTEMIC DISEASE (E.G., SYSTEMIC LUPUS ERYTHEMATOSUS, INFLAMMATION, AND DAMAGE TO VARIOUS BODY TISSUES, INCLUDING JOINTS, SKIN, KIDNEYS, HEART, LUNGS, BLOOD VESSELS, AND BRAIN), THE VALVULAR DISEASE ITSELF, AND REDUCED CARDIAC EJECTION FRACTION CAN CONTRIBUTE TO INCREASED RISK OF THROMBUS/THROMBOSIS. SUB-OPTIMAL ANTICOAGULATION DURING THE PROCEDURE AND UNDERLYING PATIENT CONDITIONS CAN ALSO RESULT IN INCREASED THROMBOGENICITY. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS PREPARATION, ASPIRATION, AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE CLOT(S). SHORT-TERM ANTICOAGULATION THERAPY MAY ALSO BE NECESSARY AFTER VALVE REPAIR, ANTICOAGULATION IS PRESCRIBED PER INSTITUTIONAL GUIDELINES. INTRA-PROCEDURAL INTRA-CARDIAC THROMBUS AND POST-PROCEDURE OR LATE VALVE THROMBOSIS ARE COMPLEX PROCESSES TRIGGERED BY THE INTERACTION BETWEEN THE HOST AND THE DEVICE WHICH ARE HIGHLY VARIABLE AMONG PATIENTS. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE AND A DEFINITIVE ROOT CAUSE CANNOT ALWAYS BE CONFIRMED. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR TO CONSIDER ALL PROCEDURAL AND ANATOMICAL FACTORS. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE DEVICE SYSTEM. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY IN CONJUNCTION WITH ECHOCARDIOGRAPHY FOR OPTIMAL VISUALIZATION DURING POSITIONING AND DEPLOYMENT AND CAUTIONS INCLUDE THE MAINTENANCE OF THE PATIENT'S ANTICOAGULATION STATUS (ACT AT >250 SECONDS) DURING THE PROCEDURE IN ORDER TO PREVENT THROMBOSIS. IN THIS CASE, DESPITE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THIS EVENT. DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. AT THE TIME OF THIS REPORT, THE INFORMATION AVAILABLE DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS ARTICLE. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO INCLUDE CORRECT H10.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105501&pc=BZD | 2518422-2025-000796 | 2021-10-21 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded; Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE THAT A USER OF THE PHILIPS CPAP DEVICE HAD PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER BECAME AWARE THAT A USER OF THE PHILIPS CPAP DEVICE HAD PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. DESPITE THREE GFE ATTEMPTS ON 04/24/2025, 04/29/2025, 05/05/2025 THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE HAS BEEN NO RESPONSE FROM THE CUSTOMER ON THE RETURN OF THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IN ADDITION, APPROPRIATE CODE UPDATED/CORRECTED IN THIS FOLLOW-UP REPORT. Manufacturer Narrative: (DEVICE) PROBLEM CODE GRID AND PATIENT OUTCOME CODE GRID HAVE BEEN CORRECTED IN THIS FOLLOW-UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105518&pc=NPT | 2015691-2025-00161 | 2017-10-01 04:00 | Death | EDWARDS LIFESCIENCES | 2025-01-08 05:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEATH VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140031 |  | 1 | Event Description: EDWARDS REVIEWED THE ARTICLE, ''COMPARABLE EFFICACY AND SAFETY FOR BICUSPID AORTIC VALVE STENOSIS PATIENTS UNDERGOING TRANSCATHETER AORTIC VALVE REPLACEMENT WITH BALLOON-EXPANDABLE OR SELF-EXPANDING VALVES USING WEI'S SIZING METHOD'', CORRESPONDING AUTHOR/S DR. WEI-HSIAN YINA ET AL. A TOTAL OF 43 PATIENTS WERE TREATED IN CHENG HSIN GENERAL HOSPITAL BETWEEN OCTOBER 2017 AND JUNE 2021 AND THE PRESENT STUDY AIMED TO INVESTIGATE WHETHER PROPER SIZING CAN IMPROVE THE PROCEDURAL AND CLINICAL OUTCOMES OF TAVR. THE FOLLOWING EVENTS WERE IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE: ONE PATIENT WITH SAPIEN 3 VALVE IMPLANTED, HAD A POST-PROCEDURAL VALVE FAILURE AND REQUIRED INTERVENTION. THE PATIENT PASSED AWAY. Manufacturer Narrative: THE DATES OF THE EVENTS ARE UNKNOWN. HOWEVER, THE ARTICLE PROVIDED THE FOLLOWING DATE RANGE, ''FROM 10/2017 UNTIL 06/2021''. THEREFORE, 01/10/2017 USED AS THE OCCURRENCE DATE. PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. VALVE REINTERVENTION MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY/LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, DESPITE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THIS EVENT. DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. AT THE TIME OF THIS REPORT, THE INFORMATION AVAILABLE DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS ARTICLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105732&pc=DQY | 2124215-2024-83129 | 2024-12-11 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | DQY | WATCHMAN FXD CURVE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT. PRIOR TO THE WDS BEING INSERTED OR DEPLOYED IN THE PATIENT A PERICARDIAL EFFUSION WAS NOTED. THE PATIENT WAS TAKEN TO OPEN HEART SURGERY TO TREAT THE EFFUSION. THE SAME DAY, THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105829&pc=NKM | 2135147-2025-00219 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-01-08 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "ASSOCIATION OF BASELINE MITRAL VALVE AREA WITH PROCEDURAL AND CLINICAL OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTICENTER STUDY, TO EVALUATE THE ASSOCIATION OF BASELINE MITRAL VALVE AREA (MVA) WITH PROCEDURAL AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TEER WITH MITRACLIP FROM THE OCEAN-MITRAL REGISTRY (OPTIMIZED CATHETER VALVULAR INTERVENTION-MITRAL). DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN PATIENTS UNDERGOING TEER, A SMALL MVA <4.0 CM MAY LIMIT THE NUMBER OF CLIPS IMPLANTED AND INCREASE THE TRANSMITRAL PRESSURE GRADIENT AFTER TEER, BUT BASELINE MVA WAS NOT ASSOCIATED WITH MITRAL REGURGITATION REDUCTION AND CLINICAL OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS SHUNSUKE KUBO AT KURASHIKI CENTRAL HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL DABA-KUBOSHUN101@HOTMAIL.CO.JP]. THE TIME FRAME OF THE STUDY WAS APRIL 2018 TO JUNE 2021. A TOTAL OF 1768 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 1768 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, HEART FAILURE HOSPITALIZATION, PRIOR CORONARY REVASCULARIZATION, PRIOR OPEN-HEART SURGERY, STROKE. (B)(6) UNK MITRACLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), TISSUE DAMAGE, HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE: ASSOCIATION OF BASELINE MITRAL VALVE AREA WITH PROCEDURAL AND CLINICAL OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLES ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "ASSOCIATION OF BASELINE MITRAL VALVE AREA WITH PROCEDURAL AND CLINICAL OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY" SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF THE MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE TITLED ¿ASSOCIATION OF BASELINE MITRAL VALVE AREA WITH PROCEDURAL AND CLINICAL OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY¿. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, HEART FAILURE HOSPITALIZATION, PRIOR CORONARY REVASCULARIZATION, PRIOR OPEN-HEART SURGERY, STROKE. SOME OF THE COMPLICATIONS REPORTED INCLUDED INGLE LEAFLET DEVICE ATTACHMENT (SLDA), TISSUE DAMAGE, HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106023&pc=DSQ | 2916596-2024-52883 | 2024-10-14 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106140&pc=DSQ | 2916596-2024-52881 | 2023-10-24 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PUMP WAS NOT EXPLANTED.Event Description: THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: A1-A4: PATIENT INFORMATION NOT PROVIDED DUE TO PATIENT PRIVACY LAWS. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106356&pc=KDJ | 1416980-2024-07897 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-08 05:00 | KDJ | MINICAP WITH POVIDONE-IODINE SOLUTION | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Unspecified Infection; Peritonitis | NI |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS MANIFESTED BY ABDOMINAL PAIN. ACCORDING TO THE REPORTER, THEY BELIEVED "THE RISK OF PERITONITIS WAS INCREASED DUE TO USING MINICAPS THAT WERE PART OF AN ¿URGENT MEDICAL DEVICE RECALL¿ LETTER REGARDING SEALED POUCHES FOR MINICAPS; HOWEVER, THE CAUSE REMAINS UNKNOWN. IT WAS REPORTED THE PATIENT WAS HOSPITALIZED AND RECEIVED UNSPECIFIED TREATMENT FOR PERITONITIS. ACCORDING TO THE REPORTER, THE PATIENT HAD MULTIPLE SUBSEQUENT HOSPITALIZATIONS DUE TO COMPLICATIONS FROM THE INITIAL PERITONITIS THAT DID NOT RESOLVE. ON AN UNREPORTED DATE, THE PD THERAPY WAS DISCONTINUED, AND THE PATIENT WAS TRANSFERRED TO HEMODIALYSIS. APPROXIMATELY, ONE YEAR AND TWO MONTHS AFTER THE ONSET OF PERITONITIS, THE PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT REPORTED; HOWEVER, THE CAREGIVER REPORTED ¿FROM MULTIPLE ABDOMINAL INFECTIONS". IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: B2: THE PATIENT PASSED AWAY ON AN UNREPORTED DATE IN (B)(6) 2024. B3: THE PERITONITIS OCCURRED ON AN UNSPECIFIED DATE IN JANUARY 2023. G1: MANUFACTURING FACILITY-THIS DEVICE WAS MANUFACTURED AT ONE OF THE TWO FOLLOWING MANUFACTURING SITES: BAXTER HEALTHCARE - CLEVELAND 911 HIGHWAY 61 NORTH PO BOX 1058 CLEVELAND MS 38732 UNITED STATES OR BAXTER HEALTHCARE - CUERNAVACA AVE. DE LOS 50 METROS NO. 2 CUERNAVACA 62578 MEXICO. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H11. H11: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. THE CAUSE OF THE CONDITION COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106366&pc=DSQ | 2916596-2024-52931 | 2023-02-21 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. ADDITIONALLY, DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT PROVIDE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106388&pc=QZI | 2124215-2024-80944 | 2024-12-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | QZI | FARAWAVE PULSED FIELD ABLATION CATHETER, 31MM | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED FOLLOWING A PROCEDURE USING THE FARAWAVE ABLATION CATHETER 31 MM WHERE THE VEINS AND POSTERIOR WALL WERE ABLATED. TREATING FOR POSTERIOR WALL IS OFF-LABEL USE OF THE DEVICE. NO DEVICE ISSUES WERE REPORTED. THE CATHETER IS NOT AVAILABLE FOR RETURN AS IT WAS DISPOSED. THE PULSE FIELD ABLATION PROCEDURE TO TREAT ATRIAL FIBRILLATION WAS COMPLETED WITHOUT INCIDENCE. AFTERWARDS, THEY THEN BEGAN THE WATCHMAN PROCEDURE BY SWITCHING OUT THE FARADRIVE SHEATH WITH A WATCHMAN SHEATH, AND DURING CONTRAST INJECTION INTO APPENDAGE THE PATIENT CODED. THEY WERE ABLE TO RECOVER PATIENT AND TRIED TO CONTINUE WATCHMAN IMPLANT, BUT PATIENT CODED AGAIN. AFTER THE SECOND EVENT A BALLOON PUMP WAS PLACED, AND PATIENT WAS TRANSPORTED TO ICU. A INTRACARDIAC ECHOCARDIOGRAPHY, FLUOROSCOPY, AND CARTO MAPPING WERE PERFORMED. THE PATIENT EXPERIENCED CARDIAC ARREST REQUIRING CHEST COMPRESSIONS. AN IMPELLA PUMP WAS PLACED IN THE PATIENT AS WELL. IT WAS ALSO NOTED THAT PATIENT HAD A NON-CARDIAC PROCEDURE DONE A FEW WEEKS PRIOR TO THIS EVENT AND HAD ARRESTED AS WELL. THE OFFICIAL DATE OF DEATH IS UNKNOWN. TWO DAYS AFTER THE PROCEDURE THE PATIENT PASSED AWAY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106391&pc=NGV | 2124215-2025-00119 | 2024-12-30 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-08 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Dyspnea; No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED USING INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING. A WATCHMAN ACCESS SYSTEM (WAS) WAS POSITIONED AT THE LAA AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS ADVANCED, DEPLOYED, AND SUBSEQUENTLY IMPLANTED AFTER MEETING ALL RELEASE CRITERIA. PROTAMINE WAS ADMINISTERED FOR ANTICOAGULATION REVERSAL AND THE PROCEDURE WAS CONCLUDED. AFTER THE PATIENT WOKE UP FROM ANESTHESIA, THE PATIENT REPORTED HAVING DIFFICULTY BREATHING. MULTIPLE TRANSTHORACIC ECHOCARDIOGRAM (TTE) IMAGING SWEEPS WERE PERFORMED AND FOUND NO PERICARDIAL EFFUSION AND SHOWED THE CLOSURE DEVICE REMAINED IN THE LAA. INTRAVENOUS LASIX, NEBULIZER TREATMENT WAS ADMINISTERED. PULSES CHECKED. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND EPINEPHRINE, VASOPRESSIN, BICARBONATE, AND CALCIUM EMERGENCY RESUSCITATION MEDICATIONS WERE ADMINISTERED. THE PATIENT DIED OF AN UNKNOWN CAUSE.Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED USING INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING. A WATCHMAN ACCESS SYSTEM (WAS) WAS POSITIONED AT THE LAA AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS ADVANCED, DEPLOYED, AND SUBSEQUENTLY IMPLANTED AFTER MEETING ALL RELEASE CRITERIA. PROTAMINE WAS ADMINISTERED FOR ANTICOAGULATION REVERSAL AND THE PROCEDURE WAS CONCLUDED. AFTER THE PATIENT WOKE UP FROM ANESTHESIA, THE PATIENT REPORTED HAVING DIFFICULTY BREATHING. MULTIPLE TRANSTHORACIC ECHOCARDIOGRAM (TTE) IMAGING SWEEPS WERE PERFORMED AND FOUND NO PERICARDIAL EFFUSION AND SHOWED THE CLOSURE DEVICE REMAINED IN THE LAA. INTRAVENOUS LASIX, NEBULIZER TREATMENT WAS ADMINISTERED. PULSES CHECKED. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND EPINEPHRINE, VASOPRESSIN, BICARBONATE, AND CALCIUM EMERGENCY RESUSCITATION MEDICATIONS WERE ADMINISTERED. THE PATIENT DIED OF AN UNKNOWN CAUSE. IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH CERTIFICATE LISTED THE CAUSE OF DEATH AS REACTION TO PROTAMINE MEDICATION POST PROCEDURE. THIS EVENT WAS FURTHER REVIEWED BY BOSTON SCIENTIFIC MEDICAL SAFETY AND WAS DETERMINED TO HAVE BEEN REPORTED AS A REPORTABLE EVENT IN ERROR; THEREFORE, THIS REPORTED EVENT IS WITHDRAWN. Manufacturer Narrative: B5: INFORMATION ADDED. H6 DEVICE CODE CORRECTED TO ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM H6 PATIENT CODE CORRECTED TO NO CLINICAL SIGNS SYMPTOMS OR CONDITIONS H6 IMPACT CODES ALL CORRECTED TO NO HEALTH CONSEQUENCES OR IMPACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106416&pc=DSQ | 2916596-2024-52880 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE CUSTOMER WILL NOT PROVIDE ANY FURTHER INFORMATION REGARDING THE CAUSE OF DEATH. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106426&pc=DSQ | 2916596-2024-52879 | 2023-05-16 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THAT THE PATIENT PASSED AWAY. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: A1-4: PATIENT INFORMATION NOT PROVIDED DUE TO PATIENT PRIVACY LAWS. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106459&pc=DSQ | 2916596-2024-52935 | 2020-09-04 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE DEATH WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(4), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(4) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106496&pc=DSQ | 2916596-2024-52946 | 2024-09-19 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106540&pc=DSQ | 2916596-2024-52944 | 2024-11-18 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106557&pc=CBK | 2518422-2025-100262 | 2021-10-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | CBK | TRILOGY DEVICE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT'S DAUGHTER (REPORTER) CALLED TO REPORT THE DEATH OF HER MOTHER. THE REPORTER STATED THAT HER "MOTHER PASSED AWAY DUE TO THE MACHINE IN 2021. SHE (THE MOTHER) WAS ON AND USING THE DEVICE IN THE MIDDLE OF THE NIGHT, WHEN SHE WENT INTO THE BATHROOM. SHE HAD TROUBLE BREATHING AND FELL OVER AND COULDN'T BREATHE UNTIL THE AMBULANCE TOOK HER TO THE HOSPITAL. THE DAUGHTER STATES SHE HAD STOPPED THE DEVICE FOR A COUPLE WEEKS AND THEN WENT BACK ON IT. THE REPORTER STATED THAT HER MOTHER TOLD HER SHE FELT LIKE SHE COULDN'T BREATHE. SHE HAD PNEUMONIA OVER AND OVER AGAIN. SHE WAS AFRAID TO USE IT. CAUSE OF DEATH WAS MYOCARDIAL INFARCTION, HEART FAILURE, AND CHRONIC KIDNEY DISEASE. THE DAUGHTER ASKED FOR AN AUTOPSY, AND SHE WAS DENIED". THE REPORTER STATES THAT SHE HAD A TRILOGY DEVICE, AND THAT THE DME (APRIA) TOOK THE MACHINE ON (B)(6) 2021. THERE IS NO DEVICE INFORMATION AVAILABLE AT THE TIME OF THIS REPORT. GOOD FAITH EFFORTS HAVE BEEN MADE TO ACQUIRE THE DEVICE DETAILS AND FACILITATE THE RETURN OF THE DEVICE FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106637&pc=NRY | 2029214-2025-00085 | 2024-01-02 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-08 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING: THROMBECTOMY WITH THE PRESET VS SOLITAIRE STENT RETRIEVERS AS FIRST-LINE LARGE VESSEL OCCLUSION STROKE TREATMENT. THE TIME FRAME OF THIS STUDY WAS: OCTOBER 2019 TO FEBRUARY 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: SOLITAIRE STENT RETRIEVER. DEATHS OCCURRED IN THE STUDY POPULATION: YES. MORTALITY OCCURRED IN 24 PATIENTS (14.4%) IN THE SOLITAIRE GROUP AT 90 DAYS. SPECIFIC CAUSES WERE NOT DETAILED, THOUGH IT WAS NOTED THERE WERE 6 CASES OF STROKE-RELATED MORTALITY IN THE SOLITAIRE GROUP. IT WAS NOTED THAT THERE WERE NO OBSERVED DEVICE DEFICIENCIES THAT COULD HAVE LED TO DEVICE-RELATED OR PROCEDURE-RELATED SERIOUS ADVERSE EVENTS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: NOGUEIRA RG, LOBSIEN D, KLISCH J, PIELENZ D, LOBSIEN E, SAUVAGEAU E, AGHAEBRAHIM N, MÖHLENBRUCH M, VOLLHERBST D, ULFERT C, BOZORGCHAMI H, CLARK W, PRIEST R, SAMANIEGO EA, ORTEGA-GUTIERREZ S, GHANNAM M. THROMBECTOMY WITH THE PRESET VS SOLITAIRE STENT RETRIEVERS AS FIRST-LINE LARGE VESSEL OCCLUSION STROKE TREATMENT A RANDOMIZED CLINICAL TRIAL. JAMA NEUROLOGY 81(2):170-178 2024. DOI:10.1001/JAMANEUROL.2023.5010 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106656&pc=BZD | 2518422-2025-000826 | 2024-06-10 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | DREAMSTATION BIPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106660&pc=DSQ | 2916596-2024-52891 | 2024-02-16 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE CUSTOMER STATED THAT NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: SECTION D4, PRIMARY UDI NUMBER: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106682&pc=DSQ | 2916596-2024-52959 | 2023-05-21 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED. IT WAS NOT PROVIDED WHETHER THIS WAS DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: THE PATIENT PRIVACY LAWS OF THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION A2-A4: PATIENT INFORMATION NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107413&pc=OUT | 2029214-2025-00087 | 2024-12-02 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-08 05:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Activation, Positioning or Separation Problem; Adverse Event Without Identified Device or Use Problem; Activation Failure | Anemia; Hemorrhage/Bleeding; Intracranial Hemorrhage; Vasoconstriction; Stenosis; Ventilator Dependent; Bronchopneumonia; Convulsion/Seizure; Thrombosis/Thrombus; Respiratory Insufficiency; Swelling/ Edema | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT A PATIENT TREATED WITH TWO PED2 PIPELINES HAD COMPLICATIONS. THE PATIENT HAD A WORSENING PULMONARY EDEMA ON (B)(6) 2024. TACHYPNEA BEGAN (B)(6) AND PATIENT WAS INTUBATED FOR AIRWAY PROTECTION. GIVEN IV LASIX. (B)(6) CXR CONCERNING FOR PULMONARY EDEMA. CXR REPEAT 12/11 WITH IMPROVED PULMONARY EDEMA 12/13: DESATURATES WHEN REPEATEDLY REPOSITIONED AND DURING PERIODS OF AGITATION, CHEST X-RAY REVIEWED, NOTABLE FOR WORSENING PULMONARY EDEMA. GIVEN LASIX. CONTINUE AGGRESSIVE PULMONARY HYGIENE. 12/16 TOLERATING GENTLE DIURESIS FOR PULMONARY EDEMA. EXTENDED HOSPITALIZATION. POSSIBLY RELATED TO INDEX PROCEDURE. 12/8: CXR WITH MULTIFOCAL CONSOLIDATION AND MEETING ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS) CRITERIA 12/9: LASIX STARTED FOR PULMONARY EDEMA AND ARDS 12/24: CT CHEST SHOWING MULTIFOCAL PNEUMONIA AND ARDS 12/25: TCD MILD INCREASED VELOCITY IN LEFT ANTERIOR CEREBRAL ARTERY (ACA). PRECEDEX WEANED. 12/26: TRANSFUSED 1U PRBC FOR LOW HGB. INCREASED FIO2 DUE TO DYSSYNCHRONOUS BREATHING ON VENT AND DESATURATING. 12/28: REQUIRE HIGH FIO2 DUE TO DESATURATIONS. 12/29: MAINTAINING FIO2 AT 70%. R RADIAL ALINE PLACED. 12/30: FIO2 REQUIREMENT INCREASED TO 100%, TIME OF DEATH 21:17. INCREASED OXYGEN DUE TO DESATURATION; FAMILY CHOSE TO WITHDRAW CARE. CAUSAL RELATIONSHIP WITH STUDY PROCEDURE, RETREATMENT PROCEDURE, AND NON-INDEX PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE CAUSE OF THE DEATH WAS ARDS. IT WAS NOTED THAT THERE WAS TREATMENT DONE FOR THE ARDS. THE PATIENT WAS HOSPITALIZED WITH RUPTURED FUSIFORM VERTEBRAL ARTERY ANEURYSM RESULTED IN SUB ARACHNOID HEMORRHAGE (SAH) IN CIRCLE OF (B)(6). OTHER MHS INCLUDE HYPERTENSION, GERD, OA, RESPIRATORY DISTRESS ON (B)(6)2024. THE INDEX PROCEDURE WAS ON (B)(6)2024. (B)(6)2024, THERE WAS MILD VASOSPASM IN THE RIGHT MCA ARTERY DUE TO ELEVATED MEAN VELOCITIES REPORTED. (B)(6)2024 DSA IMAGING REPORTED DIFFUSED VASOSPASM (DONE DUE TO DECLINE IN MENTAL STATUS). VERAPAMIL WAS GIVEN. THE PATIENT EXPERIENCED SEIZURE ON (B)(6) 2024 AND WAS INTUBATED FOR AIRWAY PROTECTION ON (B)(6)2024. ADDITIONALLY, POST PROCEDURE DEEP VEIN THROMBOSIS (DVT) WAS REPORTED WHICH WAS MANAGED BY IVC FILTER. INTRACRANIAL DOPPLER ON (B)(6)2024 REPORTED PRESUMED EVIDENCE OF MODERATE VASOSPASM IN THE BA AND L PCA, AS WELL AS MILD VASOSPASM IN THE L MCA. CTA IMAGING ON (B)(6)2024 REPORTED SEVERE VASOSPASM IN THE LEFT M1 SEGMENT AND DISTAL BRANCHES HAS PROGRESSED, SEVERE VASOSPASM IN THE ANTERIOR CEREBRAL ARTERIES HAS PROGRESSED, AND MILD VASOSPASM IN THE RIGHT MCA IS IMPROVED. ON (B)(6)2024, ANGIOGRAPHY WAS DONE; VERAPAMIL IA WAS ADMINISTERED AND BALLOON ANGIOPLASTY OF LEFT MCA FOR VASOSPASM WAS DONE. PER REPORT, FOLLOWING THESE INTERVENTION, FLOW IN THE LEFT MCA WAS NORMALIZED. DOPPLER ON (B)(6)2024 REPORTED EVIDENCE OF MILD VASOSPASM IN THE RIGHT PCA. ANOTHER F/U DOPPLER ON (B)(6)2024 REPORTED NO EVIDENCE OF VASOSPASM. ADDITIONALLY, IN COMPARISON TO THE PRIOR CT BRAIN DATED (B)(6)2024, INTERVAL DEVELOPMENT OF RECENT INFARCTS IN THE DISTRIBUTION OF THE LEFT POSTERIOR-INFERIOR CEREBELLAR AND LEFT MIDDLE CEREBRAL ARTERIES. INTERVAL DEVELOPMENT OF A SMALL MIXED DENSITY SUBDURAL HEMATOMA OVERLYING THE RIGHT FRONTOPARIETAL CONVEXITY WITH NO MIDLINE SHIFT, AND PARTIAL RESORPTION OF THE SUBARACHNOID AND INTRAVENTRICULAR HEMORRHAGE. THERE WAS MILD DECREASE IN THE DILATATION OF THE LATERAL AND THIRD VENTRICLES. THE STATUS POST PLACEMENT OF A PIPELINE STENT ACROSS THE LEFT V4 SEGMENT. THE STENT WAS PATENT WITH NO DEFINITE RESIDUAL FILLING OF THE LEFT VERTEBRAL ARTERY ANEURYSM. THE PATIENT PASSED AWAY AT THE MEDICAL CENTER. THE SITE UPDATED THEIR ASSESSMENT OF THE PULMONARY EDEMA TO POSSIBLY RELATED TO THE DISEASE UNDER STUDY, AND THE ASSESSMENT OF THE ARDS WAS NOT RELATED TO A NON-INDEX OR RETREATMENT PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT A PIPELINE PATIENT EXPERIENCED ANEMIA. ACTION WAS TAKEN: DAILY CBC, TRANSFUSION FOR HGB 7, AND TS COMPLETED. ON 26-DEC. ONE UNIT OF RBC WAS TRANSFUSED. THE EVENT WAS CAUSALLY RELATED TO THE DISEASE UNDER STUDY. IT WAS NOTED THAT THE DEVICE WAS INSERTED INTO THE MICROCATHETER AND WAS RESHEATHED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PATIENT WAS TREATED WITH A BLOOD TRANSFUSION. DAILY CBC, TRANSFUSION FOR HGB <(><<)>7, AND T <(>&<)>S WAS COMPLETED ON THE (B)(6), RBC WAS TRANSFUSED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT 40MG OF LASIX WAS ADMINISTERED FROM (B)(6) 2024, AND 355,000MG OF PACKED RED BLOOD CELLS WERE GIVEN FROM (B)(6) 2024.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT AFTER DEPLOYMENT OF THE PIPELINE SHIELD, THE DEVICE HAD PROXIMAL INCOMPLETE WALL APPOSITION DUE TO INCOMPLETE OPENING OF THE DEVICE. THIS WAS TREATED WITH BALLOON ANGIOPLASTY WITH GOOD RESULTS, AND FULL WALL APPOSITION WAS ACHIEVED. ANGIOGRAPHIC RESULTS POST PROCEDURE SHOWED COMPLETE OCCLUSION OF THE TARGET ANEURYSM. MORE THAN 50% HAD BEEN DEPLOYED WHEN IT FAILED TO OPEN. IT WAS NOTED THAT PATIENT DIED ON (B)(6) 2024 FROM ACUTE RESPIRATORY DISTRESS SYNDROME.Event Description: ADDITIONAL INFORMATION WAS RECEIVED REPORTING CT ON (B)(6) 2024 SHOWED AN INCREASING RIGHT FRONTOPARIETAL SUBDURAL HEMORRHAGE. AE DID NOT RESULT FROM A DEVICE DEFICIENCY. AE HAD A CAUSAL RELATIONSHIP WITH THE DISEASE UNDER STUDY. OUTCOME STATUS WAS NOT RECOVERED/NOT RESOLVED. SITE ASSESSED AE AS NOT RELATED TO APT OR STUDY DEVICE, CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED REPORTING POSTERIOR INFERIOR CEREBRAL ARTERY STENOSIS. PER THIRD PARTY CORE LAB, THE SUBJECT HAD 75-95% STENOSIS OF THE POSTERIOR INFERIOR CEREBELLAR ARTERY (PICA) ON DSA OBTAINED ON (B)(6) 2024. AE RESULT FROM DEVICE DEFICIENCY OF INCOMPLETE OPEN, POSSIBLY RELATED TO THE DISEASE UNDER STUDY. OUTCOME STATUS WAS NOT RECOVERED/NOT RESOLVED. THE SITE ASSESSED THIS EVENT AS POSSIBLY RELATED TO THE INDEX AND RETREATMENT PROCEDURES, POSSIBLY RELATED TO THE DEVICE VANTAGE. THESE ADVERSE EVENTS (AE) OCCURRED POST-PROCEDURE, AND WAS NOT ASSOCIATED WITH SAH IN EVOLUTION. THERE WAS NO DECLINE IN MRS OR NEW OR WORSENING HEADACHE. EVENT WAS NOT ASSOCIATED WITH NON-TARGET ANEURYSM TREATMENT. SITE ASSESSED THESE AES DID NOT LEAD TO CONGENITAL ANOMALY, DEATH, DISABILITY, HOSPITALIZATION, OR MEDICAL INTERVENTION, AND WAS NOT CONSIDERED LIFE-THREATENING.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED INFORMATION WAS CORRECTED TO FOLLOW UP DSA IMAGING OBTAINED (B)(6) 2024, THE SUBJECT HAD 75-95% STENOSIS OF THE POSTERIOR INFERIOR CEREBELLAR ARTERY (PICA). THIS WAS A BRANCH ARTERY THAT WAS COVERED BY THE STUDY DEVICE DURING THE INDEX PROCEDURE AND THEREFORE RELATED TO THE STUDY DEVICE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT ADVERSE EVENT OF WORSENING PULMONARY EDEMA RESULTED IN PROLONGED PATIENT HOSPITALIZATION FROM (B)(6) 2024. THIS SITE ASSESSED THIS EVENT AS NOT RELATED TO THE PROCEDURE OR DEVICE BUT THE STUDY SPONSOR CONSERVATIVELY ASSESSED THE EVENT AS POSSIBLY PROCEDURE RELATED, NOT DIRECTLY RELATED TO ANY MEDTRONIC DEVICE. THE ASSESSMENT OF THE INCREASING RIGHT FRONTOPARIETAL SUBDURAL HEMORRHAGE ADVERSE EVENT WAS ADJUDICATED TO HAVING A CAUSAL RELATIONSHIP TO THE ANTIPLATELET MEDICATION.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE SITE UPDATED THEIR ASSESSMENT OF THE SAH AS POSSIBLY RELATED TO THE DISEASE UNDER STUDY AND ANTIPLATELET THERAPY (APT). THE SPONSOR ASSESSED THE SAH AS POSSIBLY RELATED TO APT. THE SAH WAS NOT IN THE TREATED VASCULAR TERRITORY, AND WAS GUSTO MILD AND BARC TYPE 1. THE SPONSOR UPDATED THEIR ASSESSMENT OF THE PULMONARY EDEMA TO NOT RELATED TO THE DEVICE, PROCEDURE, OR APT. THE SPONSOR ASSESSED THE ANEMIA AS POSSIBLY RELATED TO APT BUT NOT RELATED TO INSUFFICIENT APT. THE SPONSOR ASSESSED THE DEATH WITHIN ONE MONTH OF THE PROCEDURE AS RELATED TO THE DISEASE UNDER STUDY, AND AS A NON-NEUROLOGICAL DEATH. THE PICA STENOSIS WAS NOTED TO BE POD5, POD 18 SEVERE DIFFUSE VASOSPASM. PICA INFARCT WAS NOTED, SO THE STENOSIS MAY HAVE CONTRIBUTED TO THE DELAYED CEREBRAL ISCHEMIA (DCI). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107415&pc=OUT | 2029214-2025-00088 | 2024-12-02 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-08 05:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Anemia; Hemorrhage/Bleeding; Intracranial Hemorrhage; Vasoconstriction; Stenosis; Ventilator Dependent; Bronchopneumonia; Convulsion/Seizure; Thrombosis/Thrombus; Respiratory Insufficiency; Swelling/ Edema | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT A PATIENT TREATED WITH TWO PED2 PIPELINES HAD COMPLICATIONS. THE PATIENT HAD A WORSENING PULMONARY EDEMA ON (B)(6) 2024. TACHYPNEA BEGAN (B)(6) AND PATIENT WAS INTUBATED FOR AIRWAY PROTECTION. GIVEN IV LASIX. (B)(6) CXR CONCERNING FOR PULMONARY EDEMA. CXR REPEAT (B)(6) WITH IMPROVED PULMONARY EDEMA (B)(6): DESATURATES WHEN REPEATEDLY REPOSITIONED AND DURING PERIODS OF AGITATION, CHEST X-RAY REVIEWED, NOTABLE FOR WORSENING PULMONARY EDEMA. GIVEN LASIX. CONTINUE AGGRESSIVE PULMONARY HYGIENE. (B)(6) TOLERATING GENTLE DIURESIS FOR PULMONARY EDEMA. EXTENDED HOSPITALIZATION. POSSIBLY RELATED TO INDEX PROCEDURE. (B)(6): CXR WITH MULTIFOCAL CONSOLIDATION AND MEETING ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS) CRITERIA (B)(6): LASIX STARTED FOR PULMONARY EDEMA AND ARDS (B)(6): CT CHEST SHOWING MULTIFOCAL PNEUMONIA AND ARDS (B)(6): TCD MILD INCREASED VELOCITY IN LEFT ANTERIOR CEREBRAL ARTERY (ACA). PRECEDEX WEANED. (B)(6): TRANSFUSED 1U PRBC FOR LOW HGB. INCREASED FIO2 DUE TO DYSSYNCHRONOUS BREATHING ON VENT AND DESATURATING. (B)(6): REQUIRE HIGH FIO2 DUE TO DESATURATIONS. (B)(6): MAINTAINING FIO2 AT 70%. R RADIAL ALINE PLACED. (B)(6): FIO2 REQUIREMENT INCREASED TO 100%, TIME OF DEATH 21:17. INCREASED OXYGEN DUE TO DESATURATION; FAMILY CHOSE TO WITHDRAW CARE. CAUSAL RELATIONSHIP WITH STUDY PROCEDURE, RETREATMENT PROCEDURE, AND NON-INDEX PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE CAUSE OF THE DEATH WAS ARDS. IT WAS NOTED THAT THERE WAS TREATMENT DONE FOR THE ARDS. THE PATIENT WAS HOSPITALIZED WITH RUPTURED FUSIFORM VERTEBRAL ARTERY ANEURYSM RESULTED IN SUB ARACHNOID HEMORRHAGE (SAH) IN CIRCLE OF (B)(6). OTHER MHS INCLUDE HYPERTENSION, GERD, OA, RESPIRATORY DISTRESS ON (B)(6)2024. THE INDEX PROCEDURE WAS ON (B)(6)2024. (B)(6)2024, THERE WAS MILD VASOSPASM IN THE RIGHT MCA ARTERY DUE TO ELEVATED MEAN VELOCITIES REPORTED. (B)(6)2024 DSA IMAGING REPORTED DIFFUSED VASOSPASM (DONE DUE TO DECLINE IN MENTAL STATUS). VERAPAMIL WAS GIVEN. THE PATIENT EXPERIENCED SEIZURE ON (B)(6) 2024 AND WAS INTUBATED FOR AIRWAY PROTECTION ON (B)(6)2024. ADDITIONALLY, POST PROCEDURE DEEP VEIN THROMBOSIS (DVT) WAS REPORTED WHICH WAS MANAGED BY IVC FILTER. INTRACRANIAL DOPPLER ON (B)(6)2024 REPORTED PRESUMED EVIDENCE OF MODERATE VASOSPASM IN THE BA AND L PCA, AS WELL AS MILD VASOSPASM IN THE L MCA. CTA IMAGING ON (B)(6) 2024 REPORTED SEVERE VASOSPASM IN THE LEFT M1 SEGMENT AND DISTAL BRANCHES HAS PROGRESSED, SEVERE VASOSPASM IN THE ANTERIOR CEREBRAL ARTERIES HAS PROGRESSED, AND MILD VASOSPASM IN THE RIGHT MCA IS IMPROVED. ON (B)(6)2024, ANGIOGRAPHY WAS DONE; VERAPAMIL IA WAS ADMINISTERED AND BALLOON ANGIOPLASTY OF LEFT MCA FOR VASOSPASM WAS DONE. PER REPORT, FOLLOWING THESE INTERVENTION, FLOW IN THE LEFT MCA WAS NORMALIZED. DOPPLER ON (B)(6)2024 REPORTED EVIDENCE OF MILD VASOSPASM IN THE RIGHT PCA. ANOTHER F/U DOPPLER ON (B)(6)2024 REPORTED NO EVIDENCE OF VASOSPASM. ADDITIONALLY, IN COMPARISON TO THE PRIOR CT BRAIN DATED (B)(6)2024, INTERVAL DEVELOPMENT OF RECENT INFARCTS IN THE DISTRIBUTION OF THE LEFT POSTERIOR-INFERIOR CEREBELLAR AND LEFT MIDDLE CEREBRAL ARTERIES. INTERVAL DEVELOPMENT OF A SMALL MIXED DENSITY SUBDURAL HEMATOMA OVERLYING THE RIGHT FRONTOPARIETAL CONVEXITY WITH NO MIDLINE SHIFT, AND PARTIAL RESORPTION OF THE SUBARACHNOID AND INTRAVENTRICULAR HEMORRHAGE. THERE WAS MILD DECREASE IN THE DILATATION OF THE LATERAL AND THIRD VENTRICLES. THE STATUS POST PLACEMENT OF A PIPELINE STENT ACROSS THE LEFT V4 SEGMENT. THE STENT WAS PATENT WITH NO DEFINITE RESIDUAL FILLING OF THE LEFT VERTEBRAL ARTERY ANEURYSM. THE PATIENT PASSED AWAY AT THE MEDICAL CENTER. THE SITE UPDATED THEIR ASSESSMENT OF THE PULMONARY EDEMA TO POSSIBLY RELATED TO THE DISEASE UNDER STUDY, AND THE ASSESSMENT OF THE ARDS WAS NOT RELATED TO A NON-INDEX OR RETREATMENT PROCEDURE.Event Description: 2024-DEC-18 LSH 1295031620 X3, LSH 1303461980 X2 (HCP, SDY): MEDTRONIC RECEIVED INFORMATION THAT A PATIENT TREATED WITH TWO PED2 PIPELINES HAD COMPLICATIONS. THE PATIENT HAD A WORSENING PULMONARY EDEMA ON (B)(6) 2024. TACHYPNEA BEGAN (B)(6) AND PATIENT WAS INTUBATED FOR AIRWAY PROTECTION. GIVEN IV LASIX. (B)(6) CXR CONCERNING FOR PULMONARY EDEMA. CXR REPEAT (B)(6) WITH IMPROVED PULMONARY EDEMA (B)(6): DESATURATES WHEN REPEATEDLY REPOSITIONED AND DURING PERIODS OF AGITATION, CHEST X-RAY REVIEWED, NOTABLE FOR WORSENING PULMONARY EDEMA. GIVEN LASIX. CONTINUE AGGRESSIVE PULMONARY HYGIENE. (B)(6) TOLERATING GENTLE DIURESIS FOR PULMONARY EDEMA. EXTENDED HOSPITALIZATION. POSSIBLY RELATED TO INDEX PROCEDURE. (B)(6): CXR WITH MULTIFOCAL CONSOLIDATION AND MEETING ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS) CRITERIA (B)(6): LASIX STARTED FOR PULMONARY EDEMA AND ARDS (B)(6): CT CHEST SHOWING MULTIFOCAL PNEUMONIA AND ARDS (B)(6): TCD MILD INCREASED VELOCITY IN LEFT ANTERIOR CEREBRAL ARTERY (ACA). PRECEDEX WEANED. (B)(6): TRANSFUSED 1U PRBC FOR LOW HGB. INCREASED FIO2 DUE TO DYSSYNCHRONOUS BREATHING ON VENT AND DESATURATING. (B)(6): REQUIRE HIGH FIO2 DUE TO DESATURATIONS. (B)(6): MAINTAINING FIO2 A T 70%. R RADIAL ALINE PLACED. (B)(6): FIO2 REQUIREMENT INCREASED TO 100%, TIME OF DEATH 21:17. INCREASED OXYGEN DUE TO DESATURATION; FAMILY CHOSE TO WITHDRAW CARE. CAUSAL RELATIONSHIP WITH STUDY PROCEDURE, RETREATMENT PROCEDURE, AND NON-INDEX PROCEDURE. (B)(6) 2025 MPXR 1263208, LSH 1303461980 UPDATE, LSH 1295031620 UPDATE, E1 (FOR, SDY, HCP): ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE CAUSE OF THE DEATH WAS ARDS. IT WAS NOTED THAT THERE WAS TREATMENT DONE FOR THE ARDS. THE PATIENT WAS HOSPITALIZED WITH RUPTURED FUSIFORM VERTEBRAL ARTERY ANEURYSM RESULTED IN SUB ARACHNOID HEMORRHAGE (SAH) IN CIRCLE OF WILLIS. OTHER MHS INCLUDE HYPERTENSION, GERD, OA, RESPIRATORY DISTRESS ON (B)(6) 2024. THE INDEX PROCEDURE WAS ON (B)(6) 2024, THERE WAS MILD VASOSPASM IN THE RIGHT MCA ARTERY DUE TO ELEVATED MEAN VELOCITIES REPORTED. (B)(6) 2024 DSA IMAGING REPORTED DIFFUSED VASOSPASM (DONE DUE TO DECLINE IN MENTAL STATUS). VERAPAMIL WAS GIVEN. THE PATIENT EXPERIENCED SEIZURE ON (B)(6) 2024, AND WAS INTUBATED FOR AIRWAY PROTECTION ON (B)(6) 2024. ADDITIONALLY, POST PROCEDURE DEEP VEIN THROMBOSIS (DVT) WAS REPORTED WHICH WAS MANAGED BY IVC FILTER. INTRACRANIAL DOPPLER ON (B)(6) 2024 REPORTED PRESUMED EVIDENCE OF MODERATE VASOSPASM IN THE BA AND L PCA, AS WELL AS MILD VASOSPASM IN THE L MCA. CTA IMAGING ON (B)(6) 2024 REPORTED SEVERE VASOSPASM IN THE LEFT M1 SEGMENT AND DISTAL BRANCHES HAS PROGRESSED, SEVERE VASOSPASM IN THE ANTERIOR CEREBRAL ARTERIES HAS PROGRESSED, AND MILD VASOSPASM IN THE RIGHT MCA IS IMPROVED. ON (B)(6) 2024, ANGIOGRAPHY WAS DONE; VERAPAMIL IA WAS ADMINISTERED AND BALLOON ANGIOPLASTY OF LEFT MCA FOR VASOSPASM WAS DONE. PER REPORT, FOLLOWING THESE INTERVENTION, FLOW IN THE LEFT MCA WAS NORMALIZED. DOPPLER ON (B)(6) 2024 REPORTED EVIDENCE OF MILD VASOSPASM IN THE RIGHT PCA. ANOTHER F/U DOPPLER ON (B)(6) 2024 REPORTED NO EVIDENCE OF VASOSPASM. ADDITIONALLY, IN COMPARISON TO THE PRIOR CT BRAIN DATED (B)(6) 2024, INTERVAL DEVELOPMENT OF RECENT INFARCTS IN THE DISTRIBUTION OF THE LEFT POSTERIOR-INFERIOR CEREBELLAR AND LEFT MIDDLE CEREBRAL ARTERIES. INTERVAL DEVELOPMENT OF A SMALL MIXED DENSITY SUBDURAL HEMATOMA OVERLYING THE RIGHT FRONTOPARIETAL CONVEXITY WITH NO MID LINE SHIFT, AND PARTIAL RESORPTION OF THE SUBARACHNOID AND INTRAVENTRICULAR HEMORRHAGE. THERE WAS MILD DECREASE IN THE DILATATION OF THE LATERAL AND THIRD VENTRICLES. THE STATUS POST PLACEMENT OF A PIPELINE STENT ACROSS THE LEFT V4 SEGMENT. THE STENT WAS PATENT WITH NO DEFINITE RESIDUAL FILLING OF THE LEFT VERTEBRAL ARTERY ANEURYSM. THE PATIENT PASSED AWAY AT THE MEDICAL CENTER. THE SITE UPDATED THEIR ASSESSMENT OF THE PULMONARY EDEMA TO POSSIBLY RELATED TO THE DISEASE UNDER STUDY, AND THE ASSESSMENT OF THE ARDS WAS NOT RELATED TO A NON-INDEX OR RETREATMENT PROCEDURE. (B)(6) 2025 LSH 1316120103 (REP, HCP, SDY): ADDITIONAL INFORMATION RECEIVED REPORTED THAT A PIPELINE PATIENT EXPERIENCED ANEMIA. ACTION WAS TAKEN: DAILY CBC, TRANSFUSION FOR HGB 7, AND TS COMPLETED. ON (B)(6). ONE UNIT OF RBC WAS TRANSFUSED. THE EVENT WAS CAUSALLY RELATED TO THE DISEASE UNDER STUDY. IT WAS NOTED THAT THE DEVICE WAS INSERTED INTO THE MICROCATHETER AND WAS RESHEATHED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT 40MG OF LASIX WAS ADMINISTERED FROM (B)(6) 2024, AND 355,000MG OF PACKED RED BLOOD CELLS WERE GIVEN FROM (B)(6) 2024.Event Description: ADDITIONAL INFORMATION WAS RECEIVED REPORTING CT ON (B)(6) 2024 SHOWED AN INCREASING RIGHT FRONTOPARIETAL SUBDURAL HEMORRHAGE. AE DID NOT RESULT FROM A DEVICE DEFICIENCY. AE HAD A CAUSAL RELATIONSHIP WITH THE DISEASE UNDER STUDY. OUTCOME STATUS WAS NOT RECOVERED/NOT RESOLVED. SITE ASSESSED AE AS NOT RELATED TO APT OR STUDY DEVICE, CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED REPORTING POSTERIOR INFERIOR CEREBRAL ARTERY STENOSIS. PER THIRD PARTY CORE LAB, THE SUBJECT HAD 75-95% STENOSIS OF THE POSTERIOR INFERIOR CEREBELLAR ARTERY (PICA) ON DSA OBTAINED ON (B)(6) 2024. AE RESULT FROM DEVICE DEFICIENCY OF INCOMPLETE OPEN, POSSIBLY RELATED TO THE DISEASE UNDER STUDY. OUTCOME STATUS WAS NOT RECOVERED/NOT RESOLVED. THE SITE ASSESSED THIS EVENT AS POSSIBLY RELATED TO THE INDEX AND RETREATMENT PROCEDURES, POSSIBLY RELATED TO THE DEVICE VANTAGE. THESE ADVERSE EVENTS (AE) OCCURRED POST-PROCEDURE AND WAS NOT ASSOCIATED WITH SAH IN EVOLUTION. THERE WAS NO DECLINE IN MRS OR NEW OR WORSENING HEADACHE. EVENT WAS NOT ASSOCIATED WITH NON-TARGET ANEURYSM TREATMENT. SITE ASSESSED THESE AES DID NOT LEAD TO CONGENITAL ANOMALY, DEATH, DISABILITY, HOSPITALIZATION, OR MEDICAL INTERVENTION, AND WAS NOT CONSIDERED LIFE-THREATENING.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PATIENT WAS TREATED WITH A BLOOD TRANSFUSION. DAILY CBC, TRANSFUSION FOR HGB, 7, AND T & S WAS COMPLETED ON THE 9TH, 12TH, 13TH, AND 26TH OF DECEMBER, RBC WAS TRANSFUSED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED INFORMATION WAS CORRECTED TO FOLLOW UP DSA IMAGING OBTAINED 07-DEC-2024, THE SUBJECT HAD 75-95% STENOSIS OF THE POSTERIOR INFERIOR CEREBELLAR ARTERY (PICA). THIS WAS A BRANCH ARTERY THAT WAS COVERED BY THE STUDY DEVICE DURING THE INDEX PROCEDURE AND THEREFORE RELATED TO THE STUDY DEVICE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT ADVERSE EVENT OF WORSENING PULMONARY EDEMA RESULTED IN PROLONGED PATIENT HOSPITALIZATION FROM (B)(6) 2024. THIS SITE ASSESSED THIS EVENT AS NOT RELATED TO THE PROCEDURE OR DEVICE BUT THE STUDY SPONSOR CONSERVATIVELY ASSESSED THE EVENT AS POSSIBLY PROCEDURE RELATED, NOT DIRECTLY RELATED TO ANY MEDTRONIC DEVICE. THE ASSESSMENT OF THE INCREASING RIGHT FRONTOPARIETAL SUBDURAL HEMORRHAGE ADVERSE EVENT WAS ADJUDICATED TO HAVING A CAUSAL RELATIONSHIP TO THE ANTIPLATELET MEDICATION.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE SITE UPDATED THEIR ASSESSMENT OF THE SAH AS POSSIBLY RELATED TO THE DISEASE UNDER STUDY AND ANTIPLATELET THERAPY (APT). THE SPONSOR ASSESSED THE SAH AS POSSIBLY RELATED TO APT. THE SAH WAS NOT IN THE TREATED VASCULAR TERRITORY, AND WAS GUSTO MILD AND BARC TYPE 1. THE SPONSOR UPDATED THEIR ASSESSMENT OF THE PULMONARY EDEMA TO NOT RELATED TO THE DEVICE, PROCEDURE, OR APT. THE SPONSOR ASSESSED THE ANEMIA AS POSSIBLY RELATED TO APT BUT NOT RELATED TO INSUFFICIENT APT. THE SPONSOR ASSESSED THE DEATH WITHIN ONE MONTH OF THE PROCEDURE AS RELATED TO THE DISEASE UNDER STUDY, AND AS A NON-NEUROLOGICAL DEATH. THE PICA STENOSIS WAS NOTED TO BE POD5, POD 18 SEVERE DIFFUSE VASOSPASM. PICA INFARCT WAS NOTED, SO THE STENOSIS MAY HAVE CONTRIBUTED TO THE DELAYED CEREBRAL ISCHEMIA (DCI). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107602&pc=DSQ | 2916596-2024-52841 | 2024-12-12 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN INTRA-OPERATIVE STROKE WHICH THE PATIENT DID NOT RECOVER FROM NEUROLOGICALLY. THE DEATH WAS NOT CONSIDERED DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD A STABLE LARGE LEFT CEREBRAL HEMISPHERIC INFARCTION WITHOUT HEMORRHAGIC TRANSFORMATION. THERE WAS A 4-MILLIMETER LEFT-TO-RIGHT MIDLINE SHIFT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107616&pc=NRY | 2029214-2025-00089 | 2024-03-29 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-08 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke |  |  | 1 | Event Description: THE LITERATURE REVIEWED IN THE STUDY AIMED TO COMPARE BALLOON GUIDE CATHETERS (BGCS) VERSUS NON-BALLOON GUIDE CATHETERS (NBGCS) AS A PART OF A COMBINED TREATMENT MODALITY IN PATIENTS PRESENTING WITH ACUTE ISCHEMIC STROKE. THE STUDY WAS CONDUCTED OVER A PERIOD FROM BETWEEN APRIL 2018 AND OCTOBER 2020. AMONG THE DEVICES USED, THE STUDY SPECIFICALLY MENTIONED THE UTILIZATION OF MEDTRONIC'S SOLITAIRE STENT AND CELLO AND MARKSMAN CATHETERS. NO PROCEDURE-RELATED MORTALITY WAS OBSERVED; HOWEVER, 7 (25.9%) PATIENTS IN THE BGC GROUP DIED BEFORE DISCHARGE COMPARED TO 16 (42.1%) IN THE NBGC GROUP. THE TWO GROUPS DIDN¿T DIFFER IN TERMS OF THE NIHSS SCORE CHANGE COMPARED TO BASELINE. MOREOVER, THERE WERE NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS IN TERMS OF THE SUCCESSFUL RECANALIZATION RATE, THREE-MONTH FAVORABLE FUNCTIONAL OUTCOME RATE, AND IN-HOSPITAL MORTALITY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-MARKSMAN (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/APRODUCT ID UNK-NV-CELLO (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/AG2: CITATION: AUTHORS: UYSAL E, BODELSCHWINGH B, TABAKCI O, BASARIR C, BULUT S. COMBINED ASPIRATION AND STENT RETRIEVER THROMBECTOMY FOR DISTAL CAROTID ARTERY OCCLUSION USING BALLOON GUIDE VERSUS NON-BALLOON GUIDE CATHETER. JOURNAL OF CLINICAL MEDICINE 13, 1978 2024. 10.3390/JCM13071978EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107626&pc=KRA | 2029214-2025-00090 | 2024-03-29 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-08 05:00 | KRA | MARKSMAN | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke |  |  | 1 | Event Description: THE LITERATURE REVIEWED IN THE STUDY AIMED TO COMPARE BALLOON GUIDE CATHETERS (BGCS) VERSUS NON-BALLOON GUIDE CATHETERS (NBGCS) AS A PART OF A COMBINED TREATMENT MODALITY IN PATIENTS PRESENTING WITH ACUTE ISCHEMIC STROKE. THE STUDY WAS CONDUCTED OVER A PERIOD FROM BETWEEN APRIL 2018 AND OCTOBER 2020. AMONG THE DEVICES USED, THE STUDY SPECIFICALLY MENTIONED THE UTILIZATION OF MEDTRONIC'S SOLITAIRE STENT AND CELLO AND MARKSMAN CATHETERS. NO PROCEDURE-RELATED MORTALITY WAS OBSERVED; HOWEVER, 7 (25.9%) PATIENTS IN THE BGC GROUP DIED BEFORE DISCHARGE COMPARED TO 16 (42.1%) IN THE NBGC GROUP. THE TWO GROUPS DIDN¿T DIFFER IN TERMS OF THE NIHSS SCORE CHANGE COMPARED TO BASELINE. MOREOVER, THERE WERE NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS IN TERMS OF THE SUCCESSFUL RECANALIZATION RATE, THREE-MONTH FAVORABLE FUNCTIONAL OUTCOME RATE, AND IN-HOSPITAL MORTALITY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: UNK-NV-SOLITAIRE (UNKNOWN); IMPLANT DATE N/A; EXPLANT DATE: N/A, PRODUCT ID: UNK-NV-CELLO (UNKNOWN); IMPLANT DATE N/A; EXPLANT DATE N/A. G2: CITATION: AUTHORS: UYSAL E, BODELSCHWINGH B, TABAKCI O, BASARIR C, BULUT S. COMBINED ASPIRATION AND STENT RETRIEVER THROMBECTOMY FOR DISTAL CAROTID ARTERY OCCLUSION USING BALLOON GUIDE VERSUS NON-BALLOON GUIDE CATHETER. JOURNAL OF CLINICAL MEDICINE 13, 1978 2024. 10.3390/JCM13071978 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107666&pc=DQY | 2029214-2025-00091 | 2024-03-29 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-08 05:00 | DQY | CELLO BALLOON GUIDE CATHETER | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke |  |  | 1 | Event Description: THE LITERATURE REVIEWED IN THE STUDY AIMED TO COMPARE BALLOON GUIDE CATHETERS (BGCS) VERSUS NON-BALLOON GUIDE CATHETERS (NBGCS) AS A PART OF A COMBINED TREATMENT MODALITY IN PATIENTS PRESENTING WITH ACUTE ISCHEMIC STROKE. THE STUDY WAS CONDUCTED OVER A PERIOD FROM BETWEEN APRIL 2018 AND OCTOBER 2020. AMONG THE DEVICES USED, THE STUDY SPECIFICALLY MENTIONED THE UTILIZATION OF MEDTRONIC'S SOLITAIRE STENT AND CELLO AND MARKSMAN CATHETERS. NO PROCEDURE-RELATED MORTALITY WAS OBSERVED; HOWEVER, 7 (25.9%) PATIENTS IN THE BGC GROUP DIED BEFORE DISCHARGE COMPARED TO 16 (42.1%) IN THE NBGC GROUP. THE TWO GROUPS DIDN¿T DIFFER IN TERMS OF THE NIHSS SCORE CHANGE COMPARED TO BASELINE. MOREOVER, THERE WERE NO SIGNIFICANT DIFFERENCES BETWEEN THE TWO GROUPS IN TERMS OF THE SUCCESSFUL RECANALIZATION RATE, THREE-MONTH FAVORABLE FUNCTIONAL OUTCOME RATE, AND IN-HOSPITAL MORTALITY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); IMPLANT DATE N/A; EXPLANT DATE N/A; PRODUCT ID UNK-NV-MARKSMAN (UNKNOWN); IMPLANT DATE N/A; EXPLANT DATE N/A. G2: CITATION: AUTHORS: UYSAL E, BODELSCHWINGH B, TABAKCI O, BASARIR C, BULUT S. COMBINED ASPIRATION AND STENT RETRIEVER THROMBECTOMY FOR DISTAL CAROTID ARTERY OCCLUSION USING BALLOON GUIDE VERSUS NON-BALLOON GUIDE CATHETER. JOURNAL OF CLINICAL MEDICINE 13, 1978 2024. 10.3390/JCM13071978 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107792&pc=DSQ | 2916596-2024-52919 | 2024-12-15 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Purulent Discharge; Unspecified Infection; Sepsis; Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. IT WAS STATED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND WAS DUE TO SEPSIS. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: IT WAS NOTED THAT THE PATIENT HAD MOVED FROM THEIR ORIGINAL IMPLANTATION SITE TO A NEW SITE. ONE WEEK AFTER MOVING THE PATIENT PASSED AWAY FROM SEPSIS THAT WAS MOST PROBABLY FROM URINARY INFECTIONS.Event Description: THE CAUSE OF DEATH WAS NOT CONFORMED TO BE DUE TO SEPSIS FROM URINARY INFECTION. THE CAUSE OF THE PATIENT'S ARREST AND HEMODYNAMIC COMPROMISE IN THE SETTING OF LEUKOCYTOSIS AND PURULENT DRAINAGE, FUNCTIONING VENTRICULAR ASSIST DEVICE (VAD), COLLAPSIBLE INFERIOR VENA CAVA (IVC) AND RETURN OF IMPROVED BLOOD PRESSURE AND HEART RATE WITH ADMINISTRATION OF EPINEPHRINE BOLUSES WAS NOT CLEAR. THE PATIENT MOST LIKELY EXPERIENCED DISTRIBUTIVE SHOCK. THE ONSET OF THE INFECTION WAS UNKNOWN. BLOOD CULTURES WERE ORDERED BUT THE RESULTS WERE NOT DOCUMENTED. THE PATIENT PASSED AWAY IN THE EMERGENCY DEPARTMENT BEFORE THE INFECTION WAS TREATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, A SPECIFIC CAUSE FOR THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED (INCLUDING PRODUCTION ORDERS AS WELL AS RECORDS IN THE MANUFACTURING EXECUTION SYSTEM) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS SEPSIS, DRIVELINE INFECTION, LOCALIZED INFECTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION INCLUDES INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE AS WELL AS INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION AS WELL AS INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107939&pc=DSQ | 2916596-2024-52904 | 2024-02-01 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED BY THE CUSTOMER. IT WAS NOT PROVIDED WHETHER THIS WAS DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. PATIENT PRIVACY LAWS WOULD NOT ALLOW THE COMMUNICATION OF PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT'S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF PATIENT DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE ACCOUNT DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4 PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107965&pc=DSQ | 2916596-2024-52914 | 2024-08-10 04:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOT PROVIDED WHETHER OR NOT THE DEATH WAS DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: THE PATIENT PRIVACY LAWS OF THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: A4 PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108033&pc=DSQ | 2916596-2025-00092 | 2023-12-17 05:00 | Death | THORATEC CORPORATION | 2025-01-08 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OUTCOME THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. IT STATED THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE CUSTOMER STATED THAT NO FURTHER INFORMATION WAS AVAILABLE. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108177&pc=DYB | 3007284313-2025-03698 | 2024-12-20 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-08 05:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Patient Device Interaction Problem | Perforation of Vessels; Vascular Dissection | K160254 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2024, THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT USING DEBRANCHING AND CHIMNEY TECHNIQUE FOR A DISTAL ARCH AORTIC ANEURYSM WITH GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM, GORE® EXCLUDER® AAA ENDOPROSTHESIS, AND GORE® DRYSEAL FLEX INTRODUCER SHEATH. AFTER DEBRANCHING, THE LEFT SUBCLAVIAN ARTERY WAS PLUG EMBOLISED AND A TGM343415J WAS IMPLANTED DISTALLY. A TGM404020J WAS THEN INSERTED, BUT TGM404020J WAS NOT DEPLOYED AT THIS POINT AS IT WAS TO BE DEPLOYED SIMULTANEOUSLY WITH A CONTRALATERAL LEG ENDOPROSTHESIS (PLC161200J) TO BE IMPLANTED IN THE BRACHIOCEPHALIC ARTERY. AN 8FR NON-GORE SHEATH, A RADIFOCUS¿ GUIDEWIRE, AND AN ANGIOGRAPHIC CATHETER WERE INSERTED FROM THE RIGHT SUBCLAVIAN ARTERY AND SUCCESSFULLY PASSED TO THE ASCENDING AORTA, BUT THE BRACHIOCEPHALIC ARTERY WAS HIGHLY TORTUOUS AND REMAINED TORTUOUS EVEN AFTER THE GUIDEWIRE WAS REPLACED WITH A STIFF TYPE. THE SHEATH WAS REPLACED WITH A DSF1233 AND INSERTED, WHICH PASSED THROUGH WITHOUT RESISTANCE, BUT THE BRACHIOCEPHALIC ARTERY REMAINED TORTUOUS. AN ATTEMPT WAS MADE TO UNSHEATHE THE PLC161200J AFTER DELIVERY TO THE ASCENDING AORTA THROUGH THE DSF1233, BUT THE DSF1233 COULD NOT BE FULLY WITHDRAWN, POSSIBLY BECAUSE THE DSF1233 WAS FOLLOWING THE TORTUOSITY, AND THE PLC161200J WAS RETURNED TO THE DSF1233. THE DSF1233 AND THE PLC161200J WERE PUSHED WITH FORCE AND ADVANCED AGAIN. SUBSEQUENTLY, WHILE ATTEMPTING TO UNSHEATHE AGAIN, THE DSF1233 WAS OBSERVED TO BE OUT OF THE BRACHIOCEPHALIC ARTERY. ANGIOGRAPHY SHOWED A RUPTURE OF THE BRACHIOCEPHALIC ARTERY AND HYPOTENSION WAS OBSERVED, SO A BLOOD TRANSFUSION WAS GIVEN. THE BRACHIOCEPHALIC ARTERY WAS TORN. IT WAS REPORTED THAT THE PLC161200J WAS THOUGHT TO HAVE PASSED THROUGH THE RUPTURE SITE WHILE STILL IN THE DSF1233 AND THAT THE BRACHIOCEPHALIC ARTERY WAS THOUGHT TO HAVE BEEN TORN BY THE MANUPIRATION OF THE DSF1233. BLOOD FLOW IN THE RIGHT SUBCLAVIAN ARTERY WAS UNAFFECTED. ANOTHER ATTEMPT WAS MADE TO INSERT A RADIFOCUS¿ GUIDEWIRE, BUT IT STRAYED INTO A FALSE LUMEN THAT APPEARED TO HAVE FORMED AS A RESULT OF THE VESSEL RUPTURE. AN ATTEMPT TO CREATE A PULL-THROUGH WAS ALSO UNSUCCESSFUL, SO THE PROCEDURE WAS CONVERTED TO OPEN SURGERY, BUT THE VESSEL REPAIR WAS DIFFICULT. THE HEARTBEAT DID NOT RETURN DUE TO HAEMORRHAGE AND THE CHEST WAS INEVITABLY CLOSED WITHOUT FURTHER TREATMENT. THE PATIENT DIED THE SAME DAY AFTER BEING TRANSFERRED TO INTENSIVE CARE. THE CAUSE OF DEATH WAS REPORTED TO BE EXSANGUINATION DUE TO RUPTURE OF THE BRACHIOCEPHALIC ARTERY. IT WAS NOT AVAILABLE WHETHER AN AUTOPSY WAS PERFORMED. REPORTEDLY, TGMR404020J AND PLC161200J WERE REMOVED UNDEPLOYED. THE PHYSICIAN REPORTEDLY STATED AS FOLLOWS: THERE MAY HAVE BEEN AN ERROR OF JUDGEMENT BY THE SURGEON. PERHAPS A STIFFER STIFF GUIDE WIRE SHOULD HAVE BEEN USED. Manufacturer Narrative: SECTION H6: INVESTIGATION FINDINGS AND CONCLUSION CODES WERE UPDATED TO REFLECT RESULTS OF INVESTIGATION. SECTION H6: D12: THE GORE® DRYSEAL FLEX INTRODUCER SHEATH IFU PROVIDES THE FOLLOWING WARNING: CAREFUL EVALUATION OF VESSEL SIZE, ANATOMY, TORTUOSITY, AND DISEASE STATE (INCLUDING CALCIFICATION, PLAQUE, AND THROMBUS) IN ACCORDANCE WITH THE OBSERVATION IN CT IMAGE OR ANGIOGRAPHY OBTAINED BY PRE-OPERATION / OPERATION OR VISUAL OBSERVATION OBTAINED BY OPERATION IS REQUIRED TO ENSURE SUCCESSFUL USE OF GORE® DRYSEAL FLEX INTRODUCER SHEATH (FLEX). [IF VESSEL SIZE IS SMALLER THAN THE INTRODUCER SHEATH OUTER DIAMETER OR IF VESSEL IS NOT ADEQUATE FOR ACCESS, MAJOR BLEEDING, VESSEL DAMAGE, OR EMBOLISM, MAY RESULT]. DO NO CONTINUE ADVANCEMENT OR RETRACTION OF THE DEVICE IF THERE IS RESISTANCE. DETERMINE THE LOCATION AND THE CAUSE OF RESISTANCE BEFORE PROCEEDING [IF WE DID NOT CONFIRM THE LOCATION UNDER FLUOROSCOPIC GUIDANCE OR CONTINUED ADVANCEMENT OR RETRACTION AGAINST RESISTANCE MAY RESULT IN VESSEL DAMAGE OR DAMAGE TO/BREAKAGE OF THE DEVICE]. IN ADDITION, THE IFU STATES POSSIBLE ADVERSE EVENTS INCLUDE BLOOD LOSS, BLEEDING, HEMATOMA, VASCULAR TRAUMA, DISSECTION, RUPTURE, PERFORATION, TEAR. PRODUCT INVESTIGATION REPORT CONCLUSION AN EVALUATION OF THE PROVIDED RADIOGRAPHIC IMAGE APPEARS TO SUGGEST A POSSIBLE RUPTURE OF THE BRACHIOCEPHALIC ARTERY; HOWEVER, IT CANNOT BE CONCLUSIVELY CONFIRMED. IT WAS REPORTED THAT THE BRACHIOCEPHALIC ARTERY WAS HIGHLY TORTUOUS AND AFTER ADVANCING A DEVICE THROUGH THE SHEATH THERE WAS DIFFICULTY WITHDRAWING IT FOR DEPLOYMENT, SO THE DEVICE WAS RE-SHEATHED AND THE DEVICE AND SHEATH WERE FORCIBLY ADVANCED AGAIN. THE ROOT CAUSE FOR THE REPORTED BRACHIOCEPHALIC ARTERY RUPTURE CANNOT BE ESTABLISHED WITH THE AVAILABLE INFORMATION, HOWEVER THE REPORTED TORTUOSITY AND FORCEFUL ADVANCEMENT MAY HAVE CONTRIBUTED. Manufacturer Narrative: H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICE WAS DISCARDED AT FACILITY. RETURN NOT POSSIBLE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108685&pc=BZD | 2518422-2025-000863 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-08 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21088449&pc=OZP | 2032227-2025-103496 | 2024-11-07 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-07 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS DECEASED ON (B)(6) 2024. CAUSE OF DEATH WAS INJURY FROM VEHICULAR ACCIDENT. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-7040MA, MMT-332A, MMT-397A, MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040MA. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21088561&pc=NAY | 2955842-2025-00167 |  | Death | INTUITIVE SURGICAL, INC | 2025-01-07 05:00 | NAY | DAVINCI XI | Insufficient Information | Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE PROCEDURE DATE WAS UNKNOWN, AND THE CAUSE OF DEATH WAS NOT PROVIDED. THE DA VINCI PROCEDURE WAS CONVERTED TO OPEN; THE SITE OPERATING ROOM MANAGER REPORTED THAT THE CONVERSION WAS NOT RELATED TO A DEVICE ISSUE OR A DA VINCI PRODUCT MALFUNCTION. THE REASON FOR THE CONVERT TO OPEN WAS NOT PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED; NO RESPONSE HAS BEEN RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC CHOLECYSTECTOMY WHICH REQUIRED CONVERSION TO AN OPEN PROCEDURE. NO INFORMATION AS TO THE REASONS FOR THE CONVERSION WERE PROVIDED. THE PATIENT SUBSEQUENTLY DIED BUT NO DETAILS AS TO THE EVENTS THAT LED TO THEIR DEATH OR HOW THEY DIED WERE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: SECTION B5 ADDITIONAL INFORMATION RECEIVED: THE DIRECTOR OF QUALITY FOR THIS HOSPITAL REPORTED THAT THERE WAS NO DA VINCI ISSUE AND THEY DO NOT ATTRIBUTE THIS EVENT TO A DA VINCI PRODUCT. NO ADDITIONAL EVENT INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21088563&pc=NAY | 2955842-2025-00166 | 2024-12-11 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-07 05:00 | NAY | DAVINCI SI | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Unspecified Tissue Injury | K081137 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT PULMONARY LOBECTOMY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING AND THE PROCEDURE WAS CONVERTED TO OPEN. THE PATIENT EXPIRED THE SAME DAY AS THE PROCEDURE. IT IS UNKNOWN IF THE PATIENT DIED INTRA-OPERATIVELY, OR POST-OPERATIVELY. IT WAS REPORTED THAT VASCULAR RECONSTRUCTION WAS PERFORMED TO CONTROL THE INTRA-OPERATIVE BLEEDING. THE CAUSE AND SOURCE OF THE BLEEDING WAS NOT KNOWN TO THE REPORTER. THE AMOUNT OF BLOOD LOSS WAS LESS THAN 500ML; THE PATIENT WAS ADMINISTERED UNSPECIFIED BLOOD PRODUCT(S). IT WAS REPORTED THAT THERE WAS NO DA VINCI MALFUNCTION PRIOR TO THE CONVERT TO OPEN. ATTEMPTS WERE MADE TO SPEAK WITH THE SURGEON AND HAVE BEEN SO FAR UNSUCCESSFUL.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD LIKELY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) INVOLVED WITH THIS REPORTED EVENT AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC PULMONARY LOBECTOMY AND ULTIMATELY DIED THE SAME DAY. BLEEDING OCCURRED DURING THE CASE. THE DETAILS OF HOW THAT BLEEDING OCCURRED, THE FACTORS THAT MAY HAVE CONTRIBUTED TO THAT EVENT AND HOW THE BLEEDING WAS ADDRESSED, AND HOW THIS EVENT MAY HAVE CONTRIBUTED TO THE DEATH OF THE PATIENT WERE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: MARYLAND BIPOLAR FORCEPS, CADIERE FORCEPS, PERMANENT CAUTERY SPATULA, VESSEL SEALER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21088897&pc=MIH | 3013164176-2025-02342 | 2024-05-29 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-07 05:00 | MIH | GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P020004 |  | 1 | Event Description: THE FOLLOWING ARTICLE WAS REVIEWED. SOURCE: OFFICIAL JOURNAL OF THE JAPANESE SOCIETY FOR VASCULAR SURGERY, (0918-6778)33 SUPPL. PAGE RO8-1 TITLE: MIDTERM RESULTS OF EVAR USING GORE EXCLUDER ILIAC BRANCH ENDOPROSTHESIS (IBE) ALL 42 EVAR TREATED WITH GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS(IBE) FROM SEPTEMBER 2017 TO OCTOBER 2023 WERE REVIEWED RETROSPECTIVELY; ONLY ONE PATIENT USED IBE BILATERALLY AND ALL OTHERS USED IBE ONE SIDE ONLY. MEAN AGE WAS 74±7 YEARS (59-90 YEARS), 39 WERE MALE AND 3 WERE FEMALE. MEAN FOLLOW-UP WAS 38 ± 22 MONTHS (0-72 MONTHS). FOLLOW-UP RATE WAS 100%. MORTALITY, ABDOMINAL AORTIC AND COMMON ILIAC ARTERY ANEURYSMS ENLARGEMENT MORE THAN 5 MM, AND INTERNAL ILIAC ARTERY PATENCY ON THE IBE SIDE WERE REVIEWED. TECHNICAL SUCCESS WAS 97.7%, AND ONE PATIENT HAD INTRAOPERATIVE INTERNAL ILIAC ARTERY OCCLUSION DUE TO DISSECTION OF THE INTERNAL ILIAC ARTERY. MEAN OPERATIVE TIME WAS 152±49 MINUTES (82-279 MINUTES), AND BLOOD LOSS WAS 67±74 ML (9-353 ML). 14 PATIENTS UNDERWENT INTERNAL ILIAC ARTERY EMBOLIZATION OF CONTRALATERAL SIDE. INTRAOPERATIVE COMPLICATION WAS INTERNAL ILIAC ARTERY OCCLUSION IN ONE CASE. ENDOLEAKS WERE OBSERVED IN 20 PATIENTS: TYPE 1A IN 1, TYPE 2 IN 18, AND TYPE 4 IN 1 (USING ENDURANT LEG). THERE WAS NO DEATH DURING HOSPITALIZATION. BUTTOCK CLAUDICATION WAS OBSERVED IN 3 CASES, ALL OF WHICH PERFORMED EMBOLIZATION OF CONTRALATERAL INTERNAL ILIAC ARTERIES. ALL THE CLAUDICATION OCCURRED AT THE CONTRALATERAL SIDE. 1 PATIENT HAD A MICROEMBOLIZATION OF THE IBE SIDE OF THE INTERNAL ILIAC ARTERY, WHICH GOT LESS SEVERE WITH CONSERVATIVE TREATMENT. THE MEAN POSTOPERATIVE HOSPITAL STAY WAS 7±2 DAYS (5-12 DAYS). LONG-TERM MORTALITY IN 5 CASES: 1 STENTGRAFT INFECTION, 1 RUPTURE OF THE DESCENDING AORTA DUE TO TYPE B DISSECTION, AND 3 CASES OF PNEUMONIA. POSTOPERATIVE REINTERVENTION WAS PERFORMED IN 4 PATIENTS, ANEURYSM RESECTION WITH ABDOMINAL AORTIC ANEURYSM ENLARGEMENT IN 1 CASE, AND CONTRALATERAL INTERNAL ILIAC ARTERY EMBOLIZATION AND EVAR IN 3 CASES. AT THE FINAL FOLLOW-UP, AORTIC ANEURYSM DIAMETER WAS REDUCED IN 8 CASES, UNCHANGED IN 28 CASES, AND ENLARGED IN 6 CASES (INCLUDING 1 ANEURYSM RESECTION), AND COMMON ILIAC ARTERY ANEURYSM WAS REDUCED IN 18 CASES, UNCHANGED IN 24 CASES, AND ENLARGED IN 0 CASE. 33 CASES OF INTERNAL ILIAC ARTERY ON THE IBE TREATMENT SIDE WERE PATENT. MORTALITY RATES WERE 100% AT 12 MONTHS, 96.8% AT 24 MONTHS, AND 83.7% AT 48 MONTHS. REINTERVENTION AVOIDANCE RATES WERE 97.2% AT 12 MONTHS, 97.2% AT 24 MONTHS, AND 97.2% AT 48 MONTHS. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. A2, A3: DEMOGRAPHIC INFORMATION IDENTIFIED IN THE LITERATURE HAS BEEN ENTERED AS PATIENT INFORMATION. B2: THE DATE OF PATIENT DEATH IS UNKNOWN. H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICES NOT AVAILABLE FOR RETURN. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: INFECTION TITLE: MIDTERM RESULTS OF EVAR USING GORE EXCLUDER ILIAC BRANCH ENDOPROSTHESIS (IBE) SOURCE: OFFICIAL JOURNAL OF THE JAPANESE SOCIETY FOR VASCULAR SURGERY, (0918-6778)33 SUPPL. PAGE RO8-1 W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21089206&pc=BZD | 2518422-2025-000339 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Asthma; Headache; Hypersensitivity/Allergic reaction; Nausea; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation; Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), CANCER, OTHER (UNSPECIFIED EVENT) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21090246&pc=GDW | 2647580-2025-00096 | 2024-08-09 04:00 | Death | US SURGICAL PUERTO RICO | 2025-01-07 05:00 | GDW | UNKNOWN EEA | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Necrosis; Obstruction/Occlusion |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY WAS COMPLETED INVOLVING 100 CASES HAVING PREOPERATIVE ARTERIAL CALCIFICATIONS THAT UNDERWENT MINIMALLY INVASIVE IVOR LEWIS ESOPHAGECTOMIES FROM 2013 TO 2018. AN UNKNOWN LAPAROSCOPIC LINEAR STAPLER WAS USED TO CREATE A GASTRIC CONDUIT. THE 25 MM DST XL CIRCULAR STAPLER WAS USED TO CREATE END-SIDE ESOPHAGOGASTRIC ANASTOMOSIS. POSTOPERATIVE COMPLICATIONS INCLUDE ANASTOMOTIC LEAK, GRAFT/CONDUIT NECROSIS, AND CONDUIT OBSTRUCTION. REOPERATION WAS REQUIRED FOR CONDUIT OBSTRUCTION/NECROSIS. NO REOPERATION WAS REPORTED FOR ANASTOMOTIC LEAK. READMISSIONS WERE ALSO REPORTED. OTHER COMPLICATIONS UNRELATED TO THE DEVICE INCLUDE ATRIAL FIBRILLATION, RESPIRATORY FAILURE, PLEURAL EFFUSIONS, COLONIC PSEUDO-OBSTRUCTION, CONDUIT VOLVULUS, CONDUIT REDUNDANCY, CECAL VOLVULUS, TOXIC MEGACOLON, CLOSED-LOOP OBSTRUCTION, WOUND INFECTION, AND LUNG DECORTICATION. DEATH OCCURRED IN FOUR CASES DUE TO UNSPECIFIED REASONS. TITLE: MAJOR CORONARY ARTERY CALCIFICATIONS AS PREDICTORS OF POSTOPERATIVE COMPLICATIONS IN IVOR LEWIS ESOPHAGECTOMIES: A FVE-YEAR RETROSPECTIVE ANALYSIS. KEOUNA PATHER1 · HAYTHAM ALABBAS1,3 · CARLOS GONZALEZ-BAERGA2 · MANUEL MENENDEZ2 · MAYUR K. VIRARKAR2 · IRAI SANTANA DE OLIVEIRA2 · ERIN M. MOBLEY1 · ZIAD T. AWAD1 SURGICAL ENDOSCOPY (2024) 38:6865¿6872 HTTPS://DOI.ORG/10.1007/S00464-024-11181-3. Manufacturer Narrative: KEOUNA PATHER·HAYTHAM ALABBAS·CARLOS GONZALEZ-BAERGA·MANUEL MENENDEZ·MAYUR K. VIRARKAR·IRAI SANTANA DE OLIVEIRA· ERIN M. MOBLEY· ZIA D T. AWAD. "MAJOR CORONARY ARTERY CALCIFCATIONS AS PREDICTORS OF POSTOPERATIVE COMPLICATIONS IN IVOR LEWIS ESOPHAGECTOMIES: A FVE-YEAR RETROSPECTIVE ANALYSIS", 2024, SURGICAL ENDOSCOPY (2024) 38:6865¿6872 HTTPS://DOI.ORG/10.1007/S00464-024-11181-3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21090287&pc=BZD | 2518422-2025-000438 | 2024-05-02 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21090369&pc=DTB | 2649622-2025-00432 | 2024-12-16 05:00 | Death | MPRI | 2025-01-07 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS UNDERSENSING OF ATRIAL FIBRILLATION. THE RIGHT VENTRICULAR (RV) LEAD AND RIGHT ATRIAL (RA) LEAD REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED WITH CAUSE OF DEATH UNKNOWN. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE.Event Description: IT WAS REPORTED THAT THERE WAS UNDERSENSING OF ATRIAL FIBRILLATION. THE RIGHT VENTRICULAR (RV) LEAD AND RIGHT ATRIAL (RA) LEAD REMAIN IN USE. IT WAS FURTHER REPORTED THAT THE RA LEAD EXHIBITED LOSS OF SENSING. IT WAS NOTED THAT THE PATIENT IS DECEASED WITH CAUSE OF DEATH UNKNOWN. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: CORRECTED B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21090456&pc=BZD | 2518422-2025-000445 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Headache; Nausea; Vomiting; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21090947&pc=DTB | 2649622-2025-00440 | 2024-12-16 05:00 | Death | MPRI | 2025-01-07 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS UNDERSENSING OF ATRIAL FIBRILLATION. THE RIGHT VENTRICULAR (RV) LEAD AND RIGHT ATRIAL (RA) LEAD REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED WITH CAUSE OF DEATH UNKNOWN. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21091435&pc=CBK | 9710602-2025-00039 | 2024-12-11 05:00 | Death | GE MEDICAL SYSTEMS (CHINA) CO., LTD (WUXI) | 2025-01-07 05:00 | CBK | CARESCAPE | Insufficient Flow or Under Infusion | Insufficient Information | K210384 |  | 1 | Event Description: THE HOSPITAL REPORTED A PATIENT CONNECTED TO A CARESCAPE R860 WHEN IT WAS ALLEGED THAT THE UNIT ALARMED FOR LOW TIDAL VOLUME. REPORTEDLY, THE TIDAL VOLUME WAS 120ML, BUT WAS SET FOR 450ML. MANUAL VENTILATION WAS INITIATED CONTINUOUSLY UNTIL IT EXCEEDED THE ALARM UNIT. IT WAS REPORTED THAT RESUSCITATION EFFORTS FAILED AND THE PATIENT DIED. GE HEALTHCARE WILL SUBMIT A FOLLOW-UP REPORT WHEN THE INVESTIGATION HAS BEEN COMPLETED. Manufacturer Narrative: GE HEALTHCARE'S (GEHC) INVESTIGATION HAS BEEN COMPLETED AND THE ROOT CAUSE COULD NOT BE DETERMINED. BASED ON THE REVIEW OF THE DEVICE LOGS, PART INSPECTION, LAB TESTING, THE ROOT CAUSE OF THE FAULTY EXPIRATORY FLOW SENSOR WAS DETERMINED TO BE CONTAMINATION DUE TO THE USER NOT USING AN EXPIRATORY FILTER. THIS WOULD ONLY AFFECT THE DISPLAYED VALUES AND NOT DELIVERY. THE ROOT CAUSE OF THE PATIENT DEATH IS UNDETERMINED. Manufacturer Narrative: A GE HEALTHCARE SERVICE REPRESENTATIVE PERFORMED A CHECKOUT OF THE SYSTEM BUT DID NOT CONFIRM THE REPORTED ISSUE. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. BLOCK A: NO PATIENT INFORMATION PROVIDED. D4 PRIMARY UDI NUMBER: THIS DEVICE WAS NOT MANUFACTURED FOR SALE IN THE USA AND THEREFORE A USA UDI IS NOT AVAILABLE. LEGAL MANUFACTURER: HCS MADISON - 3030 OHMEDA DR, USA MADISON, WI 53718.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21091575&pc=MKJ | 1220908-2024-04800 | 2024-12-15 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-07 05:00 | MKJ | AED PLUS | Battery Problem | Insufficient Information | P160015 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 60/70-YEAR-OLD FEMALE PATIENT, THE DEVICE PROMPTED A CHANGE BATTERIES MESSAGE. THE USER THOUGHT THIS MEANT THE DEVICE WAS NOT OPERATIONAL AND DID NOT TRY TO USE THE DEVICE FOLLOWING THE CHANGE BATTERIES PROMPT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL JAPAN FOR EVALUATION. THE CUSTOMER'S REPORT WAS VERIFIED DURING EVALUATION. A REVIEW OF THE DEVICE LOG SHOWS THE DEVICE WAS NOTIFYING A RED X CONDITION BEGINNING 10 MONTHS PRIOR TO THE DEVICE BEING DEPLOYED ON A PATIENT THAT WAS NOT ADDRESSED. THE DEVICE WAS TESTED WITH THE CUSTOMER'S RETURNED BATTERIES AND WAS ABLE TO POWER ON AND PASS THE DEVICE BATTERY INSTALLATION TEST AND BATTERY TEST. THE USER AT THE TIME OF THE EVENT WAS A LAY-USER AND DID NOT KNOW THE DEVICE WOULD STILL OPERATE DESPITE PROMPTING A "CHANGE BATTERIES" MESSAGE. ROOT CAUSE IS RELATED TO FOLLOWING ZOLL REOMMENDATIONS FOR OPERATIONAL READINESS. THE AED PLUS ADMINISTRATOR'S GUIDE, 9650-0301-28 RECOMMENDS THAT THE DEVICE BE VISUALLY INSPECTED OFTEN, AS NECESSARY, TO ENSURE THAT THE DEVICE IS DISPLAYING A GREEN CHECK AND IS RESCUE READY. THE ADMINISTRATOR'S GUIDE ALSO STATES THAT, IF THE UNIT PROMPTS "CHANGE BATTERIES" A LOW BATTERY CONDITION HAS BEEN DETECTED AND THE BATTERIES MUST BE REPLACED IMMEDIATELY. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21091589&pc=GKT | 1219343-2025-00001 | 2024-12-09 05:00 | Death | HAEMONETICS CORPORATION | 2025-01-07 05:00 | GKT | NEXSYS PLASMA COLLECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | BK180185 |  | 1 | Event Description: THIS REPORT IS BEING SUBMITTED TO ALIGN WITH CUSTOMER REPORTING REQUIREMENTS. THE USE OF THE HAEMONETICS MEDICAL DEVICE IN THE DONATION PROCEDURE IS NOT SUSPECTED TO HAVE RESULTED IN THE ADVERSE OUTCOME OF THE DONOR. CUSTOMER REPORTED AN OVER-WEIGHT 43-YEAR-OLD MALE DONOR EXPIRED SOMETIME AFTER LEAVING THE DONATION CENTER ON (B)(6) 2024 DUE TO AN UNKNOWN CAUSE. LAST DONATION OCCURRED ON (B)(6) 2024 WITH NO ISSUES REPORTED WITH THE PROCEDURE OR EQUIPMENT/DISPOSABLES USED. REVIEW OF DONOR'S CHART INDICATED NO MEDICAL PROBLEMS, NO MEDICATIONS TAKEN, NORMAL PHYSICAL EXAM, NORMAL LABS AND TEST RESULTS AND REVIEW OF VITALS WERE NORMAL. DONOR HAD NO DEFERRALS. THERE WERE NO KNOWN ALLERGIES OR PRE/POST IMMUNIZATIONS. IT IS UNKNOWN IF AN AUTOPSY WILL BE PERFORMED. THERE IS NO INFORMATION FROM THE ATTENDING PHYSICIAN, SURGEON, HOSPITAL REPRESENTATIVE OR HEALTH CARE PROFESSIONAL. Manufacturer Narrative: A HAEMONETICS FIELD SERVICE ENGINEER EVALUATED THE EQUIPMENT USED DURING THE DONATION. DEVICE WAS TESTED AND NO PROBLEM WAS FOUND. ALL OTHER PARAMETERS VERIFIED. FUNCTION TESTS COMPLETED. MACHINE MEETS MANUFACTURER'S SPECIFICATIONS AND IS READY TO USE. BASED ON THIS DESCRIPTION THE DEVICE IS EVALUATED WITH NO DEFECT. THERE ARE NO NONCONFORMANCES AGAINST THE DEVICE SERIAL NUMBER AND NO CAPAS RELATED TO THIS COMPLAINT. A REVIEW OF THE DHR SHOWS NO ISSUES DURING MANUFACTURING AND ALL TESTING PASSED. THE DISPOSABLES USED WITH THE SYSTEM WERE DISCARDED, HOWEVER THERE WERE NO RECALLS OR ADVERSE TRENDS RELATED TO THE PRODUCT LOTS USED IN THE PROCEDURE. THERE IS NO EVIDENCE TO SUGGEST THAT THE DONOR EVENT WAS RELATED TO THE DEVICE OR DISPOSABLES USED DURING THE PLASMAPHERESIS PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092195&pc=NPT | 2025587-2025-00126 | 2024-05-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-07 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY ONE MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH HEART FAILURE, THE PATIENT PASSED AWAY. PER THE PHYSICIAN, THE PATIENT'S HEART FAILURE DID NOT IMPROVE FOLLOWING THE PROCEDURE AND THE CAUSE OF DEATH WAS RELATED TO THE CARDIOVASCULAR SYSTEM.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THERE WAS NO RELATION BETWEEN THE MEDTRONIC DEVICE AND THE DEATH. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED D.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092221&pc=NVN | 2182208-2025-00059 | 2024-01-01 05:00 | Death | MPRI | 2025-01-07 05:00 | NVN | SELECTSECURE MRI SURESCAN | Failure to Capture; Device Dislodged or Dislocated; High Capture Threshold | Endocarditis; Renal Failure; Multiple Organ Failure | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRICUSPID REGURGITATION AND INFECTIVE ENDOCARDITIS AS A CAUSE OF LEAD DISLODGEMENT IN LEFT BUNDLE BRANCH AREA PACING. THE AUTHORS DESCRIBED A PATIENT WITH SEVERE PREOPERATIVE TRICUSPID REGURGITATION WHO WAS IMPLANTED WITH LEFT BUNDLE BRANCH AREA PACING. POST-PROCEDURE ECHOCARDIOGRAPHY SHOWED THE LEAD LOCATED IN CLOSE PROXIMITY OF THE TRICUSPID REGURGITATION JET. THRESHOLD AND IMPEDANCE VALUES, TRICUSPID REGURGITATION SEVERITY, AND RIGHT VENTRICULAR DIAMETERS AND FUNCTION REMAINED STABLE AT THE FOUR-MONTH FOLLOW-UP. HOWEVER, TWO MONTHS LATER DURING ROUTINE CHECK-UP, THERE WAS AN ELEVATION OF THRESHOLD AND LOSS OF LEFT BUNDLE CAPTURE. OSCILLATION MOVEMENT WAS OBSERVED IN THE LEAD AT THE IMPLANTATION AREA AND A MOBILE VEGETATION-LIKE IMAGE WAS DETECTED ON THE LEAD. THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH SUSPECTED DISLODGMENT AND POSSIBLE INFECTIVE ENDOCARDITIS. BLOOD CULTURES WERE TAKEN, AND THE PATIENT WAS STARTED ON ANTIBIOTIC TREATMENT. DISLODGMENT WAS CONFIRMED AND A LEAD EXTRACTION WAS PERFORMED. THE PATIENT SUBSEQUENTLY DIED TEN DAYS AFTER HOSPITALIZATION BECAUSE OF MULTI-ORGAN FAILURE SECONDARY TO ACUTE RENAL FAILURE, POSSIBLY DUE TO THE ANTIBIOTIC TREATMENT. THE STATUS OF THE LEAD IS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. THE EVENT ONLY OCCURRED WITH ONE PATIENT BUT SPECIFIC DETAILS ON THE PATIENT WERE NOT PROVIDED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: TRICUSPID REGURGITATION AND INFECTIVE ENDOCARDITIS AS A CAUSE OF LEAD DISLODGEMENT IN LEFT BUNDLE BRANCH AREA PACING. PACING AND CLINICAL ELECTROPHYSIOLOGY. 2024; 47:683¿687. DOI: 10.1111/PACE.14810. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092227&pc=OZD | 1220648-2025-25371 | 2024-12-18 05:00 | Death | ABIOMED, INC. | 2025-01-07 05:00 | OZD | IMPELLA CP | Pumping Stopped | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT IMPELLA CP WAS INSERTED TO SUPPORT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. WHILE ON SUPPORT THE IMPELLA HAD A FEW IMPELLA STOPPED ALARMS. THE TEAM WAS ABLE TO SUCCESSFULLY RESTART THE CP. HEPARIN HAD BEEN STOPPED BY THE DOCTOR BECAUSE OF GASTROINTESTINAL BLEED. IT WAS FURTHER REPORTED THAT THE PATIENT WAS PLACED IN COMFORT CARE AS THE PATIENT WAS NOT IMPROVING. THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION FOR THE PUMP STOP AND GASTROINTESTINAL BLEED (GI) HAS BEEN COMPLETED. PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOGS WERE INVESTIGATED AND "PUMP STOP" ALARMS WERE CONFIRMED. THERE WAS A SUDDEN MOTOR CURRENT SPIKE AND AN AUTO P-LEVEL CHANGE RESPONSE ALONG WE FLOWS DROPPING TO ZERO. PURGE PRESSURE WAS SEEN TO RISE PROMPTLY AS WELL. THESE ARE SYMPTOM SEEN IN A SUDDEN CLOT INGESTION. CLINICAL MENTIONS THAT HEPARIN WAS STOPPED DUE TO INTERNAL BLEEDING WHICH FURTHER SUPPORTS CLOT FORMATION IN THE BODY. THEREFORE, THE ROOT CAUSE OF THE PUMP STOP ISSUE WAS MOST LIKELY BIOMATERIAL. AS PER THE CLINICAL INFORMATION PROVIDED, THERE WAS A GI BLEED BUT NO ADDITIONAL INFORMATION WAS PROVIDED. THEREFORE, THE ROOT CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092456&pc=DSQ | 3007042319-2025-00028 | 2024-12-25 05:00 | Death | HEARTWARE, INC. | 2025-01-07 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Mechanical Problem; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Cardiac Arrest | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A PUMP VENTRICULAR ASSIST DEVICE IMPLANT KIT EXPIRED AT HOME, WITH A SUSPECTED CARDIAC ARREST AS THE CAUSE OF DEATH. THE PUMP WAS DECOMMISSIONED AT THE PATIENT'S HOME FOLLOWING THE DEATH, AND NO AUTOPSY WAS PERFORMED. NO EQUIPMENT IS BEING RETURNED. Manufacturer Narrative: A SUPPLEMENTAL IS BEING SUBMITTED FOR COMPLETION OF THE INVESTIGATION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO A SUSPECTED CARDIAC ARREST. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, CARDIOPULMONARY ARREST AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A REVISION TO THE PRODUCT EVENT SUMMARY. REVISED PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. TWO (2) CONTROLLERS (B)(6) WERE RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. NO PERFORMANCE ALLEGATIONS WERE MADE AGAINST THE CONTROLLERS. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT. FAILURE ANALYSIS OF (B)(6) REVEALED THAT THE DEVICES PASSED FUNCTIONAL TESTING. VISUAL INSPECTION OF (B)(6) REVEALED CONTAMINATION WITHIN BOTH POWER PORTS. INTERNAL VISUAL INSPECTION ASSOCIATED WITH (B)(6) REVEALED CRACKS IN THE STANDOFF POSTS OF THE CONTROLLER HOUSING. THESE ARE ADDITIONAL FINDINGS NOT RELATED TO THE REPORTED EVENT. THE MOST LIKELY ROOT CAUSE OF THE OBSERVED CONTAMINATION WITHIN THE CONTROLLER POWER PORTS CAN BE ATTRIBUTED TO THE HANDLING OF THE DEVICE. BASED ON AN INVESTIGATION, THE ROOT CAUSE OF THE STANDOFF POST CRACKS WAS DETERMINED TO BE DUE TO MINERAL OIL APPLIED TO THE CONTROLLER¿S INTERNAL MAIN GASKET AND/OR TO THE SILICONE ADHESIVE APPLIED AROUND THE CONTROLLER¿S INTERNAL BATTERY COMING INTO CONTACT WITH THE STANDOFF POST. THE MINERAL OIL AND/OR SILICONE ADHESIVE CONTRIBUTED TO ENVIRONMENTAL STRESS CRACKING. CAPA PR00517125 IS INVESTIGATING THIS ISSUE. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED THAT A CLOCK CHANGE WAS LOGGED, IN WHICH THE CONTROLLER'S DATE WAS SET TO (B)(6) 2010; NO PUMP ID SETTING EVENTS HAD BEEN LOGGED ON THE CONTROLLER PRIOR TO THE INITIAL CLOCK CHANGE EVENT. IF THE PUMP ID IS NOT SET WHEN THE CONTROLLER IS CONNECTED TO THE MONITOR, THE MONITOR WILL AUTOMATICALLY SET THE CONTROLLER DATE/TIME TO MATCH THE CURRENT DATE/TIME OF THE MONITOR. ADDITIONALLY, REVIEW OF THE EVENT LOG FILE REVEALED THAT THE PUMP ID WAS SET FOLLOWING THE INITIAL CLOCK CHANGE EVENT; HOWEVER, DATE/TIME SETTINGS WERE NOT ADJUSTED. DUE TO THE DATE/TIME SETTINGS ON THE CONTROLLER, THE SEQUENCE OF EVENTS IS NOT ABLE TO BE DETERMINED. THE MOST LIKELY ROOT CAUSE OF THE INCORRECT DATE AND TIME SETTINGS CAN BE ATTRIBUTED TO IMPROPER SET UP DURING INITIAL PROGRAMMING OF THE CONTROLLER. LOG FILE ANALYSIS ALSO REVEALED A DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS TO PARAMETERS BELOW NORMAL OPERATING RANGE SINCE (B)(6) 2011 AND TWO (2) LOW FLOW ALARMS LOGGED SINCE (B)(6) 2011. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED NO ANOMALIES WITHIN THE ANALYZED PERIOD. OF NOTE, (B)(6) WERE LOADED WITH A SOFTWARE CONTAINING AN UNAPPROVED PUMP START ALGORITHM. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO A SUSPECTED CARDIAC ARREST. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE OBSERVED LOW FLOW EVENT INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FIL LING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, CARDIOPULMONARY ARREST AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - CONTROLLER D4: MODEL#: 1420 / CATALOG#: 1420 / EXPIRATION DATE: DD-MMM-YYYY / SERIAL#: (B)(6) D9: YES, RETURN DATE: 21-APRIL-2025 H3: YES H4: MFG DATE: 01-NOV-2022 H5: NO D1: CONTROLLER D4: MODEL#: 1420 / CATALOG#: 1420 / EXPIRATION DATE: DD-MMM-YYYY / SERIAL#: (B)(6) D9: YES, RETURN DATE: 21-APRIL-2025 H3: YES H4: MFG DATE: 08-SEPT-2023 H5: NO INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092834&pc=PNJ | 9612164-2025-00109 | 2024-11-29 05:00 | Death | MEDTRONIC IRELAND | 2025-01-07 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092968&pc=NPT | 2025587-2025-00142 | 2024-12-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-07 05:00 | NPT | EVOLUT FX PLUS VALVE | Difficult or Delayed Activation; Device Dislodged or Dislocated; Patient Device Interaction Problem | Hemorrhage/Bleeding; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE REPLACEMENT, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS A STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING THE VALVE DEPLOYMENT, THE VALVE DISLODGED. A SNARE WAS USED TO MOVE THE VALVE INTO THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM (DCS) AND ADVANCED TO THE ANNULUS; HOWEVER, THE NOSECONE OF THE DCS CAUGHT ON THE DISLODGED VALVE, MOVING IT AND CAUSING AN AORTIC DISSECTION ABOVE THE SINOTUBULAR JUNCTION. SUBSEQUENTLY, A PERICARDIAL WINDOW WAS PERFORMED AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED. A 2-HOUR PROCEDURAL DELAY OCCURRED. DUE TO THE SEVERE BLOOD LOSS, ECMO WAS NOT ABLE TO FLOW SUCCESSFULLY, AND SUBSEQUENTLY, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE REPLACEMENT, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS A STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING THE VALVE DEPLOYMENT, THE VALVE DISLODGED. A SNARE WAS USED TO MOVE THE VALVE INTO THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM (DCS) AND ADVANCED TO THE ANNULUS; HOWEVER, THE NOSECONE OF THE DCS CAUGHT ON THE DISLODGED VALVE, MOVING IT AND CAUSING AN AORTIC DISSECTION ABOVE THE SINOTUBULAR JUNCTION. SUBSEQUENTLY, A PERICARDIAL WINDOW WAS PERFORMED AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED. A 2-HOUR PROCEDURAL DELAY OCCURRED. DUE TO THE SEVERE BLOOD LOSS, ECMO WAS NOT ABLE TO FLOW SUCCESSFULLY, AND SUBSEQUENTLY, THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS DEPLOYED OVER A NON-MEDTRONIC GUIDEWIRE WITH A DEPLOYMENT STARTING POINT AT MID PIGTAIL CATHETER. THE VALVE WAS DEPLOYED AT A DEPTH OF 2 MM ON BOTH THE LEFT AND NON-CORONARY CUSPS, AND DISLODGED TOWARDS THE AORTA. THE FIRST DCS MIGHT HAVE CONTRIBUTED TO THE DISLODGEMENT AS THE VALVE PADDLE LOOKED LIKE IT COULD HAVE GOTTEN STUCK IN THE PADDLE RECEIVER POCKET UPON RELEASE. ADDITIONALLY, THE PATIENT'S TORTUOUS AORTA COULD HAVE CONTRIBUTED TO THE DISLODGEMENT DUE TO THE ANGLE OF THE DCS. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D-EVOLUTFX-2329; LOT #: 0012447152; UBD: 2026-09-17; UDI: (B)(4). UPDATED DATA: B5. SECOND PARAGRAPH ADDED D10. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONCLUSION: THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THIS PRODUCT MET ALL MANUFACTURING SPECIFICATIONS FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. A MEDICAL SAFETY ASSESSMENT WAS COMPLETED AND DETERMINED THAT, BASED ON THE INFORMATION PROVIDED, THE RELATIONSHIP BETWEEN THE FIRST VALVE, DELIVERY CATHETER SYSTEM (DCS) AND DISLODGEMENT IS PROBABLE. PATIENT ANATOMY LIKELY CONTRIBUTED TO THE EVENT. A TORTUOUS AORTA COULD HAVE CONTRIBUTED TO THE DISLODGEMENT DUE TO THE ACUTE ANGLE OF THE DCS. THE SECOND VALVE DISLODGMENT WAS LIKELY RELATED TO THE DCS AND INTERACTION WITH THE VALVE AS THE SYSTEM WAS REMOVED THROUGH THE ANNULUS. THE DISSECTION WITH SUBSEQUENT BLOOD LOSS AND DEATH WERE LIKELY RELATED TO THE DISLODGEMENT OF THE SECOND VALVE AND INTERACTION OF THE VALVE AND THE AORTA. POTENTIAL FACTORS THAT CAN INFLUENCE A DISLODGED VALVE INCLUDE TENSION APPLIED ON THE DCS DURING POSITIONING, CALCIFICATION LEVELS IN THE NATIVE VESSEL, COMPLIANCE OF THE AORTA AND NATIVE VESSELS, AND OTHER POTENTIAL FACTORS. VASCULAR COMPLICATIONS, SUCH AS DISSECTION, ARE A KNOWN POTENTIAL ADVERSE PATIENT EFFECT PER THE DEVICE SYSTEM INSTRUCTIONS FOR USE (IFU), AND ARE TYPICALLY RELATED TO PATIENT FACTORS (ANATOMY, COMORBIDITIES, ETC.), AND/ OR PROCEDURAL EFFECTS (SHEATH USED, USER TECHNIQUE, PUNCTURE CUT LOCATION, ETC.). AS REPORTED, THE NOSECONE OF THE DCS CAUGHT ON THE DISLODGED VALVE, MOVING IT AND CAUSING AN AORTIC DISSECTION ABOVE THE SINOTUBULAR JUNCTION. BLEEDING IS A KNOWN POTENTIAL ADVERSE EFFECT PER THE IFU AND CAN OCCUR DURING THE IMPLANT PROCEDURE, WITH A VARYING RISK THAT IS DEPENDENT ON SEVERAL FACTORS SUCH AS THE ACCESS POINT FOR THE IMPLANT, THE PATIENT'S STATE OF HEALTH, AND PRE-EXISTING MEDICAL CONDITIONS. THE ISSUE IS COMMONLY RESOLVED THROUGH A BLOOD TRANSFUSION OR PROVISION OF PACKED RED BLOOD CELLS (PRBCS). PATIENT DEATH IS A POTENTIAL RISK ASSOCIATED WITH THE IMPLANTATION OF A BIOPROSTHESIS VALVE PER THE DEVICE IFU. A PROCEDURE-RELATED OR VALVE- RELATED DEATH IS AN INHERENT RISK WHEN THE PATIENT CONDITION IS SUCH THAT A TRANSCATHETER AORTIC VALVE IS NEEDED TO SUSTAIN CARDIAC FUNCTION, AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. THE REPORTED EVENT INDICATES THAT THE VALVE PADDLE LOOKED LIKE IT COULD HAVE GOTTEN STUCK IN THE PADDLE RECEIVER POCKET UPON RELEASE. THE USER SHOULD RELEASE THE TENSION IN SYSTEM JUST BEFORE FINAL RELEASE TO REDUCE POTENTIAL FOR VALVE MOVEMENT AND THE USER SHOULD SLIGHTLY PUSH ON THE DCS WHILE TUNING THE DEPLOYMENT KNOB VERY SLOWLY TO ALLOW THE PADDLES TO DETACH ONE AT A TIME. ADDITIONALLY, IF PADDLE FAILS TO IMMEDIATELY DETACH DO NOT TORQUE (TURN) THE DCS HANDLE AS THIS WILL NOT TRANSLATE TO THE DISTAL TIP. OFTEN A HUNG PADDLE WILL RESOLVE ITSELF AFTER A FEW HEART CYCLES, BUT IF IT DOESN¿T, THE OPERATOR CAN EITHER PULL SLIGHTLY ON THE GUIDEWIRE OR GENTLY PUSH THE DCS FORWARD TO RELEASE THE PADDLE. UPDATED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012449260); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: D-EVOLUTFX-2329 (0012447152); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21092979&pc=NPT | 2025587-2025-00143 | 2024-12-24 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-07 05:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem | Hemorrhage/Bleeding; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE REPLACEMENT, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS A STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING THE VALVE DEPLOYMENT, THE VALVE DISLODGED. A SNARE WAS USED TO MOVE THE VALVE INTO THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM (DCS) AND ADVANCED TO THE ANNULUS; HOWEVER, THE NOSECONE OF THE DCS CAUGHT ON THE DISLODGED VALVE, MOVING IT AND CAUSING AN AORTIC DISSECTION ABOVE THE SINOTUBULAR JUNCTION. SUBSEQUENTLY, A PERICARDIAL WINDOW WAS PERFORMED AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED. A 2-HOUR PROCEDURAL DELAY OCCURRED. DUE TO THE SEVERE BLOOD LOSS, ECMO WAS NOT ABLE TO FLOW SUCCESSFULLY, AND SUBSEQUENTLY, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE REPLACEMENT, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS A STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING THE VALVE DEPLOYMENT, THE VALVE DISLODGED. A SNARE WAS USED TO MOVE THE VALVE INTO THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM (DCS) AND ADVANCED TO THE ANNULUS; HOWEVER, THE NOSECONE OF THE DCS CAUGHT ON THE DISLODGED VALVE, MOVING IT AND CAUSING AN AORTIC DISSECTION ABOVE THE SINOTUBULAR JUNCTION. SUBSEQUENTLY, A PERICARDIAL WINDOW WAS PERFORMED AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED. A 2-HOUR PROCEDURAL DELAY OCCURRED. DUE TO THE SEVERE BLOOD LOSS, ECMO WAS NOT ABLE TO FLOW SUCCESSFULLY, AND SUBSEQUENTLY, THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS DEPLOYED OVER A NON-MEDTRONIC GUIDEWIRE WITH A DEPLOYMENT STARTING POINT AT MID PIGTAIL CATHETER. THE VALVE WAS DEPLOYED AT A DEPTH OF 2 MM ON BOTH THE LEFT AND NON-CORONARY CUSPS, AND DISLODGED TOWARDS THE AORTA. THE FIRST DCS MIGHT HAVE CONTRIBUTED TO THE DISLODGEMENT AS THE VALVE PADDLE LOOKED LIKE IT COULD HAVE GOTTEN STUCK IN THE PADDLE RECEIVER POCKET UPON RELEASE. ADDITIONALLY, THE PATIENT'S TORTUOUS AORTA COULD HAVE CONTRIBUTED TO THE DISLODGEMENT DUE TO THE ANGLE OF THE DCS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012449260); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: EVFXPLUS-26 (J359217); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093175&pc=PYX | 1220648-2025-25382 | 2024-12-20 05:00 | Death | ABIOMED, INC. | 2025-01-07 05:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hematoma | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA RP FLEX SUPPORT. WHILE ON SUPPORT, A LARGE HEMATOMA WAS NOTED AT THE IMPELLA SITE. NO HEPARIN WAS GIVEN AND FEMOSTOP WAS APPLIED. CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL INFORMATION WAS SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY ANTICOAGULATION MANAGEMENT SINCE THE HEMOSTASIS WAS ACHIEVED AFTER APPLYING FEMSTOP AND DISCONTINUING HEPARIN. B.2 REVISED SELECTIONS AS REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25382. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER. USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093474&pc=OUT | 2029214-2025-00065 | 2024-12-06 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-07 05:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Fever; Intracranial Hemorrhage; Vasoconstriction; Cognitive Changes | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A PATIENT WHO HAD TWO PIPELINES OF THE SAME MODEL/LOT IMPLANTED WHO HAD POST-OPERATIVE COMPLICATIONS. IT WAS REPORTED THAT THE TWO PIPELINE FLEX WITH SHIELD STENTS WERE IMPLANTED ON (B)(6) 2024. NO DEVICE MALFUNCTION OR INTRAOPERATIVE ISSUE WAS REPORTED. ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED AFTER EXPERIENCING FEVER AND PROGRESSIVE DECLINE IN MENTAL STATUS. IMAGING RESULTS WERE SUGGESTIVE OF CONCERN FOR CEREBRAL VASOSPASM ASSOCIATED WITH SUBARACHNOID HEMORRHAGE (SAH) IN EVOLUTION. THE PATIENT UNDERWENT A PROCEDURE FOR INTRAARTERIAL VASODILATOR (VERAPAMIL 20MG) INTERNAL CAROTID ARTERY (ICA) INFUSION WITH MILD IMPROVEMENT AND WAS MONITORED. IT WAS CONSIDERED LIKELY THAT VASOSPASM CAUSED THE SYMPTOMS OF FEVER AND SIMILAR SYMPTOMS TO PNEUMONIA AND SEPTIC SHOCK. THE PATIENT OUTCOME WAS NOTED AS "RECOVERING/RESOLVING." THE EVENT REQUIRED PATIENT HOSPITALIZATION AND INTERVENTION BUT WAS NOT LIFE-THREATENING AND DID NOT RESULT IN PATIENT DISABILITY. SITE ASSESSMENT CONCLUDED THE EVENTS WERE POSSIBLE RELATED TO THE PATIENT DISEASE UNDER STUDY AND/OR THE INDEX PROCEDURE BUT NOT RELATED TO THE IMPLANTED PIPELINES OR ACCESSORY DEVICES.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE EVENT WAS CAUSALLY RELATED TO THE DISEASE UNDER STUDY. THE SITE WAS QUERIED TO CONFIRM IF MULTIPLE DEVICES WERE USED. AVAILABLE SOURCE DOCUMENT ONLY MENTIONS ONE STUDY DEVICE IMPLANTATION.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE EVENT BECAME FATAL ON (B)(6) 2024 . SYMPTOMATIC VASOSPASM PDO 5, 8, DCI ON CT POD 18, DEATH SEVERE ADVERSE EVENT CRITERIA ALTHOUGH NOT IMMEDIATE CAUSE OF DEATH, THE VASOSPASM AND DCI MAY HAVE CONTRIBUTED TO PROLONGED INTUBATION AND ARDS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID PED2-325-20 (B621488); PRODUCT TYPE: ; IMPLANT DATE ; EXPLANT DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093492&pc=OUT | 2029214-2025-00066 | 2024-12-06 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-07 05:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Fever; Intracranial Hemorrhage; Vasoconstriction; Cognitive Changes | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A PATIENT WHO HAD TWO PIPELINES OF THE SAME MODEL/LOT IMPLANTED WHO HAD POST-OPERATIVE COMPLICATIONS. IT WAS REPORTED THAT THE TWO PIPELINE FLEX WITH SHIELD STENTS WERE IMPLANTED ON (B)(6) 2024. NO DEVICE MALFUNCTION OR INTRAOPERATIVE ISSUE WAS REPORTED. ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED AFTER EXPERIENCING FEVER AND PROGRESSIVE DECLINE IN MENTAL STATUS. IMAGING RESULTS WERE SUGGESTIVE OF CONCERN FOR CEREBRAL VASOSPASM ASSOCIATED WITH SUBARACHNOID HEMORRHAGE (SAH) IN EVOLUTION. THE PATIENT UNDERWENT A PROCEDURE FOR INTRAARTERIAL VASODILATOR (VERAPAMIL 20MG) INTERNAL CAROTID ARTERY (ICA) INFUSION WITH MILD IMPROVEMENT AND WAS MONITORED. IT WAS CONSIDERED LIKELY THAT VASOSPASM CAUSED THE SYMPTOMS OF FEVER AND SIMILAR SYMPTOMS TO PNEUMONIA AND SEPTIC SHOCK. THE PATIENT OUTCOME WAS NOTED AS "RECOVERING/RESOLVING." THE EVENT REQUIRED PATIENT HOSPITALIZATION AND INTERVENTION BUT WAS NOT LIFE-THREATENING AND DID NOT RESULT IN PATIENT DISABILITY. SITE ASSESSMENT CONCLUDED THE EVENTS WERE POSSIBLE RELATED TO THE PATIENT DISEASE UNDER STUDY AND/OR THE INDEX PROCEDURE BUT NOT RELATED TO THE IMPLANTED PIPELINES OR ACCESSORY DEVICES.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE EVENT WAS CAUSALLY RELATED TO THE DISEASE UNDER STUDY. THE SITE WAS QUERIED TO CONFIRM IF MULTIPLE DEVICES WERE USED. AVAILABLE SOURCE DOCUMENT ONLY MENTIONS ONE STUDY DEVICE IMPLANTATION.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE EVENT BECAME FATAL ON (B)(6) 2024. SYMPTOMATIC VASOSPASM PDO 5, 8, DCI ON CT POD 18, DEATH SEVERE ADVERSE EVENT CRITERIA ALTHOUGH NOT IMMEDIATE CAUSE OF DEATH, THE VASOSPASM AND DCI MAY HAVE CONTRIBUTED TO PROLONGED INTUBATION AND ARDS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: PED2-325-20 (B621488). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093700&pc=DYB | 2015691-2025-00112 | 2024-11-26 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-07 05:00 | DYB | EDWARDS ESHEATH+ INTRODUCER SET | Material Split, Cut or Torn | Stroke/CVA; Unspecified Vascular Problem | K200258 |  | 1 | Event Description: AS REPORTED FROM A VOLUNTARY MEDWATCH, TAVR PROCEDURE LEFT COMMON CAROTID ARTERY CUTDOWN PERFORMED AND A 6 FR SHEATH WAS PLACED. THE AORTIC VALVE WAS THEN CROSSED UTILIZING AN AL1 CATHETER AND A STRAIGHT TIP AMPLATZ WIRE. THIS WAS THEN EXCHANGED FOR A PRECURVED SAFARI WIRE UTILIZING A PIGTAIL CATHETER. SUBSEQUENTLY, A 29 MM EDWARDS S3 VALVE WAS DEPLOYED UNDER RAPID VENTRICULAR PACING. TRANSESOPHAGEAL ECHOCARDIOGRAPHY DEMONSTRATED NO PARAVALVULAR LEAK OR AORTIC REGURGITATION. NEXT THE TAVR DELIVERY SYSTEM WAS REMOVED FROM THE LEFT COMMON CAROTID ARTERY. DUE TO A SPLIT SHEATH WITH THE DELIVERY SYSTEM UNABLE TO BE WITHDRAWN COMPLETELY INTO THE SHEATH THIS WAS REMOVED AS ONE PIECE. THIS THEN RESULTED IN FURTHER DISRUPTION OF THE CAROTID ARTERIOTOMY. A 10X20 MM BALLOON WAS INFLATED IN THE PROXIMAL LEFT COMMON CAROTID ARTERY TO OBTAIN PROXIMAL CONTROL. DUE TO THE COMPLEXITY OF THE REPAIR WITH CONCERNS REGARDING FURTHER LOSS OF HEMOSTASIS AND THAT THE PATIENT HAD NO EVIDENCE OF IMPAIRED SIGNAL ON NEURO MONITORING, THE CAROTID ARTERY WAS LIGATED BELOW THE CLAVICLE. THE PATIENT WAS TRANSFERRED TO THE ICU FOR FURTHER CARE. IT WAS REPORTED THAT THE PATIENT WAS MOVING ALL EXTREMITIES IMMEDIATELY POST TAVR. A HEAD CT WAS PERFORMED THE DAY FOLLOWING SURGERY AND DID NOT SHOW OBVIOUS SIGNS OF STROKE, HOWEVER A REPEAT CT WAS PERFORMED 12 HOURS LATER AND SHOWED THE PATIENT HAD SUFFERED A LEFT HEMISPHERIC STROKE. THE PATIENT REMAINED INTUBATED AND SEDATED AND ALSO WENT IN TO RENAL FAILURE REQUIRING CRRT. NEUROSURGERY CONSULTED ON THE PATIENT AND DETERMINED THERE WAS NOTHING THEY COULD OFFER THE PATIENT. THE PATIENT'S FAMILY ESSENTIALLY MADE THE PATIENT A DNR AND HE WAS ALLOWED TO EXPIRE COMFORTABLY. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, SHEATH LINER TEAR EVENTS HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, AND/OR WITHDRAWAL OF AN ALTERED BALLOON PROFILE OF THE DELIVERY SYSTEM OR BAV. IN THIS CASE, THE COMPLAINT SHEATH LINER TORN WAS UNABLE TO BE CONFIRMED AS NO DEVICE AND/OR IMAGERY WERE PROVIDED. PATIENT FACTORS SUCH AS CALCIFICATION AND TORTUOSITY MAY CAUSE LINER TEAR. CALCIFICATION CAN DAMAGE THE EXPOSED PORTION OF THE SHEATH LINER, WHICH CAN LEAD TO IMMEDIATE CUTTING, TEARING, OR WEAKENING OF THE LINER. A WEAKENED LINER CAN ALSO TEAR DURING ADVANCEMENT OR RETRIEVAL OF THE DELIVERY SYSTEM OR BALLOON CATHETER. VESSEL TORTUOSITY CAN CREATE SUBOPTIMAL ANGLES DURING DELIVERY SYSTEM OR BALLOON CATHETER ADVANCEMENT/RETRIEVAL THROUGH THE SHEATH. NON-COAXIAL ALIGNMENT CAN CAUSE THE DELIVERY SYSTEM OR BALLOON CATHETER TO CATCH ONTO THE SHEATH LINER AND TEAR IT UPON ADVANCEMENT/RETRIEVAL. HOWEVER, DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY (B)(4), STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THE INFORMATION PROVIDED PROCEDURAL FACTORS (LINER TEAR AND WITHDRAW DIFFICULTY THROUGH LEFT COMMON CAROTID ARTERY) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. Manufacturer Narrative: THE INVESTIGATION IS IN PROGRESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093732&pc=DSQ | 2916596-2024-52862 | 2024-06-27 04:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS NO ADDITIONAL INFORMATION WILL BE PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093909&pc=LWQ | 2135147-2025-00177 | 2024-12-12 05:00 | Death | ABBOTT MEDICAL | 2025-01-07 05:00 | LWQ | REGENT HEART VALVE | Detachment of Device or Device Component | Low Blood Pressure/ Hypotension | P810002 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A 19MM REGENT AORTIC MECHANICAL HEART VALVE WAS SELECTED FOR IMPLANT. AFTER THE VALVE WAS PARACHUTED DOWN, THE HANDLE WAS REMOVED AND THE SUTURES WERE TIED. THE VALVE WAS NOT ATTEMPTED TO BE ROTATED. PRIOR TO PATIENT CLOSURE, THE VALVE LEAFLETS WERE INTENDED TO BE TESTED FOR MOBILITY. HOWEVER, IT WAS DISCOVERED THAT ONE OF THE LEAFLETS WAS MISSING. THE HOLDER HANDLE WAS FULLY INSERTED INTO THE ORIFICE AT 90 DEGREES. THE LEAFLET HAD DISLODGED IN ONE PIECE AND WAS RECOVERED FROM THE PATIENT. THE VALVE WAS EXPLANTED AND A NEW 21MM REGENT AORTIC MECHANICAL HEART VALVE WAS IMPLANTED AS THE SAME SIZE WAS NOT AVAILABLE. POST-OPERATIVELY, THE PATIENT REMAINED ON VENTILATOR SUPPORT. ON (B)(6) 2024, THE PATIENT STOPPED MAKING URINE AND WAS NOT HOLDING BLOOD PRESSURE. THEREFORE, INTRA-AORTIC BALLOON WAS INSERTED TO AUGMENT THE BLOOD PRESSURE. ON (B)(6) 2024, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN. THE 21MM REGENT VALVE REMAINS IMPLANTED IN THE PATIENT. HOWEVER, THE PATIENT ALLEGED THAT THE 19MM VALVE CONTRIBUTED TO THE PATIENT PASSING AND THAT THE 21MM REGENT VALVE DID NOT CONTRIBUTE TO THE PATIENT PASSING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF LEAFLET DISLODGMENT, HYPOTENSION, AND DEATH WAS REPORTED. ONE IMAGE AND TWO VIDEOS RECEIVED FROM THE FIELD SHOWED ONE LEAFLET TO BE DISLODGED FROM THE VALVE. THE VALVE WITH THE DISLODGED LEAFLET WERE RETURNED TO ABBOTT FOR A DETAILED INSPECTION AND EVALUATION. THE DEVICE CONFIRMED TO MEET FUNCTIONAL SPECIFICATION WITH NO DAMAGE OR ANOMALIES OBSERVED WITH THE DEVICE. A REVIEW OF THE DEVICE HISTORY RECORD CONFIRMED THAT THE VALVE WAS MANUFACTURED ACCORDING TO ALL REQUIRED SPECIFICATIONS AND PROCEDURES. THIS INCLUDED AN ASSESSMENT OF THE LEAFLET AND ORIFICE DIMENSIONAL MATCHING DATA, WHICH WAS FOUND TO BE COMPLIANT WITH THE ESTABLISHED REQUIREMENTS. ADDITIONALLY, THE REVIEW ENSURED THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND COMPLETED IN ACCORDANCE WITH ABBOTT SPECIFICATIONS, INCLUDING ALL CONTROL MEASURES AND FUNCTIONAL TESTING TO DETECT ANY IMPERFECTIONS IN THE VALVE¿S CONSTRUCTION. BASED ON THE AVAILABLE INFORMATION, THE EXACT MECHANISM THAT CAUSED THE LEAFLET DISLODGEMENT IN THIS CASE COULD NOT BE CONCLUSIVELY DETERMINED. THE CAUSE OF THE REPORTED DEATH AND HYPOTENSION COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21093966&pc=NRY | 2029214-2025-00071 | 2024-05-17 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-07 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING OUTCOMES OF MECHANICAL THROMBECTOMY FOR PATIENTS WITH STROKE PRESENTING WITH LOW ALBERTA STROKE PROGRAM EARLY COMPUTED TOMOGRAPHY SCORE IN EARLY AND LATE TIME WINDOWS. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION, THOUGH IT WAS NOT CLARIFIED WHAT DEVICES WERE USED. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. AMONG PATIENT ADVERSE EVENTS INCLUDED: 90-DAY MORTALITY OCCURRED IN 117/274 PATIENTS. SYMPTOMATIC INTRACRANIAL HEMORRHAGE OCCURRED IN 44/274 PATIENTS. ANY HEMORRHAGE OCCURRED IN 117/274 PATIENTS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: ELAWADY, S. S., CUNNINGHAM, C., HIDETOSHI MATSUKAWA, KAZUTAKA UCHIDA, LIN, S., MAIER, I., JABBOUR, P., JOON-TAE KIM, WOLFE, S. Q., RAI, A., STARKE, R. M., PSYCHOGIOS, M.-N., SAMANIEGO, E. A., ET AL. OUTCOMES OF MECHANICAL THROMBECTOMY FOR PATIENTS WITH STROKE PRESENTING WITH LOW ALBERTA STROKE PROGRAM EARLY COMPUTED TOMOGRAPHY SCORE IN EARLY AND LATE TIME WINDOWS. NEUROSURGERY 95(4) 2024. DOI: 10.1227/NEU.0000000000002992. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21094320&pc=BZD | 2518422-2025-000533 | 2023-12-29 05:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 1/30/2025 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G AND H IS UPDATED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21094387&pc=OZD | 1220648-2025-25386 | 2024-12-20 05:00 | Death | ABIOMED, INC. | 2025-01-07 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Arrhythmia; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT TWO DAYS INTO SUPPORT, AN IMPLANTED IMPELLA CP PUMP EXPERIENCED A SUDDEN SUCTION PROBLEM COINCIDING WITH A SELF-LIMITED VENTRICULAR TACHYCARDIA. THE MOTOR CURRENT PEAKED INTERMITTENTLY WITH PERSISTENT SUCTION ISSUES FOR A "FEW" HOURS. THE STAFF WAS ADVISED TO EXPLANT THE PUMP, BUT THE DECISION WAS MADE TO MAINTAIN SUPPORT.Event Description: DESPITE THAT DECISION TO KEEP THE PATIENT ON SUPPORT, THE PUMP WAS REMOVED LATER THAT MORNING AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO DISASSOCIATE THE USE OF THE IMPELLA WITH THE PATIENT'S PASSING. Manufacturer Narrative: B.2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386 TO REFLECT THE PATIENTS DEATH AND DATE OF DEATH. REQUIRED INTERVENTION WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386. B.5 ADDED PATIENT OUTCOME INFORMATION AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386. E.1 REVISED FACILITY NAME AS IT WAS PREVIOUSLY REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386. G.1 MANUFACTURING SITE NAME AND ADDRESS WAS CORRECT ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386. THIS SECTION WAS REPORTED INCORRECTLY ON THE FOLLOW UP 1 MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386 AND WAS REVISED BACK TO WHAT WAS REFLECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386 H.1 TYPE OF REPORTABLE EVENT WAS REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386 TO REFLECT THE PATIENTS OUTCOME. H.6 CODES 1721, 4611, 3331 WERE ALL REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386. CODES 2095, 4110 AND 1802 WERE ADDED AS THEY WERE INADVERTENTLY OMITTED FROM PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE (IFUS) WERE REPORTED ON THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25386 INCORRECTLY. NO IFUS ARE NEEDED TO BE REPORTED FOR THIS REPORT IN ACCORDANCE WITH UPDATED PROCEDURES. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9, H3, H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. DEVICE HISTORY RECORD REVIEW CONFIRMED THAT THE PUMP PASSED ALL POST-STERILE INSPECTION CHECKS. THE PUMP WAS RETURNED FOR INVESTIGATION. NO PHYSICAL ABNORMALITIES WERE REPORTED ON THE RETURNED PRODUCT. SOME SMALL CLEAR BIOMATERIAL WAS NOTED IN THE PURGE GAP, BUT IT WAS ALSO CLEARED OUT DURING THE TEST RUN IN A BUCKET OF WATER. THE PUMP OPERATED AS PER THE SPECIFICATIONS WITHOUT ANY NOTED LOW PUMP FLOW. THE DATA LOG SHOWS THAT THE PUMP FLOW REDUCED AFTER THE FIRST MOTOR CURRENT SPIKE DURING THE CASE, AND A FEW MORE MOTOR CURRENT SPIKES WERE REPORTED IN THE 2 HOURS BEFORE THE FLOW RETURNED TO THE NORMAL RANGE. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON DATA LOG REVIEW AND RETURN PRODUCT INVESTIGATION. CARDIAC ARRYTHMIA WAS REPORTED DURING IMPELLA SUPPORT. TO MAKE A DETERMINATION AS TO THE CAUSE OF THE CARDIAC ARRYTHMIA, CLINICAL INFORMATION MUST BE CONSIDERED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, AN INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21094449&pc=DSQ | 2916596-2024-52800 | 2023-09-13 04:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. NO FURTHER INFORMATION WILL BE PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095302&pc=BZD | 2518422-2025-000538 | 2024-02-23 05:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973 AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095304&pc=BZD | 2518422-2025-000537 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973 AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095516&pc=BZD | 2518422-2025-000539 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Asthma; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGATION OF ASTHMA (NEW OR WORSENING); LIVER DISEASE/TOXICITY; LUNG DISEASE; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095517&pc=BZD | 2518422-2025-000536 | 2024-04-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095630&pc=BZD | 2518422-2025-000542 | 2024-07-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON JUNE 04, 2025, AND H11 SHOULD BE REPORTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PREVIOUSLY ALLEGED VISUALIZATION OF PARTICLES. THERE WAS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS NOT OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION. EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID WAS UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095795&pc=BZD | 2518422-2025-000541 | 2024-07-29 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095914&pc=MKJ | 1220908-2024-04743 | 2024-12-17 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-07 05:00 | MKJ | R SERIES | Failure to Deliver Shock/Stimulation | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING INCLUDING STRESS, SHOCK, PATIENT IMPEDANCE, AND DEFIB CYCLE TESTS WITHOUT DUPLICATING THE REPORT. A REVIEW OF THE DEVICE LOG SUGGESTS THE EVENT DATE MAY HAVE BEEN DECEMBER 18, 2024. THERE ARE INSTANCES OF POOR PAD CONTACT, CHECK PADS, AND ECG LEAD OFF. THERE WAS ONE ATTEMPT OF ANALYSIS, BUT IT WAS HALTED WITH POOR PAD CONTACT AND CHECK PADS MESSAGES. THERE IS NO INDICATION FROM THE LOG THAT THE DEVICE HAS A MALFUNCTION. IF THERAPY ELECTRODES ARE NOT MAKING GOOD CONTACT WITH THE PATIENT'S SKIN, THE UNIT ISSUES ADVISORY MESSAGES "CHECK PADS" AND "POOR PAD CONTACT" AND DOES NOT ALLOW DELIVERY OF ENERGY. THE CUSTOMER WAS CONTACTED FOR ADDITIONAL INFORMATION ABOUT ELECTRODE PADS USED, PLACEMENT, AND PREPARATION BUT WE WERE UNSUCCESSFUL WORKING THROUGH A THIRD PARTY BIOMED. THE ANALOG BOARD WAS REPLACED AS A PRECAUTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095943&pc=BZD | 2518422-2025-100195 | 2024-11-22 05:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE PASSED AWAY AND THE PATIENT¿S DAUGHTER IS UNABLE TO RETURN THE AFFECTED DEVICE BECAUSE THEY NO LONGER HAVE IT. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21095945&pc=BZD | 2518422-2025-100197 | 2024-11-18 05:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096003&pc=DSQ | 2916596-2024-52860 | 2024-12-13 05:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cyanosis; Tachycardia; Hypovolemia; Cardiogenic Shock; Hypervolemia; Hyponatremia; Alteration in Body Temperature | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK. THE PATIENT WAS TRANSFERRED FOR HIGH LEVEL OF CARE AND INITIALLY PRESENTED WITH FLUID OVERLOAD AND HYPONATREMIA. THEY THEN DEVELOPED CARDIOGENIC SHOCK AND WERE STARTED ON DOPAMINE PRIOR TO TRANSFER. ON ARRIVAL, THE PATIENT'S CENTRAL VENOUS PRESSURE IN THE 20'S. THE PATIENT WAS COLD AND DRY ON EXAM WITH CYANOSIS IN THEIR TOES. MEAN ARTERIAL PRESSURE WAS IN THE 50S, AND HEARTRATE WAS 110 SINUS TACHYCARDIA. THE PATIENT REMAINED ON DOPAMINE AND WAS SEDATED WITH PROPOFOL. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) AT 6600 ROTATIONS PER MINUTE (RPM) AND WAS FLOWING AT 4 LITERS. THERE WERE NO SIGNIFICANT PULSATILITY INDEX (PI) EVENTS OR ALARMS. IT IS UNKNOWN IF THE DEATH WAS CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU, REV. D, IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION", LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096017&pc=DSQ | 2916596-2024-52886 | 2024-03-17 04:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096170&pc=BZD | 2518422-2025-100196 | 2024-11-18 05:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME DUE TO A USER OF DREAMSTATION CPAP. DAUGHTER CALLED IN STATING PATIENT PASSED AWAY LAST (B)(6) 2023. DID NOT PASS DUE TO THE USE OF CPAP. SHE WANTS TO STOP ALL MAIL. DON'T KNOW WHERE THE CPAP DEVICE WENT THAT WAS RECALLED AND WILL NOT BE RETURNING FOR EVALUATION. REPLACEMENT DEVICE WAS DONATED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096177&pc=DSQ | 2916596-2024-52888 | 2024-10-09 04:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A SUSPECTED CEREBROVASCULAR ACCIDENT (CVA).Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B5 NARRATIVE INFORMATION. H6 - ADVERSE EVENT PROBLEM UPDATE CODES. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096372&pc=DSQ | 2916596-2024-52908 | 2024-10-02 04:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO THROMBOSIS/EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO).Event Description: THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED AND THERE WERE NO CHANGES TO PUMP PARAMETERS. A COMPUTED TOMOGRAPHY (CT) SCAN OF THE CHEST REVEALED MASSIVE THROMBOSIS OF THE LEFT VENTRICULAR ASSIST DEVICE (LVAD), HOWEVER NO CT IMAGES WERE AVAILABLE. THE PATIENT EXPERIENCED PERSISTENT LOW FLOW ALARMS AS A RESULT OF THE PUMP THROMBOSIS. THE LVAD THROMBOSIS CONTINUED TO PROGRESS DESPITE THE PATIENT BEING ON THERAPEUTIC ANTICOAGULATION AT THE TIME OF ADMISSION. Manufacturer Narrative: SECTION B: DATE OF DEATH CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED DEVICE THROMBUS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED, AS NO LOG FILES WERE SUBMITTED FOR REVIEW. A SPECIFIC CAUSE FOR THE LOW FLOW ALARMS AS WELL AS A DIRECT CORRELATION TO THE REPORTED THROMBOSIS COULD NOT BE CONCLUSIVELY ESTABLISHED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND THE HEARTMATE 3 PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 OF THE IFU ALSO EXPLAINS THAT PI EVENTS ARE ASSUMED BY THE SYSTEM DURING CASES WHEN THERE ARE SUDDEN AND SUBSTANTIAL CHANGES IN THE PULSATILITY INDEX. THESE EVENTS MAY BE INITIATED FOR REASONS OTHER THAN TRUE PI EVENTS, INCLUDING SUDDEN CHANGES IN A PATIENT¿S VOLUME STATUS, ARRHYTHMIAS, SUDDEN CHANGES IN POWER, AND SUDDEN CHANGES IN PUMP SPEED. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION, UNDER ¿ANTICOAGULATION¿, ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5: ADDITIONAL INFORMATION. H6 HEALTH EFFECT - CLINICAL CODE UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096388&pc=DSQ | 2916596-2024-52911 | 2024-11-30 05:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Neuropathy; Coma; Ischemia Stroke; Embolism/Embolus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024 THE PATIENT HAD A CEREBRAL INFARCTION. LOG FILES CAPTURED LOW FLOW EVENTS ON (B)(6) 2024. SURGERY WAS PLANNED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT ON 18DEC2024 THE PATIENT HAD NEUROLOGICAL DYSFUNCTION. THE PATIENT HAD A BILATERAL EMBOLIC STROKE. THEY WENT INTO A COMA. THE STROKE WAS DIAGNOSED WITH A COMPUTED TOMOGRAPHY (CT) SCAN. THE PATIENT WAS ON HEPARIN ANTICOAGULATION AND RECEIVED ANTI-SEIZURE MEDICATION. THE CAUSE OF THE NEUROPATHY WAS THOUGHT TO BE COMPLEXITY OF DEVICE-RELATED AND MEDICAL MANAGEMENT. THE DEVICE WAS POSSIBLY RELATED TO THE EVENT BECAUSE OF THROMBUS AND IT WAS UNLIKELY THAT THERE WAS A CAUSE OTHER THAN THE DEVICE.Event Description: IIT WAS ADDITIONALLY REPORTED THAT THERE WAS NO INFORMATION THAT THE PATIENT DIED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED LOW FLOW ALARMS. A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 24NOV2024 THROUGH 18DEC2024. LOW FLOW HAZARD ALARMS WERE CAPTURED ON 30NOV2024. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. LOG FILES ANALYSES WERE REQUESTED TO CONFIRM IF THERE ARE ANY PROBLEMS WITH THE DEVICE DUE TO CEREBRAL INFARCTION. THE PATIENT REMAINS ONGOING ON HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND NO FURTHER RELATED EVENTS HAVE BEEN REPORTED AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. THE IFU, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THE IFU, PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE PATIENT HANDBOOK, AND THE IFU, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, THE PATIENT HANDBOOK, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B5 NARRATIVE INFORMATION: H1 - TYPE OF REPORTABLE EVENT UPDATED. H6 - ADVERSE EVENT PROBLEM CODES UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096422&pc=QFG | 3013756811-2025-00739 | 2024-12-09 05:00 | Death | TANDEM DIABETES CARE | 2025-01-07 05:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. PER THE CONTACT, THE CAUSE OF DEATH IS UNKNOWN, BUT ALLEGED THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE EVENT. NO ADDITIONAL INFORMATION WAS PROVIDED. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096623&pc=CAW | 2518422-2025-000549 | 2024-05-16 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | CAW | RECERT SIMPLYGO MINI | Degraded | Insufficient Information | K111885 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE THAT A USER OF THE RECERT SIMPLYGO MINI HAD PASSED AWAY AND THE DEVICE WAS ELIGIBLE FOR BUYBACK. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096627&pc=BZD | 2518422-2025-000545 | 2024-06-26 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21096724&pc=DSQ | 2916596-2024-52890 | 2024-10-27 04:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE MEDICAL CENTER DUE TO ACUTE ENCEPHALOPATHY SECONDARY TO INTRAVENTRICULAR HEMORRHAGE. THE PATIENT WAS INITIALLY TREATED WITH URGENT WARFARIN WITH PROTHROMBIN COMPLEX CONCENTRATE. SERIAL COMPUTED TOMOGRAPHY FOR THE HEAD (CTH) DEMONSTRATED NO PROGRESSION OF THE INTRAVENTRICULAR HEMORRHAGE. THE PATIENT'S DESIGNATED DECISION MAKERS EVENTUALLY ELECTED TO WITHDRAW LIFE SUPPORT AND PURSUE COMFORT MEASURES ONLY. THE LIFE SUPPORT WAS WITHDRAWN ON (B)(6) 2024. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: D4 - UDI: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097497&pc=NVN | 2017865-2025-10475 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-07 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-10473. RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-10474. IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097501&pc=KDI | 0001713747-2025-00025 | 2024-12-10 05:00 | Death | OGDEN MANUFACTURING PLANT | 2025-01-07 05:00 | KDI | OPTIFLUX F180NRE HF 12/CS 1.7SA EBEAM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Loss of consciousness; Insufficient Information | K162488 |  | 1 | Event Description: ON (B)(6) 2024 FRESENIUS WAS NOTIFIED THAT A PATIENT PASSED AWAY WHILE UTILIZING THE 2008T MACHINE FOR HEMODIALYSIS (HD) TREATMENT. A REQUEST FOR MACHINE AND DISPOSABLE EVALUATION WAS MADE. A FIELD SERVICE TECHNICIAN (FST) WAS REQUESTED TO PERFORM AN INSPECTION OF THE MACHINE. INSPECTION COMPLETED. AN ADVERSE EVENT PRE-SERVICE EVALUATION CHECKLIST AND ULTRAFILTRATION (UF) PROBLEM IDENTIFICATION CHECKLIST WAS PERFORMED. A SMALL LEAK WAS OBSERVED ON THE BOTTOM OF THE AIR SEPARATOR. THE FAILURE ON THE ALARM TEST WAS THE 9 VOLT BATTERY.Event Description: ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM 29/DEC/2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER.Event Description: ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM (B)(6) 2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN HEMODIALYSIS UTILIZING THE 2008T MACHINE AND THE PATIENT DEATH DURING TREATMENT. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE 2008T MACHINE. ADDITIONALLY, THERE IS NO ALLEGATION OF MACHINE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE CIRCUMSTANCES SURROUNDING THE DEATH ARE UNKNOWN. THE PATIENT¿S HEALTH STATUS AND HEALTH HISTORY ARE UNKNOWN. THE CIRCUMSTANCES SURROUNDING THE PATIENT DEATH ARE UNKNOWN. THE ASSOCIATION BETWEEN CHRONIC KIDNEY DISEASE (CKD) AND HIGH CARDIOVASCULAR MORBIDITY AND MORTALITY IS WELL KNOWN. ACCORDING TO THE UNITED STATES RENAL DATA SYSTEM (USRDS), THE LEADING CAUSE OF DEATH AMONG CKD PATIENTS UNDERGOING DIALYSIS IS RELATED TO CARDIAC ARRHYTHMIAS. BASED ON THE LIMITED INFORMATION AND NO ALLEGATION OF A MACHINE MALFUNCTION OR DEFICIENCY, THE FRESENIUS 2008T MACHINE CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: A1, A2, A3, A4, B5, D10, G2, H6, H10. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HD THERAPY UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, COMBI-SET BLOODLINES AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, LOSS OF CONSCIOUSNESS, AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON THE POST-EVENT TESTING, THERE IS EVIDENCE TO SUPPORT THE EXISTENCE OF INTERNAL DEVICE FAILURES, WHICH MAY HAVE CAUSED THE 2008T HEMODIALYSIS SYSTEM TO PERFORM OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR APPROXIMATELY 45% OF ALL ESRD DEATHS, AND APPROXIMATELY 25% OF THOSE DEATHS ARE DUE TO SUDDEN CARDIAC DEATH. BASED ON THE INFORMATION AVAILABLE, THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS ACTIVELY UNDERGOING HD THERAPY WHEN HE WAS FOUND UNRESPONSIVE, THE FACILITY¿S REQUEST FOR POST-EVENT FUNCTIONAL COMPLIANCE TESTING, THE UF PUMP FAILURE REQUIRING RECALIBRATION, THE PRESSURE HOLDING TEST FAILURE, AND A PENDING MANUFACTURER EVALUATION OF THE DIALYZER/BLOODLINES, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE COMPLAINT DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A SAMPLE FROM THE REPORTED LOT WAS RETURNED. NO DAMAGE, DEFECT, OR IRREGULARITY WAS IDENTIFIED ON THE RETURNED SAMPLE. AS THE DIALYZER WAS USED FOR TREATMENT, IT COULD NOT BE FORWARDED TO THE QUALITY SYSTEMS LAB FOR FURTHER TESTING. DURING THE LOT HISTORY REVIEW IT WAS NOTED THAT THERE WERE TWO OTHER COMPLAINTS REPORTED AGAINST THE LOT. THERE WAS ONE EXTERNAL BLOOD LEAK (CONFIRMED POLYURETHANE SLIVER) AND ONE INTERNAL BLOOD LEAK THE OTHER COMPLAINT IS THE PREVIOUSLY CLOSED NO-SAMPLE INVESTIGATION OF THE SAME EVENT. NO COMPLAINTS HAVE BEEN RECEIVED REPORTING A CARDIAC EVENT OR PATIENT DEATH AND THERE WAS NO MENTION OF A DIALYZER LEAK IN THE CURRENT INVESTIGATION. THE NUMBER OF COMPLAINTS FOR THE ASSIGNED SYMPTOM CODE ON THE LOT NUMBER WAS REVIEWED AND IT WAS DETERMINED THAT NO LOT COMPLAINT EXCURSION OCCURRED. AN INVESTIGATION OF THE DEVICE HISTORY RECORDS (DHR) WAS CONDUCTED BY THE MANUFACTURER. THERE WERE MULTIPLE APPROVED TEMPORARY DEVIATION NOTICES (DN) AND ONE NON-CONFORMANCE (NC) REPORTED ON THE LOT WHICH WERE UNRELATED TO THE COMPLAINT EVENT. THERE WAS NO INDICATION OF PRODUCT NONACCEPTANCE, DEVIATION, NON-CONFORMANCE, REWORK, LABELING OR PROCESS CONTROL FAILURE DURING THE MANUFACTURING PROCESS WHICH COULD BE ASSOCIATED WITH THE REPORTED EVENT. THE LOT MET ALL RELEASE CRITERIA. BASED ON THE INFORMATION AVAILABLE, THE OPTIFLUX 180NRE DIALYZER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENT. THERE IS CURRENTLY NO ALLEGATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. CAPAS FOR VISION SYSTEMS AND BLOOD LEAK REDUCTION ARE RECENT EXAMPLES OF LEAK RELATED INVESTIGATIONS DIRECTED AT AN OVERALL REDUCTION IN DIALYZER LEAKS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097507&pc=LWP | 2017865-2025-10473 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-07 05:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-10474. RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-10475. IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097509&pc=FJK | 0008030665-2025-00054 | 2024-12-10 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-07 05:00 | FJK | CUSTOM COMBI SET | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Loss of consciousness; Insufficient Information | K962081 |  | 1 | Event Description: ON 10/DEC/2024 FRESENIUS WAS NOTIFIED THAT A PATIENT PASSED AWAY WHILE UTILIZING THE 2008T MACHINE FOR HEMODIALYSIS (HD) TREATMENT. A REQUEST FOR MACHINE AND DISPOSABLE EVALUATION WAS MADE. A FIELD SERVICE TECHNICIAN (FST) WAS REQUESTED TO PERFORM AN INSPECTION OF THE MACHINE. INSPECTION COMPLETED. AN ADVERSE EVENT PRE-SERVICE EVALUATION CHECKLIST AND ULTRAFILTRATION (UF) PROBLEM IDENTIFICATION CHECKLIST WAS PERFORMED. A SMALL LEAK WAS OBSERVED ON THE BOTTOM OF THE AIR SEPARATOR. THE FAILURE ON THE ALARM TEST WAS THE 9 VOLT BATTERY.Event Description: ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM 29/DEC/2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER.Event Description: ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM 29/DEC/2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER.Event Description: ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6)2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM (B)(6)2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN HEMODIALYSIS UTILIZING THE 2008T MACHINE AND THE PATIENT DEATH DURING TREATMENT. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE 2008T MACHINE. ADDITIONALLY, THERE IS NO ALLEGATION OF MACHINE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE CIRCUMSTANCES SURROUNDING THE DEATH ARE UNKNOWN. THE PATIENT¿S HEALTH STATUS AND HEALTH HISTORY ARE UNKNOWN. THE CIRCUMSTANCES SURROUNDING THE PATIENT DEATH ARE UNKNOWN. THE ASSOCIATION BETWEEN CHRONIC KIDNEY DISEASE (CKD) AND HIGH CARDIOVASCULAR MORBIDITY AND MORTALITY IS WELL KNOWN. ACCORDING TO THE UNITED STATES RENAL DATA SYSTEM (USRDS), THE LEADING CAUSE OF DEATH AMONG CKD PATIENTS UNDERGOING DIALYSIS IS RELATED TO CARDIAC ARRHYTHMIAS. BASED ON THE LIMITED INFORMATION AND NO ALLEGATION OF A MACHINE MALFUNCTION OR DEFICIENCY, THE FRESENIUS 2008T MACHINE CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER AND THE LOT NUMBER WAS NOT PROVIDED. A MANUFACTURING REVIEW WAS PERFORMED ON THE PRODUCTS SHIPPED TO THE PATIENT FOR THE THREE (3) MONTH TIME FRAME WHICH IMMEDIATELY PRECEDED THE EVENT OCCURRENCE DATE. THIS REVIEW INCLUDED THE LOT NUMBERS FOR ALL FRESENIUS LIBERTY CYCLER SETS SHIPPED TO THIS ACCOUNT WITHIN THE SELECTED TIME FRAME. THE ENTIRE SET OF LOTS HAVE BEEN SOLD AND DISTRIBUTED. THERE WERE NO NON-CONFORMANCES OR ABNORMALITIES IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE ASSOCIATED WITH THE REPORTED EVENT. AN INVESTIGATION OF THE DEVICE HISTORY RECORDS (DHR) WAS CONDUCTED AND CONFIRMED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE PRODUCT LOTS INVOLVED MET ALL SPECIFICATIONS FOR RELEASE. A REVIEW OF THE DHR DID NOT REVEAL A PROBABLE CAUSE FOR THE CUSTOMER COMPLAINT. AS A PHYSICAL EVALUATION COULD NOT BE PERFORMED, A DEFINITIVE CONCLUSION REGARDING THE REPORTED INCIDENT COULD NOT BE REACHED AND A CAUSE COULD NOT BE CONFIRMED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE PLANT RECEIVED ON (B)(6)2025 A COMPLAINT PRODUCT SAMPLE FROM THE CUSTOMER WITHOUT ORIGINAL PACKAGE. THE COMPLAINT PRODUCT SAMPLE WAS VISUALLY INSPECTED AND NO DAMAGES WERE OBSERVED IN THE LINES AND CHAMBERS; THE COMBI SET IS ACCEPTABLE. IN ADDITION, THE SAMPLE WAS TESTED IN THE HEMODIALYSIS MACHINE AND WORKED AS INTENDED WITHOUT ANY ABNORMALITIES DURING THE TESTED PERIOD. NO DISCONNECTION OR LEAK OCCURRED DURING TESTED PERIOD FROM THE PATIENT VENOUS AND ARTERIAL CONNECTOR TO CATHETER. NO AIR BUBBLES INSIDE THE SYSTEM, NO LEAKS, NO LEVEL VARIATION OF VENOUS AND ARTERIAL CHAMBER OR ANY ALARM WAS NOTICED. THE SAMPLE TESTED WORKED AS INTENDED WITHOUT ANY ABNORMALITIES DURING THE TESTED PERIOD. THE LOT NUMBER WAS NOT PROVIDED. A MANUFACTURING REVIEW WAS PERFORMED ON THE PRODUCTS SHIPPED TO THE PATIENT FOR THE THREE (3) MONTH TIME FRAME WHICH IMMEDIATELY PRECEDED THE EVENT OCCURRENCE DATE. THIS REVIEW INCLUDED THE LOT NUMBERS FOR ALL FRESENIUS LIBERTY CYCLER SETS SHIPPED TO THIS ACCOUNT WITHIN THE SELECTED TIME FRAME. THE ENTIRE SET OF LOTS HAVE BEEN SOLD AND DISTRIBUTED. THERE WERE NO NON-CONFORMANCES OR ABNORMALITIES IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE ASSOCIATED WITH THE REPORTED EVENT. AN INVESTIGATION OF THE DEVICE HISTORY RECORDS (DHR) WAS CONDUCTED AND CONFIRMED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE PRODUCT LOTS INVOLVED MET ALL SPECIFICATIONS FOR RELEASE. A REVIEW OF THE DHR DID NOT REVEAL A PROBABLE CAUSE FOR THE CUSTOMER COMPLAINT. A DEFINITIVE CONCLUSION REGARDING THE REPORTED INCIDENT COULD NOT BE REACHED AND A CAUSE COULD NOT BE CONFIRMED. Manufacturer Narrative: ADDITIONAL INFORMATION: A1, A2, A3, A4, B5, G2, H6, H10 CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HD THERAPY UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, COMBI-SET BLOODLINES AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, LOSS OF CONSCIOUSNESS, AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON THE POST-EVENT TESTING, THERE IS EVIDENCE TO SUPPORT THE EXISTENCE OF INTERNAL DEVICE FAILURES, WHICH MAY HAVE CAUSED THE 2008T HEMODIALYSIS SYSTEM TO PERFORM OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR APPROXIMATELY 45% OF ALL ESRD DEATHS, AND APPROXIMATELY 25% OF THOSE DEATHS ARE DUE TO SUDDEN CARDIAC DEATH. BASED ON THE INFORMATION AVAILABLE, THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS ACTIVELY UNDERGOING HD THERAPY WHEN HE WAS FOUND UNRESPONSIVE, THE FACILITY¿S REQUEST FOR POST-EVENT FUNCTIONAL COMPLIANCE TESTING, THE UF PUMP FAILURE REQUIRING RECALIBRATION, THE PRESSURE HOLDING TEST FAILURE, AND A PENDING MANUFACTURER EVALUATION OF THE DIALYZER/BLOODLINES, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097526&pc=FJI | 9611369-2024-00239 | 2024-12-06 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-07 05:00 | FJI | POLYFLUX | Adverse Event Without Identified Device or Use Problem | Dyspnea; Hypersensitivity/Allergic reaction; Swelling/ Edema | NI |  | 1 | Event Description: IT WAS REPORTED THAT ONE HOUR INTO TREATMENT WITH POLYFLUX DIALYZER, THE PATIENT WAS REPORTED TO HAVE ¿ALLERGIC SYMPTOMS.¿ IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED DYSPNEA, SHORTNESS OF BREATH, EYELID EDEMA AND FACIAL EDEMA. DIALYSIS WAS DISCONTINUED AND THE PATIENT WAS GIVEN UNSPECIFIED ¿RESCUE TREATMENT.¿ THE PATIENT WAS TRANSFERRED TO A ¿SUPERIOR¿ HOSPITAL FOR TREATMENT AND SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. THERE IS NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: E1 INITIAL REPORTER FACILITY - (B)(6) HOSPITAL. E1 INITIAL REPORTER ADDRESS - (B)(6). E1 INITIAL REPORTER CITY - (B)(6). H11: A BATCH REVIEW WAS CONDUCTED AND THERE WERE NO DEVIATIONS FOUND RELATED TO THIS REPORTED CONDITION DURING THE MANUFACTURE OF THIS LOT. THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. THE CAUSE OF THE CONDITION COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097650&pc=BZD | 2518422-2025-000554 | 2024-06-10 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Unspecified Heart Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: IN THE PREVIOUS REPORT PATIENT OUTCOME GRID WAS NOT ADDED. IN THIS REPORT PATIENT OUTCOME CODE GRID (BOX H) HAS BEEN UPDATED OR CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097671&pc=DSQ | 2916596-2024-52885 | 2024-12-17 05:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Infusion or Flow Problem | Dyspnea; Loss of consciousness; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS FOUND UNRESPONSIVE AND HAD LOW FLOWS. THE PATIENT WAS UNABLE TO BE RESUSCITATED. THE LOG FILES WERE REQUESTED TO BE REVIEWED. THE LOG FILES WERE REVIEWED AND CAPTURED INTERMITTENT LOW FLOW EVENTS THROUGHOUT THE LOG FILE WITH FLOW NOTED AS LOW AS 1.4 LPM. THE LOWER FLOWS WERE HOVERING MAINLY AROUND THE 2.4 TO 2.5 LPM RANGE BUT DID SEE A DROP INTO THE 1 RANGE PRIOR TO THE DRIVELINE BEING DISCONNECTED ON (B)(6) 2024. THESE LOW FLOW EVENTS WERE LIKELY PATIENT RELATED AS THEY WERE ASSOCIATED WITH ELEVATED PULSATILITY INDEX (PI) VALUES.Event Description: THE PATIENT WAS FOUND DYSPNEIC AND UNRESPONSIVE AND THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS LOW FLOWING. THE PATIENT WAS ATTEMPTED TO BE RESUSCITATED WITH ADVANCED CARDIAC LIFE SUPPORT (ACLS) BUT WERE UNABLE TO ACHIEVE RETURN OF SPONTANEOUS CIRCULATION (ROSC). THERE WAS UNKNOWN CAUSE OF DYSPNEA/ALTERED MENTAL STATUS. THE DEATH WAS NOT DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. NO PRODUCTS WOULD BE RETURNED. THE DRIVELINE WAS CONFIRMED TO BE DISCONNECTED AFTER THE PATIENT'S DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE COULD NOT BE DETERMINED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2024, PER THE TIMESTAMPS. LOW FLOW ALARMS WERE CAPTURED THROUGHOUT THE FILE. A DRIVELINE DISCONNECT ALARM WAS CAPTURED, CONSISTENT WITH THE REMOVAL OF DEVICE SUPPORT FOLLOWING THE PATIENT'S DEATH. THE PUMP OPERATED AT THE SET SPEED THROUGHOUT THE ENTIRETY OF THE FILE WHILE THE DRIVELINE WAS CONNECTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDE INFORMATION ON SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097682&pc=LWP | 2017865-2025-10476 | 2024-12-18 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-07 05:00 | LWP | ASSURITY MRI | Use of Device Problem; Insufficient Information; Noise, Audible | Arrhythmia; Discomfort; Syncope/Fainting | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10478. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10480. IT WAS REPORTED THE PATIENT EXPERIENCED DISCOMFORT, SYNCOPE AND ARRYTHMIA AND PASSED AWAY. THE HEALTHCARE PROVIDER INDICATED AN ANOMALY ON THE PATIENT ELECTROCARDIOGRAM PRIOR TO DEATH AND THE PHYSICIAN ALLEGED THAT A PACER FAILURE CAUSED THE DEATH. THE DEVICE WAS NOT EXPLANTED OR RETURNED.Event Description: NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ALLEGED TO BE VENTRICULAR FIBRILLATION, AND PATIENT USER CONCERN WAS EXPRESSED LEADING UP TO THE EVENT. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097685&pc=NVN | 2017865-2025-10478 | 2024-12-18 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-07 05:00 | NVN | TENDRIL STS | Use of Device Problem; Insufficient Information | Arrhythmia; Discomfort; Syncope/Fainting | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10476. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10480. IT WAS REPORTED THE PATIENT EXPERIENCED DISCOMFORT, SYNCOPE AND ARRYTHMIA AND PASSED AWAY. THE HEALTHCARE PROVIDER INDICATED AN ANOMALY ON THE PATIENT ELECTROCARDIOGRAM PRIOR TO DEATH AND THE PHYSICIAN ALLEGED THAT A PACER FAILURE CAUSED THE DEATH. THE DEVICE WAS NOT EXPLANTED OR RETURNED.Event Description: NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ALLEGED TO BE VENTRICULAR FIBRILLATION, AND PATIENT USER CONCERN WAS EXPRESSED LEADING UP TO THE EVENT. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097691&pc=NVN | 2017865-2025-10480 | 2024-12-18 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-07 05:00 | NVN | TENDRIL STS | Use of Device Problem; Insufficient Information | Arrhythmia; Discomfort; Syncope/Fainting | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10476. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-10478. IT WAS REPORTED THE PATIENT EXPERIENCED DISCOMFORT, SYNCOPE AND ARRYTHMIA AND PASSED AWAY. THE HEALTHCARE PROVIDER INDICATED AN ANOMALY ON THE PATIENT ELECTROCARDIOGRAM PRIOR TO DEATH AND THE PHYSICIAN ALLEGED THAT A PACER FAILURE CAUSED THE DEATH. THE DEVICE WAS NOT EXPLANTED OR RETURNED.Event Description: NEW INFORMATION RECEIVED NOTES THAT THE CAUSE OF DEATH WAS ALLEGED TO BE VENTRICULAR FIBRILLATION, AND PATIENT USER CONCERN WAS EXPRESSED LEADING UP TO THE EVENT. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097823&pc=NRY | 2029214-2025-00075 | 2024-06-21 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-07 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion; Thrombosis/Thrombus |  |  | 1 | Event Description: MEDTRONIC LITERATURE REVIEW FOUND REPORTED OF DEATH DUE TO CHRONIC REJECTION, THROMBOSIS, DISSECTION, HEPATIC ARTERY STENT OCCLUSION, AND BILIARY NECROSIS REQUIRING RETRANSPLANTATION OF LIVER TRANSPLANT IN ASSOCIATION WITH SOLITAIRE THROMBECTOMY AND MARKSMAN CATHETER. THE TIME FRAME OF THIS STUDY WAS FROM OCTOBER 2015 TO NOVEMBER 2021. THE PURPOSE OF THIS ARTICLE WAS TO INVESTIGATE AND EVALUATE THE EFFECTIVENESS AND SAFETY OF MECHANICAL THROMBECTOMY, SPECIFICALLY UTILIZING STENT RETRIEVER AND/OR ASPIRATION THROMBECTOMY, FOR THE TREATMENT OF HEPATIC ARTERY THROMBOSIS (HAT) IN PATIENTS WITH ORTHOTOPIC LIVER TRANSPLANT (OLT). THE AUTHORS REVIEWED 8 CASES OF PATIENTS TREATED FOR HEPATIC ARTERY THROMBOSIS USING SOLITAIRE THROMBECTOMY AND A MARKSMAN CATHETER. OF THE 8 PATIENTS, THE AVERAGE AGE WAS 54 YEARS, 1 WAS FEMALE AND 7 WERE MALE. ALL HAD TECHNICALLY SUCCESSFUL RESTORATION OF FLOW WITH STENT PLACEMENT OF THE ANASTOMOTIC STENOSIS IN 8 CASES. THE ARTICLE DOES NOT STATE ANY TECHNICAL ISSUES DURING USE OF THE SOLITAIRE STENT OR MARKSMAN CATHETER. THE FOLLOWING INTRA- OR POST-PROCEDURAL OUTCOMES WERE NOTED: ONE PATIENT (PATIENT 1) DIED FROM LIVER FAILURE DUE TO CHRONIC REJECTION AND HAD A PATENT HEPATIC ARTERY ON EXPLANT. THIS PATIENT HAD A THROMBOSIS OCCUR AGAIN 2 WEEKS AFTER THE PRIMARY PROCEDURE AND 2 DAYS AFTER THE SECOND AND DIED SHORTLY AFTER THE SECOND PROCEDURE. DISSECTION OF THE HEPATIC ARTERY IN 1 PATIENT HEPATIC ARTERY STENT OCCLUSION IN 1 PATIENT RETRANSPLANTATION DUE TO BILIARY NECROSIS- HEPATIC ARTERY WAS PATENT A TIME OF EXPLANT.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE HEALTHCARE PROVIDER (HCP) LOOKED BACK AT OUR CASES AND THE ADVERSE EVENTS WERE NOT TIED TO EITHER THE SOLITAIRE OR MARKSMAN DEVICES. ALL COMPLICATIONS WERE FELT TO BE A RESULT OF THE OVERALL CLINICAL STATUS AND NOT RELATED TO THE DEVICES USED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097852&pc=KRA | 2029214-2025-00076 | 2024-06-21 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-07 05:00 | KRA | MARKSMAN | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion; Thrombosis/Thrombus |  |  | 1 | Event Description: MEDTRONIC LITERATURE REVIEW FOUND REPORTED OF DEATH DUE TO CHRONIC REJECTION, THROMBOSIS, DISSECTION, HEPATIC ARTERY STENT OCCLUSION, AND BILIARY NECROSIS REQUIRING RETRANSPLANTATION OF LIVER TRANSPLANT IN ASSOCIATION WITH SOLITAIRE THROMBECTOMY AND MARKSMAN CATHETER. THE TIME FRAME OF THIS STUDY WAS FROM OCTOBER 2015 TO NOVEMBER 2021. THE PURPOSE OF THIS ARTICLE WAS TO INVESTIGATE AND EVALUATE THE EFFECTIVENESS AND SAFETY OF MECHANICAL THROMBECTOMY, SPECIFICALLY UTILIZING STENT RETRIEVER AND/OR ASPIRATION THROMBECTOMY, FOR THE TREATMENT OF HEPATIC ARTERY THROMBOSIS (HAT) IN PATIENTS WITH ORTHOTOPIC LIVER TRANSPLANT (OLT). THE AUTHORS REVIEWED 8 CASES OF PATIENTS TREATED FOR HEPATIC ARTERY THROMBOSIS USING SOLITAIRE THROMBECTOMY AND A MARKSMAN CATHETER. OF THE 8 PATIENTS, THE AVERAGE AGE WAS 54 YEARS, 1 WAS FEMALE AND 7 WERE MALE. ALL HAD TECHNICALLY SUCCESSFUL RESTORATION OF FLOW WITH STENT PLACEMENT OF THE ANASTOMOTIC STENOSIS IN 8 CASES. THE ARTICLE DOES NOT STATE ANY TECHNICAL ISSUES DURING USE OF THE SOLITAIRE STENT OR MARKSMAN CATHETER. THE FOLLOWING INTRA- OR POST-PROCEDURAL OUTCOMES WERE NOTED: ONE PATIENT (PATIENT 1) DIED FROM LIVER FAILURE DUE TO CHRONIC REJECTION AND HAD A PATENT HEPATIC ARTERY ON EXPLANT. THIS PATIENT HAD A THROMBOSIS OCCUR AGAIN 2 WEEKS AFTER THE PRIMARY PROCEDURE AND 2 DAYS AFTER THE SECOND AND DIED SHORTLY AFTER THE SECOND PROCEDURE. DISSECTION OF THE HEPATIC ARTERY IN 1 PATIENT HEPATIC ARTERY STENT OCCLUSION IN 1 PATIENT RETRANSPLANTATION DUE TO BILIARY NECROSIS- HEPATIC ARTERY WAS PATENT A TIME OF EXPLANT.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE HEALTHCARE PROVIDER (HCP) LOOKED BACK AT OUR CASES AND THE ADVERSE EVENTS WERE NOT TIED TO EITHER THE SOLITAIRE OR MARKSMAN DEVICES. ALL COMPLICATIONS WERE FELT TO BE A RESULT OF THE OVERALL CLINICAL STATUS AND NOT RELATED TO THE DEVICES USED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21097955&pc=DSQ | 2916596-2024-52928 | 2024-12-16 05:00 | Death | THORATEC CORPORATION | 2025-01-07 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: E1 - REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21098150&pc=BZD | 2518422-2025-000546 | 2024-06-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-07 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY AND HE HAD LUNG CANCER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106431&pc=QBJ | MW5164442 |  | Death | UNKNOWN | 2025-01-07 05:00 | QBJ | FREESTYLE LIBRE 2 |  | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: CALLER STATED THE PATIENT IS DECEASED. CALLER STATED PATIENT PASSED AWAY IN DECEMBER AND THAT IT WAS QUITE POSSIBLE PATIENT WAS STILL USING THE PRODUCT AT THE TIME OF THEIR PASSING. - REPORTED BY HCP WHO CONSENTED TO FOLLOW UP - (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21076944&pc=BZD | 2518422-2025-000212 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-06 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21077335&pc=BZD | 2518422-2025-000155 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-06 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE IS NO REPORT OF MEDICAL INTERVENTION BEING REQUIRED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21077542&pc=OZP | 2032227-2025-102953 | 2024-11-04 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-06 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hypoglycemia; Pneumonia; Irregular Pulse; Respiratory Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS ACUTE HYPOXEMIC RESPIRATORY FAILURE. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: MENTAL SOLEMPHOMA, HYPOVOLEMIA RENAL FAILURE. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-7040A, UNOMEDICAL, UNK_RESERVOIR, MMT-1884. TROUBLESHOOTING WAS PERFORMED. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A, UNOMEDICAL, UNK_RESERVOIR AND MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21078085&pc=PNJ | 9612164-2025-00076 | 2024-12-14 05:00 | Death | MEDTRONIC IRELAND | 2025-01-06 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANTATION OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), A CARDIAC TAMPONADE OCCURRED, RESULTING IN THE PATIENT'S DEATH. FULL CODE MEASUREMENTS WERE ADMINISTERED FOR APPROXIMATELY TWO HOURS BEFORE THE PATIENT EXPIRED.Event Description: IT WAS FURTHER REPORTED THAT THERE WAS NO INDICATION THAT THE PRODUCT HAD A PERFORMANCE ISSUE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION TO H6 THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE. MC2AVR1-DEL-SYS: (B)(6): CORRECTION TO H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21078296&pc=HSB | 8030965-2025-00116 |  | Death | SYNTHES GMBH | 2025-01-06 05:00 | HSB | TFNA FEM NAIL Ø12 R 130° L400 TIMO15 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K160167 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, THE PATIENT UNDERWENT AN UNKNOWN SURGERY WITH THE PRODUCT IN QUESTION. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO SURGICAL DELAY. THE MORNING AFTER THE DAY OF THE SURGERY, THE PATIENT DIED. THE CAUSAL RELATIONSHIP WITH THE PRODUCT IN QUESTION IS UNKNOWN. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. H11 ADDITIONAL NARRATIVE: H6 COMPONENT CODES: MOST RELEVANT COMPONENT CODE IS G07002 (APPROPRIATE TERM/CODE NOT AVAILABLE) TO CAPTURE NO FINDINGS AVAILABLE DUE TO NO PRODUCT RETURNED. INVESTIGATION SUMMARY ==> PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21078722&pc=DSY | 21078722 | 2024-11-29 05:00 | Death | W. L. GORE ASSOCIATES, INC. | 2025-01-06 05:00 | DSY | GORE-TEX VASCULAR GRAFT |  | Exsanguination |  |  | 1 | Event Description: UNEXPECTED DEATH, RELATED TO AV (ARTERIOVENOUS) GRAFT HEMORRHAGE. PATIENT WAS FOUND UNRESPONSIVE IN ROOM, HEMORRHAGE EVIDENT. DNR (DO-NOT-RESUSCITATE ORDER).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21078939&pc=DSQ | 3007042319-2025-00023 | 2024-10-28 04:00 | Death | HEARTWARE, INC. | 2025-01-06 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Respiratory Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING PATIENTS WITH ANTI-MITOCHONDRIAL ANTIBODY (AMA)-POSITIVE MYOSITIS. THE AUTHORS DESCRIBED ONE PATIENT WITH AMA-POSITIVE MYOSITIS THAT REQUIRED THE IMPLANT OF A LEFT VENTRICULAR ASSIST DEVICE (VAD). THERE WERE NO SPECIFIC COMPLICATIONS NOTED WITH THIS PATIENT; HOWEVER, IT IS POSSIBLE THAT THEY HAD COMPLICATIONS SUCH AS CARDIAC FAILURE, ARRHYTHMIAS, OR RESPIRATORY FAILURE AND MAY HAVE DIED. THE CAUSE OF DEATH IS UNKNOWN. ADDITIONAL INFORMATION ABOUT THIS PATIENT COULD NOT BE OBTAINED. THE STATUS OF THE VAD IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. THE EVENT ONLY OCCURRED WITH ONE PATIENT BUT SPECIFIC DETAILS ON THE PATIENT WERE NOT PROVIDED. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PREVALENCE AND RISK SURVEILLANCE OF ANTI-MITOCHONDRIAL ANTIBODY-POSITIVE MYOSITIS: OUTCOMES OF A NATIONWIDE SURVEY. JOURNAL OF THE NEUROLOGICAL SCIENCES. 2024. 467; 123287. DOI: 10.1016/J.JNS.2024.123287 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21079034&pc=DQA | 2936999-2025-00001 |  | Death | MEDIANA CO. LTD. | 2025-01-06 05:00 | DQA | UNSPECIFIED PULSE OXIMETER | Incorrect Measurement; Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED THAT, THE SPO2 WAS SHOWING 100% TWO HOURS AFTER DEATH. THE SENSOR SHOWED 100% AFTER REMOVING THE SENSOR. THERE WAS A PATIENT DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: G3 CORRECTION: H6 INFORMATION RECEIVED SHOWS THAT THIS EVENT IS A DUPLICATE OF ANOTHER ISSUE REPORTED UNDER REGULATORY REPORT #2936999-2024-01860. THIS FILE WILL BE CLOSED AND ALL FURTHER UPDATES PROCESSED UNDER THE ABOVE-REFERENCED REPORT FOR THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21079100&pc=DYB | 9612164-2025-00087 | 2024-12-12 05:00 | Death | MEDTRONIC IRELAND | 2025-01-06 05:00 | DYB | SENTRANT INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K123990 |  | 1 | Event Description: A SENTRANT SHEATH (SENSH1428W) WAS INTENDED TO BE USED AS A CONDUIT DURING A TAVI PROCEDURE. IT WAS NOTED THAT THE SENTRANT SHEATH WAS FREE FROM ANY DEFECTS. THE PATIENT PRESENTED WITH HYPOTENSION AND PERICARDIAL EFFUSION IT WAS REPORTED DURING THE INDEX PROCEDURE, THAT THE MEDICAL TEAM INADVERTENTLY PERFORATED THE LEFT VENTRICLE WHILE POSITIONING THE GUIDE WITHIN THE VENTRICLE. DESPITE EFFORTS, INCLUDING CARDIORESPIRATORY RESUSCITATION AND PERICARDIAL DRAINAGE, THE INTERVENTIONS WERE UNSUCCESSFUL, AND THE PATIENT EXPIRED DURING THE INDEX PROCEDURE. PER THE PHYSICIAN THE CAUSE OF THE PERFORATION AND SUBSEQUENT DEATH WAS DUE TO USER ERROR. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED; THE SENTRANT SHEATH WAS INSERTED INTO THE PATIENT FOLLOWING THE PUNCTURE. IT WAS DETERMINED THAT THE SHEATH DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT¿S DEATH, AND THERE HAS BEEN NO COMPLAINTS AGAINST THE DEVICE. ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21079230&pc=MVK | 3008642652-2025-00111 | 2024-07-21 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-06 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 22:01:08 ON (B)(6)2024. AT 19:42:14 ON (B)(6)2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM DEGRADING TO VT @ 180 BPM. THE DEVICE PROPERLY DETECTED VT. CPR/ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 20:09:30 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21079242&pc=MVK | 3008642652-2025-00112 | 2024-09-22 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-06 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO NON-LIFEVEST DEFIBRILLATIONS. THE DEVICE WAS STARTED UP AT 04:41:11 ON (B)(6)2024. AT 11:20:34, THE PATIENT RECEIVED THE FIRST NON-LIFEVEST DEFIBRILLATION. RHYTHM AT TIME OF TREATMENT WAS SVT @ 190 BPM WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS SVT @ 190 BPM WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 11:21:41, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF PREVENTED. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 11:32:56, THE PATIENT RECEIVED THE SECOND NON-LIFEVEST DEFIBRILLATION. RHYTHM AT TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 11:37:53 ON (B)(6)2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21079692&pc=DRA | 2135147-2025-00112 | 2024-12-17 05:00 | Death | ABBOTT MEDICAL | 2025-01-06 05:00 | DRA | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | K190167 |  | 1 | Event Description: IT WAS REPORTED THAT STEERABLE GUIDE CATHETER (SGC) WAS PREPPED AND INSERTED PER INSTRUCTIONS FOR USE (IFU). UPON REMOVAL OF THE DILATOR, IT WAS OBSERVED THAT THE PATIENT'S BLOOD PRESSURE WAS LOW. PATIENT WAS NOTED TO BE BRADYCARDIC AND HYPOTENSIVE. CPR WAS PERFORMED. THE SGC WAS REMOVED FROM THE LEFT ATRIUM AND BACK INTO THE INFERIOR VENA CAVA. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED DEATH AND HYPOTENSION. DEATH AND HYPOTENSION ARE LISTED IN THE INSTRUCTIONS FOR USE IS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE SPECIFIC CIRCUMSTANCE AS CPR WAS PERFORMED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21079749&pc=NKM | 2135147-2025-00115 | 2008-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-01-06 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "VERY LONG-TERM OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) USING THE MITRACLIP DEVICE (ABBOTT) HAS BEEN A SIGNIFICANT ADVANCEMENT IN THE MANAGEMENT OF PATIENTS WITH SEVERE MITRAL REGURGITATION (MR) WHO ARE CONSIDERED HIGH RISK FOR CONVENTIONAL SURGERY. DESPITE ITS WIDESPREAD ACCEPTANCE AND GROWING APPLICATION, THE LONG-TERM BENEFITS AND PATIENT SELECTION CRITERIA CONTINUE TO BE EXPLORED AND DEBATED WITHIN THE CLINICAL COMMUNITY. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT TEER WITH MITRACLIP OFFERS A TREATMENT OPTION FOR PATIENTS WITH SEVERE MR WHO ARE NOT CANDIDATES FOR SURGERY. WHILE THE PROCEDURE HAS SHOWN PROMISING RESULTS, THE ONGOING CHALLENGE REMAINS IN REFINING PATIENT SELECTION AND IMPROVING TECHNIQUES TO ENSURE THE BEST POSSIBLE OUTCOMES. AS TECHNOLOGY ADVANCES AND MORE DATA BECOME AVAILABLE, THESE INSIGHTS WILL CONTINUE TO SHAPE THE FUTURE OF MR MANAGEMENT, ENSURING THAT PATIENTS RECEIVE THE MOST EFFECTIVE AND PERSONALISED CARE POSSIBLE. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARTA BARGAGNA AT SAN RAFFAELE UNIVERSITY HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL BARGAGNA.MARTA@HSR.IT]. THE TIME FRAME OF THE STUDY WAS (B)(6) 2008 TO (B)(6) 2013. A TOTAL OF 150 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 150 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, DIABETES, PREVIOUS CARDIAC SURGERY. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION, MITRAL VALVE REPLACEMENT, ADDITIONAL MITRACLIP PROCEDURE, UNEXPECTED MEDICAL INTERVENTION, SLDA (SINGLE LEAFLET DEVICE ATTACHMENT), RECURRENT MR. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE: ARTICLE TITLE: VERY LONG-TERM OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLES ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "VERY LONG-TERM OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR". SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF THE MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE TITLED ¿VERY LONG-TERM OUTCOMES OF MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR¿. COMORBIDITIES INCLUDED CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, DIABETES, PREVIOUS CARDIAC SURGERY. SOME OF THE COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION, MITRAL VALVE REPLACEMENT, ADDITIONAL MITRACLIP PROCEDURE, UNEXPECTED MEDICAL INTERVENTION, SLDA (SINGLE LEAFLET DEVICE ATTACHMENT), AND RECURRENT MR; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080183&pc=MCX | 3004742232-2025-00007 | 2024-12-16 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-01-06 05:00 | MCX | DIAMONDBACK CORONARY | Entrapment of Device; Material Separation | Perforation of Vessels; Device Embedded In Tissue or Plaque | P130005 |  | 1 | Event Description: A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A ECCENTRIC CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY. DURING THE PROCEDURE, IT WAS INDICATED THE OAD BECAME STUCK AND THE ENTIRE SHAFT OF THE ARTERY WAS RUPTURED DUE TO THE OAD. AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT EXPIRED.Event Description: A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT AN ECCENTRIC CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY. DURING THE PROCEDURE, IT WAS INDICATED THE OAD BECAME STUCK AND THE ENTIRE SHAFT OF THE ARTERY WAS RUPTURED DUE TO THE OAD. IT WAS INDICATED THE CROWN DETACHED. AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT EXPIRED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED MATERIAL SEPARATION, ENTRAPMENT OF DEVICE, DEVICE EMBEDDED IN TISSUE OR PLAQUE, PERFORATION OF VESSELS, SURGICAL INTERVENTION AND DEATH APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED MATERIAL SEPARATION, ENTRAPMENT OF DEVICE, DEVICE EMBEDDED IN TISSUE OR PLAQUE, PERFORATION OF VESSELS, SURGICAL INTERVENTION AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. MEDICAL REVIEW: SIX CINEFLUOROGRAPHY (CINE) IMAGES WERE PROVIDED ON BEHALF OF DR. (B)(6). ALL IMAGES ARE IN THE RIGHT ANTERIOR OBLIQUE (RAO) CRANIAL PROJECTION. THE FIRST IMAGE VISUALIZES PERFUSION OF THE LEFT CORONARY ARTERIES VIA GUIDE CATHETER AND CONTRAST. ALSO NOTED IS AN AORTIC VALVE AND NON-ABBOTT LEFT VENTRICULAR HEART PUMP. THE SECOND IMAGE SHOWS THE INTERVENTIONAL WIRE INSERTED PAST THE LAD LESION AND PROLAPSED WITHIN THE VESSEL. THE THIRD IMAGE SHOWS THE OAD OVER THE INTERVENTIONAL WIRE POSITIONED AT THE PROXIMAL EDGE OF THE LESION. THE FOURTH IMAGE SHOWS THE OAD STILL ON THE INTERVENTIONAL WIRE WITH AN EXTRAVASATION OF CONTRAST MEDIA, POSSIBLY INDICATIVE OF A PERFORATION. THE FIFTH IMAGE SHOWS THE PERFORATION MAGNIFIED. THE SIXTH IMAGE SHOWS OAD REMOVED, AND BALLOON OR STENT CATHETER BEING ADVANCED OVER THE WIRE IN AN ATTEMPT TO TAMPONADE THE LAD THE VESSEL HIGHLIGHTED BY CONTRAST MEDIA IN THE FIRST CINE COMPARED WITH THE WIRE TRACK REFLECTED IN THE SECOND CINE ARE SUBSTANTIALLY DIFFERENT LAD PATHWAYS. THE SECOND CINE IMAGE SHOWS THE WIRE TAKING AN UNKNOWN TURN WHILE CURLING, AS WELL AS DEPICTING A LENGTHENED WIRE PATH COMPARED TO THE FIRST CINE. THIS INDICATES THAT THE WIRE IS SUBINTIMAL SPACE INSTEAD OF TRUE LUMEN.1 THE THIRD CINE SHOWS THE WIRE TO BE REPOSITIONED IN TRUE LUMEN OF THE LAD WITH THE OAD ON THE WIRE. HOWEVER, THERE IS NO VISUAL CONFIRMATION THAT THE OPERATOR HAD REPOSITIONED THE WIRE IN TRUE LUMEN BY DOING A CONTRAST MEDIA CINE PRIOR TO OAD ADVANCEMENT. IN FACT, THERE¿S DELAYED CONTRAST MEDIA CLEARANCE AROUND THE OAD¿S BURR RATHER THAN IT CLEARING WITH BLOOD FLOW. THE OAD IS MEANT FOR ORBIT WITHIN RELATIVELY OPEN, CONCENTRIC AND CALCIFIED TRUE LUMEN. BY ROTATING THE BURR IN SUB-INTIMA, THE VESSEL COLLAPSED AROUND THE CROWN. ROTATING THE BURR MORE OR TRYING AND MEANS TO REMOVE THE OAD FROM THE BODY SHOULD HAVE BEEN IMMEDIATELY HALTED IN ORDER TO GET A SURGICAL CONSULT. BUT CINE IMAGES FOUR THROUGH SIX DEPICT CONTRAST EXTRAVASATION THAT GROWS THROUGH THE CINE SERIES, INDICATING THE PERFORATION IS GROWING LARGER. THIS WAS A CASE TO TREAT A PATIENT FOR LAD CALCIFICATION. BASED ON THE LIMITED INFORMATION REPORTED AND SUBMITTED CINE IMAGES, IT IS MOST LIKELY THAT THE OAD INDIRECTLY CONTRIBUTED OR CAUSED THE PERFORATION OR LAD RUPTURE THAT MOST LIKELY LED TO THE OUTCOME OF DEATH. SINCE THE OAD DEVICE WAS USED FOR THIS PROCEDURE, IT CANNOT BE CONFIRMED BY THE CINE IMAGES IF THE LAD RUPTURE POTENTIALLY CAUSED BY THE DEVICE. TO SUMMARIZE, THE IMAGING PROVIDED DID CONFIRM THAT THE DBACK GLIDE COR 1.25MM CL135CM OAD CE CAME TO A STOP IN THE CALCIFIED LAD DISTAL TO THE UNSPECIFIED DEPLOYED STENT AND COULD NOT BE RETRACTED BY THE OPERATOR. THE IMAGES DID NOT CONFIRM THE ¿DE-STRAPPING¿ OR MATERIAL SEPARATION OF THE CROWN PER THE INCIDENT REPORT NOR DID THE IMAGES CONFIRM THE LAD RUPTURE POTENTIALLY CAUSED BY THE DEVICE. THE INTENT OF THE OPERATOR WAS TO ATHERECTOMIZE DISTAL TO THE ECCENTRIC PLAQUE. THE POSSIBLE MALFUNCTION OF THE OAD AS REPORTED WAS CONFIRMED VIA THE IMAGING PROVIDED IN THAT IT WAS INDEED STUCK IN PLACE, UNABLE TO BE RELEASED AND WITHDRAWN SUCCESSFULLY; HOWEVER, A PROBABLE CAUSE COULD NOT BE DETERMINED VIA THE IMAGES PROVIDED. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080396&pc=CBK | 9617566-2025-00001 |  | Death | BREAS MEDICAL AB | 2025-01-06 05:00 | CBK | VIVO 45LS | Insufficient Information | Insufficient Information | K193586 |  | 1 | Event Description: 12/11/2024: BREAS RECEIVED FROM EASTERN MEDTECH NOTIFYING BREAS THAT A 'PT PASSED USING OUR VENTILATOR AND THE VENT MAY HAVE BEEN TAKEN BY A HAZMAT TEAM OR ANOTHER DME COMPANY. BREAS REACHED OUT FOR MORE INFORMATION AND WAS PROVIDED THE FOLLOWING ON 12/23/2024: WE HAD A PATIENT WHO PASSED AWAY. HE WAS USING THE VENT WHEN HE PASSED AWAY ON (B)(6) 2024 AND HIS BODY WAS NOT DISCOVERED UNTIL 2-3 DAYS LATER. WE WERE TOLD THAT THE HAZMAT TEAM TOOK THE VENT DUE TO IT BEING "HAZARDOUS" BECAUSE HE WAS HOOKED UP TO IT WHEN HE PASSED. WE HAVE CALLED MANY PLACES TRYING TO TRACK DOWN WHO TOOK THE VENT FROM THE HOUSE AND WHERE IT COULD POSSIBLY BE. THE PATIENT'S UNCLE IS IN CHARGE OF EVERYTHING AND THOUGHT THAT ANOTHER COMPANY TOOK IT THINKING IT WAS THEIRS. THEY SAID THEY LEFT IT BECAUSE IT HAD OUR NAME ON IT. WE WERE THEN TOLD THAT THE FUNERAL HOME TOOK IT. I CALLED THERE AND THEY SAID THAT THEY DO NOT TOUCH THAT TYPE OF THING AND IT WAS MORE LIKELY THAT THE HAZMAT TEAM TOOK IT. WE CALLED THE POLICE DEPARTMENT WHICH UNFORTUNATELY WAS ANOTHER DEAD END. AT THIS POINT, WE ARE NOT SURE WHERE IT IS OR WHAT HAPPENED TO IT. BREAS HAS NOT RECEIVED ANY ADDITIONAL INFORMATION TO DATE. THE VENTILATOR IS NO LONGER IN POSSESSION OF EASTERN MEDTECH. THE VENTILATOR HAS NOT BEEN RETURNED TO BREAS FOR INVESTIGATION. Manufacturer Narrative: DEVICE WAS TAKEN INTO POSSESSION BY A THIRD PARTY WHICH IS UNKNOWN AT THIS TIME. LOG FILES HAVE NOT BEEN MADE AVAILABLE TO THE MANUFACTURER AT THIS TIME. THE US AGENT HAS MADE MULTIPLE ATTEMPTS TO REQUEST ADDITIONAL INFORMATION AND TO OBTAIN POSSESSION OF THE DEVICE TO NO AVAIL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080698&pc=DSR | 3007972010-2025-00002 | 2024-12-10 05:00 | Death | CVRX, INC. | 2025-01-06 05:00 | DSR | BAROSTIM NEO2 | Adverse Event Without Identified Device or Use Problem | Dehydration; Renal Failure | P180050 |  | 1 | Event Description: A BAROSTIM SYSTEM WAS IMPLANTED ON (B)(6) 2024. ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED WITH DEHYDRATION, KIDNEY ISSUES, AND A POSSIBLE INFECTION. IT WAS NOTED DIALYSIS WAS PLANNED FOR (B)(6) 2024. IN THE OPINION OF THE PHYSICIAN, BAROSTIM WAS WORKING, SO THE PATIENT'S DIURETICS NEEDED TO BE STOPPED TO PREVENT FUTURE DEHYDRATION. THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO RENAL FAILURE. IN THE OPINION OF THE PHYSICIAN, THE DEATH WAS NOT CAUSED BY BAROSTIM. Manufacturer Narrative: WHILE ANALYSIS WAS UNABLE TO BE PERFORMED AS THE DEVICE WAS NOT RETURNED, PER THE OPINION OF THE PHYSICIAN, THE ROOT CAUSE OF THE EVENT WAS THE NEED TO ADJUST MEDICATIONS AND NOT CAUSED BY BAROSTIM THERAPY. THE DEVICE HISTORY AND STERILIZATION RECORD FOR THIS DEVICE SERIAL NUMBER HAS BEEN REVIEWED. NO ISSUES OR DISCREPANCIES WERE NOTED DURING THIS REVIEW THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE DEVICE MET MATERIAL, ASSEMBLY, AND QUALITY CONTROL REQUIREMENTS. CVRX ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080811&pc=OZD | 1220648-2025-25333 | 2024-12-18 05:00 | Death | ABIOMED, INC. | 2025-01-06 05:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Increase in Pressure; Infusion or Flow Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT MULTIPLE COMPLICATIONS WERE ENCOUNTERED DURING IMPELLA 5.5 SUPPORT. FOLLOWING IMPLANT, THE PUMP WAS BEING REPOSITIONED WHEN PUMP SATURATION WAS NOTED. A PURGE SYSTEM BLOCKED ALARM APPEARED AND THE WAVEFORMS ON THE AIC FLIPPED. THE WAVEFORMS EVENTUALLY CORRECTED THEMSELVES, BUT THE PURGE SYSTEM BLOCKED ALARM PERSISTED AND TISSUE PLASMINOGEN ACTIVATOR WAS ADMINISTERED. THE FOLLOWING DAY, THE PATIENT WAS HYPERCOAGULABLE AND THE IMPELLA EXPERIENCED LOW FLOWS. A TRANSESOPHAGEAL ECHOCARDIOGRAM SHOWED THERE WAS NO MOVEMENT OF THE HEART COINCIDING WITH LARGE CLOT FORMATIONS IN THE LEFT ATRIUM AND AORTA. IT WAS DECIDED THAT NO FURTHER ACTION WAS POSSIBLE AND THE PUMP WAS REMOVED. THE PATIENT PASSED AWAY SHORTLY AFTER. IT IS UNKNOWN IF COMPLICATIONS WITH THE PUMP CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, AN INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER "UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE SATURATED FLOW, LOW PUMP FLOW, HIGH PURGE PRESSURE, AND THROMBUS HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. WHEN THE PUMP WAS STARTED, FLOWS WERE VARIABLE AT P-4 AND THE MOTOR CURRENT WAS FLAT. THERE WERE "IMPELLA POSITION UNKNOWN" AND "SUCTION" ALARMS. ABOUT AN HOUR INTO THE CASE, THERE WAS A SMALL SHIFT IN THE MOTOR CURRENT AND INCREASE IN FLOWS. ABOUT TWO HOURS LATER, THERE WAS A SPIKE IN MOTOR CURRENT FOLLOWED BY A RAPID INCREASE OF PURGE PRESSURE TO 1200 MMHG OVER THE SPAN OF FIVE MINUTES. "PURGE SYSTEM BLOCKED" AND "HIGH PURGE PRESSURE" ALARMS WERE SEEN. SATURATED PUMP FLOWS THEN OCCURRED REACHING ALMOST 4.5 LITERS PER MINUTE AT P-4. THE HIGH PURGE PRESSURE LASTED FOR ABOUT TWO HOURS AND DECREASED ALONG WITH THE PUMP FLOWS. THE PUMP FLOWS LOWERED AND THEN REMAINED LOW THROUGHOUT THE CASE. THE PATIENT WAS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND ALSO RECEIVING LOW FLOWS ON ECMO. THERE WAS A TREND IN THE PLACEMENT SIGNAL. ONCE THE PUMP WAS REMOVED, THERE WAS A LARGE CLOT FORMATION IN THE LEFT ATRIUM AND AORTA. THE ROOT CAUSE OF THE LOW PUMP FLOWS IS MOST LIKELY DUE TO THE PATIENTS CONDITION BASED ON THE LOGS AND THE LACK OF VOLUME OF THE PATIENT. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE IS MOST LIKELY DUE TO INGESTED BIOMATERIAL BASED ON THE LOG SIGNATURE AND THAT TISSUE PLASMINOGEN ACTIVATOR RESOLVED THE HIGHER PRESSURES. THE ROOT CAUSE OF THE SATURATED FLOW IS MOST LIKELY DUE TO INGESTED BIOMATERIAL BASED ON THE LOG SIGNATURES. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. G.1 REVISED REPORTING CONTACT EMAIL SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25333 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES. H.6 ADDITIONAL CODES WERE ADDED TO HEALTH EFFECT - IMPACT CODE SECTION AS THEY WERE INADVERTENTLY OMITTED FROM INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25333.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080874&pc=KNT | 3010757606-2025-00003 | 2024-10-01 04:00 | Death | ABBVIE MEDICAL DEVICE CENTER | 2025-01-06 05:00 | KNT | DUODOPA_DUOPA | Adverse Event Without Identified Device or Use Problem | Peritoneal Laceration(s); Sepsis; Bowel Perforation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K142816 |  | 1 | Event Description: ON (B)(6) 2024 A PATIENT IN SPAIN UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY-JEJUNAL (PEG-J) TUBES, APPARENTLY WITHOUT COMPLICATIONS AND RETURNED HOME. AT HOME, THE PATIENT STARTED EXPERIENCING BLOCKAGES, THE INTESTINAL TUBE WAS REMOVED, PROMPTING A RETURN TO THE HOSPITAL ON (B)(6) 2024. IN THE HOSPITAL, AN ABDOMINAL CT SCAN REVEALED INTESTINAL PERFORATION, RESULTING IN PERITONITIS AND SEPSIS, BOTH WITH A POOR PROGNOSIS. BY (B)(6) 2024, THE PATIENT DEVELOPED A STOMACH INFECTION AND ABUNDANT DRAINAGE THROUGH THE STOMA. AFTER THE CONTINUOUS INFUSION PUMP WAS REMOVED, COMPLICATIONS PERSISTED, LEADING TO ALL PARKINSON'S TREATMENTS BEING WITHDRAWN AND REPLACED WITH MORPHIC CHLORIDE. THE PATIENT PASSED AWAY ON (B)(6) 2024. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. DESPITE MULTIPLE QUERIES, THERE IS NO CLEAR EVIDENCE LINKING THE DEVICE TO THE REPORTED EVENT, AND NO ADDITIONAL INFORMATION REGARDING DEVICE-RELATED ALLEGATIONS, MALFUNCTIONS, COMPLICATIONS, DIAGNOSTIC TESTING, OR TREATMENT IS AVAILABLE. THEREFORE, OUT OF ABUNDANCE OF CAUTION, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT WITHOUT ATTRIBUTING CAUSALITY TO THE DEVICE. Manufacturer Narrative: REFERENCE NUMBER (B)(4). THE DEVICE MANUFACTURER AND LOT NUMBER OF THE DEVICE INVOLVED IN THIS COMPLAINT WAS NOT PROVIDED. THEREFORE, IT IS UNKNOWN IF THE DEVICE INVOLVED WAS ABBVIE BRANDED TUBING. CONSERVATIVELY, ABBVIE HAS CHOSEN TO REPORT THIS COMPLAINT DUE TO THE POTENTIAL THAT THE DEVICE INVOLVED COULD HAVE BEEN ABBVIE BRANDED TUBING. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. POTENTIAL POST PROCEDURAL COMPLICATION IS A KNOWN COMPLICATION OF A PEG- J TUBE PLACEMENT. H6 CODE OF E2402 WAS CHOSEN TO CAPTURE THE EVENT OF POSSIBLE POST PROCEDURAL COMPLICATION. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080889&pc=DSQ | 2916596-2024-52835 | 2024-02-16 05:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. IT WAS STATED THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE ACCOUNT STATED THAT NO FURTHER INFORMATION WAS AVAILABLE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081094&pc=DSQ | 2916596-2024-52973 | 2024-01-24 05:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: SECTION A2-A4: PATIENT INFORMATION WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081375&pc=GDW | 3005075853-2025-00119 | 2024-11-01 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-01-06 05:00 | GDW | GST BLUE RELOAD 45MM 6 ROW | Patient Device Interaction Problem | Failure to Anastomose | K163454 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A BARIATRIC SURGERY, THE PATIENT HAD DEHISCENCE OF THE VERTICAL STAPLE LINE OF THE GASTRIC POUCH IN THE SECOND POST-OPERATIVE DAY. HE DIED. Manufacturer Narrative: (B)(4). DATE SENT: 1/15/2025. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHAT IS THE SCHEDULE OF THE EVENT? DATE OF SURGERY (B)(6) 2024, REOPERATION 48MIN- (B)(6) 2024 AT DAWN, DEATH (B)(6) 2024, WE WERE INFORMED ABOUT THE FACT ON 19/11 AT 18H:17, WE INFORMED JOHNSON ON 20/12/2024 AT 16H:15MIN. WHICH DEVICE WAS USED WITH THE RECHARGE? PSE45A, GST45B. ANY DIFFICULTY WITH THE DEVICE INTRAOPERATIVELY? NOT ANY PROBLEMS OBSERVED WITH THE FORMATION OF THE CLAMP DURING THE PROCEDURE? NOT WHAT WAS THE LOCATION OF THE LEAK? IN THE VERTICAL LINE OF THE GASTRIC POUCH. WAS THE DEHISCENCE SITE MUCH HIGHER IN THE CLIP LINE? IT WAS DIFFICULT TO IDENTIFY, A METHYLENE BLUE TEST SHOWED LEAKAGE IN THE STAPLE LINE. WHAT WAS THE CAUSE OF DEATH? SEPSIS. IS IT RELATED TO INITIAL SURGERY? YES. DOES THE SURGEON BELIEVE THAT THERE IS A SUPPOSED DEFICIENCY IN THE ETHICON DEVICE OR WERE THERE OTHER CONTRIBUTING FACTORS OF THE PATIENT? NO, A YOUNG PATIENT WITHOUT COMORBIDITIES HAD A BMI OF 40. IS THE SURGEON INTERESTED IN TALKING TO THE ETHICON TEAM ABOUT THE EVENT? YES. Manufacturer Narrative: (B)(4). DATE SENT 1/6/2025. D4 BATCH # UNK. B3 EXACT DATE IS UNK. ASSUMED FIRST DAY OF THE MONTH. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE LOT#722C37 HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE EQUIPMENT. THE CERTIFICATE RECORDS ARE ACCESSIBLE THROUGH EXTERNAL MANUFACTURING. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT IS THE TIMELINE OF THE EVENT? WHAT DEVICE WAS USED WITH THE RELOAD? ANY DIFFICULTY WITH THE DEVICE INTRAOPERATIVELY? ANY ISSUE NOTED WITH STAPLE FORMATION THROUGHOUT THE PROCEDURE? WHAT WAS THE SITE OF THE LEAK? WAS THE SITE OF THE DEHISCENCE FAR SUPERIOR ON THE STAPLE LINE? WHAT WAS THE CAUSE OF DEATH? IS IT RELATED TO THE INITIAL SURGERY? DOES THE SURGEON BELIEVE THAT THERE IS AN ALLEGED DEFICIENCY WITH THE ETHICON DEVICE OR WERE THERE OTHER CONTRIBUTING PATIENT FACTORS? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON PERSONNEL REGARDING THE EVENT? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081435&pc=DSQ | 2916596-2024-52858 | 2023-07-07 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE ACCOUNT DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081540&pc=DSQ | 2916596-2024-52806 | 2024-12-14 05:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Respiratory Acidosis; Ascites; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2024 WITH WORSENING RIGHT VENTRICULAR FAILURE AND ASCITES. LABS SHOWED WORSENING RENAL FUNCTION. THE PATIENT WAS TRANSFERRED TO THE CARDIOVASCULAR INTENSIVE CARE UNIT (CVICU) ON (B)(6) 2024 AND WAS STARTED ON A BUMNEX DRIP AND A DIURIL REGIMEN, WITH WORSENING ACIDOSIS. NEPHROLOGY WAS CONSULTED AND CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED. MULTIPLE PRESSORS, INCLUDING MILRINONE, EPINEPHRINE, LEVOPHED, VASOPRESSOR, AND INHALED FLOLAN. THE PATIENT CONTINUED TO HAVE WORSENING SYMPTOMS WITH A POOR PROGNOSIS. THE DECISION WAS MADE TO PROVIDE COMFORT CARE AND WITHDRAW VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE VAD WAS FUNCTIONING AS INTENDED, THERE WAS NO AUTOPSY, AND THE VAD WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE ORGAN DYSFUNCTIONS/FAILURES, INCLUDING RIGHT HEART FAILURE AND RENAL DYSFUNCTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE), STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. SECTION 6 (UNDER ¿CAUTION!¿) FURTHER EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LEFT VENTRICULAR ASSIST SYSTEM DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081553&pc=NRY | 2029214-2025-00047 | 2024-02-11 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-06 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING COMPARISON OF 4 MM-SIZED AND 3 MM-SIZED STENT RETRIEVERS IN MECHANICAL THROMBECTOMY FOR M2 OCCLUSION. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION. IT WAS NOT REPORTED WHICH DEVICES WERE ASSOCIATED WITH THE REPORTED COMPLICATIONS. THE FOLLOWING MEDTRONIC DEVICES WERE USED: SOLITAIRE STENT RETRIEVERS. DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. 6 TOTAL CASES OF MORTALITY WERE REPORTED AMONG PATIENT ADVERSE EVENTS ISSUES INCLUDED: 12 CASES OF DISTAL EMBOLI, 4 CASES OF DISSECTION/PERFORATION, 6 CASES OF SYMPTOMATIC HEMORRHAGE. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE RESEARCH PEROID WAS JANUARY 2015 - JULY 2023. CORRECTIONS WAS MADE THAT SRT WAS PERFORMED WITH A 4MM AND 3MM SIZED RETRIEVABLE STENT, SUCH AS SOLITAIRE (EV3/COVIDIEN, IRVINE, CA, USA) OR TREVO (STRYKER NEUROVASCULAR, FREMONT, CA, USA). THE PRODUCT (SOLITAIRE) ITSELF DID NOT CAUSE ANY PROBLEMS AT ALL. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: UNK-NV-SOLITAIRE (UNKNOWN); IMPLANT DATE: N/A; EXPLANT DATE: N/A. G2: CITATION: AUTHORS: LEE, H. G., YI, H. J., SHIN, D.-S., & KIM, B.-T. COMPARISON OF 4 MM-SIZED AND 3 MM-SIZED STENT RETRIEVERS IN MECHANICAL THROMBECTOMY FOR M2 OCCLUSION. CURRENT NEUROVASCULAR RESEARCH 21(2):157-165 2024. DOI:10.2174/0115672026303196240327053722. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081573&pc=MNS | 2518422-2025-100090 | 2024-12-30 05:00 | Death | RESPIRONICS, INC. | 2025-01-06 05:00 | MNS | BIPAP S/T C-SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM A FAMILY FRIEND INFORMING PHILIPS THAT THE PATIENT HAD A NEW BIPAP S/T C SERIES DEVICE SINCE THE ONE THEY HAD WAS SENT IN TO BE FIXED. THE FAMILY FRIEND WANTED TO KNOW IF THE NEW DEVICE WAS REAL OR FAKE BECAUSE THEY ALLEGED THE NEW DEVICE GIVEN TO THE PATIENT HAD CAUSED THE PATIENT TO BECOME "VERY SICK USING THE NEW DEVICE AND WAS IN AND OUT OF THE HOSPITAL UNTIL HE PASSED AWAY." THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION FROM A FAMILY FRIEND INFORMING PHILIPS THAT THE PATIENT HAD A NEW BIPAP S/T C SERIES DEVICE SINCE THE ONE THEY HAD WAS SENT IN TO BE FIXED. THE FAMILY FRIEND WANTED TO KNOW IF THE NEW DEVICE WAS REAL OR FAKE BECAUSE THEY ALLEGED THE NEW DEVICE GIVEN TO THE PATIENT HAD CAUSED THE PATIENT TO BECOME "VERY SICK USING THE NEW DEVICE AND WAS IN AND OUT OF THE HOSPITAL UNTIL HE PASSED AWAY." REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081604&pc=DSQ | 2916596-2024-53011 | 2024-10-29 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Anemia; Low Blood Pressure/ Hypotension; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024 WITH HYPOXIA THAT REQUIRED INTUBATION. CULTURES FOUND ENTEROBACTER CHOACAE AND PATIENT WAS PLACED ON CEFEPIME (MAXIPIME). DURING HOSPITALIZATION, PATIENT EXPERIENCED ISSUES WITH HYPOTENSION AND INCREASING CARBON DIOXIDE LEVELS. CHEST X-RAY SHOWED HAZY OPACITIES NOTED THROUGHOUT BOTH LUNGS AND SMALL PLEURAL EFFUSIONS. THE PATIENT WAS ANEMIC WITH POSITIVE STOOL OCCULT BLOOD. THE PATIENT WAS ON ELIQUIS (APIXABAN). GASTROENTEROLOGY WAS CONSULTED, AND ELIQUIS WAS STOPPED. NO SCOPES PERFORMED. DUE TO A POOR PROGNOSIS, THE FAMILY MADE THE DECISION TO WITHDRAW CARE AND PROVIDE COMFORT MEASURES. THE VENTRICULAR ASSIST DEVICE (VAD) WAS FUNCTIONING AS INTENDED AND NO AUTOPSY WAS PERFORMED.Event Description: OUTCOME WAS LATER REPORTED TO NOT HAVE BEEN DEVICE RELATED AND DEACTIVATED PER PATIENT/FAMILY WISHES. Manufacturer Narrative: SECTION H1 - TYPE OF REPORTABLE EVENT: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING BLEEDING. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081612&pc=NPT | 2015691-2025-00072 | 2024-12-17 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-06 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult to Advance; Material Integrity Problem | Great Vessel Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 23 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION VIA RIGHT TRANSFEMORAL ARTERY, WHILE ADVANCING THE DELIVERY SYSTEM, AFTER VALVE ALIGNMENT, WHEN JUST MEETING THE ANNULUS WITH THE NOSE CONE, THE PATIENT'S PRESSURE SUDDENLY DROPPED. THE IMAGING WAS PANNED OUT WITH THE C-ARM AND THE TEAM REALIZED THAT THE DELIVERY SYSTEM KINKED AND RUPTURED THROUGH THE GREATER CURVATURE OF THE THORACIC AORTIC WALL. CPR WAS STARTED AND VASCULAR SURGERY WAS CALLED FOR BACKUP. THEY WERE NOT ABLE TO CONTROL THE BLEEDING AND DISCONTINUED THEIR EFFORTS BEFORE BEING ABLE TO REPAIR. THE PATIENT EXPIRED. AFTER DEBRIEFING, THE TEAM THINKS THAT THIS WAS CAUSED BY THE SEVERE BEND IN THE THORACIC AORTA CAUSING THE SYSTEM TO BUCKLE AND RUPTURE THROUGH A THIN PORTION OF THE AORTIC WALL. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED FROM A PRODUCT EVALUATION. THE PRODUCT WAS RETURNED; THEREFORE, A PRODUCT EVALUATION INVESTIGATION WAS COMPLETED. AS A DEVICE WAS RETURNED, A VISUAL EVALUATION WAS COMPLETED. DUE TO THE NATURE OF THIS COMPLAINT, FUNCTIONAL AND DIMENSIONAL TESTING WAS NOT PERFORMED ON THE RETURNED DEVICE. THE RETURNED DEVICE WAS EVALUATED AND FOLLOWING WAS OBSERVED: KINK WAS OBSERVED ON PROXIMAL BALLOON SHAFT, LIKELY DUE TO PACKAGING. KINK WAS OBSERVED APPROXIMATELY 4.5' FROM FLEX TIP. FLEX TIP GOUGES WERE OBSERVED. PROVIDED 3MENSIO WAS EVALUATED AND FOLLOWING WAS OBSERVED: TORTUOUS AORTA (CIRCLE) AND CALCIFICATION (ARROWS) PRESENT IN THE AORTA BIFURCATION. TORTUOUS THORACIC AORTA (YELLOW LINE). PROVIDED FLUOROSCOPIC IMAGERY WAS EVALUATED AND FOLLOWING WAS OBSERVED: DELIVERY SYSTEM APPEARED TO BE KINKED/BENT WHILE NAVIGATING THROUGH THE THORACIC AORTA. AS THE RETURNED DEVICE MEETS SPECIFICATION WITH NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT HAS CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. AN ENGINEERING SUMMARY WRITTEN BY EDWARDS LIFESCIENCES APPLIES TO THIS COMPLAINT EVENT AND ESTABLISHES, THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, THAT DIFFICULTY TRACKING, CROSSING, OR ARTICULATING THE CATHETER THROUGH THE ANATOMY HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. IFU AND TRAINING MATERIALS PROVIDE ADEQUATE INSTRUCTIONS ON DEVICE USE, RISKS, AND PRECAUTIONS AND MANUFACTURING MITIGATIONS REMAIN IN PLACE. WITH NO UNIQUE ISSUES OR CONCERNS RAISED WITH THIS COMPLAINT, THIS EVENT WILL BE INCLUDED IN ONGOING MONITORING AND TRENDING. THERE IS NO EVIDENCE TO SUPPORT THAT A DESIGN/ MANUFACTURING DEFECT HAS POTENTIALLY CONTRIBUTED TO THE COMPLAINT; THEREFORE, A MANUFACTURING MITIGATIONS REVIEW IS NOT REQUIRED. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT FOR UNABLE TO TRACK SYSTEM THROUGH ANATOMY WAS CONFIRMED BASED ON PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (TORTUOUS THORACIC AORTA) LIKELY CONTRIBUTED TO THE EVENT AS 3MENSIO REVEALED THE PRESENCE OF TORTUOSITY IN THORACIC AORTA. IN THIS CASE, PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, SMALL VESSEL SIZE, PRE-EXISTING VASCULAR STENT) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. THE COMPLAINT FOR SYSTEM COMPONENT DAMAGED WAS CONFIRMED BASED ON PRODUCT EVALUATION. DURING PROCEDURE, IT WAS NOTED THAT THE DELIVERY SYSTEM WAS UNABLE TO TRACK THROUGH PATIENT'S ANATOMY DUE TO PATIENT CONDITIONS (SEVERE TORTUOUS THORACIC AORTA). TO OVERCOME THE RESISTANCE IN THE ANATOMY, THE DEVICE MAY HAVE BEEN MANIPULATED, POTENTIALLY CAUSING THE DELIVERY SYSTEM TO KINK. AS A RESULT, THE KINKED DEVICE MAY HAVE INTERACTED WITH THE VASCULATURE WHILE NAVIGATING THROUGH THE ANATOMY AND RESULTED INTO THE POTENTIAL INJURIES AS REPORTED. AS SUCH, AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (DEVICE MANIPULATION) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. PER THE ASSESSMENT, IT HAS BEEN DETERMINED THAT NO CAPA, SCAR, OR PRA IS REQUIRED AS THE ESTABLISHED TRIGGERS HAVE NOT BEEN MET. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081830&pc=KRD | 3008114965-2025-00013 | 2024-12-14 05:00 | Death | MEDOS INTERNATIONAL SARL | 2025-01-06 05:00 | KRD | CEREPAK FREEFORM | Positioning Failure; Separation Failure | Intracranial Hemorrhage; Ruptured Aneurysm | K220040 |  | 1 | Event Description: IT WAS REPORTED, VIA A HEALTHCARE PROFESSIONAL, THAT AN UNKNOWN CEREPAK COIL (PRODUCT CODE/LOT # UNKNOWN) WAS USED FOR AN ENDOVASCULAR EMBOLIZATION PROCEDURE. DURING THE PROCEDURE, THE USER REPORTED THE COIL FAILED TO DETACH. THE EVENT RESULTED IN THE PHYSICIAN USING A BALLOON GUIDE CATHETER (UNKNOWN BRAND) TO ARREST BLOOD FLOW AND REMOVE THE COIL. COMPETITOR COILS WERE USED TO COMPLETE THE SURGERY. THE REPORTED BELIEVED THE ¿PATIENT DID END UP RE RUPTURING DURING THE CASE.¿ THE EVENT WAS REPORTED AS SUCH, ¿SURGEON CALLED ME ON SATURDAY 12-14-24 AT 4:22PM EST TO LET ME KNOW HE WOULD BE USING A CEREPAK FREEFORM MINI COIL TO TREAT A RUPTURED ACOM ANEURYSM. HE MENTIONED OVER THE PHONE HE DID NOT HAVING ANY DETACHERS AVAILABLE AND THAT HE WOULD NEED TO USE MANUAL BREAK AS A MECHANISM OF DETACHMENT. DR. CALLED ME BACK AT 5:22PM TO LET ME KNOW HE WAS HAVING AN ISSUE, AND HE COULDN¿T GET THE COIL DETACHED. ONCE THE COIL WAS INSIDE THE DOME OF THE ANEURYSM, THE MANUAL BREAK DETACHMENT FAILED. WITH THE COIL ALREADY IN THE ANEURYSM PROTECTING THE RUPTURE SITE HE WAS TRYING TO AVOID TAKING THE COIL OUT BECAUSE THE PATIENT COULD END UP RE RUPTURING. DR. OPTED TO PUT UP A BALLOON AND TAKE THE COIL OUT AND PUT UP ANOTHER COIL FROM ANOTHER VENDOR. I BELIEVE THE PATIENT DID END UP RE RUPTURING DURING THE CASE. WE ARE UNAWARE OF THE SKU#¿S/LOTS OF COILS USED ON SATURDAY NIGHT AND WE ARE UNSURE IF THE OPERATOR/STAFF SAVED THE DEVICE IN QUESTION. I WAS NOT PRESENT AT THE CASE.¿ NO FURTHER INFORMATION WAS MADE AVAILABLE AT THE TIME OF THIS REVIEW. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). UPDATED SECTIONS ON THIS MED WATCH: B1, B2, B4, B5, D1, D2A, D4, G3, G6, H2, H4, H6, H11, CONCOMITANT PRODUCTS AND RELATED REPORT NUMBER. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE 31212049 NUMBER, AND NO NON-CONFORMANCES RELATED TO THE MALFUNCTION WERE IDENTIFIED. ADDITIONAL INFORMATION WAS RECEIVED ON 14-JAN-2025. SUMMARY: PER THE INFORMATION, THE CATHETER USED WAS A HEADWAY DUO 156CM (MICROVENTION). REGARDING WHETHER THE ANEURYSM RE-RUPTURE LED TO DETERIORATION OR AN ALTERED FISHER GRADE, IT WAS REPLIED, ¿PATIENT DID RE-RUPTURE DURING THE CASE. UNSURE OF DETERIORATION.¿ THE PROXIMAL COIL WIRE WAS STRAIGHTENED FROM THE GROIN TO THE DETACHMENT ZONE BEFORE ATTEMPTING THE DETACHMENT. THE PATIENT¿S BASELINE NEUROLOGICAL STATUS PRIOR TO THE PROCEDURE WAS UNKNOWN. REGARDING WHAT WAS BEING DONE PROCEDURALLY WHEN THE RE-RUPTURE OCCURRED, IT WAS REPORTED, ¿NON DETACHED COIL WAS BEING REMOVED.¿ THE RE-RUPTURED ANEURYSM WAS TREATED WITH ¿OTHER COILS.¿ THE PRODUCT CODE AND/OR LOT NUMBER FOR THE INVOLVED CEREPAK COIL WAS NOT AVAILABLE. THE SUBMITTER WAS UNSURE IF COILS WERE SAVED FOR RETURN. TWO COILS WERE PREVIOUSLY IMPLANTED DURING PROCEDURE. REGARDING HOW LONG THE EVENT PROLONGED THE PROCEDURE, THIS INFORMATION WAS UNOBTAINABLE. THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR FURTHER ANALYSIS. PATIENT DEMOGRAPHICS AND MEDICAL HISTORY WERE NOT MADE AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED ON 15-JAN-2025. SUMMARY: THIS COMPLAINT WAS SUBMITTED TO THE USFDA BY THE HEALTHCARE FACILITY. THE INFORMATION INCLUDED IN THE EVENT DESCRIPTION IS AS FOLLOWS: ¿OVER THE WEEKEND [PHYSICIAN¿S NAME EXTRACTED] HAD AN ISSUE WITH A CERENOVUS CEREPAK COIL. HE WAS COILING AND ATTEMPTED TO DETACH THE COIL USING THE COMPANY'S RECOMMENDED 2-HAND, MANUAL TECHNIQUE, THE COIL DID NOT DETACH. THE TETHER LENGTHENED APPROPRIATELY, BUT THE COIL DID NOT DETACH. HE TRIED SEVERAL TIMES TO DETACH THE COIL WITH NO SUCCESS. THE COIL NEVER DETACHED. HE WAS FORCED TO BRING A BALLOON ACROSS THE ACOMM ARTERY AND INFLATE IT, REMOVE THE COIL, DEPLOY A MICROVENTION HYDROSOFT COIL, AND DEFLATE THE BALLOON. THE MICROVENTION COIL WAS DEPLOYED SUCCESSFULLY. THIS WAS A VERY DIFFICULT CASE WHERE DETACHMENT WAS OF EXTREME IMPORTANCE. THE PHYSICIAN REPORTED HIS CONCERNS TO THE COMPANY AS A SAFETY ISSUE. THERE WAS NO HARM TO THE PATIENT HOWEVER THE LENGTH OF THE PROCEDURE WAS EXTENDED CONSIDERABLY.¿ PER THE ADDITIONAL INFORMATION RECEIVED ON 14-JAN-2025 AND 15-JAN-2025, IT WAS CLARIFIED THE TREATING PHYSICIAN WAS REQUIRED TO USE A BALLOON GUIDE CATHETER TO FACILITATE THE REMOVAL OF THE CEREPAK COIL AND THE ANEURYSM RUPTURE REQUIRED THE IMPLANTATION OF ADDITIONAL COILS; THEREFORE, IMDRF HEALTH IMPACT CODES: ¿MODIFIED SURGICAL PROCEDURE & SURGICAL INTERVENTION¿ WERE ADDED TO THE FILE. ADDITIONALLY, IT WAS DISCLOSED THAT THE PHYSICIAN ¿REPORTED HIS CONCERNS TO THE COMPANY AS A SAFETY ISSUE. THERE WAS NO HARM TO THE PATIENT HOWEVER THE LENGTH OF THE PROCEDURE WAS EXTENDED CONSIDERABLY.¿ BASED ON THIS INFORMATION, IMDRF HEALTH IMPACT CODE: ¿SURGERY PROLONGED¿ WAS ADDED TO THE FILE; THE SEVERITY IS UNKNOWN. NO FURTHER CHANGES TO THE FILE ARE REQUIRED. THE EVENT REMAINS REPORTABLE TO THE USFDA. THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED TO IDENTIFY STATISTICAL SIGNALS FOR CONSIDERATION OF FURTHER CORRECTION ACTION. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. Manufacturer Narrative: MANUFACTURER'S REF. NO: (B)(4). SECTION B5: PER THE ADDITIONAL INFORMATION RECEIVED ON 06-MAR-2025, IT WAS DISCLOSED THAT A LOOP OF THE COIL LANDED IN THE SUBARACHNOID SPACE, RUPTURED THE NECK OF THE ANEURYSM, AND RESULTED IN AN ACUTE LEAK AND SUBARACHNOID BLEEDING. THE PATIENT SUBSEQUENTLY DIED AFTER THE PROCEDURE DUE TO SUBARACHNOID HEMORRHAGE. ALTHOUGH THE PATIENT HAD SIGNIFICANT HEMORRHAGING PRIOR TO THE PROCEDURE WITH POOR PROGNOSIS (HUNT & HESS 4), THE CORRELATION BETWEEN THE ALLEGED DEVICE DEFICIENCY AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT. BASED ON THIS INFORMATION, THIS EVENT WILL BE CONSERVATIVELY REPORTED TO THE US FDA REPORTING UNDER CRITERIA 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH,¿ WITH AN AWARENESS DATE OF 26-FEB-2025. THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED TO IDENTIFY STATISTICAL SIGNALS FOR CONSIDERATION OF FURTHER CORRECTION ACTION. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). INFORMATION REGARDING PATIENT WEIGHT, HEIGHT, MEDICAL HISTORY, RACE, AND ETHNICITY WAS NOT REPORTED. SECTION D2B ¿ PROCODE: KRD/HCG SECTION D.4: THE PRODUCT CATALOG AND LOT NUMBER IS NOT AVAILABLE / NOT REPORTED. THE UNIQUE IDENTIFIER (UDI) AND EXPIRATION DATE OF THE DEVICE IS NOT KNOWN. FAILURE TO DETACH THE CEREPAK EMBOLIC COIL COULD CAUSE STRETCHING, SEPARATION AND/OR PREMATURE DETACHMENT WHICH COULD RESULT IN NON-TARGET SITE EMBOLIZATION, ISCHEMIA OR INFARCT. THERE ARE CLINICAL AND PROCEDURAL FACTORS, INCLUDING ANEURYSM/VESSEL CHARACTERISTICS, DEVICE SELECTION, DEVICE INTERACTION, AND OPERATOR TECHNIQUE, THAT MAY HAVE CONTRIBUTED TO THE REPORTED EVENT RATHER THAN THE DESIGN OR MANUFACTURE OF THE DEVICE. IN THIS CASE, THE EVENT REQUIRED THE REMOVAL OF THE COIL, AND IT WAS REPORTED THE ANEURYSM MAY HAVE RE-RUPTURED. THE SEVERITY OF THE EVENT IS UNKNOWN. THE CORRELATION BETWEEN THE EVENT TO THE USED DEVICE AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT ENTIRELY. THEREFORE, THIS EVENT DOES MEET US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF A ¿SERIOUS INJURY.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED TO IDENTIFY STATISTICAL SIGNALS FOR CONSIDERATION OF FURTHER CORRECTION ACTION. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS BLANK; THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION. THE COMPANY IS SEEKING THIS INFORMATION THROUGH THE EVENT INVESTIGATION. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY CERENOVUS, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, CERENOVUS, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21082515&pc=DSQ | 2916596-2024-53010 | 2019-10-04 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THE CAUSE IS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE CUSTOMER ALSO STATED THAT NO FURTHER INFORMATION WAS AVAILABLE. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21082958&pc=DSQ | 2916596-2024-52974 | 2023-08-11 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2023 AND THE CAUSE WAS NOT PROVIDED BY THE CUSTOMER. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE ACCOUNT DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084228&pc=KDI | 9616240-2024-00028 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-06 05:00 | KDI | PHOENIX/INNOVA MACHINES | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K070643 |  | 1 | Event Description: IT WAS REPORTED THAT DURING HEMODIALYSIS THERAPY WITH A PHOENIX MACHINE, "WHILE DRAWING BLOOD", THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: E1: INITIAL REPORTER PHONE NO.:(B)(6). SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: H11: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084391&pc=NKM | 2135147-2025-00145 | 2024-12-17 05:00 | Death | ABBOTT MEDICAL | 2025-01-06 05:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MIXED MITRAL REGURGITATION (MR) WITH A GRADE OF 4. A MITRACLIP XTW (40619A1007) WAS INSERTED AND DEPLOYED ON THE MITRAL VALVE. HOWEVER, POST DEPLOYMENT, IT WAS OBSERVED THAT THE ANTERIOR LEAFLET WAS PERFORATED. A SECOND MITRACLIP XTW (40724A1011) WAS THEN INSERTED AND DEPLOYED ON THE MITRAL VALVE. HOWEVER, POST DEPLOYMENT OF THE SECOND CLIP, IT WAS OBSERVED THAT THE POSTERIOR LEAFLET WAS PERFORATED. THE PROCEDURE WAS COMPLETED WITH TWO CLIPS IMPLANTED. BOTH CLIPS WERE NOTED TO BE ATTACHED TO BOTH LEAFLETS AND STABLE ON THE LEAFLETS. THE MR REMAINED AT A GRADE OF 4. POST PROCEDURE, ON AN UNKNOWN DATE, THE PATIENT DIED. IT WAS NOTED THAT POST PROCEDURE, THE PATIENT HAD BEEN PROPOSED FOR SURGICAL INTERVENTION. HOWEVER, THE PATIENT DID NOT UNDERGO SURGICAL INTERVENTION, AND THE PATIENT DIED. THE PATIENT HAD A PRE-EXISTING CONDITION OF HEART FAILURE, WHICH COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. HOWEVER, THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. BASED ON AVAILABLE INFORMATION, THE CAUSE OF THE REPORTED DEATH AND PERFORATION WERE UNABLE TO BE DETERMINED. THE REPORTED PATIENT EFFECTS OF TISSUE INJURY AND DEATH, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED SERIOUS INJURY/ ILLNESS/ IMPAIRMENT WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE CLIP REMAINS IN THE PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084426&pc=DSQ | 2916596-2024-52810 | 2015-09-24 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. IT STATED THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ACCOUNT ALSO COMMUNICATED THAT NO FURTHER PATIENT OUTCOME INFORMATION WAS AVAILABLE. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084608&pc=NKM | 2135147-2025-00147 | 2024-12-17 05:00 | Death | ABBOTT MEDICAL | 2025-01-06 05:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Mitral Valve Insufficiency/ Regurgitation; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MIXED MITRAL REGURGITATION (MR) WITH A GRADE OF 4. A MITRACLIP XTW (40619A1007) WAS INSERTED AND DEPLOYED ON THE MITRAL VALVE. HOWEVER, POST DEPLOYMENT, IT WAS OBSERVED THAT THE ANTERIOR LEAFLET WAS PERFORATED. A SECOND MITRACLIP XTW (40724A1011) WAS THEN INSERTED AND DEPLOYED ON THE MITRAL VALVE. HOWEVER, POST DEPLOYMENT OF THE SECOND CLIP, IT WAS OBSERVED THAT THE POSTERIOR LEAFLET WAS PERFORATED. THE PROCEDURE WAS COMPLETED WITH TWO CLIPS IMPLANTED. BOTH CLIPS WERE NOTED TO BE ATTACHED TO BOTH LEAFLETS AND STABLE ON THE LEAFLETS. THE MR REMAINED AT A GRADE OF 4.
POST PROCEDURE, ON AN UNKNOWN DATE, THE PATIENT DIED. IT WAS NOTED THAT POST PROCEDURE, THE PATIENT HAD BEEN PROPOSED FOR SURGICAL INTERVENTION. HOWEVER, THE PATIENT DID NOT UNDERGO SURGICAL INTERVENTION, AND THE PATIENT DIED. THE PATIENT HAD A PRE-EXISTING CONDITION OF HEART FAILURE, WHICH COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. HOWEVER, THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. BASED ON AVAILABLE INFORMATION, THE CAUSE OF THE REPORTED DEATH AND PERFORATION WERE UNABLE TO BE DETERMINED. THE REPORTED UNCHANGED MR SEEMS TO BE A CASCADING EFFECT OF THE REPORTED PERFORATION. THE REPORTED PATIENT EFFECTS OF TISSUE INJURY, MITRAL REGURGITATION, AND DEATH, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED SERIOUS INJURY/ ILLNESS/ IMPAIRMENT WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE CLIP REMAINS IN THE PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084629&pc=LWS | MW5164319 | 2024-09-04 04:00 | Death | MEDTRONIC, INC. | 2025-01-06 05:00 | LWS | ICD |  | Arrhythmia |  |  | 1 | Event Description: ON (B)(6) 2024, (B)(6) - DEVICE CLINIC MAILED A LETTER TO MY HUSBAND, (B)(6), LETTING HIM KNOW THAT HIS CARELINK CARDIAC DEVICE OR HOME MONITOR MISSED A REMOTE TRANSMISSION. THE LETTER, WHICH WAS RECEIVED IN THE MAIL ON OCTOBER 24, 2024, PROVIDED SEVERAL PHONE NUMBERS TO CALL FOR ASSISTANCE. AS (B)(6) SOLE CAREGIVER AND WIFE, ON OCTOBER 25, 2024, I CALLED THE LOCAL (B)(6) TEAM OFFICE AT (B)(6) AND INITIALLY SPOKE WITH (B)(6). I WAS TOLD THE LAST DEVICE TRANSMISSION IN THEIR RECORDS WAS FROM (B)(6) 2024. OBVIOUSLY, I WAS A BIT CONCERNED THAT ALMOST THREE MONTHS HAD PASSED SINCE (B)(6) LAST DEVICE TRANSMISSION AND THAT WE HAD ONLY RECEIVED A LETTER VS. RECEIVING A PHONE CALL. I PROCEEDED TO ASK (B)(6) IF ANY ALERTS HAD BEEN RECORDED AS (B)(6) PACEMAKER ALSO HAD A DEFIBRILLATOR. (B)(6) SAID, "I'LL HAVE TO TRANSFER YOU TO OUR PACEMAKER GROUP" AND TRANSFERRED ME TO (B)(6). AFTER BEING TRANSFERRED I ASKED (B)(6) "WERE THERE ANY ALERTS FROM MY HUSBAND'S DEVICE?". (B)(6) CHECKED AND SAID NO ALERTS EXISTED FROM MY HUSBAND'S PACEMAKER/DEFIBRILLATOR DEVICE. I RELAYED TO (B)(6) THAT MY HUSBAND HAD PASSED AWAY ON (B)(6) AND AT THAT TIME I UNPLUGGED HIS HOME MONITOR. (B)(6) TOLD ME, "EVEN IF THE DEFIBRILLATOR FAILED THERE SHOULD HAVE BEEN A DEFIBRILLATOR FAILURE ALERT", BUT THERE WERE NO ALERTS. I MUST ADMIT THAT I AM UNBELIEVABLY ANGRY, UPSET AND CONFUSED AS TO WHY THIS TWO CHAMBER PACEMAKER/DEFIBRILLATOR HAD BEEN SURGICALLY IMPLANTED IN MY HUSBAND AROUND (B)(6) 2023 TO REPLACE HIS PREVIOUS SINGLE CHAMBER PACEMAKER WHEN ALL CARELINK RECORDS INDICATE IT DID NOTHING TO PREVENT MY HUSBAND DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084722&pc=QBJ | MW5164320 | 2024-12-31 05:00 | Death | ABBOTT DIABETES CARE LIMITED | 2025-01-06 05:00 | QBJ | FREESTYLE LIBRE 3 |  | Insufficient Information |  |  | 1 | Event Description: HCPO USER CHATTED IN TO ADVISE PATIENT IS DECEASED AND REQUESTED THE KEY BE DELETED. I ADVISED OF RENEWAL AND GATHERED DETAILS FOR REPORT. PATIENT WAS ACTIVELY USING PRODUCT (A SENSOR) DURING THE TIME HE PASSED AWAY. SHE DID DECLINE FOLLOW UP AS THE PATIENT WAS SEEN AT ANOTHER FACILITY. NO FURTHER INFORMATION/CLARIFICATION AVAILABLE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21084740&pc=FJK | 8030638-2024-00010 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-06 05:00 | FJK | HD CARTRIDGE LINES | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NI |  | 1 | Event Description: IT WAS REPORTED THAT DURING HEMODIALYSIS THERAPY WITH A HD CARTRIDGE LINE, "WHILE DRAWING BLOOD", THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6, H11. H11: A BATCH REVIEW WAS CONDUCTED AND THERE WERE NO DEVIATIONS FOUND RELATED TO THIS REPORTED CONDITION DURING THE MANUFACTURE OF THIS LOT. THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: B3: THE PATIENT PASSED AWAY ON AN UNSPECIFIED DATE "ABOUT TWO MONTHS AGO". E1: INITIAL REPORTER PHONE NO.: (B)(6). SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21085102&pc=DSQ | 2916596-2024-53008 | 2011-09-17 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT, US, EP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THE CAUSE OF DEATH IS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE CUSTOMER ALSO STATED THAT NO FURTHER INFORMATION WAS AVAILABLE. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: D1: BRAND NAME CORRECTED. D4: MODEL, CATALOG, AND UDI NUMBER CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED BY THE CUSTOMER THAT NO FURTHER INFORMATION WAS AVAILABLE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21085203&pc=OZD | 1220648-2025-25342 | 2024-12-18 05:00 | Death | ABIOMED, INC. | 2025-01-06 05:00 | OZD | IMPELLA CP | Positioning Failure; Deformation Due to Compressive Stress | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE MD PLACE IMPELLA CP FOR SUPPORT AS THE PATIENT HAD A SPIRAL DISSECTION FROM THE LM-LAD. THE IMPELLA WAS INSERTED WITHOUT DIFFICULTY AND STARTED IN AUTO. MD ELECTED TO FIX THE LAD. REPO SHEATH INSERTED AND SUTURED TO THE PATIENT. REMOVING THE DRAPE, STAFF NOTED THAT THERE WAS A HEMATOMA ON THE LEFT GROIN WHERE THE IMPELLA AND 6FR CENTRAL LINE WAS. MANUAL PRESSURE WAS HELD ON THE GROIN. ECHO WAS PERFORMED BY THE BEDSIDE WHICH SHOWED THAT THE IMPELLA WAS TOO DEEP. MD ADJUSTED THE IMPELLA BUT THE STERILE SLEEVE WAS TORN IN THE PROCESS. MD ATTEMPTED TO REATTACH THE SLEEVE WITH STERILE TEGADERM, AND THE PATIENT WAS PREPPED AND TRANSFERRED BY CAREFLITE TO ANOTHER FACILITY. THE PATIENT CODED OVERNIGHT. RN STAFF STATED THAT PATIENT WAS ON RAPID INFUSIONS AND THAT PATIENT POSSIBLY HAD COAGULOPATHY DISORDER. RN REPORTED THAT THE PATIENT WAS BLEEDING FROM ALL ACCESS LINE POINTS, EYES, AND NOSE. NO FURTHER INFORMATION OR DETAILS WERE GIVEN. THE PATIENT EXPIRED AFTER 11 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, AN INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. AS PER THE CLINICAL INFORMATION PROVIDED, THE PATIENT WAS TRANSFERRED AFTER IMPELLA INSERTION TO ANOTHER HOSPITAL AND HEMATOMA WAS NOTICED BY THE MEDICAL DOCTOR WHILE UNDRAPING. ACTIVATED CLOTTING TIME WAS SAID TO BE AROUND 300, BUT THE TIMING OF THE TRANSFER IS UNCLEAR. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT INFORMATION WAS PROVIDED. CLINICAL INFORMATION MENTIONED THAT AFTER THE PATIENT WAS MOVED TO A DIFFERENT FACILITY THROUGH HELICOPTER, THE MEDICAL DOCTOR STATED THAT THE PUMP WAS SEATED DEEP WHICH WAS DISCOVERED DURING ECHOCARDIOGRAM. CLINICAL INFORMATION MENTIONED THAT WHILE REPOSITIONING THE PUMP, THE STERILE SLEEVE OF THE PUMP WAS ACCIDENTLY TORN BY THE DOCTOR WHICH WAS LATER CONTROLLED BY TEGADERM. THERE IS A LACK OF INFORMATION ON POSITIONING OF DEVICE PRIOR TO TRANSFER AND THE EXACT TIMING OF THE TRANSFER/DISCOVERY OF IMPROPER POSITIONING, AS WELL AS ANY PATIENT CONDITIONS. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED DUE TO LACK OF CLINICAL DETAILS AND NO PRODUCT WAS RETURNED. THE STERILE SLEEVE DAMAGE OCCURRED DURING THE REPOSITIONING AND THUS IS CAPTURED UNDER POSITIONING. B.2 REQUIRED INTERVENTION WAS SELECTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPOR NUMBER 1220648-2025-25342. G.1 REVISED REPORTING CONTACT EMAIL SINCE THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25342 WAS SUBMITTED IN ACCORDANCE WITH UPDATED PROCEDURES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21085373&pc=DSQ | 2916596-2024-52869 | 2024-12-16 05:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI ORGAN FAILURE. IT STATED THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PATIENT'S DEATH WAS RELATED TO HEALTH CONDITION. THE PUMP WAS WORKING AS INTENDED DURING THE TIME OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: H6-3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED MULTI ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH, AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21085666&pc=NPT | 2015691-2025-00088 | 2024-12-17 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-06 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING PROCEDURE, AN ANNULAR RUPTURE OCCURRED DURING VALVE DEPLOYMENT. PERICARDIUM DRAINAGE WAS PERFORMED HOWEVER, THE PATIENT EXPIRED DUE TO SUBSEQUENT HEMORRHAGE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THE CAUSE OF THE EVENT IS UNKNOWN DUE TO THE LIMITED INFORMATION AVAILABLE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21086751&pc=OZD | 1220648-2025-25345 | 2024-12-17 05:00 | Death | ABIOMED, INC. | 2025-01-06 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT AND PERIPHERAL VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WENT TO THE CARDIOVASCULAR OPERATING ROOM (CVOR) FOR AXILLARY IMPELLA 5.5 PLACEMENT. THE CP AND PERIPHERAL VA ECMO WAS GOING TO BE REMOVED. UPON ARRIVAL TO CVOR, A TRANSESOPHAGEAL ECHOCARDIOGRAM WAS PERFORMED, AND SIGNIFICANT THROMBUS FORMATION WAS NOTED IN THE LEFT VENTRICLE CAVITY. A DECISION WAS MADE TO ABORT 5.5 PLACEMENT AND REMAIN ON ECMO AND CP. THE PATIENT WAS NOT A CANDIDATE FOR LYTIC THERAPY AND WAS MADE COMFORT CARE. OVER NIGHT THE PATIENT HAD A GASTROINTESTINAL BLEED THEREFORE THE MEDICAL DOCTORS AND FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED THROMBUS HAS BEEN COMPLETED. NEITHER THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THE ROOT CAUSE OF THROMBUS COULD NOT BE DETERMINED AS THE DEVICE WAS NOT RETURNED NOR SUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES, ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21087031&pc=DSQ | 2916596-2024-52801 | 2024-12-13 05:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE 3¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO PALLIATIVE CARE END OF LIFE. IT STATED THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PATIENT AT THE END OF LIFE, PALLIATIVE CARE. PUMP PERFORMED AS INTENDED. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21087036&pc=OZD | 1220648-2025-25348 | 2024-11-25 05:00 | Death | ABIOMED, INC. | 2025-01-06 05:00 | OZD | IMPELLA CP | Increase in Pressure | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: EU COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WAS HIGH PURGE PRESSURE. TISSUE PLASMA ACTIVATOR WAS ADDED TO THE PURGE. THE ISSUE RESOLVED AND SUPPORT CONTINUED. THE PATIENT EXPIRED AFTER 99.75 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INVESTIGATION WITH ONLY THE SHORT-TERM LOG FOR THE FIRST 10 HOURS, WHICH REPORTED NO ABNORMALITIES. HIGH PURGE PRESSURE ISSUE REPORTED ON 4TH DAY OF PUMP USE. THE RETURNED PRODUCT SHOWED NO PHYSICAL ABNORMALITIES, NO BIOMATERIAL WAS OBSERVED ON THE IMPELLER OR PURGE GAP, AND THE PUMP WAS PRIMED SUCCESSFULLY WITH NO ISSUES REPRODUCED ON PURGE PRESSURE. AS PER THE GIVEN CLINICAL DETAIL, TPA ADMINISTRATION SUCCESSFULLY RESOLVED THE HIGH PURGE PRESSURE. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON RETURN PRODUCT INVESTIGATION AND PROVIDED CLINICAL DETAILS. D9 DATE DEVICE RETURNED TO MANUFACTURER ADDED. B1 SHOULD HAVE BEEN BOTH ADVERSE EVENT AND PRODUCT PROBLEM ON MANUFACTURER DEVICE REPORT 1220648-2025-25348. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25348. H6 CODE 4582 WAS REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT 1220648-2025-25348. NEW CODE WAS ADDED TO HEALTH EFFECT - CLINICAL CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21087226&pc=OZD | 1220648-2025-25355 | 2024-12-19 05:00 | Death | ABIOMED, INC. | 2025-01-06 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhagic Stroke; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THEY WERE ON IMPELLA 5.5 SUPPORT AND THE PATIENT HAD A RIGHT AND LEFT VENTRICLE THROMBUS. WHEN TRYING TO WEAN EXTRA CORPOREAL MEMBRANE OXYGENATION, THE PATIENT WOULD NOT WAKE UP. THE PATIENT WAS TAKEN FOR A COMPUTED TOMOGRAPHY SCAN AND THE PATIENT WAS FOUND TO HAVE A HEMORRHAGIC STROKE. THE FAMILY WAS MADE AWARE AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA AND THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25355. H6 CODES 1770 AND 4617 WERE REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT 1220648-2025-25355. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21087441&pc=MHX | 9610816-2025-000020 | 2024-12-12 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-01-06 05:00 | MHX | INTELLIVUE MX450 PATIENT MONITOR | Unable to Obtain Readings | Insufficient Information | K130849 |  | 1 | Event Description: IT WAS REPORTED THE ECG LEADS WERE OFF AND THE PATIENT PASSED AWAY. THE STAFF WANTED TO CHECK THE FUNCTIONALITY OF THE MONITOR. THE CLINICAL AUDIT TRAIL WAS REVIEWED WITH THE CUSTOMER, REVEALING THE FIRST INDICATION OF ECG LEADS OFF WAS AT 00:37. THE RED ECG LEADS OFF ENDED AND REPEATED NUMBERS TIMES UNTIL 03:27. AT THAT TIME, ECG LEADS OFF PERSISTED UNTIL 06:16 WHEN THE RED ECG LEADS OFF ALARM ENDED. OVER THE COURSE OF THAT SAME TIME PERIOD, THE RED ECG LEADS OFF REMINDERS WERE SILENCED NUMEROUS TIMES. THE PATIENT PASSED AWAY AT APPROXIMATELY 06:00. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED. THE FSE TESTED AND VERIFIED DEVICE FUNCTION. THE FSE FOUND THAT BOTH THE SURVEILLANCE AND THE OVERVIEW HAD AUDIO WHICH WAS SET AT VOLUME LEVEL OF 4. THE PATIENT PASSED AWAY AT APPROXIMATELY 06:00 ON (B)(6) 2024 IN BED M510. THE CLINICAL AUDIT TRAIL WAS REVIEWED WITH THE CUSTOMER, REVEALING THE FIRST INDICATION OF ECG LEADS OFF WAS AT 00:37. THE RED ECG LEADS OFF ENDED AND REPEATED NUMEROUS TIMES UNTIL 03:27. AT THAT TIME, ECG LEADS OFF PERSISTED UNTIL 06:16 WHEN THE RED ECG LEADS OFF ALARM ENDED. OVER THE COURSE OF THAT SAME TIME PERIOD, THE RED ECG LEADS OFF REMINDERS WERE SILENCED NUMEROUS TIMES FROM THE UNIT OVERVIEW HOST (B)(6). THE REPORTED PROBLEM WAS NOT CONFIRMED. THE INVESTIGATION VERIFIED THAT THE ECG LEADS WERE OFF PRIOR TO THE PATIENT EXPIRATION; ECG COULD NOT BE MEASURED ON THE PATIENT. THE DEVICE PROVIDED INOPS FOR THIS CONDITION. THERE WAS NO MALFUNCTION OF THE PHILIPS DEVICES. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21087994&pc=DSQ | 2916596-2024-52821 | 2013-09-11 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT, US, EP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. IT WAS STATED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE CUSTOMER STATED THAT NO FURTHER INFORMATION WOULD BE PROVIDED. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION D1: BRAND NAME CORRECTED. SECTION D4: MODEL NUMBER, CATALOG NUMBER, AND PRIMARY UDI CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE ACCOUNT COMMUNICATED THAT NO FURTHER INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21088082&pc=DSQ | 2916596-2024-52814 | 2024-05-08 04:00 | Death | THORATEC CORPORATION | 2025-01-06 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE CUSTOMER STATED THAT NO MORE INFORMATION WOULD BE PROVIDED. THE DEATH WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED AND THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT NO FURTHER INFORMATION WOULD BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A - ALL PATIENT INFORMATION THAT WAS PROVIDED WAS INCLUDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21088342&pc=MDP | 8010002-2024-00050 | 2024-12-11 05:00 | Death | ASAHI KASEI MEDICAL CO., LTD. | 2025-01-06 05:00 | MDP | PLASMAFLO OP | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P820033 |  | 1 | Event Description: THIS EVENT OCCURED DURING PLASMA ADSORPTION TREATMENT PERFORMED ON A PATIENT IN JAPAN. PLASMAFLO OP IS IDENTICAL MODEL TO PLASMAFLO OP-05W(A) MARKETED IN US. THE UNDERLYING DISEASE WAS TAFRO SYNDROME. ON (B)(6) 2024, THE PATIENT'S SYMPTOMS OF TAFRO SYNDROME WORSENED, AND PLASMA EXCHANGE TREATMENT WAS CONDUCTED USING PLASMAFLO OP. FUTHAN (NAFAMOSTAT MESILATE) WAS USED AS THE ANTICOAGULANT. NORADRENALINE WAS NOT ADMINISTERED CONTINUOUSLY, AND THERE WAS NO HYPOTENTION DURING THE PLASMA EXCHANGE TREATMENT. PLASMA EXCHANGE WAS PERFORMED A TOTAL OF SEVEN TIMES. ON (B)(6) 2024, PLASMA ADSORPTION TREATMENT WAS PERFORMED DUE TO HIGH BILIRUBIN LEVELS. CHDF (CONTINUOUS HEMODIAFILTRATION), WHICH HAD BEEN PERFORMED UNTIL STARTING PLASMA ADSORPTION, WAS STOPPED AND BLOOD WAS RETURNED. AT THAT TIME, THE BLOOD PRESSURE WAS 110-130 MMHG. THE BLOOD PUMP WAS STARTED IN ORDER TO PERFORM PLASMA ADSORPTION TREATMENT USING PLASMAFLO OP AND PLASORBA BRS-350, WHICH IS NOT MARKETED IN US. THE BLOOD PRESSURE AT THE START WAS 111 MMHG. THE BLOOD PRESSURE BEGAN TO DECREASE SLOWLY. LACTATE WAS 2 MMOL/L. AT 11:23 AM, IMMEDIATELY AFTER THE START OF PLASMA ADSORPTION TREATMENT, THE PATIENT'S BLOOD PRESSURE DROPPED TO BELOW 80 MMHG. THE CONTINUOUS ADMINISTRATION OF NORADRENALINE WAS CHANGED FROM 1.5ML/H TO 1.8ML/H. THE PATIENT'S BLOOD PRESSURE ROSE FROM THE 70 MMHG RANGE TO THE 80 MMHG RANGE, BUT THEN DROPPED BACK TO THE 70 MMHG RANGE. LACTEC WAS ADMINISTERED TO MAINTAIN THE BLOOD PRESSURE AT THE 80 MMHG RANGE. AT THE TIME, THE PATIENT WAS CONSCIOUS AND ABLE TO SPEAK. AT 12:55 PM, THE PATIENT'S CONDITION SUDDENLY WORSENED. THE BLOOD PRESSURE DROPPED TO THE 50-60 MMHG RANGE. PLASMA ADSORPTION TREATMENT WAS STOPPED AND ONLY THE BLOOD PUMP WAS OPERATED. 100 ML OF REPLACEMENT FLUID WAS ADMINISTERED. THE PATIENT WENT INTO SHOCK, AND EMERGENCY TREATMENT WAS PERFORMED. (ONE SHOT OF NORADRENALINE, NEOSYNEDINE, ETC.) PH WAS 7.1. AT 13:04 PM, BLOOD WAS RETURNED TO THE PATIENT. BLOOD PRESSURE DID NOT RETURN TO NORMAL. PH WAS 7.2, AND LACTATE WAS 10 MMOL/L. AT 17:00 PM, BLOOD PRESSURE WAS IN THE 30-40 MMHG RANGE. AT 18:30 PM, THE PATIENT WAS CONFIRMED DEAD. Manufacturer Narrative: SURVEY OF THE RETURNED PRODUCT AND MANUFACTURING RECORDS. THE COMPLAINT PRODUCT WAS DISCARDED. THE LOT NUMBER WAS UNKNOWN. CAUSE INVESTIGATION COULD NOT BE CONDUCTED. PHYSICIAN'S COMMENT: THE PHYSICIAN DETERMINED THAT THE DEATH WAS DUE TO SEPTIC SHOCK CAUSED BY PORTAL VENOUS GAS PROVOKED BY THE HYPOTENSION. THE PHYSICIAN COMMENTED CAUSAL RELATIONSHIP BETWEEN HYPOTENSION AND THE USED OF PLASMAFLO OP WAS UNCLEAR BECAUSE IT IS UNABLE TO DETERMINE WHETHER THE HYPOTENSION WAS CAUSED BY THE PRODUCT OR WAS DUE TO EXTRACORPOREAL CIRCULATION. COMPANY'S COMMENT: THE CAUSE OF THE HYPOTENSION MAY BE DUE TO THE MEMBRANE OR EXTRACORPOREAL CIRCULATION; HOWEVER, THE CAUSE OF THIS EVENT COULD NOT BE IDENTIFIED. THE CAUSE OF DEATH WAS REPORTED TO BE SEPTIC SHOCK CAUSED BY PORTAL VENOUS GAS PROVOKED BY THE HYPOTENSION; HOWEVER, THE PATIENT ORIGINALLY HAD AN INTRACTABLE DISEASE CALLED TAFRO SYNDROME, AND IT CAN BE ASSUMED THAT THE SEPTIC SHOCK WAS CAUSED BY THE WORSENING OF THE UNDERLYING DISEASE. FURTHERMORE, BECAUSE NO SYMPTOMS OF HYPOTENSION OCCURRED WHEN PLASMA EXCHANGE WAS TREATED ALONE, THE RELEVANCE OF PLASMAFLO OP TO THE CONDITION IS CONSIDERED TO BE EXTREMELY LOW. HYPOTENSION IS DESCRIBED AS ONE OF THE SIDE EFFECTS IN "F. WARNINGS" OF IFU. WE WILL CONTINUE TO MONITOR THE OCCURRENCE OF SIMILART EVENTS CAREFULLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21076239&pc=DSQ | 2916596-2024-52747 | 2024-11-26 05:00 | Death | THORATEC CORPORATION | 2025-01-05 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE PATIENT OUTCOME; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21076494&pc=LWS | 2182208-2025-00030 | 2024-09-24 04:00 | Death | MEDTRONIC, INC. | 2025-01-05 05:00 | LWS | SPRINT FIDELIS | Signal Artifact/Noise; Over-Sensing; Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s) | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NOISE AND OVERSENSING WHICH LED TO INAPPROPRIATE SHOCKS. IT WAS FURTHER REPORTED THAT THE LEAD WENT OUT OF SERVICE APPROXIMATELY EIGHT YEARS AFTER IMPLANT DUE TO THE PATIENT PASSING AWAY FOR UNRELATED REASONS. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21076496&pc=EOQ | 2955842-2024-23997 | 2024-12-13 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-05 05:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pneumothorax | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT SUBSEQUENTLY EXPIRED SEVERAL HOURS AFTER THE PROCEDURE. IN THE POST-ANESTHESIA CARE UNIT, A TENSION PNEUMOTHORAX WAS DIAGNOSED WHEN THE PATIENT BECAME UNSTABLE WITH NO BREATH SOUNDS HEARD FROM THE RIGHT LUNG. A CHEST TUBE WAS PLACED. THE PHYSICIAN STATED THAT THE TENSION PNEUMOTHORAX WAS NOT RELATED TO THE ION PROCEDURE BUT DUE TO THE LESION'S LOCATION AND THE PATIENT'S CONDITION. THE PHYSICIAN BELIEVED THAT THE OUTCOME WAS NOT DEVICE-RELATED, AND THERE WAS NO DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. Manufacturer Narrative: THE PHYSICIAN REPORTED THAT THE ADVERSE EVENT AND PATIENT OUTCOME WERE NOT DEVICE-RELATED AND WOULD HAVE LIKELY OCCURRED VIA ANOTHER MODALITY DUE TO THE LOCATION OF THE LESION AND PATIENT CONDITION. A REVIEW OF THE SITE'S SYSTEM LOGS FOUND THAT NO RELATED SYSTEM ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH COPD, PRIOR BREAST CANCER, DESCRIBED AS ¿FRAIL¿ UNDERWENT AN ION LUNG BIOPSY FOR A RIGHT UPPER LOBE CAVITARY NODULE. THE BIOPSY WAS COMPLETED WITHOUT ISSUE. POST PROCEDURE A RIGHT CHEST TUBE WAS PLACED FOR HEMODYNAMIC INSTABILITY AND NO BREATH SOUNDS ON THE RIGHT ON EXAM. CONFIRMATION OF A PNEUMOTHORAX VIA IMAGING WAS NOT REPORTED. POST CHEST TUBE PLACEMENT CXR CONFIRMED CORRECT CHEST TUBE PLACEMENT. SHORTLY AFTER, THE PATIENT SUFFERED A CARDIAC ARREST WITH RETURN OF SPONTANEOUS CIRCULATION AFTER 2 ROUNDS OF CPR. THE PATIENT WAS INTUBATED AND TRANSFERRED TO THE ICU BUT THEN EXTUBATED WITHIN AN HOUR OF TRANSFER. SEVERAL HOURS LATER THE PATIENT SUFFERED ANOTHER CARDIAC ARREST. EFFORTS AT RESUSCITATION WERE UNSUCCESSFUL AND THE PATIENT DIED. BASED ON THE AVAILABLE DATA THE CAUSE OF DEATH IS UNCLEAR AS THE SUSPECTED PNEUMOTHORAX WAS TREATED WITH A CHEST TUBE. PNEUMOTHORAX RARELY LEADS TO DEATH WHEN TREATED. HOWEVER, IT IS NOTABLE THAT A PNEUMOTHORAX WAS SUSPECTED BUT NOT CONFIRMED WITH IMAGING PRIOR TO CHEST TUBE PLACEMENT AND IT IS POSSIBLE THAT THE UNDERLYING CAUSE OF DEATH WAS SOMETHING OTHER THAN PNEUMOTHORAX. IT IS ALSO INFREQUENT FOR A PATIENT TO BE LIBERATED FROM MECHANICAL VENTILATORY SUPPORT SO QUICKLY AFTER A CARDIAC ARREST WHICH MAY HAVE CONTRIBUTED TO THE SECOND ARREST AND DEATH. REGARDLESS THE ETIOLOGY OF THE INITIAL DECOMPENSATION AND SUBSEQUENT EVENTS, BASED ON THE AVAILABLE INFORMATION THE REPORTED ADVERSE EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. PNEUMOTHORAX IS A KNOWN COMPLICATION OF BRONCHOSCOPY AND BIOPSY. THE PROSPECTIVE MULTICENTER REGISTRY STUDY OF 581 BRONCHOSCOPIC BIOPSIES OF PERIPHERAL LUNG LESIONS REPORTED 10 PNEUMOTHORACES REFLECTING A RATE OF 1.7%. THE PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF NAVIGATIONAL BRONCHOSCOPY IN 1,215 SUBJECTS REPORTED A TOTAL PNEUMOTHORAX RATE OF 4.3% AND A RATE OF 2.9% REQUIRING HOSPITALIZATION OR INTERVENTION. THE RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED A PNEUMOTHORAX RATE OF 2.5%. THE SUBSEQUENT CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES INCLUDING 415 CASES REPORTED A PNEUMOTHORAX RATE OF 3.1% WITH A TOTAL OF 7 (1.7%) PATIENTS REQUIRING A CHEST TUBE. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21064228&pc=BZD | 2518422-2025-000102 | 2024-07-31 04:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE NEEDED TO SEND BACK HER RECALLED DEVICE AND WAS NOT SURE WHICH DEVICE SHE NEEDED TO SEND BACK. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21064303&pc=GDW | 2955842-2025-00036 | 2024-12-02 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-03 05:00 | GDW | SUREFORM | Failure to Form Staple; Failure to Cut | Failure to Anastomose; Unspecified Tissue Injury; Insufficient Information | K173721 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BARIATRIC GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION. THE PATIENT EXPERIENCED AN UNSPECIFIED COMPLICATION AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY 24. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY THE SURGEON'S PHYSICIAN ASSISTANT THAT DURING THE ROBOTIC PROCEDURE, TWO SUREFORM 60 BLUE STAPLER RELOADS WERE FIRED. ON THE SECOND STAPLER FIRING, A TISSUE PUSH EVENT OCCURRED TOWARDS THE END OF THE STAPLE LINE. THE SURGEON OVERSEWED THE STAPLE LINE AT THE TIME. THERE WERE NO OTHER ISSUES REPORTED DURING THE DA VINCI PROCEDURE. ON AN UNSPECIFIED LATER DATE, A POST-OPERATIVE LEAK WAS DISCOVERED FROM THIS STAPLE LINE. THE PATIENT UNDERWENT A RE-OPERATION TO REPAIR THE LEAK. ON A LATER DATE, THE PATIENT DEVELOPED A COMPLICATION SECONDARY TO THE STAPLE LINE LEAK AND WAS REINTUBATED. THE PATIENT'S FAMILY ULTIMATELY ELECTED TO TAKE THE PATIENT OFF OF LIFE SUPPORT AND THE PATIENT PASSED. NO FURTHER SPECIFIC DETAILS WERE PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: SECTION B5 ADDITIONAL INFORMATION: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED GASTRIC REVISION AND LAP BAND REMOVAL PROCEDURE, A STAPLE LINE LEAK OCCURRED REQUIRING RE-OPERATION THE PATIENT EXPIRED. THE SURGEON STATED THAT THE PROCEDURE WAS A COMPLEX REVISION OF AN OLD VERTICAL-BANDED GASTROPLASTY COMPLICATED BY SEVERE OUTLET STENOSIS AND A CHRONICALLY-ERODED, FIXED GASTRIC BAND. AFTER REMOVING THE OLD GASTRIC BAND, THE SURGEON INTENDED TO CREATE A LONGITUDINAL GASTROGASTROSTOMY TO CONNECT THE OBSTRUCTED VERTICAL GASTRIC COMPARTMENT AND THE GASTRIC REMNANT. THE SURGEON STATED THAT THE COMPLICATION WITH THE STAPLER RELOAD OCCURRED WITH THE FIRST OF THE TWO BLUE RELOADS. THE BLUE RELOAD AND STAPLER INSTRUMENT WERE INSERTED INTO THE GASTRIC LUMEN VIA GASTROTOMY AND FIRED ON THE GASTRIC COMPARTMENTS AT THE OLD STAPLE LINE FROM THE ORIGINAL GASTRIC BAND. THIS RELOAD FIRE HAD NO PAUSES, BUT AT THE END OF THE FIRING SEQUENCE, THE SYSTEM INDICATED THAT THE FIRING SEQUENCE COULD NOT BE COMPLETED. THE MESSAGE OCCURRED WITH ABOUT 1CM LEFT UNTIL COMPLETION OF THE STAPLE LINE. THE STAPLE LINE WAS INSPECTED AND APPEARED UNREMARKABLE. WHEN POSITIONING THE SECOND BLUE RELOAD AND STAPLER A DEHISCENCE WAS NOTED AT THE END OF THE PREVIOUS STAPLE LINE; THE REST OF THE STAPLE LINE APPEARED IN GOOD CONDITION. THE SECOND BLUE RELOAD FIRED WITHOUT ISSUE TO EXTEND THE DESIRED GASTROGASTROSTOMY. BESIDES THE DEHISCENCE AT THE END OF THE FIRST BLUE RELOAD FIRED, THE TWO STAPLE LINES LOOKED IN GOOD CONDITION AND THERE WAS NO BLEEDING. THE DEHISCENCE AND THE GASTROTOMY SITE WERE SUTURED IN TWO LAYERS AND CLOSED WITH A DRAIN IN PLACE OVER THE SUTURES. AN UPPER GI CONTRAST EXAM ON POST-OPERATIVE DAY (POD) 1 FOUND NO OBSERVED LEAKS. THE PATIENT WAS DISCHARGED ON POD 4 IN GOOD CONDITION. THE PATIENT RETURNED TO THE EMERGENCY ROOM FIVE DAYS AFTER DISCHARGE (POD 9) WITH SEPSIS AND RENAL FAILURE. THE PATIENT RECEIVED FLUID RESUSCITATION, PRESSORS, ANTIBIOTICS, INTENSIVE CARE UNIT (ICU) ADMISSION, PERCUTANEOUS DRAINAGE OF INTRA-ABDOMINAL PROCESS, AND CONTINUOUS RENAL REPLACEMENT THERAPY. A GASTRIC LEAK WAS CONFIRMED AND THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM ON POD 11 WHERE A GASTRIC DEHISCENCE EXTENDING FROM THE FIRST BLUE RELOAD TO THE REPAIRED GASTROTOMIES AT EITHER END. MALFORMED STAPLES WERE OBSERVED ALONG THE DEHISCENCE. A TRANSLUMINAL INSPECTION DID NOT SHOW ANY PROBLEM WITH THE BACK WALL OF THE STAPLE LINE. THERE WAS NO APPARENT ISSUE RELATED TO THE SECOND BLUE RELOAD STAPLE LINE. THE DEHISCENCE WAS DEBRIDED AND REPAIRED AND THE ABDOMINAL CAVITY WAS IRRIGATED AND LAVAGED. AN OMENTAL PATCH WAS SECURED OVER THE GASTRIC REPAIR AND DRAINS WERE PLACED. THE REPAIR WAS SUCCESSFUL AND THE PATIENT GRADUALLY IMPROVED; HOWEVER, THE PATIENT'S AIRWAYS WERE COMPLICATED AND RE-INTUBATION WAS REQUIRED. THE FAMILY ELECTED FOR COMFORT MEASURES ONLY AND PATIENT EXPIRED ON POD 24. THE DISCHARGE SUMMARY NOTED "ASYSTOLE" AS THE CAUSE OF DEATH. THE SURGEON STATED A CONCERN THAT THE FIRST BLUE RELOAD'S DEHISCENCE COULD HAVE CONTRIBUTED TO THE PATIENT'S CLINICAL COURSE. RE-REVIEW OF THE EVENT BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED FOLLOWING A COMPLEX BARIATRIC REVISION PROCEDURE. ALTHOUGH THERE WAS A STAPLER ISSUE DURING THE PROCEDURE, THE ISSUE WAS ADDRESSED AND THE ANASTOMOSIS CREATED BY ANOTHER CARTRIDGE OF THE STAPLER APPEARED INTACT. THE ANASTOMOSIS WAS CONFIRMED TO HAVE NO LEAK ON A SUBSEQUENT POSTOPERATIVE CONTRAST STUDY ON POST OPERATIVE DAY ONE. DESPITE THESE REASSURANCES, A LEAK EVENTUALLY DEVELOPED A FEW DAYS LATER. MALFORMED STAPLES WERE NOTED ON A SUBSEQUENT PROCEDURE. OTHER CONTRIBUTING CAUSES OF DELAYED LEAKS, SUCH AS DESCRIBED HERE, CAN BE MULTIFACTORIAL AND INCLUDE TISSUE ISCHEMIA, COMPROMISED TISSUE INTEGRITY DUE TO PRIOR PROCEDURES, AND OTHER POSSIBILITIES. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: AN INVESTIGATION WAS COMPLETED TO DETERMINE THE CAUSE OF THIS REPORTED EVENT. INTUITIVE SURGICAL, INC. (ISI) DID NOT RECEIVE THE DA VINCI PRODUCTS FOR INVESTIGATION. ALTHOUGH THE COMPLAINT COULD NOT BE CONFIRMED BY FAILURE ANALYSIS, THE INFORMATION GATHERED INDICATES THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE CUSTOMER REPORTED ISSUE. ADVANCED FAILURE ANALYSIS REVIEW OF THE SUREFORM 60 STAPLER (MODEL 480460-09, LOT NUMBER K11240927-0395) USED DURING THE PROCEDURE FOUND THAT IT WAS INSTALLED ON THE SYSTEM 2 TIMES AND FIRED 2 BLUE RELOADS. ON INSTALL 1, THE FIRST CLAMP WAS SUCCESSFUL, BUT WAS FOLLOWED BY A FIRING FAILURE. THE FIRING STOPPED AT ABOUT 99% COMPLETION, AFTER A DURATION OF ABOUT 11 SECONDS. ON INSTALL 2, THE FIRST CLAMP WAS ABORTED BY THE USER. THE NEXT 3 CLAMP ATTEMPTS WERE SUCCESSFUL, AND THE FIRING WAS COMPLETED WITH 1 PAUSE FOR COMPRESSION. THE INSTRUMENT WAS THEN REMOVED AND NOT USED IN THE PROCEDURE AGAIN. NO ERRORS WERE OBSERVED IN THE LOGS. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) INVOLVED WITH THIS REPORTED EVENT AND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THIS COMPLAINT. A PROBABLE ROOT CAUSE CANNOT BE DETERMINED BASED ON THE INFORMATION PROVIDED AS THE INSTRUMENT AND RELOAD WERE NOT RETURNED FOR FURTHER EVALUATION. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS EVENT. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21064433&pc=BZD | 2518422-2024-105577 | 2024-11-21 05:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | BZD | RP PERFORMAX SE ELBOW KIT, 5 PACK | Adverse Event Without Identified Device or Use Problem |  |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE RP PERFORMAX SE ELBOW KIT, 5 PACK. AN ATTENDEE/CUSTOMER OF THE TRADE SHOW ALLEGED THAT HER HUSBAND HAD TAKEN THE MASK HOME AND MAY HAVE SUFFOCATED AND LATER PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21064764&pc=BZD | 2518422-2025-000118 | 2023-07-25 04:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | BZD | REP DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE RECEIVED A REPLACEMENT AND REQUESTED TO STOP SENDING ANOTHER REPLACEMENT TO HER ADDRESS. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21065771&pc=BZD | 2518422-2025-000144 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21066222&pc=OZD | 1220648-2025-25291 | 2024-12-12 05:00 | Death | ABIOMED, INC. | 2025-01-03 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Decrease in Pressure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THAT A PATIENT WAS TRANSFERRED FROM (B)(6) HOSPITAL. THE BEDSIDE TEAM WAS CHANGING OVER CONSOLES TO (B)(6) AIC 10000. AFTER TRANSFERRING THE ALARM SIGNALED FOR "AIR IN PURGE LINE" WHEN RE-PRIMING THE PURGE CASSETTE THROUGH THE MENU, THE SCREEN FROZE AND DID NOT ADVANCE. THE IMPELLA WAS THEN TRANSFERRED TO A NEW CONSOLE AND WAS FUNCTIONING APPROPRIATELY. NO PATIENT HARM HAS BEEN REPORTED.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. PER MEDICAL SAFETY OFFICER REVIEW THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE AUTOMATED IMPELLA CONTROLLER (AIC) AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: B1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM B2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 TO REFLECT THE PATIENT OUTCOME DEATH AND DATE OF DEATH ADDED. B5 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 TO REFLECT THE MSO STATEMENT AND PATIENT OUTCOME E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H1 TYPE OF REPORTABLE EVENT REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 TO REFLECT THE PATIENT OUTCOME DEATH H5 SELECTION WAS INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 H6 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291 TO REFLECT THE PATIENT OUTCOME DEATH H6 CODES 925 WAS REPORTED INCORRECTLY THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291. H8 USAGE OF DEVICE SELECTION WAS INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25291.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21066899&pc=OZD | 1220648-2025-25295 | 2024-12-14 05:00 | Death | ABIOMED, INC. | 2025-01-03 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPLANT, THE PATIENT BECAME AGITATED AND WAS INTUBATED. AN ECHOCARDIOGRAM WAS PERFORMED AND A LARGE MURAL CLOT WAS DISCOVERED IN THE PATIENT'S LEFT VENTRICLE. THE PATIENT WAS RETURNED TO THE CATHETER LAB AND THE IMPELLA CP WAS EXPLANTED IN RESPONSE TO THE CONDITION. IT WAS DOCUMENTED SHORTLY AFTER THAT THE PATIENT PASSED AWAY. IT IS UNKNOWN IF THE COMPLICATION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED THROMBUS HAS BEEN COMPLETED. NEITHER THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THE ROOT CAUSE OF THROMBUS COULD NOT BE DETERMINED AS THE DEVICE WAS NOT RETURNED NOR SUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21066930&pc=GDW | 1219930-2025-00041 | 2024-08-19 04:00 | Death | COVIDIEN LP LLC NORTH HAVEN | 2025-01-03 05:00 | GDW | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE SOURCE OF A PROSPECTIVE AND RETROSPECTIVE STUDY PERFORMED BETWEEN 2013 AND 2023, INVOLVING 116 PATIENTS WHO UNDERWENT MINIMALLY INVASIVE DISTAL PANCREATECTOMY (MIDP) WITH OR WITHOUT SPLENECTOMY BETWEEN 2012 AND 2023. PANCREATIC TRANSECTION WAS PERFORMED USING A 60MM STAPLER. SPLENIC VESSELS WERE SECTIONED USING EITHER AN UNKNOWN VASCULAR STAPLER OR COMPETITOR CLIPS. DEVICE RELATED COMPLICATIONS INCLUDED POSTOPERATIVE PANCREATIC FISTULA (POPF), POSTOPERATIVE BLEEDING, AND ABDOMINAL FLUID COLLECTION. ONE DEATH WAS REPORTED DUE TO LATE POSTOPERATIVE UNCONTROLLED BLEEDING. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN EGIA SU - UNKNOWN ENDO GIA SULU, LOT# UNKNOWN LITERATURE EVENTS: AUTHOR: MICHELE MAZZOLA1 · ANTONIO BENEDETTI1· ALESSANDRO GIANI1· PIETRO CALCAGNO1 · ANDREA ZIRONDA1 · MICHELE PATERNO1 · ALESSANDRO GIACOMONI1 · PAOLO DE MARTINI1· GIOVANNI FERRARI1 TITLE: ABDOMINAL DRAINAGE AFTER MINIMALLY INVASIVE DISTAL PANCREATECTOMY: OUT OF SIGHT, OUT OF MIND?: MICHELE MAZZOLA; 2024; SURGICAL ENDOSCOPY SOURCE: SURGICAL ENDOSCOPY (2024) 38:6396¿6405 / HTTPS://DOI.ORG/10.1007/S00464-024-11217-8. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067001&pc=DSQ | 2916596-2024-53019 | 2024-12-25 05:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL HEMORRHAGE (ICH) ON (B)(6)2024. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED OR RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067006&pc=DSQ | 2916596-2024-52921 | 2024-12-09 05:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AND WAS RELATED TO HEALTH CONDITION. THE PUMP WAS WORKING AS INTENDED DURING THE TIME OF SUPPORT. THE CAUSE OF DEATH WAS NOT ACCESSIBLE IN THE PATIENT'S RECORDS. THIS WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED AND THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU), REVISION B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067067&pc=OZD | 1220648-2025-25296 | 2024-12-15 05:00 | Death | ABIOMED, INC. | 2025-01-03 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD ACCESS SITE BLEEDING. ONE UNIT OF PACKED RED BLOOD CELLS WAS TRANSFUSED AND THE DOCTOR REDID THE STITCH TO THE GROIN. THE PATIENT WAS ON IMPELLA SUPPORT FOR 202.15 HOURS BEFORE THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25296. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067292&pc=DSQ | 2916596-2024-52833 | 2012-11-19 05:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: A4: WEIGHT WAS NOT PROVIDED. SECTION D4: DEVICE LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067459&pc=LWS | 2649622-2025-00170 | 2023-09-19 04:00 | Death | MPRI | 2025-01-03 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Tachycardia | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT REVIEW OF THE LAST REMOTE FOLLOW-UP SHOWED THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DELIVERED A SHOCK F OR VENTRICULAR TACHYCARDIA (VT). THE PATIENT WAS INSTRUCTED TO PRESENT TO THEIR LOCAL EMERGENCY ROOM, AND THE ICD WAS REPROGRAMMED AS DEVICE DISCRIMINATORS HAD DELAYED THERAPY FOR VT DUE TO AN INITIAL CLASSIFICATION OF THE ARRHYTHMIA AS SUPRA VENTRICULAR TACHYCARDIA (SVT). THE PATIENT'S MEDICATIONS WERE ADJUSTED AS WELL. IT WAS FURTHER REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING DURING VENTRICULAR TACHYCARDIA (VT) WHICH CONTRIBUTED TO THE DELAY IN TACHYARRHYTHMIA THERAPY. THE PATIENT SUBSEQUENTLY EXPIRED TWO WEEKS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067530&pc=DSQ | 2916596-2024-52820 | 2017-06-21 04:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM AN UNKNOWN CAUSE ON (B)(6)2017. IT WAS STATED THAT THE DEATH WAS NOT DEVICE OR THERAPY RELATED, AND THAT AN AUTOPSY WAS NOT PERFORMED. THE PUMP WAS NOT EXPLANTED AND WAS NOT RETURNED. THE CUSTOMER STATED NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: CORRECTED DATA: CORRECTED INFORMATION IN D4 (ADDITIONAL DEVICE INFORMATION): REMOVED INCORRECT INFORMATION AND UPDATED PRIMARY UDI NUMBER. CORRECTED INFORMATION IN H4 (DEVICE MANUFACTURE DATE): REMOVED INCORRECT INFORMATION. D4: THE DEVICE SERIAL AND LOT NUMBERS WERE NOT PROVIDED, AND THE MANUFACTURE (H4) AND EXPIRATION DATES COULD NOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4: PATIENT WEIGHT WAS NOT PROVIDED. D4: DEVICE MANUFACTURING AND EXPIRATION DATES (H4) COULD NOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067562&pc=DSQ | 3007042319-2025-00016 | 2024-12-18 05:00 | Death | HEARTWARE, INC. | 2025-01-03 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DIED OF RESPIRATORY FAILURE WHILE IN HOSPICE CARE. THE PATIENT HAD BEEN IMPLANTED WITH A VENTRICULAR ASSIST DEVICE. Manufacturer Narrative: ###A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RESPIRATORY FAILURE WHILE IN HOSPICE CARE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067757&pc=OZD | 1220648-2025-25301 | 2024-12-17 05:00 | Death | ABIOMED, INC. | 2025-01-03 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Bradycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. THE PATIENT WAS DEFIBRILLATED ONCE AND RETURNED TO BRADYCARDIA WITH HEART BLOCK. ARRHYTHMIA OCCURRED AN HOUR AFTER REPOSITIONING THE IMPELLA. IT WAS NOTED THAT ARRHYTHMIA WAS UNRELATED TO THE IMPELLA CP. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ATRIAL ARRHYTHMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ATRIAL ARRHYTHMIA COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067914&pc=DSQ | 3007042319-2025-00018 | 2024-11-01 04:00 | Death | HEARTWARE, INC. | 2025-01-03 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Renal Failure; Respiratory Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE IMPACT OF HEART TRANSPLANT ALLOCATION CHANGE ON OUTCOMES AMONG PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES (VADS). THE AUTHORS DESCRIBED PATIENT DEATHS WHILE ON THE WAITLIST FOR HEART TRANSPLANTATION; HOWEVER, THE CAUSES OFDEATH WERE UNKNOWN. THERE WERE PATIENTS WHO REQUIRED DIALYSIS OR MECHANICAL VENTILATION AT THE TIME OF LISTING OR AT THE TIME OF TRANSPLANT. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/55 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPACT OF HEART TRANSPLANT ALLOCATION CHANGE ON WAITLIST MORTALITY AND POSTTRANSPLANT MORTALITY IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. CIRCULATION: HEART FAILURE. 2024. 17:E011621. DOI: 10.1161/CIRCHEARTFAILURE.124.011621 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067938&pc=NPT | 2025587-2025-00046 | 2023-07-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-03 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Patient Device Interaction Problem | Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION APPROXIMATELY 6 YEARS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF THE CARDIAC DEATH WAS REPORTED AS ACUTE ON CHRONIC DIASTOLIC HEART FAILURE. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067998&pc=KDI | 0008030665-2025-00017 | 2024-12-23 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-03 05:00 | KDI | 2008T HD SYS. CDX W/BIBAG BLUE STAR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K173972 |  | 1 | Event Description: A BIOMEDICAL TECHNICIAN (BIOMED) REPORTED TO FRESENIUS THAT AN UNKNOWN PATIENT EXPIRED WHILE CONNECTED TO THE FRESENIUS 2008T HEMODIALYSIS MACHINE. IT WAS REPORTED THAT THERE WERE NO DIAGNOSTIC MESSAGES OR AUDIBLE ALARMS, AND THE DEATH WAS NOT SUSPECTED TO HAVE BEEN CAUSED BY THE HEMODIALYSIS MACHINE. THE BIOMED CONTACTED FRESENIUS TECHNICAL SERVICES TO INQUIRE IF THERE WERE ANY INTERNAL MACHINE LOGS THAT CAN BE DOWNLOADED FROM THE MACHINE. LIMITED INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE BIOMED. IT WAS STATED THAT THE PATIENT WAS RECEIVING HEMODIALYSIS WHILE HOSPITALIZED. THE BIOMED DID NOT HAVE ANY PATIENT INFORMATION OR TREATMENT DETAILS AVAILABLE AND ATTEMPTS TO FOLLOW-UP WITH A MEDICAL PROFESSIONAL WERE UNSUCCESSFUL. PER THE BIOMED, POST EVENT THE FRESENIUS HEMODIALYSIS MACHINE WAS PUT THROUGH FUNCTIONAL TESTING POST, PER HOSPITAL PROTOCOL. THE BIOMED WAS COLLECTING DATA FOR THE MACHINE FOR POST EVENT TESTING, PER PROTOCOL ONLY. IT WAS REPORTED THAT THERE WERE NO MALFUNCTIONS WITH THE 2008T HEMODIALYSIS MACHINE AND THAT ALL FUNCTIONAL TESTS PASSED. THE HEMODIALYSIS MACHINE HAS BEEN RETURNED TO SERVICE WITHOUT ANY ISSUES, ACCORDING TO THE BIOMEDICAL TECHNICIAN. IT WAS STATED THAT THE REPORTED DEATH IS NOT SUSPECTED TO BE RELATED IN ANY WAY TO THE FRESENIUS HEMODIALYSIS MACHINE. THE BIOMED CONFIRMED THEY WERE COLLECTING DATA FOR THE MACHINE FOR POST EVENT TESTING PER PROTOCOL ONLY.Event Description: A BIOMEDICAL TECHNICIAN (BIOMED) REPORTED TO FRESENIUS THAT AN UNKNOWN PATIENT EXPIRED WHILE CONNECTED TO THE FRESENIUS 2008T HEMODIALYSIS MACHINE. IT WAS REPORTED THAT THERE WERE NO DIAGNOSTIC MESSAGES OR AUDIBLE ALARMS, AND THE DEATH WAS NOT SUSPECTED TO HAVE BEEN CAUSED BY THE HEMODIALYSIS MACHINE. THE BIOMED CONTACTED FRESENIUS TECHNICAL SERVICES TO INQUIRE IF THERE WERE ANY INTERNAL MACHINE LOGS THAT CAN BE DOWNLOADED FROM THE MACHINE. LIMITED INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE BIOMED. IT WAS STATED THAT THE PATIENT WAS RECEIVING HEMODIALYSIS WHILE HOSPITALIZED. THE BIOMED DID NOT HAVE ANY PATIENT INFORMATION OR TREATMENT DETAILS AVAILABLE AND ATTEMPTS TO FOLLOW-UP WITH A MEDICAL PROFESSIONAL WERE UNSUCCESSFUL. PER THE BIOMED, POST EVENT THE FRESENIUS HEMODIALYSIS MACHINE WAS PUT THROUGH FUNCTIONAL TESTING POST, PER HOSPITAL PROTOCOL. THE BIOMED WAS COLLECTING DATA FOR THE MACHINE FOR POST EVENT TESTING, PER PROTOCOL ONLY. IT WAS REPORTED THAT THERE WERE NO MALFUNCTIONS WITH THE 2008T HEMODIALYSIS MACHINE AND THAT ALL FUNCTIONAL TESTS PASSED. THE HEMODIALYSIS MACHINE HAS BEEN RETURNED TO SERVICE WITHOUT ANY ISSUES, ACCORDING TO THE BIOMEDICAL TECHNICIAN. IT WAS STATED THAT THE REPORTED DEATH IS NOT SUSPECTED TO BE RELATED IN ANY WAY TO THE FRESENIUS HEMODIALYSIS MACHINE. THE BIOMED CONFIRMED THEY WERE COLLECTING DATA FOR THE MACHINE FOR POST EVENT TESTING PER PROTOCOL ONLY. Manufacturer Narrative: PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. ADDITIONALLY, NO ON-SITE EVALUATION WAS PERFORMED BY A FRESENIUS FIELD SERVICE TECHNICIAN (FST). A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HEMODIALYSIS TREATMENT IN THE HOSPITAL WITH THE FRESENIUS 2008T HEMODIALYSIS MACHINE AND THE PATIENT¿S DEATH. BASED ON THE LIMITED INFORMATION AVAILABLE, THE PATIENT¿S CAUSE OF DEATH CANNOT BE DETERMINED. HOWEVER, THERE IS NO ALLEGATION THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS 2008T MACHINE. THE BIOMEDICAL TECHNICIAN REPORTED THE FRESENIUS 2008 T MACHINE PASSED ALL FUNCTIONAL TESTS POST EVENT AND THE MACHINE HAS BEEN RETURNED TO SERVICE. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068010&pc=LYJ | 1644487-2025-00013 | 2024-09-30 04:00 | Death | LIVANOVA USA, INC. | 2025-01-03 05:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY BUT THE CAUSE OF DEATH IS UNKNOWN. PER THE PHYSICIAN IT MAY HAVE BEEN RELATED TO A SEIZURE AND THE PATIENT'S MOM BELIEVED THE BATTERY MAY HAVE BEEN DEPLETED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTING THAT THE DEATH IS NOT BELIEVED TO BE RELATED TO THE VNS DEVICE. THE DEATH FORM NOTED THAT THE VNS PROVIDED SEIZURE REDUCTION FOR THE PATIENT AND THE DEVICE WAS NOT ON AT THE TIME OF DEATH. THE PATIENT PASSED AWAY ON (B)(6) 2024 AND THE VNS WAS NOT EXPLANTED. THE CAUSE OF DEATH WAS NOTED TO BE "SEVERE HIE FOLLOWING CARDIAC ARREST" AND IT IS NOT RELATED. THE DEATH WAS NOT WITNESSED AND NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068222&pc=NPT | 2025587-2025-00049 | 2024-12-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-03 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Stroke/CVA; Endocarditis; Unspecified Infection; Asystole; Renal Impairment | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 29 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, DIAGNOSTIC TESTS WERE PERFORMED AND STAPHYLOCOCCUS LUGDUNENSIS WAS IDENTIFIED. PER THE PHYSICIAN, THE CAUSE OF THE INFECTION WAS UNKNOWN. INTRAVENOUS (IV) MEDICATION WAS ADMINISTERED. PER THE PHYSICIAN, THE EVENT WAS POSSIBLY RELATED TO THE VALVE, DELIVERY CATHETER SYSTEM (DCS), COMPRESSION LOADING SYSTEM (CLS) AND VALVE IMPLANT PROCEDURE. THE EVENT WAS NOT RESOLVED.Event Description: IT WAS REPORTED THAT 29 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, DIAGNOSTIC TESTS WERE PERFORMED AND STAPHYLOCOCCUS LUGDUNENSIS WAS IDENTIFIED. PER THE PHYSICIAN, THE CAUSE OF THE INFECTION WAS UNKNOWN. INTRAVENOUS MEDICATION WAS ADMINISTERED. PER THE PHYSICIAN, THE EVENT WAS POSSIBLY RELATED TO THE VALVE, DELIVERY CATHETER SYSTEM, COMPRESSION LOADING SYSTEM, AND VALVE IMPLANT PROCEDURE. THE EVENT WAS NOT RESOLVED. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY ENDOCARDITIS WAS DIAGNOSED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT A RAPID ONSET OF FOCAL OR GLOBAL NEUROLOGICAL DEFICIT OCCURRED. A STROKE WAS CONFIRMED AND ALTHOUGH INITIALLY REPORTED UNRELATED TO THE DEVICE OR THERAPY, THE STROKE WAS NOW REPORTED AS PROBABLY RELATED TO THE INFECTIVE ENDOCARDITIS. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED LARGE IRREGULAR VEGETATIONS AND AN ABSCESS OF THE MITRAL VALVE. HOSPITALIZATION WAS PROLONGED AS RESULT OF THIS EVENT. THE PHYSICIAN NOW INDICATED A CAUSAL RELATIONSHIP OF THE EVENT TO THE VALVE, DELIVERY CATHETER SYSTEM, COMPRESSION LOADING SYSTEM, AND VALVE IMPLANT PROCEDURE.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED THAT APPROXIMATELY 1 MONTHS AND 15 DAYS FOLLOWING THE ONSET OF ENDOCARDITIS THE PATIENT DIED. SUDDEN CARDIAC DEATH WAS REPORTED.Event Description: ADDITIONAL INFORMATION INDICATED THAT THE OFFICIAL CAUSE OF DEATH WAS ASYSTOLE CAUSED BY ACUTE RENAL FAILURE AND ADDITIONAL NEUROLOGIC DETERIORATION AFTER THE STROKE. Manufacturer Narrative: DEVICE EVALUATED VIA IMAGE REVIEW: TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGES FROM 2024-DEC-19 WERE PROVIDED. SUBOPTIMAL IMAGE QUALITY IDENTIFIED. THE TTE SHOWED A LOW EJECTION FRACTION (EF) OF 30-35%. THE EVOLUT IMPLANT APPEARED TO BE AT A GOOD DEPTH. A MOBILE ECHOGENIC STRUCTURE WAS OBSERVED ON THE ATRIAL SIDE OF THE ANTERIOR MITRAL VALVE LEAFLET. MILD TO MODERATE MITRAL REGURGITATION AND MILD TO MODERATE TRICUSPID REGURGITATION WERE ALSO IDENTIFIED. UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX LS; PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012272646); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012193975); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5, H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2, B5, H1, H6. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. THE PATIENT CODE WAS UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2, B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068351&pc=LWS | 2649622-2025-00180 | 2024-12-27 05:00 | Death | MPRI | 2025-01-03 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD HAD A FEW UNDERSENSING BEATS WITH CONSECUTIVE COUNTER RESETTING AND BEING UNDER THE TREATMENT ZONE. IT WAS NOTED THAT STORED EPISODES SUGGEST ONE ONGOING VENTRICULAR FIBRILLATION (VF) EPISODE. THE PATIENT ULTIMATELY DIED AFTER THE ARRHYTHMIA. Manufacturer Narrative: CONTINUATION OF D10: DTPB2QQ CRTD IMPLANTED: (B)(6)2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068385&pc=NKM | 2135147-2025-00052 | 2024-12-09 05:00 | Death | ABBOTT MEDICAL | 2025-01-03 05:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Failure/Congestive Heart Failure | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MIXED MITRAL REGURGITATION (MR) GRADE 4. XTW (40916R1023) AND XTW (40916R1105) WERE IMPLANTED SUCCESSFULLY, REDUCING MR TO GRADE 1. THE PROCEDURE ENDED WITH THE PATIENT STABLE. FIVE HOURS POST PROCEDURE, IT WAS REPORTED THAT THE PATIENT DIED. THE CAUSE OF DEATH WAS SUDDEN HEART FAILURE. RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL. THE MITRACLIPS WERE CONFIRMED TO BE STABLE ON BOTH LEAFLETS. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED RECURRENT HEART FAILURE. THE REPORTED CARDIAC ARREST AND SUBSEQUENT APPEAR TO BE RELATED TO THE HEART FAILURE. HEART FAILURE, DEATH, AND CARDIAC ARREST ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068421&pc=BZD | 2518422-2025-000178 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | BZD | REMSTAR AUTOA-FLEX | Degraded | Septic Shock; Respiratory Failure; Cancer; Unspecified Mental, Emotional or Behavioural Problem; Unspecified Heart Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF CANCER PANCREATIC AND PROSTATE, CORONARY ARTERY DISEASE, ACUTE HYPOXIC RESPIRATORY FAILURE, ALTERED MENTAL STATUS, AND SEPTIC SHOCK, DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON 11/08/2024 . H11 UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF CANCER PANCREATIC AND PROSTATE, CORONARY ARTERY DISEASE, ACUTE HYPOXIC RESPIRATORY FAILURE, ALTERED MENTAL STATUS, AND SEPTIC SHOCK, DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H6 WAS UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068574&pc=DSQ | 2916596-2024-52827 | 2023-05-02 04:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. NO FURTHER INFORMATION WOULD BE PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. THE PATIENT OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE ACCOUNT COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068664&pc=OZD | 1220648-2025-25305 | 2024-12-16 05:00 | Death | ABIOMED, INC. | 2025-01-03 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT ON DAY TWO OF IMPELLA CP INSERTION THE PATIENT¿S LABS WERE COMING BACK HEMOLYZED. THE PLAN WAS TO DECREASE THE P-LEVEL TO SEE IF THE HEMOLYSIS CLEARS BUT THE NEXT DAY THE PATIENT DECOMPENSATED AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: COMPLAINT DEVICE WAS NOT RECEIVED FOR ANALYSIS. DEVICE HISTORY RECORD REVIEW CONFIRMED THE PUMP PASSED ALL POST STERILE INSPECTION CHECKS. DATA LOGS WERE REVIEWED: NO MOTOR CURRENT SPIKE OUTSIDE OF P-LEVEL CHANGES OBSERVED. SUCTIONS OCCURRED. MANY IMPELLA UNKNOWN POSITION ALARMS OBSERVED. NO CLEAR INDICATIONS OBSERVED FROM REVIEWING THE LOGS TO CONFIRM THE ROOT CAUSE OF THE REPORTED ISSUE. THE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED SINCE THE COMPLAINT DEVICE WAS NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21068730&pc=DSQ | 2916596-2024-52779 | 2024-12-06 05:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO BEING ON HOSPICE AND WITHDREW SUPPORT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. WHETHER THE PATIENT¿S OUTCOME WAS DEVICE OR THERAPY-RELATED WAS NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21069291&pc=BZD | 2518422-2025-000166 | 2024-04-26 04:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | BZD | DREAMSTATION AUTO CPAP | Insufficient Information | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. ADDITIONALLY, THE DEVICE HAS BEEN RETURNED TO THE THIRD-PARTY SERVICE CENTER. THERE WAS NO FOAM PARTICLES WERE OBSERVED DURING THE EVALUATION. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21069709&pc=OZD | 1220648-2025-25311 | 2022-02-10 05:00 | Death | ABIOMED, INC. | 2025-01-03 05:00 | OZD | IMPELLA 5.0 | Difficult to Insert | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ADMITTED TO THE HOSPITAL WITH COMPLAINTS OF SHORTNESS OF BREATH, ELEVATED TROPONIN AND IN ST ELEVATION MYOCARDIAL INFARCTION. IN THE CARDIAC CATHETERIZATION LAB, THE PATIENT WAS FOUND TO HAVE OCCLUSIONS IN BOTH LEFT ANTERIOR DESCENDING (LAD) AND RIGHT CORONARY (RCA) ARTERIES. THE PATIENT WAS URGENTLY INTUBATED DUE TO SHORTNESS OF BREATH AND DIFFICULTY LAYING FLAT. AN IMPELLA CP WAS IMPLANTED FOR MECHANICAL CIRCULATORY SUPPORT, AND THE PATIENT HAD STENTS PLACED IN THEIR LAD AND RCA. OVERNIGHT, THE PATIENT DECOMPENSATED AND REQUIRED AN INCREASE OF PRESSOR SUPPORT. THE MEDICAL TEAM DECIDED TO ESCALATE THE PATIENT SUPPORT TO AN IMPELLA 5.0. DURING IMPELLA 5.0 IMPLANT, THE MEDICAL TEAM EXPERIENCED DIFFICULTY ADVANCING THE IMPELLA 5.0 PAST THE ILIAC. A PERIPHERAL BALLOON WAS USED, ALONG WITH VIPERGLIDE TO LUBRICATE THE SHEATH AND GRAFT. THE IMPELLA 5.0 WAS THEN SUCCESSFULLY DELIVERED INTO THE LEFT VENTRICLE. ON (B)(6) 2022, IT WAS REPORTED A LARGE AMOUNT OF BLEEDING WAS FOUND COMING FROM THE JACKSON-PRATT (JP) DRAIN AT THE IMPELLA INSERTION SITE. HEPARIN WAS DECREASED, AND SIX UNITS OF BLOOD WERE TRANSFUSED TO THE PATIENT. HEMOSTASIS WAS UNABLE TO BE ACHIEVED, AND THE PATIENT WAS TAKEN FOR A COMPUTED TOMOGRAPHY SCAN TO ASSESS THE SOURCE AND THE EXTENT OF THE BLEEDING. A FEM STOP WAS PLACED AT THE ACCESS SITE, HOWEVER OOZING PERSISTED AND THE PATIENT REQUIRED AN ADDITIONAL 4 UNITS OF BLOOD TO BE TRANSFUSED. THERE WAS NO REPORT THAT HEMOSTASIS WAS EVER ACHIEVED. IT WAS ADDITIONALLY NOTED THAT THE PATIENT EXPIRED ON SUPPORT THE FOLLOWING DAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21069761&pc=DSQ | 2916596-2024-52816 | 2024-12-12 05:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Fungal Infection; Hemorrhagic Stroke; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A CEREBRAL VASCULAR ACCIDENT (CVA) STATED TO BE A HEMORRHAGIC STROKE. THE PATIENT WAS MOVED TO THE INTENSIVE CARE UNIT (ICU). THERE WERE NO REPORTED PROBLEMS WITH ANTICOAGULATION. THE PATIENT WAS NOTED TO HAVE SUSPECTED THROMBUS WITHIN THE OUTFLOW GRAFT. THEY HAD A LOT OF LOW FLOW ALARMS. THE CUSTOMER REQUESTED LOW FLOW SOFTWARE.Event Description: NO SYMPTOMS OF CVA WERE NOTED. IT WAS NOTED THAT THE PATIENT HAD AN ASPERGILLUS INFECTION WHICH COULD CONTRIBUTE TO BOTH BLEEDING AND CLOTTING. ANTICOAGULATION WAS STABLE UNTIL THE BRAIN HEMORRHAGE AND THEN WAS STOPPED. THE PATIENT WAS NOTED TO HAVE SUSPECTED THROMBUS WITHIN THE OUTFLOW GRAFT (OFG). LOW FLOW 2.0 SOFTWARE WAS INSTALLED. SYMPTOMS OF THROMBUS INCLUDED THE PATIENT BEING SEDATED AND LOW FLOWS ON THE PUMP. THE PUMP SPEED HAD BEEN GRADUALLY INCREASED DUE TO INCREASING LOW FLOW ALARMS AND DECREASING FLOW. AN ECHOCARDIOGRAM AND COMPUTED TOMOGRAPHY (CT) SCAN WERE CONDUCTED; HOWEVER, CT IMAGES WERE UNAVAILABLE. THE PATIENT WAS GIVEN A HEPARIN INFUSION DUE TO THROMBUS IN THE OFG. THE THROMBUS WAS REPORTEDLY UNRESOLVED. IT WAS NOTED THAT THE CVA RESOLVED FROM A BLEEDING STANDPOINT AFTER AN INITIAL DRAINAGE OF BLOOD FROM THE BRAIN, AND A CRANIECTOMY TO DECREASE PRESSURE. THE PATIENT PASSED AWAY ON (B)(6)2024 WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED AND WOULD NOT BE PROVIDED. IT WAS UNKNOWN IF THE PUMP WOULD BE EXPLANTED AND RETURNED. Manufacturer Narrative: A4 PATIENT WEIGHT NOT PROVIDED. B2: DEATH DATE PROVIDED. B5: ADDITIONAL INFORMATION. H1 REPORT TYPE UPDATED. H6: HEALTH EFFECT-CLINICAL, HEALTH EFFECT- IMPACT AND MEDICAL DEVICE PROBLEM CODES, UPDATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED SUSPECTED THROMBUS COULD NOT BE CONFIRMED AS NO IMAGES WERE PROVIDED AND NO PRODUCT WAS RETURNED. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED DATA FROM 16:27:31 THROUGH 20:14:09 ON (B)(6) 2024. THE FILE CAPTURED THE PUMP OPERATING AT A FIXED SPEED OF 5400 REVOLUTIONS PER MINUTE (RPM) WITH A LOW SPEED LIMIT OF 5000 RPM. THE FILE CONSISTED OF INTERMITTENT LOW FLOW EVENTS, MANY OF WHICH RESULTED IN LOW FLOW ALARMS. THESE LOW FLOW ALARMS LASTED UP TO TWO HOURS, WITH FLOW CAPTURED AT 2.3-2.4 LITERS PER MINUTE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED AND WOULD NOT BE PROVIDED. IT WAS UNKNOWN IF THE PUMP WOULD BE EXPLANTED AND RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, STROKE, BLEEDING, INFECTION, AND DEVICE THROMBOSIS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AND THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. UNDER ¿ANTICOAGULATION," THIS SECTION ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21070040&pc=NKM | 2135147-2025-00071 | 2024-12-09 05:00 | Death | ABBOTT MEDICAL | 2025-01-03 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "INTRAPROCEDURAL 3D-VENA CONTRACTA AREA PREDICTS SURVIVAL AFTER TRANSCATHETER EDGE-TO-EDGE REPAIR: RESULTS FROM MITRA-PRO REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTICENTER STUDY, TO EVALUATE THE MITRA-PRO REGISTRY REVEALED RESIDUAL MITRAL REGURGITATION (MR) TO BE AN IMPORTANT PREDICTOR OF SURVIVAL FOLLOWING TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER). INTRAPROCEDURAL MR ASSESSMENT USING 3D-VENA CONTRACTA AREA (VCA) MIGHT BE A FEASIBLE TOOL TO GUIDE MITRAL TEER PROCEDURES. THE STUDY AIMED TO ASSESS THE IMPACT OF RESIDUAL MR ASSESSED BY 3D-VCA ON 1-YEAR MORTALITY. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT RESIDUAL MR ASSESSED BY 3D-VCA AFTER TEER IS ASSOCIATED WITH 1-YEAR MORTALITY. THEREFORE, 3D-VCA IS A VALUABLE ECHOCARDIOGRAPHIC TOOL FOR INTRAPROCEDURAL MR ASSESSMENT DURING MITRAL TEER AND ACHIEVING A LOWER 3D-VCA IMPROVE PATIENT SURVIVAL. (GERMAN CLINICAL TRIALS REGISTER: DRKS00012288). [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS PETER BOEKSTEGERS AT UNIVERSITY OF WITTEN-HERDECKE INSTITUTION WITH CORRESPONDING EMAIL PETERBOEKSTEGERS@GMX.DE]. THE TIME FRAME OF THE STUDY WAS 2016 TO 2024 NOT CONFIRMED. A TOTAL OF 823 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 823 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, CORONARY ARTERY DISEASE, PREVIOUS TAVR, PREVIOUS CAGB, ISCHEMIC CARDIOMYOPATHY, ATRIAL FIBRILLATION. (B)(4), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED PROLONGED HOSPITALIZATION, HEART FAILURE, DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "INTRAPROCEDURAL 3D-VENA CONTRACTA AREA PREDICTS SURVIVAL AFTER TRANSCATHETER EDGE-TO-EDGE REPAIR: RESULTS FROM MITRA-PRO REGISTRY." SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP IMPLANTATION WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, CORONARY ARTERY DISEASE, PREVIOUS TAVR, PREVIOUS CAGB, ISCHEMIC CARDIOMYOPATHY, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED PROLONGED HOSPITALIZATION, HEART FAILURE, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21070720&pc=NIK | 2017865-2025-00163 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-00167, 2017865-2025-00169, 2017865-2025-00171. IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND PASSED AWAY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, LEFT VENTRICULAR (LV) LEAD AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR PATIENT DECEASED. INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. THE DEVICE¿S SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE DEVICE¿S FUNCTION WAS NORMAL. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21070789&pc=NVY | 2017865-2025-00167 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-00163, 2017865-2025-00169, 2017865-2025-00171. IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND PASSED AWAY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, LEFT VENTRICULAR (LV) LEAD AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DECEASED. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. DURING ANALYSIS, A FAILURE EVENT WAS OBSERVED WHICH WAS UNRELATED TO THE REPORTED EVENT. VISUAL EXAMINATION FOUND ONE EXTERNAL INSULATION ABRASION BREACHING THE OPTIM SHEATH ONLY WITH INTACT SILICONE INSULATION BENEATH AT PROXIMAL REGION CONSISTENT WITH FRICTION TO THE DEVICE CAN. Manufacturer Narrative: CORRECTION: SECTION H6- COMPONENT CODE NUMBER 2 IN FOLLOW UP 1 SHOULD BE "3079 - PATIENT LEAD" RATHER THAN "525 - TUBE." Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21070842&pc=NVN | 2017865-2025-00169 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | NVN | OPTISENSE LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-00163, 2017865-2025-00167, 2017865-2025-00171. IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND PASSED AWAY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, LEFT VENTRICULAR (LV) LEAD AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21070857&pc=NIK | 2017865-2025-00171 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-00163, 2017865-2025-00167, 2017865-2025-00169. IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND PASSED AWAY. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) LEAD, LEFT VENTRICULAR (LV) LEAD AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21071490&pc=GKT | 1219343-2024-00025 | 2024-12-05 05:00 | Death | HAEMONETICS CORPORATION | 2025-01-03 05:00 | GKT | NEXSYS PLASMA COLLECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | BK180185 |  | 1 | Event Description: THIS REPORT IS BEING SUBMITTED TO ALIGN WITH CUSTOMER REPORTING REQUIREMENTS. THE USE OF THE HAEMONETICS MEDICAL DEVICE IN THE DONATION PROCEDURE IS NOT SUSPECTED TO HAVE RESULTED IN THE ADVERSE OUTCOME OF THE DONOR. CUSTOMER MADE HAEMONETICS AWARE THAT A 26-YEAR-OLD MALE EXPIRED ON (B)(6) 2024 FROM AN UNKNOWN CAUSE AT HIS HOME SOMETIME AFTER LEAVING A SUCCESSFUL DONATION. THE DONATION BEGAN AT 8:30 AM ON (B)(6) 2024 FOLLOWING A SUCCESSFUL VENIPUNCTURE IN THE LEFT ARM. NO APHERESIS COMPLICATIONS OR DONOR ADVERSE EVENTS WERE NOTED DURING OR AFTER THE DONATION. THE DONOR CENTER LEARNED OF THE DONOR'S DEATH ON (B)(6) 2024 WHEN THE (B)(6) MEDICAL EXAMINER'S OFFICE CONTACTED THE DONOR CENTER TO REQUEST RECORDS. NO OTHER INFORMATION WAS PROVIDED BY THE MEDICAL EXAMINER'S OFFICE. THE CUSTOMER HAS NO INFORMATION FROM THE ATTENDING PHYSICIAN, SURGEON HOSPITAL REPRESENTATIVE OR HEALTHCARE PROFESSIONAL. Manufacturer Narrative: A HAEMONETICS FIELD SERVICE ENGINEER EVALUATED THE EQUIPMENT USED DURING THE DONATION. DEVICE WAS TESTED AND NO PROBLEM WAS FOUND. ALL OTHER PARAMETERS VERIFIED. FUNCTION TESTS COMPLETED. MACHINE MEETS MANUFACTURER'S SPECIFICATIONS AND IS READY TO USE. BASED ON THIS DESCRIPTION THE DEVICE IS EVALUATED WITH NO DEFECT. THERE ARE NO NONCONFORMANCES AGAINST THE DEVICE SERIAL NUMBER AND NO CAPAS RELATED TO THIS COMPLAINT. A REVIEW OF THE DHR SHOWS NO ISSUES DURING MANUFACTURING AND ALL TESTING PASSED. THE DISPOSABLES USED WITH THE SYSTEM WERE DISCARDED, HOWEVER THERE WERE NO RECALLS OR ADVERSE TRENDS RELATED TO THE PRODUCT LOTS USED IN THE PROCEDURE. THERE IS NO EVIDENCE TO SUGGEST THAT THE DONOR EVENT WAS RELATED TO THE DEVICE OR DISPOSABLES USED DURING THE PLASMAPHERESIS PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21071636&pc=DYB | 3012120746-2025-00002 | 2023-09-03 04:00 | Death | ACUTUS MEDICAL, INC. | 2025-01-03 05:00 | DYB | FLEXCATH CROSS TRANSSEPTAL SOLUTION | Device Dislodged or Dislocated | Cardiac Arrest; Low Blood Pressure/ Hypotension; Vascular Dissection | K210685 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A ROUTINE CRYOBALLOON ABLATION PROCEDURE FOR ATRIAL FIBRILLATION, THE PATIENT WAS TRANSFERRED FROM THE OPERATING ROOM TO THE WARD IN STABLE CONDITION, WITH NORMAL BLOOD PRESSURE AND HEART RATE, AND FEELING WELL. HOWEVER, SHORTLY THEREAFTER, THE PATIENT'S CONDITION DETERIORATED RAPIDLY, RESULTING IN A DROP IN BLOOD PRESSURE AND EVENTUAL ASYSTOLE. DESPITE RESUSCITATION EFFORTS, THE PATIENT PASSED AWAY. AN AUTOPSY REVEALED A SIGNIFICANT ACCUMULATION OF BLOOD IN THE ABDOMINAL CAVITY, CAUSED BY A VENOUS TEAR. ACCORDING TO THE PHYSICIAN, DURING THE INSERTION OF THE FLEXCATH ADVANCE SYSTEM (WITH THE FLEXCATH CROSS NEEDLE/DILATOR) INTO THE VENOUS SYSTEM, THE FLEXCATH CROSS NEEDLE/DILATOR SYSTEM DETACHED AND INADVERTENTLY RETRACTED FROM THE FLEXCATH, CAUSING THE VENOUS TEAR THAT ULTIMATELY LED TO THE PATIENT'S DEATH ON THE WARD. DEVICE WAS DISCARDED.Event Description: IT WAS REPORTED THAT AFTER A ROUTINE CRYOBALLOON ABLATION PROCEDURE FOR ATRIAL FIBRILLATION, THE PATIENT WAS TRANSFERRED FROM THE OPERATING ROOM TO THE WARD IN STABLE CONDITION, WITH NORMAL BLOOD PRESSURE AND HEART RATE, AND FEELING WELL. HOWEVER, SHORTLY THEREAFTER, THE PATIENT'S CONDITION DETERIORATED RAPIDLY, RESULTING IN A DROP IN BLOOD PRESSURE AND EVENTUAL ASYSTOLE. DESPITE RESUSCITATION EFFORTS, THE PATIENT PASSED AWAY. AN AUTOPSY REVEALED A SIGNIFICANT ACCUMULATION OF BLOOD IN THE ABDOMINAL CAVITY, CAUSED BY A VENOUS TEAR. ACCORDING TO THE PHYSICIAN, DURING THE INSERTION OF THE FLEXCATH ADVANCE SYSTEM (WITH THE FLEXCATH CROSS NEEDLE/DILATOR) INTO THE VENOUS SYSTEM, THE FLEXCATH CROSS NEEDLE/DILATOR SYSTEM DETACHED AND INADVERTENTLY RETRACTED FROM THE FLEXCATH, CAUSING THE VENOUS TEAR THAT ULTIMATELY LED TO THE PATIENT'S DEATH ON THE WARD. DEVICE WAS DISCARDED. PER FOLLOW-UP, WHEN PREPARED OUTSIDE OF THE PATIENT, THE 900304 WAS INSERTED INTO THE FLEXCATH AND A CLICK WAS HERD INDICATING SOLID CONNECTION. DURING INTRODUCTION OF THE SYSTEM INTO THE PATIENT THE 900304 DETACHED FROM THE FLEXCATH ("UN-CLICKED") AND RETRACTED, EXPOSING THE SHARP DISTAL EDGE OF THE FLEXCATH WHICH WAS IN THE PATIENT'S VEIN. THE PHYSICIAN BELIEVES THAT THE VEIN TEAR COMPLICATION OCCURRED AT THIS MOMENT. Manufacturer Narrative: PENDING FOLLOW-UP INFORMATION TO COMPLETE INVESTIGATION. Manufacturer Narrative: CORRECTED DATA: D1, D2A, H6. ADDITIONAL DATA: B5. ROOT CAUSE COULD NOT BE DETERMINED. LOT NUMBER WAS NOT PROVIDED AND DEVICE WAS DISCARDED; THEREFORE, NO FURTHER INVESTIGATION COULD BE PERFORMED. TEAR IS A POTENTIAL CLINICAL RISK ASSOCIATED WITH THE USE OF THE THE FLEXCATH CROSS DEVICE, WHICH IS THE SAME AS THE RISK OF A PROCEDURE PERFORMED WITH SIMILAR, COMPETITIVE DEVICES. THE REPORTED ADVERSE EVENT IS KNOWN INHERENT RISK OF THE DEVICE THAT IS DOCUMENTED IN THE LABELING. NO FURTHER INVESTIGATION IS POSSIBLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21072004&pc=NKM | 2135147-2025-00080 | 2010-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-01-03 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "THE IMPACT OF PREPROCEDURAL BLOOD PRESSURE ON OUTCOME AFTER TEER: THE PARADOX OR SOMETHING ELSE" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE EFFECT OF SYSTOLIC BLOOD PRESSURE (SBP) AND ARTERIAL HYPERTENSION ON 1-YEAR OUTCOMES IN PATIENTS WITH FUNCTIONAL AND STRUCTURAL MITRAL REGURGITATION WHO UNDERWENT MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER). DEVICES MENTIONED INCLUDE MITRACLIP AND PASCAL. THE ARTICLE CONCLUDED THAT HIGHER SBP ON ADMISSION AND PREPROCEDURAL LVEF WERE PREDICTORS OF BETTER 1-YEAR SURVIVAL, BUT DID NOT IMPACT 1-YEAR HOSPITALIZATION RATE OR MAJOR ADVERSE CARDIOVASCULAR EVENTS (MACE) COMPOSITE OUTCOME IN ALL PATIENTS WHO UNDERWENT M-TEER. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MARIJANA TADIC (B)(6)] THE TIME FRAME OF THE STUDY WAS JANUARY 2010 TO OCTOBER 2020. A TOTAL OF 858 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. COMORBIDITIES INCLUDED: OBESITY, HEART FAILURE, PREVIOUS CARDIAC INTERVENTION, CORONARY ARTERY DISEASE, HYPERTENSION, DYSLIPIDEMIA, DIABETES, SMOKING, ATRIAL FIBRILLATION, PERIPHERAL ARTERY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, PEPTIC ULCER DISEASE, RENAL FAILURE, ACUTE RENAL FAILURE, HEPATIC CIRRHOSIS, PREVIOUS CANCER, PACEMAKER, CARDIAC RESYNCHRONIZATION THERAPY, INTRA CARDIAC DEVICE, HEART FAILURE MEDICATION, FUNCTIONAL MITRAL REGURGITATION, DEGENERATIVE MITRAL REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, MAJOR ADVERSE CARDIOVASCULAR EVENTS (MYOCARDIAL INFARCTION, STROKE), MITRAL STENOSIS. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLES ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "THE IMPACT OF PREPROCEDURAL BLOOD PRESSURE ON OUTCOME AFTER TEER: THE PARADOX OR SOMETHING ELSE". SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING OBESITY, HEART FAILURE, PREVIOUS CARDIAC INTERVENTION, CORONARY ARTERY DISEASE, HYPERTENSION, DYSLIPIDEMIA, DIABETES, SMOKING, ATRIAL FIBRILLATION, PERIPHERAL ARTERY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, PEPTIC ULCER DISEASE, RENAL FAILURE, ACUTE RENAL FAILURE, HEPATIC CIRRHOSIS, PREVIOUS CANCER, PACEMAKER, CARDIAC RESYNCHRONIZATION THERAPY, INTRA CARDIAC DEVICE, HEART FAILURE MEDICATION, FUNCTIONAL MITRAL REGURGITATION, DEGENERATIVE MITRAL REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, MAJOR ADVERSE CARDIOVASCULAR EVENTS (MYOCARDIAL INFARCTION, STROKE), MITRAL STENOSIS; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21072091&pc=DSQ | 2916596-2024-52906 | 2024-05-06 04:00 | Death | THORATEC CORPORATION | 2025-01-03 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT WAS PLACED ON HOSPICE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21072398&pc=NPT | 2025587-2025-00064 | 2023-12-22 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-03 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Intracranial Hemorrhage | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 1 MONTH AND 20 DAYS FOLLOWING IMPLANT OF A TRANSCATHETER VALVE, A CEREBRAL HEMORRHAGE OCCURRED AND THE PATIENT SUBSEQUENTLY DIED. PER THE PHYSICIAN, THE VALVE AND THE VALVE IMPLANT PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21072740&pc=NPT | 2025587-2025-00068 | 2023-11-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-03 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE VALVE AND THE VALVE IMPLANT PROCEDURE DID NOT CONTRIBUTE TO THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21072857&pc=NPT | 2025587-2025-00069 | 2024-12-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-03 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Degraded; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 6 YEARS AND 5 MONTHS FOLLOWING IMPLANT OF A TRANSCATHETER VALVE, THE VALVE FAILED DUE TO AN UNKNOWN REASON. TREATMENT FOR THE FAILED VALVE WAS PLANNED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT DIED PRIOR TO TREATMENT FOR THE FAILED VALVE. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. H1. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21073208&pc=MVK | 3008642652-2025-00091 | 2024-09-28 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SEVEN INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 02:40:03 ON (B)(6) 2024. AT 07:14:15, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION ARTIFACT. BETWEEN 07:15:06 AND 07:32:23, THE PATIENT RECEIVED SEVEN INAPPROPRIATE TREATMENTS DURING ASYSTOLE. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENTS. THE RESULT OF THESE TREATMENTS WAS THAT THE PATIENT REMAINED IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 07:42:23 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21073212&pc=MVK | 3008642652-2025-00092 | 2024-10-05 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND FIFTEEN INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 21:13:04 ON (B)(6) 2024. AT 00:42:51 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 210 BPM. AT 00:43:28, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 250 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE FOR 19 SECONDS TRANSITIONING TO SINUS BRADYCARDIA/AF FROM 30-100 BPM WITH PVCS. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 00:44:14, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM. AT 00:44:44, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM WITH PVCS AND HB. AT 00:45:21, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS RHYTHM @ 80 BPM WITH PVCS AND PAUSES. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM WITH HB/UNCONDUCTED P WAVES. AT 00:50:45 AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE TRANSITIONING TO AF @ 100 BPM THAT DEGRADES TO VF. AT 00:51:35, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY . POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. BETWEEN 00:55:06 AND 03:29:39, THE PATIENT RECEIVED FOURTEEN INAPPROPRIATE TREATMENTS DURING ASYSTOLE. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENTS. THE RESULT OF THESE TREATMENTS WAS THAT THE PATIENT REMAINED IN ASYSTOLE. THE DEVICE WAS SHUT DOWN AT 03:53:00 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21073214&pc=MVK | 3008642652-2025-00093 | 2024-10-15 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm; Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE APPROPRIATE TREATMENTS AND THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 09:09:13 ON (B)(6) 2024. AT 20:16:07 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT. AT 20:17:32, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS AF @ 80 BPM WITH PVCS. AT 21:52:13, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT. AT 21:53:59, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 90 BPM. AT 23:43:47, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT WITH CPR/ELECTRODE LEAD FALL OFF. AT 23:44:54, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS NSVT WITH CPR/ELECTRODE LEAD FALL OFF. AT 01:22:25 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT WITH CPR/ELECTRODE LEAD FALL OFF. AT 01:23:32, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 170 BPM WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS VF WITH CPR/ELECTRODE LEAD FALL OFF. AT 01:24:17, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS VF WITH CPR/ELECTRODE LEAD FALL OFF. AT 01:24:38, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS VF WITH CPR/ELECTRODE LEAD FALL OFF/TACTILE ARTIFACT. AT 01:25:15, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/ELECTRODE LEAD FALL OFF/TACTILE ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/ELECTRODE LEAD FALL OFF/TACTILE ARTIFACT. AT 01:25:52, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/ELECTRODE LEAD FALL OFF/TACTILE ARTIFACT. POST SHOCK RHYTHM WAS VF WITH MOTION ARTIFACT AND VARYING AMPLITUDES. THE DEVICE WAS SHUT DOWN AT 02:22:17 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21073242&pc=MVK | 3008642652-2025-00094 | 2024-09-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWENTY-EIGHT APPROPRIATE TREATMENTS.
THE DEVICE WAS STARTED UP AT 08:04:58 ON (B)(6) 2024.
AT 15:12:51, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF.
AT 15:13:27, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY FOR 11 SECONDS TRANSITIONING TO SINUS BRADYCARDIA @ 50 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY.
BETWEEN 20:07:43 ON (B)(6) 2024 AND 02:53:40 ON (B)(6) 2024, THE PATIENT RECEIVED TWENTY-SEVEN ADDITIONAL APPROPRIATE TREATMENTS IN RESPONSE TO VF RHYTHM. THE RESULTING HEART RHYTHM OF THE FINAL SHOCK WAS IDIOVENTRICULAR RHYTHM @ 30 BPM.
THE ELECTRODE BELT WAS DISCONNECTED AT 03:22:45 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21073248&pc=MVK | 3008642652-2025-00095 | 2024-09-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 22:21:02 ON (B)(6) 2024. AT 119:43:59 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH MOTION ARTIFACT. AT 19:43:59, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS ASYSTOLE. THE ELECTRODE BELT WAS DISCONNECTED AT 20:09:16 ON (B)(6) 2024 . Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21073251&pc=MVK | 3008642652-2025-00096 | 2024-10-04 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-03 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 00:48:55 ON (B)(6) 2024. AT 00:57:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 00:57:48, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 00:58:17, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 01:04:06 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21074439&pc=NVY | 2017865-2025-00229 | 2024-12-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | NVY | ISOFLEX OPTIM LEAD | Failure to Capture | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P960030 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-00227. RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-00228. IT WAS REPORTED THAT THE DEVICE WAS EXPLANTED AND REPLACED AS IT IS SUBJECT TO THE ZENEX, ASSURITY, ENDURITY LASER ADHESION PREPARATION FIELD SAFETY CORRECTION ACTION WHICH APPLIES TO A SUBSET OF DEVICES DISTRIBUTED AND IMPLANTED OUTSIDE OF THE UNITED STATES. PACING FAILURE BEGAN TO OCCUR, REDUCING THE PATIENT¿S RATE TO 30 BPM AND THE DEVICE WAS UNABLE TO INTERROGATE. CLINICAL ENGINEER INFORMED THAT ALTHOUGH INTERROGATION WAS NOT POSSIBLE, THE RATE HAD RISEN TO ABOUT 70 BPM, AND IT WAS SUSPECTED THAT THE DEVICE HAD BEEN ENTERING BACKUP VVI MODE. AFTER THAT, THE HOSPITAL RECOMMENDED THE PATIENT TO REPLACE THE DEVICE, WHICH WAS REFUSED, AND THE TEMPORARY PACING CATHETER WAS ALSO REFUSED. AND THEN, PACING FAILURE WAS CONFIRMED AGAIN. THE PATIENT DIED IN THE EARLY MORNING OF (B)(6). FURTHER INFORMATION WAS REQUESTED HOWEVER, WAS NOT PROVIDED.Event Description: NEW INFORMATION NOTED THAT THERE WERE NO COMPLAINTS OR DEFECTS AGAINST THE LEADS. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21074440&pc=LWP | 2017865-2025-00227 | 2024-12-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | LWP | ZENEX MRI | Failure to Interrogate; No Pacing | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-00228. RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-00229. IT WAS REPORTED THAT THE DEVICE WAS EXPLANTED AND REPLACED AS IT IS SUBJECT TO THE ZENEX, ASSURITY, ENDURITY LASER ADHESION PREPARATION FIELD SAFETY CORRECTION ACTION WHICH APPLIES TO A SUBSET OF DEVICES DISTRIBUTED AND IMPLANTED OUTSIDE OF THE UNITED STATES. PACING FAILURE BEGAN TO OCCUR, REDUCING THE PATIENT¿S RATE TO 30 BPM AND THE DEVICE WAS UNABLE TO INTERROGATE. CLINICAL ENGINEER INFORMED THAT ALTHOUGH INTERROGATION WAS NOT POSSIBLE, THE RATE HAD RISEN TO ABOUT 70 BPM, AND IT WAS SUSPECTED THAT THE DEVICE HAD BEEN ENTERING BACKUP VVI MODE. AFTER THAT, THE HOSPITAL RECOMMENDED THE PATIENT TO REPLACE THE DEVICE, WHICH WAS REFUSED, AND THE TEMPORARY PACING CATHETER WAS ALSO REFUSED. AND THEN, PACING FAILURE WAS CONFIRMED AGAIN. THE PATIENT DIED IN THE EARLY MORNING OF (B)(6). FURTHER INFORMATION WAS REQUESTED HOWEVER, WAS NOT PROVIDED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: THE REPORTED EVENTS OF NO OUTPUT AND INABILITY TO INTERROGATE WERE CONFIRMED. THE DEVICE WAS RECEIVED WITH NO TELEMETRY COMMUNICATION AND NO OUTPUT. VISUAL INSPECTION OF THE HEADER ATTACHMENT AREA DETECTED AN ANOMALY BETWEEN THE PRE-MOLDED HEADER AND TITANIUM CASE. THE DEVICE WAS CUT OPEN TO ENABLE FURTHER TESTING AND THE BATTERY WAS FOUND DEPLETED. FEEDTHROUGH LEAK TEST WAS PERFORMED, INDICATING A DEVICE HERMETICITY BREACH. THIS IS CONSISTENT WITH FEEDTHROUGH DAMAGE AS A RESULT OF FLUID INTRUSION BETWEEN THE HEADER AND CASE, AND SUBSEQUENT FLUID INGRESS TO INTERNAL ELECTRONICS. HYBRID CIRCUITRY WAS TESTED BY CONNECTING TO AN EXTERNAL POWER SOURCE. TEST RESULTS INDICATED ELEVATED CURRENT DRAIN, CONSISTENT WITH MOISTURE DAMAGE, RESULTING IN THE REPORTED EVENT. Manufacturer Narrative: THE DEVICE IS INCLUDED IN THE ZENEX, ASSURITY, ENDURITY LASER ADHESION PREPARATION ADVISORY ISSUED BY ABBOTT ON (B)(6) 2022 FOR A SUBSET OF DEVICES DISTRIBUTED AND IMPLANTED OUTSIDE OF THE UNITED STATES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21074443&pc=NVN | 2017865-2025-00228 | 2024-12-09 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-03 05:00 | NVN | TENDRIL STS | Failure to Capture | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-00227 RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-00229. IT WAS REPORTED THAT THE DEVICE WAS EXPLANTED AND REPLACED AS IT IS SUBJECT TO THE ZENEX, ASSURITY, ENDURITY LASER ADHESION PREPARATION FIELD SAFETY CORRECTION ACTION WHICH APPLIES TO A SUBSET OF DEVICES DISTRIBUTED AND IMPLANTED OUTSIDE OF THE UNITED STATES. PACING FAILURE BEGAN TO OCCUR, REDUCING THE PATIENT¿S RATE TO 30 BPM AND THE DEVICE WAS UNABLE TO INTERROGATE. CLINICAL ENGINEER INFORMED THAT ALTHOUGH INTERROGATION WAS NOT POSSIBLE, THE RATE HAD RISEN TO ABOUT 70 BPM, AND IT WAS SUSPECTED THAT THE DEVICE HAD BEEN ENTERING BACKUP VVI MODE. AFTER THAT, THE HOSPITAL RECOMMENDED THE PATIENT TO REPLACE THE DEVICE, WHICH WAS REFUSED, AND THE TEMPORARY PACING CATHETER WAS ALSO REFUSED. AND THEN, PACING FAILURE WAS CONFIRMED AGAIN. THE PATIENT DIED IN THE EARLY MORNING OF DECEMBER 18. FURTHER INFORMATION WAS REQUESTED HOWEVER, WAS NOT PROVIDED.Event Description: NEW INFORMATION NOTED THAT THERE WERE NO COMPLAINTS OR DEFECTS AGAINST THE LEADS. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21074662&pc=MNS | 2518422-2025-100054 | 2024-12-24 05:00 | Death | RESPIRONICS, INC. | 2025-01-03 05:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION STATING THAT A DREAMSTATION DEVICE IS RETURNING AS A PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINICAL ASSESSMENT STATES THAT THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION STATING THAT A DREAMSTATION DEVICE IS RETURNING AS A PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINICAL ASSESSMENT STATES THAT THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE ORIGINAL REPORT WAS SUBMITTED AS A FINAL REPORT BUT IT HAS SINCE BEEN CHANGED TO REFLECT AN INITIAL ONLY REPORT. THE DEVICE SERVICED BY 3RD PARTY, DEVICE EVALUATED BY MANUFACTURER, REASON DEVICE NOT EVALUATED, EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE, AND CONCLUSION CODE FIELDS HAVE ALL BEEN UPDATED TO REFLECT THIS CHANGE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION STATING THAT A DREAMSTATION DEVICE IS RETURNING AS A PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINICAL ASSESSMENT PREVIOUSLY STATED THAT THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. DURING EVALUATION OF THE DEVICE AT A THIRD PARTY SERVICE CENTER, NO ERRORS WERE FOUND IN THE DEVICE'S ERROR LOG. NO FAULTS WERE FOUND. THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE PASSED ALL TESTING. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK. NO FURTHER INVESTIGATION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21052373&pc=MNS | 2518422-2025-000037 | 2023-12-29 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | MNS | BIPAP AUTOSV ADV | Degraded | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF BIPAP AUTOSV ADV. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF BIPAP AUTOSV ADV. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTRE FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF VISIBLE FOAM PARTICLES. THE TECHNICIAN CONFIRMED NO PARTICLES IN THE AIR PATH. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTRE CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21052447&pc=BZD | 2518422-2025-000039 | 2023-11-29 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | REMSTAR PRO C-FLEX | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REMSTAR PRO C-FLEX. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REMSTAR PRO C-FLEX. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21052666&pc=QMG | 3015053858-2025-00001 | 2024-12-02 05:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-01-02 05:00 | QMG | SHOCKWAVE C2/C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Difficult to Advance | Perforation of Vessels; Thrombosis/Thrombus |  |  | 1 | Event Description: SHOCKWAVE C2+ INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A LESION IN THE RIGHT CORONARY ARTERY (RCA) AND LEFT ANTERIOR DESCENDING ARTERY (LAD). THE RCA WAS SUCCESSFULLY TREATED WITH NO COMPLICATIONS. AN ATTEMPT TO TREAT THE LAD WAS INITIATED AND THE PROXIMAL SEGMENT WAS MEASURED AT 3.5, BUT THE CATHETER COULD NOT BE ADVANCED TO THE DISTAL SEGMENT. A PREDILATION BALLOON WAS PLACED IN THE PROXIMAL SEGMENT BUT FAILED. A 3.5X12 IVL DEVICE WAS USED, DELIVERING APPROXIMATELY 60 PULSES TO THE LESION AND 40-50 PULSES TO THE PROXIMAL AREA. A 3.25 NON-COMPLIANT BALLOON WAS INFLATED TO 18ATM RESULTING IN PERFORATION. A COVERED STENT WAS DEPLOYED AT THE SITE OF THE PERFORATION, AND PERICARDIAL TAPE WAS APPLIED. POST INTERVENTION ANGIOGRAM SHOWED SHUTDOWN OF THE RCA AND CLOTTING IN THE LEFT SYSTEM. THE PATIENT EXPIRED ON THE TABLE FOLLOWING THE PROCEDURE. A PERICARDIOCENTESIS WAS PERFORMED, REVEALING THROMBOSIS IN BOTH THE RCA AND LAD, WITH CONFIRMED PERFORATION IN THE LAD. THE CLOT WAS NOT ADDRESSED DUE TO RAPID PATIENT DECOMPENSATION. THE PHYSICIAN NOTED THE PATIENT WAS HEPARINIZED, AND ACTIVATED CLOTTING TIMES (ACT'S) WERE MONITORED THROUGHOUT THE PROCEDURE. THE PHYSICIAN NOTED THAT THE THROMBOTIC EVENTS IN THE RCA AND LAD, PARTICULARLY THE PERFORATION OF THE LAD, CONTRIBUTED TO THE PATIENT OUTCOME. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. DIFFICULTIES TO ADVANCE THE CATHETER WAS REPORTED. HOWEVER, IT CANNOT BE CONFIRMED WITH THE INFORMATION AVAILABLE IF THE DIFFICULTY TO ADVANCE WAS ASSOCIATED WITH THE IVL CATHETER OR THE OTHER NON-SHOCKWAVE DEVICES USED DURING THE PROCEDURE. THERE WAS NO MALFUNCTION REPORTED ON THE IVL CATHETER. THE CAUSE OF THE PERFORATION, THROMBOSIS, AND DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. AFTER IVL TREATMENT, A NC BALLOON WAS INFLATED TO 18ATM WHICH COULD HAVE BEEN A CONTRIBUTING FACTOR TO THE PERFORATION; HOWEVER, COULD NOT BE DEFINITELY CONFIRMED. THE PHYSICIAN NOTED THAT THE THROMBOTIC EVENTS IN THE RCA AND LAD, PARTICULARLY THE PERFORATION OF THE LAD, CONTRIBUTED TO THE PATIENT OUTCOME. THE LOT NUMBER OF THE DEVICE PROVIDED WAS INCORRECT. THEREFORE, REVIEW OF THE DEVICE MANUFACTURING AND TEST DOCUMENTATION COULD NOT BE COMPLETED. HOWEVER, SHOCKWAVE MEDICAL INC. HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053310&pc=NVN | 2649622-2025-00016 | 2024-12-10 05:00 | Death | MPRI | 2025-01-02 05:00 | NVN | CAPSURE FIX MRI | Intermittent Capture; Capturing Problem; Unstable Capture Threshold | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P090013 |  | 1 | Event Description: IIT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED INTERMITTENT LOSS OF CAPTURE ON CURRENT ELECTROGRAMS (EGM) AND VARYING RISING THRESHOLDS. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: W1DR01 IPG DOI: (B)(6) 2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053381&pc=OZP | 2032227-2025-100730 | 2022-01-27 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-02 05:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Heart Block | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2022. THE CAUSE OF DEATH WAS HEART ATTACK COMPLICATION FROM HEART SURGERY. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: DIABETES AND HEART PROBLEM. THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE SMARTGUARD/AUTO MODE FEATURE OF THE INSULIN PUMP. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KPK, UNOMEDICAL. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE THE USE OF THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053542&pc=ILJ | 3007420694-2025-00005 | 2023-12-01 05:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-01-02 05:00 | ILJ | SYSTEM 2000 | Use of Device Problem; Insufficient Information | Respiratory Arrest; Insufficient Information |  |  | 1 | Event Description: ARJO WAS NOTIFIED OF AN EVENT INVOLVING SYSTEM 2000 BATH. ARJO SALES REPRESENTATIVES WAS INFORMED OF A PATIENT'S DEATH IN A SYSTEM 2000 BATHTUB. THIS OCCURRED OVER A YEAR AGO. THE CAREGIVER HAD LEFT THE PATIENT (A KID) ALONE TO GO GET CLOTHES AND THE KID HAD PROBABLY DROWNED. Manufacturer Narrative: PROCESS OF COLLECTING AND ANALYZING INFORMATION IS ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED UPON INVESTIGATION CONCLUSION. Manufacturer Narrative: ARJO WAS NOTIFIED OF AN EVENT INVOLVING A SYSTEM 2000 BATH. ARJO SALES REPRESENTATIVE WAS INFORMED OF A PATIENT'S DEATH IN THE SYSTEM 2000 BATHTUB. THIS OCCURRED OVER A YEAR AGO (IN 2023). IT WAS INDICATED THAT THE CAREGIVER HAD LEFT THE PATIENT ALONE TO GO GET CLOTHES AND THE PATIENT HAD DROWNED WHILST THE CAREGIVER WAS GONE. THE INFORMATION COLLECTED DID NOT REVEAL ANY CORRELATION BETWEEN THE REPORTED PATIENT DEATH AND THE ARJO BATH INVOLVED IN THE EVENT. THERE WAS NO ALLEGATION OF ARJO PRODUCT MALFUNCTION. THE PERSON WHO REPORTED THE INCIDENT DIDN¿T PROVIDE DETAILED INFORMATION AS THE EVENT OCCURRED OVER A YEAR AGO (IN 2023). IN THE REPORTED EVENT, THE PATIENT WAS LEFT UNATTENDED, WHICH IS INCONSISTENT WITH THE SYSTEM 2000 INSTRUCTIONS FOR USE. THE SYSTEM 2000 BATH IFU (04.AR.12_22) STATES THE FOLLOWING; "WARNING: TO AVOID INJURY, MAKE SURE THAT THE PATIENT IS NOT LEFT UNATTENDED AT ANY TIME." IT IS ALSO UNKNOWN IF THE PATIENT WAS PLACED IN THE BATH ON THE LIFT, WHICH SHALL BE USED AS A SYSTEM WITH AN APPROPRIATE TRANSFER DEVICE. ACCORDING TO THE INFORMATION INCLUDED IN THE IFU: "COMPLETELY BEDRIDDEN PATIENT/RESIDENTS CAN USE THE BATHTUB. CORRECT TRANSFER EQUIPMENT SHALL BE USED." "THE BATHING MUST BE ASSISTED BY APPROPRIATELY TRAINED CAREGIVERS WITH AN ADEQUATE KNOWLEDGE OF THE CARE ENVIRONMENT AND ITS COMMON PRACTICE AND PROCEDURE, AND IN ACCORDANCE WITH THE GUIDELINES IN THE INSTRUCTION FOR USE (IFU). TRANSFER TO THE BATHTUB SHALL TAKE PLACE BY MEANS OF AN APPROPRIATE TRANSFER DEVICE." TO SUM UP, THE SYSTEM 2000 BATH WAS USED WITH THE PATIENT WHEN THE EVENT OCCURRED. CURRENTLY, ARJO IS NOT AWARE OF ANY DEVICE MALFUNCTION THAT COULD CONTRIBUTE TO THE EVENT. TAKING A CONSERVATIVE APPROACH, THIS COMPLAINT WAS REPORTED FOLLOWING AN INDICATION OF DEATH WHILE USING ARJO BATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053546&pc=ILJ | 1419652-2025-00001 | 2023-12-01 05:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-01-02 05:00 | ILJ | SYSTEM 2000 | Use of Device Problem; Insufficient Information | Respiratory Arrest; Insufficient Information |  |  | 1 | Event Description: ARJO WAS NOTIFIED OF AN EVENT INVOLVING SYSTEM 2000 BATH. ARJO SALES REPRESENTATIVES WAS INFORMED OF A PATIENT'S DEATH IN A SYSTEM 2000 BATHTUB. THIS OCCURRED OVER A YEAR AGO. THE CAREGIVER HAD LEFT THE PATIENT (A KID) ALONE TO GO GET CLOTHES AND THE KID HAD PROBABLY DROWNED.Event Description: ARJO WAS NOTIFIED OF AN EVENT INVOLVING A SYSTEM 2000 BATH. ARJO SALES REPRESENTATIVE WAS INFORMED OF A PATIENT'S DEATH IN THE SYSTEM 2000 BATHTUB. THIS OCCURRED OVER A YEAR AGO (IN 2023). IT WAS INDICATED THAT THE CAREGIVER HAD LEFT THE PATIENT ALONE TO GO GET CLOTHES AND THE PATIENT HAD DROWNED WHILST THE CAREGIVER WAS GONE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053739&pc=PYX | MW5164077 | 2024-12-07 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | PYX | IMPELLA RP SMART- ASSIST CATHETER |  | Insufficient Information |  |  | 1 | Event Description: IMPELLA CONSOLE DISPLAY UNABLE TO RECEIVE PLACEMENT SIGNAL FROM IMPELLA CATHETER DUE TO FAULTY CATHETER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053746&pc=DTE | MW5164078 | 2024-12-19 05:00 | Death | MEDTRONIC, INC. | 2025-01-02 05:00 | DTE | DUAL CHAMBER TEMPORARY PACEMAKER |  | Cardiac Arrest; Ventricular Fibrillation |  |  | 1 | Event Description: CODE BLUE EVENT IN CRITICAL CARE ON A POST OPEN HEART SURGICAL PATIENT ON (B)(6) 2024. THE PATIENT HAD A R ON T VENTRICULAR FIBRILLATION ARREST. CONCERN FOR POSSIBLE DEVICE FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053760&pc=KNT | 3010757606-2025-00001 |  | Death | ABBVIE - MEDICAL DEVICE CENTER | 2025-01-02 05:00 | KNT | DUODOPA_DUOPA | Adverse Event Without Identified Device or Use Problem | Peritoneal Laceration(s); Pneumonia | K142793 |  | 1 | Event Description: ON AN UNKNOWN DATE A PATIENT IN SPAIN UNDERWENT A PROCEDURE FOR THE PLACEMENT OF A PERCUTANEOUS ENDOSCOPIC GASTROSTOMY(PEG) TUBE. ON (B)(6) 2024, IT WAS REPORTED THE PATIENT HAD PERITONITIS DUE TO THE TRANSLOCATED PEG TUBE THAT WAS COMPLICATED WITH PNEUMONIA AND PATIENT PASSED AWAY 2 YEARS AGO. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. DESPITE MULTIPLE QUERIES, THERE IS NO CLEAR EVIDENCE LINKING THE DEVICE TO THE REPORTED EVENT, AND NO ADDITIONAL INFORMATION REGARDING DEVICE-RELATED ALLEGATIONS, MALFUNCTIONS, COMPLICATIONS, DIAGNOSTIC TESTING, OR TREATMENT IS AVAILABLE. THEREFORE, OUT OF ABUNDANCE OF CAUTION, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT WITHOUT ATTRIBUTING CAUSALITY TO THE DEVICE. Manufacturer Narrative: REFERENCE NUMBER (B)(4). THE DEVICE MANUFACTURER AND LOT NUMBER OF THE DEVICE INVOLVED IN THIS COMPLAINT WAS NOT PROVIDED. THEREFORE, IT IS UNKNOWN IF THE DEVICE INVOLVED WAS ABBVIE BRANDED TUBING. CONSERVATIVELY, ABBVIE HAS CHOSEN TO REPORT THIS COMPLAINT DUE TO THE POTENTIAL THAT THE DEVICE INVOLVED COULD HAVE BEEN ABBVIE BRANDED TUBING. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053795&pc=FNL | 1419652-2025-00003 | 2024-10-24 04:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-01-02 05:00 | FNL | ENTERPRISE 5000X | Use of Device Problem | Strangulation |  |  | 1 | Event Description: ARJO WAS NOTIFIED BY (B)(6) OF THE EVENT REGARDING THE PATIENT'S DEATH FROM INTENTIONALLY LOWERING THE BACK REST ONTO HIS/HER NECK RESULTING IN ASPHYXIATION. THE FOLLOWING INFORMATION WAS PROVIDED: ¿COMPLETED SUICIDE. THE MECHANICAL AGENTS OF THE BED WERE USED. THE BACK REST WAS ELEVATED, BED RAILS DOWN. THE PATIENT HAD ACCESSED THE BED FROM THE SIDE PLACING THEIR HEAD ON THE BAR, TAKING A KNEELING POSITION FACING DOWN AND USING THE PATIENT CONTROLLER BROUGHT THE BACK REST DOWN ON THE BACK OF THEIR NECK TO COMPRESS AIRWAY. BED HAS BEEN ASSESSED FULLY BY MEDICAL PHYSICS POST INCIDENT REPORT INDICATES IN FULL WORKING ORDER.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21053796&pc=FNL | 3007420694-2025-00006 | 2024-10-24 04:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-01-02 05:00 | FNL | ENTERPRISE 5000X | Use of Device Problem | Strangulation |  |  | 1 | Event Description: ARJO WAS NOTIFIED BY MHRA OF THE EVENT REGARDING THE PATIENT'S DEATH FROM INTENTIONALLY LOWERING THE BACK REST ONTO HIS/HER NECK RESULTING IN ASPHYXIATION. THE FOLLOWING INFORMATION WAS PROVIDED: ¿COMPLETED SUICIDE. THE MECHANICAL AGENTS OF THE BED WERE USED. THE BACK REST WAS ELEVATED, BED RAILS DOWN. THE PATIENT HAD ACCESSED THE BED FROM THE SIDE PLACING THEIR HEAD ON THE BAR, TAKING A KNEELING POSITION FACING DOWN AND USING THE PATIENT CONTROLLER BROUGHT THE BACK REST DOWN ON THE BACK OF THEIR NECK TO COMPRESS AIRWAY. BED HAS BEEN ASSESSED FULLY BY MEDICAL PHYSICS POST INCIDENT REPORT INDICATES IN FULL WORKING ORDER.¿ Manufacturer Narrative: ARJO WAS NOTIFIED BY THE MHRA (MEDICINES AND HEALTHCARE PRODUCTS REGULATORY AGENCY) OF AN EVENT REGARDING A PATIENTS' DEATH CAUSED BY INTENTIONALLY LOWERING THE BACK REST ONTO THE NECK, RESULTING IN ASPHYXIATION. THE FOLLOWING INFORMATION WAS PROVIDED: ¿COMPLETED SUICIDE. THE MECHANICAL AGENTS OF THE BED WERE USED. THE BACK REST WAS ELEVATED, BED RAILS DOWN. THE PATIENT HAD ACCESSED THE BED FROM THE SIDE PLACING THEIR HEAD ON THE BAR, TAKING A KNEELING POSITION FACING DOWN AND USING THE PATIENT CONTROLLER BROUGHT THE BACK REST DOWN ON THE BACK OF THEIR NECK TO COMPRESS AIRWAY. BED HAS BEEN ASSESSED FULLY BY MEDICAL PHYSICS POST INCIDENT REPORT INDICATES IN FULL WORKING ORDER.¿ IRIC (INCIDENT REPORTING AND INVESTIGATION CENTRE) IS INVESTIGATING THIS CASE AS A SUICIDE. BASED ON THE INFORMATION RECEIVED, THE INVOLVED ENTERPRISE 5000X BED WAS INSPECTED BY MEDICAL PHYSICS AFTER THE EVENT. NO DEVICE MALFUNCTION WAS INDICATED, AND THE DEVICE WAS FOUND TO BE IN FULL WORKING ORDER. IT IS NOT KNOWN WHY THE PATIENT PERFORMED SUCH AN ACTION. THIS IS CONSIDERED INTENTIONAL DEVICE USE FOR A NON-APPROVED PURPOSE. THE INSTRUCTIONS FOR USE FOR ENTERPRISE 5000X BED (746-577-EN) INCLUDES THE FOLLOWING INFORMATION RELATED TO THE ENTRAPMENT HAZARDS: - "BEFORE OPERATING THE BED, MAKE SURE THE PATIENT IS POSITIONED CORRECTLY TO AVOID ENTRAPMENT." - "TO ENSURE THE PATIENT CAN USE THE BED SAFELY, THEIR AGE, SIZE AND CONDITION SHOULD BE ASSESSED BY A CLINICALLY QUALIFIED PERSON." THE ENTERPRISE 5000X BED WAS WORKING PER MANUFACTURERS¿ SPECIFICATION WHEN THE EVENT OCCURRED. THE DEVICE WAS USED BY A PATIENT AT THE TIME OF THE EVENT. TAKING A CONSERVATIVE APPROACH, ARJO DECIDED TO REPORT THE EVENT DUE TO THE PATIENT DEATH REPORTED WHILE USING THE ARJO BED. Manufacturer Narrative: PROCESS OF COLLECTING AND ANALYZING INFORMATION IS ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED UPON INVESTIGATION CONCLUSION. THE UDI NUMBER IS NOT AVAILABLE AS THE DEVICE WAS MANUFACTURED BEFORE UDI IMPLEMENTATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054087&pc=OZD | 1220648-2025-25272 | 2024-12-12 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA 5.5 | Failure to Advance | Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. DURING PLACEMENT, DIFFICULTY WITH THE PUMP AND GUIDEWIRE WAS ENCOUNTERED, FOLLOWED BY A STROKE WITH A DEMISED OUTCOME. THE PUMP WAS UNABLE TO PASS ON .018 GUIDEWIRE. THE PHYSICIAN STATED THE GUIDEWIRE WAS TOO LAX AND BUCKLING UNDER PRESSURE OF THE PUMP ADVANCEMENT. THE DECISION WAS MADE TO USE .027 WIRE FROM AN IMPELLA RP FLEX KIT, AND THE PUMP SUCCESSFULLY ADVANCED. NEXT DAY AFTER START OF SUPPORT, IT WAS NOTED THE PATIENT EXPERIENCED A BILATERAL ACUTE MIDDLE CEREBRAL ARTERY (MCA) STROKE. PROGNOSIS AND CHANGE IN NEUROLOGIC STATUS WERE INDICATED INCOMPATIBLE WITH RECOVERY. CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE DELIVERY ISSUES AND THE THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL DETAILS PROVIDED WERE SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, THE GUIDEWIRE WAS UNABLE TO PASS THROUGH THE VASCULATURE. THE ROOT CAUSE OF THE THROMBUS COULD NOT BE DETERMINED WITHOUT ADDITIONAL INFORMATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054153&pc=NPT | 2025587-2025-00005 | 2024-07-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-02 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Non specific EKG/ECG Changes | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, A PERMANENT PACEMAKER WAS IMPLANTED DUE TO AN UNS PECIFIED SEVERE CONDUCTION DISTURBANCE. THE PATIENT CONTINUED TO HAVE DELIRIUM AND WAS UNABLE TO PERFORM REHABILITATION AND DIALYSIS ADEQUATELY. SUBSEQUENTLY, A DECREASE IN CONSCIOUSNESS LEVEL WAS OBSERVED, AND A HEAD COMPUTED TOMOGRAPHY WAS PERFORMED AND WAS NEGATIVE FOR AN INTRACRANIAL HEMORRHAGE; HOWEVER, THE PATIENT STOPPED BREATHING AND DIED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS RENAL, AND THE VALVE AND IMPLANT PROCEDURE WERE NOT RELATED. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054178&pc=MKJ | 3004123209-2025-00002 | 2024-12-19 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-01-02 05:00 | MKJ | UNKNOWN_MEDICAL - BELFAST_PRODUCT | Unexpected Shutdown | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED THEIR DISTRIBUTOR TO REPORT THAT THEIR DEVICE POWERED OFF UNEXPECTEDLY DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED.Event Description: THE CUSTOMER CONTACTED THEIR DISTRIBUTOR TO REPORT THAT THEIR DEVICE POWERED OFF UNEXPECTANTLY DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. THE HDF-3500 DEVICE IS NOT AVAILABLE FOR DISTRIBUTION IN THE UNITED STATES BUT IS SIMILAR TO THE SAM 350P AND 450P WHICH IS DISTRIBUTED IN THE UNITED STATES. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. NO RESPONSE HAS BEEN RECEIVED FROM THE CUSTOMER. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. A CLINICAL REVIEW OF THE REPORTED EVENT WILL BE PERFORMED. THE CUSTOMER WILL BE PROVIDED WITH A REPLACEMENT DEVICE. Manufacturer Narrative: HEARTSINE EVALUATED THE CUSTOMER'S DEVICE AND FOUND NO FAULT WITH THE DEVICE. A REVIEW OF THE ELECTRONIC RECORDS FROM THE DEVICE REVEALED THAT DURING ALL 9 POWER-ONS, THE DEVICE POWERED OFF SUDDENLY WITHOUT ENTERING THE SHUTDOWN SEQUENCE, WHICH WOULD INDICATE POWER HAD BEEN REMOVED FROM THE DEVICE. THIS INDICATES USER ERROR FOR THE DEVICE POWERING OFF, POTENTIALLY AN ISSUE WITH PAD-PAK SEATING. A CLINICAL REVIEW OF THE EVENT WAS PERFORMED AND OBSERVED THE PATIENT PRESENTED WITH A NON-SHOCKABLE RHYTHM FOR THE DURATION OF THE EVENT. THE DEVICE PERFORMED TO SPECIFICATION. THE CUSTOMER WAS SENT A REPLACEMENT DEVICE, AND THE RETURNED DEVICE WAS SCRAPPED BY HEARTSINE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054408&pc=NKM | 21054408 | 2024-04-15 04:00 | Death | ABBOTT | 2025-01-02 05:00 | NKM | MITRACLIP G4 |  | Low Blood Pressure/ Hypotension; Pericardial Effusion; Mitral Valve Insufficiency/ Regurgitation |  |  | 1 | Event Description: SL1 AND BRK NEEDLE SET WAS USED FOR TRANSSEPTAL PUNCTURE. IT WAS A RATHER DIFFICULT TRANSSEPTAL GIVEN SIGNIFICANT ENLARGEMENT OF THE LEFT ATRIUM. SUBSEQUENTLY, AN AMPLATZ WIRE WAS PLACED IN THE LEFT UPPER PULMONARY VEIN. A MITRAL CLIP DELIVERY SHEATH WAS THEN INSERTED INTO THE LEFT ATRIUM WITHOUT. NO PERICARDIAL EFFUSION IS OBSERVED. INITIAL STRATEGY WAS TO PLACE AN NTW DEVICE ON THE MEDIAL SIDE OF THE A2/P2 SEGMENT WITH HOPE TO ADD A 2ND DEVICE TO FURTHER REDUCTION OF THE RESIDUAL MITRAL REGURGITATION. HOWEVER, THE 1ST GRASP DID NOT IMPACT THE MITRAL REGURGITATION. WE THEN ATTEMPTED MIDPORTION OF A2/P2 AND THEN EVENTUALLY HAD TO USE THE LATERAL SIDE OF THE A2/P2 GIVEN SIGNIFICANT REGURGITATION JET ORIENTED FROM THE LATERAL SIDE. THE INITIAL GRASP ON THE MEDIAL SIDE OF A2/P2 ALSO GENERATED SIGNIFICANT TRANSMITRAL GRADIENT UP TO 9 MMHG AT A HEART RATE OF 80 BEATS PER MINUTES. EVENTUALLY, WE HAVE TO ATTEMPT ON THE LATERAL SIDE OF THE A2/P2 SEGMENT FOR BETTER GRASP, HOWEVER, DUE TO SIGNIFICANT CHORDAE INTERACTION, WE HAD TO RETRACT SEVERAL TIMES INTO THE LEFT ATRIUM TO REPOSITION THE CLIP INTO THE LATERAL SIDE OF A2/P2 SEGMENT. SURPRISINGLY, WE OBSERVED SUDDEN HYPOTENSION AFTER THE LAST GRASP OF A2/P2 SEGMENT BEFORE WE FURTHER TEST OF THE OPTIMAL GRASP AND REGURGITATION JET IN THE REDUCTION. THERE IS A LARGE PERICARDIA! EFFUSION ACCUMULATED QUICKLY. THE PT WAS RESUSCITATED BUT LATER DIED IN THE CRITICAL CARE UNIT. MULTIDISCIPLINARY REVIEW COMPLETED INCLUDING REVIEW OF TECHNIQUES EMPLOYED (CONDUCTED BY SECTION HEAD). INVESTIGATION DID NOT IDENTIFY DEVIATION FROM STANDARD PERFORMANCE. (B)(6) 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054454&pc=NGV | 21054454 |  | Death | BOSTON SCIENTIFIC | 2025-01-02 05:00 | NGV | WATCHMAN FLX PRO |  | Insufficient Information |  |  | 1 | Event Description: PATIENT WAS SCHEDULED FOR A WATCHMAN DEVICE IMPLANT AS AN OUTPATIENT IN THE CARDIAC CATHETERIZATION LAB. THE PROVIDER WAS UNABLE TO SEAT THE WATCHMAN PROPERLY AND REMOVED THE DEVICE. THE PROVIDER THEN PLACED AN AMULET DEVICE, WHICH DID EVENTUALLY SEAT CORRECTLY. WHILE THE PATIENT WAS IN THE RECOVERY AREA, SHE EXPERIENCED A RAPID CLINICAL DETERIORATION, AND WAS TAKEN URGENTLY BACK TO THE CARDIAC CATHETERIZATION LAB, WHERE THE AMULET WAS REMOVED WITH DIFFICULTY, AND AN ATTEMPT OF MITRA! CLIP WAS UNSUCCESSFUL. THE PATIENT WAS THEN TAKEN URGENTLY TO THE CRITICAL CARE UNIT, WHERE SHE CONTINUED TO DECLINE, AND SHE DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054513&pc=BZD | 2518422-2025-00006 | 2024-11-08 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF DS2ADV AUTO CPAP, PATIENT'S WIFE CALLED IN STATING THAT THE PATIENT HAS PASSED AWAY AND IS ASKING FOR A RETURN LABEL AS SHE RECEIVED A REPLACEMENT DEVICE BUT DOES NOT NEED ONE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054530&pc=OZD | 1220648-2025-25274 | 2024-03-06 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED GASTROINTESTINAL (GI) BLEED WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA 5.5 DEVICE USED: ¿3/6. LARGE BLOOD BM X 2 WITH HYPOTENSION, HEPARIN STOPPED, VASOPRESSIN INITIATED ALONG WITH BLOOD TRANSFUSIONS 3/7. ADDITIONAL GI BLEED 7 UNITS OF PRBCS AND 1-UNIT FFP¿. THE PATIENT RECOVERED WITH NO SEQUELAE. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. THE CAUSE OF THE VASCULAR DAMAGE- PERIPHERAL VESSEL ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED, UNKNOWN RELATION TO IMPELLA WITH GASTROINTESTINAL BLEED BEING OBSERVED AND DUE TO INSUFFICIENT INFORMATION. G.2 REVISED AS STUDY WAS MISSING FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25274.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054792&pc=MIH | 2247858-2025-00003 | 2024-12-09 05:00 | Death | BOLTON MEDICAL, INC. | 2025-01-02 05:00 | MIH | CUSTOM-MADE TREO ABDOMINAL STENT-GRAFT SYSTEM | Separation Failure; Mechanical Jam | Foreign Body In Patient | P190015 |  | 1 | Event Description: "DURING THE RELEASE OF THE MAIN BODY THE CONSTRICTION WIRE WAS TORN INSIDE THE DELIVERY SYSTEM APPROX. 10 CM TO THE DISTAL END OF THE SYSTEM. THE DELIVERY SYSTEM HAD TO BE DISMANTLED TO REACH THE REST OF THE WIRE TO BE ABLE TO RELEASE THE BODY. AFTER SEVERAL ATTEMPTS THE WIRE COULD BE PULLED BACK. THE CONE TIP COULD ALSO NOT BE RETRACTED. FINALLY, THE GRAFT WAS NOT IN AN OPTIMAL POSITION BUT THE PHYSICIAN MANAGED TO CANNULATE ALL TARGET VESSELS AND PLACED THE BRIDGING STENTS. FINAL ANGIOGRAM SHOWED A SATISFYING RESULT. THE WHOLE PROCEDURE TOOK ABOUT 7 HOURS." PATIENT OUTCOME: "PATIENT CONDITION IS STABLE. ADDITIONAL INFORMATION RECEIVED ON 12/16/2024: PATIENT PASSED AWAY ON POD 2."Event Description: "DURING THE RELEASE OF THE MAIN BODY THE CONSTRICTION WIRE WAS TORN INSIDE THE DELIVERY SYSTEM APPROX. 10CM TO THE DISTAL END OF THE SYSTEM. THE DELIVERY SYSTEM HAD TO BE DISMANTLED TO REACH THE REST OF THE WIRE TO BE ABLE TO RELEASE THE BODY. AFTER SEVERAL ATTEMPTS THE WIRE COULD BE PULLED BACK. THE CONE TIP COULD ALSO NOT BE RETRACTED. FINALLY, THE GRAFT WAS NOT IN AN OPTIMAL POSITION BUT THE PHYSICIAN MANAGED TO CANNULATE ALL TARGET VESSELS AND PLACED THE BRIDGING STENTS. FINAL ANGIOGRAM SHOWED A SATISFYING RESULT. THE WHOLE PROCEDURE TOOK ABOUT 7 HOURS." PATIENT OUTCOME: "PATIENT CONDITION IS STABLE. ADDITIONAL INFORMATION RECEIVED ON 12/16/2024: PATIENT PASSED AWAY ON POD 2." Manufacturer Narrative: BOLTON MEDICAL IS VOLUNTARILY REPORTING AN EVENT RELATED TO A TREO CUSTOM-MADE DEVICE. THE CUSTOM MADE TREO DEVICES ARE NOT MARKETED IN THE US, HOWEVER THEY ARE SIMILAR TO THE TREO ABDOMINAL STENT GRAFT DELIVERY SYSTEM APPROVED FOR SALE IN THE US (P190015). THE EVENT OCCURRED IN GERMANY. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: BOLTON MEDICAL IS VOLUNTARILY REPORTING AN EVENT RELATED TO A TREO CUSTOM-MADE DEVICE. THE CUSTOM-MADE TREO DEVICES ARE NOT MARKETED IN THE US, HOWEVER THEY ARE SIMILAR TO THE TREO ABDOMINAL STENT GRAFT DELIVERY SYSTEM APPROVED FOR SALE IN THE US (P190015). THE EVENT OCCURRED IN GERMANY. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054799&pc=MIH | 2247858-2025-00005 | 2024-12-09 05:00 | Death | BOLTON MEDICAL, INC. | 2025-01-02 05:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Retraction Problem; Material Deformation; Insufficient Information | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P200045 |  | 1 | Event Description: "PATIENT HAD PREVIOUS TEVAR WITH CHRONIC TYPE B DISSECTION AND 7.2 CM ANEURYSM; HE RUPTURED AND WAS STABLE AT TIME OF IMPLANTATION. HIS PREVIOUS TEVAR EXTENDED TO 5 CM PROXIMAL TO CELIAC ARTERY; THE RUPTURE WAS AT THE DISTAL END OF THE EXISTING GRAFT. WE OPENED N4-42-159-42U REF#(B)(4) AND "BAREBACKED IT" INTO THE R CIA. DEVICE WENT UP WITH NO ISSUE; IN THE VISCERAL SEGMENT WE TOOK THE FLEXIBLE INNER SHEATH OUT OF HARD OUTER SHEATH "IN STEP 1" AND PLACED IT IN LANDING ZONE; WE MOVED TO STEP 2 AND DR (B)(6) ATTEMPTED TO TURN HANDLE COUNTER CLOCKWISE TO LINE UP D MARKER WITH GRAFT MARKER. DR (B)(6) HAS DEPLOYED OVER 50 RELAY PROS AND IS A TRAINER FOR TA. HE COMPLAINED THAT THE HANDLE KEPT RECOILING BACK TO ORIGINAL POSITION EVERY TIME HE TURNED IT; UNDER FLUORO, IT WAS CLEAR THAT THE D MARKER WASN'T MOVING, AND EVERY TIME HE TURNED THE HANDLE, THE NOSE CONE WOULD RECOIL DOWN TOWARD THE PATIENTS FEET. AFTER A FEW ATTEMPTS IT WAS CLEAR THE GRAFT WAS GETTING SCRUNCHED WITH EVERY ATTEMPTED TURN OF THE HANDLE. WE MOVED THE CONTROL PANEL TO STEP 1 IN AN ATTEMPT TO RECAPTURE THE GRAFT (STILL CONSTRAINED BY INNER SHEATH) AND WERE UNABLE TO (EVEN THOUGH WE HAVE DONE THIS BEFORE IN THE PAST). WE WERE FORCED TO PULL THE ENTIRE DEVICE AS-IS OUT OF THE BODY, RESULTING IN INTIMA COMING OUT WITH THE GRAFT. THE PATIENT REMAINED STABLE AND WE PROCEEDED WITH AN ADDITIONAL GRAFT. WE SEALED THE ANEURYSM, CLOSED THE PATIENT, AND PATIENT WAS MOVED FROM OPERATING ROOM TO PACU. PATIENT CODED IN PACU AND BECAME HYPOTENSIVE AND DIED." PATIENT OUTCOME: "DEATH APPROX. 1 HR AFTER EVENTUAL IMPLANT."Event Description: "PATIENT HAD PREVIOUS TEVAR WITH CHRONIC TYPE B DISSECTION AND 7.2 CM ANEURYSM; HE RUPTURED AND WAS STABLE AT TIME OF IMPLANTATION. HIS PREVIOUS TEVAR EXTENDED TO 5 CM PROXIMAL TO CELIAC ARTERY; THE RUPTURE WAS AT THE DISTAL END OF THE EXISTING GRAFT. WE OPENED N4-42-159-42U REF#(B)(4) AND "BAREBACKED IT" INTO THE R CIA. DEVICE WENT UP WITH NO ISSUE; IN THE VISCERAL SEGMENT WE TOOK THE FLEXIBLE INNER SHEATH OUT OF HARD OUTER SHEATH "IN STEP 1" AND PLACED IT IN LANDING ZONE; WE MOVED TO STEP 2 AND DR (B)(6) ATTEMPTED TO TURN HANDLE COUNTER-CLOCKWISE TO LINE UP D MARKER WITH GRAFT MARKER. DR (B)(6) HAS DEPLOYED OVER 50 RELAY PROS AND IS A TRAINER FOR TA. HE COMPLAINED THAT THE HANDLE KEPT RECOILING BACK TO ORIGINAL POSITION EVERY TIME HE TURNED IT; UNDER FLUORO, IT WAS CLEAR THAT THE D MARKER WASN'T MOVING, AND EVERY TIME HE TURNED THE HANDLE, THE NOSE CONE WOULD RECOIL DOWN TOWARD THE PATIENTS' FEET. AFTER A FEW ATTEMPTS IT WAS CLEAR THE GRAFT WAS GETTING SCRUNCHED WITH EVERY ATTEMPTED TURN OF THE HANDLE. WE MOVED THE CONTROL PANEL TO STEP 1 IN AN ATTEMPT TO RECAPTURE THE GRAFT (STILL CONSTRAINED BY INNER SHEATH) AND WERE UNABLE TO (EVEN THOUGH WE HAVE DONE THIS BEFORE IN THE PAST). WE WERE FORCED TO PULL THE ENTIRE DEVICE AS-IS OUT OF THE BODY, RESULTING IN INTIMA COMING OUT WITH THE GRAFT. THE PATIENT REMAINED STABLE AND WE PROCEEDED WITH AN ADDITIONAL GRAFT. WE SEALED THE ANEURYSM, CLOSED THE PATIENT, AND PATIENT WAS MOVED FROM OR TO PACU. PATIENT CODED IN PACU AND BECAME HYPOTENSIVE AND DIED." PATIENT OUTCOME: "DEATH APPROX. 1 HR AFTER EVENTUAL IMPLANT." Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054898&pc=DTB | 3008973940-2025-00031 | 2024-12-11 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-02 05:00 | DTB | CAPSURE SENSE LEAD MRI SURESCAN | Impedance Problem | Insufficient Information | P830061 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED SUDDEN INCREASE IN IMPEDANCE.THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055121&pc=DSQ | 3007042319-2025-00001 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-01-02 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING BRIDGING TO HEART TRANSPLANTATION WITH THE USE OF A TEMPORARY OR DURABLE LEFT VENTRICULAR ASSIST DEVICE (VAD). THERE WERE TWO PATIENTS WITH A DURABLE LEFT VAD THAT DID NOT SURVIVE TO HEART TRANSPLANTATION; THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO WERE TRANSPLANTED AFTER PUMP THROMBOSIS OR DRIVELINE INFECTION. AT THE TIME OF TRANSPLANTATION, THERE WERE MAJOR AND MINOR COMPLICATIONS IN PATIENTS WITH A LEFT VAD. MAJOR COMPLICATIONS INCLUDED STROKE, INFECTION, RIGHT VENTRICULAR (RV) FAILURE, UNKNOWN BLEEDING WHICH REQUIRED MORE THAN TWO UNITS OF PACKED RED BLOOD CELLS (PRBC), AND PUMP THROMBOSIS. MINOR COMPLICATIONS INCLUDED SURGICAL SITE HEMATOMA AND HEMOLYSIS. THERE WERE ALSO UNKNOWN FAILURES OF CONTROLLERS. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/50 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: COMPARISON OF INTRAOPERATIVE BLOOD PRODUCT USE DURING HEART TRANSPLANTATION IN PATIENTS BRIDGED WITH IMPELLA 5.5 VERSUS DURABLE LEFT VENTRICULAR ASSIST DEVICES. JOURNAL OF CARDIOTHORACIC AND VASCULAR ANESTHESIA. 2024. 38; 2567-2575. DOI: 10.1053/J.JVCA.2024.04.047 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055359&pc=OZP | 2032227-2025-100924 | 2024-12-10 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-02 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hyperglycemia; Hypoglycemia; Renal Failure; Dizziness; Malaise | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024 DATE OF DEATH. CAUSE OF DEATH WERE DIABETES AND KIDNEY AND WAS TREATED IN THE EMERGENCY ROOM. REASON FOR ADMISSION TO HOSPITAL WERE NOT FEELING WELL, DIZZINESS. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-7841ZN, MMT-7040A, MMT-332A, MMT-242A. TROUBLESHOOTING WAS PERFORMED AND THE CUSTOMER WAS USING A SENSOR AT TIME OF EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-242A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055550&pc=OZD | 1220648-2025-25279 | 2024-12-14 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA CP | Increase in Pressure | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT ON THE SAME DAY OF IMPELLA CP INSERTION SOME OOZING WAS NOTED AT THE PATIENT¿S GROIN SITE. MANUAL PRESSURE WAS HELD, FEMSTOP, AND SUTURE WAS ADDED. THE BLEEDING WAS RESOLVED. ADDITIONALLY, THE PATIENT FOLLOWING A COUGHING EPISODE HAD POSITIONING ISSUES, FOLLOWED BY PURGE PRESSURE HIGH ALARM WITH FLOWS AROUND 1.5ML/HR. DURING THIS TIME THE PATIENT ALSO EXPERIENCED STROKE LIKE SYMPTOMS. COMPUTED TOMOGRAPHY SCAN WAS NEGATIVE, AND SYMPTOMS IMPROVED WITHIN AN HOUR. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL, BASED ON DATA LOG REVIEW. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055627&pc=BZD | 2518422-2025-000087 | 2024-06-28 04:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | REMSTAR PLUS | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055628&pc=BZD | 2518422-2025-000072 | 2024-05-06 04:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | DREAMSTATION AUTO CPAP | Insufficient Information | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055654&pc=NPT | 2025587-2025-00015 | 2024-10-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-02 05:00 | NPT | EVOLUT FX VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information; Patient Device Interaction Problem | Aneurysm; Bradycardia; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S BLOOD PRESSURE DROPPED TO 20 MM HG WITH A HEART RATE OF 40 BEATS PER MINUTE. TEN MINUTES LATER, ASYSTOLE WAS NOTED. DEATH WAS CONFIRMED 7 MINUTES AFTER THE ASYSTOLE BEGAN. PER THE PHYSICIAN, THE PRIMARY CAUSE OF DEATH WAS INTESTINAL NECROSIS DUE TO SEPSIS. CONTRIBUTING FACTORS TO THE PATIENT'S DEATH WERE A RIGHT VERTEBRAL ANEURYSM AND A HEMORRHAGIC INFARCTION.Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S BLOOD PRESSURE DROPPED TO 20 MM HG WITH A HEART RATE OF 40 BEATS PER MINUTE. TEN MINUTES LATER, ASYSTOLE WAS NOTED. DEATH WAS CONFIRMED 7 MINUTES AFTER THE ASYSTOLE BEGAN. PER THE PHYSICIAN, THE PRIMARY CAUSE OF DEATH WAS INTESTINAL NECROSIS DUE TO SEPSIS. CONTRIBUTING FACTORS TO THE PATIENT'S DEATH WERE A RIGHT VERTEBRAL ANEURYSM AND A HEMORRHAGIC INFARCTION. ADDITIONAL INFORMATION WAS RECEIVED TO CONFIRM PER THE PHYSICIAN, NEITHER A MEDTRONIC DEVICE NOR THE TRANSCATHETER AORTIC VALVE REPL ACEMENT PROCEDURE WERE CONTRIBUTING FACTORS TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. DEVICE CODE, EVAL CODE RESULT, AND EVAL CODE CONCLUSION WERE UPDATED. CORRECTED DATA: E. THE FULL FACILITY ADDRESS WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. H6. PATIENT CODE WAS INADVERTENTLY OMITTED FROM THE INITIAL MEDWATCH REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2: DATE IS APPROXIMATE, YEAR IS VALID. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055862&pc=BZD | 2518422-2025-00015 | 2024-12-09 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING METASTATIC ESOPHAGEAL CANCER AND DEATH. NO DEVICE INFORMATION WAS PROVIDED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056014&pc=FDI | 3005099803-2024-05657 | 2017-10-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-02 05:00 | FDI | ULTRATOME | Adverse Event Without Identified Device or Use Problem | Perforation; Pancreatitis |  |  | 1 | Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF TWO DEVICES MENTIONED IN THE LITERATURE ARTICLE. REFER TO MANUFACTURER REPORT # 3005099803-2024-05659 FOR THE ASSOCIATED DEVICE INFORMATION. BOSTON SCIENTIFIC BECAME AWARE OF EVENTS INVOLVING ULTRATOME SPHINCTEROTOME AND DREAMTOME SPHINCTEROTOME DEVICES THROUGH THE ARTICLE, TRAINEE INVOLVEMENT INCREASES PRECUT RATES AND DELAYS ACCESS TO THE COMMON BILE DUCT WITHOUT AN INCREASE IN PROCEDURE-RELATED ADVERSE EVENTS: A BRAVE NEW WORLD OF ERCP TRAINING? BY THEODOR VOIOSU MD, PHD, ET AL. PER THE ARTICLE, PATIENTS UNDERWENT AN ENDOSCOPIC RETROGRADE. CHOLANGIOPANCREATOGRAPHY (ERCP) PROCEDURE, AND THE FOLLOWING OCCURRED: SIX PATIENT DEATHS OCCURRED IN THE 30-DAY INTERVAL FOLLOWING THE PROCEDURE, THREE OF THESE DEATHS WERE RELATED TO PROCEDURAL COMPLICATIONS (TWO CASES OF SEVERE PANCREATITIS AND ONE PERFORATION). PLEASE SEE THE REFERENCED ARTICLE FOR FULL DETAILS. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENT DEATH WAS NOT REPORTED. THE RETROSPECTIVE STUDY DATE (B)(6) 2017, IS USED FOR THE ESTIMATED DATE AS IT IS THE ONLY DATE LISTED IN THE ARTICLE. BLOCK B3: THE EXACT DATE OF EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY DATE (B)(6) 2017, IS USED FOR THE ESTIMATED DATE AS IT IS THE ONLY DATE LISTED IN THE ARTICLE. BLOCK D4, H4: THE LITERATURE ARTICLE DID NOT PROVIDE THE SUSPECT DEVICE UPN AND LOT NUMBER; THEREFORE, THE LOT EXPIRATION AND DEVICE MANUFACTURE DATES ARE UNKNOWN. BLOCK G2: LITERATURE SOURCE JOURNAL ARTICLE: VOIOSU T, ET AL. "TRAINEE INVOLVEMENT INCREASES PRECUT RATES AND DELAYS ACCESS TO THE COMMON BILE DUCT WITHOUT AN INCREASE IN PROCEDURE-RELATED ADVERSE EVENTS: A BRAVE NEW WORLD OF ERCP TRAINING?" ROM. J. INTERN. MED., 2018, DOI: 10.1515/RJIM-2017-0041. BLOCK H6: IMDRF PATIENT CODE E1021 CAPTURES THE REPORTABLE EVENT OF PANCREATITIS. IMDRF PATIENT CODE E2114 CAPTURES THE REPORTABLE EVENT OF PERFORATION. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056120&pc=BZD | 2518422-2025-000089 | 2024-06-24 04:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056153&pc=DSQ | 3007042319-2025-00006 | 2024-10-13 04:00 | Death | HEARTWARE, INC. | 2025-01-02 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT DIED. Manufacturer Narrative: THIS EVENT IS A DUPLICATE OF FDA REPORT NUMBER (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056156&pc=DSQ | 3007042319-2025-00007 | 2022-10-01 04:00 | Death | HEARTWARE, INC. | 2025-01-02 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Purulent Discharge; Erythema; Fever; Hemorrhage/Bleeding; Unspecified Infection; Pain; Drug Resistant Bacterial Infection | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SURGICAL TREATMENT FOR LEFT VENTRICULAR ASSIST DEVICE (VAD) INFECTION. THE AUTHORS DESCRIBED PATIENTS WHO DEVELOPED A VAD SPECIFIC INFECTION (VSI). VSI WAS DEFINED AS AN INFECTION WITH CLINICAL SYMPTOMS, SUCH AS FEVER, LOCALIZED PAIN, ERYTHEMA, AND PURULENT DRAINAGE, SUPPORTED BY DIAGNOSTIC FINDINGS, SUCH AS POSITIVE WOUND CULTURE AND IMAGING EVIDENCE OF VSI. WOUND AND BLOOD CULTURES WERE TAKEN, AND COMPUTERIZED TOMOGRAPHY (CT) SCANS OF THE CHEST, ABDOMEN, AND PELVIS WERE PERFORMED. ALL VSI WERE INITIALLY TREATED WITH BROAD-SPECTRUM INTRAVENOUS (IV) ANTIBIOTICS. THERE WERE PATIENTS WITH A HISTORY OF PUMP EXCHANGE FOR UNKNOWN REASONS, HISTORY OF EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AFTER VAD IMPLANT, OR BLEEDING WHICH REQUIRED REOPERATION AFTER VAD IMPLANT. INFECTIONS INCLUDED DRIVELINE, PUMP, AND OUTFLOW GRAFT INFECTIONS. GRAM-POSITIVE SPECIES WERE FOUND AND THERE WERE POSITIVE BLOOD CULTURES. THE MOST HIGHLY REPRESENTED ORGANISMS INCLUDED METHICILLIN-SENSITIVE STAPHYLOCOCCUS AUREUS (MSSA), CORYNEBACTERIUM SPECIES, METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA), AND STAPHYLOCOCCUS EPIDERMIDIS. PATIENTS WERE TREATED WITH PUMP REMOVAL OR TRANSPLANTATION; IF THAT WAS NOT AN OPTION, THEN THEY WERE TREATED WITH MEDICAL MANAGEMENT, SURGICAL DEBRIDEMENT, OR FLAP RECONSTRUCTION. ONE YEAR POST VSI, THERE WERE PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THERE WERE OTHER PATIENTS WITHIN A YEAR OF INITIAL VSI WHO WERE READMITTED WITH RECURRENT INFECTION, SECONDARY INFECTION WITH A NEW ORGANISM, OR INCREASED ANTIMICROBIAL RESISTANCE OF THE PRIMARY PATHOGEN. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/56 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ESCALATING SURGICAL TREATMENT FOR LEFT VENTRICULAR ASSIST DEVICE INFECTION AND EXPECTED MORTALITY: CLINICAL RISK PREDICTION SCORE. JOURNAL OF THE AMERICAN COLLEGE OF SURGEONS. 2024; 239:263¿275. DOI: 10.1097/XCS.0000000000001096. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ OUTFLOW GRAFT D4: MODEL #: UNKNOWN/ CATALOG #: UNKNOWN/ EXPIRATION DATE: UNK / SERIAL OR LOT#: UNKNOWN D10: NO H4: MFG DATE: UNK H5: YES D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056482&pc=MCX | 3004742232-2025-00002 | 2024-12-10 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-01-02 05:00 | MCX | DIAMONDBACK CORONARY | Material Separation | Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED IN A CORONARY VESSEL. AFTER THREE LOW SPEED TREATMENTS AND THREE HIGH SPEED TREATMENT, THE VIPERWIRE FRACTURED AND CAUSED A PERFORATION IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE OAD CROWN DID NOT MAKE CONTACT WITH THE VIPERWIRE SPRING TIP. A STENT WAS IMPLANTED TO TREAT THE PERFORATION. THE PATIENT WAS STABLE FROM CIRCULATION FOR 20 MINUTES BEFORE REQUIRING RESUSCITATION. IT WAS INDICATED BLOOD WAS RUNNING OUT INTO THE VENTRICLE WHICH LED TO THE PATIENT'S TEMPORARY STABILITY. A PACEMAKER WAS ATTEMPTED TO BE IMPLANTED. THE PATIENT EXPIRED. IT WAS INDICATED 2.5 CENTIMETERS OF THE GUIDE WIRE SPRING TIP REMAINED IN-VIVO AND 15 CENTIMETERS OF THE REST OF THE GUIDE WIRE REMAINED IN-VIVO.Event Description: IT WAS REPORTED A DIAMONDBACK CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED IN A CORONARY VESSEL. AFTER THREE LOW SPEED TREATMENTS AND THREE HIGH SPEED TREATMENT, THE VIPERWIRE FRACTURED AND CAUSED A PERFORATION IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY. THE OAD CROWN DID NOT MAKE CONTACT WITH THE VIPERWIRE SPRING TIP. A STENT WAS IMPLANTED TO TREAT THE PERFORATION. THE PATIENT WAS STABLE FROM CIRCULATION FOR 20 MINUTES BEFORE REQUIRING RESUSCITATION. IT WAS INDICATED BLOOD WAS RUNNING OUT INTO THE VENTRICLE WHICH LED TO THE PATIENT'S TEMPORARY STABILITY. A PACEMAKER WAS ATTEMPTED TO BE IMPLANTED. THE PATIENT EXPIRED. IT WAS INDICATED 2.5 CENTIMETERS OF THE GUIDE WIRE SPRING TIP REMAINED IN-VIVO AND 15 CENTIMETERS OF THE REST OF THE GUIDE WIRE REMAINED IN-VIVO. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED MATERIAL SEPARATION THAT LEAD TO THE PERFORATION OF VESSELS, UNEXPECTED MEDICAL INTERVENTION, AND EVENTUAL PATIENT DEATH APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. DUE TO THE VERY LIMITED INFORMATION AND DETAILS PROVIDED, IT CAN BE STATED THAT THE MAIN CAUSE OF DEATH WAS DUE TO PERFORATION OF AN UNKNOWN VESSEL. THE VIPERWIRE MOST LIKELY CONTRIBUTED TO THE PERFORATION IN THAT IT FRACTURED IN THE VESSEL AND PARTS WERE LEFT IN THE VESSEL WHICH HAVE ALSO CAUSED THE PATIENT TO BE HEMODYNAMICALLY UNSTABLE. IT IS UNKNOWN IF THERE WAS ANY EXCESSIVE FORCE OR MANIPULATION THAT COULD HAVE LED TO THE FRACTURE IN THIS CASE, IT IS LIKELY THAT THE GUIDE WIRE FRACTURE IS A RESULT OF USE TECHNIQUES EMPLOYED OR DEVICE PLACEMENT; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. PRODUCT PERFORMANCE ENGINEERING REVIEWED THE INCIDENT INFORMATION; HOWEVER, THE PRODUCT WAS NOT RETURNED TO ABBOTT VASCULAR FOR ANALYSIS. FACTORS THAT MAY CONTRIBUTE TO THE REPORTED MATERIAL SEPARATION THAT LEAD TO THE PERFORATION OF VESSELS, UNEXPECTED MEDICAL INTERVENTION, AND EVENTUAL PATIENT DEATH, INCLUDE, BUT ARE NOT LIMITED TO PATIENT ANATOMICAL MORPHOLOGY, PATIENT DISEASE STATE, OR DEVICE PLACEMENT OR USE TECHNIQUES EMPLOYED. DUE TO THE VERY LIMITED INFORMATION AND DETAILS PROVIDED, IT CAN BE STATED THAT THE MAIN CAUSE OF DEATH WAS DUE TO PERFORATION OF AN UNKNOWN VESSEL. THE VIPERWIRE MOST LIKELY CONTRIBUTED TO THE PERFORATION IN THAT IT FRACTURED IN THE VESSEL AND PARTS WERE LEFT IN THE VESSEL WHICH HAVE ALSO CAUSED THE PATIENT TO BE HEMODYNAMICALLY UNSTABLE. IT IS UNKNOWN IF THERE WAS ANY EXCESSIVE FORCE OR MANIPULATION THAT COULD HAVE LED TO THE FRACTURE. THE REPORTED PATIENT EFFECT OF PERFORATION OF VESSELS AND CARDIAC DEATH ARE LISTED IN THE DIAMONDBACK 360¿ CORONARY ORBITAL ATHERECTOMY SYSTEM INSTRUCTIONS FOR USE AS A KNOWN POTENTIAL PATIENT EFFECT ASSOCIATED WITH THE USE OF THE DEVICE. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056485&pc=BZD | 2518422-2025-000090 | 2024-06-14 04:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056877&pc=NZU | MW5164149 | 2024-08-24 04:00 | Death | C. R. BARD, INC. | 2025-01-02 05:00 | NZU | PUREWICK INCONTENANCE PRODUCT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Lethargy; Cancer; Fecal Incontinence |  |  | 1 | Event Description: WELL ADMITTED TO (B)(6). THE MEDICAL STAFF USED THE PURE WICK AT HOME INCONTINENCE PRODUCT ON MY MOTHER, (B)(6) WITHOUT CONCERN FOR HARM OR INJURY TO HER. AFTER LETHARGIC AND BEDRIDDEN FOR FOUR DAYS WHEN THE PRODUCT RULES REMOVED, IT CAUSED SEVERE TRAUMA TO THE SKIN. IT CONTINUED FOR YEARS THROUGH WOUND EXPERT TREATMENT AND EVENTUALLY DEVELOPED INTO CANCER. FOLLOWED BY MANY MONTHS OF TREATMENT, FETAL INCONTINENCE AND BLEEDING TO THE POINT OF ALMOST HEMORRHAGING. Event Description: ADDITIONAL INFORMATION RECEIVED ON 1/6/2024 FOR REPORT MW5164149 TO UPDATE THE MANUFACTURER TO C.R. BARD, INC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056980&pc=BZD | 2518422-2025-00026 | 2024-11-19 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER INQUIRY ABOUT THE RETURN OF A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057017&pc=QBJ | MW5164151 |  | Death | DEXCOM, INC. | 2025-01-02 05:00 | QBJ | DEXCOM SENSOR |  | Insufficient Information |  |  | 1 | Event Description: PATIENT PASSED AWAY - REPORTED BY HCP WHO DID NOT CONSENT TO FOLLOW UP - (B)(4). INSERTED (B)(6) 2024. ACTIVATED ON (B)(6) 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057026&pc=MGQ | MW5164152 |  | Death | SOLEVA PHARMA LLC. | 2025-01-02 05:00 | MGQ | MUGARD ORAL RINSE 240ML/BTL |  | Insufficient Information |  |  | 1 | Event Description: EVENT: THE PATIENT WAS MARKED DECEASED IN (B)(6) RETAIL SYSTEM WHICH FLOWED TO (B)(6) SPECIALTY SYSTEM, AND WE ARE REPORTING PER CONTRACTUAL REQUIREMENTS AND STANDARD OPERATING PROCEDURE WITHIN ONE BUSINESS DAY; HOWEVER, NO CAUSE OF DEATH KNOWN. DATE OF DEATH UNKNOWN AND NO FURTHER INFORMATION AVAILABLE. FREQ: SWISH AND SWALLOW 5ML AS DIRECTED 4-6 TIMES PER DAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057075&pc=BZD | 2518422-2025-00024 | 2024-11-19 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON A COMPENSATION CHECK ISSUED TO HER HUSBAND WHO IS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057251&pc=OZD | 1220648-2025-25282 | 2024-12-12 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P140003 |  | 1 | Event Description: THE US COMPLAINANT HAD A CP PUMP PLACED TO SUPPORT A PATIENT. AFTER A DAY OF SUPPORT, WHICH INCLUDED AND ACCESS SITE BLEED, THE TEAM MADE CHOICE TO EXPLANT AND ESCALATE TO A 5.5 PUMP. AFTER EXPLANT AND GROIN CLOSURE THE TEAM DISCOVERED A VESSEL DAMAGE THAT PROMPTED THE TEAM TO INTERVENE AND REPAIR THE VESSEL.Event Description: IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: ¿BE CAREFUL NOT TO PULL ON THE IMPELLA CATHETER WHEN TRANSFERRING A PATIENT FROM ONE BED TO ANOTHER.¿ ¿POTENTIAL ADVERSE EVENTS (UNITED STATES): ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND VASCULAR INJURY¿. Manufacturer Narrative: THE INVESTIGATION FOR THE VASCULAR DAMAGE HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION, AND THE DATA LOG WAS NOT REQUIRED FOR THIS CASE. ADDITIONALLY, THE LEFT FEMORAL ARTERY WAS REPAIRED AFTER IT WAS DISSECTED DURING THE REMOVAL OF THE PUMP. THE CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. CORRECTIONS TO THE INITIAL MFR REPORT: B2-SELECTED DEATH AND ADDED DATE OF DEATH. B5-PATIENT OUTCOME AND MSO REVIEW. G1 MDR REPORTING CONTACT EMAIL. H1-CHANGED TO DEATH. H6-IMPACT CODE 1802.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057340&pc=BZD | 2518422-2025-00032 | 2024-11-07 05:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | DREAMSTATION CPAP PRO | Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION CPAP PRO DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057363&pc=OZD | 1220648-2025-25284 | 2024-12-09 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT IN JAPAN REPORTED THAT THE PATIENT ON DAY ONE OF IMPELLA CP SUPPORT DEVELOPED SEVERE LOWER LIMB ISCHEMIA. THE IMPELLA WAS REMOVED. IT WAS ALSO REPORTED THAT PATIENT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. B.2 REVISED AS REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25284. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25284 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.1 REVISED REPORTING CONTACT FAX NUMBER AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25284. H.6 COMPONENT CODE 925 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25284 AND A NEW CODE WAS ADDED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057402&pc=DSQ | 2916596-2024-52838 | 2018-08-29 04:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE CUSTOMER STATED THAT NO MORE INFORMATION WOULD BE PROVIDED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED AND THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057549&pc=DSQ | 2916596-2024-52854 | 2024-03-20 04:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE DEATH WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND THE PUMP WAS NOT EXPLANTED. NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057900&pc=LWS | 2649622-2025-00091 | 2024-12-26 05:00 | Death | MPRI | 2025-01-02 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Signal Artifact/Noise; Under-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A NOISE WARNING AND POSSIBLE UNDERSENSING ON THE RIGHT VENTRICULAR (RV) LEAD. IT WAS NOTED THAT EVEN WITH THE POSSIBILITY OF UNDERSENSING, VENTRICULAR FIBRILLATION (VF) WAS DETECTED AND APPROPRIATELY TREATED THE ARRHYTHMIA. THE SHOCK TERMINATED THE ARRHYTHMIA, HOWEVER, POST SHOCK THE PATIENT WAS NOTED TO BE IN A 100% PACED RHYTHM WITH NO MECHANICAL RESPONSE. THE PATIENT SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: DTPA2QQ CRT-D IMPLANTED (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21057902&pc=BZD | 2518422-2025-000094 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-02 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Septic Shock; Cancer; Unspecified Mental, Emotional or Behavioural Problem; Unspecified Heart Problem; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PANCREATIC AND PROSTATE CANCER, CORONARY ARTERY DISEASE, ACUTE HYPOXIC RESPIRATORY FAILURE WITH ASSOCIATED ALTERED MENTAL STATUS, SEPTIC SHOCK AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21059431&pc=NPT | 2025587-2025-00028 | 2024-12-18 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-02 05:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PLACEMENT OF A TRANSCATHETER AORTIC VALVE IN A PATIENT WITH EFFACED SINUSES AND A LOW LEFT CORONARY ARTERY, A DROP IN BLOOD PRESSURE WAS NOTED AS THE CATHETER TO CANNULATE THE LEFT MAIN ARTERY CROSSED THE VALVE. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. SUBSEUQENTLY, THE EVOLUT FX 26MM VALVE WAS DEPLOYED. THE VALVE APPEARED CONSTRAINED, HOWEVER, AND IT WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. A NON-MEDTRONIC BALLOON WAS USED TO PERFORM AN BALLOON AORTIC VALVULOPLASTY. A SECOND VALVE, THE EVOLUT 26MM FX+, WAS INSERTED INTO THE PATIENT USING A NEW DCS AND SUCCESSFULLY DEPLOYED. A STENT WAS ALSO DEPLOYED IN THE LEFT MAIN ARTERY. DESPITE THESE INTERVENTIONS, THE PATIENT CONTINUED TO EXPERIENCE LOW BLOOD PRESSURE. CHEST COMPRESSIONS WERE PERFORMED AND A BALLOON PUMP WAS INSERTED; HOWEVER, THE PATIENT DID NOT RECOVER AND SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS NOTED AS POSSIBLY RELATED TO ANESTHESIA.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOTED THAT THE PATIENT¿S EJECTION FRACTION WAS 55%, BUT THE PATIENT STARTED DETERIORATING AT INDUCTION. PER THE PHYSICIAN, THE VALVE AND THE DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-26 (J285559); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21059702&pc=NPT | 2025587-2025-00030 | 2024-12-18 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-02 05:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PLACEMENT OF A TRANSCATHETER AORTIC VALVE IN A PATIENT WITH EFFACED SINUSES AND A LOW LEFT CORONARY ARTERY, A DROP IN BLOOD PRESSURE WAS NOTED AS THE CATHETER TO CANNULATE THE LEFT MAIN ARTERY CROSSED THE VALVE. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. SUBSEUQENTLY, THE EVOLUT FX 26MM VALVE WAS DEPLOYED. THE VALVE APPEARED CONSTRAINED, HOWEVER, AND IT WAS RECAPTURED AND WITHDRAWN FROM THE PATIENT. A NON-MEDTRONIC BALLOON WAS USED TO PERFORM AN BALLOON AORTIC VALVULOPLASTY. A SECOND VALVE, THE EVOLUT 26MM FX+, WAS INSERTED INTO THE PATIENT USING A NEW DCS AND SUCCESSFULLY DEPLOYED. A STENT WAS ALSO DEPLOYED IN THE LEFT MAIN ARTERY. DESPITE THESE INTERVENTIONS, THE PATIENT CONTINUED TO EXPERIENCE LOW BLOOD PRESSURE. CHEST COMPRESSIONS WERE PERFORMED AND A BALLOON PUMP WAS INSERTED; HOWEVER, THE PATIENT DID NOT RECOVER AND SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS NOTED AS POSSIBLY RELATED TO ANESTHESIA.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOTED THAT THE PATIENT¿S EJECTION FRACTION WAS 55%, BUT THE PATIENT STARTED DETERIORATING AT INDUCTION. PER THE PHYSICIAN, THE VALVE AND THE DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS. Manufacturer Narrative: UPDATED DATA: B5. CORRECTED DATA: D1. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060103&pc=DSQ | 2916596-2024-52721 | 2024-12-07 05:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL HEMORRHAGE. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS REQUESTED BUT NOT YET PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060230&pc=DSQ | 2916596-2024-52687 | 2024-12-07 05:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 FROM CONGESTIVE HEART FAILURE. Manufacturer Narrative: A4: PATIENT WEIGHT WAS REQUESTED BUT NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060427&pc=NIQ | 2124215-2024-79201 | 2024-12-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-02 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY. IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA, SILENT ISCHEMIA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO DISTAL LAD WITH 80% STENOSIS AND WAS 40 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF 2.50 MM X 32 MM OVERLAPPED WITH 2.75 MM X 16 MM SYNERGY STENT SYSTEMS WITH RESIDUAL STENOSIS OF 0%. POST-DILATATION WAS NOT PERFORMED. TWO DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND TICAGRELOR. IN (B)(6) 2023, THE SUBJECT DIED, AND THE PRIMARY CAUSE OF DEATH IS UNKNOWN. AT THE TIME OF EVENT, THE SUBJECT WAS ON CLOPIDOGREL AND OTHER ANTIPLATELET MEDICATION. THERE WAS NO ACTION WAS TAKEN TO TREAT THE EVENT, AND IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B2: DATE OF DEATH: USED (B)(6) 2023, AS THE EXACT DEATH DATE WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060433&pc=NIQ | 2124215-2024-79203 | 2023-12-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-02 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY. IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA, SILENT ISCHEMIA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO DISTAL LAD WITH 80% STENOSIS AND WAS 40 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF 2.50 MM X 32 MM OVERLAPPED WITH 2.75 MM X 16 MM SYNERGY STENT SYSTEMS WITH RESIDUAL STENOSIS OF 0%. POST-DILATATION WAS NOT PERFORMED. TWO DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND TICAGRELOR. IN (B)(6) 2023, THE SUBJECT DIED, AND THE PRIMARY CAUSE OF DEATH IS UNKNOWN. AT THE TIME OF EVENT, THE SUBJECT WAS ON CLOPIDOGREL AND OTHER ANTIPLATELET MEDICATION. THERE WAS NO ACTION WAS TAKEN TO TREAT THE EVENT, AND IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B2: DATE OF DEATH: USED (B)(6) 2023, AS THE EXACT DEATH DATE WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060477&pc=NIK | 2017865-2025-00050 | 2024-12-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-02 05:00 | NIK | QUADRA ASSURA MP ICD | Failure to Capture | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR REMOTE FOLLOW UP VIA MERLIN.NET. A REVIEW OF THE TRANSMISSION REVEALED LOSS OF VENTRICULAR CAPTURE EXHIBITED BY THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). IT WAS LATER REVEALED THAT THE PATIENT EXPIRED DUE TO UNKNOWN CAUSE. THE DEATH WAS REPORTED WITHOUT CLEAR INFORMATION AS TO WHETHER THE DEATH WAS DEVICE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060743&pc=NRY | 2029214-2025-00004 | 2024-05-06 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-02 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem; Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING "EFFICACY AND SAFETY OF A NEW DELIVERY ASSIST CATHETER WITH A FLEXIBLE, SPINDLE-SHAPED SHAFT IN MECHANICAL THROMBECTOMY". THE PURPOSE OF THIS STUDY WAS TO EVALUATE WHETHER NAVIGATING AN ASPIRATION CATHETER TOWARD THE OCCLUSION SITE COULD BE DONE SAFELY, STEADILY, AND RAPIDLY USING A DELIVERY ASSIST INNER CATHETER. THE TIME FRAME OF THIS STUDY WAS JANUARY 2018 AND MAY 2022. DEVICE MANUFACTURERS: MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 68 ASPIRATION CATHETER AND SOLITAIRE STENT RETRIEVER SIX DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSE OF DEATH WAS NOT NOTED AMONG PATIENT ADVERSE EVENTS INCLUDED: -1 ARTERY DISSECTION -8 SYMPTOMATIC INTRACRANIAL HEMORRHAGES -MEDIAN MRS AT 90 DAYS WAS 2, RANGE 1-4 NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: G2: CITATION: AUTHORS: TAKESHITA, S., NII, K., TSUGAWA, J., ISHII, A., FUKUMOTO, H., HANADA, H., INOUE, R., SAKAMOTO, K., & HIGASHI, T. EFFICACY AND SAFETY OF A NEW DELIVERY ASSIST CATHETER WITH A FLEXIBLE, SPINDLE-SHAPED SHAFT IN MECHANICAL THROMBECTOMY. WORLD NEUROSURGERY 187, E997¿E1003. 2024. DOI:10.1016/J.WNEU.2024.05.025 EARLIEST DATE OF ACCEPTANCE USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060745&pc=DQY | 2029214-2025-00005 | 2024-05-06 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-02 05:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem; Insufficient Information | K180715 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING "EFFICACY AND SAFETY OF A NEW DELIVERY ASSIST CATHETER WITH A FLEXIBLE, SPINDLE-SHAPED SHAFT IN MECHANICAL THROMBECTOMY". THE PURPOSE OF THIS STUDY WAS TO EVALUATE WHETHER NAVIGATING AN ASPIRATION CATHETER TOWARD THE OCCLUSION SITE COULD BE DONE SAFELY, STEADILY, AND RAPIDLY USING A DELIVERY ASSIST INNER CATHETER. THE TIME FRAME OF THIS STUDY WAS JANUARY 2018 AND MAY 2022. DEVICE MANUFACTURERS: MULTIPLE MANUFACTURER¿S DEVICES WERE USED INTHE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 68 ASPIRATION CATHETER AND SOLITAIRE STENT RETRIEVER SIX DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSE OF DEATH WAS NOT NOTED AMONG PATIENT ADVERSE EVENTS INCLUDED: -1 ARTERY DISSECTION -8 SYMPTOMATIC INTRACRANIAL HEMORRHAGES -MEDIAN MRS AT 90 DAYS WAS 2, RANGE 1-4 NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060776&pc=DSQ | 2916596-2024-52842 | 2024-11-12 05:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT HAD A HISTORY OF BIVENTRICULAR ASSIST DEVICE (BIVAD) IMPLANTED ON (B)(6) 2023. THE PATIENT HAD A HISTORY OF RIGHT HEART FAILURE PRIOR TO IMPLANT WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE LVAD DID NOT CAUSE OF CONTRIBUTE TO THE PATIENT'S REQUIREMENT FOR RIGHT-SIDED SUPPORT. THE PATIENT HAD 0 FLOW ON THE RIGHT VENTRICULAR ASSIST DEVICE (RVAD) AND THE CAUSE WAS NOT DETERMINED. THE SITE DISCONNECTED THE DRIVELINE. THE ZERO FLOW EVENT DID NOT RESOLVE, HOWEVER THE PATIENT WAS HEMODYNAMICALLY STABLE AND DID NOT EXPERIENCE ANY SYMPTOMS AS A RESULT OF THE EVENT. FOLLOWING REVIEW OF THE SUBMITTED LOG FILES, IT APPEARED THERE WERE ONGOING LOW FLOW EVENTS WITH FLOWS NOTED AT ZERO. THE DATA CAPTURE WAS SHORTENED TO ABOUT 5 HOURS DUE TO THE AMOUNT OF INCOMING DATA SO THE BEGINNING OF THE EVENTS WAS NOT AVAILABLE. THERE APPEARED TO BE A DROP IN FLOW AND PUMP POWER ON THE PERIODIC LOG FILE. THE PULSATILITY INDEX (PI) VALUES WERE NON-VARIABLE AND SETTLED INTO THE 2-3 RANGE. THE PUMP POWERS APPEARED ON THE LOW SIDE IN THE 2 W RANGE. IT WAS LATER REPORTED THAT THE PATIENT WAS DECLARED DO NOT RESUSCITATE (DNR) ON PALLIATIVE CARE DUE TO COMORBIDITIES AND PASSED AWAY.Event Description: IT WAS REPORTED THAT THE SITE DISCONNECTED THE DRIVELINE BECAUSE NO FLOW RETURNED TO THE PUMP. THE PATIENT TOLERATED ONLY LVAD SUPPORT AFTER THE DISCONNECT UNTIL THE WITHDRAWAL OF CARE. Manufacturer Narrative: SECTION B3 - DATE OF DEATH UNKNOWN NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060805&pc=DSQ | 2916596-2024-52793 | 2024-12-11 05:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD CARDIAC ARREST AND RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED. LOG FILES WERE SENT FOR REVIEW. THE LOG FILE CAPTURED A SERIES OF LOW FLOW EVENTS BETWEEN 12:17 AM AND 12:19 AM (B)(6) 2024. ON (B)(6) 2024 THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST.Event Description: ON (B)(6) 2024 THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS IN BED WHEN ALARMS WERE HEARD. NURSES FOUND PATIENT WITH LOW FLOWS AND CPR WAS INITIATED. RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS OBTAINED BUT THE PATIENT WAS LATER DETERMINED TO BE BRAIN DEAD. THE FAMILY WITHDREW CARE AND THE PATIENT PASSED AWAY. THE DEVICE REPORTEDLY OPERATED AS EXPECTED AND THE DEATH WAS NOT CONSIDERED DEVICE RELATED Manufacturer Narrative: B5: ADDITIONAL EVENT INFORMATION. H6: HEALTH EFFECT -IMPACT, HEALTH EFFECT-CLINICAL, MEDICAL DEVICE PROBLEM CODES, UPDATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 1802 - DEATH REMOVED. MANUFACTURER'S INVESTIGATION CONCLUSION: ANALYSIS OF THE SUBMITTED LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED CARDIAC ARREST AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 27NOV2024 THROUGH 11DEC2024. LOW FLOW HAZARD ALARMS WERE CAPTURED ON 11DEC2024. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21060831&pc=DSQ | 2916596-2024-52723 | 2024-09-28 04:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock; Multiple Organ Failure; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A CHRONIC DRIVELINE INFECTION WITH EXTRA RESISTING PSEUDOMONAS AERUGINOSA. THEY TESTED POSITIVE FOR COVID-19 AT THE END OF (B)(6) 2024. THE PATIENT WAS HOSPITALIZED AND RECEIVED THERAPY FOR COVID-19. POST-COVID THE INFECTION MIGRATED INTO THE BLOODSTREAM AND THE PATIENT DEVELOPED SEPTIC SHOCK AND MULTISYSTEM ORGAN FAILURE. THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND WAS DUE TO THE PATIENT CONDITION.Event Description: IT WAS FURTHER COMMUNICATED ON 13JAN2025 THAT THE PATIENT WAS ADMITTED ON (B)(6) 2024. Manufacturer Narrative: B3, B5 NARRATIVE INFORMATION RE EVENT DATE - DATE OF EVENT UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP WAS NOT RETURNED. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21061167&pc=DSQ | 2916596-2024-52739 | 2024-12-11 05:00 | Death | THORATEC CORPORATION | 2025-01-02 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. PRELIMINARY CASE OF DEATH WAS SEPSIS. AUTOPSY RESULTS WERE PENDING. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21061304&pc=OZD | 1220648-2025-25289 | 2024-12-14 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE US COMPLAINANT HAD A CP PUMP SUPPORT INITIATED FOR A PATIENT ADMITTED IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE PUMP WAS SUPPORTING WHEN THE TEAM OBSERVED SIGNS OF HEMOLYSIS. THE PATIENT EXPIRED WITH CARE WITHDRAWAL PRIOR TO ANY INTERVENTION FOR THE HEMOLYSIS. THE SUPPORT HAD BEEN FOR 37 HOURS.Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT PRESENTED SIGNS OF HEMOLYSIS. THE PATIENT HAD DECREASED URINE PRODUCTION AND THE COLOR WAS NOTED TO BE CHANGED. THE TEAM DISCUSSED CARE ESCALATION AND SWAN PLACEMENT. SWAN PLACEMENT WOULD HAVE ALLOWED FOR BETTER HEMODYNAMIC MONITORING. NEITHER OF THESE MEASURES (ESCALATE/SWAN) WERE TAKEN. THE PUMP WAS REPOSITIONED AS THE PUMP WAS VERY CLOSE TO THE MITRAL. NO OTHER INTERVENTIONS WERE NOTED FOR THE HEMOLYSIS. IT WAS NOTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED AFTER 37 HOURS OF SUPPORT AND WITHOUT INTERVENTION TAKEN FOR THE HEMOLYSIS. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. DEVICE HISTORY RECORD REVIEW CONFIRMED THE PUMP PASSED ALL POST STERILE INSPECTION CHECKS. THE DEVICE WAS NOT RECEIVED FOR EVALUATION. DATA LOGS WERE REVIEWED. NO MOTOR CURRENT SPIKE WAS OBSERVED. IMPELLA POSITION IN VENTRICLE ALARMED AT P-2 AND END OF THE CASE. NO CLEAR EVIDENCE WAS OBSERVED FROM REVIEWING THE LOGS. AS CLINICAL DETAILS STATED THAT THE DEVICE WAS VERY CLOSE TO THE MITRAL VALVE AND LEFT SHALLOW IN ATTEMPT TO FREE PUMP FROM PAPILLARY MUSCLE, THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY IMPROPER POSITIONING DUE TO IMPROPER TECHNIQUE. CORRECTED INFORMATION PROVIDED IN B5 AND H10. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, AN INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE HEMOLYSIS ISSUE MOST LIKELY WAS IMPROPER POSITIONING DUE TO IMPROPER TECHNIQUE. G1 REPORTING CONTACT EMAIL WAS INADVERTENTLY REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT 1220648-2025-25289. G1 MANUFACTURING SITE FAX NUMBER WAS INADVERTENTLY REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT 1220648-2025-25289.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21062279&pc=NPT | 2135147-2025-00032 | 2018-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-01-02 05:00 | NPT | PORTICO TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P190023 |  | 1 | Event Description: THE ARTICLE, "PREVALENCE AND PROGRESSION OF LV DYSFUNCTION AND DESYNCHRONY IN PATIENTS WITH NEW-ONSET LBBB POST TAVR", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO ASSESS THE PREVALENCE AND PROGRESSION OF LEFT VENTRICULAR (LV) DYSFUNCTION AND MECHANICAL DESYNCHRONY IN PATIENTS WITH NEW-ONSET LEFT BUNDLE BRANCH BLOCK (N-LBBB) AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). DEVICES INCLUDED PORTICO, EDWARDS SAPIEN 3, AND ALLEGRA VALVES. N-LBBB AFTER TAVR RESULTS IN AN IMMEDIATE REDUCTION OF CARDIAC FUNCTION, IN SPITE OF ONLY MILD DESYNCHRONY. WHEN LBBB PERSISTS, PATIENTS WITH BETTER CARDIAC FUNCTION BEFORE TAVR ARE MORE LIKELY TO HAVE LBBB-INDUCED LV DYSFUNCTION AFTER TAVR. [THE PRIMARY AND CORRESPONDING AUTHOR WAS ANDREI MARGULESCU, DEPARTMENT OF CARDIOLOGY, MORRISTON REGIONAL CARDIAC CENTRE, MORRISTON HOSPITAL, HEOL MAES EGLWYS, SWANSEA SA6 6NL, UK, WITH CORRESPONDING EMAIL: ANDREI.MARGULESCU@WALES.NHS.UK]. THE TIME FRAME OF THE STUDY WAS FROM OCTOBER 2018 TO SEPTEMBER 2021. A TOTAL OF 55 PATIENTS WERE INCLUDED IN THE FINAL ANALYSIS, OF WHICH 33 (60%) RECEIVED AN ABBOTT DEVICE. FOR PATIENTS IN N-LBBB GROUP, THE AVERAGE AGE WAS 83.2 YEARS AND THE MAJORITY GENDER WAS FEMALE. IN THE C-LBBB GROUP, THE AVERAGE AGE WAS 82.7 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ISCHEMIC HEART DISEASE, CORONARY ARTERY BYPASS GRAFT, STROKE, DIABETES MELLITUS, HYPERTENSION, CHRONIC KIDNEY DISEASE, HEART BLOCK. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF PORTICO VALVES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ISCHEMIC HEART DISEASE, CORONARY ARTERY BYPASS GRAFT, STROKE, DIABETES MELLITUS, HYPERTENSION, CHRONIC KIDNEY DISEASE, AND HEART BLOCK. SOME OF THE COMPLICATIONS REPORTED WERE DEATH AND HEART BLOCK. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON ALL AVAILABLE INFORMATION, CAUSES FOR THE REPORTED HEART BLOCK AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE: ARTICLE TITLE "PREVALENCE AND PROGRESSION OF LV DYSFUNCTION AND DYSSYNCHRONY IN PATIENTS WITH NEW-ONSET LBBB POST TAVR." Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE "PREVALENCE AND PROGRESSION OF LV DYSFUNCTION AND DYSSYNCHRONY IN PATIENTS WITH NEW-ONSET LBBB POST TAVR" INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21063029&pc=KRH | 9612164-2025-00042 |  | Death | MEDTRONIC MEXICO | 2025-01-02 05:00 | KRH | CONTOUR 3D | Insufficient Information | Insufficient Information | K101212 |  | 1 | Event Description: IT WAS REPORTED THAT A POSTOPERATIVE DEATH OCCURRED FOLLOWING THE IMPLANTATION OF A 30MM CONTOUR 3D TRICUSPID ANNULOPLASTY RING. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO ALLEGED PRODUCT ISSUES REPORTED. THE DEVICE REMAINS IN THE PATIENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21063177&pc=OZD | 1220648-2025-25290 | 2024-12-15 05:00 | Death | ABIOMED, INC. | 2025-01-02 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND HAD LOW PUMP FLOW AND WOULD ULTIMATELY EXPIRE. THE PATIENT ARRIVED AT THE EMERGENCY DEPARTMENT AND FELL WHILE AMBULATING IN THE PARKING LOT. EMERGENCY ROOM STAFF STARTED CARDIOPULMONARY RESUSCITATION (CPR) AND OBTAINED RETURN OF SPONTANEOUS CIRCULATION. THE PATIENT CODED AGAIN, ST-ELEVATION MYOCARDIAL INFARCTION WAS CALLED, AND THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB FOR PERCUTANEOUS CORONARY INTERVENTION WHICH WAS PERFORMED TO THE RIGHT CORONARY ARTERY. THE IMPELLA CP DEVICE WAS THEN PLACED VIA THE RIGHT FEMORAL ARTERY. A TEMPORARY TRANSVENOUS PACER WAS PLACED AS WELL. THE PATIENT WAS TRANSPORTED TO THE UNIT ON MCS. IT WAS NOTED THE PATIENT CONTINUED TO CODE MULTIPLE TIMES FROM ARRIVAL TO THROUGH THE ICU ADMISSION. IMPELLA WAS TURNED TO SUPPORT LEVEL P-0 FOR APPROXIMATELY ONE MINUTE AS THE TEAM MADE THE DECISION TO PRONOUNCE THE PATIENT. THE PATIENT BEGAN SPONTANEOUSLY BREATHING, AND THE IMPELLA CP WAS RESTARTED WITH CPR EFFORTS. THE NURSE STATED THE WAS STILL BEING CODED AND SUCTION ALARMS REMAINED, POSSIBLY RELATED TO A CLOT. SUSPICION OF CLOT WAS NOTED RELATED TO SPIKE IN MOTOR CURRENT WHEN IMPELLA WAS RESTARTED. PATIENT EVENTUALLY SUCCUMBED TO CARDIOVASCULAR DEMISE; CARE WAS WITHDRAWN AFTER APPROXIMATELY NINE HOURS OF MCS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION FOR THE LOW PUMP FLOW HAS BEEN COMPLETED. DATA LOGS WERE REVIEWED WHICH CONFIRMED THE FAILURE MODE AND CLINICAL NARRATIVE. LOGS SHOWED LOW FLOW OCCURRED AFTER PUMP STARTED THAT TRIGGERED AUTO SUCTION RESPONSE AND SUCTION ALARMS. NEXT, PUMP WAS RUN WITH LOW FLOW FOR 6 HOURS AND STOPPED MANUALLY. PUMP THEN WAS RESTARTED WITH MOTOR CURRENT (MC) SPIKED FOLLOWED FLOW DROPPED TO 0.5 L/MIN. THIS EVENT REMAINED TO THE REST OF THE CASE. PRODUCT WAS NOT RETURNED FOR REVIEW. PATIENT WAS DETERIORATING THE ENTIRE CASE WITH CPR REQUIRED. DATA LOGS SHOW FLOWS ARE ON THE LOWER SIDE THE ENTIRE CASE AND TRENDING DOWNWARDS WHILE ON P-6. SUCTION ALARMS ARE CONTINUOUS THROUGHOUT. THE MC SPIKE OCCURS AT THE END OF THE CASE WITH AN ADDITIONAL DROP IN FLOW. THE OVERALL CAUSE OF THE LOW FLOWS WAS MOST LIKELY PATIENT CONDITION RELATED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21063371&pc=NIQ | 2124215-2024-79205 | 2024-11-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-02 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2020, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE DISTAL RIGHT CORONARY ARTERY (RCA) WITH 95% STENOSIS AND WAS 18 MM LONG, WITH REFERENCE VESSEL DIAMETER OF 2.75 MM. THE TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF 2.75 MM X 20 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. FIVE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2024, THE PATIENT DIED, AND THE PRIMARY CAUSE OF DEATH IS UNKNOWN. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. THERE WAS NO ACTION WAS TAKEN TO TREAT THE EVENT, AND IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR NOT. Manufacturer Narrative: B2: DATE OF DEATH: USED (B)(6) 2024, AS THE EXACT DEATH DATE WAS UNKNOWN. B3: DATE OF EVENT: USED (B)(6) 2024, AS THE EXACT DEATH DATE WAS UNKNOWN. E1: INITIAL REPORTER FACILITY NAME: (B)(6) HOSPITAL. E1: INITIAL REPORTER PHONE: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21063480&pc=DSY | 2242352-2025-00002 | 2023-05-01 04:00 | Death | INTERVASCULAR SAS | 2025-01-02 05:00 | DSY | CARDIOROOT WOVEN | Insufficient Information | Insufficient Information | K103347 |  | 1 | Event Description: IT WAS REPORTED TO INTERVASCULAR THAT DURING A CONVERSATION BETWEEN THE SALES REPRESENTATIVE AND THE HOSPITAL PHARMACIST ON DECEMBER 06, 2024, INFORMATION ABOUT A PATIENT'S DEATH WAS COMMUNICATED, PRESUMABLY FOLLOWING THE IMPLANTATION OF A CARDIOROOT PROSTHESIS IN (B)(6) 2023 (EXACT DAY NOT SPECIFIED). AT THE TIME OF THIS EMDR, THERE IS NO FURTHER INFORMATION ON THE CASE AND AN APPOINTMENT BETWEEN THE SALES REPRESENTATIVE AND THE SURGEON IS PENDING. IT SHOULD BE NOTED THAT THE DATE OF EVENT IN BLOCK B3 IS SET TO MAY 1, 2023, AS AN APPROXIMATION OF THE EVENT DATE. THE DATE WILL BE UPDATED LATER IN A FOLLOW-UP REPORT AS MORE INFORMATION WILL BECOME AVAILABLE.Event Description: COMPLAINT # (B)(4). Manufacturer Narrative: (4117) AT THE TIME OF THIS REPORT, IT IS UNKNOWN WHETHER THE PRODUCT IS AVAILABLE FOR ANALYSIS. (4118/3221) THE SERIAL NUMBER AND STERILIZATION LOT NUMBER OF THE INVOLVED PRODUCT ARE NOT AVAILABLE. THE PRODUCT IDENTIFICATION AND FURTHER INFORMATION OF THE EVENT ARE BEING SOUGHT TO INITIATE THE INVESTIGATION. (11) THE INVESTIGATION IS PENDING ADDITIONAL INFORMATION. A FOLLOW-UP REPORT WILL BE SENT UPON COMPLETION OF THE INVESTIGATION. Manufacturer Narrative: CORRECTED DATA: IN BLOCK D4, THE SUSPECTED MEDICAL DEVICE INFORMATION HAS BEEN UPDATED AFTER A THOROUGH PRODUCT TRACEABILITY REVIEW TO IDENTIFY THE INVOLVED CARDIOROOT, AS DETAILED BELOW. ADDITIONAL MANUFACTURER NARRATIVE: (4115) THE INVOLVED DEVICE WAS EXPLANTED AND DISCARDED. (4112/213) THE MANUFACTURER CONDUCTED A THOROUGH PRODUCT TRACEABILITY INVESTIGATION TO IDENTIFY THE INVOLVED CARDIOROOT WITH MODEL HEWROOT0028. THE GETINGE SALES OPERATIONS MANAGER (FRENCH SUBSIDIARY) CONFIRMED THAT BETWEEN JAN 2022 AND MAY 2023, ONLY ONE CARDIOROOT WITH MODEL HEWROOT0030, SERIAL NUMBER (B)(6) AND LOT NUMBER 21F10 WAS IMPLANTED AT THE INVOLVED HOSPITAL ON (B)(6) 2022. NOTE THAT THERE WAS NO CARDIOROOT CHARGED FOR THIS HOSPITAL IN 2023; GRAFTS ARE CHARGED TO HOSPITALS ONCE THEY ARE IMPLANTED. THEREFORE, IT IS NOT EXCLUDED THAT THE INITIAL INFORMATION REGARDING CARDIOROOT REFERENCE (HEWROOT0028) PROVIDED BY THE HOSPITAL REGARDING THE COMPLAINT IS NOT ENTIRELY ACCURATE. (4121/213) THE DEVICE HISTORY REVIEW (DHR) OF THE HYPOTHETICALLY INVOLVED CARDIOROOT HEWROOT0030 (SN (B)(6) AND LOT# 21F10) WAS PERFORMED BY THE QUALITY ASSURANCE SUPERVISOR IN ORDER TO DETECT ANY RELEVANT INFORMATION THAT MAY HAVE IMPACTED THE PRODUCT AND BE RELATED TO THE COMPLAINT. IT WAS CONCLUDED THAT NO DEVIATION WAS IDENTIFIED IN RELATION WITH THE REPORTED EVENT. (4109/213) THE REVIEW OF HISTORICAL DATA INDICATED THAT NO OTHER SIMILAR COMPLAINT WAS REPORTED FOR THE SAME STERILIZATION LOT NUMBER 21F10. (4110/213) THE OCCURRENCE RATE WAS CALCULATED FOR THE SIMILAR EVENTS REPORTED (INCLUDING THIS COMPLAINT) FOR GRAFTS AND PATCHES ON THE SAME MANUFACTURING LINE (INTERGARD / HEMAGARD) FOR 1-YEAR PERIOD FROM 01-JAN-24 TO 31-JAN-25. THE OCCURRENCE RATE IS (B)(4), WHICH IS BELOW THE ANTICIPATED OCCURRENCE RATE (MAXIMUM) OF (B)(4) IN THE PRODUCT RISK ASSESSMENT. (4112/4315) BASED ON THE LIMITED PROVIDED INFORMATION AND THE MEDICAL REVIEW, NO CONCLUSION CAN BE DRAWN ON THE EXACT ORIGIN OF THE ADVERSE EVENT AND THE INVOLVEMENT OF THE DEVICE IN THE SERIOUS INCIDENT. AS CONCLUDED IN THE MEDICAL REVIEW, DUE TO THE LACK OF ANY INFORMATION REGARDING THE PATIENT, OR THE PROCEDURE, IT IS NOT POSSIBLE TO COMPLETE A MEDICAL ASSESSMENT OF THE EVENT. MOREOVER, THE CONDUCTED INVESTIGATION ON THE HYPOTHETICALLY INVOLVED CARDIOROOT HEWROOT0030 SUGGESTS THAT THE PRODUCT WAS NOT DEFECTIVE AT THE TIME OF MANUFACTURING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21063488&pc=ODG | 3002808148-2025-00105 | 2024-01-20 05:00 | Death | SHIRAKAWA OLYMPUS CO., LTD. | 2025-01-02 05:00 | ODG | EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K093395 |  | 1 | Event Description: OLYMPUS MEDICAL SYSTEMS CORP. (OMSC) RECEIVED A LITERATURE TITLED "OUTCOMES OF ENDOSCOPIC ULTRASOUND-GUIDED GALLBLADDER DRAINAGE: A MULTICENTER STUDY FROM INDIA (WITH VIDEO)." ABSTRACT: BACKGROUND ENDOSCOPIC ULTRASOUND-GUIDED GALLBLADDER DRAINAGE (EUS-GBD) OFFERS A SAFE AND MINIMALLY INVASIVE ALTERNATIVE FOR PERCUTANEOUS CHOLECYSTOSTOMY (PCC) IN ACUTE CHOLECYSTITIS PATIENTS WITH HIGH-SURGICAL RISK. ADDITIONALLY, EUS-GBD SERVES AS A RESCUE BILIARY DRAINAGE IN MALIGNANT DISTAL BILIARY OBSTRUCTION. DESPITE ITS WIDESPREAD APPLICATION, DATA WITHIN THE INDIAN CONTEXT REMAINS SPARSE. THIS STUDY AIMS TO REPORT THE OUTCOMES OF EUS-GBD THROUGH THE FIRST MULTI-CENTER STUDY FROM INDIA. METHODS: WE RETROSPECTIVELY ANALYZED PATIENTS UNDERGOING EUS-GBD AT SIX TERTIARY CARE CENTERS OF INDIA FROM MARCH 2022 TO NOVEMBER 2023. EUS-GBD WAS PERFORMED BY FREE HAND OR OVER-THE-GUIDEWIRE TECHNIQUE WITH LUMEN-APPOSING METAL STENT. (LAMS) OR LARGE CALIBER METAL STENT (LCMS). THE PRIMARY OUTCOME WAS TECHNICAL SUCCESS (DEFINED AS SUCCESSFUL DEPLOYMENT. OF STENT BETWEEN GALLBLADDER AND STOMACH/DUODENAL LUMEN). THE SECONDARY OUTCOMES WERE CLINICAL SUCCESS (DEFINED. AS RESOLUTION OF SYMPTOMS OF ACUTE CHOLECYSTITIS AND MORE THAN > 50% REDUCTION IN BILIRUBIN LEVEL WITHIN TWO WEEKS IN DISTAL. BILIARY OBSTRUCTION), ADVERSE EVENT RATE, 30-DAY MORTALITY RATE AND 90-DAY REINTERVENTION RATE. RESULTS: TOTAL 29 PATIENTS (MEAN AGE 65.86 ± 12. (B)(6), 11 FEMALE) UNDERWENT EUS-GBD. THE INDICATION FOR EUS-GBD WERE ACUTE CHOLECYSTITIS (79.31%) AND RESCUE BILIARY DRAINAGE FOR MALIGNANT DISTAL BILIARY OBSTRUCTION (20.69%). LAMS WAS DEPLOYED IN 92.86%, PREDOMINANTLY BY FREE-HAND TECHNIQUE (78.57%). TECHNICAL AND CLINICAL SUCCESS RATES WERE 96.55% AND 82.75%, RESPECTIVELY. ADVERSE EVENTS OCCURRED IN 27.59% PATIENTS, WITH SEVERE ADVERSE EVENTS (BILE LEAK AND BLEEDING) BEING UNCOMMON (10%). BOTH 30-DAY MORTALITY RATE AND 90-DAY REINTERVENTION RATE WERE 13.79% IN PATIENTS. CHOLECYSTO-DUODENAL FISTULA FACILITATED CHOLECYSTOSCOPIC INTERVENTION AND STONE REMOVAL IN ONE PATIENT AND TRANSGASTRIC EUS-GBD DID NOT HAMPER BILIO-ENTERIC ANASTOMOSIS DURING WHIPPLE SURGERY IN TWO PATIENTS. ADVERSE EVENTS: ADVERSE EVENTS OCCURRED IN EIGHT PATIENTS (27.59%), WHILE EARLY ADVERSE EVENTS WERE SEEN IN SEVEN PATIENTS (24.13%). THESE INCLUDED MILD ABDOMINAL PAIN IN TWO PATIENTS, CARDIOVASCULAR EVENT (LIKELY ANESTHESIA RELATED) IN TWO PATIENTS, FEVER IN ONE PATIENT ON DAY 3 DUE TO LAMS CLOGGING, BILE LEAK IN ONE NECESSITATING SURGERY AND SEVERE BLEEDING IN ONE PATIENT FROM ABERRANT CYSTIC ARTERY REQUIRING ANGIOEMBOLIZATION. DELAYED EVENT WAS SEEN IN ONE PATIENT WITH MULTIPLE COMORBIDITIES RECEIVING DUAL ANTI-PLATELETS, PRESENTED WITH SEVERE BLEEDING ON DAY 29 OF EUS-GBD. ALSO REFER TO TABLE 3. AT AN AVERAGE FOLLOW-UP OF 102.76 ± 113.71 DAYS, 15 PATIENTS (51.72%) SURVIVED, 10 PATIENTS SUCCUMBED TO UNDERLYING DISEASES (MALIGNANCY, CHRONIC LIVER DISEASE, CHRONIC KIDNEY DISEASE, ETC.) AND FOUR PATIENTS WERE LOST TO FOLLOW-UP. CONCLUSION: EUS-GBD IS A SAFE AND EFFECTIVE TECHNIQUE FOR MANAGING ACUTE CHOLECYSTITIS IN HIGH-RISK PATIENTS AND FOR BILIARY DRAINAGE IN CASES WITH MALIGNANT DISTAL BILIARY OBSTRUCTION. TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS] TABLE 3 ABDOMINAL PAIN (2 PATIENTS) BILE LEAK (1 PATIENT WITHIN 1 WEEK REQUIRING SURGERY) SEVERE BLEEDING (2 PATIENTS, 1 PATIENT REQUIRING ANGIOEMBOLIZATION WITHIN 1 WEEK, 1 PATIENT WITH DELAYED BLEEDING AT DAY 29 REQUIRING TRANSFUSION AND SURGICAL INTERVENTION) DEATH (14 PATIENTS, 4 PATIENT DEATHS WITHIN 30 DAYS).Event Description: NO ADDITIONAL INFORMATION WAS RECEIVED. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIER: (B)(6). THE SUSPECT DEVICE HAS NOT BEEN RETURNED TO OLYMPUS FOR EVALUATION. THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE APPROVED FINAL INVESTIGATION. UPDATED FIELDS: D8, H6. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067841&pc=MOZ | MW5164234 |  | Death | GENZYME CORPORATION | 2025-01-02 05:00 | MOZ | SYNVISC PFS 8MG/ML (3X2ML) |  | Insufficient Information |  |  | 1 | Event Description: CALLER'S NAME: (B)(6). CALLER'S RELATIONSHIP TO THE CUSTOMER: PT'S NEPHEW. CALLER'S MAILING ADDRESS: (B)(6). CALLER'S PHONE NUMBER: (B)(6). CUSTOMER'S DECEASED DATE: (B)(6) 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067855&pc=QBJ | MW5164235 |  | Death | UNK | 2025-01-02 05:00 | QBJ | FREESTYLE LIBRE 3 SENSOR |  | Insufficient Information |  |  | 1 | Event Description: THE PATIENT PASSED ON (B)(6) 2024 AND ALSO STATED THAT THIS MEDICATION DID NOT CAUSE THE INCIDENT. THE PATIENT WAS ON THIS MEDICATION AT THE TIME OF PASSING. (B)(6). HCP DID CONSENT TO FOLLOW UP. (B)(6). NO FURTHER INFORMATION PROVIDED OR AVAILABLE REGARDING THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21067881&pc=NZU | MW5164236 | 2024-11-21 05:00 | Death | BD/C.R. BARD, INC. | 2025-01-02 05:00 | NZU | BD PUREWICK MALE EXTERNAL CATHETER |  | Renal Failure; Sepsis; Urinary Tract Infection; Fungal Infection; Liver Failure |  |  | 1 | Event Description: HELLO, MY HUSBAND CONTRACTED A MALE YEAST INFECTION THAT WENT SEPTIC AFTER USING THE "MALE EXTERNAL BD PUREWICK CATHETER". AT THE TIME HE WAS MEDICALLY FRAGILE AND HE DIED ON (B)(6) 2024 FROM LIVER AND KIDNEY FAILURE CAUSED PRIMARILY BY THE TOXIC ANTIFUNGALS GIVEN TO HIM TO TREAT THAT SEPTIC INFECTION. THE PRODUCT IS SUPPOSED TO WICK AWAY URINE AND IS SUPPOSED TO BE SUCTIONED INTO A VACUUM BOTTLE. WE HAD A GREAT EXPERIENCE WITH THE FEMALE CATHETER AND HOPED THE MALE VERSION WOULD BE EVEN BETTER. INSTEAD, THE MALE VERSION KEPT URINE ON MY HUSBAND'S SKIN OVERNIGHT AND WE WERE AT THE EMERGENCY ROOM WITHIN 24 HRS WITH A UTI (URINARY TRACT INFECTION) THAT TURNED SEPTIC. THE TESTS SHOWED CANDIDIASIS AS THE MAIN PROBLEM, AND HE WAS SUBSEQUENTLY TREATED WITH STRONG ANTIFUNGAL MEDS THAT BASICALLY OVERWHELMED HIS LIVER AND KIDNEYS. HE DIED ON (B)(6) 2024 OF LIVER AND KIDNEY FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21050645&pc=LJT | 3006260740-2024-07980 | 2022-12-14 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-01-01 05:00 | LJT | UNK MRI POWER PORT | Adverse Event Without Identified Device or Use Problem | Anemia; Atrial Fibrillation; Bacterial Infection; Fever; Sepsis | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT APPROXIMATELY SEVEN MONTHS AND TWENTY FOUR DAYS LATER POST PORT PLACEMENT PROCEDURE FOR THE ADMINISTRATION OF CHEMOTHERAPY, THE PATIENT REPORTEDLY DEVELOPED WITH FEVER, ANEMIA AND ATRIAL FIBRILLATION. IT WAS FURTHER REPORTED THAT PATIENT WAS DIAGNOSED WITH BACTEREMIA, BACTERIAL INFECTION AND SEPSIS. REPORTEDLY, THE INFECTED PORT WAS REMOVED. HOWEVER THE PATIENT WAS EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED.Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT APPROXIMATELY SEVEN MONTHS AND TWENTY FOUR DAYS LATER POST PORT PLACEMENT PROCEDURE FOR THE ADMINISTRATION OF CHEMOTHERAPY, THE PATIENT REPORTEDLY DEVELOPED WITH FEVER, ANEMIA AND ATRIAL FIBRILLATION. IT WAS FURTHER REPORTED THAT PATIENT WAS DIAGNOSED WITH BACTEREMIA, BACTERIAL INFECTION AND SEPSIS. REPORTEDLY, THE INFECTED PORT WAS REMOVED. HOWEVER THE PATIENT WAS EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: H10: MANUFACTURING REVIEW: A MANUFACTURING REVIEW WAS NOT REQUESTED AS THE LOT NUMBER REPORTED IS UNKNOWN. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORT OF THE PATIENT ALLEGEDLY DEVELOPING WEAKNESS, FEVER, ANEMIA, ATRIAL FIBRILLATION, BACTERIAL INFECTION, AND SEPSIS, NOR THE REPORTED PATIENT DEATH, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED SYMPTOMS AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3. H11: G2. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21050966&pc=BZD | 2518422-2025-000024 | 2023-09-26 04:00 | Death | RESPIRONICS, INC. | 2025-01-01 05:00 | BZD | REP DREAMSTATION AUTO BIPAP | Degraded | Heart Failure/Congestive Heart Failure | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO BIPAP, DOM-RECRT DOM DEVICE'S EVENT. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE DEVICE HAS HEART FAILURE AND PASSED AWAY. THERE WAS NO SERIOUS PATIENT HARM OR INJURY. THE PATIENT 'S REQUIRED NO MEDICAL INTERVENTION. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21051139&pc=BZD | 2518422-2025-000027 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-01 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Headache; Dizziness; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, CANCER, AND DEATH. NO FURTHER CLINICAL DETAILS OR MEDICAL INTERVENTION WERE REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21051325&pc=BZD | 2518422-2025-000028 | 2023-12-15 05:00 | Death | RESPIRONICS, INC. | 2025-01-01 05:00 | BZD | REMSTAR PRO C-FLEX+ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21051357&pc=BZD | 2518422-2025-000029 | 2023-09-21 04:00 | Death | RESPIRONICS, INC. | 2025-01-01 05:00 | BZD | REMSTAR PRO C-FLEX+ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. THE MANUFACTURER RECEIVED NEW AND RELEVANT INFORMATION ON 02/14/2024. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THERE WAS NO VISIBLE FOAM PARTICLES FOUND. DUST CONTAMINATION FOUND DURING THE DEVICE EVALUATION. THE DEVICE WAS REPAIRED. SECTION G3 HAS BEEN UPDATED. IN ADDITION, APPROPRIATE CODE UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188362&pc=BZD | 2518422-2025-004659 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188369&pc=BZD | 2518422-2025-004643 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LEUKEMIA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188527&pc=BZD | 2518422-2025-004680 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER PROSTATE AND ESOPHAGEAL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188662&pc=BZD | 2518422-2025-004691 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188698&pc=BZD | 2518422-2025-004686 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188713&pc=BZD | 2518422-2025-004726 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188803&pc=BZD | 2518422-2025-004769 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188964&pc=BZD | 2518422-2025-004805 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Nausea; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189004&pc=BZD | 2518422-2025-004844 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Respiratory Failure | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189007&pc=BZD | 2518422-2025-004755 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, BRAIN AND THROAT. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189009&pc=BZD | 2518422-2025-004825 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189062&pc=BZD | 2518422-2025-004885 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189072&pc=BZD | 2518422-2025-004851 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189189&pc=BZD | 2518422-2025-004888 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Hepatic or Biliary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, DIZZINESS, HEADACHE, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189198&pc=BZD | 2518422-2025-004884 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Nausea; Vomiting; Dizziness; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, NAUSEA, VOMITING, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189523&pc=CAF | 2518422-2025-004950 | 2023-10-31 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | CAF | INNOSPIRE DELUXE RENTAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K170853 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF INNOSPIRE DELUXE RENTAL. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189549&pc=BZD | 2518422-2025-004934 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND RIGHT LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189550&pc=BZD | 2518422-2025-004954 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189552&pc=BZD | 2518422-2025-004944 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD) | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189686&pc=BZD | 2518422-2025-004940 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Chronic Obstructive Pulmonary Disease (COPD); Fibrosis; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, DEATH, COPD AND PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189693&pc=BZD | 2518422-2025-004952 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND NSC LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189740&pc=BZD | 2518422-2025-005017 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND ESOPHAGEAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189747&pc=BZD | 2518422-2025-005023 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER, DEATH AND COLON CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189773&pc=MIH | 9612164-2025-00336 | 2024-12-26 05:00 | Death | MEDTRONIC IRELAND | 2025-01-20 05:00 | MIH | VALIANT CAPTIVIA - FF | Adverse Event Without Identified Device or Use Problem; Positioning Problem | Cardiac Arrest | P100040 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A THORACIC AORTIC ANEURYSM MEASURING 80.5MM IN DIAMETER, THE DEVICE VAMF3232C150TE WAS NOT TRACKING ACROSS THE PROXIMAL DESCENDING THORACIC AORTA (DTA). THE PATIENT EXPERIENCED CARDIO RESPIRATORY ARREST DURING THE PROCEDURE AND SUBSEQUENTLY DIED. THE PROCEDURE INVOLVED THE PLACEMENT OF VAMF2622C150TE PROXIMALLY AT THE LEVEL OF THE CELIAC ARTERY, FOLLOWED BY THE DEPLOYMENT OF VAMF3232C200TE. ATTEMPTS WERE MADE TO TRACK VAMF3232C150TE OVER A LUNDERQUIST WIRE WITH A BUDDY LUNDERQUIST WIRE, BUT THE DEVICE WAS NOT ABLE TO CROSS THE PROXIMAL DTA, EVEN WITH TWO BUDDY WIRES. WHILE CONTINUING TO ATTEMPT TO TRACK THE DEVICE, THE PATIENT WENT INTO CARDIO PULMONARY ARREST AND DIED ON THE TABLE. THE HEALTHCARE PROFESSIONAL COULD NOT DETERMINE THE CAUSE OF THE EVENT. Manufacturer Narrative: PRODUCT ANALYSIS CONCLUSION THE REPORTED DIFFICULT TO POSITION EVENT COULD NOT BE CONFIRMED ON THE PRE-IMPLANT FILMS PROVIDED; THEREFORE THE CAUSE OF THE EVENT COULD NOT BE DETERMINED. THE REPORTED CARDIAC ARREST AND DEATH COULD NOT BE ASSESSED. PROCEDURAL ANGIOGRAMS SHOWING ATTEMPTS TO POSITION THE VALIANT CAPTIVIA STENT GRAFT WERE NOT RECEIVED FOR A THOROUGH ANALYSIS OF THE EVENT. LACK OF FULL PRE-IMPLANT CT¿S DID NOT PROVIDE OPPORTUNITY FOR A MORE IN DEPTH ANALYSIS OF THE PATIENT¿S ANATOMY. IT IS LIKELY THAT PATIENT ANATOMY CONTRIBUTED TO THE DIFFICULTIES POSITIONING THE DEVICE, BUT THIS COULD NOT BE CONFIRMED. B5: ADDITIONAL INFORMATION IT WAS CONFIRMED THAT THE PROXIMAL DTA WAS TORTUOUS AND WAS A TYPE 3 AORTIC ARCH. THE DEVICE ORDER WAS CONFIRMED AS VAMF2622C150TE / (B)(6) FOLLOWED BY VAMF3232C200TE / (B)(6) AND VAMF3232C150TE/ (B)(6). NO DAMAGE WAS OBSERVED TO DEVICE OR PACKAGING PRIOR TO PROCEDURE. THE PHYSICIAN ADHERED TO THE DEPLOYMENT STEPS AS OUTLINED IN THE INSTRUCTIONS FOR USE (IFU). THE TORTUOSITY OF PROXIMAL DTA, ALONG WITH THE PATIENTS ANATOMY MAY HAVE CONTRIBUTED FOR TRACKING DIFFICULTY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: DEVICE ANALYSIS THE DEVICE RETURNED WITH THE EXTERNAL SLIDER AND TIP CAPTURE RELEASE HANDLE IN THE HOME POSITION. THE TAPER TIP WAS CURVED WITH A LONGITUDINAL SCRATCH VISIBLE ALONG THE SURFACE. THERE WAS NO FURTHER DAMAGE OBSERVED TO THE DEVICE. THE REPORTED POSITIONING DIFFICULTIES WERE CONFIRMED THROUGH ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189800&pc=MIH | 9612164-2025-00337 | 2024-12-26 05:00 | Death | MEDTRONIC IRELAND | 2025-01-20 05:00 | MIH | VALIANT CAPTIVIA - FF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P100040 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO TREAT A THORACIC AORTIC ANEURYSM MEASURING 80.5MM IN DIAMETER, THE DEVICE VAMF3232C150TE WAS NOT TRACKING ACROSS THE PROXIMAL DESCENDING THORACIC AORTA (DTA). THE PATIENT EXPERIENCED CARDIO RESPIRATORY ARREST DURING THE PROCEDURE AND SUBSEQUENTLY DIED. THE PROCEDURE INVOLVED THE PLACEMENT OF VAMF2622C150TE PROXIMALLY AT THE LEVEL OF THE CELIAC ARTERY, FOLLOWED BY THE DEPLOYMENT OF VAMF3232C200TE. ATTEMPTS WERE MADE TO TRACK VAMF3232C150TE OVER A LUNDERQUIST WIRE WITH A BUDDY LUNDERQUIST WIRE, BUT THE DEVICE WAS NOT ABLE TO CROSS THE PROXIMAL DTA, EVEN WITH TWO BUDDY WIRES. WHILE CONTINUING TO ATTEMPT TO TRACK THE DEVICE, THE PATIENT WENT INTO CARDIO PULMONARY ARREST AND DIED ON THE TABLE. THE HEALTHCARE PROFESSIONAL COULD NOT DETERMINE THE CAUSE OF THE EVENT. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE:VAMF3232C200TE : SERIAL NO UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS CONCLUSION THE REPORTED DIFFICULT TO POSITION EVENT COULD NOT BE CONFIRMED ON THE PRE-IMPLANT FILMS PROVIDED; THEREFORE THE CAUSE OF THE EVENT COULD NOT BE DETERMINED. THE REPORTED CARDIAC ARREST AND DEATH COULD NOT BE ASSESSED. PROCEDURAL ANGIOGRAMS SHOWING ATTEMPTS TO POSITION THE VALIANT CAPTIVIA STENT GRAFT WERE NOT RECEIVED FOR A THOROUGH ANALYSIS OF THE EVENT. LACK OF FULL PRE-IMPLANT CT¿S DID NOT PROVIDE OPPORTUNITY FOR A MORE IN DEPTH ANALYSIS OF THE PATIENT¿S ANATOMY. IT IS LIKELY THAT PATIENT ANATOMY CONTRIBUTED TO THE DIFFICULTIES POSITIONING THE DEVICE, BUT THIS COULD NOT BE CONFIRMED. B5: ADDITIONAL INFORMATION RECEIVED IT WAS CONFIRMED THAT THE PROXIMAL DTA WAS TORTUOUS AND WAS A TYPE 3 AORTIC ARCH. THE DEVICE ORDER WAS CONFIRMED AS VAMF2622C150TE / (B)(6) FOLLOWED BY VAMF3232C200TE / (B)(6) AND VAMF3232C150TE/ (B)(6). NO DAMAGE WAS OBSERVED TO DEVICE OR PACKAGING PRIOR TO PROCEDURE. THE PHYSICIAN ADHERED TO THE DEPLOYMENT STEPS AS OUTLINED IN THE INSTRUCTIONS FOR USE (IFU). THE TORTUOSITY OF PROXIMAL DTA, ALONG WITH THE PATIENTS ANATOMY MAY HAVE CONTRIBUTED FOR TRACKING DIFFICULTY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189882&pc=BZD | 2518422-2025-005049 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ACINAR PREDOMINANT LUNG CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189914&pc=BZD | 2518422-2025-004865 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189942&pc=MIH | 9612164-2025-00340 | 2025-01-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-20 05:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Leak/Splash | Hemorrhage/Bleeding; Insufficient Information | P100021 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A PRE-OPERATIVE RUPTURE OF AN 80MM ABDOMINAL AORTIC ANEURYSM, WHICH WAS TREATED USING AN ENDURANT STENT GRAFT SYSTEM. DURING THE POST-OPERATIVE PERIOD, THE PATIENT EXPERIENCED ANEURYSMAL BLEEDING AND WAS HYPOTENSIVE IN THE INTENSIVE CARE UNIT. ACCORDING TO THE PHYSICIAN, THE POST-OPERATIVE BLEEDING WITHIN THE ANEURYSM WAS DUE TO POROSITY OF THE ENDURANT PRODUCTS WHICH WERE INITIALLY UTILIZED. THE HEALTHCARE PROFESSIONAL INDICATED THAT THE CAUSE OF THE EVENT WAS DUE TO A TYPE IV ENDOLEAK NOTING OOZING FROM THE ENDURANT MAIN BODY STENT GRAFT AND ITS ENDURANT LIMBS. THE TYPE IV ENDOLEAK WAS DESCRIBED AS JETTING THROUGH THE FABRIC, CAUSING THE PHYSICIAN TO EXPRESS DISSATISFACTION. THE PATIENT WAS SENT BACK TO SURGERY FOR TREATMENT. TO ADDRESS THE TYPE IV ENDOLEAK, THE PHYSICIAN RELINED THE ENTIRE ENDURANT STENT GRAFT SYSTEM USING A NON-MDT CUFF AND LIMBS. SUBSEQUENTLY, DUE TO SIGNIFICANT ABDOMINAL ENLARGEMENT, THE PHYSICIAN PERFORMED AN OPEN PROCEDURE TO EVACUATE THE BLEEDING AND EXPLANT THE DEVICES. THE PATIENT DID NOT SURVIVE THE PROCEDURE AND EXPIRED ON THE SAME DATE OF THE PROCEDURE. Manufacturer Narrative: FILM ANALYSIS THE REPORTED TYPE IV ENDOLEAK COULD NOT BE CONFIRMED BASED ON THE PRE-IMPLANT FILMS PROVIDED; THEREFORE, THE CAUSE OF THE EVENT COULD NOT BE DETERMINED. PROCEDURAL ANGIOGRAMS AND POST-IMPLANT IMAGING WERE NOT AVAILABLE FOR ASSESSMENT OF THE ENDOLEAK AND STENT GRAFT IN VIVO CONFIGURATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E TLW1616C124E, SERIAL/LOT #: (B)(6), UBD: 17-SEP-2026, UDI#: (B)(4); PRODUCT ID: ETLW1616C156E, SERIAL/LOT #: (B)(6), UBD: 12-AUG-2026 AND UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21189973&pc=MNS | 2518422-2025-005081 | 2022-05-16 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | MNS | REP DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION ST30. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190034&pc=BZD | 2518422-2025-005085 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, CHOLANGIOCARCINOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190051&pc=NVZ | 3004209178-2025-01060 | 2024-11-29 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-20 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF CAPTURE. IT WAS FURTHER REPORTED THAT THE RV LEAD EXHIBITED U NDERSENSING. THE RV LEAD REMAINS IN USE. IT WAS FURTHER REPORTED THAT THE PATIENT HAD DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190078&pc=DTB | 2649622-2025-02094 | 2024-11-29 05:00 | Death | MPRI | 2025-01-20 05:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF CAPTURE. IT WAS FURTHER REPORTED THAT THE RV LEAD EXHIBITED UNDERSENSING. THE RV LEAD REMAINS IN USE. IT WAS FURTHER REPORTED THAT THE PATIENT HAD DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190110&pc=BZD | 2518422-2025-004878 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD) | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CHRONIC OBSTRUCTIVE PULMONARY DISEASE PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190116&pc=BZD | 2518422-2025-005117 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190150&pc=BZD | 2518422-2025-005086 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Pulmonary Hypertension; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS, HEADACHE, LUNG DISEASE AND PULMONARY HYPERTENSION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H: TYPE OF REPORTED COMPLAINT WAS CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190153&pc=BZD | 2518422-2025-005139 | 2022-09-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190301&pc=BZD | 2518422-2025-005174 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190403&pc=BZD | 2518422-2025-005184 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Heart Failure/Congestive Heart Failure | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CONGESTIVE HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191010&pc=BZD | 2518422-2025-005270 | 2022-02-23 05:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191089&pc=DSQ | 2916596-2025-00331 | 2021-07-01 04:00 | Death | THORATEC CORPORATION | 2025-01-20 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Aortic Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿RIGHT VENTRICULAR¿PULMONARY ARTERIAL COUPLING IN PATIENTS WITH IMPLANTED LEFT VENTRICULAR ASSIST DEVICES¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT VENTRICULAR FAILURE, STROKE, ANEMIA, ATRIAL FIBRILLATION/FLUTTER, MITRAL/TRICUSPID/AORTIC REGURGITATION AND DEATH. BASIC CLINICAL, ECHOCARDIOGRAPHIC, AND RIGHT HEART CATHETERIZATION CHARACTERISTICS WERE REVIEWED IN THE SCREENED PATIENTS, AND THE INFORMATION INCLUDED PREVIOUS STROKE, ANEMIA, AF/FLUTTER, MR, AR, TR. ADDITIONALLY, EARLY POST-IMPLANT RHF OCCURRED IN 11 SUBJECTS, AND LATE IN 17. THE GOAL OF THE STUDY WAS TO DETERMINE THE PROGNOSTIC IMPACT OF RIGHT VENTRICULAR (RV)- PULMONARY ARTERIAL (PA) COUPLING IN 72 PATIENTS IMPLANTED BETWEEN OCT2015 AND JUL2021 WITH LEFT VENTRICULAR ASSIST DEVICE¿S (LVAD). THE STUDY CLAIMED THAT THE ADDITIONAL VOLUME LOAD SUPPLIED BY A LVAD PUMP AFFECTS RV CONTRACTILITY THROUGH LEFT VENTRICULAR REMODELING, PULMONARY VASCULAR REMODELING, POSTCAPILLARY PULMONARY HYPERTENSION, AND INCREASED RV AFTERLOAD. RV-PA IS BASED ON INDEXING RV PERFORMANCE TO RV AFTERLOAD. WHEN RV ¿ PA COUPLING IS INTACT, THE MECHANICAL STROKE WORK IS TRANSFERRED TO THE PULMONARY CIRCUIT, MEANING THAT AN INCREASE IN RV AFTERLOAD WILL MAKE THEE CONTRACTIBILITY OF RV INCREASE. IN RV ¿ PA UNCOUPLING, RV IS NOT ABLE TO MAINTAIN ADEQUATE FLOW ACROSS THE PULMONARY VASCULATURE FOR THE AMOUNT OF RV AFTERLOAD. AS A RESULT, RV - PA UNCOUPLING MAY BE ASSOCIATED WITH ADVERSE OUTCOMES IN PATIENTS WITH IMPLANTED LVAD¿S. THE RV-PA COUPLING WAS ASSESSED PREOPERATIVELY BY THE RATIO OF RV FREE WALL STRAIN (RVFWS) DERIVED FROM SPECKLE-TRACKING ECHOCARDIOGRAPHY AND NONINVASIVELY MEASURED PEAK RV SYSTOLIC PRESSURE (RVSP). ONLY PATIENTS WITH POOR RV MYOCARDIAL IMAGING WERE EXCLUDED FROM THE STUDY. THE MEDIAN AGE OF THE STUDY GROUP WAS 57 YEARS. (P=0.399) THE STUDY FOUND THAT 19 PATIENTS HAD RV-PA UNCOUPLING. WHEN COMPARING OUTCOMES AFTER 12 MONTHS, 47.37% OF PATIENTS (9 PATIENTS) WITH RV-PA UNCOUPLING PASSED AWAY, WHILE 13.21% OF PATIENTS (7 PATIENTS) WHERE RV-PA REMAINED COUPLED ALSO PASSED AWAY, P=0.003. FURTHERMORE AFTER 12 MONTHS, 89.47% OF PATIENTS (17 PATIENTS) WITH RV-PA UNCOUPLING EITHER PASSED AWAY OR WERE HOSPITALIZED FOR RIGHT HEART FAILURE (RHF), WHILE 30.19% OF PATIENTS (16 PATIENTS) WHERE RV-PA REMAINED COUPLED EITHER PASSED AWAY OR WERE HOSPITALIZED FOR RHF, P<0.001. IN TOTAL 16 PATIENTS PASSED AWAY AND 17 PATIENTS WERE HOSPITALIZED FOR RHF BUT DID NOT DIE IN THE 12-MONTH STUDY PERIOD. THE PAPER CONCLUDED THAT EVALUATION OF THE RV FUNCTION BY INDEXING RV PERFORMANCE TO PA AFTERLOAD MAY PROVIDE OPTIMAL PROGNOSTIC INFORMATION FOR PATIENTS CONSIDERED FOR LVAD SURGERY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN, DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01JUL2021 AS DATA WAS COLLECTED BETWEEN OCT2012 AND JUL2021. AUTHOR INFORMATION: STAPÓR, M., ET AL. (2024). RIGHT VENTRICULAR-PULMONARY ARTERIAL COUPLING IN PATIENTS WITH IMPLANTED LEFT VENTRICULAR ASSIST DEVICES. HELLENIC JOURNAL OF CARDIOLOGY: HJC = HELLENIKE KARDIOLOGIKE EPITHEORESE, 77, 44¿53. HTTPS://DOI.ORG/10.1016/J.HJC.2023.06.002. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THIS DOCUMENT ALSO WARNS THE USER THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191109&pc=NGV | 2124215-2024-84130 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEVICE EXPLANTATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 TIME TO EVENT: 10 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191110&pc=NGV | 2124215-2024-84126 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR IMPLANT EMBOLIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 TIME TO EVENT: 10 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191112&pc=NGV | 2124215-2024-84134 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CARDIAC ARREST. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191114&pc=NGV | 2124215-2024-84187 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191116&pc=NGV | 2124215-2024-84158 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 TIME TO EVENT: 0 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191120&pc=NGV | 2124215-2024-84131 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HOSPITALIZATION OR PROLONGED HOSPITALIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 10 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191124&pc=NGV | 2124215-2024-84133 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 10 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191127&pc=NGV | 2124215-2024-84169 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HEART FAILURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191208&pc=BZD | 2518422-2025-005278 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191315&pc=BZD | 2518422-2025-005275 | 2023-10-08 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Headache; Dizziness; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS AND/OR HEADACHE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191449&pc=NGV | 2124215-2024-84128 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR CARDIOVASCULAR PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 27 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS. (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191470&pc=BZD | 2518422-2025-005165 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION RELATED TO A REP DREAMSTATION AUTO CPAP DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE, LIVER DISEASE, CANCER, AND DEATH. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191477&pc=NGV | 2124215-2024-84129 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 2 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191521&pc=NGV | 2124215-2024-84135 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 8 | Event Description: THIS REPORT SUMMARIZES 8 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 67 TO 88 YEARS. PATIENT SEX: 4 FEMALE; 4 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191522&pc=NGV | 2124215-2024-84132 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 10 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191571&pc=BZD | 2518422-2025-005303 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191693&pc=DSQ | 2916596-2025-00084 | 2024-10-27 04:00 | Death | THORATEC CORPORATION | 2025-01-20 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Ischemia; Tachycardia; Fungal Infection; Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. INTESTINAL ISCHEMIA WAS SUGGESTED, BUT NO OBVIOUS THROMBI WERE OBSERVED WITHIN THE VISIBLE RANGE OF THE AORTIC VALVE, ARTIFICIAL BLOOD VESSELS, OR THE EXCISED LEFT VENTRICULAR ASSIST DEVICE.Event Description: AFTER A FALL AND SPRAIN AT HOME, THE PATIENT'S ACTIVITIES OF DAILY LIVING (ADL) DECLINED, LEADING TO POOR FOOD INTAKE. ON (B)(6) 2024, THE PATIENT WAS URGENTLY READMITTED DUE TO DEHYDRATION, LOW FLOW, AND VT (VENTRICULAR TACHYCARDIA). AT THAT TIME, MODERATE AORTIC REGURGITATION (AR) WAS NOTED. POST-ADMISSION, THE PATIENT EXHIBITED SYMPTOMS OF RIGHT HEART FAILURE (ASCITES ACCUMULATION) AND SHOWED POOR RESPONSE TO MEDICAL TREATMENT. GRADUALLY, THE PATIENT DEVELOPED ANOREXIA AND MALNUTRITION, LIKELY DUE TO LOW PERFUSION CAUSED BY AR AND LOW OUTPUT SYNDROME (LOS). ON (B)(6) 2024, AN AORTIC VALVE REPLACEMENT (AVR) SURGERY WAS PLANNED. HOWEVER, ON (B)(6) 2024, THE PATIENT DEVELOPED A CENTRAL VENOUS (CV) CATHETER INFECTION AND FUNGAL BACTEREMIA, LEADING TO THE POSTPONEMENT OF THE SURGERY FOR TREATMENT INITIATION. ENTERAL NUTRITION WAS STARTED ON (B)(6) 2024, BUT ON (B)(6) 2024, THE PATIENT DEVELOPED ABDOMINAL PAIN. A CT SCAN ON (B)(6) 2024 REVEALED PORTAL VENOUS GAS AND PNEUMATOSIS INTESTINALIS, INDICATING INTESTINAL NECROSIS. LOW FLOW PERSISTED, AND THE PATIENT WAS PRONOUNCED DEAD AT 1:08 PM ON THE SAME DAY.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE SITE DID NOT BELIEVE THE FALL OR DEHYDRATION WAS LVAD RELATED. IT WAS NOTED THAT THE PATIENTS AR WAS TRIVIAL PRE-SURGERY BUT PROGRESSED TO MODERATE 6 MONTHS POST IMPLANT. THE DEVICE COULD NOT BE RULED OUT AS A POSSIBLE CONTRIBUTORY FACTORY TO THE AR SINCE THE DOWNWARD FORCE FROM THE PUMP FLOW MAY HAVE AFFECTED THE AORTIC VALVE, HOWEVER THE PATIENTS BACKGROUND AS AN ELDERLY PERSON MAY HAVE ALSO CONTRIBUTED. BASED ON A MYOCARDIAL BIOPSY, A DIAGNOSIS OF CARDIAC SARCOIDOSIS WAS MADE. GIVEN THE PREOPERATIVE CATECHOLAMINE DEPENDENCE AND REFRACTORY ARRHYTHMIAS, IT IS PRESUMED THAT RIGHT HEART FAILURE HAD ALREADY DEVELOPED PRE-LVAD, HOWEVER IT WAS NOTED THAT THE DEVICE COULD NOT BE RULED OUT AS A POSSIBLE CONTRIBUTORY FACTOR DUE TO THE FACT THAT LOW PERFUSION DUE TO POSTOPERATIVE AORTIC REGURGITATION (AR) AND DECREASED LEFT HEART FUNCTION MAY HAVE LED TO REDUCED PUMP EFFICIENCY, INDIRECTLY INCREASING THE LOAD ON THE RIGHT HEART, AND THUS POSSIBLY CONTRIBUTING TO THE RIGHT HEART FAILURE. THE CENTER BELIEVED THE CAUSE OF THE LOW FLOW ALARMS WAS DEHYDRATION DUE TO DECREASED ADL POST FALL, ALONG WITH POOR ORAL INTAKE. THE PATIENT SHOWED TEMPORARY FLOW IMPROVEMENT AFTER HOSPITALIZATION THROUGH FLUID REPLACEMENT. DUE TO THE PATIENT'S LOSS OF APPETITE, THE PATIENT REMAINED DEPENDENT ON INTRAVENOUS FLUIDS. NO LOG FILES WERE AVAILABLE. Manufacturer Narrative: CORRECTION: SECTION D9, H6 ADDITIONAL CODES. HEALTH EFFECT - CLINICAL CODES: 1807 - DEHYDRATION. 2596 - ASCITES. 4569 - DECREASED APPETITE. 2501 - LOW CARDIAC OUTPUT. 4868 - HEMODYNAMIC INSTABILITY. 1685 - ABDOMINAL PAIN. 1971 - NECROSIS. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES. DECREASED FLOW WAS CONFIRMED THROUGH ANALYSIS OF THE RETRIEVED LOG FILES. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 7¿ FROM THE PUMP HEADER. THE REMAINDER OF THE PUMP CABLE AND THE MODULAR CABLE WERE ALSO RETURNED. APPROXIMATELY 1¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE BEND RELIEF ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED DETACHED FROM THE PUMP WITH THE CUFF LOCK DISENGAGED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE CAPTURED DECREASED FLOW ON (B)(6) 2024 FROM 12:58:08 THROUGH 13:09:19. NO OTHER NOTABLE EVENTS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED FOR THE DURATION OF THE LOG FILE. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS RIGHT HEART FAILURE, AORTIC INSUFFICIENCY, THROMBOEMBOLISM, ARRHYTHMIA, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO PROVIDES AN EXPLANATION OF EACH OF THE PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, AND OUTLINES INDICATIONS OF THROMBUS AS WELL AS HOW TO RESPOND TO SUCH EVENTS. THIS SECTION ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191695&pc=NGV | 2124215-2024-84156 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR ADDITIONAL DEVICE REQUIRED. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024, AVERAGE TIME TO EVENT: 0 DAYS, DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191697&pc=NGV | 2124215-2024-84157 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 0 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191722&pc=NGV | 2124215-2024-84160 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR CARDIAC TAMPONADE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191724&pc=NGV | 2124215-2024-84161 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 78 TO 91 YEARS PATIENT SEX: 2 FEMALE; 1 MALE TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 41 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191725&pc=NGV | 2124215-2024-84170 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR PERICARDIAL EFFUSION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. OF THE REPORTED 2 PERICARDIAL EFFUSION DEATH EVENTS: 1 INVOLVED SURGICAL INTERVENTION. 2 INVOLVED CARDIAC TAMPONADE. 2 INVOLVED AN ADDITIONAL DEVICE (DRAINAGE). TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191726&pc=NGV | 2124215-2024-84159 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-20 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR CARDIAC ARREST. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191750&pc=LWP | 2017865-2025-11168 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-20 05:00 | LWP | ENDURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-11169, 2017865-2025-11172. IT WAS REPORTED THAT THE PACEMAKER SYSTEM WAS SUCCESSFULLY IMPLANTED ON (B)(6), 2025. ON (B)(6), 2025, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191761&pc=NVN | 2017865-2025-11169 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-20 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-11168, 2017865-2025-11172. IT WAS REPORTED THAT THE PACEMAKER SYSTEM WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191763&pc=NVN | 2017865-2025-11172 | 2025-01-10 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-20 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-11169, 2017865-2025-11168. IT WAS REPORTED THAT THE PACEMAKER SYSTEM WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21192692&pc=NPT | 2015691-2025-00401 | 2024-12-27 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-20 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED FROM CANADA, ANNULAR RUPTURE OCCURRED POST DEPLOYMENT OF A 26MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION. THE RUPTURE WAS THOUGHT TO BE CAUSED BY CALCIUM IN THE SINUS. NO ACTIONS WERE TAKEN, AND THE PATIENT EXPIRED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (PRESENCE OF CALCIUM IN THE SINUS) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21193028&pc=BZD | 2518422-2025-100925 | 2024-11-05 05:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21193128&pc=BZD | 2518422-2025-100926 | 2024-12-20 05:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: DESPITE MULTIPLE ATTEMPTS BY EMAIL TO THE CUSTOMER ON 01/31/2025, 02/06/2025 AND 02/11/2025, THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE HAS BEEN NO RESPONSE FROM THE CUSTOMER ON THE RETURN OF THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. SECTION H HAS BEEN UPDATED IN THE CURRENT REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21193199&pc=BZD | 2518422-2025-005319 | 2024-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | BIPAP PLUS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K061034 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP PLUS DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 08/09/2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP PLUS DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. REPEATED ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. BOX H WAS UPDATED IN THIS REPORT. Manufacturer Narrative: THIS DEVICE BIPAP PLUS WAS MANUFACTURED IN 03/28/2006 WHICH IS PRIOR TO UDI REQUIREMENT IMPLEMENTATION. THIS DEVICE DOES NOT HAVE A UDI AND NO UDI WILL BE LISTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194405&pc=BZD | 2518422-2025-005322 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-20 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21187689&pc=BZD | 2518422-2025-004510 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Pulmonary Dysfunction; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/ OR HEADACHE, HYPERSENSITIVITY, LIVER DISEASE/ TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, LUNG CANCER AND LIVER CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21187691&pc=BZD | 2518422-2025-004507 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Bronchitis; Pneumonia; Pulmonary Dysfunction | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CHRONIC BRONCHITIS, RECURRING PNEUMONIA, COPD. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21187963&pc=BZD | 2518422-2025-004518 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAM STATION CPAP | Degraded; Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: IN THE PREVIOUS REPORT IN SECTION D1 BRAND NAME AND H6 MEDICAL DEVICE PROBLEM CODE WAS CAPTURED INCORRECTLY, AND BOTH WERE CORRECTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21187965&pc=BZD | 2518422-2025-004521 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAM STATION CPAP | Degraded; Adverse Event Without Identified Device or Use Problem | Chronic Obstructive Pulmonary Disease (COPD); Aspiration Pneumonitis | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, ASPIRATION PNEUMONITIS AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN THE PREVIOUS REPORT IN SECTION D1 BRAND NAME AND H6 MEDICAL DEVICE PROBLEM CODE WAS CAPTURED INCORRECTLY, AND BOTH WERE CORRECTED IN THIS REPORT. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188183&pc=NVN | 2017865-2025-11111 | 2024-12-24 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-19 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11109. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11110. FOLLOWING THE ELECTRONICS PERFORMANCE INDICATOR (EPI), AN ALERT WAS RECEIVED BY THE CLINICIAN AND AWAITING DEVICE EXPLANT. HOWEVER, PATIENT ENDED UP PASSING AWAY DUE TO COMPLICATIONS FROM PNEUMONIA BEFORE DEVICE COULD BE EXPLANTED. IT WAS UNKNOWN IF THERE WERE ANY ALLEGATIONS THAT INFECTION WAS RELATED TO THE ABBOTT SYSTEM OR ANY PROCEDURES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188184&pc=NVZ | 2017865-2025-11109 | 2024-12-24 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-19 05:00 | NVZ | ASSURITY RF DR | Premature Discharge of Battery; Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P880086 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11110, RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11111. FOLLOWING THE ELECTRONICS PERFORMANCE INDICATOR (EPI), AN ALERT WAS RECEIVED BY THE CLINICIAN AND AWAITING DEVICE EXPLANT. HOWEVER, PATIENT ENDED UP PASSING AWAY DUE TO COMPLICATIONS FROM PNEUMONIA BEFORE DEVICE COULD BE EXPLANTED. IT WAS UNKNOWN IF THERE WERE ANY ALLEGATIONS THAT INFECTION WAS RELATED TO THE ABBOTT SYSTEM OR ANY PROCEDURES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188185&pc=NVN | 2017865-2025-11110 | 2024-12-24 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-19 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11109. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11111. FOLLOWING THE ELECTRONICS PERFORMANCE INDICATOR (EPI), AN ALERT WAS RECEIVED BY THE CLINICIAN AND AWAITING DEVICE EXPLANT. HOWEVER, PATIENT ENDED UP PASSING AWAY DUE TO COMPLICATIONS FROM PNEUMONIA BEFORE DEVICE COULD BE EXPLANTED. IT WAS UNKNOWN IF THERE WERE ANY ALLEGATIONS THAT INFECTION WAS RELATED TO THE ABBOTT SYSTEM OR ANY PROCEDURES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188318&pc=BZD | 2518422-2025-004627 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188320&pc=BZD | 2518422-2025-004623 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL ADAMOU CARCINOMA, EYE IRRITATION, SKIN IRRITATION CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188321&pc=BZD | 2518422-2025-004630 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-19 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186476&pc=BZD | 2518422-2025-004467 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-18 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186505&pc=NPT | 2025587-2025-00483 | 2025-01-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-18 05:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Gradient Increase; Perivalvular Leak; Insufficient Information; Device Stenosis; Central Regurgitation | Aortic Valve Stenosis; High Blood Pressure/ Hypertension; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT NINE YEARS AND FIVE MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, WORSENING OF AORTIC STENOSIS (AS) OCCURRED. NO TREATMENT WAS PROVIDED. FURTHER MEDICAL EVALUATION WAS RECOMMENDED BY THE PHYSICIAN. AS WAS ONGOING. NO PATIENT SYMPTOMS WERE REPORTED.Event Description: IT WAS REPORTED THAT NINE YEARS AND FIVE MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, WORSENING OF AORTIC STENOSIS (AS) OCCURRED. NO TREATMENT WAS PROVIDED. FURTHER MEDICAL EVALUATION WAS RECOMMENDED BY THE PHYSICIAN. AS WAS ONGOING. NO PATIENT SYMPTOMS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED MULTIPLE JETS OF UNSPECIFIED PARAVALVULAR LEAK AND A MEAN GRADIENT OF 42 MM HG. A COMPUTED TOMOGRAPHY SCAN WAS ORDERED.Event Description: ADDITIONAL INFORMATION INDICATED THAT A TRANSTHORACIC ECHOCARDIOGRAM (TTE) APPROXIMATELY 9 YEARS AND 5 MONTHS FOLLOWING THE VALVE IMPLANT IDENTIFIED THE SEVERE AORTIC STENOSIS WITH AN ATRIOVENTRICULAR (AV) VMAX 4.0-4.5 METERS PER SECOND (M/S). MILD PARAVALVULAR LEAK (PVL) AND SOME TRANSVALVULAR AORTIC REGURGITATION WHICH WAS QUESTIONED AS MODERATE SEVERITY. THE LEFT VENTRICULAR EJECTION FRACTION (LVEF) WAS REPORTED AS NORMAL. MODERATE RESTING PULMONARY HYPERTENSION WAS ALSO PRESENT. ON AN UNKNOWN DATE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2, B5, H1, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186646&pc=LWS | 2017865-2025-11083 | 2024-12-27 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-18 05:00 | LWS | FORTIFY ASSURA VR ICD, US | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P910023 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-11084. IT WAS REPORTED BY THE CLINICIAN ON (B)(6) 2024 THAT A REMOTE TRANSMISSION WAS RECEIVED NOTING THE PATIENT EXPERIENCED A VENTRICULAR TACHYCARDIA (VT) STORM AND RECEIVED MULTIPLE ANTI TACHYCARDIA PACING (ATP) THERAPY AND SHOCKS. THE CLINIC REACHED OUT TO THE FAMILY AND THE PATIENT WAS REPORTED TO HAVE DIED (B)(6) 2024. THE FAMILY REPORTED EMERGENCY SERVICES HAD BEEN CONTACTED BUT THE PATIENT HAD BEEN CONFIRMED TO BE DECEASED UPON THEIR ARRIVAL. TECHNICAL SERVICES WAS CONTACTED TO ENSURE SOME UNDERSENSING DID NOT IMPACT DELIVERY OF THERAPY. THE UNDER SENSING WAS NOTED TO BE RIGHT AFTER A VERY LARGE COMPLEX AND SO THE FOLLOWING BEAT WAS NOT SENSED BUT THE NEXT BEATS WERE. THE ISSUE WAS NOTED NOT TO BE DEVICE RELATED BUT DUE TO THE PATIENTS INTRINSIC QRS. NOTHING WAS NOTED TO INDICATE THE DEVICE DID NOT PERFORM AS EXPECTED. THE ARRHYTHMIA/VENTRICULAR TACHYCARDIA (VT) STORM OBSERVED WERE FELT BY THE PHYSICIAN TO BE INTRINSIC. THE ANTI-TACHYCARDIA THERAPY AND SHOCKS OBSERVED WERE APPROPRIATE. THE PATIENT HAD MULTIPLE CO-MORBIDITIES. THE PHYSICIAN AGREED THE DEVICE DID NOT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186647&pc=NVY | 2017865-2025-11084 | 2024-12-27 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-18 05:00 | NVY | DURATA STS OPTIM PASSIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-11083. IT WAS REPORTED BY THE CLINICIAN ON (B)(6) 2024 THAT A REMOTE TRANSMISSION WAS RECEIVED NOTING THE PATIENT EXPERIENCED A VENTRICULAR TACHYCARDIA (VT) STORM AND RECEIVED MULTIPLE ANTI TACHYCARDIA PACING (ATP) THERAPY AND SHOCKS. THE CLINIC REACHED OUT TO THE FAMILY AND THE PATIENT WAS REPORTED TO HAVE DIED (B)(6) 2024. THE FAMILY REPORTED EMERGENCY SERVICES HAD BEEN CONTACTED BUT THE PATIENT HAD BEEN CONFIRMED TO BE DECEASED UPON THEIR ARRIVAL. TECHNICAL SERVICES WAS CONTACTED TO ENSURE SOME UNDERSENSING DID NOT IMPACT DELIVERY OF THERAPY. THE UNDER SENSING WAS NOTED TO BE RIGHT AFTER A VERY LARGE COMPLEX AND SO THE FOLLOWING BEAT WAS NOT SENSED BUT THE NEXT BEATS WERE. THE ISSUE WAS NOTED NOT TO BE DEVICE RELATED BUT DUE TO THE PATIENTS INTRINSIC QRS. NOTHING WAS NOTED TO INDICATE THE DEVICE DID NOT PERFORM AS EXPECTED. THE ARRHYTHMIA/VENTRICULAR TACHYCARDIA (VT) STORM OBSERVED WERE FELT BY THE PHYSICIAN TO BE INTRINSIC. THE ANTI-TACHYCARDIA THERAPY AND SHOCKS OBSERVED WERE APPROPRIATE. THE PATIENT HAD MULTIPLE CO-MORBIDITIES. THE PHYSICIAN AGREED THE DEVICE DID NOT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21176033&pc=BZD | 2518422-2025-003952 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21176466&pc=OZP | 2032227-2025-111735 | 2024-12-26 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-17 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hyperglycemia; Cancer | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS LUNG CANCER AND HYPERGLYCEMIA. THE PUMP WAS NOT WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-387A, MMT-1884, MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-387A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A.Event Description: UPDATED SUMMARY: CUSTOMER DIDN'T ALLEGE ANY ALLEGATION ON THE DEVICE. ALSO, THE CAUSE OF DEATH WAS DUE TO LUNG CANCER AND THE INSULIN PUMP WAS NOT WORN AT THE TIME OF DEATH. SO, THE EVENT SUBMITTED UNDER REGULATORY REPORT NUMBER 2032227-2025-111735 IS NOT REPORTABLE ON THE INSULIN PUMP. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21176844&pc=PNJ | 9612164-2025-00313 | 2024-12-30 05:00 | Death | MEDTRONIC IRELAND | 2025-01-17 05:00 | PNJ | MICRA¿ AV2 | Coagulation in Device or Device Ingredient; Difficult to Flush; High Capture Threshold | Dyspnea | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT THE PATIENT PASSED DIED. AFTER THE FIRST ATTEMPT WAS MADE AND THE LEADLESS IPG WAS RECAPTURED DUE TO ELEVATED THRESHOLDS. CLOTTING WAS SUSPECTED IN THE DELIVERY SYSTEM DUE TO DIFFICULTY IN FLUSHING. THE CATHETER APPEARED TO HAVE THICK BLOOD IN THE DELIVERY SYSTEM CUP. SEVERAL FLUSHING TO REMOVE BLOOD WAS DONE AND WAS SUCCESSFULLY PREPARED FOR IMPLANT. AFTER THE DELIVERY SYSTEM WAS FLUSHED IT WAS ABOUT TO BE INSERTED INTO THE INTRODUCER. THE PATIENT STOPPED BREATHING AND IMPLANT WAS ABORTED. THE PATIENT WAS A DO NOT RESUSCITATE AND PATIENT WAS TRANSFERRED TO HOLDING AREA. THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177188&pc=BZD | 2518422-2025-004152 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | PHILIPS CPAP | Degraded | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND TUMOR. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177499&pc=MXD | 3008973940-2025-00479 | 2025-01-02 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-17 05:00 | MXD | LINQ II¿ INSERTABLE CARDIAC MONITOR | Signal Artifact/Noise; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | K200795 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIAC MONITOR (ICM) FAILED TO DETECT AN ASYSTOLE EPISODE DUE TO NOISE. IT WAS ALSO REPORTED THAT A PATIENT WAS MISTAKENLY DISCONTINUED FROM THE CARELINK SYSTEM. THE ISSUE WAS IDENTIFIED, AND IT WAS NOTED THAT THE CLINIC RESPONSIBLE FOR THE DISCONTINUATION COULD REVERSE THE ACTION WITHIN 30 DAYS BY FOLLOWING SPECIFIC STEPS IN THE SYSTEM. THE ICM REMAINS IN USE. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PATIENT DIED IN A MANNER UNRELATED TO THE DEVICE SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED FALSE ASYSTOLE DUE TO UNDERSENSING. ANALYSIS OF THE DEVICE MEMORY HAD AN OBSERVATION RELATING TO THE DETECTION. ANALYSIS OF THE DEVICE MEMORY INDICATED NOISE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177507&pc=DSQ | 2916596-2025-00118 | 2024-12-30 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF ANY AUTOPSY WAS PERFORMED AND PROVIDED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED AND RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: D4: SERIAL NUMBER, LOT NUMBER, EXPIRATION DATE, AND UDI CORRECTED. H4: MANUFACTURE DATE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE PATIENT OUTCOME; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177682&pc=BZD | 2518422-2025-004247 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Gastroesophageal Burn | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND ESOPHAGEAL . THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177690&pc=BZD | 2518422-2025-004251 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177719&pc=MVK | 3008642652-2025-00596 | 2024-12-07 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing; Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 17:59:37 ON (B)(6)2024. AT 16:30:17, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 16:31:24, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 20 BPM WITH PVC¿S. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH PVC¿S. AT 16:31:51, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS ASYSTOLE. AT 16:32:18, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM WITH PVC¿S. AT 16:57:46, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 16:58:20, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 70 BPM DEGRADING TO ASYSTOLE WITH CPR/MOTION ARTIFACT . THE DEVICE WAS SHUT DOWN AT 18:20:21 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177724&pc=MVK | 3008642652-2025-00598 | 2024-12-14 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 02:30:10 ON (B)(6)2024. AT 03:28:58 AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 180 BPM. AT 03:29:35, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 190 BPM. POST SHOCK RHYTHM WAS VT @ 170 BPM. AT 03:30:02, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 170 BPM. POST SHOCK RHYTHM WAS VT @ 210 BPM. AT 03:30:26, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS VT @ 240 BPM WITH CPR/MOTION ARTIFACT. AT 03:31:01, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 03:32:16, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 03:32:51, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 220 BPM DEGRADING TO VF WITH LOW/VARYING AMPLITUDES. THE DEVICE WAS SHUT DOWN AT 03:33:42 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177725&pc=MVK | 3008642652-2025-00597 | 2024-11-19 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 14:15:19 ON (B)(6) 2024. AT 00:57:25 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 00:58:00, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM. AT 00:58:00, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS IDIOVENTRICULAR RHYTHM @ 20 BPM. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 40 BPM. AT 01:03:16, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 01:03:57, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM. AT 01:05:33, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 01:06:40, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 01:07:07, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM. AT 01:07:33, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 01:40:07 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177731&pc=MVK | 3008642652-2025-00600 | 2024-11-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 18:56:36 ON (B)(6)2024. AT 22:43:37, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 280 BPM WITH CPR/ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VT. CPR/ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. BETWEEN 22:43:40 ON (B)(6)2024 AND 00:05:36 ON (B)(6)2024, THE PATIENT RECEIVED SEVEN NON-LIFEVEST DEFIBRILLATIONS. THE PATIENT'S LAST SEEN RHYTHM AFTER THE DEFIBRILLATIONS WAS AF/SINUS RHYTHM FROM 90-60 BPM WITH CPR/ELECTRODE LEAD FALL OFF. THE DEVICE WAS SHUT DOWN AT 00:53:18 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177760&pc=FSA | 9681684-2025-00001 | 2024-12-20 05:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-01-17 05:00 | FSA | MAXI SKY 600 | Use of Device Problem | Bone Fracture(s); Laceration(s); Hip Fracture |  |  | 1 | Event Description: ARJO WAS NOTIFIED OF AN INCIDENT INVOLVING A MAXI SKY 600 CEILING LIFT (B)(6) AND AN ARJO LOOP SLING (MLAAS2000-L). IT WAS REPORTED THAT A RESIDENT WAS BEING TRANSFERRED FROM A BED TO A WHEELCHAIR WITH THE ASSISTANCE OF 2 CAREGIVERS. AS THE RESIDENT WAS BEING LIFTED OVER THE WHEELCHAIR, THE RESIDENT SLIPPED OUT OF THE SLING AND FELL (HEADFIRST TO THE FLOOR). THE RESIDENT SUSTAINED VISIBLE LACERATION THEREFORE WAS TAKEN TO THE EMERGENCY DEPARTMENT FOR ASSESSMENT. A COMPUTED TOMOGRAPHY SCAN AND AN X-RAY WERE PERFORMED, WHICH SHOWED NO INJURIES. THE RESIDENT RETURNED TO THE FACILITY THE FOLLOWING DAY. THE RESIDENT'S CONDITION DETERIORATED OVERNIGHT AND SHE WAS SENT BACK TO THE EMERGENCY DEPARTMENT. THE MEDICAL DIRECTOR REVIEWED THE RESIDENT'S MEDICAL RECORDS AND REPORTED THAT THE RESIDENT SUSTAINED A RIGHT HIP FRACTURE AND A LEFT DISTAL CLAVICLE FRACTURE. THESE INJURIES HAD NOT BEEN IDENTIFIED AT THE INITIAL EMERGENCY DEPARTMENT VISIT. THE PATIENT SUFFERED A FAT EMBOLISM SYNDROME. THE RESIDENT PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: ADDITIONAL INFORMATION WILL BE PROVIDED UPON INVESTIGATION CONCLUSION. Manufacturer Narrative: ADDITIONAL INFORMATION WILL BE PROVIDED UPON INVESTIGATION CONCLUSION. Manufacturer Narrative: THE UDI NUMBER IS NOT AVAILABLE AS THE DEVICE WAS MANUFACTURED BEFORE UDI IMPLEMENTATION. IT WAS REPORTED THAT A PATIENT WAS LIFTED FROM A BED TO A WHEELCHAIR IN A SITTING POSITION USING THE MAXI SKY 600 CEILING LIFT AND LOOP SLING (MLAAS200-L, 3000223267023). WHEN THE PATIENT WAS LIFTED OVER THE WHEELCHAIR, THE RESIDENT SLIPPED OUT OF THE SLING AND LANDED ON HER HEAD. THE RESIDENT WAS TAKEN TO THE EMERGENCY DEPARTMENT FOR FURTHER ASSESSMENT OF A SUSTAINED HEAD LACERATION. THE COMPUTED TOMOGRAPHY AND X-RAY WERE UNREMARKABLE. THE RESIDENT'S CONDITION DETERIORATED OVERNIGHT AND SHE WAS RETURNED TO THE EMERGENCY DEPARTMENT. THE MEDICAL DIRECTOR REVIEWED THE RESIDENT'S MEDICAL RECORDS AND REPORTED THAT THE RESIDENT HAD SUSTAINED A RIGHT HIP FRACTURE AND A LEFT DISTAL CLAVICLE FRACTURE. THESE INJURIES HAD NOT BEEN IDENTIFIED AT THE INITIAL EMERGENCY DEPARTMENT VISIT. THE PATIENT SUFFERED A FAT EMBOLISM SYNDROME. THE RESIDENT PASSED AWAY ON (B)(6) 2024. AFTER THE EVENT, THE LIFT AND SLING WERE TESTED AND NO MALFUNCTION WAS FOUND. THE CUSTOMER REPORTED THAT STAFF WERE TRYING TO RECREATED THE EVENT BUT WAS NOT ABLE TO. THE CUSTOMER STAFF STATED THAT THEY ARE NOT ABLE TO EXPLAIN HOW THE PATIENT FELL. THEY STATED THAT "IT HAPPENED IN A BLINK OF AN EYE". IN THE COURSE OF THE INVESTIGATION TWO HYPOTHESIS WERE CONSIDERED: SLING LOOP ACCIDENTALLY REMOVED FROM THE SPREADER BAR OR INCORRECT SIZE OF THE SLING WAS USED. THE SLING LOOP IS ATTACHED TO THE LIFT SPREADER BAR BY THE LOADER LATCH, WHICH IS EQUIPPED WITH A SPRING MECHANISM THAT CLOSES THE LATCH TO PREVENT ACCIDENTAL REMOVAL OF THE LOOP. THEREFORE IF THE LATCH IS CLOSED THE SLING LOOP WILL NOT DETACH FROM THE SPREADER BAR. THE CUSTOMER CONFIRMED THAT THE SLING LOOP DID NOT DETACH FROM THE LIFT SPREADER BAR. THE FIRST HYPOTHESIS CAN THEREFORE BE RULED OUT. IN REGARDS TO INCORRECT SIZE OF THE SLING. THE INSTRUCTION FOR USE (IFU) FOR PASSIVE CLIP SLINGS (04.SL.00) INCLUDES SECTION 'APPLYING THE SLING,' WHICH PROVIDES STEP-BY-STEP INSTRUCTIONS AND GRAPHICAL ILLUSTRATIONS ON HOW TO APPLY SLING ON PATIENT¿S BODY. ADDITIONALLY SECTION "SELECTING SLING SIZE¿, DESCRIBES HOW TO SELECT PROPER SLING SIZE TO FIT THE PATIENT¿S PHYSIQUE. THE IFU INDICATES TO MEASURE THE PATIENT AND THEN FOLLOW THE CHART TO PICK THE CORRECT SIZE. THE CUSTOMER STATED THAT THE SLING HAD CORRECT SIZE, BUT DID NOT PROVIDE INFORMATION HOW THE SLING WAS APPLIED. THIS HYPOTHESIS CANNOT BE CONFIRMED. TAKING THE ABOVE INTO ACCOUNT, THE ROOT CAUSE OF THE REPORTED PATIENT FALL IS IMPOSSIBLE TO DEFINE. TO SUM UP, THE ARJO SLING WAS USED DURING THE RESIDENT TRANSFER. THE LIFT AND SLING SYSTEM IS DESIGNED TO PROVIDE PATIENT TRANSFERS. SINCE THE RESIDENT SLIPPED OUT OF THE DEVICE IT IS CONSIDERED THAT THE LIFT DID NOT MEET ITS PERFORMANCE SPECIFICATION. THE COMPLAINT WAS DECIDED TO BE REPORTABLE DUE TO RESIDENT'S SLIPPING OUT OF THE SLING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177763&pc=FSA | 1419652-2025-00002 | 2024-12-20 05:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-01-17 05:00 | FSA | MAXI SKY 600 | Use of Device Problem | Bone Fracture(s); Laceration(s); Hip Fracture |  |  | 1 | Event Description: ARJO WAS NOTIFIED OF AN INCIDENT INVOLVING A MAXI SKY 600 CEILING LIFT (B)(6) AND AN ARJO LOOP SLING (B)(6). IT WAS REPORTED THAT A RESIDENT WAS BEING TRANSFERRED FROM A BED TO A WHEELCHAIR WITH THE ASSISTANCE OF 2 CAREGIVERS. AS THE RESIDENT WAS BEING LIFTED OVER THE WHEELCHAIR, THE RESIDENT SLIPPED OUT OF THE SLING AND FELL (HEADFIRST TO THE FLOOR). THE RESIDENT SUSTAINED VISIBLE LACERATION THEREFORE WAS TAKEN TO THE EMERGENCY DEPARTMENT FOR ASSESSMENT. A COMPUTED TOMOGRAPHY SCAN AND AN X-RAY WERE PERFORMED, WHICH SHOWED NO INJURIES. THE RESIDENT RETURNED TO THE FACILITY THE FOLLOWING DAY. THE RESIDENT'S CONDITION DETERIORATED OVERNIGHT AND SHE WAS SENT BACK TO THE EMERGENCY DEPARTMENT. THE MEDICAL DIRECTOR REVIEWED THE RESIDENT'S MEDICAL RECORDS AND REPORTED THAT THE RESIDENT SUSTAINED A RIGHT HIP FRACTURE AND A LEFT DISTAL CLAVICLE FRACTURE. THESE INJURIES HAD NOT BEEN IDENTIFIED AT THE INITIAL EMERGENCY DEPARTMENT VISIT. THE PATIENT SUFFERED A FAT EMBOLISM SYNDROME. THE RESIDENT PASSED AWAY ON (B)(6) 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177879&pc=OZP | 2032227-2025-112279 | 2024-12-20 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-17 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Myocardial Infarction | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-326A, MMT-1884, MMT-397A, MMT-7040A, MMT-7841ZN. TROUBLESHOOTING WAS PERFORMED. PRODUCT RETURN STATUS FOR MMT-326A, MMT-1884 IS UNKNOWN. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. PRODUCT RETURN STATUS FOR MMT-7040A IS UNKNOWN. PRODUCT RETURN STATUS FOR MMT-7841ZN IS UNKNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21177972&pc=BZD | 2518422-2025-004193 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LIVER AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178000&pc=DSQ | 2916596-2025-00343 | 2024-04-01 04:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE ARTICLE, "DEGREE OF PRE-IMPLANT LV DILATION DOES NOT PRECLUDE FAVORABLE OUTCOME POST LVAD WEANING: A VAD WEAN REGISTRY ANALYSIS", THAT MYOCARDIAL RECOVERY DURING SUPPORT WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS MORE LIKELY IN PATIENTS WITH A SMALLER VENTRICLE SIZE AT TIME OF IMPLANT. HOWEVER, IT IS UNKNOWN IF LONG TERM DURABILITY OF RESPONSE IS INFERIOR FOR THOSE PATIENTS WITH LARGER LEFT VENTRICLE SIZE THAT ENOUGH HAD REVERSE REMODELED DURING LVAD SUPPORT AND WERE ABLE TO HAVE MYOCARDIAL RECOVERY WEANING. SURVIVAL RATES WERE COMPARED BETWEEN PATIENTS WITH BASELINE LEFT VENTRICULAR INTERNAL DIASTOLIC DIMENSIONS (LVIDD) =65MM AND THOSE WITH LVIDD <65MM. A LINEAR REGRESSION WAS USED TO CORRELATE THE BASELINE LVIDD WITH TIME FROM IMPLANT TO EXPLANT, AS WELL AS LVIDD CHANGE FROM BASELINE TO PRIOR TO EXPLANT. OF THE 318 PATIENTS STUDIED, 22 WERE REPORTED TO HAVE A HEARTMATE 3 LVAD. OF THE REMAINING PATIENTS, THERE WERE 169 HEARTMATE II PATIENTS, 97 HEARTWARE VENTRICULAR ASSIST DEVICE (HVAD) PATIENTS, AND 30 WITH PRODUCT TYPE "OTHER". THE PATIENTS WITH LVIDD > 65 MM WERE YOUNGER, MORE LIKELY TO HAVE BEEN MALE, HAD LESS ATRIAL FIBRILLATION, AND MORE HAD AN IMPLANTABLE DEFIBRILLATOR. LVIDD DID NOT CORRELATE WITH TIME FROM IMPLANT TO EXPLANT (R =0.11, P = 0.073). A LARGER LVIDD CORRELATED WITH GREATER CHANGE IN LVIDD FROM PRE-LVAD TO PRE-EXPLANT (R = 0.7, P < 0.001). POST LVAD WEANING SURVIVAL WAS NOT DIFFERENT BETWEEN THOSE WITH LVIDD = 65MM OR THOSE WITH < 65MM. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. INVOLVED HOSPITALS: MEDICINE, UNIVERSITY OF PENNSYLVANIA, PHILADELPHIA, PA; UNIVERSITY OF UTAH, SALT LAKE CITY, UT; MONTEFIORE-EINSTEIN, NEW YORK, NY; WEILL CORNELL MEDICINE-NEW YORK PRESBYTERIAN, NEW YORK, NY; MINNEAPOLIS HEART INSTITUTE, MINNEAPOLIS, MN; HOUSTON METHODIST HOSPITAL, HOUSTON, TX; MEDICAL UNIVERSITY OF VIENNA, VIENNA; UNIVERSITY OF DUISBURG-ESSEN, WEST GERMAN HEART AND VASCULAR CENTER, ESSEN; UNIVERSITY OF PENNSYLVANIA, PHILADELPHIA, PA. SCHURR, J., DRAKOS, S., PATEL, S., SOBOL, I., SUN, B., FIDA, N., ZIMPFER, D., PAPATHANASIOU, M., & WALD, J. (2024). DEGREE OF PRE-IMPLANT LV DILATION DOES NOT PRECLUDE FAVORABLE OUTCOME POST LVAD WEANING: A VAD WEAN REGISTRY ANALYSIS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S172¿S173. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.349. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178286&pc=MVK | 3008642652-2025-00602 | 2024-11-28 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 17:57:56 ON (B)(6)2024. AT 10:20:03 ON (B)(6)2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 10:20:45, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 10:40:05, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 10:40:43, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 10:40:47 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178303&pc=MVK | 3008642652-2025-00604 | 2024-10-15 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 22:00:55 ON 1(B)(6)2024. AT 13:42:21 ON (B)(6)024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 60 BPM DEGRADING TO VF. AT 13:43:09, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 13:49:39, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 60 BPM DEGRADING TO VF. AT 13:50:23, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 13:52:23, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 13:53:41, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 14:16:08 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178326&pc=BZD | 2518422-2025-004300 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178568&pc=BZD | 2518422-2025-004320 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178607&pc=DQY | 2124215-2025-01793 | 2025-01-14 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | DQY | WATCHMAN FXD CURVE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion; Retroperitoneal Hemorrhage | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PERICARDIAL EFFUSION. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN PERFORMED THE TRANSSEPTAL PUNCTURE USING THE VERSACROSS CONNECT AND FXD CURVE WAS. FOLLOWING THE PUNCTURE THE PHYSICIAN NOTED THAT THERE WAS A DECREASE IN SYSTOLIC BLOOD PRESSURE. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING SHOWED THAT THERE WAS A PERICARDIAL EFFUSION PRESENT. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS DRAINING BLOOD FROM THE PATIENT. THE PATIENT STABILIZED AND THE PROCEDURE WAS CONTINUED. THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. POST PROCEDURE THE PERICARDIOCENTESIS DRAIN NEEDED TO BE EXCHANGED FOR A LARGER DRAIN AND MORE FLUID WAS REMOVED FROM THE PATIENT. THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY AND WILL BE MONITORED OVERNIGHT.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PERICARDIAL EFFUSION. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN PERFORMED THE TRANSSEPTAL PUNCTURE USING THE VERSACROSS CONNECT AND FXD CURVE WAS. FOLLOWING THE PUNCTURE THE PHYSICIAN NOTED THAT THERE WAS A DECREASE IN SYSTOLIC BLOOD PRESSURE. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING SHOWED THAT THERE WAS A PERICARDIAL EFFUSION PRESENT. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS DRAINING BLOOD FROM THE PATIENT. THE PATIENT STABILIZED AND THE PROCEDURE WAS CONTINUED. THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. POST PROCEDURE THE PERICARDIOCENTESIS DRAIN NEEDED TO BE EXCHANGED FOR A LARGER DRAIN AND MORE FLUID WAS REMOVED FROM THE PATIENT. THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY AND WILL BE MONITORED OVERNIGHT. IT WAS FURTHER REPORTED THAT THE PATIENT DIED AS A RESULT OF THIS EVENT. THE BLEEDING FROM THE EFFUSION DID STOP BUT THEY SUSPECTED A CLOT WAS COMPRESSING THE HEART. THE PATIENT ALSO HAD ALL KINDS OF OTHER ISSUES WITH SEVERE HYPERKALEMIA AND THEN ASPIRATED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PERICARDIAL EFFUSION. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN PERFORMED THE TRANSSEPTAL PUNCTURE USING THE VERSACROSS CONNECT AND FXD CURVE WAS. FOLLOWING THE PUNCTURE THE PHYSICIAN NOTED THAT THERE WAS A DECREASE IN SYSTOLIC BLOOD PRESSURE. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) IMAGING SHOWED THAT THERE WAS A PERICARDIAL EFFUSION PRESENT. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS DRAINING BLOOD FROM THE PATIENT. THE PATIENT STABILIZED AND THE PROCEDURE WAS CONTINUED. THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. POST PROCEDURE THE PERICARDIOCENTESIS DRAIN NEEDED TO BE EXCHANGED FOR A LARGER DRAIN AND MORE FLUID WAS REMOVED FROM THE PATIENT. THE PATIENT IS EXPECTED TO MAKE A FULL RECOVERY AND WILL BE MONITORED OVERNIGHT. IT WAS FURTHER REPORTED THAT THE PATIENT DIED AS A RESULT OF THIS EVENT. THE BLEEDING FROM THE EFFUSION DID STOP BUT THEY SUSPECTED A CLOT WAS COMPRESSING THE HEART. THE PATIENT ALSO HAD ALL KINDS OF OTHER ISSUES WITH SEVERE HYPERKALEMIA AND THEN ASPIRATED. THE PHYSICIAN REPORTED THAT THE CAUSE OF DEATH WAS RETROPERITONEAL BLEEDING. Manufacturer Narrative: H6 PATIENT CODES: RETROPERITONEAL HEMORRHAGE E1027 ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178638&pc=OZD | 1220648-2025-25638 | 2024-08-31 04:00 | Death | ABIOMED, INC. | 2025-01-17 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED HEMOLYSIS WITH A "DEFINITE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿THE PATIENT DEVELOPED HEMOLYSIS. ADMISSION TOTAL HEMOGLOBIN 11.1 G/DL, NADIR 7.2 G/DL ((B)(6) 2024); PEAK LDH 1314 U/L ((B)(6) 2024); ADMISSION TOTAL BILIRUBIN 0.9 MG/DL, PEAK 1.5 MG/DL ((B)(6) 2024).¿. THE PATIENT RECOVERED WITH SEQUELAE. LOQI ALSO REPORTED THAT THE PATIENT ENDURED BLEEDING AT IMPELLA CP ACCESS SITE RIGHT FEMORAL ARTERY WITH A "DEFINITE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿THE PATIENT DEVELOPED BLEEDING FROM THE RIGHT FEMORAL ARTERY AT THE IMPELLA CP ACCESS SITE. ADMISSION TOTAL HGB 11.1 G/DL, NADIR 7.2 ((B)(6) 2024). HEPARIN WAS HELD DUE TO BLEEDING, FEMOSTOP WAS APPLIED, AND DRESSINGS WERE CHANGED MULTIPLE TIMES.¿. THE PATIENT RECOVERED WITH NO SEQUELAE. IT WAS NOTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUES HAS BEEN COMPLETED. DEVICE HISTORY RECORD REVIEW CONFIRMED THAT THE PUMP PASSED ALL POST STERILE INSPECTION CHECKS. DATA LOGS WERE REVIEWED, AND IT WAS CONFIRMED THAT NO SUCTION ALARMS OCCURRED. NO MOTOR CURRENT SPIKE WAS OBSERVED OUT OF P-LEVEL CHANGE. MANY IMPELLA POSITION IN AORTA ALARMS WERE OBSERVED AROUND THE TIME ISSUE REPORTED FOLLOWED HIGH PURGE PRESSURE ALARMS AND FLOW WAS ABOUT 3L/MIN AT P-2. THE ROOT CAUSE OF THE REPORTED ISSUES CANNOT BE DETERMINED SINCE THE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT DATA PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178851&pc=DSQ | 2916596-2025-00237 | 2021-02-01 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿PREOPERATIVE PASSIVE VENOUS PRESSURE-DRIVEN CARDIAC FUNCTION DETERMINES LEFT VENTRICULAR ASSIST DEVICE OUTCOMES¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT VENTRICULAR FAILURE, BLEEDING, STROKE, DEATH, AND TRANSPLANT. THIS SINGLE-CENTER RETROSPECTIVE STUDY EVALUATED 139 PATIENTS WHO WERE IMPLANTED WITH HM3 BETWEEN 12FEB2012 TO 10FEB2021. DATA FOR THIS STUDY WAS RECOVERED THROUGH DETAILED CLINICAL FOLLOW-UP. THE STUDY HYPOTHESIZED BY SEPARATING RIGHT-SIDED OUTPUT INTO PASSIVE (VENOUS PRESSURE) AND ACTIVE (RV CONTRACTION) COMPONENTS, THEY COULD BETTER PREDICT THE DEGREE OF END ORGAN DAMAGE FROM BACKPRESSURE AND THUS GAIN INSIGHT INTO LEFT VENTRICULAR ASSIST DEVICE (LVAD) SURVIVAL AS WELL AS ACTUAL RIGHT VENTRICULAR CONTRACTILITY. THE CONCEPT OF PASSIVE VENOUS OR FONTAN TYPE CIRCULATION WAS USED AS A PREDICTOR OF RVF AND LVAD MORTALITY IN THIS STUDY. PRE-LVAD CARDIAC OUTPUT DRIVEN BY VENOUS PRESSURE WAS DETERMINED BY DIVIDING RIGHT ATRIAL PRESSURE BY MEAN PULMONARY ARTERY PRESSURE, MULTIPLIED BY TOTAL CARDIAC OUTPUT. NORMALIZATION TO BODY SURFACE AREA LED TO THE PASSIVE CARDIAC INDEX (PASCI). THE YOUDEN J STATISTIC WAS THEN USED TO LOCATE THE PASCI THRESHOLD, IN ORDER TO BETTER PREDICT LVAD DEATH. THE STUDY FOUND THE THRESHOLD TO BE 0.5. RELEVANT PATIENT INFORMATION FOR THIS STUDY WAS DIVIDED BY WHERE THE HM3 PATIENTS FELL UPON THE PASCI THRESHOLD. THERE WERE 63 (49 MALE, 14 FEMALE) HM3 PATIENTS WITH A PASCI <0.5 AND 76 (35.3 MALE, 40.7 FEMALE) WITH PASCI >= 0.5 (P=0.720). THE AVERAGE AGE OF THIS GROUP WAS 53.48 FOR PATIENTS WITH PASCL <0.5 AND 54.54 WITH PASCI >= 0.5 (P=0.663). THE AVERAGE WEIGHT WAS 83.00 KG FOR PASCI <.5 AND 91.40 KG FOR PASCI >= 0.5 (P=0.07). OUT OF THE TOTAL 139 EVALUATED PATIENTS, 7 HM3 PATIENTS SUFFERED POST OPERATION RVF (P = 0.090). SEVEN PATIENTS WERE IMPLANTED WITH A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) POST LVAD IMPLANTATION (P = 0.090). 11 PATIENTS SUFFERED A STROKE/TIA (P = 0.993). 21 PATIENTS EXPERIENCED GASTROINTESTINAL BLEEDING (P = 0.231). FIVE PATIENTS DIED WITHIN 30 DAYS OF IMPLANT DATE (P = 0.38). IN TOTAL, 18 PATIENTS DIED AFTER LVAD IMPLANTATION (P = 0.009). 36 PATIENTS UNDERWENT HEART TRANSPLANT (P = 0.069). THE STUDY CONCLUDED THAT RAISING VENOUS PRESSURE NEGATIVELY IMPACTS END-ORGAN FUNCTION AND ALSO INCREASES HEART FAILURE READMISSION DAYS. PATIENTS WHO HAD A PRE-LVAD PASCI >= 0.5 HAD WORSE POST-LVAD SURVIVAL AND INCREASED RVF. PASCI METRIC MAY IMPROVE THE MANAGEMENT OF LVAD CANDIDATES WITH RIGHT VENTRICLE DYSFUNCTION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01FEB2021 AS DATA WAS COLLECTED BETWEEN 12FEB2012 AND 10FEB2021. AUTHOR INFORMATION: TANG, P. C., MILLAR, J., NOLY, P. E., SICIM, H., LIKOSKY, D. S., ZHANG, M., PAGANI, F. D., & MICHIGAN CONGESTIVE HEART FAILURE INVESTIGATORS (2024). PREOPERATIVE PASSIVE VENOUS PRESSURE-DRIVEN CARDIAC FUNCTION DETERMINES LEFT VENTRICULAR ASSIST DEVICE OUTCOMES. THE JOURNAL OF THORACIC AND CARDIOVASCULAR SURGERY, 168(1), 133¿144.E5. HTTPS://DOI.ORG/10.1016/J.JTCVS.2023.07.019. UNIVERSITY OF MICHIGAN FRANKEL CARDIOVASCULAR CENTER, ANN ARBOR, MICHIGAN. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179104&pc=CAW | 21179104 | 2025-01-04 05:00 | Death | COMPASS HEALTH BRANDS | 2025-01-17 05:00 | CAW | COMPASS TRUAIRE | Adverse Event Without Identified Device or Use Problem | Partial thickness (Second Degree) Burn |  |  | 1 | Event Description: PATIENT SUFFERED FLASH BURN TO FACE WHILE SMOKING AND USING OXYGEN CONCENTRATOR. PATIENT SUSTAINED 2ND DEGREE BURNS TO FACE. PATIENT WAS TRANSFERRED TO TRAUMA CENTER AND WAS PLACED ON A VENTILATOR. PATIENT WAS EXTUBATED SAME DAY AND SUBSEQUENTLY DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179117&pc=OZD | 1220648-2025-25643 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-01-17 05:00 | OZD | IMPELLA CP | Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT HAD AN ATTEMPTED IMPELLA CP IMPLANT BUT THE WIRE KEPT COMING OUT OF THE LEFT VENTRICLE DURING CARDIOPULMONARY RESUSCITATION (CPR). THE PATIENT WENT INTO PULSELESS ELECTRONIC ACTIVITY AND ALL ATTEMPTS TO SAVE THE PATIENT STOPPED. THE PATIENT NEVER REGAINED A PULSE AND THE DOCTOR WAS UNABLE TO RECROSS THE AORTIC VALVE. THE PATIENT EXPIRED AFTER 15-20 MINUTES OF CPR. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, THE IMPELLA WAS UNABLE TO PASS THROUGH THE VASCULATURE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK. PERCUTANEOUS CORONARY INTERVENTION. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179196&pc=QEZ | 1319211-2024-00076 | 2024-12-23 05:00 | Death | ANGIODYNAMICS, INC. | 2025-01-17 05:00 | QEZ | ALPHAVAC | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | K240397 |  | 1 | Event Description: A TERRITORY MANAGER (TM) REPORTED AN ADVERSE EVENT DURING A PROCEDURE USING AN ALPHAVAC MULTIPURPOSE MECHANICAL ASPIRATOR F18 C85. THE FOLLOWING WAS REPORTED: " A PATIENT EXPIRED FROM, WHAT IS BELIEVED TO BE, A PERICARDIAL EFFUSION. AFTER REMOVING MATERIAL FROM THE RIGHT SIDE OF THE LUNG AND SWITCHING TO THE LEFT, THE PHYSICIAN LOST ACCESS AND THE CANNULA FELL OUT OF THE LUNGS AND INTO THE RVOT/RV. THE PHYSICIAN TRIED TO REGAIN ENTRY INTO THE MAIN PULMONARY ARTERY BY ADVANCING THE CANNULA. WHEN HE WAS STEERING IN THE HEART, THE FUNNEL DIDN'T HAVE THE FULL BEND OF THE CANNULA PLUS 2 CM. IT WAS BARELY EXPOSED OUT OF THE OUTER SHEATH. THE TM RECOMMENDED THEY REMOVE THE WHOLE SYSTEM AND START OVER WITH THE SWAN OR SOME OTHER DEVICE. THE PHYSICIAN STEERED AROUND FOR JUST A LITTLE AND TOOK A PICTURE WHEN THE PERICARDIAL EFFUSION WAS NOTED. ONCE CONTRAST WAS SHOT INTO THE CANNULA AND THE PERICARDIAL EFFUSION WAS SHOWING, THE PATIENT STARTED DECOMPENSATING. AFTER AN HOUR OF CPR AND EPI ADMINISTERED, THE PATIENT EXPIRED. IT WAS THOUGHT THAT BECAUSE THE CANNULA WAS BARELY EXPOSED OUT OF THE OUTER SHEATH AND THE PHYSICIAN WAS TRYING TO REGAIN ACCESS, THAT THIS MAY HAVE CONTRIBUTED TO THE EVENT; HOWEVER, IT WAS LIKELY RELATED TO THE PROCEDURE ITSELF AND SOME PATIENTS NOT BEING ABLE TO HANDLE A LARGE BORE SYSTEM. THE PHYSICIAN HAS BEEN IN-SERVICED A FEW TIMES FROM THE TM FOR THIS DEVICE AND HAS BEEN USING INARI FOR A LONG TIME. " Manufacturer Narrative: THE REPORTED DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4). Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT EXPIRED FROM A PERICARDIAL EFFUSION DURING THE ALPHAVAC PROCEDURE IS NOT CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). THERE WERE NO REPORTS OF ALPHAVAC DEVICE MALFUNCTION DURING THE PROCEDURE, THEREFORE, NO DEVICE WAS RETURNED FOR EVALUATION. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. THERE HAVE BEEN NO OTHER REPORTED COMPLAINTS AGAINST THE INDICATED PACKAGING/CANNULA LOTS FOR SIMILAR VESSEL PERFORATION ISSUE. IT WAS REPORTED THAT THE CANNULA FELL OUT OF THE LUNGS AND INTO THE RVOT/RV. THE PHYSICIAN TRIED TO REGAIN ENTRY INTO THE MAIN PULMONARY ARTERY BY ADVANCING THE CANNULA, HOWEVER, THE OBTURATOR AND GUIDEWIRE WERE NOT IN THE SHEATH/CANNULA. THE TERRITORY MANAGER WAS PRESENT FOR THIS CASE AND RECOMMENDED THEY REMOVE THE WHOLE SYSTEM AND START OVER WITH THE SWAN OR SOME OTHER DEVICE. PER DIRECTIONS FOR USE (DFU): "CAUTION: THE ALPHAVAC SHEATH SHOULD NOT BE ADVANCED/NAVIGATED UNLESS THE OBTURATOR AND GUIDEWIRE ARE IN PLACE." IT IS UNKNOWN IF THIS ADVANCEMENT WAS THE ROOT CAUSE FOR THE PERICARDIAL EFFUSION. THE PHYSICIAN WAS ASKED IF THERE WAS ANY CAUSAL RELATIONSHIP BETWEEN THE ALPHAVAC DEVICE AND THE PATIENT'S EXPIRATION AND THEIR RESPONSE WAS IT'S POSSIBLE, BUT THINKS IT'S LIKELY RELATED TO THE PROCEDURE ITSELF SINCE SOME PATIENTS NOT BEING ABLE TO HANDLE A LARGE BORE SYSTEM. CORRECTION/CORRECTIVE ACTIONS: NO CORRECTION IS REQUIRED SINCE NO ALPHAVAC DEVICE WAS RETURNED TO ANGIODYNAMICS FOR EVALUATION - THERE WAS NO REPORT OF DEVICE MALFUNCTION DURING THE PROCEDURE. THE ROOT CAUSE OF THE PERICARDIAL EFFUSION COULD NOT BE DEFINITIVELY DETERMINED BUT IS POTENTIALLY RELATED TO THE PROCEDURE ITSELF SINCE SOME PATIENTS NOT BEING ABLE TO HANDLE A LARGE BORE SYSTEM. PERICARDIAL EFFUSION AND DEATH ARE POTENTIAL ADVERSE EVENTS OF EMBOLECTOMY SYSTEM USE; THIS IS CAUTIONED IN THE DIRECTIONS FOR USE. THIS EVENT IS CAPTURED IN THE POST MARKET SURVEILLANCE OF THE ALPHAVAC PRODUCT FAMILY. BASED ON NO ADVERSE TRENDS, NO ISSUES NOTED IN THE DEVICE HISTORY RECORD REVIEW, AND ADEQUATE PROCESS CONTROLS IN PLACE, NO CORRECTIVE ACTION TO BE TAKEN BY ANGIODYNAMICS AT THIS TIME. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE: THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. WARNINGS: SELECTION OF THE PATIENT AS A CANDIDATE FOR USE WITH THIS DEVICE AND FOR SUCH PROCEDURES AS IT IS INTENDED IS THE PHYSICIANS' SOLE RESPONSIBILITY. THE OUTCOME IS DEPENDENT ON MANY VARIABLES INCLUDING, PATIENT PATHOLOGY, SURGICAL PROCEDURE, AND PROCEDURE/TECHNIQUE. THE BENEFITS OF USE OF THIS DEVICE MUST BE WEIGHED AGAINST THE RISKS INCLUDING RISKS OF SYSTEMIC ANTICOAGULATION AND MUST BE ASSESSED BY THE PRESCRIBING PHYSICIAN. PRECAUTIONS: CANNULA PLACEMENT AND POSITIONING SHOULD BE GUIDED AND CONFIRMED USING STANDARD GUIDEWIRE, FLUOROSCOPIC, AND OTHER APPROPRIATE IMAGING TECHNIQUES. TO MINIMIZE THE POTENTIAL FOR VASCULAR TRAUMA, ENSURE THAT THE CANNULA IS PLACED IN AN APPROPRIATELY SIZED VESSEL. CAUTION: THE ALPHAVAC SHEATH SHOULD NOT BE ADVANCED/NAVIGATED UNLESS THE OBTURATOR AND GUIDEWIRE ARE IN PLACE. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: CARDIAC ARREST, CARDIAC PERFORATION, DEATH, PERFORATION, PERICARDIAL EFFUSION, PULMONARY EMBOLISM. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179207&pc=MVK | 3008642652-2025-00612 | 2024-10-17 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 18:26:55 ON (B)(6) 2024. AT 03:01:17 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 03:19:24 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179293&pc=OZD | 1220648-2025-25642 | 2024-12-21 05:00 | Death | ABIOMED, INC. | 2025-01-17 05:00 | OZD | IMPELLA 5.5 | Patient-Device Incompatibility | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT. WHILE IN THE INTENSIVE CARE UNIT, AN ECHOCARDIOGRAM WAS DONE AND NOTED TAMPONADE. THE PATIENT WAS EMERGENTLY TAKEN BACK TO CARDIOVASCULAR OPERATING ROOM FOR CHEST EXPLORATION AND DIRECT AORTIC APPROACH 5.5. THE 0.035 WIRE WAS REMOVED AND WAS REPLACED WITH 0.018 WIRE. DIAGNOSTIC CATHETER WAS REMOVED AND THE IMPELLA WAS BACKLOADED ONTO 0.018 GUIDEWIRE AND PLACED INTO LEFT VENTRICLE USING DOUBLE BARREL TECHNIQUE. AFTER SUCCESSFUL REPAIR OF LEFT VENTRICLE PERFORATION AND SUCCESSFUL IMPLANT OF IMPELLA 5.5 DIRECT AORTIC APPROACH CHEST PACKED STERNUM AND SKIN CLOSED. THE PATIENT EVENTUALLY EXPIRED AFTER CARE WAS WITHDRAWN. Manufacturer Narrative: THE DEVICE OR THE DATA LOGS WERE NOT RETURNED FOR ANALYSIS. THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE THE INSERTION WAS ABORTED DUE TO SMALL ANATOMY OF THE PATIENT. THE ROOT CAUSE OF VENTRICLE PERFORATION IS NOT DETERMINED BECAUSE THE TIMING OF THE EVENT IS UNCLEAR AS PER THE PROVIDED CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER: SECTION: DIRECT AORTIC INSERTION: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST: SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179461&pc=BZD | 2518422-2025-100824 | 2024-08-27 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | REMSTAR PRO C-FLEX+ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179507&pc=OZD | 1220648-2025-25645 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-17 05:00 | OZD | IMPELLA CP | Increase in Pressure | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP DEVICE WAS PRE-OPERATIVELY IMPLANTED IN A PATIENT FOR MECHANICAL CIRCULATORY SUPPORT DURING A HIGH-RISK SURGERY AND HIGH PURGE PRESSURE WAS ENCOUNTERED. THE PATIENT WAS PRESENTED TO THE HOSPITAL DUE TO SUSPECTED ACUTE CORONARY SYNDROME AND MULTIPLE BRANCH LESIONS. VENTRICULAR SEPTAL PERFORATION WAS DISCOVERED AND THE IMPELLA CP DEVICE WAS IMPLANTED. IMMEDIATELY AFTER THE IMPELLA CP PUMP WAS TURNED ON, THE PURGE BLOCKAGE ALARM SOUNDED. THE PURGE CASSETTE WAS REPLACED WITHOUT RESOLUTION. CONSEQUENTLY, THE IMPELLA CP DEVICE WAS EXPLANTED AND REPLACED. THE PATIENT REMAINED ON SUPPORT FOR APPROXIMATELY NINE DAYS BEFORE EXPIRING DUE TO MULTIPLE ORGAN FAILURE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE FIRST 30 MINUTES TEMPORARY IMPEDE OF ADEQUATE SUPPORT WITH THE INITIALLY PLACED PUMP CANNOT BE RULED OUT AS A CONTRIBUTING FACTOR. HOWEVER, IT IS NOTED, RECOGNIZED THE REPLACEMENT PUMP WAS INSERTED (USED WITHOUT ISSUES), AND THE PATIENT SUBSEQUENTLY DIED NINE DAYS LATER BECAUSE OF MULTI-ORGAN FAILURE DUE TO OTHER PREMORBID CONDITIONS. Manufacturer Narrative: THE INVESTIGATION FOR THE HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE DATA LOGS INDICATE THERE WERE MOTOR CURRENT SPIKES ONE MINUTE AFTER STARTING THE PUMP AND THE PURGE PRESSURE IMMEDIATELY RISES. IT RISES FOR 8 MINUTES UNTIL IT TRIGGERS HIGH PURGE PRESSURE ALARMS. TROUBLESHOOTING IS DONE BY REPLACING THE CASSETTE, BUT IT REMAINS HIGH. THE CASE LASTS FOR 32 MINUTES AND THEN THE PUMP IS EXPLANTED. DATA LOGS CONFIRM PRIMING WAS COMPLETE AND THE PUMP IS STARTED A MINUTE AFTER. THE RETURNED PUMP HAD BIOMATERIAL IN THE PURGE GAP. HIGH PURGE PRESSURE REPRODUCED. THERE WAS A BLOCKAGE ISOLATED TO THE MOTOR AND THE PURGE LINE WAS SEEN TO BE AN AMBER COLOR INDICATING THERE WAS BLOOD IN THE LINE. THE PUMP WAS TORN DOWN AND BIOMATERIAL WAS SEEN INSIDE THE MOTOR. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE IS DUE TO BIOMATERIAL BASED ON DATA LOG AND PRODUCT ANALYSIS. INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: TROUBLESHOOTING THE PURGE SYSTEM: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179520&pc=DSQ | 2916596-2025-00366 | 2021-10-01 04:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿IMPACT OF HEART TRANSPLANT ALLOCATION CHANGE ON WAITLIST MORTALITY AND POSTTRANSPLANT MORTALITY IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES¿ THAT VIEWED THE EFFECT ON WAITLIST MORTALITY RATE THAT THE CHANGES TO THE HEART TRANSPLANT ALLOCATION SYSTEM HAD ON LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS. THERE WAS A TOTAL OF 4,406 LVAD PATIENTS BETWEEN OCTOBER 2016 AND OCTOBER 2021 LOOKED AT THAT WERE DIVIDED INTO TWO GROUPS, LISTING GROUP AND TRANSPLANT GROUP. PATIENTS COULD BE IN EITHER OR BOTH GROUPS. THE LISTING GROUP CONTAINED 3,835 PATIENTS (1,232 WERE THE OLD LISTING COHORT AND 2,603 WERE THE NEW LISTING COHORT) AND THE TRANSPLANT GROUP CONTAINED 3,486 PATIENTS. OF NOTE THERE WERE 3 LVAD DEVICES IN THE STUDY HEARTWARE HVAD, HEARTMATE II AND HEARTMATE 3. OF THE LISTING GROUP 1,515 PATIENTS HAD HEARTMATE 3 AND IN THE TRANSPLANT GROUP 891 PATIENTS HAD HEARTMATE 3. THE WAITLIST MORTALITY RATES WERE 50 DEATHS FOR THE OLD LISTING COHORT AND 87 DEATHS FOR THE NEW LISTING COHORT. LVAD BRAND WAS SIGNIFICANTLY ASSOCIATED WITH WAITLIST MORTALITY (OVERALL P=0.004). COMPARED WITH PATIENTS WITH HEARTMATE II, PATIENTS WITH HEARTMATE 3 HAD A SIGNIFICANTLY LOWER RISK OF WAITLIST MORTALITY (HR, 0.38 95% CI, 0.21¿0.69]; P=0.002). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. TIBREWALA, A., CHUZI, S., WU, T., BALDRIDGE, A. S., HARAP, R., BRYNER, B., . . . WILCOX, J. E. (2024). IMPACT OF HEART TRANSPLANT ALLOCATION CHANGE ON WAITLIST MORTALITY AND POSTTRANSPLANT MORTALITY IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. CIRCULATION.HEART FAILURE, 17(11), E011621. DOI:HTTP://DX.DOI.ORG/10.1161/CIRCHEARTFAILURE.124.011621. DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE (A.T., S.C., J.E.W.), NORTHWESTERN UNIVERSITY, CHICAGO, IL. DIVISION OF CARDIAC SURGERY, DEPARTMENT OF SURGERY (T.W., A.S.B., R.H., B.B., D.T.P.), NORTHWESTERN UNIVERSITY, CHICAGO, IL. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179676&pc=DSQ | 2916596-2025-00326 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO HOSPITAL POLICY, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179677&pc=NIQ | 2124215-2024-82121 | 2023-11-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY. IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE 1ST OBTUSE MARGINAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE 1ST DIAGONAL WITH 80% STENOSIS AND WAS 39 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM AND 2.50 MM X 20 MM SYNERGY STENT SYSTEMS. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. ONE WEEK LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN. IN (B)(6) 2023, THE SUBJECT DIED. AT THE TIME OF THE EVENT THE SUBJECT WAS ON ASPIRIN, AND OTHER ACTION WAS TAKEN TO TREAT THE EVENT. THE PRIMARY CAUSE OF DEATH IS UNKNOWN, AND IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179678&pc=NIQ | 2124215-2024-82118 | 2023-11-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY. IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE 1ST OBTUSE MARGINAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE 1ST DIAGONAL WITH 80% STENOSIS AND WAS 39 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM AND 2.50 MM X 20 MM SYNERGY STENT SYSTEMS. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. ONE WEEK LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN. IN (B)(6) 2023, THE SUBJECT DIED. AT THE TIME OF THE EVENT THE SUBJECT WAS ON ASPIRIN, AND OTHER ACTION WAS TAKEN TO TREAT THE EVENT. THE PRIMARY CAUSE OF DEATH IS UNKNOWN, AND IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179684&pc=NIQ | 2124215-2024-82123 | 2023-11-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE 1ST OBTUSE MARGINAL WITH 90% STENOSIS AND WAS 21 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. THE SECOND TARGET LESION WAS LOCATED IN THE 1ST DIAGONAL WITH 80% STENOSIS AND WAS 39 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 24 MM AND 2.50 MM X 20 MM SYNERGY STENT SYSTEMS. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. ONE WEEK LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN. IN (B)(6) 2023, THE SUBJECT DIED. AT THE TIME OF THE EVENT THE SUBJECT WAS ON ASPIRIN, AND OTHER ACTION WAS TAKEN TO TREAT THE EVENT. THE PRIMARY CAUSE OF DEATH IS UNKNOWN, AND IT IS UNKNOWN IF AUTOPSY WAS PERFORMED OR NOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179705&pc=MVK | 3008642652-2025-00613 | 2024-09-28 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-17 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 19:55:33 ON (B)(6) 2024. AT 006:59:27 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 07:00:38, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 07:10:38 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179947&pc=MKJ | 1218058-2025-00003 | 2024-11-23 05:00 | Death | BIO-DETEK INCORPORATED | 2025-01-17 05:00 | MKJ | CPR-D PADZ | Patient Device Interaction Problem | Cardiac Arrest | P160015 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 75-YEAR-OLD MALE PATIENT, THE CPR-D PADZ WOULD NOT ADHERE TO THE PATIENT'S SKIN. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE CPR-D PADZ FROM THE EVENT WERE RETURNED TO ZOLL PAWTUCKET FOR EVALUATION. THE LEVEL OF ADHESION COULD NOT BE EVALUATED DUE TO THE CONDITION OF THE PADS. UPON VISUAL INSPECTION THERE WERE PORTIONS OF THE APEX AND STERNUM PAD THAT WERE COVERED WITH HAIR. THE IFU FOR THE CPR-D PADZ (AW R2001-96 REV L) CAUTIONS THE USER TO REMOVE EXCESSIVE CHEST HAIR. THE PADS WERE SCRAPPED AFTER TESTING. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: THIS PRODUCT HAS A UDI; SOME INFORMATION WAS NOT PROVIDED AND IS UNKNOWN; THEREFORE, A COMPLETE UDI CANNOT BE PROVIDED. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180123&pc=NIQ | 2124215-2024-82126 | 2024-03-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY. IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2019, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE MIDDLE LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH 50% STENOSIS AND WAS 20 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 3.50 MM. THE TARGET LESION WAS TREATED WITH PRE-DILATION AND FOLLOWED BY THE PLACEMENT OF A 3.50 MM X 32 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. SIX DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN. IN (B)(6) 2024, THE SUBJECT DIED, AND THE PRIMARY CAUSE OF DEATH IS UNKNOWN. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN. THERE WAS NO ACTION TAKEN TO TREAT THE EVENT AND IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: B2 - DATE OF EVENT: USED THE FIRST DATE OF THE MONTH OF THE EVENT DATE AS NO EXACT DATE WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180124&pc=NIQ | 2124215-2024-83746 | 2022-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY: IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE MIDDLE RIGHT CORONARY ARTERY (RCA) WITH 90% STENOSIS AND WAS 60 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE TARGET LESION #1 WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.50 MM X 38 MM SYNERGY STENTS SYSTEM OVERLAPPED WITH A 2.50 MM X 24 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. FOUR DAYS AFTER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2022, THE SUBJECT DIED. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. MEDICATION WAS GIVEN TO TREAT THE EVENT. THE PRIMARY CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: B3: DATE OF EVENT: USED FIRST DATE OF THE MONTH SINCE EXACT EVENT DATE WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180126&pc=NGV | 2124215-2024-83870 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 6 | Event Description: THIS REPORT SUMMARIZES 6 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 69 TO 83 YEARS. PATIENT SEX: 3 FEMALE; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 441 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180138&pc=NIQ | 2124215-2024-83745 | 2022-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY: IN (B)(6) 2019, THE SUBJECT PRESENTED WITH UNSTABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE MIDDLE RIGHT CORONARY ARTERY (RCA) WITH 90% STENOSIS AND WAS 60 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.5 MM. THE TARGET LESION #1 WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 2.50 MM X 38 MM SYNERGY STENTS SYSTEM OVERLAPPED WITH A 2.50 MM X 24 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. FOUR DAYS AFTER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2022, THE SUBJECT DIED. AT THE TIME OF EVENT, THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. MEDICATION WAS GIVEN TO TREAT THE EVENT. THE PRIMARY CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: B3: DATE OF EVENT: USED FIRST DATE OF THE MONTH SINCE EXACT EVENT DATE WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180151&pc=NPS | 2135147-2025-00340 | 2020-11-01 04:00 | Death | ABBOTT MEDICAL | 2025-01-17 05:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "SEX DIFFERENCES AMONG PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR USING THE EDGE-TO-EDGE TECHNIQUE" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE SEC DIFFERENCES AMONG PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR (TTVR). DEVICES MENTIONED INCLUDE TRICLIP AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT IN OUR REAL-WORLD COHORT, MORE FEMALES UNDERWENT TTVR THAN MALES. NO DIFFERENCE WAS OBSERVED IN OUTCOMES BETWEEN MALES AND FEMALES AT ONE-YEAR FOLLOW-UP. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MHD NAWAR ALACHKAR AT DEPARTMENT OF CARDIOLOGY AND ANGIOLOGY, KLINIKUM COBURG INSTITUTION WITH CORRESPONDING EMAIL NAWAR-ALASHKAR@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS NOVEMBER 2020 TO MARCH 2023. A TOTAL OF 105 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 101 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, CHRONIC RENAL FAILURE, CORONARY ARTERY DISEASE, PREVIOUS CARDIAC SURGERY. (B)(6), UNK TRICLIP: PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF THE TRICLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE TITLED ¿SEX DIFFERENCES AMONG PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR USING THE EDGE-TO-EDGE TECHNIQUE¿. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, CHRONIC RENAL FAILURE, CORONARY ARTERY DISEASE, PREVIOUS CARDIAC SURGERY. SOME OF THE COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2- DATE OF DEATH: DEATH DATE WAS ESTIMATED TO THE END OF STUDY RANGE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. LITERATURE: ARTICLE TITLE: SEX DIFFERENCES AMONG PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR USING THE EDGE-TO-EDGE TECHNIQUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180218&pc=NPT | 2025587-2025-00461 | 2025-01-14 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-17 05:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO CROSS THE AORTIC ARCH DUE TO THE NOSE CONE GETTING CAUGHT ON CALCIUM. THE DCS WAS MANIPULATED BY TURNING THE HANDLE TO REALIGN THE SPINE OF THE DCS AND WAS READVANCED. FORWARD FORCE WAS APPLIED TO ADVANCE THE DCS HOWEVER IT CONTINUED TO GET CAUGHT ON THE CALCIUM. THE PATIENT BECAME HYPOTENSIVE. UPON INVESTIGATION OF THE AORTIC ARCH, AN AORTOGRAM REVEALED A PERFORATION. THE PATIENT WAS CONVERTED TO OPEN HEART SURGERY. THE LOCATION OF THE PERFORATION WAS TOO DIFFICULT TO ACCESS AND CONTROL AND THE PATIENT DIED.Event Description: IT WAS REPORTED THAT THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO CROSS THE AORTIC ARCH DUE TO THE NOSE CONE GETTING CAUGHT ON CALCIUM. THE DCS WAS MANIPULATED BY TURNING THE HANDLE TO REALIGN THE SPINE OF THE DCS AND WAS READVANCED. FORWARD FORCE WAS APPLIED TO ADVANCE THE DCS HOWEVER IT CONTINUED TO GET CAUGHT ON THE CALCIUM. THE PATIENT BECAME HYPOTENSIVE. UPON INVESTIGATION OF THE AORTIC ARCH, AN AORTOGRAM REVEALED A PERFORATION. THE PATIENT WAS CONVERTED TO OPEN HEART SURGERY. THE LOCATION OF THE PERFORATION WAS TOO DIFFICULT TO ACCESS AND CONTROL AND THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DCS CAUSED THE PERFORATION AND THE CAUSE OF DEATH WAS BLEEDING DUE TO BEING UNABLE TO PATCH THE PERFORATION DUE TO THE LOCATION. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012421132); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVOLUTFX-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180249&pc=NGV | 2124215-2024-83933 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 670 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180290&pc=NGV | 2124215-2024-83936 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 15 | Event Description: THIS REPORT SUMMARIZES 15 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 65 TO (B)(6) YEARS. PATIENT SEX: 8 FEMALE; 7 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 578 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180384&pc=NPT | 2015691-2025-00376 | 2024-12-26 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-17 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN SAUDI ARABIA, THIS WAS A CASE WITH A 26 MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. AFTER DEPLOYMENT OF THE VALVE IN THE INTENDED POSITION, IT EMBOLIZED INTO THE AORTA. A SECOND 26 MM SAPIEN 3 ULTRA VALVE WAS PREPARED AND DEPLOYED IN THE INTENDED POSITION, BUT AGAIN THE VALVE EMBOLIZED INTO THE AORTA. BOTH VALVES WERE DEPLOYED IN THE AORTA. THE PATIENT WAS PUT ON ECMO AND TRANSFERRED FOR SURGICAL VALVE REPLACEMENT. THE PATIENT EXPIRED. AS PER MEDICAL OPINION, THE MOST PROBABLE ROOT CAUSE OF THE EVENT WAS LOSS OF PACING CAPTURE DURING THE DEPLOYMENT OF THE VALVES. Manufacturer Narrative: THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT NO: 2015691-2025-00375. PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. BASED ON THE INFORMATION RECEIVED, THE PROBABLE ROOT CAUSE OF THE VALVE EMBOLIZATION WAS LOSS OF PACING CAPTURE DURING DEPLOYMENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180405&pc=NGV | 2124215-2024-83978 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 69 TO (B)(6) YEARS. PATIENT SEX: 3 FEMALE; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 600 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180406&pc=NGV | 2124215-2024-83962 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 8 | Event Description: THIS REPORT SUMMARIZES 8 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 70 TO 87 YEARS. PATIENT SEX: 4 FEMALE; 4 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 509 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180409&pc=NGV | 2124215-2024-83937 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Transient Ischemic Attack |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR TRANSIENT ISCHEMIC ATTACK (TIA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 489 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180604&pc=DQY | 2029214-2025-00164 | 2024-10-21 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-17 05:00 | DQY | REACT CATHETER | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | K182097 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A CLINICAL STUDY PATIENT WHO HAD UNDERGONE A THROMBECTOMY PROCEDURE ON (B)(6) 2024 IN WHICH A REACT-71 ASPIRATION CATHETER WAS USED. THERE WAS NO REPORTED DEVICE MALFUNCTION. THE PATIENT'S TICI AT ADMISSION WAS 1 AND A TICI 3 WAS ACHIEVED WITH THE PROCEDURE. IT WAS NOTED THE ADVERSE EVENT BEGAN ON (B)(6) 2024 AND THE PATIENT HAD FATAL OUTCOME THOUGH THE EXACT DATE OF DEATH WAS NOT SPECIFIED IN THE CLINICAL STUDY REPORT. THE SITE ASSESSMENT CONCLUDED THE EVENT HAD A CAUSAL RELATIONSHIP TO THE PATIENT DISEASE UNDER STUDY AND WAS NOT RELATED TO THE REACT CATHETER, ANCILLARY DEVICES, OR THE PROCEDURE. HOWEVER, THE STUDY SPONSOR ASSESSED CONSERVATIVELY THAT THE EVENT COULD POSSIBLY BE PROCEDURE RELATED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED MESENCEPHALIC HEMORRHAGIC REMODELING WITH MAJOR EDEMA RESPONSIBLE FOR EXCLUSIONARY TRIVENTRICULAR HYDROCEPHALUS WITH ASSOCIATED (B)(6) INVOLVEMENT RAPID PROGRESSION TO CLINICAL BRAIN DEATH DIABETES INSIPIDUS TREATED BY COMPENSATORY CRYSTALLOID FILLING. THE PATIENT PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6. CONCLUSION CODE UPDATED/CORRECTED BASED ON AVAILABLE INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180638&pc=CAF | 2518422-2025-002279 | 2024-03-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | CAF | I-NEB CF UK UNIT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102454 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT. THE PATIENT HAS PASSED AWAY. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: CORRECTION TO BOX B5: EVENT DESCRIPTION: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT. THE PATIENT HAS PASSED AWAY. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180639&pc=CAF | 2518422-2025-004378 | 2022-02-18 05:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | CAF | INSPIRATION ELITE NEB | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information | K042655 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A INSPIRATION ELITE NEB. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED THAT A PATIENT PASSED AWAY THAT WAS USING A INSPIRATION ELITE NEBULIZER. UPON REVIEW, THIS INFORMATION WAS INCORRECT. THIS REPORT CORRECTS THE PREVIOUS FILED INFORMATION. THE MANUFACTURER RECEIVED INFORMATION FROM A RELATIVE OF A DECEASED PATIENT INDICATING THAT A BRAND-NEW NEBULIZER WAS FOUND UNUSED IN THE PATIENT'S CLOSET BY THE PATIENT'S MOTHER-IN-LAW. ACCORDING TO FDA 21 CFR PART 803, A DEATH OR SERIOUS INJURY IS CONSIDERED REPORTABLE IF THE MEDICAL DEVICE IS DETERMINED TO HAVE "CAUSED OR CONTRIBUTED" TO THE EVENT, MEANING THE DEVICE MUST HAVE BEEN A FACTOR IN THE DEATH OR SERIOUS INJURY. BASED ON THE INFORMATION PROVIDED, THE NEBULIZER IN QUESTION WAS UNUSED, AND AS SUCH, IT COULD NOT HAVE CONTRIBUTED TO OR CAUSED THE PATIENT'S DEATH. THEREFORE, THERE IS INSUFFICIENT EVIDENCE TO SUGGEST THAT THIS EVENT IS REPORTABLE TO ANY REGULATORY AGENCY UNDER THE CRITERIA. ADDITIONALLY, THE PRIOR REPORT SUBMITTED WAS FILED IN ERROR, AND THIS EVENT HAS BEEN DEEMED NOT REPORTABLE. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180640&pc=BZD | 2518422-2025-004331 | 2023-06-21 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180797&pc=BZD | 2518422-2025-004344 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21180994&pc=BZD | 2518422-2025-100856 | 2025-01-14 05:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. AS A RESULT OF A THIRD-PARTY INSPECTION, ONE ERROR, E004, WAS FOUND IN THE DEVICE LOG BUT WAS UNABLE TO REPLICATE. IN ADDITION, NO FAULTS WERE FOUND, AND THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO THE LOAN STOCK. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. AS A RESULT OF A THIRD-PARTY INSPECTION, ONE ERROR, E004, WAS FOUND IN THE DEVICE LOG BUT WAS UNABLE TO REPLICATE. IN ADDITION, NO FAULTS WERE FOUND, AND THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO THE LOAN STOCK. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21181000&pc=DSQ | 2916596-2025-00243 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21181063&pc=DSQ | 2916596-2025-00241 | 2022-12-01 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Electrical /Electronic Property Problem; Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Multiple Organ Failure; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "OUTCOME AFTER LEFT VENTRICULAR ASSIST DEVICE EXCHANGE" AS LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY HAS EVOLVED FROM SHORT-TERM BRIDGE BRIDGE-TO-TRANSPLANT STRATEGY INTO A LONG-TERM AND OFTEN CHRONIC THERAPY AND APPLICATION AS DESTINATION THERAPY THAT PATIENTS ARE AT RISK OF DEVELOPING COMPLICATIONS SUCH AS PUMP THROMBOSIS, DEVICE INFECTIONS, TECHNICAL PUMP FAILURE, GASTROINTESTINAL (GI) BLEEDING AND LATE RIGHT HEART FAILURE NECESSITATING LVAD EXCHANGE. SIXTY-ONE PATIENTS WHO UNDERWENT LVAD EXCHANGE BETWEEN JANUARY 2010 AND DECEMBER 2022 WERE INCLUDED. OF THESE 61 PATIENTS 18 HAD HEARTWARE DEVICES, 38 WERE HEARTMATE II DEVICES AND 5 WERE HEARTMATE 3 DEVICES RESULTING IN A TOTAL OF 80 LVAD EXCHANGES. THE MEAN AGE AT THE TIME OF EXCHANGE WAS 49.7 YEARS AND 27.7% WERE FEMALE. FOR THE 61 PATIENTS¿ SURVIVAL WAS 83% AFTER 1 YEAR AND 67% 6 YEARS AFTER EXCHANGE. AFTER 5 YEARS 25% HAD BEEN TRANSPLANTED, 23.8% HAD UNDERGONE RE-EXCHANGE AND 32.5% WERE ALIVE WITHOUT NEW INTERVENTION. THE MEDIAN DURATION FROM INDEX IMPLANTATION TO LVAD EXCHANGE WAS 787 DAYS FOR HEARTMATE II AND 134 DAYS FOR HEARTMATE 3. INDICATIONS FOR LVAD EXCHANGE WERE 88.5% DEVICE MALFUNCTION, 8.2% DEVICE INFECTIONS, 1.6% PROPHYLACTIC EXCHANGE FOR THE RISK OF HVAD DYSFUNCTION AND 1.6 % RECURRING GI BLEEDING. OF THE PATIENTS IN THE STUDY SOME PATIENTS HAD TRICUSPID REGURGITATION (TRIVIAL - 42.6%, MILD 18%, MODERATE - 6.6%, SEVERE - 4.9%), SOME HAD AORTIC REGURGITATION (TRIVIAL - 31.1%, MILD -13.1%, MODERATE - 4.9%), AND SOME HAD MITRAL REGURGITATION (TRIVIAL - 18%, MILD - 18%, MODERATE - 6.6%). THERE WERE ALSO SOME PATIENT'S THAT HAD EXCHANGED DUE TO OUTFLOW GRAFT OBSTRUCTION (3.3%) AND INLET OBSTRUCTION (3.3%). RENAL DYSFUNCTION WAS SEEN IN 9.8% OF PATIENTS PRIOR TO LVAD EXCHANGE, IT OCCURRED 10% IN THE SHORT TERM AFTER EXCHANGE AND 4% IN THE LONG TERM. A TOTAL OF 15 PATIENTS DIED AFTER LVAD EXCHANGE DUE TO ISCHEMIC CEREBROVASCULAR ACCIDENTS (13.3%), MULTI-ORGAN FAILURE (13.3%) AND PUMP FAILURE (6.7%) WERE THE CAUSES OF DEATH IN THE SHORT-TERM. THE MOST COMMON CAUSES FOR LONG-TERM MORTALITY WERE INTRA-CRANIAL HEMORRHAGE (26.7%), MULTI-ORGAN FAILURE (20.0%) AND PUMP FAILURE (6.7%). THE MOST FREQUENTLY OBSERVED SHORT-TERM COMPLICATIONS AFTER LVAD EXCHANGE PROCEDURES WERE RIGHT HEART FAILURE (16.3%), RESPIRATORY FAILURE (16.3%), PULMONARY INFECTION (16.3%) AND VENTRICULAR TACHYCARDIA (13.8%). RE-STERNOTOMY DUE TO BLEEDING OR PERICARDIAL EFFUSION WAS PERFORMED AFTER 10 PROCEDURES (12.5%). AFTER 9 (11.3%) PROCEDURES, A SEPTIC EPISODE WAS ENDURED. IN THE LONG-TERM, EXIT-SITE INFECTIONS (34.7%), DEVICE MALFUNCTIONS (25.3%), HEPATIC DYSFUNCTION (16.0%) AND MAJOR HEMOLYSIS (16.0%) WERE THE MOST FREQUENTLY OBSERVED COMPLICATIONS AFTER AN EXCHANGE PROCEDURE. THE COMPLICATIONS AND ADVERSE EVENTS WERE NOT BROKEN DOWN BY LVAD TYPE. ALTHOUGH LVAD EXCHANGE CAN BE PERFORMED WITH A RELATIVELY LOW MORTALITY, OTHER POST-OPERATIVE ADVERSE EVENTS WERE COMMON. Manufacturer Narrative: SECTION A AND D - NO SPECIFIC PATIENT AND DEVICE INFORMATION IS KNOWN. NIAMAT, J., RAMJANKHAN, F., VAN DER KAAIJ, N., GIANOLI, M., VAN LAAKE, L.,W., & MOKHLES, M. M. (2024). OUTCOME AFTER LEFT VENTRICULAR ASSIST DEVICE EXCHANGE. EUROPEAN JOURNAL OF CARDIO-THORACIC SURGERY : OFFICIAL JOURNAL OF THE EUROPEAN ASSOCIATION FOR CARDIO-THORACIC SURGERY, 66(4) DOI:HTTP://DX.DOI.ORG/10.1093/EJCTS/EZAE317. DEPARTMENT OF CARDIOTHORACIC SURGERY, UNIVERSITY MEDICAL CENTER UTRECHT, UTRECHT, NETHERLANDS. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED "OUTCOME AFTER LEFT VENTRICULAR ASSIST DEVICE EXCHANGE" REPORTED THAT AS LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY HAS EVOLVED FROM SHORT-TERM BRIDGE-TO-TRANSPLANT STRATEGY INTO A LONG-TERM AND OFTEN CHRONIC THERAPY, PATIENTS ARE AT RISK OF DEVELOPING COMPLICATIONS NECESSITATING LEFT VENTRICULAR ASSIST DEVICE (LVAD) EXCHANGE. THE AIM OF THE STUDY WAS TO ASSESS PATIENT OUTCOMES AFTER LVAD EXCHANGE. 61 PATIENTS WHO UNDERWENT LVAD EXCHANGE BETWEEN JANUARY 2010 AND DECEMBER 2022 WERE INCLUDED. OF THESE 61 PATIENTS, 18 HAD HEARTWARE DEVICES, 38 HAD HEARTMATE II DEVICES, AND 5 HAD HEARTMATE 3 DEVICES RESULTING IN A TOTAL OF 80 LVAD EXCHANGES. ALL 5 HEARTMATE 3 DEVICES WERE EXCHANGED FOR NEW HEARTMATE 3 DEVICES. THE REMAINING DEVICES WERE EXCHANGED FOR EITHER HEARTWARE, HEARTMATE II, OR HEARTMATE 3 DEVICES. ADVERSE EVENTS PRIOR TO LVAD EXCHANGE INCLUDED STROKE, CHRONIC KIDNEY DISEASE, ACUTE RENAL DYSFUNCTION, RESPIRATORY FAILURE, THROMBOEMBOLISM, RIGHT HEART FAILURE, VENTRICULAR TACHYCARDIA, SUPRAVENTRICULAR TACHYCARDIA, HEPATIC DYSFUNCTION, INTRACARDIAC THROMBUS, GASTROINTESTINAL (GI) BLEEDING, ENDOCARDITIS, HEMOLYSIS, PULMONARY EMBOLISM, PERICARDIAL FLUID COLLECTION, SEPSIS, AND HEMOTHORAX. PATIENT CHARACTERISTICS ALSO INCLUDED TRICUSPID REGURGITATION, AORTIC REGURGITATION, MITRAL REGURGITATION. 17 PATIENTS WENT ON TO RECEIVE A SECOND EXCHANGE, WITH 1 PATIENT RECEIVING A THIRD AND FOURTH EXCHANGE. INDICATIONS FOR EXCHANGE INCLUDED DEVICE INFECTION, DEVICE MALFUNCTION, PUMP THROMBOSIS, OUTFLOW GRAFT OBSTRUCTION, INLET OBSTRUCTION, AND GI BLEEDING. THE MEAN AGE AT THE TIME OF EXCHANGE WAS 49.7 YEARS AND 27.7% WERE FEMALE. FOR THE 61 PATIENTS¿ SURVIVAL WAS 83% AFTER 1 YEAR AND 67% 6 YEARS AFTER EXCHANGE. AFTER 5 YEARS, 25% HAD BEEN TRANSPLANTED, 23.8% HAD UNDERGONE RE-EXCHANGE AND 32.5% WERE ALIVE WITHOUT NEW INTERVENTION. THE MEDIAN DURATION FROM INDEX IMPLANTATION TO LVAD EXCHANGE WAS 134 DAYS FOR HEARTMATE 3. SHORT-TERM CAUSES OF DEATH INCLUDED MULTI-ORGAN FAILURE, ISCHEMIC CEREBROVASCULAR ACCIDENT (CVA), AND PUMP FAILURE. LONG-TERM CAUSES OF DEATH INCLUDED MULTI-ORGAN FAILURE, INTRACRANIAL HEMORRHAGE, AND PUMP FAILURE. ADDITIONALLY, THERE WERE 2 LONG TERM CASES WHERE CAUSE OF DEATH WAS NOT REPORTED. SHORT-TERM ADVERSE EVENTS OBSERVED AFTER LVAD EXCHANGE INCLUDED ISCHEMIC CVA, HEMORRHAGIC CVA, TRANSIENT ISCHEMIC ATTACK (TIA), GI BLEEDING, POST-OPERATIVE BLEEDING, PULMONARY INFECTION, EXIT-SITE INFECTION, PUMP THROMBOSIS, RESPIRATORY FAILURE, ARTERIAL THROMBOEMBOLISM, VENOUS THROMBOEMBOLISM, HEPATIC DYSFUNCTION, RENAL DYSFUNCTION, CARDIAC ARRHYTHMIA, RIGHT HEART FAILURE, SEPSIS, PERICARDIAL FLUID COLLECTION, HEMOLYSIS, AND PSYCHIATRIC EPISODE. LONG-TERM ADVERSE EVENTS INCLUDED ISCHEMIC CVA, HEMORRHAGIC CVA, TIA, PULMONARY INFECTION, EXIT-SITE INFECTION, DEVICE INFECTION, GI BLEEDING, PUMP THROMBOSIS, RENAL DYSFUNCTION, CARDIAC ARRHYTHMIA, SEPSIS, ENDOCARDITIS, RESPIRATORY FAILURE, RIGHT HEART FAILURE, ARTERIAL THROMBOEMBOLISM, HEPATIC DYSFUNCTION, PERICARDIAL FLUID COLLECTION, HEMOLYSIS, HEMOTHORAX, AND PSYCHIATRIC EPISODE. THE COMPLICATIONS AND ADVERSE EVENTS WERE NOT BROKEN DOWN BY LVAD TYPE. THE ARTICLE CONCLUDED THAT LVAD EXCHANGE IS A VIABLE OPTION WITH AN ACCEPTABLE SURVIVAL RATE; HOWEVER, THE RISK OF COMPLICATIONS SHOULD BE CAREFULLY WEIGHED. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. G, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH, MULTIPLE ORGAN FAILURES/DYSFUNCTIONS, STROKE, AND BLEEDING. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, LISTS EXAMPLES OF EMERGENCIES AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THESE EMERGENCIES. FURTHERMORE, THE PATIENT HANDBOOK INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21181069&pc=DSQ | 3007042319-2025-00071 | 2025-01-09 05:00 | Death | HEARTWARE, INC. | 2025-01-17 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT, WHO HAD A HISTORY OF DILATED CARDIOMYOPATHY (DCMP) AND WAS IN PALLIATIVE CARE, ALONG WITH THEIR S OUSE AND PALLIATIVE CARE TEAM, MADE THE DECISION TO DEACTIVATE THE VENTRICULAR ASSIST DEVICE-SYSTEM. THE PATIENT SUBSEQUENTLY PASSED AWAY. THERE WERE NO ALLEGATIONS AGAINST THE VENTRICULAR ASSIST DEVICE PUMP, ASSOCIATED DEVICES, OR THERAPY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS IN PALLIATIVE CARE AND SUBSEQUENTLY EXPIRED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, WORSENING HEART FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21181089&pc=DSQ | 2916596-2025-00205 | 2024-11-15 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO FAILURE TO THRIVE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. C. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21181259&pc=DSQ | 2916596-2025-00143 | 2024-12-23 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PUMP WAS NOT EXPLANTED, AND IT WAS UNKNOWN IF IT WOULD BE RETURNED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21181262&pc=DSQ | 2916596-2025-00356 | 2023-06-01 04:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THROUGH THE RESEARCH ARTICLE TITLED ¿OUTCOME AND EXPENSE BETWEEN THE GROUP OF PATIENTS WITH HEARTMATE 3 AND RECEIVING HEART TRANSPLANT - ANALYSIS FROM THE EXPENSE POINT¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH DEATH. THIS STUDY RETROSPECTIVELY ENROLLED PATIENTS WHO RECEIVED THE HM3 OR A HEART TRANSPLANT BETWEEN OCT2018 AND JUN2023. HEART TRANSPLANT PATIENTS WHO WERE BRIDGED WITH THE HM3 WERE EXCLUDED. THE HM3 GROUPS ENROLLED THE PATIENTS RECEIVING THE HM3 WITHOUT RECEIVING A HEART TRANSPLANT (HT). A TOTAL OF 114 PATIENTS (47 IN THE HM3 GROUP AND 67 IN THE HT GROUP) WERE INCLUDED. THE MEAN FOLLOW-UP WAS 46.1 ± 10.2 MONTHS. THE WHOLE GROUP SURVIVAL AT 2-YEAR WAS 73.9 % AND 3-YEAR WAS 72.0%. THE HT GROUP HAD THE BETTER 90-DAY SURVIVAL (P=0.028) THAN HM3 GROUP. THERE WERE NO STATISTICAL DIFFERENCES IN 180-DAY,1-YEAR, 2-YEAR AND 3-YEAR SURVIVAL. THE HM3 GROUP MEAN EXPENSE 1 YEAR BEFORE SURGERY WAS LOWER THAN HT (US $(B)(6) VS. US$(B)(6), P=0.000). THE POST HT EXPENSE WAS SIGNIFICANTLY HIGHER IF THE PATIENT DIED (P < 0.05). THE MEAN COST AFTER THE FOLLOW-UP COMPARISON BETWEEN THE 2 GROUP SHOWED NO SIGNIFICANT DIFFERENCE WHEN PATIENTS SURVIVAL WAS MORE THAN 90 DAYS (P=0.328). IN CONCLUSION, ALTHOUGH RECEIVING A HEART TRANSPLANT WAS CONSIDERED PRICELESS, HM3 THERAPY MAY BE COST-EFFECTIVE COMPARED TO HT IN THE CONTEXT. EXPENSE ANALYSIS OF ANY TREATMENT WAS CHALLENGING AND HAD INHERENT LIMITATIONS. MORE DETAILED DATA ON CORRELATING CLINICAL OUTCOME AND THE COST IN FUTURE WAS NEEDED. Manufacturer Narrative: D4: MODEL NUMBER CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. TSAO, C., YUNG, Y., WANG, C., CHI, N., HUANG, S., YU, H., HSU, R., & CHEN, Y. (2024). OUTCOME AND EXPENSE BETWEEN THE GROUP OF PATIENTS WITH HEARTMATE 3 AND RECEIVING HEART TRANSPLANT - ANALYSIS FROM THE EXPENSE POINT. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S270. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1193. (B)(6) UNIVERSITY HOSPITAL. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182180&pc=DSQ | 2916596-2025-00453 | 2024-12-28 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CONTROLLER FAULT ALARM. THEY REPORTED THE GREEN LIGHT WAS ON AT THAT TIME AND THEY FELT FINE. UPON ARRIVAL TO THE EMERGENCY DEPARTMENTS (ED), THE GREEN LIGHT WAS OFF AND THE PATIENT WAS BEGINNING TO HAVE SYMPTOMS OF HEART FAILURE. THE REGISTERED NURSE (RN) STATED THAT THE CONTROLLER WOULD NOT CONNECT TO THE HEARTMATE TOUCH MONITOR, THERE WAS NO GREEN LIGHT AND THE DISPLAY SHOWED "CONTROLLER FAULT". IT WAS ESTIMATED THAT THE PUMP COULD HAVE BEEN OFF FOR POSSIBLY 1 HOUR. HEPARIN WAS SUGGESTED FOR THE PATIENT AS THEY WERE NOT ANTICOAGULATED THERAPEUTICALLY. A PUMP EXCHANGE WAS CONSIDERED. THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) AND PLACED ON IMPELLA 5.5 DUE TO WORSENING RENAL FUNCTION. THEY WERE AWAITING A POSSIBLE PUMP EXCHANGE. THE ALARM WAS DESCRIBED AS A CONTROLLER FAULT THAT PROGRESSED TO THE GREEN LIGHT BEING OFF AND THE PUMP BEING OFF. THIS WAS VERIFIED BY AUSCULTATION OF THE PATIENT'S CHEST. NO VAD SOUNDS WERE HEARD. THE CONTROLLER AND MODULAR CABLE WERE EXCHANGED. LOG FILES WERE UNAVAILABLE AS THE CONTROLLER WAS UNABLE TO BE CONNECTED TO SYSTEM MONITOR OR HEARTMATE TOUCH. THE PATIENT UNDERWENT A HEARTMATE 3 EXCHANGE ON (B)(6)2024. THE PATIENT SURVIVED THE OPERATING ROOM BUT PASSES AWAY 4 HOURS POSTOPERATIVELY ON (B)(6)2025 IN THE INTENSIVE CARE UNIT (ICU).Event Description: IT WAS ADDITIONALLY REPORTED THAT THE OUTCOME WAS NOT CONSIDERED DEVICE RELATED AND THE DEATH WAS DUE TO BLEEDING COMPLICATIONS FROM THE PUMP EXCHANGE. THE NEWLY IMPLANTED PUMP WAS FUNCTIONING AS INTENDED. THE LVAD WAS LEFT OFF FROM(B)(6) 2024 TO THE PUMP EXCHANGE ON (B)(6) 2025. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182182&pc=NGV | 2124215-2024-84025 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR A SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182385&pc=BZD | 2518422-2025-100870 | 2024-08-15 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | REMSTAR AUTO A-FLEX W/HUMID, SD CARD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K091319 |  | 1 | Event Description: IT WAS REPORTED THAT THE REMSTAR AUTO A-FLEX DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182411&pc=NPT | 2025587-2025-00468 | 2025-01-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-17 05:00 | NPT | EVOLUT FX DCS | Positioning Problem; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Rupture; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE IMPLANTATION WITH A 26MM EVOLUT FX+ VALVE, A PRE-BALLOON AORTIC VALVULOPLASTY (BAV) USING A 22MM TRUE BALLOON WAS PERFORMED DUE TO CALCIFICATION. DURING THE PROCEDURE, TWO RECAPTURES WERE REQUIRED BEFORE FINAL RELEASE OF THE VALVE. THE FIRST RECAPTURE WAS DUE TO THE VALVE BEING DEPLOYED TOO HIGH, AND THE SECOND WAS BECAUSE THE VALVE DISLODGED DUE TO A LACK OF PACING. THE VALVE WAS SUCCESSFULLY DEPLOYED ON THE THIRD ATTEMPT. HOWEVER, DURING THE REMOVAL OF THE DELIVERY CATHETER SYSTEM, THERE WAS A DROP IN BLOOD PRESSURE. DESPITE CARDIOPULMONARY RESUSCITATION AND ANESTHESIA EFFORTS, THE PATIENT DID NOT RECOVER. AN EFFUSION WAS IDENTIFIED AT THE LEVEL OF THE ANNULUS AND LEFT VENTRICULAR OUTFLOW TRACT. AN ANNULAR RUPTURE WAS NOTED. PER THE PHYSICIAN, THE RUPTURE WAS THOUGHT TO HAVE POSSIBLY OCCURRED AT THE TIME OF THE PRE-IMPLANT BAV. THE PATIENT'S ANATOMY AND CALCIFICATION WERE NOTED AS CONTRIBUTING FACTORS TO THE EVENT. THE PATIENT ULTIMATELY DIED FOLLOWING THE ANNULAR RUPTURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER AND THE VALVE DISLODGED TOWARDS THE VENTRICLE. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE DELIVERY CATHETER SYSTEM CONTRIBUTED TO THE ANNULAR RUPTURE. Manufacturer Narrative: UPDATED DATA: H2 H3 H6. IMAGE REVIEW: VIDEOS OF THE CINE IMAGES PROVIDED FOR THE EVENT. THERE WAS NO COMPUTED TOMOGRAPHY (CT) IMAGES PROVIDED FOR REVIEW OF ANATOMICAL CONSIDERATIONS. THERE WAS AN EXECUTIVE SUMMARY PROVIDED TO SHOW PATIENTS ANATOMY SIZE, AND CALCIFICATION. FLUOROSCOPIC VALVE LOAD INSPECTION WAS NOT PROVIDED THUS IT IS NOT POSSIBLE TO VALIDATE A GOOD LOAD. A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANTATION. EVIDENCE SHOWS THE BALLOON MOVED DURING INFLATION. YOU CAN SEE THE PIGTAIL IN THE ANNULUS LOCATION. PER MEDTRONIC BEST PRACTICES AN ADEQUATE PRE-DILATION CAN HELP REDUCE THE POTENTIAL NEED FOR POST DILATION, AND MAY ALSO BE USEFUL TO PREPARE THE VALVE FOR CROSSING WITH THE DELIVERY SYSTEM. IN THIS EVENT A 22 MILLIMETER (MM) TRUE BALLOON WAS USED TO PERFORM A PRE-DILATION. PER MEDTRONIC BEST PRACTICED WHEN USING A NON-COMPLIANT BALLOON, SUBTRACT 1 WHEN DETERMINING THE BALLOON SIZE FROM THE PERIMETER DERIVED ANNULAR DIAMETER IN THE ABSENCE OF ANNULAR OR LEFT VENTRICULAR OUTFLOW TRACT (LVOT) CALCIFICATION. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND DEPTH ASSESSMENT WAS PERFORMED. DEPTH APPEARED TO BE APPROXIMATELY 3MM ON THE NON-CORONARY CUSP(NCC) IN THE CUSP OVERLAP VIEW, AND 3MM ON THE LEFT CORONARY CUSP IN THE LEFT ANTERIOR OBLIQUE (LAO) VIEW. AT THE POINT OF NO RECAPTURE THERE IS EVIDENCE OF AN AORTIC DISSECTION IN ALL VIEWS. THE VALVE WAS RELEASED AND AN AORTIC ANGIOGRAPHY WAS PERFORMED. EVIDENCE SHOWS THE LEFT CORONARY CUSP (LCC) FILLING AND NO CONTRAST FILLING THE NCC. EVIDENCE ALSO SHOWS A POSSIBLE DISSECTION IN THE AORTA IN THE IMAGES HIGHLIGHTED BELOW. PER THE EVENT DESCRIPTION THE PATIENT DIED FROM AN ANNULAR RUPTURE. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE DELIVERY SYSTEM CONTRIBUTED TO THE ANNULAR RUPTURE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012449278); PRODUCT TYPE: 0195-HEART VALVES. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME: EVOLUT FX PLUS VALVE; PRODUCT ID: EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182426&pc=DSQ | 2916596-2025-00050 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-17 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiovascular Insufficiency; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO VASCULOPATHY, RIGHT VENTRICULAR (RV) FAILURE, AND PULSELESS ELECTRICAL ACTIVITY (PA) ARREST. Manufacturer Narrative: SECTION A3A: PATIENT SEX WAS NOT YET PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO RELEVANT DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS RIGHT HEART FAILURE AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS RIGHT HEART FAILURE AS A POTENTIAL RISK/ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182466&pc=NPT | 2025587-2025-00469 | 2025-01-12 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-17 05:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Patient Device Interaction Problem | Hemorrhage/Bleeding | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT SEVERAL ATTEMPTS WERE MADE TO ACCESS TO THE NATIVE VALVE HOWEVER IT WAS DIFFICULT DUE TO THE DELIVERY CATHETER SYSTEM (DCS) GETTING STUCK IN THE ASCENDING AORTA. THE GUIDEWIRE WAS CHANGED, A BUDDYWIRE WAS ATTEMPTED AS WELL AS A SNARE. AN INTRODUCER WAS PLACED AND THE DCS VISUALLY WAS GOOD. IT WAS DETERMINED TO CHANGE THE DCS, BUT THE DCS THEN ADVANCED TO THE NATIVE VALVE. THE PATIENT WAS GOOD LEAVING THE OPERATING ROOM. SEVERAL HOURS LATER, HEMORRHAGE WAS NOTED AND THE PATIENT DIED. NO AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH IS UNKNOWN. THERE WAS FEMORAL BLEEDING AT THE END OF THE PROCEDURE BUT IT WAS RESOLVED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF THE FEMORAL BLEEDING WAS NOT KNOWN AND WAS TREATED WITH COMPRESSION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT A 14 FRENCH (FR) INTRODUCER SHEATH WAS USED DURING THE PROCEDURE. IT WAS CONFIRMED THAT THE VALVE WAS FULLY DEPLOYED WITHIN THE PATIENT. IT WAS REPORTED THAT THE HEMORRHAGING WHICH OCCURRED SEVERAL HOURS AFTER THE PROCEDURE WAS FROM THE FEMORAL PUNCTURE SITE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS: UPON RECEIPT OF THE SUSPECT COMPLAINT DEVICE AT MEDTRONIC¿S QUALITY LABORATORY, VISUAL ANALYSIS SHOWED THE DELIVERY CATHETER SYSTEM (DCS) WAS RECEIVED WITHOUT A VALVE LOADED WITHIN THE CAPSULE. THE DEVICE WAS RECEIVED WITH THE CAPSULE PARTIALLY OPENED. THERE WAS A SLIGHT BEND TO THE NOSECONE. DELAMINATION WAS OBSERVED OVER THE NITINOL REINFORCING FRAME THE DISTAL END AND THE PROXIMAL END OF THE CAPSULE. THERE WAS DEFORMATION TO THE DISTAL END OF THE STABILITY SHAFT. THERE WAS A BEND TO THE STABILITY SHAFT. THE HANDLE APPEARED INTACT. THE DEVICE WAS RETURNED WITH THE END CAP/SCREW GEAR SNAP FIT CONNECTED. THE HANDLE APPEARED TO RETRACT AND ADVANCE THE CAPSULE. THE TRIGGER MOVED TO FULLY ADVANCED AND RETRACTED POSITIONS AND LOCKED IN PLACE WHEN RELEASED. THE TIP-RETRIEVAL MECHANISM APPEARED INTACT. THE INNER MEMBER SHAFT AND SPINDLE HUB APPEARED INTACT WITH NO EVIDENCE OF DAMAGE. THERE WAS DAMAGE OBSERVED ON THE THREADING OF THE SCREW GEAR. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183116&pc=NGV | 2124215-2024-84081 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Migration | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEVICE MIGRATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183141&pc=BZD | 2518422-2025-100868 | 2025-01-14 05:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A SERVICE CENTER. DURING THE EVALUATION, ERROR E004 WAS FOUND IN THE DEVICE'S EVENT LOG, AND THE PCA WAS REPLACED. IN ADDITION, THE UPPER ENCLOSURE WAS REPLACED DUE TO COSMETIC DAMAGE, AND THE BLOWER ASSEMBLY WAS REPLACED DUE TO CONTAMINATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A SERVICE CENTER. DURING THE EVALUATION, ERROR E004 WAS FOUND IN THE DEVICE'S EVENT LOG, AND THE PCA WAS REPLACED. IN ADDITION, THE UPPER ENCLOSURE WAS REPLACED DUE TO COSMETIC DAMAGE, AND THE BLOWER ASSEMBLY WAS REPLACED DUE TO CONTAMINATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183177&pc=NGV | 2124215-2024-84083 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR IMPLANT EMBOLIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183496&pc=NGV | 2124215-2024-84089 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 63 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183499&pc=NGV | 2124215-2024-84085 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ADDITIONAL DEVICE REQUIRED. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 0 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS. (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183501&pc=NGV | 2124215-2024-84086 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR PROCEDURE CANCELLED/RESCHEDULED - POST SEDATION/SEDATION UNKNOWN. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183504&pc=NGV | 2124215-2024-84088 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183507&pc=NGV | 2124215-2024-84090 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-17 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183576&pc=CBK | 2518422-2025-100879 | 2024-12-24 05:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY WHILE ON A TRILOGY EVO DEVICE. ADDITIONAL INFORMATION PROVIDED STATES THE DEVICE WAS PLACE IN STAND-BY MODE BY THE FACILITY NURSE. THE DEVICE WAS NOT TAKEN OUT OF STAND-BY MODE FOR APPROXIMATELY AND HOUR AND A HALF. WHEN THE STAFF ENTERED THE ROOM, THE PATIENT WAS NOT BREATHING. RESUSCITATION WAS STARTED, BUT WAS UNSUCCESSFUL. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM ANY DEVICE MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT PASSED AWAY WHILE ON A TRILOGY EVO DEVICE. ADDITIONAL INFORMATION PROVIDED STATES THE DEVICE WAS PLACE IN STAND-BY MODE BY THE FACILITY NURSE. THE DEVICE WAS NOT TAKEN OUT OF STAND-BY MODE FOR APPROXIMATELY AND HOUR AND A HALF. WHEN THE STAFF ENTERED THE ROOM, THE PATIENT WAS NOT BREATHING. RESUSCITATION WAS STARTED BUT WAS UNSUCCESSFUL. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE'S DOWNLOADED EVENT LOG WAS REVIEWED, AND NO REPORTABLE ALARMS WERE OBSERVED. THE DEVICE WAS TESTED, AND NO MALFUNCTION OF THE DEVICE WAS FOUND. THE DEVICE WAS CLEANED AND TESTED AND PASSED ALL FINAL TESTING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183791&pc=FTM | 1000306051-2025-00003 | 2024-12-16 05:00 | Death | LIFECELL | 2025-01-17 05:00 | FTM | 20 X 20 FIRM STRATTICE FOR EU DISTRIBUTION | Adverse Event Without Identified Device or Use Problem | Impaired Healing; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K070560 |  | 1 | Event Description: A SURGEON REPORTED VIA THE DISTRIBUTOR THE FOLLOWING: ¿WE RECENTLY HAD A COMPLEX PAEDIATRIC KIDNEY TRANSPLANT, BUT UNFORTUNATELY THE KIDNEY THROMBOSED AND NEEDED TO BE REMOVED A FEW DAYS LATER. THE CHILD WAS UNWELL AND A STRATTICE RTM WAS INSERTED AS A BRIDGE BETWEEN THE ABDOMINAL FASCIA TO ENABLE SKIN CLOSURE. I TOOK THE PATIENT BACK TO THEATRE A WEEK OR SO LATER BECAUSE THEY HAD DEVELOPED A SUPERFICIAL WOUND BREAKDOWN. AT THAT POINT I NOTICED TWO HOLES IN THE MIDDLE OF THE STRATTICE MESH, EACH APPROXIMATELY 3 TO 4 CM IN LENGTH, WHICH I THOUGHT WAS VERY UNUSUAL. THERE SEEMED TO BE SOME ENZYMATIC DEGRADATION AND SOFTENING OF THE MESH ALTHOUGH THERE WAS NO OBVIOUS BOWEL PERFORATION OR CONTAMINATION WITHIN THE PERITONEAL CAVITY. BECAUSE THE PATIENT WAS UNWELL, I REPAIRED THE DEFECTS WITH INTERRUPTED PDS SUTURES. UNFORTUNATELY, THE CHILD DIED A FEW DAYS LATER. I HAVE NEVER SEEN A STRATTICE MESH DEGRADE OR RIP OR TEAR IN THIS WAY." THE SURGEON ALSO REPORTED THAT "CHILD DIED." "THIS DOES NOT APPEAR TO BE DIRECTLY ATTRIBUTABLE TO MESH DEFECTS." THE SURGEON SUBSEQUENTLY REPORTED THAT "AT THE TIME OF MESH INSERTION THE WOUND WAS CLEAN CONTAMINATED (ANASTOMOSIS TO THE BLADDER), BUT THE SUPERFICIAL WOUND SUBSEQUENTLY BROKE DOWN WITH THE MESH VISIBLE UNDERNEATH, WITH SOME MUCKY FLUID APPARENT, SO DIRTY-INFECTED AT THE TIME OF WASHOUT." THE IMPLANTING SURGEON REPORTED THAT THE TISSUE WAS "SOAKED IN A BUCKET WITH NORMAL SALINE FOR 2-3 MINUTES BEFORE USING." THE SURGEON ALSO REPORTED VIA A USER REPORT TO THE MHRA THAT "CHILD DIED 2 DAYS LATER, CAUSE YET TO BE IDENTIFIED (?MULTI-ORGAN FAILURE DUE TO POSSIBLE SEPSIS?), WITH MESH STILL IN SITU."Event Description: THIS MEDWATCH IS BEING FILED TO REPORT THE FOLLOWING CORRECTED INFORMATION. OUT OF AN ABUNDANCE OF CAUTION, WE ARE CONSERVATIVELY REPORTING THE PATIENT'S DEATH GIVEN THE SURGEON'S STATEMENT THAT "MESH DEFECTS NOT APPARENTLY DIRECTLY CAUSALLY RELATED TO DEATH, BUT YET TO BE DETERMINED." AS ADDITIONAL INFORMATION, THE SURGEON REPORTED THE PATIENT'S DIAGNOSES INCLUDE "-POSTERIOR URETHRAL VALVES -RENAL FAILURE AND LIMITED DIALYSIS ACCESS -KIDNEY TRANSPLANTATION AND GRAFT THROMBOSIS AND MULTIPLE LAPAROTOMIES." THE PATHOLOGIST ALSO REPORTED THAT THE TISSUE HAD BEEN RETRIEVED FROM THE PATIENT. THE TISSUE WILL BE SENT FOR ANALYSIS.Event Description: THE PATHOLOGIST REPORTED THAT THE TISSUE HAD BEEN RETRIEVED FROM THE PATIENT. THE TISSUE WAS RECEIVED AND UNDERWENT HISTOLOGICAL, PATHOLOGICAL, AND GROSS VISUAL ANALYSIS. Manufacturer Narrative: THIS SAFETY EVENT IS BEING REPORTED AS SERIOUS INJURY DUE TO THE REPORTED WOUND HEALING AND HOLES IN THE STRATTICE MESH WITH SURGICAL INTERVENTION. THE EVENT OF DEATH HAS BEEN EVALUATED AND DETERMINED TO BE NOT DEVICE-RELATED BASED ON THE SURGEON'S STATEMENT. A REVIEW OF THE DEVICE HISTORY RECORD FOR STRATTICE LOT SP300022 HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS IS A KNOWN POTENTIAL ADVERSE EVENT ADDRESSED IN THE PRODUCT LABELING. WITHOUT RELEVANT PATIENT FACTORS, A RELATIONSHIP BETWEEN THE STRATTICE AND THIS EVENT COULD NOT BE DETERMINED. IF ADDITIONAL INFORMATION IS MADE AVAILABLE DURING FOLLOW-UP, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL AND/OR CORRECTED DATA: B5, D3, D9, H3, H6. THE DEVICE WAS ANALYZED, AND THE RESULT OF THE INVESTIGATION IS AS FOLLOWS: "GROSS DESCRIPTION: AN 18 X 6 X .2 CM FLEXIBLE OVOID WHITE AND GRAY IRREGULARLY COLORED FLAT TISSUE FRAGMENT. THERE ARE SEVERAL BLUE SYNTHETIC SUTURES AROUND THE EDGE OF THE FRAGMENT, AND A 6 CM SUTURED WAVY INCISION TOWARDS ONE EDGE OF THE TISSUE. ADJACENT TO THIS SUTURE LINE A 3 CM ATTENUATED AREA IS SEEN, WHICH APPEARS ALMOST TRANSPARENT. REPRESENTATIVE SECTIONS ARE SUBMITTED FOR MICROSCOPIC EXAMINATION IN 3 CASSETTES. MICROSCOPIC DESCRIPTION: ON H&E STAIN, DENSE COLLAGENOUS BUNDLES ARE SEEN ON CROSS-SECTION OF THE FLAT TISSUE FRAGMENT, WITHOUT INFLAMMATION OR FIBROBLASTIC INGROWTH. THERE ARE FOCI WHICH APPEAR MORE HOMOGENEOUSLY FIBROTIC WITHOUT DISCRETE COLLAGEN BUNDLES, ALSO LARGELY ACELLULAR, FOCALLY, THERE ARE A FEW CONFIRMED ZONES OF APPARENT BLUISH FIBRILLAR CHANGE, WITHOUT INFLAMMATORY INFILTRATE. THROUGHOUT MANY AREAS, SCATTERED INERT SYNTHETIC-TYPE SPHERULES ARE NOTED, ALSO WITHOUT ACCOMPANYING CELLULAR REACTION, PRESUMABLY THE SYNTHETIC SUTURE MATERIAL, SEEN IN CROSS-SECTION. SPECIAL STAIN FOR VVG-ELASTIN SHOWS NO RESIDUAL ELASTIN FIBERS: THESE WOULD BE MORE PLENTIFUL IN NATIVE HUMAN DERMIS THAN IN PORCINE DERMIS. B&B GRAM STAIN FOR BACTERIA IS NEGATIVE FOR ORGANISMS, AS IS GRIDLEY STAIN FOR YEAST/FUNGI." "ADDENDUM: FURTHER EXAMINATION OF THE H&E SECTIONS SUGGESTS THAT THE FOREIGN PARTICLES NOTED ABOVE ARE SPHERULES, OVOIDS AND SHORT SEPARATED CHAINS. ALTHOUGH SYNTHETIC SUTURE PARTICLES ARE A POSSIBILITY, THESE STRUCTURES MAY ALSO BE CONSISTENT WITH FOREIGN VEGETABLE MATERIAL. THIS MAY REPRESENT GLOVE CONTAMINANTS, POLLEN, OR STOMACH/ABDOMINAL CONTENTS, ETC. THERE IS NO TISSUE REACTION OR INFLAMMATION ASSOCIATED WITH THESE STRUCTURES, AND THEY ARE ALSO SEEN EXTENSIVELY AT THE OUTER TISSUE EDGES, SUGGESTING NON-PATHOGENIC CONTAMINATION OF SOME TYPE." THE BREAK IN THE TISSUE DESCRIBED BY THE SURGEON WAS CONSISTENT WITH THE VISUAL ANALYSIS FINDINGS; HOWEVER, IT WAS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE. THE REPORTED EVENTS OF DEATH, WOUND HEALING, AND UNSATISFACTORY RESULT ARE PHYSIOLOGICAL PHENOMENA, AND THEREFORE CAUSALITY CANNOT BE DETERMINED BY PRODUCT ANALYSIS. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME TO DETERMINE ANY OTHER PROBABLE CAUSE AND/OR THE EXACT CAUSE OF THE EVENT. THE REPORTED EVENTS ARE ADDRESSED IN THE LABELLING. INTERNAL INVESTIGATION OF COMPLAINT LOT SP300022 INCLUDED COMPLAINT HISTORY REVIEW FOR THE REPORTED LOT, LOT PROCESSING REVIEW, STORAGE REVIEW AND DISTRIBUTION/IMPLANTATION RECORDS REVIEW. QA INVESTIGATION RESULTED IN NO REMARKABLE FINDINGS INCLUDING NO OTHER RELATED COMPLAINTS AGAINST THE LOT AND NO RELATED DEVIATIONS OR NON-CONFORMANCES WERE REVEALED. THE COMPLAINT RELATED LOT MET RELEASE CRITERIA FOR THE FINISHED PRODUCT. THIS CONCLUDES THE INTERNAL INVESTIGATION. IF ADDITIONAL INFORMATION IS PROVIDED TO US, THE RECORD WILL BE RE-EVALUATED. Manufacturer Narrative: ADDITIONAL AND/OR CORRECTED DATA: B2, B5, B6, B7, D3, D6A, G1, H1, H6. A REVIEW OF THE DEVICE HISTORY RECORD FOR STRATTICE LOT SP300022 HAS BEEN COMPLETED. NO DEVIATIONS OR NON-CONFORMANCES NOTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183840&pc=MSX | 3009156722-2025-00002 | 2021-12-20 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-01-17 05:00 | MSX | BENEVISION | Device Alarm System | Insufficient Information | K220058 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, APPROXIMATELY BETWEEN 11:00 AM AND 12:30 PM, IN ROOM ID (B)(6), ON A PATIENT PLACED ON THE TM80 TELEMETRY MONITOR (ID TELE50) THE CENTRAL MONITORING UNIT (CMU) DEPARTMENT MONITOR TECHNICIAN NOTED THE PATIENT WAS IN PACED RHYTHM AT RATE 70 AT 11:00 AM. THE CMU WAS NOTIFIED BY GATEWAY STAFF THAT THE PATIENT WAS BEING RESUSCITATED. THE MONITOR TECHNICIAN NOTED THAT QRS HAD WIDENED AND THAT THERE WAS NO ALARM. THE PATIENT EXPIRED. Manufacturer Narrative: SYSTEM LOGS AND DATA ASSOCIATED WITH THE OCCURRENCE WERE EVALUATED, AND NO MALFUNCTION WAS CONFIRMED FOR ANY DEVICES INVOLVED IN THE REPORT. THE REPORTED WIDENED QRS COMPLEXES ARE THE REGULAR VENTRICULAR-PACED RHYTHM AT THE TIME OF THE OCCURRENCE, WHICH DIDN'T MEET THE ALARM THRESHOLDS. THE ARRHYTHMIA ALARM FUNCTION IS PERFORMED PER SPECIFICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21183934&pc=MSX | 2221819-2025-00002 | 2021-12-20 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-01-17 05:00 | MSX | BENEVISION | Device Alarm System | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, APPROXIMATELY BETWEEN 11:00 AM AND 12:30 PM, IN ROOM ID (B)(6), ON A PATIENT PLACED ON THE TM80 TELEMETRY MONITOR (ID TELE50) THE CENTRAL MONITORING UNIT (CMU) DEPARTMENT MONITOR TECHNICIAN NOTED THE PATIENT WAS IN PACED RHYTHM AT RATE 70 AT 11:00 AM. THE CMU WAS NOTIFIED BY GATEWAY STAFF THAT THE PATIENT WAS BEING RESUSCITATED. THE MONITOR TECHNICIAN NOTED THAT QRS HAD WIDENED AND THAT THERE WAS NO ALARM. THE PATIENT EXPIRED. Manufacturer Narrative: SYSTEM LOGS AND DATA ASSOCIATED WITH THE OCCURRENCE WERE EVALUATED, AND NO MALFUNCTION WAS CONFIRMED FOR ANY DEVICES INVOLVED IN THE REPORT. THE REPORTED WIDENED QRS COMPLEXES AT THE TIME OF THE OCCURRENCE ARE THE REGULAR VENTRICULAR-PACED RHYTHM, WHICH DIDN'T MEET THE ALARM THRESHOLDS. THE ARRHYTHMIA ALARM FUNCTION PERFORMED PER SPECIFICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186018&pc=MNS | 2518422-2025-100890 | 2025-01-13 05:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN CONTACTED BY A USER CLINIC WITH AN ALLEGATION THAT A PATIENT OF THEIRS HAS PASSED AWAY. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED "THE MANUFACTURER HAS BEEN CONTACTED BY A USER CLINIC WITH AN ALLEGATION THAT A PATIENT OF THEIRS HAS PASSED AWAY. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH." THE DEVICE HAS BEEN RETURNED AND EVALUATED BY THE MANUFACTURER. THE DEVICE EVALUATION FOUND NO EVIDENCE OF PRODUCT MALFUNCTION IN WHICH THE ABILITY OF THE DEVICE TO DELIVER THE PRESCRIBED THERAPY TO THE PUBLISHED SPECIFICATIONS WOULD BE IMPACTED. IN ADDITION, THERE IS NO EVIDENCE THAT THE DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY. THIS COMPLAINT IS REPORTABLE DUE TO THE ALLEGATION OF THE DEATH ONLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186106&pc=BZD | 2518422-2025-004435 | 2023-07-05 04:00 | Death | RESPIRONICS, INC. | 2025-01-17 05:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REP DREAMSTATION AUTO CPAP, DOM - RECRT. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21165023&pc=BZD | 2518422-2025-003129 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Nausea; Vomiting; Dizziness; Chronic Obstructive Pulmonary Disease (COPD); Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, KIDNEY DISEASE/TOXICITY, DEATH, OTHER AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21165226&pc=BZD | 2518422-2025-003132 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN BOX B. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN BOX B: ADVERSE EVENT/PRODUCT PROBLEM, OUTCOMES ATTRIBUTED TO AE AND DATE OF DEATH HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21165230&pc=BZD | 2518422-2025-002796 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Asthma; Headache; Dizziness | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, ASTHMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN BOX B. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN BOX B: ADVERSE EVENT/PRODUCT PROBLEM AND OUTCOMES ATTRIBUTED TO AE HAS BEEN CORRECTED. 510K NUMBER HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21165600&pc=BZD | 2518422-2025-003224 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21165615&pc=BZD | 2518422-2025-003251 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21165710&pc=BZD | 2518422-2025-003258 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166075&pc=BZD | 2518422-2025-003255 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Nausea; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER, DEATH AND RIGHT KIDNEY X2. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166087&pc=BZD | 2518422-2025-003350 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, DIZZINESS AND/OR HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166093&pc=BZD | 2518422-2025-003329 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Fibrosis; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, PULMONARY FIBROSIS AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166163&pc=BZD | 2518422-2025-003355 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166182&pc=BZD | 2518422-2025-003392 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND THROAT CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166379&pc=BZD | 2518422-2025-003434 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166504&pc=BZD | 2518422-2025-003422 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166642&pc=BZD | 2518422-2025-003528 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Nausea; Vomiting; Dizziness | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166647&pc=BZD | 2518422-2025-003516 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166731&pc=BZD | 2518422-2025-003204 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN BOX B. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN BOX B: OUTCOMES ATTRIBUTED TO AE HAS BEEN CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166974&pc=BZD | 2518422-2025-003628 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21166985&pc=BZD | 2518422-2025-003621 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Gastroesophageal Burn | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND ESOPHAGEAL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167072&pc=BZD | 2518422-2025-003606 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Bacterial Infection; Headache; Inflammation; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Fungal Infection; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE AND REDUCED CARDIOPULMONARY RESERVE. SHE HAD MOLD AND ESCHERICHIA COLI IN HER LUNGS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167157&pc=BZD | 2518422-2025-003602 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Hypersensitivity/Allergic reaction; Inflammation; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, AND/OR HEADACHE, HYPERSENSITIVITY, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167172&pc=MNS | 2518422-2025-003662 | 2022-02-15 05:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | MNS | BIPAP AUTOSV ADV | Degraded | Insufficient Information | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER FOAM PARTICLES WERE NOT FOUND. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167252&pc=BZD | 2518422-2025-003677 | 2023-02-03 05:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167276&pc=DTZ | 2184009-2025-00038 | 2024-12-27 05:00 | Death | PERFUSION SYSTEMS | 2025-01-16 05:00 | DTZ | FUSION HOLLOW FIBER OXYGENATOR | Increase in Pressure; Obstruction of Flow | Thrombosis/Thrombus | K240666 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A FUSION OXGENATOR AND EOPA 3D ARTERIAL CANNULA, THERE WAS AN INCREASE IN PRESSURE. A THROMBUS WAS IDENTIFIED AT THE OUTLET OF THE OXYGENATOR AND IN THE ARTERIAL CANNULA. AT THE START OF CARDIOPULMONARY BYPASS, THE FLOW INCREASED FROM THE START TO 3 L/MIN, BUT IT WAS NOT POSSIBLE TO INCREASE FURTHER, LEADING TO THE CESSATION OF CARDIOPULMONARY BYPASS. THE PRODUCT WAS REPLACED TO COMPLETE THE PROCEDURE. IT WAS NOTED THAT THE PATIENT WAS OPERATED ON AN AORTIC VALVE AND A BYPASS AND WAS THEN DISCHARGED TO THE INTENSIVE CARE UNIT FOR MONITORING (MAINLY NEUROLOGICAL). NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE HEALTH OF THE PATIENT DETERIORATED, THE PATIENT DIED AFTER A FEW DAYS OF REANIMATION. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE THROMBOSIS WAS LOCATED ON THE OUTSIDE OF THE FUSION MEMBRANE OVER APPROXIMATELY ¼ OF ITS SURFACE. THE DURATION OF THE CIRCULATION EXTRA-CORPORELLE (CEC) WAS APPROXIMATELY 2MINUTES AND 30 SECONDS, WITH A GRADUAL START TO HOMOGENIZE THE PRIMING, THEN IT WAS IMPOSSIBLE TO INCREASE THE FLOW RATE DESPITE ROTATIONS OF 3800 RPM ON THE MAQUET ROTAFLOW CENTRIFUGAL PUMP (MAXIMUM ROTATION 5000 RPM). THE PRIMING SOLUTION USED WAS ISOFUNDINE, 1 CGR, CEFAZOLIN 1G, HEPARIN 5000IU. THE PATIENT HAD AN INTRACAVITARY THROMBUS REQUIRING THE INSERTION OF CURATIVE DOSE HEPARIN 48 HOURS BEFORE THE PROCEDURE AND HAD A DISTURBED COAGULATION PROFILE WITH NOTABLY A FIBRINOGEN AT 6.4G. ANTICOAGULATION WAS ASSESSED WITH A JUNIOR HEMOCHRON AT 414S, AN ADDITIONAL DOSE OF HEPARIN WAS ADDED. THE PATIENT HAD NUMEROUS THROMBOSES (SPLEEN, LIVER) LEADING TO MULTI-ORGAN FAILURE AND THE PATIENT'S DEATH. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT "REANIMATION" MEANS "POST-OPERATIVE DEPARTMENT INTENSIVE CARE". Manufacturer Narrative: DEVICE EVALUATION SUMMARY: VISUAL INSPECTION SHOWS EVIDENCE OF THROMBUS/CLOTS. PRESSURE INTEGRITY TESTING SHOWS NO INTERNAL OR EXTERNAL LEAKS WHEN RUN AT 3 LPM WITH 23 PSI, (1189 MMHG) OF BACK PRESSURE FOR 10 MINUTES. THE DEVICE WAS SENT TO THE BLOOD LAB FOR PERFORMANCE TESTING. TESTING WAS CONDUCTED AT A 1:1 RATIO (7LPM BLOOD & GAS FLOWS) THE RESULTS ARE AS FOLLOWS: 189 MMHG BLOOD SIDE PRESSURE DROP REASON FOR RETURN WAS NOT CONFIRMED. ADDITIONAL INFO B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT HAD AN INTRACAVITARY THROMBUS. THE PATIENT WAS ON PREOPERATIVE ANTICOAGULANT THERAPY WITH HOME BLOOD PRESSURE MONITORING (HBPM) AND HEPARIN. THERE WAS NO FURTHER INVESTIGATION FOR COAGULATION ABNORMALITIES PRIOR TO THE SURGERY. THE PATIENT WAS CONSIDERED HIGH RISK FOR AN AORTIC VALVE SURGERY AND BYPASS. A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE (TAVI) WAS CONSIDERED. THE THROMBUS DID NOT INCREASE BUT WAS STILL PRESENT. THE BREAST GRAFT WAS HARVESTED. THE DEPARTMENT HEAD WAS CONSULTED AND THE DECISION TO PROCEED WITH THE PROCEDURE WAS MADE COLLEGIALLY. THE PATIENT PASSED AWAY A FEW DAYS AFTER INTENSIVE CARE. ADDITIONAL INFO D4 (D): ADDED THE SERIAL NUMBER ADDITIONAL INFO H6: UPDATED THE ANNEX B AND ANNEX C CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE DIFFERENT HEPARIN INJECTIONS ADMINISTERED WERE A HEPARIN CHOAY INITIAL BOLUS OF 24000 IU, AN ADDITION BOLUS OF 2500 IU OF HEPARIN THROUGH CENTRAL-VENOUS CATHETERISATION (CVC) AND BOLUS OF 5000 IU OF HEPARIN IN THE PRIMING OF THE CIRCUIT. THE ACT MEASUREMENT BEFORE THE START OF THE CEC WAS 417 SECONDS. THE INTERVENTION STARTED, THE PATIENT ENTERED INTO THE ROOM AT 7:50 AM AND INDUCTION WAS AT 8:30 AM. THE SURGICAL TIME STARTED AT 9:08 AM. A HEPARIN INJECTION WAS GIVEN AT 10:1 7AM. THE ACT CONTROL WAS AT 10:22 AM. CEC WAS STARTED AT 10:51 AM. THE INCIDENT OCCURRED AT 10:53 AM. HEPARINEMIA WAS GREATER THAN 1.1IU/ML AT 3.95%. THE PATIENT WAS HEPARIN-INDUCED THROMBOCYTOPENIA (HIT) NEGATIVE. CORRECTION H6.3 (EVAL CODE RESULT (FDR/ANNEX C)): THIS CODE HAS BEEN UPDATED. B2 DATE OF DEATH: THE DATE OF DEATH IS MONTH AND YEAR VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFO B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CGR STANDS FOR "CONCENTRATED RED BLOOD CELL". MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: NOTE ON D4: THERE ARE THREE POSSIBLE LOT NUMBERS FOR THE BB811 OXYGENATOR: (B)(6). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167285&pc=DWF | 2184009-2025-00039 | 2024-12-27 05:00 | Death | PERFUSION SYSTEMS | 2025-01-16 05:00 | DWF | EOPA 3D CANNULA | Increase in Pressure; Obstruction of Flow | Thrombosis/Thrombus | K150422 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A FUSION OXYGENATOR AND EOPA 3D ARTERIAL CANNULA, THERE WAS AN INCREASE IN PRESSURE. A THROMBUS WAS IDENTIFIED AT THE OUTLET OF THE OXYGENATOR AND IN THE ARTERIAL CANNULA. AT THE START OF CARDIOPULMONARY BYPASS, THE FLOW INCREASED FROM THE START TO 3 L/MIN, BUT IT WAS NOT POSSIBLE TO INCREASE FURTHER, LEADING TO THE CESSATION OF CARDIOPULMONARY BYPASS. THE PRODUCT WAS REPLACED TO COMPLETE THE PROCEDURE. IT WAS NOTED THAT THE PATIENT WAS OPERATED ON AN AORTIC VALVE AND A BYPASS, AND WAS THEN DISCHARGED TO THE INTENSIVE CARE UNIT FOR MONITORING (MAINLY NEUROLOGICAL). NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE HEALTH OF THE PATIENT DETERIORATED, THE PATIENT DIED AFTER A FEW DAYS OF REANIMATION. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE THROMBOSIS WAS LOCATED ON THE OUTSIDE OF THE FUSION MEMBRANE OVER APPROXIMATELY ¼ OF ITS SURFACE. THE DURATION OF THE CIRCULATION EXTRA-CORPORELLE (CEC) WAS APPROXIMATELY 2MINUTES AND 30 SECONDS, WITH A GRADUAL START TO HOMOGENIZE THE PRIMING, THEN IT WAS IMPOSSIBLE TO INCREASE THE FLOW RATE DESPITE ROTATIONS OF 3800 RPM ON THE MAQUET ROTAFLOW CENTRIFUGAL PUMP (MAXIMUM ROTATION 5000 RPM). THE PRIMING SOLUTION USED WAS ISOFUNDINE, 1 CGR, CEFAZOLIN 1G, HEPARIN 5000IU. THE PATIENT HAD AN INTRACAVITARY THROMBUS REQUIRING THE INSERTION OF CURATIVE DOSE HEPARIN 48 HOURS BEFORE THE PROCEDURE AND HAD A DISTURBED COAGULATION PROFILE WITH NOTABLY A FIBRINOGEN AT 6.4G. ANTICOAGULATION WAS ASSESSED WITH A JUNIOR HEMOCHRON AT 414S, AN ADDITIONAL DOSE OF HEPARIN WAS ADDED. THE PATIENT HAD NUMEROUS THROMBOSES (SPLEEN, LIVER) LEADING TO MULTI-ORGAN FAILURE AND THE PATIENT'S DEATH. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT "REANIMATION" MEANS "POST-OPERATIVE DEPARTMENT INTENSIVE CARE". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFO B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT HAD AN INTRACAVITARY THROMBUS. THE PATIENT WAS ON PREOPERATIVE ANTICOAGULANT THERAPY WITH HOME BLOOD PRESSURE MONITORING (HBPM) AND HEPARIN. THERE WAS NO FURTHER INVESTIGATION FOR COAGULATION ABNORMALITIES PRIOR TO THE SURGERY. THE PATIENT WAS CONSIDERED HIGH RISK FOR AN AORTIC VALVE SURGERY AND BYPASS. A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE (TAVI) WAS CONSIDERED. THE THROMBUS DID NOT INCREASE BUT WAS STILL PRESENT. THE BREAST GRAFT WAS HARVESTED. THE DEPARTMENT HEAD WAS CONSULTED AND THE DECISION TO PROCEED WITH THE PROCEDURE WAS MADE COLLEGIALLY. THE PATIENT PASSED AWAY A FEW DAYS AFTER INTENSIVE CARE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE DIFFERENT HEPARIN INJECTIONS ADMINISTERED WERE A HEPARIN CHOAY INITIAL BOLUS OF 24000 IU, AN ADDITION BOLUS OF 2500 IU OF HEPARIN THROUGH CENTRAL-VENOUS CATHETERISATION (CVC) AND BOLUS OF 5000 IU OF HEPARIN IN THE PRIMING OF THE CIRCUIT. THE ACT MEASUREMENT BEFORE THE START OF THE CEC WAS 417 SECONDS. THE INTERVENTION STARTED, THE PATIENT ENTERED INTO THE ROOM AT 7:50 AM AND INDUCTION WAS AT 8:30 AM. THE SURGICAL TIME STARTED AT 9:08 AM. A HEPARIN INJECTION WAS GIVEN AT 10:1 7AM. THE ACT CONTROL WAS AT 10:22 AM. CEC WAS STARTED AT 10:51 AM. THE INCIDENT OCCURRED AT 10:53 AM. HEPARINEMIA WAS GREATER THAN 1.1IU/ML AT 3.95%. THE PATIENT WAS HEPARIN-INDUCED THROMBOCYTOPENIA (HIT) NEGATIVE. B2 DATE OF DEATH: THE DATE OF DEATH IS MONTH AND YEAR VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFO B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CGR STANDS FOR "CONCENTRATED RED BLOOD CELL". MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167456&pc=DSQ | 2916596-2025-00274 | 2020-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-16 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "PRESSURE-DIMENSION INDEX AND LEFT VENTRICULAR SPHERICITY INDEX FOLLOWING HEARTMATE II AND HEARTMATE 3 IMPLANTATION" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT VENTRICULAR FAILURE, INFECTION, BLEEDING, STROKE, PUMP THROMBOSIS, AND DEATH. THIS STUDY WAS A RETROSPECTIVE ANALYSIS OF PROSPECTIVELY COLLECTED DATA FROM 50 PATIENTS IMPLANTED WITH HEARTMATE II (HMII) OR HM3 FROM 2012 TO 2020. 50 PATIENTS WITH HM3 IMPLANTATION WERE SCREENED, OF WHICH ONLY 24 WERE USED FOR DATA COLLECTING DUE TO THE FOLLOWING EXCLUSION CRITERIA: POOR-QUALITY ECHOCARDIOGRAPHY (ECHO) STUDIES AND THE ABSENCE OF FOLLOW-UP ECHO. A TOTAL OF 24 PATIENTS (20 (83%) MEN, 4 (17%) WOMEN), HAD A MEAN AGE OF 60.8 ± 9.7 YEARS AND A BODY MASS INDEX (BMI) OF 27.4 ± 6.5 KG/M2. PRE-EXISTING CONDITIONS AND CO-MORBIDITIES INCLUDED: DILATED CARDIOMYOPATHY (8 PATIENTS), ISCHEMIC CARDIOMYOPATHY (13 PATIENTS), NYHA > CLASS II (19 PATIENTS), INSULIN-DEPENDENT DIABETES MELLITUS (4 PATIENTS), PERIPHERAL ARTERY DISEASE (2 PATIENTS), TRANSIENT ISCHEMIC ATTACK (1 PATIENT), KAWASAKI DISEASE (13 PATIENTS), DIALYSIS (2 PATIENTS), CARDIOVASCULAR DISEASE (4 PATIENTS), PULMONARY HYPERTENSION (12 PATIENTS), PREVIOUS CARDIAC SURGERY (4 PATIENTS), STENT (11 PATIENTS), CHRONIC OBSTRUCTIVE PULMONARY DISEASE (4 PATIENTS), AND INTERMACS CLASS I (2 PATIENTS), CLASS II (3 PATIENTS), CLASS III (7 PATIENTS), CLASS IV (11 PATIENTS). HM3 PATIENTS HAD BOTH A SIGNIFICANTLY LONGER ICU STAY (15 ± 15 DAYS, P = 0.031) AND A SIGNIFICANTLY HIGHER SURVIVAL RATE AT 24 MONTHS (LOG-RANK TEST, P = 0.042). THERE WAS A WEAK CORRELATION BETWEEN THE PRESSURE-DIMENSION INDEX (PDI) AT 6 AND 12 MONTHS AFTER IMPLANTATION AND THE MORTALITY RATE (6 M: R = 0.3, P = 0.043; 12 M: R = 0.312, P = 0.035), WHEREBY A LOWER PDI WAS ASSOCIATED WITH A HIGHER SURVIVAL RATE. THE CORRELATION ANALYSIS ALSO REVEALED A WEAK CORRELATION BETWEEN THE SPHERICITY INDEX AT 6 MONTHS POST-OPERATIVELY AND THE NEED FOR DIALYSIS AFTER IMPLANTATION (R = 0.324, P = 0.0026). THERE WAS ALSO A WEAK NEGATIVE CORRELATION WITH THE OCCURRENCE OF SEPSIS (R = -00.355, P = 0.031) AND RE-THORACOTOMY (R = -0.362, P = 0.028) 6 MONTHS POST-OPERATIVELY IN RELATION TO THE SYSTOLIC PULMONARY ARTERY PRESSURE (SPAP). ADVERSE EVENTS INCLUDED: PNEUMONIA (11 PATIENTS), SEPSIS (5), KIDNEY DISEASE (8 PATIENTS), ISCHEMIC STROKE (2 PATIENTS), RIGHT HEART FAILURE (2 PATIENTS), TRANSIENT ISCHEMIC ATTACK (1 PATIENT), GASTROINTESTINAL BLEEDING (2 PATIENTS), RE-THORACOTOMY (7 PATIENTS), DEATH (1 PATIENT). LEFT VENTRICULAR END-DIASTOLIC DIAMETER (LVEDD) [PRE-OP] 6.7 ± 1.3 CM VS. [POST OP] 5.5 ± 0.9 CM; RIGHT VENTRICULAR FRACTIONAL AREA CHANGE (RVFAC) 35 ± 12%, TRICUSPID ANNULAR PLANE SYSTOLIC EXCURSION (TAPSE) 13.9 ± 1.9 MM, TISSUE DOPPLER ESTIMATED TRICUSPID ANNULAR SYSTOLIC VELOCITY (TASV) 11.5 ± 2.7 MM/S, LEFT VENTRICULAR SPHERICITY INDEX (LVSI) 0.8 ± 0.2, PDI [PRE-OP] 3.4 ± 1.4 MMHG/CM2 VS. [POST-OP 6 MONTHS] 2.4 ± 0.7 MMHG/CM2, P = 0.012; [PRE-OP] 3.4 ± 1.4 MMHG/CM2 VS. [POST-OP 12 MONTHS] 2.0 ± 0.8 MMHG/CM2, P < 0.01. THE ROC ANALYSIS OF PDI AS A PREDICTOR OF POST-OPERATIVE RIGHT VENTRICULAR FAILURE (RVF) SHOWED AN AREA UNDER THE CURVE OF 0.921 WITH A SPECIFICITY OF 0.976 BUT LOW SENSITIVITY 0.400. THE DESIGN OF HM3 OFFERED BETTER GEOMETRICAL PRESERVATION OF THE LV AND ENABLED NORMAL PDI VALUES, LEADING TO IMPROVED RV FUNCTION, AS INDICATED BY BETTER RVFAC, TAPSE, AND TASV VALUES. THE USE OF PRE-OPERATIVE PDI AS AN ADDITIONAL TOOL FOR ESTABLISHED RISK SCORES MIGHT OFFER A BETTER PRE-OPERATIVE PREDICTOR OF RVF. THE ANALYSIS OF POST-OPERATIVE COMPLICATIONS SHOWED A SIGNIFICANTLY LONGER ICU STAY FOR PATIENTS WITH HM3 (HMII: 7 ± 7 DAYS VS. HM3: 15 ± 15 DAYS, P = 0.031) WITH A SIGNIFICANTLY HIGHER SURVIVAL RATE AT 24 MONTHS OF HM3 PATIENTS (LOG-RANK TEST, P = 0.042). LVEDD SIGNIFICANTLY DECREASED 12 MONTHS POST-OPERATIVELY COMPARED WITH THE PRE-OPERATIVE VALUES. PDI DECREASED SIGNIFICANTLY 12 MONTHS POST-OPERATIVELY AND WAS ASSOCIATED WITH BETTER SURVIVAL. TAPSE, TASV, AND RVFAC WERE SIGNIFICANTLY HIGHER 12 MONTHS AFTER SURGERY, REFLECTING BETTER RV FUNCTION. FROM ALL PRE-OPERATIVE ECHOCARDIOGRAPHIC PARAMETERS ONLY PDI WAS A PREDICTOR OF POST-OPERATIVE RVF. A HIGH PDI INDICATES HIGH RV AFTERLOAD AND RV DILATION WITH HIGH PULMONARY VASCULAR RESISTANCE AND LOW RIGHT ATRIAL PRESSURE. THE STUDY SHOWED THAT THE PDI DECREASED SIGNIFICANTLY 12 MONTHS AFTER HM3 IMPLANTATION AND WAS IN THE LOWER TERTILE (2.0 ± 0.8 MMHG/ M2). Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01JAN2020. DEPARTMENT OF CARDIAC SURGERY, DEPARTMENT OF THORACIC SURGERY, RWTH UNIVERSITY HOSPITAL AACHEN, RWTH AACHEN UNIVERSITY, AACHEN, GERMANY; DEPARTMENT OF CARDIOTHORACIC SURGERY, HEART CENTRE TRIER, BARMHERZIGEN BRUEDER HOSPITAL, TRIER, GERMANY. AUTHOR INFORMATION: L. SCHNOERING. SCHNOERING, L., KHATTAB, M. A., AKHYARI, P., MOZA, A., HANEYA, A., DIAB, A. H., ABUGAMEH, A., LOTFI, S., & ZAYAT, R. (2024). PRESSURE-DIMENSION INDEX AND LEFT VENTRICULAR SPHERICITY INDEX FOLLOWING HEARTMATE II AND HEARTMATE 3 IMPLANTATION. ESC HEART FAILURE, 11(5), 3012¿3022. HTTPS://DOI.ORG/10.1002/EHF2.14839. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167544&pc=BZD | 2518422-2025-003704 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167550&pc=BZD | 2518422-2025-003701 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167553&pc=BZD | 2518422-2025-003699 |  | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH . THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167558&pc=BZD | 2518422-2025-003695 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Asthma; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA NEW OR WORSENING, LIVER DISEASE TOXICITY, LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167810&pc=OZD | 1220648-2025-25611 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-16 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT WAS DECOMPENSATING AS IMPELLA WAS NOT OFFLOADING THE LEFT VENTRICLE. THE CP WAS EXCHANGED FOR THE IMPELLA 5.5. THE PATIENT IS STABLE AND SURVIVED THE EVENT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO THE PATIENT'S DEATH. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION FOR THE LOW PUMP FLOW HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. THERE WAS NO LOW PUMP FLOW SEEN IN THE DATA LOGS. NO MOTOR CURRENT SPIKE, PLACEMENT SIGNAL TREND, OR POSITIONING ALARMS SEEN. THE LOGS ONLY GO FROM 12/25 2:44 TO 12:25 3:19. DUE TO INSUFFICIENT INFORMATION, THE ROOT CAUSE OF THE LOW PUMP FLOW CANNOT BE DETERMINED. Manufacturer Narrative: B.1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611 TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM. B.2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611 TO REFLECT THE PATIENT OUTCOME DEATH AND ADDED DATE OF DEATH. B.5 ADDED PATIENT OUTCOME INFORMATION AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611. D.6B REVISED EXPLANT DATE WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.1 TYPE OF REPORTABLE EVENT WAS REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611 TO REFLECT THE PATIENT OUTCOME OF DEATH. H.6 HEALTH EFFECT - IMPACT CODE 1802 WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25611.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21167923&pc=BZD | 2518422-2025-003734 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21168061&pc=BZD | 2518422-2025-003765 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21168746&pc=MKJ | 1220908-2025-00186 | 2024-12-17 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-16 05:00 | MKJ | AED PLUS | Failure to Deliver Shock/Stimulation | Insufficient Information | P160015 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL UNITED KINGDOM FOR EVALUATION. THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE WAS PUT THROUGH EXTENSIVE FUNCTIONAL TESTING WITHOUT DUPLICATING THE REPORT. THE CLINICAL FILE FROM THE EVENT HAS BEEN OVERWRITTEN AND IS NOT AVAILABLE FOR REVIEW. THE DEVICE HISTORY FILE WAS REVIEWED AND OBSERVED MULTIPLE SELF-TEST PASSED INDICATING THE DEVICE WOULD BE READY FOR USE. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21168797&pc=CBK | MW5164810 | 2022-02-01 05:00 | Death | UNK | 2025-01-16 05:00 | CBK | VENTILATOR |  | Insufficient Information |  |  | 1 | Event Description: MY MOTHER PASSED AWAY DUE TO DEFECTIVE VENTILATOR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21168828&pc=NHA | MW5164811 | 2015-03-24 04:00 | Death | GLIDEWELL LABORATORIES | 2025-01-16 05:00 | NHA | ABUTMENT, IMPLANT, DENTAL, ENDOSSEOUS |  | Abscess; Failure of Implant |  |  | 1 | Event Description: I AM LOOKING TO FILE A FORMAL COMPLAINT AGAINST GLIDEWELL LABORATORIES ON BEHALF OF MY LATE FATHER, (B)(6). MY FATHER WAS A PATIENT OF (B)(6), DMD ("DR. (B)(6)") ((B)(6)) WHO USED GLIDEWELL LABORATORIES (NEWPORT BEACH, CA) TO MANUFACTURE A STRAUMANN BRAND ABUTMENT+CROWN FOR MY FATHER'S TOOTH #30 STRAUMANN DENTAL IMPLANT POST (HARDCOPY LAB ORDER WAS SENT ON (B)(6) 2015). GLIDEWELL LABORATORIES PROVIDED AN INFERIOR ABUTMENT WHICH PROVED TO FAIL MULTIPLE TIMES IN MY FATHER'S MOUTH. THEY DID NOT FOLLOW THE (B)(6) 2015 LAB ORDERS WHICH SPECIFICALLY MENTIONED "USE STRAUMANN PARTS". THE ABUTMENT REPEATEDLY BECAME LOOSE OVER THE COURSE OF APPROX. 2.5 YEARS WHICH CAUSED FRICTION TO THE GUMLINE. BREAKING OFF THE CROWN MULTIPLE TIMES, OVER-TORQUING THE ABUTMENT AND THE OVERALL TRAUMA CENTRALLY LOCATED TO THAT TOOTH AREA CAUSED 2 ABSCESSES IN THE INTERPROXIMAL TISSUE OF #30/31 TO DEVELOP AS A RESULT (FIRST IDENTIFIED ON (B)(6) 2018). AFTER REPEATED FAILURE, STRAUMANN (ABUTMENT MANUFACTURER), SENT THEIR REPRESENTATIVE ((B)(4)) TO DR. (B)(6) OFFICE TO VISUALLY INSPECT THE COMPONENT ((B)(6) 2017). EXTRACTED FROM P. (B)(6) OF DENTAL CHART - "PATIENT PRESENTS FOR A NEW FINAL IMPLANT IMPRESSION OF TOOTH #30, IMPLANT TYPE STRAUMANN. STRAUMANN REP (B)(4) HERE BECAUSE OF CONTINUED ISSUES WITH THIS CROWN... PATIENT STATES HE DOES NOT WANT TO BE NUMB FOR THIS PROCEDURE. ATTEMPT TO ACCESS SCREW LOOSENED BY DR (B)(6). REMOVE BRUXZIR CROWN. ACCESS AND REMOVE SCREW AND ABUTMENT. (B)(6) INFORM DR NOT AUTHENTIC STRAUMANN ABUTMENT. (B)(6) WILL SEND ABUTMENT FOR SCREW RETAINED IMPLANT TO SEND WITH CASE. PLACE IMPRESSION USING LIGHT AND MEDIUM BODY MATERIAL..." ON (B)(6) 2017 A PHONE CALL NOTE IN CLINICAL NOTES INDICATED "SPOKE WITH DAT IN IMPLANT DEPARTMENT REUSING GENUINE STRAUMANN PARTS. INFORM HIM SENDING VARIOBASE SENT TO US FROM THE COMPANY. (B)(6) CASE ID # IS (B)(4)". (P. (B)(6) - DENTAL CHART). NOTE: NARRATIVE OF PHONE CONVERSION BETWEEN DR. (B)(6) AND (B)(6) WAS NEVER INCLUDED IN DENTAL RECORD. ON (B)(6) 2017 - DR. (B)(6) LAB ORDER TO GLIDEWELL LABORATORIES FOR TOOTH #30 INDICATED "ATTN: (B)(6) #30 IMPLANT REDO DR. EMAIL YOU ABOUT THIS CASE. REMAKE AS SCREW RETAINED ABUT+CRN WITH ENCLOSED STRAUMANN ABUTMENT X-RAY AND SURG REPORT THANKS" (STRAUMANN IS MISSPELLED ON LAB ORDER). (P. (B)(6) - LAB ORDERS). NOTE: COPY OF EMAIL COMMUNICATION BETWEEN DR. (B)(6) AND (B)(6) WAS NEVER INCLUDED IN DENTAL RECORD. ON (B)(6) 2017, FINAL STRAUMANN ABUTMENT+CROWN FOR TOOTH #30 WAS PLACED BY DR. (B)(6). NOTE: ATTORNEY (B)(6) HAS THIS PHYSICAL IMPLANTED TOOTH (IT FELL OUT PRIOR TO SURGERY), HOWEVER, WE BELIEVE IT IS A STRAUMANN ABUTMENT+CROWN AS (B)(6) PROVIDED IT. THE INFERIOR ABUTMENT WAS DISCARDED BY DR. (B)(6). AT THIS TIME, THE ONLY PERSON WHO WAS ABLE TO CONFIRM THE AUTHENTICITY OF THE ABUTMENT WAS (B)(6). THERE WERE NO ABUTMENT MANUFACTURER DETAILS IN THE DENTAL RECORD (NO IDENTIFIER STICKERS, LOT OR CATALOG NUMBERS WERE FOUND). ATTACHED PDF TITLED "(B)(6) - FORMAL COMPLAINT AGAINST GLIDEWELL LABORATORIES.PDF" OUTLINES THE COMPREHENSIVE SEQUENCE OF EVENTS THAT OCCURRED UNDER DR. (B)(6) CARE. I'VE ALSO ATTACHED PHOTOS AND OTHER DOCUMENTS TO SUPPORT THE EXTRACTED REFERENCED PAGES. (B)(6) DATE OF DEATH - (B)(6) 2022 (UNABLE TO INPUT IN #2 ABOVE).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169176&pc=MIH | 9612164-2025-00306 | 2024-09-06 04:00 | Death | MEDTRONIC IRELAND | 2025-01-16 05:00 | MIH | ENDURANT STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿ ENDOVASCULAR REPAIR OF RUPTURED ABDOMINAL AORTIC ANEURYSMS USING THE ENDURANT¿ ENDOGRAFT¿. THE TIME FRAME OF THIS STUDY WAS OVER A FIFTEEN-YEAR PERIOD. 88 PATIENTS WHO UNDERWENT ENDOVASCULAR TREATMENT FOR RUPTURED ABDOMINAL AORTIC ANEURYSMS WERE INCLUDED. THE MEAN ANEURYSM DIAMETER WAS 73.3MM. ENDURANT STENT GRAFTS WERE IMPLANTED IN THE PATIENT POPULATION. PATIENT MORTALITY WAS REPORTED OF WHICH ONE OF THE DEATHS WAS REPORTED AS RELATED TO A MALFUNCTION OF THE ENDURANT STENT GRAFT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿ ENDOVASCULAR REPAIR OF RUPTURED ABDOMINAL AORTIC ANEURYSMS USING THE ENDURANT¿ ENDOGRAFT¿ NANA, P.; VOLAKAKIS, G.; SPANOS, K.; KOUVELOS, G.; BAREKA, M.; ARNAOUTOGLOU, E.; GIANNOUKAS, A.; MATSAGKAS, M JOURNAL OF CLINICAL MEDICINE. 2024, 13, 5282. HTTPS://DOI.ORG/10.3390/JCM13175282 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169651&pc=DSQ | 2916596-2025-00312 | 2024-09-20 04:00 | Death | THORATEC CORPORATION | 2025-01-16 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169655&pc=OZD | 1220648-2025-25614 | 2025-01-03 05:00 | Death | ABIOMED, INC. | 2025-01-16 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE IMPELLA SHEATH WAS INSERTED WITH LARGE AMOUNT OF BLEEDING WHEN ATTEMPTING TO DO SINGLE ACCESS TECHNIQUE. THE INTERVENTION FOR THE BLEED IS UNKNOWN. CARE WAS WITHDRAWN AFTER 24.05 HOURS OF SUPPORT, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170238&pc=BZD | 2518422-2025-003810 | 2024-01-04 05:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | REP DREAMSTATION AUTO BIPAP | Degraded | Asthma; Headache; Hypersensitivity/Allergic reaction; Inflammation; Dizziness; Cancer; Unspecified Heart Problem; Unspecified Respiratory Problem; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, ASTHMA, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170303&pc=HCG | 3014162263-2024-00107 | 2024-12-04 05:00 | Death | BALT USA | 2025-01-16 05:00 | HCG | OPTIMA COIL SYSTEM | Stretched; Migration; Premature Separation | Ischemia Stroke; No Clinical Signs, Symptoms or Conditions | K200030 |  | 1 | Event Description: IT WAS REPORTED THAT: " WHEN NAVIGATING TO RELEASE WHAT WOULD BE THE LAST COIL IN THE ANEURYSM LOCATED IN THE MIDDLE CEREBRAL ARTERY, THE COIL DETACHED ON ITS OWN, BEFORE THE DESIRED LOCATION. ACCORDING TO DR. MARCIO, THE COIL WAS STRETCHED ALONG WITH THE BOVINE ARCH. THE DOCTOR WAS VERY UNCOMFORTABLE WITH THE SITUATION." INCIDENT REPORT FORM SUBMITTED FOR THIS COMPLAINT INDICATES THAT NO PATIENT INJURY WAS SUSTAINED. UPDATE 27DEC2024 - ADDITIONAL INFORMATION RECEIVED FROM THE ISSUER: "1. DID THE PROCEDURE INVOLVE TORTUOUS ANATOMY? NOT VERY MUCH. 2. WILL THE MICROCATHETER BE AVAILABLE FOR RETURN? NO. 3. WHERE DID THE IMPLANT PREMATURELY DETACH? (INSIDE THE AN, INSIDE THE MC?, ETC) THE COIL DETACHED AND BECOME LODGED IN THE BOVINE ARCH. 4. WERE ANY ATTEMPTS MADE TO DETACH THE IMPLANT COIL USING THE DETACHMENT CONTROLLER? NO 5. CAN YOU CONFIRM IF THE DEVICE WAS IMPLANTED? ACCORDING TO INFORMATION FROM DR. MARCIO, THE COIL WAS STRETCHED OUT OF THE INTENDED LOCATION BECAUSE IT DETACHED PREMATURELY."Event Description: IT WAS REPORTED THAT: " WHEN NAVIGATING TO RELEASE WHAT WOULD BE THE LAST COIL IN THE ANEURYSM LOCATED IN THE MIDDLE CEREBRAL ARTERY, THE COIL DETACHED ON ITS OWN, BEFORE THE DESIRED LOCATION. ACCORDING TO DR. MARCIO, THE COIL WAS STRETCHED ALONG WITH THE BOVINE ARCH. THE DOCTOR WAS VERY UNCOMFORTABLE WITH THE SITUATION." INCIDENT REPORT FORM SUBMITTED FOR THIS COMPLAINT INDICATES THAT NO PATIENT INJURY WAS SUSTAINED. UPDATE 27DEC2024 - ADDITIONAL INFORMATION RECEIVED FROM THE ISSUER: "1. DID THE PROCEDURE INVOLVE TORTUOUS ANATOMY? NOT VERY MUCH. 2. WILL THE MICROCATHETER BE AVAILABLE FOR RETURN? NO. 3. WHERE DID THE IMPLANT PREMATURELY DETACH? (INSIDE THE AN, INSIDE THE MC?, ETC) THE COIL DETACHED AND BECOME LODGED IN THE BOVINE ARCH. 4. WERE ANY ATTEMPTS MADE TO DETACH THE IMPLANT COIL USING THE DETACHMENT CONTROLLER? NO 5. CAN YOU CONFIRM IF THE DEVICE WAS IMPLANTED? ACCORDING TO INFORMATION FROM DR. MARCIO, THE COIL WAS STRETCHED OUT OF THE INTENDED LOCATION BECAUSE IT DETACHED PREMATURELY." UPDATE 20FEB2025 - ADDITIONAL INFORMATION RECEIVED FROM THE ISSUER: "1. IS THE EMBOLIZATION COIL THAT BECAME LODGED IN THE BOVINE ARCH STILL REMAINING IN THE PATIENT'S BODY? THE PATIENT DIED A FEW WEEKS AFTER THE PROCEDURE. THE COIL HAD MIGRATED TO THE MIDDLE CEREBRAL ARTERY! 2. ARE THERE ANY PLANS FOR ADDITIONAL PROCEDURES REGARDING THIS? NO LONGER... THERE WAS A SCHEDULE TO REVIEW THE PROCEDURE, BUT THE PATIENT DIED BEFORE THAT." UPDATE 19MAR2025 - ADDITIONAL INFORMATION RECEIVED FROM THE ISSUER: "AS THE PATIENT PASSED A FEW WEEKS FOLLOWING THE PROCEDURE, IS THE CAUSE KNOWN? THE PATIENT EVOLVED TO BRAIN DEATH AFTER EXTENSIVE STROKE CAUSED BY OCCLUSION OF THE LEFT MIDDLE CEREBRAL ARTERY FROM THE MIGRATION OF THE DETACHED COIL INSIDE THE MICROCATHETER. I HAVE TO CONFIRM THE DATES, BUT IT WAS A FEW DAYS AFTER THE PROCEDURE." Manufacturer Narrative: BALT USA REFERENCE #(B)(4). AN EVALUATION OF THE ACTUAL COMPLAINT SAMPLE COULD NOT BE PERFORMED AS DEVICE WAS UNAVAILABLE TO BE RETURN. IF THE DEVICE BECOMES AVAILABLE IN THE FUTURE, THE COMPLAINT FILE WILL BE REVISITED AND THE INVESTIGATION RESULTS UPDATED ACCORDINGLY. BASED ON THE PROVIDED INFORMATION, ROOT CAUSE COULD NOT BE DEFINITIVELY DETERMINED. LACK OF DEVICE RETURN PREVENTED DEEPER EVALUATION OF THE REPORTED ISSUE. REVIEW OF THE LOT HISTORY RECORDS DID NOT REVEAL ANY IN-PROCESS OR LOT-SPECIFIC ISSUE THAT COULD ACCOUNT FOR THE OBSERVATION. NO ADDITIONAL COMPLAINTS AGAINST LOT NUMBER F220400526 HAS BEEN MADE FOR THE SAME ISSUE. COMPREHENSIVE ANALYSIS OF THIS FAILURE MODE HAS REMAINED SUBJECT TO MONITORING FOR ANY UNACCEPTABLE INCREASE IN TREND. THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA, AND ITS RELEASE BY FDA, DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OF THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER, THEIR EMPLOYEES, OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT. Manufacturer Narrative: BALT USA REFERENCE # (B)(4). AN EVALUATION OF THE ACTUAL COMPLAINT SAMPLE COULD NOT BE PERFORMED AS DEVICE WAS UNAVAILABLE TO BE RETURN. BASED ON THE PROVIDED INFORMATION, ROOT CAUSE COULD NOT BE DEFINITIVELY DETERMINED. LACK OF DEVICE RETURN PREVENTED DEEPER EVALUATION OF THE REPORTED ISSUE. REVIEW OF THE LOT HISTORY RECORDS DID NOT REVEAL ANY IN-PROCESS OR LOT-SPECIFIC ISSUE THAT COULD ACCOUNT FOR THE OBSERVATION. NO ADDITIONAL COMPLAINTS AGAINST LOT NUMBER F220400526 HAS BEEN MADE FOR THE SAME ISSUE. COMPREHENSIVE ANALYSIS OF THIS FAILURE MODE HAS REMAINED SUBJECT TO MONITORING FOR ANY UNACCEPTABLE INCREASE IN TREND. THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA, AND ITS RELEASE BY FDA, DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OF THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER, THEIR EMPLOYEES, OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170375&pc=OZD | 1220648-2025-25616 | 2025-01-01 05:00 | Death | ABIOMED, INC. | 2025-01-16 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD OOZING/BLEEDING AT THE ACCESS SITE. A FEMSTOP WAS APPLIED TO THE ACCESS SITE AND ANTICOAGULANT WAS WITHHELD. CARE WAS WITHDRAWN AFTER 141.18 HOURS OF IMPELLA SUPPORT AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170573&pc=NPT | 2025587-2025-00441 | 2024-08-22 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-16 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Appropriate Term/Code Not Available | Cognitive Changes | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 2 MONTHS AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. ALTHOUGH SENILITY WAS NOTED IN RELATION TO THE DEATH, THE CAUSED OF DEATH WAS LISTED AS UNKNOWN. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170615&pc=DSQ | 2916596-2025-00072 | 2024-12-24 05:00 | Death | THORATEC CORPORATION | 2025-01-16 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY WHILE UNDER THE CARE OF HOSPICE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE CAUSE OF DEATH WAS HEART FAILURE AND DEMENTIA. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170628&pc=OZD | 1220648-2025-25619 | 2025-01-04 05:00 | Death | ABIOMED, INC. | 2025-01-16 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD LIMB ISCHEMIA. THE PATIENT WAS UNSTABLE AND EXPIRED AFTER 70.70 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170911&pc=BZD | 2518422-2025-003828 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Asthma |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGES DIAGNOSIS OF ASTHMA AND DEATH. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170987&pc=PYX | 1220648-2025-25622 | 2025-01-03 05:00 | Death | ABIOMED, INC. | 2025-01-16 05:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA RP FLEX SUPPORT AND DURING THE IMPLANT, SEQUENTIAL DILATION WAS PERFORMED AND WHILE INSERTING THE 23 FRENCH SHEATH, RESISTANCE WAS MET. THE SHEATH WOULD NOT ADVANCE SO A 23 FRENCH DILATOR WAS USED AND THE SHEATH WAS INSERTED. THE SHEATH HAD PERFORATED THE INTERNAL JUGULAR VEIN. THE SHEATH WAS REMOVED AND MANUAL PRESSURE WAS HELD. INSERTION WAS THEN PERFORMED THROUGH THE LEFT FEMORAL ACCESS. THE PATIENT EXPIRED AFTER 34.24 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. AN INTRODUCER AND A DILATOR WERE RETURNED FOR EVALUATION. UPON VISUAL INSPECTION, DAMAGE TO THE TIP OF THE INTRODUCER SHEATH WAS OBSERVED. ADDITIONALLY, THE TIP OF THE DILATOR ALSO SHOWED DAMAGE. CLINICAL DETAILS NOTED THAT WHEN INSERTING THE SHEATH INTO THE RIJ, RESISTANCE WAS MET AND THE PHYSICIAN PUSHED IN ORDER TO ATTEMPT TO PLACE THE INTRODUCER SHEATH. THE PATIENT WAS NOTED TO HAVE UNIQUE NECK ANATOMY WITH A SHORT NECK AND DEEP TISSUE. THE SHEATH WAS ABLE TO BE PLACED SUCCESSFULLY IN LEFT IJ. THE ROOT CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS MOST LIKELY DUE TO INTRODUCER SHEATH TIP SPLIT. D.9 DATE DEVICE RETURNED/INFORMATION WAS ADDED. Manufacturer Narrative: G6 TYPE OF REPORT; 30 DAY WAS INADVERTENTLY NOT SELECTED MANUFACTURER DEVICE REPORT 1220648-2025-25622-1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21171004&pc=NPT | 2025587-2025-00446 | 2024-12-10 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-16 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Intracranial Hemorrhage | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 11 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS HOSPITALIZED DYE TO A SUBARACHNOID HEMORRHAGE. 12 DAYS LATER, THE PATIENT DIED. THE CAUSE OF DEATH WAS RE-HEMORRHAGING. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THERE WAS NO RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED EVENT. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21171308&pc=QEW | 3005168196-2025-00016 | 2024-12-06 05:00 | Death | PENUMBRA, INC. | 2025-01-16 05:00 | QEW | INDIGO SYSTEM FLASH ASPIRATION CATHETER | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K241399 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE TO TREAT A DEEP VEIN THROMBOSIS (DVT) USING AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER). DURING THE PROCEDURE, WHILE THE PHYSICIAN WAS TREATING THE DVT WITH THE FLASH CATHETER, A MASSIVE PULMONARY EMBOLISM (PE) WAS OBSERVED. THE PHYSICIAN PERFORMED THROMBO-ASPIRATION USING A NON-PENUMBRA DEVICE TO TREAT THE PE AND NOTICED THAT THE OBSTRUCTIVE SHOCK WAS CLEARED. THE PROCEDURE ENDED AT THIS POINT. ON (B)(6) 2024, THE PATIENT DETERIORATED AND EXPIRED DUE TO MULTIPLE ORGAN FAILURE FOLLOWING THE MASSIVE PE. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT WAS IN A CRITICAL CONDITION THAN INITIALLY ASSESSED PRIOR TO THE PROCEDURE.Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE TO TREAT A DEEP VEIN THROMBOSIS (DVT) USING AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER). IT WAS NOTED THAT THE PATIENT HAD MORBID OBESITY, BAD RENAL FUNCTION, DEHYDRATION, AND WAS IN OVERALL POOR CONDITION. DURING THE PROCEDURE, WHILE THE PHYSICIAN WAS TREATING THE DVT WITH THE FLASH CATHETER, A PULMONARY EMBOLISM (PE) WAS OBSERVED. IT WAS REPORTED THE PE MAY HAVE BEEN PRESENT PRIOR TO THE PROCEDURE BUT WAS OVERLOOKED. THE PHYSICIAN PERFORMED THROMBO-ASPIRATION USING A NON-PENUMBRA DEVICE TO TREAT THE PE AND NOTICED THAT THE OBSTRUCTIVE SHOCK WAS CLEARED. THE PATIENT¿S CONDITION BRIEFLY IMPROVED AFTER TREATMENT. THE PROCEDURE ENDED AT THIS POINT. ON (B)(6) 2024, THE PATIENT DETERIORATED AND EXPIRED DUE TO MULTIPLE ORGAN FAILURE FOLLOWING THE PE. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT WAS IN A MORE CRITICAL CONDITION THAN INITIALLY ASSESSED PRIOR TO THE DVT PROCEDURE. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTIONS WERE INCORRECTLY REPORTED ON THE INITIAL MFR REPORT AND ARE BEING CORRECTED ON THIS FOLLOW-UP #02 MFR REPORT: 3005168196-2025-00016. 1. SECTION B. BOX 2. OUTCOMES ATTRIBUTED TO ADVERSE EVENT, 2. SECTION B. BOX 2. DATE OF DEATH, 3. SECTION B. BOX 5. DESCRIBE EVENT OR PROBLEM, 4. SECTION H. BOX 1. TYPE OF REPORTABLE EVENT, 5. SECTION H. BOX 6. HEALTH EFFECT IMPACT CODE 1, 6. SECTION H. BOX 6. HEALTH EFFECT IMPACT CODE 2. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTION WAS INCORRECTLY REPORTED ON THE INITIAL MFR REPORT AND IS BEING CORRECTED ON THIS FOLLOW-UP #01 MFR REPORT: 3005168196-2025-00016. 1. SECTION H. BOX. 11. ADDITIONAL MANUFACTURER NARRATIVE. THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21171849&pc=DSQ | 2916596-2025-00176 | 2025-01-04 05:00 | Death | THORATEC CORPORATION | 2025-01-16 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock; Fungal Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE DEATH WAS NOT CAUSED BY PUMP MALFUNCTION. THE CAUSE OF DEATH WAS SEPTIC SHOCK SECONDARY TO ENTEROCOCCUS FAECALIS BACTEREMIA AND CANDIDA ALBICANS ON BLOOD CULTURES. THE DEVICE OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21171874&pc=DSQ | 3007042319-2025-00068 | 2025-01-04 05:00 | Death | HEARTWARE, INC. | 2025-01-16 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Dyspnea; Pneumonia; Renal Failure; Respiratory Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY WAS ADMITTED WITH PNEUMONIA AND ACUTE RESPIRATORY FAILURE. THE PATIENT EXPERIENCED SHORTNESS OF BREATH AND DEVELOPED WORSENING RIGHT VENTRICULAR FAILURE AND KIDNEY FAILURE. THE PATIENT WAS TREATED WITH MECHANICAL VENTILATION AND CONTINUED MEDICATION. DESPITE THESE INTERVENTIONS, THE PATIENT DIED DUE TO MULTI-SYSTEM ORGAN FAILURE. Manufacturer Narrative: ### INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION.### MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTI-SYSTEM ORGAN FAILURE AFTER EXPERIENCING SHORTNESS OF BREATH IN ADDITION TO DEVELOPING WORSENING RIGHT VENTRICULAR FAILURE AND KIDNEY FAILURE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION, WORSENING HEART FAILURE, RENAL DYSFUNCTION, MULTI-ORGAN FAILURE, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF WORSENING HEART FAILURE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21171913&pc=BZD | 2518422-2025-003845 | 2023-01-18 05:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | REP DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21171954&pc=OZD | 1220648-2025-25582 | 2024-12-31 05:00 | Death | ABIOMED, INC. | 2025-01-16 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND FOLLOWING THE IMPLANT WOULD EXPERIENCE AN ACCESS SITE BLEEDING COMPLICATION. THE PATIENT CAME IN FOR SEPTIC SHOCK VERSUS CARDIOGENIC SHOCK AND CODED IN THE EMERGENCY DEPARTMENT. THE INTUBATED PATIENT CODED AND WAS SENT TO THE CATHETERIZATION LAB AND CODED AGAIN ON THE TABLE. ONE SHOCK WAS DELIVERED. THE PATIENT WAS IN CARDIOGENIC SHOCK AND WAS FOUND TO HAVE MULTIPLE VESSEL DISEASE. AN IMPELLA CP DEVICE WAS IMMEDIATELY INSERTED VIA THE RIGHT FEMORAL ARTERY. A TEMPORARY PACER WIRE WAS PLACED VIA THE RIGHT FEMORAL VEIN, AND THE PHYSICIAN INTERVENED THE CIRCUMFLEX WITH A STENT, RIGHT CORONARY WITH TWO STENTS AND RAMUS (TRANSLUMINAL ANGIOPLASTY) ARTERIES. HOWEVER, IT WAS REPORTED THE PHYSICIAN WAS UNABLE TO INTERVENE THE CHRONIC TOTALLY OCCLUDED LEFT ANTERIOR DESCENDING ARTERY. FOLLOWING INTERVENTION, THE TEMPORARY PACER WIRE WAS REMOVED, A SWAN-GANZ CATHETER WAS INSERTED, AND THE PATIENT WAS TRANSPORTED TO THE UNIT REMAINING ON IMPELLA MCS. A SHORT TIME THEREAFTER, THE RIGHT GROIN WAS OOZING. MANUAL PRESSURE WAS APPLIED, AND A FEMOSTOP WAS ADDED TO CONTROL THE BLEEDING. OF NOTE, ANTICOAGULATION WAS NOT USED DURING SUPPORT, AND IT IS UNKNOWN IF THIS WAS DUE TO THE PATIENT BLEEDING. THE NEXT DAY, THE FEMOSTOP REMAINED APPLIED, OOZING STOPPED, AND THE PATIENT HAD DEVELOPED A FEVER. THE PHYSICIAN AND FAMILY DISCUSSION OCCURRED AND THE DECISION WAS MADE TO CHANGE THE PATIENT'S STATUS TO DO NOT RESUSCITATE. SUPPORTIVE CARE WAS CONTINUED. THE IMPELLA WAS RUNNING AT SUPPORT LEVEL P-7. HOWEVER, THE FOLLOWING DAY THE FAMILY DECIDED ON COMFORT CARE, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE STATEMENT IN THE B5 NARRATIVE "THE USER FACILITY REPORTED A PATIENT" REFERS TO A PHYSICIAN IN A HOSPITAL SETTING WITH A PATIENT THAT EXPERIENCED AN ADVERSE EVENT WHILE AN IMPELLA DEVICE WAS IN USE. THE PHYSICIAN REPORTED THE ADVERSE EVENT DIRECTLY TO ABIOMED. THERE WAS NO USER REPORT THAT WAS SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172539&pc=FKX | 0008030665-2025-00138 | 2025-01-10 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-16 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6)2025 WHILE CONNECTED TO A LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY ON THE MORNING OF (B)(6)2025. THE PATIENT¿S DAUGHTER DISCOVERED THE PATIENT HAD EXPIRED WHEN SHE WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT REPORTEDLY PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6)2025 INDICATED THE PATIENT HAD COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PATIENT WAS UNDERGOING HOSPICE CARE AT HOME DUE TO TERMINAL CANCER AND WAS PERFORMING CCPD FOR COMFORT MEASURES ONLY. THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST SECONDARY TO HOSPICE CARE AND A SIGNIFICANT CARDIAC HISTORY. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION.Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE CONNECTED TO A LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY ON THE MORNING OF (B)(6) 2025. THE PATIENT¿S DAUGHTER DISCOVERED THE PATIENT HAD EXPIRED WHEN SHE WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT REPORTEDLY PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2025 INDICATED THE PATIENT HAD COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PATIENT WAS UNDERGOING HOSPICE CARE AT HOME DUE TO TERMINAL CANCER AND WAS PERFORMING CCPD FOR COMFORT MEASURES ONLY. THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST SECONDARY TO HOSPICE CARE AND A SIGNIFICANT CARDIAC HISTORY. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN SHE EXPIRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HOSPICE CARE AND A SIGNIFICANT CARDIAC HISTORY. SUDDEN CARDIAC DEATH (SCD) IS RESPONSIBLE FOR THE MAJORITY OF CARDIOVASCULAR-RELATED DEATHS IN PATIENTS WITH ESRD. ADDITIONALLY, HOSPICE CARE IS CLINICALLY DEFINED AS A MEDICAL INTERVENTION, WITH THE EXPECTED OUTCOME BEING THE PATIENT WILL EXPIRE AS A RESULT. PER THE PDRN, THE SERIOUS ADVERSE EVENT(S) WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). BASED ON THE INFORMATION AVAILABLE, THE LIBERTY CYCLER SET CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. ADDITIONALLY, THERE WAS NO REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172627&pc=NGV | 2124215-2024-83963 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Endocarditis |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ENDOCARDITIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 391 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172629&pc=NGV | 2124215-2024-83964 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 TIME TO EVENT: 374 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172632&pc=NGV | 2124215-2024-83934 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CARDIOVASCULAR PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 341 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172639&pc=NGV | 2124215-2024-83869 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. TIME TO EVENT: 157 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172735&pc=NPT | 2015691-2025-00362 | 2024-10-10 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-01-16 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Myocardial Contusion | P140031 |  | 1 | Event Description: AS REPORTED BY PATIENT SUPPORT, A FAMILY MEMBER OF THE PATIENT RECEIVED AN IMPLANT CARD. THE PATIENT PASSED AWAY 3 DAYS POST TAVR. AFTER SURGERY PATIENT WAS DIAGNOSED WITH ACUTE COMPARTMENT SYNDROME AND KIDNEY FAILURE AND WAS TRANSITIONED TO HOSPICE CARE. PATIENT PASSED AWAY DUE TO CARDIAC ARREST, COMPARTMENT SYNDROME, AND AORTIC SYNDROME. CARE ADVOCATE DID NOT ASK TO SPEAK WITH EW PHYSICIANS AND DID NOT MAKE ANY ALLEGATIONS TO OUR DEVICE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, MYOCARDIAL CONTUSION, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE (THE PATIENT'S ADVANCE AGE) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21172796&pc=DSQ | 3007042319-2025-00070 | 2024-12-16 05:00 | Death | HEARTWARE, INC. | 2025-01-16 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) EXHIBITED ABOVE NORMAL POWER AND HIGH WATT ALARMS AND VAD THROMBUS. THE SPECIFIC PATIENT SYMPTOMS OR COMPLICATIONS ASSOCIATED WITH THE EVENT WERE UNKNOWN. THE DEVICE REMAINS IN SERVICE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT NO LONGER ON SUPPORT AND SUBSEQUENTLY EXPIRED DUE TO BAD CLINICAL CONDITION. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED AN INCREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS STARTING ON (B)(6) 2024, LEADING TO PARAMETERS ABOVE NORMAL OPERATING RANGE, AND EIGHTEEN (18) HIGH WATT ALARMS LOGGED ON (B)(6) 2024. AS A RESULT, THE REPORTED HIGH POWER EVENT WAS CONFIRMED. OF NOTE, INFORMATION FROM THE SITE INDICATED THAT, AFTER EXPERIENCING THROMBUS, THE PATIENT EXPIRED "DUE TO BAD CLINICAL CONDITION." BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE AVAILABLE INFORMATION AND HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF THE REPORTED HIGH POWER EVENT CAN BE ATTRIBUTED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF THROMBUS EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES REL ATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173052&pc=BZD | 2518422-2025-003866 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Nausea; Pulmonary Dysfunction; Vomiting; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: IN THE PREVIOUS REPORT, BOX B: OUTCOME ATTRIBUTED TO AE WAS MISTAKENLY RECORDED AS SERIOUS INJURY, AND BOX H: TYPE OF REPORTED COMPLAINT WAS ALSO MARKED AS SERIOUS INJURY. THE FOLLOWING CORRECTIONS HAVE BEEN MADE IN THIS REPORT: BOX B: OUTCOME ATTRIBUTED TO AE HAS BEEN CORRECTED. BOX H: TYPE OF REPORTED COMPLAINT HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173185&pc=MHX | 1218950-2025-000038 | 2024-12-20 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-01-16 05:00 | MHX | PATIENT INFORMATION CENTER IX | Use of Device Problem | Insufficient Information | K183387 |  | 1 | Event Description: IT WAS REPORTED AFTER THE PATIENT DEATH, THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHEN THE DEVICE WAS PLACED IN STANDBY. THE PATIENT PASSED AWAY WHILE CONNECTED TO THE TELEMETRY DEVICE AFTER THE DEVICE WAS PLACED IN STANDBY. THE CUSTOMER INDICATED THE INCIDENT WAS NOT RELATED TO AN EQUIPMENT MALFUNCTION BUT TO AN INTERNAL WORK PROCESS ISSUE. IN ADDITION, CUSTOMER INDICATED IT WAS REVIEWING ITS POLICY ON THE USE OF STANDBY WITH THE MONITORS. Manufacturer Narrative: A PHILIPS REMOTE APPLICATION SPECIALIST (RAS) INTERVIEWED THE CUSTOMER AND SHOWED THEM HOW TO SEARCH FOR SEVEN DAYS FOR MONITORS IN AND OUT OF STANDBY. THE CUSTOMER DID NOT KNOW THE DATE AND TIME OF THE EVENT, SO LOGGING INTO PRS TO PULL LOG DATA WAS NOT WORTHWHILE. THE CUSTOMER REPORTED THIS WAS NOT A MONITOR ISSUE. THE STAFF HAD PUT THE MONITOR INTO STANDBY FOR TOO LONG. THE PATIENT HAD AN EVENT WHILE THE MONITOR WAS IN STANDBY. THE CUSTOMER HAS TAKEN ACTION TO LEARN FROM THIS SITUATION BY INITIATING A PROCESS REVIEW FOR USE OF STANDBY FUNCTIONS. THE CUSTOMER EXPLAINED THIS INCIDENT WAS NOT RELATED TO AN EQUIPMENT MALFUNCTION BUT CONNECTED TO PROCEDURES FOR STANDBY MODE USAGE. THE CUSTOMER CONFIRMED THIS WAS AN INTERNAL WORK PROCEDURE ISSUE. THE DEVICE WAS WORKING TO SPECIFICATION. NO FURTHER INVESTIGATION OR ACTION IS WARRANTED AT THIS TIME. BASED ON THE UPDATED INFORMATION, THIS RECORD HAS IS DEEMED NOT A REPORTABLE EVENT WITH PHILIPS DUE TO WHEN THE DEVICE ENTERS STANDBY SPONTANEOUSLY, MONITORING IS STOPPED HOWEVER THERE ARE CLEAR INDICATORS THAT ACTIVE MONITORING IS NOT OCCURRING AND THE MESSAGE "STANDBY" IS PROVIDED AT THE BEDSIDE/CENTRAL WITH NO OTHER MONITORING PARAMETERS. THE USER WOULD DETECT THIS DURING CLOSE PERSONAL SURVEILLANCE AND ACT TO RESUME MONITORING. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173200&pc=LWQ | 2135147-2025-00330 | 2024-12-23 05:00 | Death | ABBOTT MEDICAL | 2025-01-16 05:00 | LWQ | MASTERS SERIES HEMODYNAMIC PLUS VALVE | Fracture | Insufficient Information | P810002 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A 19MM SJM MASTERS SERIES HEMODYNAMIC PLUS VALVE WAS CHOSEN FOR A PROCEDURE. DURING PROCEDURE, WHILE THE VALVE WAS POSITIONED WITH ITS OWN HOLDER, BOTH LEAFLETS WERE SEPARATED AND FELL INTO THE LEFT VENTRICLE. ONE OF THE LEAFLETS WAS REMOVED, BUT WHEN THE OTHER LEAFLET WAS NOT SEEN, THE VALVE THAT WAS PLACED AND THE LEAFLET WAS REMOVED WAS REMOVED AGAIN. WHEN THE OTHER LEAFLET COULD NOT BE FOUND IN THE LEFT VENTRICLE FROM THE AORTIC ANGLE, THE ARTIFICIAL MITRAL VALVE THAT HAD BEEN PLACED WAS ALSO REMOVED AGAIN. DESPITE ALL SEARCHES, A LEAFLET COULD NOT BE FOUND. WHEN A LEAFLET COULD NOT BE FOUND WITH BOTH STERILE CAMERA AND SCOPY, A NEW 19MM SJM MASTERS SERIES HEMODYNAMIC PLUS VALVE WAS IMPLANTED. AFTER THE AIR IN THE CARDIAC CHAMBERS WAS REMOVED, THE AORTIC CROSS-CLAMP WAS REMOVED AND THE HEART WAS STARTED. THE PATIENT HEMODYNAMICS WERE STABILIZED WITH INOTROPIC SUPPORT, WAS TAKEN OFF THE PUMP, DECANNULATED, AND AFTER OTHER PROCEDURES, THE PATIENT WAS TAKEN TO THE INTENSIVE CARE UNIT. AFTER THE PROCEDURE, THE PATIENT HAD MULTIORGAN FAILURE DEVELOPING AFTER PROLONGED CARDIOPULMONARY BYPASS PERIOD. ON (B)(6) 2024, THE PATIENT WAS REPORTED PASSED AWAY. THE CAUSE OF DEATH WAS MULTIORGAN FAILURE THAT DEVELOPED AFTER A VERY LONG AORTIC CROSS CLAMP PERIOD AND CARDIOPULMONARY BYPASS PERIOD DUE TO INTRACARDIAC SEARCH OF THE REMOVED LEAFLET.Event Description: N/A. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: AN EVENT OF A LEAFLET FRACTURE DURING IMPLANT AND PATIENT DEATH WAS REPORTED. THE DEVICE WAS RETURNED TO ABBOTT FOR INVESTIGATION AND FOUND THE DEVICE WAS RETURNED WITH A TORN CUFF, ONE FRACTURED LEAFLET AND CHIPPED /FRACTURE DAMAGE ON THE BOTTOM RIM AND BUTTERFLY WALL AREA. IT WAS INDICATED THAT LEAFLET COULD NOT BE FOUND IN THE LEFT VENTRICLE FROM THE AORTIC ANGLE, THE ARTIFICIAL MITRAL VALVE THAT HAD BEEN PLACED WAS ALSO REMOVED AGAIN. REGARDLESS, THE OTHER LEAFLET WAS NOT FOUND. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. THE CAUSE OF THE REPORTED EVENT OF LEAFLET FRACTUREAND TORN CUFF IS CONSISTENT WITH ROTATING THE VALVE THROUGH RESISTANCE BUT COULD NOT BE CONFIRMED. THE CAUSE OF DEATH WAS DUE TO PROLONGED OPERATIVE CONDITIONS ON INTRACARDIAC SEARCH OF THE REMOVED LEAFLET. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FROM MEDICAL REVIEW: "THE PATIENT DEATH IS PROCEDURE RELATED DUE TO A PROLONGED PROCEDURE CAUSING A PROLONGED CARDIOPULMONARY BYPASS AND LEADING TO MULTI-SYSTEM ORGAN FAILURE POST-OPERATIVELY. IT IS UNCLEAR WHY THE TWO LEAFLETS DISLODGED FROM THE VALVE AT THE TIME OF POSITIONING, BUT MAY BE A RESULT OF APPLYING EXCESSIVE FORCE DURING SEATING CAUSING THE CARBON ORIFICE TO BECOME OVALIZED WITH BOTH LEAFLETS FALLING OUT."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173320&pc=NVN | 2124215-2025-01880 | 2024-12-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NVN | INGEVITY+ | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS LEFT BUNDLE BRANCH (LBB) WAS PART OF A SYSTEM REVISION DUE TO POSSIBLE INFECTION. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS LBB LEAD WAS TURNED OFF BUT STILL IMPLANTED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD SYSTEMIC INFECTION. FURTHERMORE, THIS LEAD WAS NOT EXPLANTED BUT WAS MADE OOS DUE TO THE PATIENT PASSING AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS LEFT BUNDLE BRANCH (LBB) LEAD HAD EXHIBITED POSSIBLE INFECTION. THE PHYSICIAN DECIDED TO SWITCH OFF THE CRT-D AS WELL AS THE LBB AND RIGHT VENTRICULAR (RV) LEAD. SUBSEQUENTLY, A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD SYSTEMIC INFECTION. FURTHERMORE, THIS LEAD WAS NOT EXPLANTED BUT WAS MADE OOS DUE TO THE PATIENT PASSING AWAY. Manufacturer Narrative: (B)(6)2025 - AFTER REVIEW, IT WAS DEEMED THAT IT COULD NOT BE CONCLUSIVELY RULED OUT THAT THE DEATH WAS NOT RELATED TO THE PRODUCT OR PROCEDURE. (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173321&pc=NVY | 2124215-2025-01874 | 2024-12-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NVY | RELIANCE 4-FRONT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P910073 |  | 1 | Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD WAS PART OF A SYSTEM REVISION DUE TO POSSIBLE INFECTION. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS RV LEAD WAS TURNED OFF BUT STILL IMPLANTED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD SYSTEMIC INFECTION. FURTHERMORE, THIS LEAD WAS NOT EXPLANTED BUT WAS MADE OOS DUE TO THE PATIENT PASSING AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS RIGHT VENTRICULAR (RV) LEAD HAD EXHIBITED POSSIBLE INFECTION. THE PHYSICIAN DECIDED TO SWITCH OFF THE CRT-D AS WELL AS THE LEFT BUNDLE BRANCH (LBB) AND RV LEAD. SUBSEQUENTLY, A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD SYSTEMIC INFECTION. FURTHERMORE, THIS LEAD WAS NOT EXPLANTED BUT WAS MADE OOS DUE TO THE PATIENT PASSING AWAY. Manufacturer Narrative: AFTER REVIEW, IT WAS DEEMED THAT IT COULD NOT BE CONCLUSIVELY RULED OUT THAT THE DEATH WAS NOT RELATED TO THE PRODUCT OR PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173334&pc=NGV | 2124215-2024-83827 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-16 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 5 | Event Description: THIS REPORT SUMMARIZES 5 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 69 TO 87 YEARS. PATIENT SEX: 3 FEMALES; 2 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 453 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173338&pc=DTB | 3008973940-2025-00473 | 2024-12-31 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-16 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT SIX DAYS POST IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM IMPLANT, THE PATIENT EXPIRED, AND A CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173346&pc=NVN | 2649622-2025-01869 | 2024-12-31 05:00 | Death | MPRI | 2025-01-16 05:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT SIX DAYS POST IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM IMPLANT, THE PATIENT EXPIRED, AND A CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173349&pc=NVZ | 9614453-2025-00279 | 2024-12-31 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-01-16 05:00 | NVZ | ASTRA S DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT SIX DAYS POST IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM IMPLANT, THE PATIENT EXPIRED, AND A CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173500&pc=MIH | 2017233-2025-05743 | 2024-12-18 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-16 05:00 | MIH | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Insufficient Information | Insufficient Information | P230023 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2024, A PATIENT WAS TREATED USING GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. AFTER SUCCESSFUL DEPLOYMENT, THE RENAL PORTALS WERE NOT FILLING, BUT THE PHYSICIAN REPORTEDLY EXPECTED THIS BASED ON PATIENT CONDITION. ON (B)(6) 2024, THE PATIENT DIED OF SPLENIC/HEPATIC INFARCT. REPORTEDLY, THE STENTS WERE ALL OPEN. NO INFORMATION HAS BEEN MADE AVAILABLE REGARDING THE CAUSE OF THE SPLENIC/HEPATIC INFARCT.Event Description: ON (B)(6) 2024, A PATIENT WAS TREATED USING GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. AFTER SUCCESSFUL DEPLOYMENT, THE RENAL PORTALS WERE NOT FILLING, BUT THE PHYSICIAN REPORTEDLY EXPECTED THIS BASED ON PATIENT CONDITION. ON (B)(6) 2024, THE PATIENT DIED. IT WAS ORIGINALLY REPORTED BY A FELLOW THAT THE PATIENT DIED OF SPLENIC/HEPATIC INFARCT. HOWEVER, AFTER CONTACTING THE PHYSICIAN, THIS WAS DETERMINED TO BE UNTRUE. THE PATIENT ASPIRATED WHEN THEY WERE EXTUBATED, AND THEN BECAME SEPTIC. ALL OF THE STENTS WERE OPEN, AND THERE WERE NO ALLEGATIONS OF DEFICIENCY AGAINST THE GORE DEVICE. Manufacturer Narrative: B3/B4/B5/B6: DESCRIPTION OF EVENT UPDATED. H.6. INVESTIGATION CONCLUSIONS CODE D14: AFTER COMMUNICATION WITH THE PHYSICIAN, IT WAS LEARNED THAT THE INITIALLY REPORTED CAUSE OF DEATH WAS INCORRECT. THE PATIENT DIED WHILE BEING EXTUBATED, AND THERE IS NO ALLEGATION OF DEFICIENCY AGAINST THE GORE DEVICE. THEREFORE, THIS MEDWATCH IS BEING RETRACTED. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173508&pc=DSQ | 2916596-2025-00112 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-16 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: THE CAUSE OF THE ZERO FLOW WAS FOUND TO BE DUE TO THROMBUS WITHIN THE LVAD OUTFLOW GRAFT. THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO TRAUMA WITH 0 FLOW IN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE PATIENT'S SON PERFORMED A CONTROLLER EXCHANGE WHEN THEY HEARD THE FIRST ALARM. THE ORIGINAL PRIMARY CONTROLLER'S LOG FILE APPEARED NORMAL UNTIL 01JAN2025 AT 7:34:53 WHEN FLOW BEGAN TO DECREASE. AT 7:35:03 A LOW FLOW ALARM WAS CAPTURED AND AT 7:48:20 FLOW DROPS TO 0.0 LITERS PER MINUTE (LPM). THE FLOW REMAINS AT 0.0 LPM FOR THE REMAINDER OF THE LOG FILE. THE PUMP WAS DISCONNECTED ON 01JAN2025 @ 7:58:37. PRIOR TO THE PUMP BEING DISCONNECTED THE PUMP WAS OPERATING AS INTENDED AND MAINTAINING THE SET SPEED OF 5600 RPM. THE PATIENT'S NEW PRIMARY CONTROLLER LOG ONLY CAPTURED 1 LINE OF DATA AND IT SHOWED FLOW REMAINED AT 0.0 LPM WITH THE PUMP OPERATING AT THE SET SPEED OF 5600 RPM. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS AND A DROP IN FLOW TO 0.0 LITERS PER MINUTE (LPM) WHICH THE ACCOUNT ATTRIBUTED TO THROMBUS WITHIN THE OUTFLOW GRAFT. HOWEVER, THE REPORTED OUTFLOW GRAFT THROMBUS COULD NOT BE CONFIRMED AS NO IMAGES WERE SUBMITTED FOR REVIEW AND THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6)2024 ¿ (B)(6)2025, PER THE TIMESTAMPS. A CONTINUOUS LOW FLOW HAZARD ALARM WAS CAPTURED BEGINNING AT 07:35:03 ON (B)(6)2025, FOLLOWED BY A DROP IN FLOW TO 0.0 LPM FROM 07:48:20 THROUGH THE END OF THE FILE AT 09:58:54. NO EXTERNAL POWER ALARMS BECAME ACTIVE FROM 07:57:38 DUE TO A DOUBLE POWER CABLE DISCONNECT, FOLLOWED BY THE DISCONNECTION OF THE DRIVELINE AT 07:58:37 RESULTING IN A PUMP STOP. THESE EVENTS APPEAR CONSISTENT WITH THE REPORTED CONTROLLER EXCHANGE PERFORMED BY THE PATIENT'S SON AFTER HEARING THE FIRST ALARM. ADDITIONAL LOG FILES WERE SUBMITTED FOLLOWING THE CONTROLLER EXCHANGE, AND EVENTS FOLLOWING THE CONNECTION OF THE DRIVELINE ON (B)(6)2000 AT 00:01:18 THROUGH 01:52:05 (PER THE TIMESTAMPS) WERE REVIEWED. ONCE THE DRIVELINE WAS CONNECTED, THE PUMP APPEARED TO RAMP UP TO THE STORED PATIENT SPEED WITHOUT ISSUE. HOWEVER, ESTIMATED FLOW REMAINED AT ZERO FOR THE DURATION OF THE FILE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED IN THE LOG FILES, AND THE SYSTEM APPEARED TO BE OPERATING AS INTENDED AT THE STORED PATIENT SPEED WHILE THE DRIVELINE WAS CONNECTED. THE PATIENT ULTIMATELY EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE PATIENT¿S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW. THIS SECTION EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. ADDITIONALLY, THIS SECTION, UNDER ¿ANTICOAGULATION¿ PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B1 ¿ TYPE OF REPORT UPDATED. B2: DEATH DATE NOT PROVIDED. B5: ADDITIONAL INFORMATION. H1 - TYPE OF REPORTABLE EVENT UPDATED. H6: HEALTH EFFECT CLINICAL, HEALTH EFFECT IMPACT, DEVICE PROBLEM CODES, UPDATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173531&pc=DSQ | 2916596-2025-00159 | 2024-12-13 05:00 | Death | THORATEC CORPORATION | 2025-01-16 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED.Event Description: AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173545&pc=PNJ | 9612164-2025-00310 | 2023-06-01 04:00 | Death | MEDTRONIC IRELAND | 2025-01-16 05:00 | PNJ | MICRA AV | Material Integrity Problem | Insufficient Information | P150033 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THIS LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). THE AUTHORS DESCRIBED ONE PATIENT WHO UNDERWENT AN IMPLANT OF A LEADLESS IPG AND WAS PLACED ADJACENT TO THEIR PREVIOUS LEADLESS DEVICE. APPROXIMATELY THIRTEEN DAYS POST IMPLANT, THE PATIENT'S CONDITION DETERIORATED AND THEY SUBSEQUENTLY DIED. THE SPECIFIC CAUSE OF DEATH WAS UNKNOWN. AN AUTOPSY REVEALED THAT THE NEWLY IMPLANTED DEVICE ONLY HAD TWO TINES HOOKED IN THE TRABECULAE AND THE BODY OF THE DEVICE APPEARED TO BE STUCK BETWEEN THE TRABECULAE. IT WAS NOTED ON FLUOROSCOPY THAT ONE OF THE TINES WAS FLOATING IN THE CAVITY. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. THE EVENT ONLY OCCURRED WITH ONE PATIENT BUT SPECIFIC DETAILS ON THE PATIENT WERE NOT PROVIDED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: AUTOPSY OF TWO MICRA LEADLESS PACEMAKERS IMPLANTED IN PARALLEL. ANNUAL SCIENTIFIC MEETING OF THE JAPANESE CIRCULATION SOCIETY. 2024. (JCS2024) VOL.88TH, PAGE.2129. CRDJ15-1. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173908&pc=KDJ | 1416980-2025-00089 | 2024-12-19 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-16 05:00 | KDJ | NI | Improper or Incorrect Procedure or Method | Abdominal Pain; Low Blood Pressure/ Hypotension; Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. THE BREACH IN ASEPTIC TECHNIQUE WAS NOT FURTHER DESCRIBED. THE PERITONITIS WAS MANIFESTED BY ABDOMINAL PAIN AND CLOUDY PD EFFLUENT. THE SAME DAY AS THE EVENT ONSET, THE PATIENT WAS HOSPITALIZED AND TREATED WITH VANCOMYCIN INJECTION (1GM, INTRAPERITONEAL, 72 HOURLY, ONGOING), CEFTAZIDIME INJECTION (1G, INTRAPERITONEAL, ONCE DAILY, ONGOING), AND AMIKACIN INJECTION (125MG, INTRAPERITONEAL, ONCE DAILY, ONGOING). IT WAS REPORTED THAT RETRAINING ON PROPER ASEPTIC TECHNIQUE WAS PERFORMED. PD THERAPY WAS ONGOING. AT THE TIME OF THIS REPORT, THE PATIENT WAS RECOVERING FROM PERITONITIS AND REMAINED HOSPITALIZED. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: B5: UPON FOLLOW-UP, IT WAS REPORTED THAT THE PATIENT EXPERIENCED LOW BLOOD (BP) PRESSURE AND SUBSEQUENTLY THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED AS DUE TO LOW BP. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. PD THERAPY WAS DISCONTINUED, THE PD CATHETER WAS REMOVED AND THE PATIENT WAS TRANSFERRED TO HEMODIALYSIS PRIOR TO PATIENT DEATH. IT WAS REPORTED THE PATIENT WAS NOT RECOVERED FROM THE PERITONITIS PRIOR TO DEATH. HEMODIALYSIS WAS ONGOING AT THE TIME OF DEATH. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS REPORT IS FOR A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. PER BAXTER LABELING, USERS ARE INSTRUCTED TO USE ASEPTIC TECHNIQUE WHEN PERFORMING PERITONEAL DIALYSIS THERAPY. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21174840&pc=NIK | 9614453-2025-00282 | 2025-01-04 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-01-16 05:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Tachycardia; Ventricular Fibrillation; Heart Block | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT NINE DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM IMPLANT THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT)/ VENTRICULAR FIBRILLATION (VF) WITH APPROPRIATE SENSING AND DETECTION. HIGH VOLTAGE THERAPIES WERE DELIVERED APPROPRIATELY AND THE ARRHYTHMIA WAS TERMINATED, HOWEVER, THE PATIENT WENT BACK INTO THE ARRHYTHMIA AND SUBSEQUENT HIGH VOLTAGE THERAPY WAS UNSUCCESSFUL AND THERAPIES WERE EXHAUSTED. VT/VF CONTINUED, THE PATIENT HAD SUBSEQUENT PULSELESS ELECTRICAL ACTIVITY NOTED AND THE PATIENT WAS DECEASED. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21174844&pc=DTB | 2649622-2025-01877 | 2025-01-04 05:00 | Death | MPRI | 2025-01-16 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Tachycardia; Ventricular Fibrillation; Heart Block | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT NINE DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM IMPLANT THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT)/ VENTRICULAR FIBRILLATION (VF) WITH APPROPRIATE SENSING AND DETECTION. HIGH VOLTAGE THERAPIES WERE DELIVERED APPROPRIATELY AND THE ARRHYTHMIA WAS TERMINATED, HOWEVER, THE PATIENT WENT BACK INTO THE ARRHYTHMIA AND SUBSEQUENT HIGH VOLTAGE THERAPY WAS UNSUCCESSFUL AND THERAPIES WERE EXHAUSTED. VT/VF CONTINUED, THE PATIENT HAD SUBSEQUENT PULSELESS ELECTRICAL ACTIVITY NOTED AND THE PATIENT WAS DECEASED. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21174850&pc=OJX | 2649622-2025-01878 | 2025-01-04 05:00 | Death | MPRI | 2025-01-16 05:00 | OJX | ATTAIN PERFORMA STRAIGHT MRI SURESCAN | Appropriate Term/Code Not Available | Tachycardia; Ventricular Fibrillation; Heart Block | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT NINE DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM IMPLANT THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT)/ VENTRICULAR FIBRILLATION (VF) WITH APPROPRIATE SENSING AND DETECTION. HIGH VOLTAGE THERAPIES WERE DELIVERED APPROPRIATELY AND THE ARRHYTHMIA WAS TERMINATED, HOWEVER, THE PATIENT WENT BACK INTO THE ARRHYTHMIA AND SUBSEQUENT HIGH VOLTAGE THERAPY WAS UNSUCCESSFUL AND THERAPIES WERE EXHAUSTED. VT/VF CONTINUED, THE PATIENT HAD SUBSEQUENT PULSELESS ELECTRICAL ACTIVITY NOTED AND THE PATIENT WAS DECEASED. THE PATIENT WAS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21174913&pc=FKX | 1416980-2025-00091 | 2024-12-23 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-16 05:00 | FKX | HOMECHOICE CLARIA | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NI |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE PATIENT WAS NOT HOSPITALIZED PRIOR TO DEATH. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. PD THERAPY WAS ONGOING AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: H11: THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED; HOWEVER, THE SHARESOURCE LOG FILES ASSOCIATED WITH HOMECHOICE CLARIA WERE REVIEWED. THE REVIEW SHOWED THE MOST RECENT TREATMENTS FROM (B)(6) 2024. THERE WAS NO INDICATION OF AN IIPV EVENT OR EVIDENCE OF A DEVICE MALFUNCTION THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THERE WAS NO EVIDENCE OF A HARDWARE DEFECT OR DEVICE MALFUNCTION THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED PATIENT DEATH. THE CAUSE OF THE CONDITION COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21175235&pc=MIH | 3007284313-2025-03717 | 2024-12-19 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-16 05:00 | MIH | GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P20004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2024, THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT FOR AN ABDOMINAL AORTIC ANEURYSM USING GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS AND GORE® EXCLUDER® AAA ENDOPROSTHESES. AFTER RETURNING TO THE HOSPITAL ROOM (OR INTENSIVE CARE UNIT), THE PATIENT COMPLAINED OF PAIN. A CT SCAN REVEALED RTAD (RETROGRADE STANFORD TYPE A AORTIC DISSECTION). THE PATIENT EXPIRED BEFORE ANY INTERVENTION WAS PERFORMED. PHYSICIAN¿S COMMENT: ¿IT IS IMPOSSIBLE TO DETERMINE THE EXACT CAUSE- WHETHER IT WAS DUE TO THE USE OF GUIDEWIRE, PUSHING THE DELIVERY CATHETER TOO HARD, EXCESSIVE BALLOON EXPANSION, OR A COMBINATION OF DIFFERENT FACTORS.¿ Manufacturer Narrative: H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICE REMAINS IMPLANTED. RETURN NOT POSSIBLE. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: DEATH, DISSECTION OF THE AORTIC VESSEL AND SURROUNDING VASCULATURE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21175366&pc=BZD | 2518422-2025-003896 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-16 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178459&pc=QBJ | MW5164857 |  | Death | DEXCOM, INC. | 2025-01-16 05:00 | QBJ | DEXCOM G7 RECEIVER DEVICE |  | Insufficient Information |  |  | 1 | Event Description: PATIENT PASSED AWAY AND WAS SUPPOSED TO COME TO OFFICE IN DEC BUT CANCELLED THEN DIDN'T RESCHEDULE. NO ADDITIONAL INFORMATION REPORTED OR AVAILABLE REGARDING THIS EVENT. PAE REPORTED BY HCP, WHO DOES NOT CONSENT TO FOLLOW UP. CAUSE AND DATE OF DEATH NOT PROVIDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21155360&pc=BZD | 2518422-2025-002486 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21155422&pc=BZD | 2518422-2025-002488 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21155425&pc=BZD | 2518422-2025-002498 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21155588&pc=BZD | 2518422-2025-002521 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21155692&pc=BZD | 2518422-2025-002525 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21155871&pc=FMF | 9681821-2025-00009 | 2024-12-01 05:00 | Death | NOVO NORDISK A/S | 2025-01-15 05:00 | FMF | NOVOPEN 6 | Break | Diabetic Ketoacidosis | K123766 |  | 1 | Event Description: EVENT VERBATIM [PREFERRED TERM] (RELATED SYMPTOMS IF ANY SEPARATED BY COMMAS). KETOACIDOSIS (PATIENT HAS DIABETES) [DIABETIC KETOACIDOSIS]. BROKEN NOVOPEN 6 [DEVICE BREAKAGE]. CASE DESCRIPTION: THIS SERIOUS SPONTANEOUS CASE FROM FRANCE WAS REPORTED BY A CONSUMER AS "KETOACIDOSIS (PATIENT HAS DIABETES) (DIABETIC KETOACIDOSIS)" WITH AN UNSPECIFIED ONSET DATE, "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" WITH AN UNSPECIFIED ONSET DATE, AND CONCERNED A 42-YEAR-OLD MALE PATIENT WHO WAS TREATED WITH NOVOPEN 6 (INSULIN DELIVERY DEVICE) FROM UNKNOWN START DATE FOR "DEVICE THERAPY", THE PATIENT'S HEIGHT, WEIGHT AND BODY MASS INDEX WERE NOT REPORTED. CURRENT CONDITION: DIABETES (TYPE AND DURATION NOT REPORTED). ON AN UNKNOWN DATE THREE WEEKS AGO (FROM THE TIME OF REPORTING, (B)(6) 2024) THE PATIENT WITH DIABETES DIED OF KETOACIDOSIS. HE WAS USING THE NOVOPEN 6 CONNECTED PEN. THIS PEN, WHICH WAS GIVEN TO PATIENT'S MOTHER ALONG WITH BELONGINGS, WAS BROKEN. THIS WAS JUST AN OBSERVATION IT WAS REPORTED TO TAKE A CLOSER LOOK AT ITS DESIGN, TO PREVENT OTHERS FROM "GOING AWAY", "BECAUSE OF HIS DIABETES". EVENT AUTOPSY DETAILS: UNKNOWN. BATCH NUMBERS: NOVOPEN 6: NOT AVAILABLE. ACTION TAKEN TO NOVOPEN 6 WAS REPORTED AS NOT APPLICABLE. THE OUTCOME FOR THE EVENT "KETOACIDOSIS (PATIENT HAS DIABETES) (DIABETIC KETOACIDOSIS)" WAS FATAL. THE OUTCOME FOR THE EVENT "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" WAS UNKNOWN. EVENT ONSET DATE IS NOT REPORTED IN THE CASE. HOWEVER, THE INCIDENT DATES ARE CAPTURED TO ENSURE THE MIR FORM IS GENERATED. NO FURTHER INFORMATION AVAILABLE. PRELIMINARY MANUFACTURER'S COMMENT: 06-DEC-2025: THE SUSPECTED DEVICE NOVOPEN 6 HAS NOT BEEN RETURNED TO NOVO NORDISK FOR EVALUATION. NO CONCLUSION REACHED. RELEVANT INFORMATION ON MEDICATION USED WITH NOVOPEN 6, PRECIPITATING FACTORS FOR DIABETIC KETOACIDOSIS, TREATMENT GIVEN IN HOSPITAL AND INVESTIGATION RESULTS OF DEFECTIVE DEVICE ARE UNAVAILABLE FOR COMPLETE CAUSALITY AS-SESSMENT. PATIENT'S UNDERLYING MEDICAL CONDITION OF DIABETES MELLITUS IS A SIGNIFICANT CONFOUNDING FACTOR FOR THE DEVELOPMENT OF ACUTE COMPLICATION OF DIABETIC KETOACIDOSIS LEAD-ING TO FATAL OUTCOME. "THIS REPORT IS FOR A FOREIGN DEVICE THAT IS ASSESSED AS "SIMILAR" TO US MARKETED NOVOPEN ECHO".Event Description: CASE DESCRIPTION: THIS SERIOUS SPONTANEOUS CASE FROM FRANCE WAS REPORTED BY A CONSUMER AS "KETOACIDOSIS (PATIENT HAS DIABETES) (DIABETIC KETOACIDOSIS)" BEGINNING ON (B)(6) 2024, "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" BEGINNING ON (B)(6) 2024, AND CONCERNED A 42-YEAR-OLD MALE PATIENT WHO WAS TREATED WITH NOVOPEN 6 (INSULIN DELIVERY DEVICE) FROM UNKNOWN START DATE FOR "DEVICE THERAPY", BATCH NUMBERS: NOVOPEN 6: NVGAB17-2. SINCE LAST SUBMISSION THE CASE HAS BEEN UPDATED WITH THE FOLLOWING: -BATCH NUMBER WAS UPDATED. -SAMPLE RECEIVED DATE UPDATED. -DEVICE NARRATIVE UPDATED IN DEVICE ADDENDUM TAB. -NARRATIVE UPDATED ACCORDINGLY. REFERENCES INCLUDED: REFERENCE TYPE: E2B COMPANY NUMBER. REFERENCE ID#: (B)(4). REFERENCE NOTES: REFERENCE TYPE: MW 3500A MFR. RPT. #. REFERENCE ID#: (B)(4). REFERENCE NOTES: MEDWATCH 3500A MFR. REPORT NUMBER. PRELIMINARY MANUFACTURER'S COMMENT: (B)(6) 2025: THE SUSPECTED DEVICE NOVOPEN 6 HAS BEEN RETURNED TO NOVO NORDISK FOR EVALUATION. INVESTIGATIONS ARE ONGOING. RELEVANT INFORMATION ON MEDICATION USED WITH NOVOPEN 6, PRECIPITATING FACTORS FOR DIABETIC KETOACIDOSIS, TREATMENT GIVEN IN HOSPITAL AND INVESTIGATION RESULTS OF DEFECTIVE DEVICE ARE UNAVAILABLE FOR COMPLETE CAUSALITY ASSESSMENT. PATIENT'S UNDERLYING MEDICAL CONDITION OF DIABETES MELLITUS IS A SIGNIFICANT CONFOUNDING FACTOR FOR THE DEVELOPMENT OF ACUTE COMPLICATION OF DIABETIC KETOACIDOSIS LEADING TO FATAL OUTCOME. H3 CONTINUED: NO (ATTACH PAGE TO EXPLAIN WHY NOT) OR PROVIDE CODE. 02 DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN.Event Description: CASE DESCRIPTION: THIS SERIOUS SPONTANEOUS CASE FROM FRANCE WAS REPORTED BY A CONSUMER AS "KETOACIDOSIS (PATIENT HAS DIABETES)(DIABETIC KETOACIDOSIS)" BEGINNING ON (B)(6)2024, "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" BEGINNING ON (B)(6)2024, AND CONCERNED A 42 YEARS OLD MALE PATIENT WHO WAS TREATED WITH NOVOPEN 6 (INSULIN DELIVERY DEVICE) FROM UNKNOWN START DATE FOR "DEVICE THERAPY", NOVORAPID PENFILL (INSULIN ASPART 100 U/ML) (DOSE, FREQUENCY & ROUTE USED- UNK) FROM UNKNOWN START DATE FOR "PRODUCT USED FOR UNKNOWN INDICATION". BATCH NUMBERS: NOVOPEN 6: UNKNOWN. NOVORAPID PENFILL: NR7KV74. ACTION TAKEN TO NOVORAPID PENFILL WAS NOT REPORTED. ON (B)(6)2024 THE OUTCOME FOR THE EVENT "KETOACIDOSIS (PATIENT HAS DIABETES) (DIABETIC KETOACIDOSIS)" WAS FATAL. ON (B)(6)2024 THE OUTCOME FOR THE EVENT "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" WAS FATAL. EVENT ONSET DATE IS NOT REPORTED IN THE CASE. HOWEVER, THE INCIDENT DATES ARE CAPTURED TO ENSURE THE MIR FORM IS GENERATED. SINCE LAST SUBMISSION THE CASE HAS BEEN UPDATED WITH THE FOLLOWING: NEW SUSPECT PRODUCT ADDED(NOVORAPID). EXPECTED DATE OF FOLLOW UP REPORT EXTENDED . NARRATIVE UPDATED ACCORDINGLY. PRELIMINARY MANUFACTURER'S COMMENT: (B)(6)2025: THE SUSPECTED DEVICE NOVOPEN 6 HAS BEEN RETURNED TO NOVO NORDISK FOR EVALUATION. INVESTIGATIONS ARE ONGOING. RELEVANT INFORMATION ON MEDICATION USED WITH NOVOPEN 6, PRECIPITATING FACTORS FOR DIABETIC KETOACIDOSIS, TREATMENT GIVEN IN HOSPITAL AND INVESTIGATION RESULTS OF DEFECTIVE DEVICE ARE UNAVAILABLE FOR COMPLETE CAUSALITY ASSESSMENT. PATIENT'S UNDERLYING MEDICAL CONDITION OF DIABETES MELLITUS IS A SIGNIFICANT CONFOUNDING FACTOR FOR THE DEVELOPMENT OF ACUTE COMPLICATION OF DIABETIC KETOACIDOSIS LEADING TO FATAL OUTCOME. "THIS REPORT IS FOR A FOREIGN DEVICE THAT IS ASSESSED AS "SIMILAR" TO US MARKETED NOVOPEN ECHO".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156018&pc=BZD | 2518422-2025-002629 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156100&pc=BZD | 2518422-2025-002652 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156110&pc=BZD | 2518422-2025-002592 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156167&pc=BZD | 2518422-2025-002644 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156276&pc=OZP | 2032227-2025-109865 | 2024-12-13 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-15 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS UNKNOWN. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: NO CONTRIBUTING FACTORS WERE IDENTIFIED. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1884, UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156410&pc=BZD | 2518422-2025-002738 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156420&pc=BZD | 2518422-2025-002745 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156512&pc=BZD | 2518422-2025-002752 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, KIDNEY DISEASE/TOXICITY AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156693&pc=DSQ | 2916596-2025-00075 | 2024-12-23 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS LATER REPORTED THAT THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5: ADDITIONAL INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: H6-3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156712&pc=DSQ | 2916596-2025-00081 | 2024-12-17 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. POST IMPLANTATION, THE PATIENT REQUIRED DIALYSIS AND VENOVENOUS (VV) EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT DEVELOPED A KLEBSIELLA INFECTION, WHICH WAS TREATED BY ANTIBIOTICS, BUT DID NOT IMPROVE.Event Description: THE PATIENT'S RENAL FAILURE EXISTED PRE-LVAD IMPLANT WITH A BASELINE CREATININE OF 1.7. THE LVAD DID NOT CONTRIBUTE OR WORSEN THE PATIENT'S RENAL FAILURE. THE INFECTION WAS MOSTLY HOSPITAL ACQUIRED. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM}. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: B5: ADDITIONAL INFORMATION. H6: HEALTH EFFECT -CLINICAL CODES, UPDATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156727&pc=BZD | 2518422-2025-002710 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156730&pc=BZD | 2518422-2025-002695 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156918&pc=BZD | 2518422-2025-002815 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Nausea; Vomiting; Dizziness; Cancer; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, VOMITING, LIVER DISEASE/TOXICITY AND CANCER, DEATH AND NAUSEA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21156987&pc=BZD | 2518422-2025-002537 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Pneumonia; Pulmonary Dysfunction; Unspecified Heart Problem; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), REDUCED CARDIOPULMONARY RESERVE, DEATH, HEART CONDITION, AND PNEUMONIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157061&pc=BZD | 2518422-2025-002836 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973 AND Z-1974.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157064&pc=BZD | 2518422-2025-002840 | 2024-06-14 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157400&pc=OZP | 2032227-2025-110239 | 2025-01-02 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-15 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Hypoglycemia; Coma | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS LOW BLOOD GLUCOSE, WENT INTO A DIABETIC COMA. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER'S BLOOD GLUCOSE VALUE AT THE TIME OF THE EVENT WAS UNKNOWN. THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOW WHETHER THE CUSTOMER USED THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-332A, MMT-7841ZN, UNOMEDICAL, AND MMT-7040A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157431&pc=LWS | 3008973940-2025-00410 | 2025-01-04 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-15 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED AND PASSED ON (B)(6) 2025. THERE WAS UNDERSENSED VENTRICULAR FIBRILLATION (VF) ON THE RIGHT VENTRICULAR (RV) LEAD THAT LED TO THE PATIENT¿S DEATH. IT WAS NOTED THAT THE UNDERSENSING ON THE RV LEAD RESULTED IN A DELAY TO TREAT THE VF.Event Description: IT WAS FURTHER REPORTED THAT THE VF (VENTRICULAR FIBRILLATION) UNDERSENSING WAS CAUSED BY THE DEGREE OF ABERRANCY AND POLYMORPHIC OR IGIN OF THIS ARRHYTHMIA ITSELF. ADDITIONALLY, IT WAS NOTED THAT ALTHOUGH UNDERSENSING WAS PRESENT IN THE VF EPISODE DETECTION WAS REACHED AND A SHOCK WAS DELIVERED, HOWEVER IT WAS NOT SUCCESSFUL TO REVERT THE RHYTHM. Manufacturer Narrative: CONTINUATION OF D10: DVPB3D4 ICD IMPLANTED: (B)(6)2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. ANALYSIS OF THE DEVICE MEMORY ALSO INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET, OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER, AND OVERSENSING ASSOCIATED WITH THE RIGHT VENTRICULAR LEAD. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157502&pc=BZD | 2518422-2025-002873 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER REPORTED INCORRECT INFORMATION IN BOX B: COMMON DEVICE NAME. AFTER REVIEWED, IT WAS CAPTURED CORRECTLY. IN BOX B: COMMON DEVICE NAME HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157505&pc=BZD | 2518422-2025-002857 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157531&pc=OZD | 1220648-2025-25577 | 2024-12-29 05:00 | Death | ABIOMED, INC. | 2025-01-15 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED A GASTROINTESTINAL (GI) BLEED DURING IMPELLA CP SUPPORT. HEPARIN WAS IN THE PURGE LINE LEADING UP TO THE CONDITION. THE PATIENT WAS GIVEN THREE UNITS OF PACKED RED BLOOD CELLS TO COUNTERACT THE BLEED AND RESULTING LOW HEMOGLOBIN LEVELS. THREE DAYS LATER, TWO UNITS OF PACKED RED BLOOD CELLS AND TWO UNITS OF PLATELETS WERE GIVEN TO THE PATIENT. CARE WAS WITHDRAWN THE FOLLOWING DAY UPON THE FAMILY'S REQUEST AND THE PATIENT PASSED AWAY. IT IS UNKNOWN IF THE GI BLEED CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION FOR THE GASTROINTESTINAL BLEED (GI BLEED) HAS BEEN COMPLETED. DATA LOGS ARE NOT RELEVANT TO THE ISSUE AND PRODUCT WAS NOT RETURNED FOR INVESTIGATION. CLINICAL DETAILS DO NOT REPORT ANY EXCESSIVE FORCE OR PATIENT CONDITIONS THAT WOULD EXPLAIN THE CAUSE FOR A GI BLEED ON THE 11TH DAY OF SUPPORT. SINCE CLINICAL DETAILS WERE LIMITED AND PRODUCT WASN'T RETURNED, THE ROOT CAUSE OF THE GI BLEED COULD NOT BE DETERMINED. THE PATIENT WAS ON CP SUPPORT FOR 11 DAYS WHEN THE COMPLICATION OCCURRED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157652&pc=BZD | 2518422-2025-002889 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157778&pc=BZD | 2518422-2025-002827 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, THROAT, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157788&pc=BZD | 2518422-2025-002868 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21157790&pc=BZD | 2518422-2025-002898 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158257&pc=BZD | 2518422-2025-002894 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Respiratory Failure; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158326&pc=NGV | 2124215-2024-83550 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Difficult to Open or Close | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEVICE LEAK. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 125 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158328&pc=NGV | 2124215-2024-83553 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 180 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158346&pc=NGV | 2124215-2024-83551 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 125 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158380&pc=BZD | 2518422-2025-002939 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158714&pc=OZD | 1220648-2025-25587 | 2025-01-03 05:00 | Death | ABIOMED, INC. | 2025-01-15 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED LIMB ISCHEMIA, HEMOLYSIS, AND BLEEDING. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT AFTER HAVING VENTRICULAR FIBRILLATION ARREST AT HOME. IN THE EMERGENCY DEPARTMENT, THE PATIENT CODED MULTIPLE TIMES, WAS INTUBATED, AND TRANSFERRED TO THE CATHETERIZATION LAB FOR AN ANGIOGRAM. RIGHT CORONARY ARTERY DISEASE WAS NOTED. CARDIOGENIC SHOCK WITH UNSTABLE HEMODYNAMICALLY. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE PRIOR TO PERCUTANEOUS CORONARY INTERVENTION. THE XEXT DAY FOLLOWING THE IMPLANT, PULSES IN THE RIGHT LEG BECAME WEAK AND THE RIGHT FOOT APPEARED MOTTLED. THE LEG BEGAN TO SWELL, AND A FASCIOTOMY WAS SUBSEQUENTLY COMPLETED TO RESTORE PULSE TO THE LEG. FOUR DAYS LATER, PULSE WAS ONCE AGAIN ABSENT IN THE PATIENT'S LEG. THE PATIENT HAD BEEN GIVEN TEN UNITS OF BLOOD AS A RESULT OF THE CONDITION AND INTERVENTION. COINCIDING WITH THE ISCHEMIA, THE PATIENT ALSO DEVELOPED HEMOLYSIS WITH FOUR UNITS OF BLEED BEING GIVEN TO HELP MANAGE THE HEMOLYSIS. THE SUPPORT LEVEL WAS LOWERED TO REDUCE SUCTION THAT MAY HAVE BEEN CONTRIBUTING TO THE HEMOLYSIS. AS A RESULT OF THE PRIOR INTERVENTIONS, THE PATIENT EXPERIENCED BLEEDING COMPLICATIONS FROM THE FASCIOTOMY SITES. MULTIPLE BLOOD PRODUCTS WERE GIVEN TO COUNTERACT THE BLEED AND HEPARIN LEVELS WERE REDUCED. THE PATIENT LATER PASSED AWAY AFTER SEVEN DAYS OF SUPPORT. IT IS UNKNOWN IF THE COMPLICATIONS CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: B.2 LIFE-THREATENING SHOULD HAD NOT BEEN SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25587.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158898&pc=NPT | 2025587-2025-00379 | 2024-11-11 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-15 05:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE PATIENT OUTCOMES OF CONCOMITANT TRANSCATHETER AORTIC VALVE REPLACEMENT AND PERCUTANEOUS CORONARY INTERVENTION (TAVR + PCI) COMPARED TO COMBINED SURGICAL AORTIC VALVE REPLACEMENT AND CORONARY ARTERY BYPASS GRAFTING (SAVR + CABG). THE MANUFACTURERS AND/OR BRAND NAMES OF THE DEVICES USED IN THE SAVR + CABG GROUP (N = 713) WERE NOT DISCLOSED; THEREFORE, THE REMAINDER OF THIS SUMMARY ONLY PERTAINS TO THE OUTCOMES THAT OCCURRED IN THE TAVR + PCI GROUP (N = 625). MEDTRONIC EVOLUT (N = 195) AND VARIOUS NON-MEDTRONIC (N = 430) VALVE TYPES WERE IMPLANTED IN THE PATIENT POPULATION. THE AUTHORS OBSERVED 5 PROCEDURAL DEATHS, 21 IN-HOSPITAL DEATHS WITH STROKE, AND A TOTAL OF 56 DEATHS WITH STROKE AT ONE-YEAR POST-PROCEDURE. OTHER ADVERSE OUTCOMES INCLUDED: VALVE EMBOLIZATION; CORONARY OCCLUSION; ANNULAR RUPTURE; TAMPONADE; REINTERVENTION; NEW PERMANENT PACEMAKER IMPLANTATION; NEW ONSET ATRIAL FIBRILLATION; MAJOR BLEEDING; MAJOR VASCULAR COMPLICATIONS; STROKE; TRANSIENT ISCHEMIC ATTACK; MYOCARDIAL INFARCTION; AORTIC REGURGITATION (MILD, MODERATE, OR SEVERE); HEART FAILURE; PROSTHESIS-PATIENT MISMATCH; NEED FOR REHOSPITALIZATION; NEED FOR NEW REVASCULARIZATION; AND ACUTE KIDNEY INJURY. Manufacturer Narrative: CITATION: AMAT-SANTOS IJ, GARCÍA-GÓMEZ M, AVANZAS P, ET AL. SURGICAL VS TRANSCATHETER TREATMENT IN PATIENTS WITH CORONARY ARTERY DISEASE AND SEVERE AORTIC STENOSIS. JACC CARDIOVASC INTERV. 2024;17(21):2472-2485. DOI:10.1016/J.JCIN.2024.09.003 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. MEDTRONIC TAVR PRODUCT REFERENCED IN THE ARTICLE: ¿EVOLUT¿. BASED ON THE TIMEFRAME OF THE STUDY (2018 ¿ 2022), THE EVOLUT R, EVOLUT PRO, AND EVOLUT PRO+ ITERATIONS OF THE EVOLUT PRODUCT FAMILY MAY HAVE BEEN USED. EARLIEST APPROVED PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158908&pc=OZD | 1220648-2025-25588 | 2025-01-02 05:00 | Death | ABIOMED, INC. | 2025-01-15 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD AN ACUTE BRAIN INFARCTION. THE PATIENT WAS DO-NOT-RESUSCITATE AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE ROOT CAUSE OF STROKE CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159000&pc=BZD | 2518422-2025-002943 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159062&pc=NPT | 2015691-2025-00337 | 2024-12-23 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-15 05:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Burst Container or Vessel; Difficult to Remove | Unspecified Vascular Problem; No Clinical Signs, Symptoms or Conditions | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS FIELD CLINICAL SPECIALIST, DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE BY TRANSFEMORAL APPROACH, A 26MM SAPIEN 3 ULTRA VALVE DEPLOYED SUCCESSFULLY. HOWEVER, THE PHYSICIAN NOTED PERIVALVULAR LEAK AND DECIDED TO POST DILATE WITH +1CC OF VOLUME. UPON POST DILATION, A BALLOON RUPTURE WAS NOTED. THE PHYSICIAN BELIEVED THAT THE BALLOON HAD RUPTURED DUE TO SINOTUBULAR JUNCTION CALCIUM. THE PHYSICIAN PULLED THE BALLOON BACK INTO THE SHEATH AND NOTED THAT THE SHEATH TIP HAD SPLIT DUE TO THIS ACTION. THE PHYSICIAN ATTEMPTED TO REMOVE THE DELIVERY SYSTEM AND SHEATH AS ONE UNIT. THE PHYSICIAN NOTED THAT THE ILIAC ARTERY WAS ATTACHED TO THE SHEATH AND STOPPED. THE CARDIOVASCULAR SURGEON IN ROOM BEGAN SURGERY TO REPAIR AND CONTROL BLEEDING. BLEEDING WAS ABLE TO BE CONTROLLED, BUT THE PATIENT WAS STILL IN VERY CRITICAL CONDITION DUE TO ISSUES WITH COAGULATION. THE PATIENT PASSED AWAY FROM THE PREVIOUSLY NOTED COMPLICATIONS. Manufacturer Narrative: UPDATED SECTION H6- TYPE OF INVESTIGATION, FINDINGS, AND CONCLUSIONS. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT NO: 2015691-2025-00337. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT WAS DISCARDED. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE CUSTOMER PROVIDED IMAGERY AND THE FOLLOWING WAS OBSERVED: CALCIFICATION IS PRESENT IN THE PATIENT'S LANDING ZONE. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, BALLOON BURST EVENTS DURING VALVE DEPLOYMENT HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO CALCIFICATION IN THE LANDING ZONE OR OVER-INFLATION OF THE BALLOON. THE COMPLAINT FOR BALLOON BURST WAS UNABLE TO BE CONFIRMED AS NEITHER THE DEVICE NOR A RELEVANT PROCEDURAL IMAGERY WAS RETURNED. AVAILABLE INFORMATION SUGGESTS CALCIFICATION LIKELY CONTRIBUTED TO THE EVENT AS THE PROVIDED 3MENSIO SHOWS PRESENCE OF CALCIFICATION IN THE LANDING ZONE. THE PRESENCE OF CALCIFICATION CAN CREATE A CHALLENGING ANATOMY FOR BALLOON INFLATION. WHILE THE BALLOONS ARE SUFFICIENTLY DESIGNED AND TESTED TO ENSURE THE BURST PRESSURE IS AT OR ABOVE THE RATED BURST PRESSURE, CALCIFIED NODULES CAN COMPROMISE THE STRUCTURE OF THE BALLOON WALL VIA FOLLOWING MECHANISMS SUCH AS PUNCTURE, LOCAL OVERSTRETCHING, OPEN CELL IMPINGEMENT, OR STRESS CONCENTRATION. THE COMPLAINT FOR INABILITY TO WITHDRAW SYSTEM THROUGH SHEATH WAS UNABLE TO BE CONFIRMED AS NEITHER THE DEVICE NOR A RELEVANT PROCEDURAL IMAGERY WAS RETURNED. AVAILABLE INFORMATION SUGGESTS THAT WITHDRAWAL OF BURST BALLOON LIKELY CONTRIBUTED TO THE EVENT. AS THE BALLOON WAS BURST, THE ALTERED BALLOON PROFILE COULD HAVE MADE IT MORE SUSCEPTIBLE TO CATCH ON THE DISTAL END OF SHEATH TIP WHICH WOULD HAVE THEN LED TO THE EXPERIENCED RETRIEVAL DIFFICULTY. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159242&pc=BZD | 2518422-2025-002981 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND ESOPHAGEAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159250&pc=OZD | 1220648-2025-25590 | 2024-01-16 05:00 | Death | ABIOMED, INC. | 2025-01-15 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. PRIOR TO IMPELLA SUPPORT, THE PATIENT WAS CANNULATED WITH EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND WAS INTUBATED. IT WAS REPORTED THAT AFTER THE IMPELLA CP WITH SERIAL NUMBER (B)(6) (PUMP 1) WAS IMPLANTED AND CONNECTED TO THE AUTOMATED IMPELLA CONTROLLER (AIC) WITH SERIAL NUMBER (B)(6) (AIC 1), AN ALARM OCCURRED FOR IMPELLA DEFECTIVE. THE IMPELLA CP WAS EXPLANTED AND REPLACED WITH A NEW IMPELLA CP PUMP WITH SERIAL NUMBER (B)(6) (PUMP 2), AND THE IMPELLA DEFECTIVE ALARM WAS NOT RESOLVED. AN ADDITIONAL AIC WITH SERIAL NUMBER (B)(6) (AIC 2) WAS CONNECTED TO THE SECOND PUMP, AND THE IMPELLA DEFECTIVE ALARM WAS NOT RESOLVED. THE AIC WAS REPLACED AGAIN WITH A THIRD UNIT, THE ALARM WAS RESOLVED, AND SUPPORT WAS INITIATED. AFTER 6 DAYS OF ECMO AND IMPELLA SUPPORT, A CEREBRAL BLEED WAS DISCOVERED, AND THERAPY WAS DISCONTINUED. IT WAS NOTED THAT THE CEREBRAL BLEED WAS NOT IMPELLA RELATED. THE PATIENT WAS GIVEN SYSTEMIC HEPARIN DURING IMPELLA USE. IT WAS NOTED THAT THE PATIENT EXPIRED WHILE ON IMPELLA CP SUPPORT. THERE WAS A DELAY IN NEEDED SUPPORT WHEN PUMP 1 WOULD NOT BE RECOGNIZED BY AIC 1. ADDITIONALLY, THERE WAS A DELAY OF SUPPORT AS THE AIC 1 AND AIC 2 WOULD NOT RECOGNIZE PUMP 1 OR PUMP 2. IT CANNOT BE SAID THE PERIOD OF TIME WITH LOSS OF HEMODYNAMIC SUPPORT DID NOT CONTRIBUTE TO THE PATIENT'S OVERALL OUTCOME. THE PATIENT EXPERIENCED A CEREBRAL VASCULAR ACCIDENT WHILE ON SUPPORT WITH PUMP 2. WHILE THE PHYSICIAN STATED IT WAS NOT IMPELLA RELATED, BUT THERE IS NO ALTERNATIVE CAUSE IDENTIFIED OR PROVIDED AS THE LIKELY CAUSE OF SUCH EVENT. ADDITIONALLY, THE PATIENT WAS GIVEN SYSTEMIC HEPARIN FOR IMPELLA USE WHICH MAY HAVE CONTRIBUTED TO THE CEREBRAL BLEEDING. PUMP 1, AIC 1, AIC 2, AND PUMP 2 CANNOT BE RULED OUT AS CONTRIBUTING FACTORS IN THE OVERALL EVENT. Manufacturer Narrative: THIS IS ONE OF FOUR MEDICAL DEVICE REPORTS ASSOCIATED WITH THE EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2024-06292 FOR AUTOMATED IMPELLA CONTROLLER 1 WITH SERIAL NUMBER (B)(6). REFER TO MANUFACTURING REPORT NUMBER 1220648-2024-06294 FOR AUTOMATED IMPELLA CONTROLLER 2 WITH SERIAL NUMBER (B)(6). REFER TO MANUFACTURING REPORT NUMBER 1220648-2024-06290 FOR PUMP 1 IMPELLA CP WITH SERIAL NUMBER (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159327&pc=OZD | 1220648-2025-25592 | 2024-12-30 05:00 | Death | ABIOMED, INC. | 2025-01-15 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED ACCESS BLEEDING REQUIRING BLOOD TRANSFUSIONS. THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB AND A 99% OCCLUSION IN THE RIGHT CORONARY ARTERY WAS OBSERVED. TWO STENTS WERE PLACED, AND AN INTRA-AORTIC BALLOON PUMP WAS UNSUCCESSFULLY ATTEMPTED. IT WAS THEN DECIDED TO PLACE THE IMPELLA CP, WHICH WAS SUCCESSFULLY IMPLANTED. HOWEVER, THREE DAYS AFTER THE START OF SUPPORT, ACCESS SITE BLEEDING WAS OBSERVED, REQUIRING TWO UNITS OF PACKED RED BLOOD CELLS. NEXT STEPS WERE TO STENT THE LEFT ANTERIOR DESCENDING ARTERY AND LOOK AT THE LEFT CIRCUMFLEX ARTERY WITH THEREAFTER NEXT STEPS TO WEAN AND EXPLANT THE IMPELLA CP. OVERNIGHT, THE PATIENT'S URINARY OUTPUT DECREASED, AND THE PATIENT CONTINUED TO DECOMPENSATE. NO INTERVENTION WAS TAKEN TO CULPRIT LESION, AND THE PATIENT WOULD EXPIRE THE SAME DAY. FAMILY DECIDED TO WITHDRAW CARE. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEED COULD NOT BE DETERMINED. B.2 REVISED AS A THRID SELECTION WAS INCORRECTLY REPORTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25592.
E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25592 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA. SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159664&pc=NGV | 2124215-2024-83712 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR A LATE DEVICE THROMBOSIS (30 DAYS - 1 YEAR POST IMPLANT). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 124 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159741&pc=NKM | 2135147-2025-00319 | 2024-12-24 05:00 | Death | ABBOTT MEDICAL | 2025-01-15 05:00 | NKM | MITRACLIP | Difficult to Remove | Hemolysis; Low Blood Pressure/ Hypotension; Mitral Valve Stenosis; Heart Failure/Congestive Heart Failure; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4+, PROLAPSED POSTERIOR LEAFLET, PRE-EXISTING CHORDAE TENDINEAE RUPTURE ON THE POSTERIOR LEAFLET SIDE, MITRAL ANNULAR CALCIFICATION (MAC), AND A MEAN PRESSURE GRADIENT OF 4MMHG. A MITRACLIP NT WAS INSERTED, AND GRASPING WAS PERFORMED. THE LEAFLETS WERE GRASPED, BUT THE CLIP WAS INVERTED AND RETURNED TO THE LEFT ATRIUM (LA) FOR REPOSITIONING. HOWEVER, THE CLIP BECAME CAUGHT IN CHORDAE. IT WAS OBSERVED THAT THE CHORDAE TENDINEAE OF THE ANTERIOR LEAFLET RUPTURED, THE PATIENT'S BLOOD PRESSURE DECREASED, AND THE MEAN PRESSURE GRADIENT INCREASED TO 7MMHG. THE CLIP WAS REPOSITIONED CENTRALLY, BUT THE MR WAS UNABLE TO BE REDUCED. AFTER UNGRASPING THE LEAFLETS, THE PRESSURE GRADIENT RETURNED TO BASELINE. THE CLIP WAS REMOVED FROM THE PATIENT AND THE PROCEDURE WAS DISCONTINUED WITH NO CLIPS IMPLANTED. ON THE FOLLOWING DAY, CARDIAC FAILURE LIKELY DUE TO THE CHORDAE RUPTURE OCCURRED, HEMOLYSIS OF AN UNKNOWN ETIOLOGY DUE TO CARDIAC FAILURE OCCURRED, AND THE PATIENT DIED. THE CAUSE OF DEATH WAS DUE TO HEART FAILURE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4+, PROLAPSED POSTERIOR LEAFLET, PRE-EXISTING CHORDAE TENDINEAE RUPTURE ON THE POSTERIOR LEAFLET SIDE, MITRAL ANNULAR CALCIFICATION (MAC), AND A MEAN PRESSURE GRADIENT OF 4MMHG. A MITRACLIP NT WAS INSERTED, AND GRASPING WAS PERFORMED. THE LEAFLETS WERE GRASPED, BUT THE CLIP WAS INVERTED AND RETURNED TO THE LEFT ATRIUM (LA) FOR REPOSITIONING. HOWEVER, THE CLIP BECAME CAUGHT IN CHORDAE. IT WAS OBSERVED THAT THE CHORDAE TENDINEAE OF THE ANTERIOR LEAFLET RUPTURED, THE PATIENT'S BLOOD PRESSURE DECREASED, AND THE MEAN PRESSURE GRADIENT INCREASED TO 7MMHG. THE CLIP WAS REPOSITIONED CENTRALLY, BUT THE MR WAS UNABLE TO BE REDUCED. AFTER UNGRASPING THE LEAFLETS, THE PRESSURE GRADIENT RETURNED TO BASELINE. THE CLIP WAS REMOVED FROM THE PATIENT AND THE PROCEDURE WAS DISCONTINUED WITH NO CLIPS IMPLANTED. ON THE FOLLOWING DAY, CARDIAC FAILURE LIKELY DUE TO THE CHORDAE RUPTURE OCCURRED, HEMOLYSIS OF AN UNKNOWN ETIOLOGY DUE TO CARDIAC FAILURE OCCURRED, AND THE PATIENT DIED. THE CAUSE OF DEATH WAS DUE TO HEART FAILURE. SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: PRIOR TO THE MITRACLIP PROCEDURE, THE PATIENT WAS HEMOLYZED, HAD LOW BLOOD PRESSURE, AND WAS PLACED ON AN INTRA-AORTIC BALLOON PUMP (IABP) AND A VENTILATOR. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON AVAILABLE INFORMATION, THE REPORTED DIFFICULT TO REMOVE FROM ANATOMY WAS DUE TO PROCEDURAL CIRCUMSTANCES (USER TECHNIQUE OF POSITIONING). THE REPORTED TISSUE INJURY WAS A CASCADING EFFECT OF THE REPORTED DIFFICULT TO REMOVE FROM ANATOMY. THE REPORTED MITRAL STENOSIS APPEARS TO BE A CASCADING EFFECT OF THE REPORTED TISSUE INJURY. THE CAUSE OF THE REPORTED DEATH, HEART FAILURE, HEMOLYSIS, AND HYPOTENSION WERE UNABLE TO BE DETERMINED. THE REPORTED PATIENT EFFECTS OF DEATH, HYPOTENSION, MITRAL STENOSIS, HEART FAILURE, AND TISSUE INJURY, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED SERIOUS INJURY/ ILLNESS/ IMPAIRMENT WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159862&pc=NGV | 2124215-2024-83710 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Endocarditis |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ENDOCARDITIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 124 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159871&pc=NGV | 2124215-2024-83711 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. PATIENT SEX: 1 FEMALE; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 100 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159901&pc=NGV | 2124215-2024-83709 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 9 | Event Description: THIS REPORT SUMMARIZES 9 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 73 TO 92 YEARS. PATIENT SEX: 5 FEMALES; 4 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 147 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21159993&pc=BZD | 2518422-2025-002911 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cough; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, COUGH, SINUS INFECTION AND UPPER AIRWAY IRRITATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160053&pc=FRN | 3012307300-2025-00756 | 2024-12-14 05:00 | Death | SMITHS MEDICAL ASD, INC. | 2025-01-15 05:00 | FRN | CADD-SOLIS VIP AMBULATORY INFUSION PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K111275 |  | 1 | Event Description: A REPORT OF A PATIENT¿S DEATH WAS RECEIVED BY ICU MEDICAL ON (B)(6)2024 WITH UPDATE ON (B)(6)2024, THIS PATIENT IS REPORTED TO HAVE A INFUSION ONGOING AT THE TIME OF THE REPORTED INCIDENT. LACK OF DETAILS REGARDING THIS CLINICAL EVENT PRECLUDES A COMPREHENSIVE MEDICAL ASSESSMENT. THERE IS A REPORT FROM COLORADO BLOOD CANCER INSTITUTE OF A CADD INFUSION PUMP SYSTEM WAS DELIVERING A MEDICATION AT 5 ML/HR AT HOME WITH NO REPORTED ISSUE. THERE IS NO HISTORY OF WHAT MEDICATION WAS BEING DELIVERED TO THE PATIENT, NO PAST MEDICAL HISTORY AND NO HISTORY OF OTHER MEDICATIONS THAT THE PATIENT WAS CONCURRENTLY TAKING. IT IS REPORTED THAT THIS WAS A FEMALE PATIENT, AND THE PATIENT WAS FOUND BY HER SISTER AT HOME AND THE DATE OF THE PATIENT¿S DEATH WAS REPORTED AS (B)(6)2024. NO FURTHER DETAILS OF THE FUNCTIONALITY OF THE PUMP OR THE CONDITION OF THE PATIENT ARE PROVIDED. THERE IS A GENERAL LACK OF INFORMATION PROVIDED IN THIS CASE TO MAKE A DETERMINATION OF CAUSALITY. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THIS MEDICAL ASSESSMENT WILL BE FURTHER REVISED AS NEEDED. Manufacturer Narrative: E1 INITIAL REPORTER FULL PHONE NUMBER: (B)(6). H3, H6: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160057&pc=NGV | 2124215-2024-83708 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 76 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160246&pc=NPT | 2025587-2025-00388 | 2024-06-09 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-15 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 1 WEEK FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WAS FOUND UNRESPONSIVE BY A FAMILY MEMBER AT AROUND 23:00. THE PATIENT WAS TRANSFERRED TO THE HOSPITAL BY EMERGENCY SERVICES WHERE CARDIOPULMONARY ARREST WAS NOTED. PER THE PATIENT'S WISHES, RESUSCITATION WAS NOT PERFORMED. THE PATIENT'S TIME OF DEATH WAS CONFIRMED THE NEXT DAY. PER THE PHYSICIAN, THERE IS NO CAUSAL RELATIONSHIP BETWEEN THE DEVICE OR THE PROCEDURE AND THE PATIENT'S CAUSE OF DEATH. THE CAUSE OF DEATH WAS DESCRIBED AS BEING RELATED TO CARDIOVASCULAR EVENTS. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160327&pc=NGV | 2124215-2024-83572 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 102 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160385&pc=DSQ | 2916596-2025-00192 | 2024-09-11 04:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Device Difficult to Setup or Prepare; Gas/Air Leak; Adverse Event Without Identified Device or Use Problem | Air Embolism; Stroke/CVA; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN AS IT WAS NOT REPORTED BY THE CLINIC. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS WORKING AS INTENDED DURING TIME OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT DURING IMPLANTATION OF HEARTMATE 3 ON (B)(6) 2024 AN INTRA-OPERATIVE AIR ENTRAINMENT OCCURRED BETWEEN APICAL CUFF AND THE PUMP, LEADING TO STROKE.Event Description: IT WAS REPORTED THAT AIR WAS DETECTED SEVERAL TIMES, FIRST DURING LVAD RAMP-UP AND THEN DURING WEANING FROM CARDIOPULMONARY BYPASS (CBP). THE CPB WAS THEN RAMPED UP AGAIN, THE LVAD SPEED REDUCED, AND THE DE-AIRING REPEATED. THIS SEQUENCE OCCURRED THREE TIMES IN TOTAL. DE-AIRING WAS PERFORMED, ALL ANESTHESIA CATHETERS WERE CHECKED FOR POSSIBLE AIR ENTRAINMENT, AND A BRONCHOSCOPY WAS PERFORMED TO RULE OUT PULMONARY VASCULAR INJURY AND A POSSIBLE PULMONARY FISTULA. REPERFUSION WAS REPEATED WITH A CAREFUL DE-AIRING MANEUVER UNTIL ABSOLUTELY NO AIR WAS VISIBLE ON THE ECHOCARDIOGRAM. IT WAS ATTEMPTED AGAIN TO WEAN FROM THE HEART LUNG MACHINE, WHICH WAS SUCCESSFUL. THERE WERE NO SIGNS OF AIR FOUND IN THE LEFT VENTRICLE. VENOUS DECANNULATION WAS PERFORMED WITH THE START OF PROTAMINE. A SUDDEN RIGHT HEART PROBLEM OCCURRED WITH SIGNS OF AIR IN THE AORTA EXTENDING INTO THE AORTIC CANNULA. THE SUPERIOR VENA CAVA (SVC) WAS IMMEDIATELY RE-CANNULATED. THE VENT AND CPB WERE RESTARTED. ADDITIONAL SUTURES WERE MADE AT THE APEX TO COUNTERACT ANY AIR INTAKE BY THE PUMP. A LARGE AMOUNT OF FIBRIN GLUE WAS THEN APPLIED TO SUPPORT HERMETICALLY SEALING. FIBRILLATION WAS INDUCED AGAIN, AND THE PUMP WAS DISCONNECTED AND RECONNECTED. AFTER THESE MANEUVERS, NO AIR WAS VISIBLE AND ANOTHER ATTEMPT TO WEAN WAS STARTED. THIS WAS ULTIMATELY SUCCESSFUL, AND NO FURTHER AIR PHENOMENA WERE FOUND IN THE AORTA.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE RIGHT HEART PROBLEM WAS DESCRIBED AS ACUTE RIGHT HEART FAILURE. THE REASON FOR THE TEMPORARY RIGHT HEART PROBLEM WAS NOTED AS PROBABLY CAUSED BY AIR IN THE RIGHT CORONARY ARTERY. Manufacturer Narrative: SECTION B3: DATE OF EVENT CORRECTED. SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE AIR ENTRAINMENT, REPORTED TO BE BETWEEN THE APICAL CUFF AND PUMP, DURING IMPLANT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED ADVERSE EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE DEVICE WILL NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS STROKE, RIGHT HEART FAILURE, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5 ¿SURGICAL PROCEDURES¿ PROVIDES INFORMATION ON ALL SURGICAL PROCEDURES, INCLUDING PRIMING THE PUMP (SUBSECTION ENTITLED ¿PREPARING, RUNNING, AND PRIMING THE PUMP¿) AND DE-AIRING THE PUMP (SUBSECTION ENTITLED ¿DE-AIRING THE PUMP¿). SECTION 5 CONTAINS SEVERAL STEPS TO PREVENT ENTRAPPED AIR AND ENSURE ENTRAPPED AIR IS REMOVED. THE IFU WARNS THAT ALL ENTRAPPED AIR MUST BE REMOVED FROM THE DEVICE TO MINIMIZE THE RISK OF AIR EMBOLUS. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SHOULD BE UTILIZED TO MONITOR FOR AIR EMBOLI. SECTION 5 WARNS THAT PROLONGED DE-AIRATION MAY BE DUE TO INADEQUATE BLOOD SUPPLY TO THE LEFT VENTRICULAR ASSIST DEVICE OR A LEAK IN THE SEALED OUTFLOW GRAFT OR SEALED INFLOW CANNULA. IF AIR PERSISTS IN THE PUMP SEALED OUTFLOW GRAFT FOR A PROLONGED PERIOD (MORE THAN 5¿10 MINUTES), RULE OUT LEAKS AT THE SEALED INFLOW CANNULA/PUMP CONNECTION. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ ALSO PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. ADDITIONALLY, SECTION 6 LISTS RIGHT HEART FAILURE AS A POTENTIAL RISK/ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. THE DEVICE IS INCLUDED IN THE HM3 INFLOW SEAL INTERFACE WITH APICAL CUFF NOTICE ISSUED BY ABBOTT ON 19MAR2024. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B1: CORRECTED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: A4 - PATIENT WEIGHT NO PROVIDED, B3: EVENT DATE, UPDATED, B5: ADDITIONAL INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160396&pc=NGV | 2124215-2024-83573 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENTS FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 410 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160397&pc=NGV | 2124215-2024-83575 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 361 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160402&pc=NGV | 2124215-2024-83574 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 11 | Event Description: THIS REPORT SUMMARIZES 11 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: (B)(6) YEARS. PATIENT SEX: 8 FEMALE; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 241 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160504&pc=DSQ | 2916596-2025-00339 | 2024-09-15 04:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hematoma | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO MOTOR VEHICLE ACCIDENT (MVA) WITH SUBDURAL HEMATOMA. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT RETURNED FOR EVALUATION.Event Description: THERE WAS NO EVIDENCE THAT SUGGESTED THE DEVICE MALFUNCTION. THERE WERE NO LOG FILES COLLECTED AS THERE WERE NO ALARMS OR ABNORMALITIES NOTED UPON INTERROGATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "ANTICOAGULATION"), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160650&pc=JOL | 1423537-2025-00033 | 2024-06-25 04:00 | Death | CARDINAL HEALTH, INC. | 2025-01-15 05:00 | JOL | ARGYLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | EXEMPT |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THERE WAS A PARTIAL DISLODGE OF CHEST TUBE WITH EXPOSED EYELET AND SUTURE FELL OFF. NO UNUSUAL CIRCUMSTANCES NOTED. THE CLINICAL STAFF NOTICED WORSENING AIR LEAK AND INFORMED THE TEAM. THE NURSE PRACTITIONER (NP) AND SURGEON CAME TO ASSESS AND DISCOVERED PARTIAL DISLODGE. THE CHEST TUBE WAS REMOVED. THE PATIENT DETERIORATED OVER THE NEXT FEW DAYS, TRANSFERRED TO ICU AND INTUBATED. THIS WAS MOST LIKELY DUE TO UNDERLYING CLINICAL FACTORS, BUT THE CHEST TUBE DISLODGES LIKELY CONTRIBUTED TO DETERIORATION AS WELL (<50%). IT IS VERY UNUSUAL TO HAVE A PARTIAL DISLODGE WITH EYELET EXPOSURE, SO THEY HAD TO PROVIDE ADDITIONAL STAFF EDUCATION AND UPDATE THEIR POLICY TO INCLUDE THIS SITUATION. PER CUSTOMER, THE CHEST TUBE WAS IN FOR 5 DAYS. ADDITIONAL INFORMATION WAS RECEIVED FROM THE CUSTOMER AND STATED THAT THERE WAS NO DETACHMENT, SPLITS OR BREAKS. DUE TO THE TUBE BEING VERY RIGID AND DIFFICULT TO SUTURE TO THE CHEST WALL, AN EYELET THAT SHOULD HAVE BEEN BELOW THE DERMIS IN THE CHEST WALL TO FACILITATE DRAINAGE WAS EXPOSED. THE PATIENT HAD MULTIPLE COMORBIDITIES; DETERIORATION NOT LIKELY DUE TO THE PARTIAL DISLODGE. PATIENT DETERIORATION WAS ATTRIBUTED TO PATIENT SEDATION AND RESPIRATORY ACIDOSIS. THE PATIENT WAS INTERMITTENTLY CONFUSED AND REFUSING HIGH FLOW NON-REBREATHER MASK, BIPAP AND OPTIFLOW WITH RT. THE PATIENT DIED IN ICU. Manufacturer Narrative: SECTION D4 (MODEL NUMBER AND CATALOG NUMBER): ORIGINALLY REPORTED AS 8888571042 AND SHOULD BE 8888570549. SECTION G3/G4 (PMA/510(K)NUMBER): ORIGINALLY REPORTED AS E223935 AND SHOULD BE EXEMPT. Manufacturer Narrative: THE DEVICE HISTORY RECORD FOR LOT 2327900464 WAS REVIEWED AND NO ANOMALIES RELATED TO THE REPORTED ISSUE WERE OBSERVED. THE DEVICE WAS MANUFACTURED ON OCTOBER 10, 2023. THERE WERE NO PHOTOS OR PHYSICAL SAMPLES RECEIVED FOR EVALUATION. BECAUSE A SAMPLE WAS NOT RETURNED, WE WERE UNABLE TO PERFORM A FOLLOW UP INVESTIGATION TO INCLUDE FUNCTIONAL AND VISUAL EVALUATIONS TO CONFIRM THE ISSUE AND ROOT CAUSE ANALYSIS. BASED ON THIS INFORMATION, NO CORRECTIVE ACTIONS ARE WARRANTED AT THIS TIME. WE WILL CONTINUE TO TRACK AND TREND THROUGH OUR MONITORING PROGRAMS AND TAKE ACTIONS AS APPLICABLE. Manufacturer Narrative: AN INVESTIGATION IS CURRENTLY UNDERWAY. UPON COMPLETION, THE RESULTS WILL BE FORWARDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160707&pc=DSQ | 2916596-2025-00078 | 2024-12-29 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Tachycardia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 FOLLOWING A VENTRICULAR TACHYCARDIA STORM. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED.Event Description: CARDIOVERSION WAS PERFORMED AS TREATMENT. THE PATIENT'S ARRHYTHMIA WAS SUSTAINED. IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF ARRHYTHMIA PRE-LAVD, HOWEVER THEY DID NOT HAVE AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR PRIOR TO THE LVAD. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: CORRECTED DATA: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING CARDIAC ARRYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATIONS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161091&pc=DSQ | 2916596-2025-00255 | 2020-10-18 04:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "THE IMPACT OF THE 2018 ALLOCATION SYSTEM CHANGE ON PATIENTS BRIDGED WITH DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD): AN UPDATED UNOS REGISTRY ANALYSIS" THAT THE HEARTMATE 3 (HM3) PATIENTS EXPERIENCED HEART TRANSPLANT (HT) AND WAITLIST MORTALITY. OF NOTE, PATIENTS IMPLANTED WITH HEARTMATE II AND HEARTWARE WERE INCLUDED IN THE ANALYSIS. THIS ARTICLE LEVERAGED THE UNITED NETWORK FOR ORGAN SHARING (UNOS) REGISTRY TO MEASURE THE IMPACT OF THE 2018 ALLOCATION SYSTEM CHANGE (ASC) PRIMARILY BY OBSERVING 1 YEAR POST HT MORTALITY PRIMARILY, AND SECONDARILY, WAITLIST MORTALITY AND HEART TRANSPLANT RATES. RESULTS FOUND THAT ALTHOUGH LISTING POST-ASC WAS ASSOCIATED WITH HIGHER HT RATES, THE LIKELIHOOD OF HT WAS LOWER AMONG PATIENTS WITH VS WITHOUT LVAD IN THE POST-ASC: HAZARD RATIO (HR) 0.97 VS POST-ASC: HR 0.75 PINTERACTION <0.001. LVAD DEVICE SUPPORT WAS ASSOCIATED WITH LOWER HAZARD OF WAITLIST DEATH AT 1 YEAR FOLLOW UP IN BOTH ERAS, BUT THE SURVIVAL BENEFIT FROM LVAD SUPPORT TWAS SIGNIFICANTLY LOWER IN THE POST-ASC PERIOD (PRE-ASC: HR, 0.28 VS. POST-ASC: HR, 0.55; P INTERACTION = 0.006). ONE YEAR POST-HT MORTALITY WAS NOT DIFFERENT BETWEEN PATIENTS WITH VERSUS WITHOUT LVAD IN THE PRE-ASC ERA (HR, 1.19; P = 0.11, TABLE 1), WHEREAS IT WAS 58% HIGHER AMONG PATIENTS WITH VERSUS WITHOUT LVAD AFTER THE ASC (HR, 1.58; P < 0.001; PINTERACTION = 0.043). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. FROM THE *DEPARTMENT OF INTERNAL MEDICINE, CLEVELAND CLINIC FOUNDATION, CLEVELAND, OHIO; ¿DEPARTMENT OF BIOSTATISTICS, UNIVERSITY OF WASHINGTON, SEATTLE, WASHINGTON; ¿DEPARTMENT OF INTERNAL MEDICINE, UNIVERSITY OF MIAMI MILLER SCHOOL OF MEDICINE, PALM BEACH CAMPUS, ATLANTIS, FLORIDA; AND §DIVISION OF CARDIOLOGY, DEPARTMENT OF INTERNAL MEDICINE, UNIVERSITY OF WASHINGTON MEDICAL CENTER, SEATTLE, WASHINGTON. BRAGHIERI, L., ET AL. (2024). THE IMPACT OF THE 2018 ALLOCATION SYSTEM CHANGE ON PATIENTS BRIDGED WITH DURABLE LEFT VENTRICULAR ASSIST DEVICE: AN UPDATED UNOS REGISTRY ANALYSIS. ASAIO JOURNAL, 70(5), 456¿459. HTTPS://DOI.ORG/10.1097/MAT.0000000000002115. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), DOCUMENT REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161103&pc=BZD | 2518422-2025-003026 | 2023-08-14 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | REP DREAMSTATION AUTO BIPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANTED TO RETURN THE REPLACEMENT DEVICE BACK AS IT WAS NO LONGER NEEDED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161198&pc=LWS | 9614453-2025-00251 | 2024-10-17 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-01-15 05:00 | LWS | COBALT¿ DR MRI SURESCAN¿ | Defibrillation/Stimulation Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED. AN INTERNAL REVIEW OF DATA TRANSMITTED FROM THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) INDICATED THAT A SECOND PHASE SHORT CIRCUIT PROTECTION WAS TRIGGERED DURING HIGH VOLTAGE THERAPY RESULTING IN LESS THAN THE PROGRAMMED HIGH VOLTAGE ENERGY BEING DELIVERED. THE HIGH VOLTAGE SHOCKS WERE CLASSIFIED AS SUCCESSFUL BY THE DEVICE, HOWEVER ELECTROGRAMS (EGM) RECORDINGS SUGGEST THAT THE RHYTHM CONTINUED BELOW THE DETECTION INTERVAL. IN ADDITION, IT WAS NOTED THAT THE RIGHT VENTRICULAR (RV) LEAD HAD EXHIBITED INTERMITTENT T-WAVE OVERSENSING (TWOS) AND ELEVATED SENSING INTEGRITY COUNTER (SIC) APPROXIMATELY ONE YEAR PRIOR. THE DISPOSITION OF THE ICD AND RV LEAD REMAIN UNKNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED THAT THE VENTRICULAR TACHYARRHYTHMIA THERAPY ENERGY WAS LOW. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 6947M62 (TDK105870V); PRODUCT TYPE: 0264-LEAD-HV; IMPLANT DATE (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161336&pc=BZD | 2518422-2025-003025 | 2022-12-07 05:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161600&pc=NGV | 2124215-2024-83671 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 120 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161624&pc=NVY | 2017865-2025-10932 | 2024-12-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-15 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10931. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10933. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10935. IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC ARREST. THERE WAS NO ALLEGATION OF MALFUNCTION. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: CORRECTION TO REPORT TYPE. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. ELECTRICAL TESTS THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL AND X-RAY INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161633&pc=NVN | 2017865-2025-10933 | 2024-12-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-15 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10931. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10932. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10935. IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC ARREST. THERE WAS NO ALLEGATION OF MALFUNCTION. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF A PARTIAL LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TEST DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES HOWEVER AN INTERNAL SHORT WAS NOTED BETWEEN CONDUCTOR PATHS DUE TO PROCEDURAL DAMAGE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161656&pc=NIK | 2017865-2025-10935 | 2024-12-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-15 05:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10931. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10932. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10933. IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC ARREST. THERE WAS NO ALLEGATION OF MALFUNCTION. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY CONNECTOR PORTION OF A PARTIAL LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTS DID NOT FIND ANY INDICATION OF ANY CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161695&pc=NGV | 2124215-2024-83670 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 13 | Event Description: THIS REPORT SUMMARIZES 13 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 66 TO 87 YEARS. PATIENT SEX: 4 FEMALE; 9 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 203 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161719&pc=NGV | 2124215-2024-83669 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. PATIENT SEX: 1 FEMALE; 1 MALE TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 280 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161720&pc=NGV | 2124215-2024-83673 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. PATIENT SEX: 1 FEMALE; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 280 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161722&pc=NGV | 2124215-2024-83667 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. PATIENT SEX: 2 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 401 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161982&pc=DSQ | 2916596-2025-00077 | 2024-12-29 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage; Renal Failure; Cognitive Changes | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN UNWITNESSED FALL AT HOME ON (B)(6) 2024 AND REFUSED TO COME IN FOR EVALUATION UNTIL (B)(6) 2024 WHEN THEIR PARTNER CALLED EMERGENCY MEDICAL SERVICES. THEY EXPERIENCED ACUTE KIDNEY INJURY AND INTRAVENTRICULAR HEMORRHAGE AND SUBARACHNOID HEMORRHAGE WITH ALTERED MENTAL STATUS. THE PATIENT ULTIMATELY PASSED AWAY ON (B)(6) 2024 DUE TO RENAL FAILURE AND HEAD BLEED. FURTHER INFORMATION PROVIDED THAT LVAD THERAPY DID NOT WORSEN RENAL DISEASE, AS THE ACUTE KIDNEY INJURY WAS ACUTE DURING HOSPITALIZATION. THE PATIENT LOG FILES WERE NOT AVAILABLE, THE DEVICE OPERATED AS EXPECTED, AND THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND RENAL DYSFUNCTION, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS DOCUMENT ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND RENAL DYSFUNCTION, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21162019&pc=NIK | 2017865-2025-10931 | 2024-12-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-15 05:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10932. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10933. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-10935. IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC ARREST. THERE WAS NO ALLEGATION OF MALFUNCTION. FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: INTERROGATION OF THE DEVICE REVEALED IT WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) WHEN RECEIVED. DEVICE SENSING, PACING, LEAD IMPEDANCE, HIGH VOLTAGE CHARGING, AND HIGH VOLTAGE SHOCK IMPEDANCE WERE TESTED, AND THE FOUND TO BE NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21162051&pc=DSQ | 2916596-2025-00128 | 2024-12-25 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Sepsis; Aspiration Pneumonitis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO SEPSIS. IT WAS STATED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE PATIENT GOT CHOCKING THEN BECAME A SEPTIC SHOCK.Event Description: ADDITIONAL INFORMATION STATED THAT THE PATIENT EXPERIENCED CHOKING WHICH THEN BECAME ASPIRATION PNEUMONIA AND SEPTIC SHOCK ON (B)(6) 2024. THE PATIENT EXPERIENCED SHORTNESS OF BREATH AND WAS INTUBATED WITH ENDOTRACHEAL TUBE. THE SOURCE OF INFECTION AND CULTURES WERE NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE OF THE PATIENT¿S INFECTION AND SEPSIS COULD NOT BE DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET) AND SEPSIS, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21162096&pc=DSQ | 2916596-2025-00129 | 2024-12-27 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO INTRACEREBRAL HEMORRHAGE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21162344&pc=BZD | 2518422-2025-003039 | 2023-01-27 05:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21162562&pc=BZD | 2518422-2025-003078 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Asthma; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21163510&pc=NIK | 2124215-2025-01607 | 2024-12-27 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | NIK | MOMENTUM CRT-D | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS PART OF A SYSTEM REVISION DUE TO POSSIBLE INFECTION. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS CRT-D WAS TURNED OFF BUT STILL IMPLANTED.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) HAD EXHIBITED POSSIBLE INFECTION. THE PHYSICIAN DECIDED TO SWITCH OFF THE CRT-D AS WELL AS THE LEFT BUNDLE BRANCH (LBB) AND RIGHT VENTRICULAR (RV) LEAD. SUBSEQUENTLY, A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD SYSTEMIC INFECTION. FURTHERMORE, THIS DEVICE WAS NOT EXPLANTED BUT WAS MADE OUT OF SERVICE (OOS) DUE TO THE PATIENT PASSING AWAY.Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) WAS PART OF A SYSTEM REVISION DUE TO POSSIBLE INFECTION. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED; HOWEVER, THE INFECTION'S CAUSE WAS NOT IDENTIFIED. IN THE INTENSIVE CARE UNIT (ICU), THE PATIENT RECEIVED PRESSORS AND INTRAVENOUS ANTIBIOTICS. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS CRT-D WAS TURNED OFF BUT STILL IMPLANTED. ADDITIONAL INFORMATION INDICATED THAT THE PATIENT HAD SYSTEMIC INFECTION. FURTHERMORE, THIS DEVICE WAS NOT EXPLANTED BUT WAS MADE OOS DUE TO THE PATIENT PASSING AWAY. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT WAS CREATED TO UPDATE H6: PATIENT CODES WITH SEPSIS. Manufacturer Narrative: AFTER REVIEW, IT COULD NOT BE RULED OUT THAT THE DEVICE OR PROCEDURE HAD INVOLVEMENT IN THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21163544&pc=DXE | 2124215-2025-01298 | 2025-01-02 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-15 05:00 | DXE | ANGIOJET SOLENT OMNI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. AN ANGIOJET SOLENT OMNI WAS SELECTED FOR USE IN A THROMBECTOMY PROCEDURE. THE OCCLUSION WAS LOCATED IN AN ARTERIOVENOUS FISTULA LOOP IN THE LEFT UPPER ARM. GENERAL ANESTHESIA WAS ADMINISTERED. A PUNCTURE WAS MADE IN THE RIGHT COMMON FEMORAL ARTERY FOR VASCULAR ACCESS. A 5F SHEATH WAS PLACED AND A 4F VERTEBRAL CATHETER WAS ADVANCED INTO THE LEFT AXILLARY ARTERY. ANGIOGRAPHY WAS PERFORMED. A NON-BSC GUIDEWIRE WAS INSERTED THROUGH THE FISTULA LOOP, AND THE TIP WAS PUSHED INTO THE DRAINING BRACHIAL VEIN. THE CATHETER WAS FURTHER ADVANCED TO THE PROXIMAL SEGMENT OF THE THROMBOSED LOOP. ANOTHER PUNCTURE FOR VASCULAR ACCESS WAS MADE IN THE RIGHT FEMORAL COMMON VEIN, AND A NON-BSC 6F SHEATH WAS INSERTED INTO THE DISTAL SEGMENT OF THE LOOP. THE ANGIOJET SOLENT OMNI WAS INSERTED, AND THROMBOASPIRATION WAS PERFORMED FIRST IN THE ARTERIAL SEGMENT OF THE LOOP, AND THEN OF THE DISTAL VENOUS SEGMENT. ASPIRATION TIME WAS APPROXIMATELY FOUR TO FOUR AND A HALF MINUTES, AND VERY SLOWLY AT 2MM TO 3MM PER SECOND. ANGIOGRAPHY WAS AGAIN PERFORMED. VIA THE VENOUS ACCESS SHEATH, A NON-BSC STENT WAS THEN IMPLANTED INTO THE AFFECTED PROXIMAL SEGMENT WHERE THE STENT EXTENDS FROM THE BRACHIAL ARTERY TO JUST DISTAL TO THE PRESENT STENOSIS. POST-DILATION WAS PERFORMED WITH A 6MM NON-BSC BALLOON. A SECOND NON-BSC STENT WAS THEN IMPLANTED INTO THE AFFECTED VENOUS SEGMENT OF THE FISTULA. THE PATIENT EXPERIENCED SUDDEN HYPOTENSION UP TO 40/20MMHG AND THEN ASYSTOLE OCCURRED WHICH REQUIRED RESUSCITATION. AFTER SUCCESSFUL RESUSCITATION, ANGIOGRAPHIC INSPECTION SHOWS AN OBSTRUCTIVE EMBOLISM IN THE LEFT BRACHIAL ARTERY, DISTAL TO AND OPPOSITE THE FIRST-PLACED STENT. IT WAS PRESUMED TO BE A PROTRUDING CHRONIC CLOT FROM THE PROXIMAL SEGMENT OF THE FISTULA. A NON-BSC STENT WAS IMPLANTED AT THE LEVEL OF THE OBSTRUCTION. ANGIOGRAPHIC INSPECTION SHOWS PERFECT POSITION AND GOOD EXPANSION OF ALL STENTS AND ALSO OF THE HEMODIALYSIS LOOP. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED TO LOOK FOR A POSSIBLE PULMONARY EMBOLISM OR OTHER EMBOLI, GIVEN THE RESUSCITATION, BUT IT WAS NEGATIVE. WHILE STAYING IN INTENSIVE CARE, THE PATIENT EXPERIENCED ANOTHER CARDIAC ARREST AND PASSED AWAY.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. AN ANGIOJET SOLENT OMNI WAS SELECTED FOR USE IN A THROMBECTOMY PROCEDURE. THE OCCLUSION WAS LOCATED IN AN ARTERIOVENOUS FISTULA LOOP IN THE LEFT UPPER ARM. GENERAL ANESTHESIA WAS ADMINISTERED. A PUNCTURE WAS MADE IN THE RIGHT COMMON FEMORAL ARTERY FOR VASCULAR ACCESS. A 5F SHEATH WAS PLACED AND A 4F VERTEBRAL CATHETER WAS ADVANCED INTO THE LEFT AXILLARY ARTERY. ANGIOGRAPHY WAS PERFORMED. A NON-BSC GUIDEWIRE WAS INSERTED THROUGH THE FISTULA LOOP, AND THE TIP WAS PUSHED INTO THE DRAINING BRACHIAL VEIN. THE CATHETER WAS FURTHER ADVANCED TO THE PROXIMAL SEGMENT OF THE THROMBOSED LOOP. ANOTHER PUNCTURE FOR VASCULAR ACCESS WAS MADE IN THE RIGHT FEMORAL COMMON VEIN, AND A NON-BSC 6F SHEATH WAS INSERTED INTO THE DISTAL SEGMENT OF THE LOOP. THE ANGIOJET SOLENT OMNI WAS INSERTED, AND THROMBOASPIRATION WAS PERFORMED FIRST IN THE ARTERIAL SEGMENT OF THE LOOP, AND THEN OF THE DISTAL VENOUS SEGMENT. ASPIRATION TIME WAS APPROXIMATELY FOUR TO FOUR AND A HALF MINUTES, AND VERY SLOWLY AT 2MM TO 3MM PER SECOND. ANGIOGRAPHY WAS AGAIN PERFORMED. VIA THE VENOUS ACCESS SHEATH, A NON-BSC STENT WAS THEN IMPLANTED INTO THE AFFECTED PROXIMAL SEGMENT WHERE THE STENT EXTENDS FROM THE BRACHIAL ARTERY TO JUST DISTAL TO THE PRESENT STENOSIS. POST-DILATION WAS PERFORMED WITH A 6MM NON-BSC BALLOON. A SECOND NON-BSC STENT WAS THEN IMPLANTED INTO THE AFFECTED VENOUS SEGMENT OF THE FISTULA. THE PATIENT EXPERIENCED SUDDEN HYPOTENSION UP TO 40/20MMHG AND THEN ASYSTOLE OCCURRED WHICH REQUIRED RESUSCITATION. AFTER SUCCESSFUL RESUSCITATION, ANGIOGRAPHIC INSPECTION SHOWS AN OBSTRUCTIVE EMBOLISM IN THE LEFT BRACHIAL ARTERY, DISTAL TO AND OPPOSITE THE FIRST-PLACED STENT. IT WAS PRESUMED TO BE A PROTRUDING CHRONIC CLOT FROM THE PROXIMAL SEGMENT OF THE FISTULA. A NON-BSC STENT WAS IMPLANTED AT THE LEVEL OF THE OBSTRUCTION. ANGIOGRAPHIC INSPECTION SHOWS PERFECT POSITION AND GOOD EXPANSION OF ALL STENTS AND ALSO OF THE HEMODIALYSIS LOOP. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED TO LOOK FOR A POSSIBLE PULMONARY EMBOLISM OR OTHER EMBOLI, GIVEN THE RESUSCITATION, BUT IT WAS NEGATIVE. WHILE STAYING IN INTENSIVE CARE, THE PATIENT EXPERIENCED ANOTHER CARDIAC ARREST AND PASSED AWAY. IT WAS FURTHER REPORTED THAT THE ANGIOJET PERFORMED AS INTENDED WITHOUT ANY PROCEDURAL MISSTEPS. HYPOTENSION BEGAN AFTER THE PLACEMENT OF THE FIRST STENT. THE HYPOTENSION WAS NOTED TO BE DUE TO THE LACK OF CARDIAC OUTPUT, NOT BRADYCARDIA. IT WAS THOUGHT IT WAS A POSSIBLY PULMONARY ISSUE. THERE WAS NO OBSTRUCTIVE EMBOLISM IN THE LEFT BRACHIAL ARTERY, RATHER STENOSIS WAS OBSERVED. THE PROXIMAL STENOSIS ON THE ARTERIAL SIDE WAS DIFFICULT, AND THE STENT REQUIRED POST-DILATION DUE TO THE AMOUNT OF THE STENOSIS. BLOOD PRESSURE CONTINUED TO DROP, AND THE PATIENT WAS ADMINISTERED MEDICATION. THE PATIENT WAS RESUSCITATED WHICH TOOK A WHILE. A SECOND STENT WAS PLACED, AND THEN A THIRD STENT WAS PLACED IN THIS LESION TO PREVENT ISCHEMIA SYMPTOMS OF THE PATIENTS ARM. THE PATIENT REMAINED INTUBATED POST-RESUSCITATION AND WAS MOVED FROM THE ANGIOJET LAB TO CT. THE CT SCAN WAS NEGATIVE. THE PATIENT BECAME UNSTABLE IN THE ICU WITH SYMPTOMS OF HYPOTENSION AND BRADYCARDIA. IT WAS DETERMINED THE PATIENT WAS NOT POSSIBLE TO TREAT; THE PATIENT PASSED THE NEXT DAY. Manufacturer Narrative: B2 DATE OF DEATH, B5 DESCRIBE EVENT OR PROBLEM, B7 OTHER RELEVANT HISTORY: UPDATED WITH ADDITIONAL INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21163714&pc=NRY | 2029214-2025-00148 | 2023-12-06 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-15 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Perforation of Vessels; Neurogenic Shock; Multiple Organ Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE SAFETY AND EFFICACY OF ENDOVASCULAR COILS AND NON-FLOW-DIVERTING STENTS FOR MANAGEMENT OF UNRU PTURED INTRACRANIAL ANEURYSMS (UIAS). THE TIME FRAME OF THIS STUDY WAS BETWEEN 2012 AND 2023. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: SOLITAIRE STENT. A DEATH OCCURRED IN THE STUDY POPULATION. THE CAUSE OF DEATH WAS REFRACTORY NEUROGENIC SHOCK. AMONG PATIENT ADVERSE EVENTS INCLUDED THROMBOEMBOLIC COMPLICATIONS, HEMORRHAGIC COMPLICATIONS, ISCHEMIC STROKE, TRANSIENT ISCHEMIC ATTACK, AND VARIOUS COMPLICATIONS RELATED TO SPECIFIC ANATOMICAL LOCATIONS OF THE UNRUPTURED INTRACRANIAL ANEURYSMS. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. THE STUDY SHOWS THAT ENDOVASCULAR TREATMENT WITH STENTS AND COILS IS EFFECTIVE AND SAFE FOR MANAGING UIAS IN VARIOUS ANATOMICAL LOCATIONS. THE INCIDENCE OF THROMBOEMBOLIC COMPLICATIONS WAS SIGNIFICANTLY HIGHER FOR UIAS LOCATED IN THE ANTERIOR CEREBRAL ARTERY. DURING THE RELEASE PROCEDURE USING THE SOLITAIRE STENT, A RUPTURE OCCURRED IN THE DISTAL SEGMENT OF THE RIGHT MCA, RESULTING IN SUBA RACHNOID HEMORRHAGE (HUNT-HESS CLASS II). TWO PATIENTS EXPERIENCED BOTH THROMBOEMBOLIC AND HEMORRHAGIC COMPLICATIONS. THE FIRST PATIENT HAD A LESION AT THE BIFURCATION OF THE RIGHT MCA. DURING THE RELEASE PROCEDURE USING THE SOLITAIRE STENT, A RUPTURE OCCURRED IN THE DISTAL SEGMENT OF THE RIGHT MCA, RESULTING IN SUBARACHNOID HEMORRHAGE (HUNT-HESS CLASS II). 5 DAYS AFTER ANEURYSM EMBOLIZATION, THE PATIENT DEVELOPED REFRACTORY NEUROGENIC SHOCK, ACCOMPANIED BY MULTIPLE ORGAN DYSFUNCTION, AND SUBSEQUENTLY DIED. THE SECOND PATIENT PRESENTED WITH ACUTE THROMBOSIS OF THE TEMPORAL BRANCH OF THE MIDDLE CEREBRAL ARTERY AS A RESULT OF SPASMS CAUSED BY POSTTREATMENT SUBARACHNOID HEMORRHAGE. THIS REQUIRED MECHANICAL THROMBECTOMY WITH ASPIRATION AND CHEMICAL THROMBOLYSIS WITH TIROFIBAN, ACHIEVING A TICI (THROMBOLYSIS IN CEREBRAL INFARCTION) CLASS III OUTCOME. SUBSEQUENTLY, THE PATIENT¿S CONDITION IMPROVED AND THEY RECOVERED TEN PATIENTS EXPERIENCED THROMBOEMBOLIC COMPLICATIONS. AMONG THE PATIENTS WITH A UIA IN THE MCA, THE FIRST PATIENT DEVELOPED INTRASTENT THROMBOSIS, WHICH WAS SUCCESSFULLY TREATED WITH TIROFIBAN. THIS INTERVENTION RESULTED IN COMPLETE RESOLUTION OF THE THROMBOSIS AND NOTABLE CLINICAL IMPROVEMENT. THE SECOND PATIENT EXPERIENCED TRANSIENT DYSARTHRIA AND RIGHT UPPER LIMB WEAKNESS, WHICH RESOLVED SPONTANEOUSLY AND WAS DIAGNOSED AS A TRANSIENT ISCHEMIC ATTACK. THE PATIENT RECEIVED ANTIPLATELET THERAPY (TICAGRELOR 90 MG), WHICH LED TO SYMPTOM RESOLUTION THE FOLLOWING DAY. THE LAST CASES INVOLVED PATIENTS WITH MULTIPLE COMPLICATIONS. AMONG THE FOUR PATIENTS WITH ACA ANEURYSMS, THE FIRST PATIENT PRESENTED WITH OCCLUSION OF THE LEFT PERICALLOSAL ARTERY SECONDARY TO STENT THROMBOSIS. THE PATIENT EXHIBITED NO NEUROLOGICAL DEFICITS AND DID NOT REQUIRE FURTHER INTERVENTION, WHICH WAS ATTRIBUTED TO THE PRESENCE OF AN ADEQUATE NETWORK OF LEPTOMENINGEAL COLLATERAL VESSELS (MRS SCORE AT 12-MONTH FOLLOW-UP, 0). THE SECOND PATIENT EXHIBITED COMPLETE OCCLUSION OF THE ACA AND HEUBNER¿S RECURRENT ARTERY AS A RESULT OF THROMBUS MIGRATION. INTRAVENOUS AND INTRA-ARTERIAL INFUSION OF TIROFIBAN WAS INITIATED, LEADING TO SUBSEQUENT RECANALIZATION (MRS SCORE AT 12- MONTH FOLLOW-UP, 0). THE THIRD PATIENT PRESENTED WITH AN ISCHEMIC STROKE 30 MINUTES AFTER THE PROCEDURE, CHARACTERIZED BY PROGRESSIVE DETERIORATION OF LANGUAGE AND MOVEMENT ON THE RIGHT SIDE OF THE BODY. CEREBRAL ANGIOGRAPHY WAS PERFORMED, REVEALING OCCLUSION OF THE LEFT ACA DUE TO STENT THROMBOSIS. THROMBOLYSIS WAS PERFORMED USING TIROFIBAN, RESULTING IN SUCCESSFUL RECANALIZATION (MRS SCORE AT 12- MONTH FOLLOW-UP, 0). THE FOURTH PATIENT HAD SLIGHT MOTOR DEFICITS IN THE LEFT UPPER AND LOWER LIMBS. BRAIN MAGNETIC RESONANCE IMAGING SHOWED A FILLING DEFECT AT THE ORIGIN OF M1, WITHOUT EVIDENCE OF ISCHEMIC OR HEMORRHAGIC AREAS. PROGRESSIVE RESOLUTION OF THE DISABILITY OCCURRED WITH PHYSIOTHERAPY SUPPORT (MRS SCORE AT 12-MONTH FOLLOW-UP, 0). FINALLY, 1 PATIENT WITH AN ACOMA UIA DEVELOPED PERIOPERATIVE STENT THROMBOSIS, WHICH WAS SUCCESSFULLY RESOLVED USING TIROFIBAN (MRS SCORE AT 12-MONTH FOLLOW-UP, 0). THE SECOND PATIENT PRESENTED WITH STENT THROMBOSIS RESULTING IN OCCLUSION OF THE RIGHT ICA. THE CONDITION WAS SUCCESSFULLY MANAGED USING TIROFIBAN, INTRA-ARTERIAL RECOMBINANT TISSUE PLASMINOGEN ACTIVATOR, AND MECHANICAL THROMBECTOMY, RESULTING IN A TICI 2B OUTCOME. SEVEN PATIENTS PRESENTED PERIOPERATIVE HEMORRHAGIC COMPLICATIONS, 4 CASES IN THE MCA (7.02%), 2 IN THE ACA (10%), AND 1 CASE IN THE ACOMA (3.45%). THE FIRST PATIENT, WITH AN UIA IN THE MCA, EXPERIENCED SELF-LIMITING HEMATEMESIS, LIKELY ATTRIBUTABLE TO ANTIPLATELET THERAPY. THE SECOND PATIENT PRESENTED WITH A RUPTURED ANEURYSM DOME, WHICH WAS SUCCESSFULLY MANAGED THROUGH COIL EMBOLIZATION. TWO COMPLICATIONS WERE OBSERVED IN PATIENTS WITH A UIA IN THE ACA. THE FIRST PATIENT EXPERIENCED SELF-LIMITING MILD EPISTAXIS. THE SECOND PATIENT DEVELOPED A LEFT FRONTOPARIETAL INTRAPARENCHYMAL HEMORRHAGE, RESULTING IN MASS EFFECT AND SUBFALCINE HERNIATION, 1 DAY AFTER THE PROCEDURE. SURGICAL DRAINAGE WAS PERFORMED, LEADING TO SUBSEQUENT RECOVERY. THIS COMPLICATION WAS LIKELY ASSOCIATED WITH ANTIPLATELET THERAPY. FURTHERMORE, 1 PATIENT WITH A UIA IN THE ACOMA DEVELOPED A HEMATOMA IN THE ABDOMINAL WALL THAT EXTENDED TO THE PERITONEAL CAVITY, LIKELY ATTRIBUTABLE TO ANTIPLATELET THERAPY. THREE CASES OF HEMORRHAGIC COMPLICATIONS WERE RECORDED AT 12 MONTHS. ONE PATIENT WITH AN UIA IN THE OPHTHALMIC SEGMENT OF THE ICA PRESENTED WITH MULTIPLE SPONTANEOUS ECCHYMOSES ON THE TRUNK AND EXTREMITIES, LIKELY ATTRIBUTABLE TO A HYPERRESPONSE TO CLOPIDOGREL (VERIFYNOW PRETREATMENT TEST RESULT, 13). THE DOSAGE OF CLOPIDOGREL WAS ADJUSTED, RESULTING IN THE SUBSEQUENT DISAPPEARANCE OF SYMPTOMS. ANOTHER PATIENT WITH A UIA IN THE MCA EXPERIENCED HEMATOCHEZIA LASTING FOR 8 DAYS, ACCOMPANIED BY SEVERE ANEMIA (HEMOGLOBIN, 6.2 G/DL). THIS CONDITION REQUIRED TRANSFUSION OF 2 U OF RED BLOOD CELLS, LEADING TO RESOLUTION OF THE SYMPTOMS. A THIRD PATIENT WITH UIAS IN THE ACOMA EXHIBITED FECAL OCCULT BLOOD WITHOUT HEMODYNAMIC CHANGES. NONE OF THE PATIENTS EXPERIENCED THROMBOEMBOLIC COMPLICATIONS DURING THE 12-MONTH FOLLOW-UP PERIOD. THE INCIDENCE OF PERIOPERATIVE THROMBOEMBOLIC COMPLICATIONS WAS HIGHER IN PATIENTS WITH UIAS IN THE ACA (20%) COMPARED WITH OTHER LOCATIONS (5.51%). SIMILARLY, BLEEDING COMPLICATIONS DURING THE PERIOPERATIVE PERIOD WERE OBSERVED IN 15% OF PATIENTS WITH ACA ANEURYSMS COMPARED WITH ONLY 11% OF PATIENTS WITH UIAS IN OTHER LOCATIONS. DURING THE 12-MONTH FOLLOW-UP PERIOD, NO COMPLICATIONS WERE OBSERVED IN PATIENTS WITH ACA ANEURYSMS. IN CONTRAST, BLEEDING COMPLICATIONS WERE OBSERVED IN ONLY 3 PATIENTS (2.52%) WITH UIAS IN OTHER LOCATIONS THE MRS SCORES INDICATE THAT PATIENTS WITH UIAS IN DIFFERENT LOCATIONS HAVE VARIOUS DEGREES OF DISABILITY AFTER THE PROCEDURE. AN MRS SCORE OF 0-1, WHICH INDICATES NO OR MINIMAL DISABILITY, WAS MOST COMMON FOR THOSE WITH UIAS LOCATED IN THE COMMUNICATING, CHOROIDAL, AND OPHTHALMIC SEGMENTS OF THE ICA AND IN THE BASILAR ARTERY, WITH A RATE OF 85%E100%. FEWER PATIENTS WITH UIAS IN THE PCA HAD AN MRS SCORE OF 0-1 AT 50% COMPARED WITH 85% OF THOSE WITH UIAS IN THE ACA, 80.71% IN THE MCA, AND 85.7% IN THE T CAROTID ARTERY. AN MRS SCORE OF 2-3, INDICATING MODERATE DISABILITY, WAS PRESENT IN A SMALL PERCENTAGE OF PATIENTS, WITH A RATE OF 3.45% IN THE ACOMA AND 5.26% IN THE MCA. AN MRS SCORE OF 4E5, INDICATING SEVERE DISABILITY, WAS MOST COMMON IN PATIENTS WITH UIAS LOCATED IN THE PCA, WITH A RATE OF 50%. ONLY 1 PATIENT (1.75%) HAD AN MRS SCORE OF 6, INDICATING DEATH, AND THIS PATIENT HAD A UIA IN THE MCA, WHICH WAS ALSO DESCRIBED IN THE COMPLICATIONS SECTION. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-SOLITAIRE (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A. G2: CITATION: AUTHORS: VERA D, CARDENAS S, ORTIZ A, RODRIGUEZ A, FERREIRA C, SERRANO S, REYES A, GALVIS M, VARGAS O, MANTILLA D. . SAFETY AND EFFICACY OF ENDOVASCULAR COILS AND NONEFLOW-DIVERTING STENTS FOR MANAGEMENT OF UNRUPTURED INTRACRANIAL ANEURYSMS: A LOCATION-SPECIFIC OUTCOMES ANALYSIS. WORLD NEUROSURGERY. 182:E734-E741. 2024. . HTTPS://DOI.ORG/10.1016/J.WNEU.2023.12 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21163859&pc=NPW | 2015691-2025-00348 | 2024-12-16 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-15 05:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Patient Device Interaction Problem | Heart Block | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN GERMANY, A PATIENT RECEIVED AN EVOQUE VALVE IN TRICUSPID POSITION WHERE ON POST-OPERATIVE DAY 4, THE PATIENT PASSED AWAY. AS REPORTED, CAUSE OF DEATH WAS ATTRIBUTED TO AN ATRIOVENTRICULAR BLOCK GRADE III BASED ON TELEMETRY DATA. THE VALVE HAD BEEN DEPLOYED ON ANNULUS HEIGHT AND THE ANCHORS WERE AT THE LEAFLET HINGE POINTS. THE PATIENT DID NOT HAVE ANY PRE-EXISTING CONDUCTION DISTURBANCES.Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN GERMANY, A PATIENT RECEIVED AN EVOQUE VALVE IN TRICUSPID POSITION WHERE ON POST-OPERATIVE DAY 4, THE PATIENT PASSED AWAY. AS REPORTED, CAUSE OF DEATH WAS ATTRIBUTED TO AN ATRIOVENTRICULAR BLOCK GRADE III BASED ON TELEMETRY DATA. ALTHOUGH THE PATIENT WAS RESUSCITATED WITH A RETURN OF CIRCULATION WITHIN 2 MINUTES, THEY DID NOT SUSTAIN A CARDIAC OUTPUT AND DEVELOPED PULSELESS ELECTRICAL ACTIVITY, ULTIMATELY LEADING TO DEATH. THE PATIENT WAS IN ATRIAL FIBRILLATION (AF) WITH NO ATRIOVENTRICULAR ALTERATION POST-PROCEDURE. THE VALVE HAD BEEN DEPLOYED ON ANNULUS HEIGHT AND THE ANCHORS WERE AT THE LEAFLET HINGE POINTS. THE PATIENT DID NOT HAVE ANY PRE-EXISTING CONDUCTION DISTURBANCES. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, B5, G3, G6, H2, H6 AND H11. THE COMPLAINT FOR VALVE INTERACTS WITH CONDUCTION SYSTEM WAS CONFIRMED WITH OTHER EMPIRICAL EVIDENCE VIA INFORMATION PROVIDED BY THE CLINICAL SPECIALIST. FROM INFORMATION RECEIVED, THE PATIENT WAS ON TELEMETRY FOR 4 DAYS. WHILE THE PATIENT WAS DISCOVERED TO HAVE THIRD-DEGREE AV BLOCK ON POST-OP DAY 4, IT IS UNCLEAR WHEN THE AV BLOCK HAD PRECISELY OCCURRED. FURTHERMORE, RESUSCITATION WAS ATTEMPTED, AND THE PATIENT WAS RETURNED TO CIRCULATION WITHIN 2 (TWO) MINUTES; HOWEVER, THEY DID NOT SUSTAIN A CARDIAC OUTPUT AND DEVELOPED PULSELESS ELECTRICAL ACTIVITY, ULTIMATELY LEADING TO DEATH. THE PATIENT WAS IN ATRIAL FIBRILLATION (AF) WITH NO ATRIOVENTRICULAR ALTERATION POST-PROCEDURE. THE REPORTED CONDUCTION DISTURBANCE DOES NOT ALLEGE A MALFUNCTION THAT COULD BE RELATED TO AN EDWARDS MANUFACTURING DEFICIENCY. THE EVENT DOES NOT ALLEGE A LABELING ISSUE OR DEVICE RELATED INFECTION; THEREFORE, NO LOT HISTORY REVIEW, OR MANUFACTURING MITIGATIONS REVIEW ARE REQUIRED. SINCE NO MANUFACTURING NON-CONFORMANCES WERE FOUND AND NO LABELING, TRAINING, OR IFU DEFICIENCIES WERE IDENTIFIED, NO PRA NOR PREVENTATIVE OR CORRECTIVE ACTIONS WERE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164169&pc=BZD | 2518422-2025-100689 | 2022-10-14 04:00 | Death | RESPIRONICS, INC. | 2025-01-15 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. IT IS UNKNOWN IF THE PATIENT WAS USING THE DEVICE AT THE TIME OF DEATH. THE DEVICE WAS THROWN OUT BY THE PATIENT'S WIFE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164443&pc=DSQ | 2916596-2025-00085 | 2024-12-28 05:00 | Death | THORATEC CORPORATION | 2025-01-15 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ascites; Heart Failure/Congestive Heart Failure; Intra-Abdominal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON COMFORT CARE RECEIVING TREATMENT FOR RIGHT VENTRICULAR FAILURE, ACUTE KIDNEY INJURY AND HEMOPERITONEUM. THE PATIENT PASSED AWAY ON (B)(6) 2024.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS ADMITTED WITH VOLUME OVERLOAD AND ACUTE KIDNEY INJURY THOUGHT TO BE CAUSED BY RIGHT VENTRICULAR FAILURE. THE RIGHT VENTRICULAR FAILURE WAS NOTED TO HAVE DEVELOPED OVER THE COURSE OF THE PATIENT'S TIME WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE PATIENT HAD PARACENTESIS FOR ASCITES WHICH LED TO HEMOPERITONEUM. THE FAMILY THEN DECIDED TO PURSUE COMFORT MEASURES. THE DEATH WAS NOT CONSIDERED DEVICE RELATED, THE DEVICE OPERATED AS EXPECTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU). SECTION 1, ¿INTRODUCTION¿, LISTS BLEEDING, RIGHT HEART FAILURE, RENAL DYSFUNCTION, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING THE INTERNATIONAL NORMALIZED RATIO RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6: KIDNEY INJURY E1017 NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21146949&pc=LTI | 3012638928-2024-03250 | 2024-12-12 05:00 | Death | SPATZ FGIA INC. | 2025-01-14 05:00 | LTI | SPATZ3 ADJUSTABLE BALLOON SYSTEM | Appropriate Term/Code Not Available | Abdominal Pain; Hemorrhage/Bleeding; Pneumonia | P190012 |  | 1 | Event Description: ON (B)(6) 2024, SPATZ3 INTRAGASTRIC BALLOON WAS IMPLANTED WITH 400ML VOLUME- NO COMPLICATIONS OR ASPIRATION RECORDED. THE PATIENT DID NOT INFORM THE DOCTOR THAT HE/SHE WAS UNDER TREATMENT WITH ANTIBIOTIC FOR PNEUMONIA AT HOME PRIOR TO THE PROCEDURE. ON (B)(6) 2024, POSTOPERATIVE DAY 1: THE PATIENT EXPERIENCED SEVERE PAIN RADIATING TO THE BACK AND RETURNED TO THE HOSPITAL. IN THE HOSPITAL, THE PATIENT FOR THE FIRST TIME NOTIFIED THE DOCTORS THAT SHE WAS BEING TREATED FOR PNEUMONIA. IN THE HOSPITAL THE PNEUMONIA PROGRESSED TO SEPTIC SHOCK, ARDS, AND MODS, WHICH SUBSEQUENTLY RESOLVED. THE PATIENT WAS KEPT IN THE HOSPITAL FOR OBSERVATION, DURING WHICH A CT SCAN WAS CONDUCTED, CONFIRMING PNEUMONIA. THE PATIENT'S CONDITION AND VENTILATION SUBSEQUENTLY DETERIORATED, AND AN EMERGENCY THORACENTESIS WAS PERFORMED. AN EMPYEMA WAS DRAINED, FOLLOWED BY 5,000 ML OF BLOOD, WHICH LED TO HEMORRHAGIC SHOCK. A THORACOTOMY WAS PERFORMED IN THE OPERATING ROOM, DURING WHICH BLEEDING WAS CONTROLLED. HOWEVER, THE PATIENT DEVELOPED REFRACTORY SHOCK, SEVERE ACIDOSIS, AND COAGULOPATHY, WITH SUBSEQUENT DEATH. ON (B)(6) 2024, THE PATIENT PASSED AWAY AT 9:25 A.M. DESPITE ALL MEASURES TAKEN. THE PHYSICIAN WAS IN CONTACT WITH THE HUSBAND AND SON IN THE ICU. Manufacturer Narrative: THE PNEUMONIA AND ONGOING ANTIBIOTIC TREATMENT WAS NOT DISCLOSED TO THE DOCTOR. THE PHYSICIAN WAS NOT AWARE OF THE PATIENT'S ACTIVE INFECTION. ON (B)(6) 2024, 1 DAY AFTER DEVICE IMPLANTATION THE PATIENT COMPLAINED OF SEVERE PAIN RADIATING TO HER BACK. THE PATIENT RETURNED TO THE HOSPITAL AND THE BALLOON WAS DOWN ADJUSTED TO 300CC AS THESE ARE THE COMMON SYMPTOMS OF INTOLERANCE. ON (B)(6) 2024, THE PATIENT PASSED AWAY. (MISTAKENLY INITIALLY DEATH DATE REPORTED ON (B)(6) 2024). THE CONDITIONS THAT LED TO DEATH (PNEUMONIA, SEPTIC SHOCK, HEMOTHORAX, HEMORRHAGIC SHOCK), THE ENDOSCOPY AND SEDATION ITSELF, WITH OR WITHOUT BALLOON IMPLANTATION, CAN BE DIRECTLY LINKED TO THE EXACERBATION OF PNEUMONIA OR SUBSEQUENT COMPLICATIONS WHICH CAUSED THE PNEUMONIA TO PROGRESS. LABELLING REVIEW: THE INSTRUCTIONS FOR USE (IFU) FOR THE SPATZ3 DEVICE WERE REVIEWED, WITH PARTICULAR FOCUS ON CONTRAINDICATIONS. THE IFU CONTRAINDICATION LIST INCLUDES: "TREATMENT REPRESENTS AN UNREASONABLE RISK TO THE PATIENT" AND ALSO "ANY OTHER MEDICAL CONDITION THAT MAY INCREASE THE RISK OF ELECTIVE ENDOSCOPY OR WEIGHT LOSS." THESE COVER ANY ACTIVE MEDICAL ISSUE AS A CONTRAINDICATION TO AN ENDOSCOPY UNDER SEDATION, WITH OR WITHOUT A BALLOON PROCEDURE. CLINICAL PRACTICE DICTATES DELAYING ELECTIVE INTERVENTIONS THAT REQUIRE SEDATION SUCH AS ENDOSCOPY, COLONOSCOPY, TEE (TRANSESOPHAGEAL ECHOCARDIOGRAM) AND OTHERS IF THE CLINICAL STATUS PUTS THE PATIENT AT RISK FOR UNDERGOING A PROCEDURE THAT REQUIRES SEDATION, UNTIL FULL RECOVERY IS ACHIEVED. FOR INSTANCE, PATIENTS PRESENTING WITH ABNORMAL VITAL SIGNS (E.G., A PULSE OF 35 OR 150 BPM) ARE NOT CANDIDATES FOR ENDOSCOPIC PROCEDURES UNTIL THEIR CONDITION STABILIZES, BUT WE CANNOT STIPULATE THE INNUMERABLE INSTANCES. SO, WE PLACE IT UNDER THE TERMS OF "TREATMENT REPRESENTS AN UNREASONABLE RISK TO THE PATIENT" AND ALSO "ANY OTHER MEDICAL CONDITION THAT MAY INCREASE THE RISK OF ELECTIVE ENDOSCOPY OR WEIGHT LOSS" WHICH DEPEND ON THE DOCTOR TO USE CLINICAL JUDGEMENT. FURTHERMORE, THE PATIENT FAILED TO DISCLOSE CRITICAL MEDICAL INFORMATION REGARDING THEIR ONGOING TREATMENT FOR PNEUMONIA. THIS OMISSION DIRECTLY INFLUENCED THE PHYSICIAN'S ABILITY TO MAKE AN INFORMED DECISION REGARDING THE TIMING OF THE PROCEDURE. ENSURING THAT PATIENTS PROVIDE A COMPLETE AND ACCURATE MEDICAL HISTORY IS ESSENTIAL TO SAFEGUARDING THE APPROPRIATENESS OF ANY ELECTIVE MEDICAL INTERVENTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147140&pc=MKJ | 3004123209-2025-00013 | 2024-12-14 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-01-14 05:00 | MKJ | PACKAGE,360P,PP03,INTL EN,360-BAS-UK-10 | Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A PATIENT INVOLVED EVENT. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED.Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A PATIENT INVOLVED EVENT. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: HEARTSINE'S INVESTIGATION WAS UNABLE TO FIND FAULT WITH THE RETURNED DEVICE. THE INVESTIGATION CAN ONLY SUGGEST THAT THE REPORTED COMPLAINT WAS DUE TO SITUATIONAL FACTORS. E.G., A PAD-PAK FITMENT ISSUE OR INCORRECT PLACEMENT OF PADS, AS NO FAULT WAS FOUND WITH THE RETURNED SAM 350P OR RETURNED PAD-PAK. A REVIEW OF DEVICE MEMORY SHOWED THE DEVICE WAS POWERED ON MULTIPLE TIMES ON THE DATE OF THE REPORTED EVENT. DURING EACH POWER ON, THE DEVICE ISSUED THE NORMAL ADVISORY SPEECH PROMPTS TO APPLY THE PADS TO PATIENT. THE DEVICE CONTINUED TO LOG ¿APPLY PADS¿ PROMPTS THROUGHOUT THE EVENT, AS AN IN-RANGE PATIENT IMPEDANCE WAS NOT DETECTED. UPON RECEIPT, THE DEVICE UNDERWENT EXTENSIVE TESTING OF ALL CRITICAL FUNCTIONS, PERFORMING TO SPECIFICATION THROUGHOUT. THE DEVICE WAS FOUND TO ACCURATELY MEASURE IMPEDANCE THROUGHOUT THE SPECIFIED RANGE, EVEN AFTER THE STRESS OF ELEVATED TEMPERATURE. THE DEVICE DELIVERED THE FULL SHOCK THERAPY SEQUENCE WITHOUT FAULT USING THE TEST POWER SUPPLY AND THE DEVICE RECORDED ECG DATA ACCURATELY WITHOUT NOISE OR OTHER ABNORMALITIES. THE DEVICE HAS BEEN RETAINED BY HEARTSINE AND THE CUSTOMER WAS PROVIDED WITH A REPLACEMENT DEVICE. Manufacturer Narrative: HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. HEARTSINE REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED HEARTSINE WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147255&pc=OCE | 3002836339-2024-00138 | 2024-12-06 05:00 | Death | VERICEL | 2025-01-14 05:00 | OCE | EPICEL | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Coagulation Disorder; Hemorrhage/Bleeding; Sepsis | 990200 |  | 1 | Event Description: PULMONARY HEMORRHAGE [PULMONARY HAEMORRHAGE] MEDICATION REFRACTORY ATRIAL FIBRILLATION [ATRIAL FIBRILLATION] TRANSFUSION-ASSOCIATED RELATED THROMBOCYTOPENIA [THROMBOCYTOPENIA] SEPSIS [SEPSIS]. Manufacturer Narrative: CASE REFERENCE NUMBER (B)(4) IS A SPONTANEOUS REPORT *INITIALLY* RECEIVED FROM A NON-HEALTHCARE PROFESSIONAL ON (B)(6) 2024, CONCERNING A 32-YEAR-OLD MALE PATIENT WHO EXPIRED AFTER GRAFTING WITH EPICEL. *ON (B)(6) 2024, QC LOT RELEASE ASSAYS (GRAFT INSPECTION QC2-027, DUAL STAIN QC2-094, ENDOTOXIN QC2-010, STERILITY PRE-RELEASE QC2-005, STERILITY FINAL PRODUCT QC2-012) AND ENVIRONMENTAL RESULTS (MANUFACTURING: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT) AND GRADE B (VIABLE AIR PARTICULATES AND NONVIABLE AIR PARTICULATES) AND QC STERILITY: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT)) WERE PERFORMED. ALL RESULTS WERE REPORTED AS ''PASS''. QC LOT RELEASE ASSAY (IMPLEMENTED 29-SEP-2019) 3T3 RESIDUAL ASSAY Q2C-136 WAS REPORTED AS LESS THAN 0.1 - PASS. * THE PATIENT'S MEDICAL HISTORY INCLUDED VON WILLEBRAND'S DISEASE. THE PATIENT'S CONCOMITANT MEDICATION DETAILS WERE NOT REPORTED. ON (B)(6) 2024, THE PATIENT RECEIVED EPICEL GRAFTS (CULTURED EPIDERMAL AUTOGRAFTS) (LOT NUMBER: EE03344, EXPIRATION DATE: 14-NOV-2024) FOR BURN OVER 30% TOTAL BODY SURFACE AREA (TBSA). ON (B)(6) 2024, GRAFT TAKEDOWN WAS PERFORMED AND 98 PERCENT (%) OF GRAFT TAKE WAS ACHIEVED ON BACK AND RIGHT LOWER ANTERIOR THIGH. SURGICAL PROCEDURE WAS TOLERATED WELL BY PATIENT WITH VITAL SIGNS WITHIN NORMAL LIMITS. ON (B)(6) 2024, PATIENT WAS TAKEN TO OPERATING ROOM (OR) FOR CHEST DEBRIDEMENT. ON (B)(6) 2024, THE PATIENT DEVELOPED *TRANSFUSION-ASSOCIATED RELATED* THROMBOCYTOPENIA (PT: THROMBOCYTOPENIA) AND SEPSIS (PT: SEPSIS), PER BURN TEAM. THE PLAN WAS TO GET PLATELET COUNT UP BEFORE THE NEXT GRAFT DATE. HOWEVER, THE PATIENT'S PLATELET COUNT CONTINUED TO STAY LOW, EVEN WITH MULTI BLOOD PRODUCTS ADMINISTERED. ON (B)(6) 2024, THE PATIENT WAS PLACED ON COMFORT CARE. *ON AN UNKNOWN DATE, THE PATIENT EXPERIENCED PULMONARY HEMORRHAGE (PT: PULMONARY HEMORRHAGE) SECONDARY TO MULTIPLE SYNCHRONIZED CARDIOVERSIONS FOR MEDICATION REFRACTORY ATRIAL FIBRILLATION (PT: ATRIAL FIBRILLATION) IN THE SETTING OF TRANSFUSION-ASSOCIATED RELATED THROMBOCYTOPENIA.* ON (B)(6) 2024, PATIENT PASSED AWAY *DUE TO PULMONARY HEMORRHAGE*. ON (B)(6) 2024, HOSPITAL TEAM WAS NOTIFIED. THE EVENTS OF *TRANSFUSION-ASSOCIATED RELATED* THROMBOCYTOPENIA AND SEPSIS WERE ASSESSED AS SERIOUS DUE TO HOSPITALIZATION AND MEDICAL SIGNIFICANCE. *THE EVENTS OF PULMONARY HEMORRHAGE AND MEDICATION REFRACTORY ATRIAL FIBRILLATION WERE ASSESSED AS SERIOUS DUE TO MEDICAL SIGNIFICANCE AND FATAL OUTCOME. * THE REPORTER DID NOT PROVIDE A CAUSALITY ASSESSMENT FOR THE EVENTS. NO FURTHER INFORMATION WAS PROVIDED. ALL FOLLOW-UP INFORMATION IS INCLUDED IN THE CASE NARRATIVE ABOVE, WITH THE LATEST INFORMATION PRESENTED BETWEEN ASTERISKS (*). FOLLOW-UP INFORMATION WAS RECEIVED FROM THE REPORTER ON (B)(6) 2024 AND INCLUDED QC RESULTS. ADDITIONAL INFORMATION WAS RECEIVED ON 07-JAN-2025 AND PROCESSED TOGETHER WITH THE FOLLOW-UP. COMPANY COMMENT: VERICEL HAS ASSESSED THE EVENT OF PULMONARY HEMORRHAGE AS SERIOUS, NOT RELATED AND UNEXPECTED, THE EVENT OF ATRIAL FIBRILLATION AS SERIOUS, NOT RELATED AND UNEXPECTED, THE EVENT OF THROMBOCYTOPENIA AS SERIOUS, NOT RELATED AND EXPECTED, AND THE EVENT OF SEPSIS AS SERIOUS, POSSIBLY RELATED AND EXPECTED PER EPICEL IFU. THE CASE DOES NOT QUALIFY AS AN MDR AND QUALIFIES AS A 15-DAY-REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147413&pc=CAF | 2518422-2025-002295 | 2023-08-15 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | CAF | DELUXE 230V 50HZ INT. UK PLUG | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K170853 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DELUXE 230V 50HZ INT. UK PLUG THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147414&pc=BZD | 2518422-2025-002290 | 2024-06-06 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | DREAMSTATION BIPAP PRO | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES NO VISIBLE FOAM DEGRADATION. EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, AND CONCLUSION CODE GRID.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147465&pc=CAF | 2518422-2025-002297 | 2023-07-24 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | CAF | INNOSPIRE DELUXE RENTAL UNIT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K170853 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF INNOSPIRE DELUXE RENTAL UNIT THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147516&pc=BZD | 2518422-2025-002346 | 2024-04-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WAS UPSET THAT INSTEAD OF A RESOLUTION THAT SHE WAS ONLY GIVEN $100 FOR THE DEVICES SHE HAD PUT INTO THE ECONOMIC LOSS SETTLEMENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147752&pc=OZD | 1220648-2025-25542 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION PRESENTED TO THE CATHETERIZATION LAB FOR IMPELLA CP IMPLANT AND UPON INSERTION THE PATIENT WENT INTO SUSTAINED VENTRICULAR FIBRILLATION WITH A DEMISED OUTCOME. LEFT HEART CATHETERIZATION WAS COMPLETED AND SHOWED CLOT FROM LEFT MAIN (LM) THROUGH THE ENTIRE LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH A HEAVILY CALCIFIED RIGHT CORONARY ARTERY AND PULMONARY ARTERIAL END-DIASTOLIC PRESSURE OF 71. FOR THE IMPELLA CP IMPLANT, THE GROIN SIZED UP TO 14FX13CM PEEL AWAY. PIGTAIL AND J WIRE WERE USED TO GAIN LEFT VENTRICLE BEFORE ABIOMED .018 WIRE PLACED AND IMPELLA INSERTED WITH FLOWS RANGING BETWEEN .5-1.5 IN AUTO. UPON INSERTION, THE PATIENT WENT INTO SUSTAINED VENTRICULAR FIBRILLATION, WAS SHOCKED AND RECEIVED CARDIOPULMONARY RESUSCITATION. AT THIS POINT, THE DECISION WAS MADE TO PLACE AN IMPELLA RP. AFTER THE IMPELLA RP HAD BEEN OPENED AND PREPPED, THE PHYSICIAN OPTED TO PLACE THE PATIENT ON VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA-ECMO) DUE TO CARDIAC STANDSTILL AND OXYGENATION REQUIREMENTS. MINIMAL FLOWS WERE ABLE TO BE ACHIEVED AFTER PLACEMENT AND AT THIS POINT THE PATIENT HAD EXHIBITED SIGNS OF PULMONARY HEMORRHAGE. SINGLE ACCESS WAS PERFORMED AFTER ECMO PLACEMENT WHERE PHYSICIAN HAD ATTEMPTED TO USE THE PENUMBRA AND SERIES OF BALLOONS ON LM-LAD BUT WAS UNSUCCESSFUL. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT IN CATHETERIZATION LAB; THE TEAM DETERMINED THERE WAS NO OTHER INTERVENTION THAT COULD BE DELIVERED TO THE PATIENT. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED VENTRICULAR FIBRILLATION (VF) HAS BEEN COMPLETED. AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED, THE ROOT CAUSE OF THE VF WAS NOT DETERMINED. B.7 INFORMATION WAS ADDED THAT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25542. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25542 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 SEVERAL HEALTH EFFECT - IMPACT CODES WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25542 AND WERE ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147767&pc=OZD | 1220648-2025-25545 | 2024-12-30 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE AND WAS ESCALATED TO AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. APPROXIMATELY FOUR DAYS AFTER START OF SUPPORT THE PATIENT EXPERIENCED RETROPERITONEAL BLEEDING. THE SOURCE OF THE BLEED COULD NOT BE ASCERTAINED. INFORMATION RECEIVED FROM THE ATTENDING PHYSICIAN INDICATED THAT THERE WAS A POSSIBILITY OF AORTIC WALL DAMAGE WHEN THE IMPELLA CP WAS REMOVED, BUT THE COMPUTED TOMOGRAPHY SCAN COULD NOT DETERMINE EITHER A LOCATION OF DISSECTION OR BLEEDING. THE CAUSAL RELATIONSHIP, HOWEVER, BETWEEN THE BLEEDING AND THE IMPELLA 5.5 IS UNKNOWN. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED VASCULAR DAMAGE HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THE CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND THERE WAS INSUFFICIENT CLINICAL INFORMATION. B.2 REVISED AS SOME SELECTIONS WERE REPORTED INCORRECTLY ON THE INITIAL MANUACTURER DEVICE REPORT NUMBER 1220648-2025-25545. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25545 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4614 WAS REPORTED INCORRECTLY ON THE INITIAL MANUACTURER DEVICE REPORT NUMBER 1220648-2025-25545. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21147913&pc=PNJ | 9612164-2025-00252 | 2024-10-01 04:00 | Death | MEDTRONIC IRELAND | 2025-01-14 05:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING OUTCOMES WITH LEADLESS VERSUS SINGLE-CHAMBER TRANSVENOUS IMPLANTABLE PULSE GENERATORS (IPGS) IN HE MODIALYSIS PATIENTS. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED ACUTE COMPLICATIONS WHICH WERE DEFINED AS LESS THAN NINETY DAYS POST IMPLANT. ACUTE COMPLICATIONS INCLUDED CARDIAC ARREST, HEMOPERICARDIUM, HEMORRHAGE, ACCESS SITE COMPLICATIONS, ENDOCARDITIS/DEVICE-RELATED INFECTIONS, DEEP VEIN THROMBOSIS, OR PULMONARY EMBOLISM. THE STATUS OF THE LEADLESS IPGS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/76 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPROVED OUTCOMES WITH LEADLESS VS. SINGLE-CHAMBER TRANSVENOUS PACEMAKER IN HEMODIALYSIS PATIENTS. EUROPACE. 2024. 26, EUAE257. DOI: 10.1093/EUROPACE/EUAE257. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21148215&pc=NVN | 2649622-2025-01369 | 2024-12-25 05:00 | Death | MPRI | 2025-01-14 05:00 | NVN | SELECTSECURE MRI SURESCAN | Over-Sensing; Under-Sensing | No Clinical Signs, Symptoms or Conditions | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDER SENSING DURING VENTRICULAR ARRHYTHMIA. IT WAS ALSO REPORTED THAT THE RV LEAD EXHIBITED OVER SENSING OF SHORT V-V INTERVALS. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21148376&pc=OZD | 1220648-2025-25552 | 2024-12-16 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED ANEMIA WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿HGB 9.3 GM/DL ON ADMISSION. POST IMPELLA IMPLANT (B)(6) 2024- HGB 4.4. NO CLEAR EVIDENCE OF BLEEDING. TRANSFUSED 3 UNITS PRBCS AND 2 UNITS PLASMA. THE PATIENT EXPIRED.¿ THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21148847&pc=DSQ | 2916596-2025-00227 | 2022-09-01 04:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "LONG TERM USE OF IMPLANTABLE LEFT VENTRICULAR ASSIST DEVICES IN PEDIATRIC PATIENTS" THAT THE HM3 MAY BE RELATED TO DEATH DUE TO MULTIORGAN FAILURE AND DEATH DUE TO UNREPORTED REASONS. A STUDY FROM MAY2008 -SEPT2022 WAS PERFORMED ON PATIENTS THAT WERE <19 YEARS OF AGE LIVING WITH AN LVAD FOR MORE THAN 1 YEAR. THE SUMMARY STATISTICS WERE USED TO DESCRIBE PATIENT DEMOGRAPHICS, LENGTH OF SUPPORT, FREQUENCY OF NON-ELECTIVE READMISSIONS, ADVERSE EVENTS (PER PEDIMACS CRITERIA), AND OUTCOMES. 90 PATIENTS RECEIVED 91 CONTINUOUS FLOW VENTRICULAR ASSIST DEVICE (CFVAD) DURING THE STUDY, OF THESE 29 PATIENTS HAD VAD FOR MORE THAN 1 YEAR. 20 PATIENTS FOR 1-3 YEARS, 5 PATIENTS FOR 3-5 YEARS, 4 PATIENTS FOR >5 YEARS, AND THE LONGEST SUPPORTED PATIENTS FOR 12 YEARS. VAD SUPPORT WAS ON HM2 (N=1), HVAD (N=23), AND HM3 (N=5). THERE WERE 81 READMISSIONS OVER 29,812 VAD SUPPORT DAYS. THE OUTCOMES SHOWED INCLUDED 1 (3%) PATIENT WHO NEVER DISCHARGED AND DIED DUE TO MULTI-ORGAN FAILURE AND 4 PATIENTS (14%, 4/29) WHO DIED ON VAD SUPPORT AFTER DISCHARGE. Manufacturer Narrative: SECTIONS A AND D4: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. TUNUGUNTLA, H., CHO, J., CHOUDHRY, S., SPINNER, J., PURI, K., HOPE, K., WATANABE, K., DREYER, W., PRICE, J., BRODA, C., WILSON, D., ELIAS, B., & ADACHI, I. (2024). LONG TERM USE OF IMPLANTABLE LEFT VENTRICULAR ASSIST DEVICES IN PEDIATRIC PATIENTS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S173. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.350. TEXAS CHILDREN'S HOSPITAL, HOUSTON,TX; BAYLOR COLLEGE OF MEDICINE, HOUSTON,TX; MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH AND MULTIPLE ORGAN FAILURES/DYSFUNCTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149186&pc=NVZ | 2182208-2025-00221 | 2006-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-01-14 05:00 | NVZ | AT500 | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P980035 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG TERM MANAGEMENT OF ATRIAL ARRHYTHMIAS IN YOUNG PATIENTS WITH SICK SINUS SYNDROME UNDERGOING EARLY OPERATION TO CORRECT CONGENITAL HEART DISEASE. THE AUTHORS DESCRIBED THREE PATIENT DEATHS; THE CAUSE OF THE DEATHS WAS CONGESTIVE HEART FAILURE (CHF). THERE WERE PATIENTS WHO WERE HOSPITALIZED DUE TO ATRIAL TACHYARRHYTHMIA (AT) EPISODES WITH SOME DUE TO INAPPROPRIATE DETECTIONS FOR ATRIAL LEAD NOISE AND/OR FAR FIELD R-WAVE (FFRW) OVERSENSING. IN ONE PATIENT, ATRIAL LEAD NOISE INDUCED INAPPROPRIATE AT DETECTION THAT RESULTED IN ANTITACHYCARDIA PACING (ATP) DELIVERY. SEVERAL AT EPISODES WERE NOT TREATED OWING TO ATRIAL UNDERSENSING OR 1:1 ATRIOVENTRICULAR (AV) CONDUCTION. IN TWO PATIENTS, AT EPISODES CONDUCTED ONE-TO-ONE VIA THE AV NODE WERE INAPPROPRIATELY CLASSIFIED AS VENTRICULAR TACHYCARDIA (VT) BY THE DEVICE. NO ADJUSTMENT OF PACING OR SENSING WAS NECESSARY. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/17 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG TERM MANAGEMENT OF ATRIAL ARRHYTHMIAS IN YOUNG PATIENTS WITH SICK SINUS SYNDROME UNDERGOING EARLY OPERATION TO CORRECT CONGENITAL HEART DISEASE. EUROPACE. 2006. 8, 488¿494. DOI: 10.1093/EUROPACE/EUL069. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149307&pc=DTB | 2182208-2025-00222 | 2006-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-01-14 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P930039 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG TERM MANAGEMENT OF ATRIAL ARRHYTHMIAS IN YOUNG PATIENTS WITH SICK SINUS SYNDROME UNDERGOING EARLY OPERATION TO CORRECT CONGENITAL HEART DISEASE. THE AUTHORS DESCRIBED THREE PATIENT DEATHS; THE CAUSE OF THE DEATHS WAS CONGESTIVE HEART FAILURE (CHF). THERE WERE PATIENTS WHO WERE HOSPITALIZED DUE TO ATRIAL TACHYARRHYTHMIA (AT) EPISODES WITH SOME DUE TO INAPPROPRIATE DETECTIONS FOR ATRIAL LEAD NOISE AND/OR FAR FIELD R-WAVE (FFRW) OVERSENSING. IN ONE PATIENT, ATRIAL LEAD NOISE INDUCED INAPPROPRIATE AT DETECTION THAT RESULTED IN ANTITACHYCARDIA PACING (ATP) DELIVERY. SEVERAL AT EPISODES WERE NOT TREATED OWING TO ATRIAL UNDERSENSING OR 1:1 ATRIOVENTRICULAR (AV) CONDUCTION. IN TWO PATIENTS, AT EPISODES CONDUCTED ONE-TO-ONE VIA THE AV NODE WERE INAPPROPRIATELY CLASSIFIED AS VENTRICULAR TACHYCARDIA (VT) BY THE DEVICE. NO ADJUSTMENT OF PACING OR SENSING WAS NECESSARY. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/17 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG TERM MANAGEMENT OF ATRIAL ARRHYTHMIAS IN YOUNG PATIENTS WITH SICK SINUS SYNDROME UNDERGOING EARLY OPERATION TO CORRECT CONGENITAL HEART DISEASE. EUROPACE. 2006. 8, 488¿494. DOI: 10.1093/EUROPACE/EUL069. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149357&pc=BZD | 2518422-2025-100565 | 2024-11-20 05:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE REP DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WAS NO REPORTED MEDICAL INTERVENTION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149370&pc=DTB | 2182208-2025-00224 | 2006-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-01-14 05:00 | DTB | CAPSURE EPICARDIAL LEAD | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P950024 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG TERM MANAGEMENT OF ATRIAL ARRHYTHMIAS IN YOUNG PATIENTS WITH SICK SINUS SYNDROME UNDERGOING EARLY OPERATION TO CORRECT CONGENITAL HEART DISEASE. THE AUTHORS DESCRIBED THREE PATIENT DEATHS; THE CAUSE OF THE DEATHS WAS CONGESTIVE HEART FAILURE (CHF). THERE WERE PATIENTS WHO WERE HOSPITALIZED DUE TO ATRIAL TACHYARRHYTHMIA (AT) EPISODES WITH SOME DUE TO INAPPROPRIATE DETECTIONS FOR ATRIAL LEAD NOISE AND/OR FAR FIELD R-WAVE (FFRW) OVERSENSING. IN ONE PATIENT, ATRIAL LEAD NOISE INDUCED INAPPROPRIATE AT DETECTION THAT RESULTED IN ANTITACHYCARDIA PACING (ATP) DELIVERY. SEVERAL AT EPISODES WERE NOT TREATED OWING TO ATRIAL UNDERSENSING OR 1:1 ATRIOVENTRICULAR (AV) CONDUCTION. IN TWO PATIENTS, AT EPISODES CONDUCTED ONE-TO-ONE VIA THE AV NODE WERE INAPPROPRIATELY CLASSIFIED AS VENTRICULAR TACHYCARDIA (VT) BY THE DEVICE. NO ADJUSTMENT OF PACING OR SENSING WAS NECESSARY. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/17 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG TERM MANAGEMENT OF ATRIAL ARRHYTHMIAS IN YOUNG PATIENTS WITH SICK SINUS SYNDROME UNDERGOING EARLY OPERATION TO CORRECT CONGENITAL HEART DISEASE. EUROPACE. 2006. 8, 488¿494. DOI: 10.1093/EUROPACE/EUL069. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149452&pc=OZD | 1220648-2025-25559 | 2024-12-30 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA 5.5 SUPPORT, THE PURGE PRESSURE SUDDENLY INCREASED WHILE THE PURGE FLOW DECREASED. A HIGH PURGE PRESSURE AND PURGE FLOW BLOCKED ALARM MESSAGE APPEARED ON THE AUTOMATED IMPELLA CONTROLLER IN RESPONSE TO THE SUDDEN CHANGE IN SUPPORT. RESOLUTION OPTIONS WERE DISCUSSED WITH THE SITE AND THE DECISION WAS MADE TO INITIATE A TISSUE PLASMINOGEN ACTIVATOR PROTOCOL. SUPPORT WAS CONTINUED FOLLOWING THE INTERVENTION. UPON FOLLOW UP, IT WAS DOCUMENTED THAT THE PATIENT PASSED AWAY THE FOLLOWING DAY DUE TO THEIR OVERALL CONDITION. IT IS UNKNOWN IF THE IMPELLA RELATED MALFUNCTION CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA 5.5 WAS RETURNED FOR EVALUATION. THE PUMP WAS RETURNED. CUT AT CATHETER AND ONLY PUMP CANNULA AND MOTOR HOUSING SIDE WAS RETURNED. UPON VISUAL INSPECTION, BIOMATERIAL WAS PRESENT IN THE PURGE GAP AFTER PUMP WAS SOAKED IN WARM WATER TO REMOVE DRIED BLOOD. DATA LOGS WERE REVIEWED AND LONG TERM LOG SHOWED AN INCREASE IN PURGE PRESSURE OVER A PERIOD OF APPROXIMATELY FOUR HOURS WITH A SIMULTANEOUS DECREASE IN PURGE FLOW. PURGE PRESSURE REMAINED HIGH FOR THE REMAINDER OF THE CASE. CLINICAL DETAILS NOTED THAT "HIGH PURGE PRESSURE" ALARM OCCURRED ON DAY FIVE OF SUPPORT, THE PUMP HAD DEXTROSE 5 PERCENT IN WATER WITH SODIUM BICARBONATE RUNNING IN THE PURGE SOLUTION. TISSUE PLASMINOGEN ACTIVATOR PROTOCOL WAS INITIATED IN ATTEMPT TO RESOLVE ALARMS. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS DUE TO BIOMATERIAL. B.1 REVISED AS PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25559. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149658&pc=NPT | 2025587-2025-00349 | 2023-04-21 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-14 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Gradient Increase; Patient Device Interaction Problem; Device Stenosis | Dyspnea; Hypoxia; Stenosis; Respiratory Failure; Pericardial Effusion; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE MEAN AORTIC VALVE GRADIENT WAS 20MM HG AND/OR PEAK VELOCITY 3 METERS PER SECOND (M/SEC). HOWEVER, EXACT VALUES WERE NOT PROVIDED. APPROXIMATELY SEVEN YEARS AND FIVE MONTHS FOLLOWING THE VALVE IMPLANT SHORTNESS OF BREATH PRESENTED. AN ECHOCARDIOGRAM IDENTIFIED AN ELEVATED GRADIENT AND VELOCITY. SEVERE MITRAL STENOSIS BY CONTINUITY EQUATION WITH A MITRAL VALVE AREA OF 1.2 CM2. THE MEAN TRANSMITTAL GRADIENT WAS 9MM HG WITH A HEART RATE OF 62 BPM. A SMALL PERICARDIAL EFFUSION WAS NOTED, NO EVIDENCE OF CARDIAC TAMPONADE. THE AORTIC VALVE VELOCITIES AND GRADIENT AND THE MITRAL GRADIENT WERE HIGHER IN COMPARISON TO PREVIOUS ASSESSMENT PERFORMED APPROXIMATELY TWO AND A HAVE YEARS PRIOR. NO TREATMENT REPORTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED THAT CORRECTED THE PREVIOUS DOCUMENT OF MITRAL STENOSIS. THE ECHOCARDIOGRAM WAS PERFORMED AND SHOWED MODERATE MITRAL STENOSIS. NO TREATMENT OR INTERVENTION WAS PERFORMED. THE PATIENT WOULD CONTINUE TO BE MONITORED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: APPROXIMATELY 7 YEARS 6 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, IMAGING WAS PERFORMED FOR A VALVE-IN -VALVE PROCEDURE DUE TO SEVERE AORTIC STENOSIS. NO TREATMENT WAS REPORTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED AT THAT TIME. APPROXIMATELY 8 YEARS AND 1 MONTH FOLLOWING THE VALVE IMPLANT HOSPITALIZATION OCCURRED AS RESULT OF ACUTE HYPOXIC RESPIRATORY FAILURE. THE PHYSICIAN INDICATED THE EVENT WAS UNRELATED THE VALVE. THE CAUSE WAS NOT REPORTED. THE PATIENT WAS DISCHARGED 7 DAYS FOLLOWING ADMISSION. APPROXIMATELY 16 DAYS LATER, THE PATIENT WAS HOSPITALIZED AGAIN FOR SHORTNESS OF BREATH. ACUTE HEART FAILURE WAS REPORTED. ELEVATED TROPONIN WAS MEASURED AND REPORTED AS LIKELY DEMAND ISCHEMIA FROM THE HEART FAILURE. A MYOCARDIAL INFARCTION DID NOT OCCUR. THE EVENT WAS TREATED WITH OXYGEN AND DIURETICS. STENOSIS WAS REPORTED. DISCHARGE OCCURRED 8 DAY FOLLOWING ADMISSION. THE PHYSICIAN INDICATED THIS EVENT WAS NOT RELATED TO THE VALVE. HOWEVER, THE CAUSE WAS NOT REPORTED. APPROXIMATELY 3.5 MONTHS LATER THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT WITH HYPOXIA. BILEVEL POSITIVE AIRWAY PRESSURE (BPAP) AND 100% HIGH-FLOW NASAL CANNULA (HFNC) THERAPY WERE INITIATED WITHOUT IMPROVEMENT. COMFORT MEASURES WERE PROVIDED, AND A NON-CARDIOVASCULAR DEATH OCCURRED. HYPOXIA WAS THE REPORTED CAUSE OF DEATH. PER THE PHYSICIAN, THE HYPOXIA AND DEATH WERE NOT RELATED TO THE VALVE. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: UPDATED DATA: B1, B2, B5, B7, G2, H1 (UPDATED REPORT TYPE), H6 NOTE: A DUPLICATE EVENT WAS IDENTIFIED IN PREVIOUS REGULATORY REPORT # 2025587-2025-00145. GOING FORWARD NEW INFORMATION PERTAINING TO THE REPORTED DETAILS OF THESE EVENTS WILL CONTINUE TO BE FILED WITH CURRENT REGULATORY REPORT # 2025587-2025-00349. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS REGULATORY REPORT IS BEING SUBMITTED DUE TO RETROSPECTIVE REVIEW THROUGH CAPA 624392. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21149855&pc=DSQ | 2916596-2025-00015 | 2024-11-22 05:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Thrombosis/Thrombus; Peripheral Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL HEMORRHAGE (ICH). THE ICH APPEARED TO BE SPONTANEOUS, THE EXACT CAUSE WAS UNKNOWN. THERE WAS NO TRAUMA PROCEEDING THE EVENT. PATIENT WAS STABLE PRIOR TO THE EVENT. THE ONLY CLINICAL CHANGE PRIOR WAS A CHANGE TO THE INTERNATIONAL NORMALIZED RATIO (INR) GOAL BECAUSE OF AN ACUTE DEEP VEIN THROMBOSIS (DVT). HER INR GOAL WAS CHANGED FROM 1.8-2.5 (PRIOR GI BLEED) TO 2.0-3.0 BECAUSE OF THE NEW FINDING. A WEEK PRIOR TO THE ICH OCCURRING THE PATIENT DID HAVE A SUPRATHERAPEUTIC INR, BUT THIS WAS CORRECTED, AND THE INR WAS 2 AT THE TIME THAT HER ICH OCCURRED. THE DEATH WAS NOT BEING CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT BEING RETURNED FOR EVALUATION. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: IT WAS ADDITIONALLY REPORTED THAT COMPUTED TOMOGRAPHY (CT) IMAGING OF THE PATIENT'S HEAD AND NECK WAS USED TO CONFIRM THE DIAGNOSIS OF ICH. ULTRASOUND IMAGING WAS PERFORMED ON (B)(6) 2024 AND THE DVT WAS SEEN. PATIENT HAD REACHED OUT PRIOR TO COMPLAINTS OF SWELLING TO LEG AND WAS THERAPEUTIC ON WARFARIN AT THE ONSET OF DVT. NO TREATMENT WAS RENDERED FOR THE BLEEDING. THE PATIENT'S FAMILY CHANGED CODE STATUS TO DO-NOT-RESUSCITATE GIVEN FINDINGS OF CT IMAGING AND PROGNOSIS OF PATIENT BY NEUROSURGERY GIVEN FINDINGS. PRIOR GASTROINTESTINAL BLEED IS REPORTED UNDER MFR 2916596-2025-02227. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION G2: REPORT SOURCE CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS STROKE, VENOUS THROMBOEMBOLISM, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATION. UNDER ¿ANTICOAGULATION¿, THIS SECTION OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE FOR PATIENT USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21150562&pc=DSQ | 3007042319-2025-00059 | 2025-01-09 05:00 | Death | HEARTWARE, INC. | 2025-01-14 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE, WHICH WAS IDENTIFIED AS THE CAUSE OF DEATH. THE PUMP IS NO LONGER ACTIVE DUE TO THE PATIENT'S DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING CEREBRAL HEMORRHAGE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF NEUROLOGICAL DYSFUNCTION EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21150601&pc=OZD | 1220648-2025-25566 | 2024-12-31 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Fluid/Blood Leak; Increase in Pressure; Pumping Stopped | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE IMPELLA CP WAS IMPLANTED TO SUPPORT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION / CARDIOGENIC SHOCK. DURING THE IMPLANT, THE CP HAD PURGE PRESSURE LOW AND MOTOR CURRENT HIGH ALARMS. THE IMPELLA STOPPED AND THEN RESTARTED. WHILE THE IMPELLA WAS BEING PLACED THE PATIENT CODED AND SHORTLY AFTER PLACEMENT THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF PUMP STOP AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE FAILURE MODE "HIGH PURGE PRESSURE" IS CHANGED TO "PURGE LEAK - PUMP" AS THE CLINICAL DETAILS STATED THERE WERE LOW PURGE PRESSURE ALARMS. THE DATA LOGS SHOW THAT THE CASE LASTED FOR 20 MINUTES. THE PUMP RAN FOR 17 MINUTES AND THEN EXPERIENCED A HIGH MOTOR CURRENT PUMP STOP. LOW PURGE PRESSURE WAS SEEN AT THE BEGINNING OF THE CASE BUT RESOLVED AFTER 1 MINUTE AND REMAINED STABLE. NO OTHER PURGE ISSUES ARE NOTED. THERE WERE POSITIONING/SUCTION ALARMS SEEN AND THE MOTOR CURRENT/PLACEMENT SIGNAL ARE NOT PULSATILE BEFORE THE PUMP STOPS. THE MOTOR CURRENT DROPS FROM 400MA TO 200MA AND PUMP AND THE PUMP WAS TURNED OFF, AND THEN 8 SECONDS LATER THERE IS A HIGH MOTOR CURRENT PUMP STOP AS IT WAS ATTEMPTED TO BE TURNED ON. THE PUMP IS RESTARTED AND THEN THERE ARE SUCTION ALARMS AFTER. IT RUNS FOR 28 MORE SECONDS AND THEN THE PUMP IS REMOVED. THE CLINICAL DETAILS STATE THAT THE PATIENT WAS CODING UPON IMPELLA PLACEMENT. THE ROOT CAUSE OF THE TEMPORARY PUMP STOP COULD NOT BE DETERMINED. THE ROOT CAUSE OF THE PURGE LEAK - PUMP COULD NOT BE DETERMINED SINCE PRODUCT WAS NOT RETURNED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK CPI & IMPELLA RP WITH SMARTASSIST SYSTEM SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA 5.5 WITH SMARTASSIST & IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK CPI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151425&pc=NGV | 2124215-2024-83729 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-14 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024 AVERAGE TIME TO EVENT: 432 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151430&pc=BZD | 2518422-2024-66025 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151492&pc=NGV | 2124215-2024-83728 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-14 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 77 TO (B)(6) YEARS. PATIENT SEX: 3 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 209 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151498&pc=NGV | 2124215-2024-83726 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-14 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q4 2024. AVERAGE TIME TO EVENT: 432 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151500&pc=CAF | 2518422-2025-002389 | 2023-09-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | CAF | DELUXE 230V 50HZ IN | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K170853 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DELUXE 230V THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151711&pc=BZD | 2518422-2025-002392 | 2021-12-06 05:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151806&pc=DSQ | 2916596-2024-53023 | 2024-12-23 05:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO VENTRICULAR FIBRILLATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS INTENDED, AND NO DEVICE ISSUE OR MALFUNCTION MAY HAVE CONTRIBUTED TO THE PATIENT'S CAUSE OF DEATH. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED ARRHYTHMIA AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151911&pc=DSQ | 2916596-2024-53032 | 2024-12-08 05:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Pumping Stopped | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS RELOCATED TO A NEW CLINIC IN (B)(6) AFTER BEING IMPLANTED IN (B)(6). IT WAS REPORTED TO THE VENTRICULAR ASSIST DEVICE (VAD) CENTER THAT THE PATIENT WAS FOUND DECEASED AT HOME WITH THE VAD PUMP OFF ON (B)(6) 2024. THE PATIENT'S SYSTEM CONTROLLER WAS SHIPPED BACK TO THE IMPLANTING CENTER FOR INTERROGATION TO IDENTIFY POSSIBLE CAUSE OR TIMING OF THE INCIDENT. THE LOG FILES WERE REVIEWED AND CAPTURED RANDOM POWER CABLE DISCONNECT AND LOW VOLTAGE ADVISORY EVENTS ON 08DEC2024 AT 03:05 TO 04:10 WHILE ON THE MOBILE POWER UNIT (MPU) POWER. THERE LOOKED TO BE A WEAK CONNECTION BETWEEN THE WHITE SYSTEM CONTROLLER POWER LEAD AND PATIENT CABLE. ON 08DEC2024 0552 IT APPEARED THAT THERE WAS A LOSS OF POWER TO THE MPU RESULTING IN LOW VOLTAGE HAZARD/NO EXTERNAL POWER EVENTS WHICH ENABLED THE EMERGENCY BACKUP BATTERY (EBB) IN THE SYSTEM CONTROLLER. THESE EVENTS LASTED UNTIL 08:02 WHEN THE VAD TURNED OFF WHEN THE EBB WAS DEPLETED.Event Description: THE PATIENT'S NEPHEW FOUND THEM UNRESPONSIVE AND THE CLINIC WAS NOTIFIED. THE PATIENT WAS FOUND WITH NO POWER CONNECTED, THE MOBILE POWER UNIT (MPU) WAS PLUGGED INTO THE OUTLET AND A BATTERY WAS IN THE BATTERY CLIP ON THE NIGHTSTAND WITH ANOTHER BATTERY ON THE FLOOR. THE CLINIC WAS UNABLE TO SEND LOG FILES WHEN ASKED TO INTERROGATE THE SYSTEM CONTROLLER. THE PATIENT'S NEPHEW THEN RETURNED THE PATIENT'S SYSTEM CONTROLLERS TO A DIFFERENT CLINIC WHO SUBMITTED THE LOG FILES FOR ANALYSIS. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS UNKNOWN IF THE DEATH WAS DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED PER THE LOG FILE REPORT. NO PRODUCTS WOULD BE RETURNED. IT WAS UNKNOWN IF THE MPU HAD A V-LOCK CONNECTOR ON THE AC POWER CORD. IT WAS UNKNOWN IF THE AC POWER CORD CAME LOOSE FROM THE BACK OF THE UNIT. IT WAS UNKNOWN IF ANY ENVIRONMENTAL FACTORS INVOLVED AS THE INFORMATION WAS NOT PROVIDED BY THE FAMILY. THE ONLY INFORMATION THAT WAS KNOWN WAS PROVIDED BY THE FAMILY. THE ONLY THING PROVIDED WAS THE SYSTEM CONTROLLER WITH LOG FILES BECAUSE THE OTHER CENTER DECLINED TO ACCESS THE LOG FILES. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED A FULL BATTERY DEPLETION RESULTING IN A PUMP STOP. A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 08DEC2024. THE MOBILE POWER UNIT (MPU) POWERING THE DEVICE EVENTUALLY DEPLETED IN BATTERY, AND THE CONTROLLER BACKUP BATTERY (EBB) BEGAN TO SUPPORT THE SYSTEM. A NO EXTERNAL POWER HAZARD ALARM BECAME ACTIVE FOR THE DURATION OF THE FILE. THE EBB POWERED THE SYSTEM UNTIL IT ALSO DEPLETED, AND THE PUMP POWERED OFF AND REMAINED OFF FOR THE DURATION OF THE FILE. THERE WAS NO USER RESPONSE CAPTURED IN THE FILE, AND A SPECIFIC CAUSE FOR THE BATTERIES BEING ALLOWED TO DEPLETE COULD NOT BE CONCLUSIVELY DETERMINED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED PRIOR TO THE BATTERY DEPLETION EVENT. IT WAS REPORTED THAT THE PATIENT WAS FOUND DECEASED AT HOME WITH THE VENTRICULAR ASSIST DEVICE PUMP OFF ON (B)(6) 2024. IT WAS COMMUNICATED THAT THE CAUSE OF DEATH WAS UNKNOWN, AND THE DEVICE OPERATED AS EXPECTED PRIOR TO THE EVENT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 2 OF THE IFU, ¿SYSTEM OPERATIONS¿, STATES THAT THE 11 VOLT LITHIUM-ION BACKUP BATTERY SHOULD BE USED ONLY FOR TEMPORARY SUPPORT DURING A POWER-LOSS EMERGENCY. THE 11 VOLT LITHIUM-ION BACKUP BATTERY INSIDE THE HEARTMATE 3 SYSTEM CONTROLLER PROVIDES ENOUGH POWER TO RUN THE IMPLANTED HEARTMATE 3 PUMP FOR AT LEAST 15 MINUTES IF THE MAIN POWER SOURCE (EITHER THE POWER MODULE, MOBILE POWER UNIT, OR TWO HEARTMATE 14 VOLT LITHIUM-ION BATTERIES) IS DISCONNECTED OR FAILS. INAPPROPRIATE USE OF THE 11 VOLT LITHIUM-ION BACKUP BATTERY MAY RESULT IN DIMINISHED RUN TIME DURING A POWER-LOSS EMERGENCY. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON SYSTEM CONTROLLER ALARMS, AS WELL AS THE PROPER ACTIONS ASSOCIATED WITH THEM. THESE DOCUMENTS ALSO WARN OF EVENTS THAT MAY CAUSE THE PUMP TO STOP AND HOW TO PREVENT PUMP STOPS FROM OCCURRING. ADDITIONALLY, THE ¿HANDLING EMERGENCIES¿ SECTION OF THE PATIENT HANDBOOK LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE. FURTHERMORE, THE PATIENT HANDBOOK CAUTIONS THE USER TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151968&pc=PNJ | 9612164-2025-00263 | 2024-10-22 04:00 | Death | MEDTRONIC IRELAND | 2025-01-14 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Pericardial Effusion |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THE AUTHORS DESCRIBED ONE PATIENT DEATH; HOWEVER, THE SPECIFIC CAUSE OF DEATH WAS UNKNOWN BUT OCCURRED AFTER THE LEADLESS IPG IMPLANT DUE TO A PERI-POCKET TISSUE INFECTION FROM A TRANSVENOUS SYSTEM. THE PATIENT HAD SEVERE COMORBIDITIES AND SUBSEQUENTLY DIED. THERE WAS ONE PATIENT WHO DEVELOPED PERICARDIAL EFFUSION HOURS AFTER THE IMPLANTATION, AND WAS DISCHARGED FOLLOWING PERICARDIOCENTESIS FOR DRAINAGE. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/78 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EFFICACY AND SAFETY OF LEADLESS VENTRICULAR PACEMAKER: A SINGLE-CENTER RETROSPECTIVE OBSERVATIONAL STUDY. CARDIO VASCULAR DIAGNOSIS AND THERAPY. 2024;14(5):878-889. DOI: 10.21037/CDT-24-181. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152040&pc=OZD | 1220648-2025-25569 | 2024-12-31 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Defective Device | Arrhythmia; Hematoma; Hemorrhage/Bleeding; Ventricular Fibrillation | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING THE IMPLANT, THE REPOSITIONING SHEATH KINKED AND IT WAS CAUSING DIFFICULTY REPOSITIONING THE PUMP. ALSO, DURING INSERTION, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS DEFIBRILLATED AND SUPPORT CONTINUED. FURTHERMORE, THE PATIENT DEVELOPED A HEMATOMA WITH OOZING AT THE IMPELLA ACCESS SITE. THE DOCTOR THOUGHT THIS OCCURRED FROM THE SHEATH EXCHANGE. MANUAL PRESSURE WAS HELD AND HEPARIN WAS WITHHELD. THE PATIENT EXPIRED AFTER 25.50 HOURS OF IMPELLA SUPPORT DUE TO HAVING VENTRICULAR TACHYCARDIA AND THE PATIENT'S FAMILY WITHDRAWING CARE. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR INVESTIGATION. A DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. AS PER THE GIVEN CLINICAL INFORMATION, THE PUMP WAS PLACED SUCCESSFULLY USING A LONG PEEL-AWAY SHEATH. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AFTER PUMP PLACEMENT. ADDITIONALLY, THE DOCTOR REPORTED DIFFICULTY REPOSITIONING THE CATHETER DUE TO A KINK IN THE REPOSITIONING SHEATH CAUSED BY THE ANGLE OF INSERTION. ALSO, HEMATOMA WERE NOTED AT THE IMPELLA INSERTION SITE, ATTRIBUTED TO SHEATH EXCHANGES, AND THE DOCTOR APPLIED MANUAL PRESSURE AND RE-DRESSED THE SITE. THE INTRODUCER DAMAGE WAS REMOVED AND ADDED TO THE POSITIONING ISSUE AFTER THE INVESTIGATION OF CLINICAL DETAILS. THE CAUSE OF THE POSITIONING ISSUE WAS USE RELATED TO STEEP ANGLE OF INSERTION, WHICH CAUSE THE KINK IN REPOSITION SHEATH. THE CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS USE RELATED SINCE RE-DRESSED THE SITE AND MAINTAINED THE ANGLE MATCH. RESOLVE THE ISSUE OF GROWING HEMATOMA. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152221&pc=KDI | 3003464075-2025-00029 | 2024-12-15 05:00 | Death | NXSTAGE MEDICAL, INC. | 2025-01-14 05:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K170469 |  | 1 | Event Description: A REPORT WAS RECEIVED ON 19 DEC 2024 FROM A HEALTH CARE PROFESSIONAL (HCP) OF A 46 YEAR OLD MALE PATIENT APPROVED FOR PERFORMING SOLO HOME HEMODIALYSIS TREATMENT WITH A MEDICAL HISTORY INCLUDING DIABETES AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 27 DEC 2024 FROM THE HOME THERAPY MANAGER (HTM) WHO STATED THE PATIENT WAS TREATING SOLO AT THE TIME OF THE EVENT AND WAS FOUND DECEASED ON (B)(6) 2024 CONNECTED TO THE DEVICE. PER THE HTM THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152251&pc=BZD | 2518422-2025-002413 | 2024-05-29 04:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | REMSTAR AUTOA-FLEX | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152304&pc=OZD | 1220648-2025-25570 | 2025-01-01 05:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, PUMP STOP WAS NOTED ON CONSOLE. NURSE REPORTED THEY WERE IN THE PROCESS OF TRANSITIONING PATIENT TO COMFORT MEASURES ONLY. PATIENT EXPIRED SHORTLY AFTER. NURSE REPORTED PUMP STOP WAS NOT THE CAUSE.Event Description: IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK CPI & IMPELLA RP WITH SMARTASSIST SYSTEM SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA 5.5 WITH SMARTASSIST & IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK CPI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR ANALYSIS. THE DATA LOGS SHOWED A MOTOR CURRENT HIGH PUMP STOP. THERE WAS A MOTOR CURRENT SPIKE WITH A CONCURRENT DROP IN MOTOR SPEED, DROP IN PURGE FLOW, AND A RISE IN PURGE PRESSURE. THE LOGS SHOWED FAILED RESTART ATTEMPTS. THE CLINICAL DETAILS DO NOT PROVIDE ANY ADDITIONAL INFORMATION ON PUMP STOP BUT REPORT THAT IT DID NOT OCCUR DURING PATIENT MOVEMENT. THE ROOT CAUSE OF THE PUMP STOP WAS NOT DETERMINED AS NO PRODUCT WAS RETURNED AND LIMITED CLINICAL INFORMATION RELATED TO FAILURE WAS PROVIDED. Manufacturer Narrative: B.1 REVISED TO ADD ADVERSE EVENT DUE TO MEDICAL SAFETY OFFICER REVIEW SINCE THE PREVIOUS MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25570 WAS SUBMITTED. B.2 REVISED TO ADD DEATH AND DATE OF DEATH DUE TO MEDICAL SAFETY OFFICER REVIEW SINCE THE PREVIOUS MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25570 WAS SUBMITTED. B.5 ABIOMED'S MEDICAL SAFETY OFFICER REVIEW WAS COMPLETED AND INFORMATION WAS ADDED SINCE THE PREVIOUS MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25570 WAS SUBMITTED. H.1 REVISED TO REFLECT DEATH DUE TO MEDICAL SAFETY OFFICER REVIEW SINCE THE PREVIOUS MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25570 WAS SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152329&pc=MNB | 1222780-2025-00018 | 2024-12-13 05:00 | Death | HOLOGIC, INC | 2025-01-14 05:00 | MNB | NOVASURE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P010013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON (B)(6) 2024. INITIALLY THE PHYSICIAN RESECTED A FIBROID WITH THE FLUENT PRO. THE PHYSICIAN WAS ABLE TO RESECT THE FIBROID AND ALMOST HIT THE 2500 DEFICIT LIMIT. MYOSURE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PHYSICIAN PERFORMED A NOVASURE PROCEDURE WITH A SUCCESSFUL ABLATION. THE PATIENT WAS DISCHARGED AT THE TIME AND APPEARED FINE. 2 DAYS LATER THE PHYSICIAN WAS EXAMINING THE PATHOLOGY SAMPLES AND NOTED THAT THE PATIENT HAD DECEASED DUE TO A CARDIAC ARREST ON (B)(6) 2024. NO OTHER INFORMATION IS AVAILABLE, MULTIPLE ATTEMPTS AT OBTAINING INFORMATION WERE NOT RESPONDED. Manufacturer Narrative: HOLOGIC IS SUBMITTING THIS REPORT IN ACCORDANCE WITH 21. CFR PART 803. THE SUBMISSION IS BASED UPON THE CURRENTLY AVAILABLE INFORMATION. THE SUBMISSION OF THIS REPORT DOES NOT REPRESENT A CONFIRMATION OF DEVICE MALFUNCTION INVOLVING THE HOLOGIC PRODUCTS IN THE EVENT DESCRIBED. LOT AND SERIAL NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152355&pc=HIH | 1222780-2025-00019 | 2024-12-13 05:00 | Death | HOLOGIC, INC | 2025-01-14 05:00 | HIH | MYOSURE REACH TISSUE REMOVAL DEVICE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K152723 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON (B)(6) 2024. INITIALLY THE PHYSICIAN RESECTED A FIBROID WITH THE FLUENT PRO. THE PHYSICIAN WAS ABLE TO RESECT THE FIBROID AND ALMOST HIT THE 2500 DEFICIT LIMIT. MYOSURE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PHYSICIAN PERFORMED A NOVASURE PROCEDURE WITH A SUCCESSFUL ABLATION. THE PATIENT WAS DISCHARGED AT THE TIME AND APPEARED FINE. 2 DAYS LATER THE PHYSICIAN WAS EXAMINING THE PATHOLOGY SAMPLES AND NOTED THAT THE PATIENT HAD DECEASED DUE TO A CARDIAC ARREST ON (B)(6) 2024. NO OTHER INFORMATION IS AVAILABLE, MULTIPLE ATTEMPTS AT OBTAINING INFORMATION WERE NOT RESPONDED. Manufacturer Narrative: HOLOGIC IS SUBMITTING THIS REPORT IN ACCORDANCE WITH 21. CFR PART 803. THE SUBMISSION IS BASED UPON THE CURRENTLY AVAILABLE INFORMATION. THE SUBMISSION OF THIS REPORT DOES NOT REPRESENT A CONFIRMATION OF DEVICE MALFUNCTION INVOLVING THE HOLOGIC PRODUCTS IN THE EVENT DESCRIBED. LOT AND SERIAL NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152356&pc=DSQ | 2916596-2024-53026 | 2024-12-24 05:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem | Bacterial Infection; Hemorrhage/Bleeding; Mitral Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE HOSPITAL ON (B)(6) 2024 WITH BLEEDING FROM THE DRIVELINE WOUND VACUUM SITE AND LOW FLOW ALARMS. THE PATIENT'S BLOOD AND VOLUME WERE RESUSCITATED, AND THE ALARMS PERSISTED. THE MEAN ARTERIAL PRESSURE WAS IN THE 70'S-80'S WHEN THE PATIENT WAS ON 7.5 MCG/KG/MIN DOBUTAMINE. AN ECHOCARDIOGRAM REVEALED A DILATED LEFT VENTRICLE WITH SEVERE MITRAL VALVE REGURGITATION DESPITE INCREASES IN VENTRICULAR ASSIST DEVICE (VAD) SPEED. A COMPUTED TOMOGRAPHY (CT) WAS COMPLETED FOR OUTFLOW GRAFT OBSTRUCTION EVALUATION. IT WAS NOTED THAT THERE WAS DECREASED DENSITY THROUGHOUT THE PROXIMAL ASPECT OF THE LEFT VAD (LVAD) OUTFLOW CANNULA THAT WAS CONNECTED TO THE PUMP. LOG FILES WERE SENT FOR REVIEW. THE EVENT LOG CAPTURED PERSISTENT LOW FLOW EVENTS WITH THE FLOW NOTED TO BE CONSISTENTLY LESS THAN 2.0 LITERS PER MINUTE (LPM). IT WAS NOTED THAT THE PULSATILITY INDEX (PI) VALUES WERE NON-VARIABLE AND SETTLED INTO THE 2 RANGE. THESE PATTERNS WERE TYPICALLY ASSOCIATED WITH AN OBSTRUCTION IN THE LVAD CIRCUIT. THE CT SHOWED POTENTIAL OBSTRUCTION WHICH WOULD COINCIDE WITH THE FINDINGS IN THE LOG FILE DATA. THE PATIENT WAS NOT A SURGICAL CANDIDATE AND PASSED AWAY ON (B)(6) 2024 DUE TO OUTFLOW GRAFT OBSTRUCTION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED OUTFLOW GRAFT OBSTRUCTION WAS UNABLE TO BE CONFIRMED AS NO IMAGES WERE SUBMITTED FOR REVIEW AND THE DEVICE WAS NOT RETURNED FOR EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS THAT APPEARED CONSISTENT WITH A POTENTIAL OBSTRUCTION. HOWEVER, A SPECIFIC CAUSE FOR THE LOW FLOW EVENTS, AS WELL AS A DIRECT CORRELATION TO THE REPORTED OBSTRUCTION, COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHERMORE, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS ON (B)(6) 2024 FROM 01:09:59 ¿ 09:30:45, PER THE TIMESTAMPS. A SUSTAINED LOW FLOW ALARM WAS CAPTURED FOR THE DURATION OF THE FILE. IN ADDITION, REVIEW OF THE SUBMITTED CONTROLLER PERIODIC LOG FILE REVEALED A GRADUAL DECLINE IN FLOW OVER TIME FROM (B)(6) 2024 TO (B)(6) 2024. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE SYSTEM APPEARED TO BE OPERATING AS INTENDED AT THE STORED PATIENT SPEED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, DRIVELINE INFECTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5 OF THE IFU, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING" OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152379&pc=DSQ | 2916596-2025-00036 | 2024-12-10 05:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Heart Failure/Congestive Heart Failure; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED MASSIVE BLEEDING FROM THE THORACIC PLURAL SPACE ON (B)(6) 2024. THE PATIENT HAD A BLOOD TRANSFUSION ON (B)(6) 2024, AND THE APPROXIMATE NUMBER OF UNITS OF RED CELL CONCENTRATE (RCC) TRANSFUSED PER DAY WAS 16 UNITS OR MORE. LABORATORY VALUES WERE UNKNOWN. IT WAS STATED THAT AFTER PUMP IMPLANTATION, INTRAOPERATIVELY FOR CANNULATION CHANGE OF RIGHT VENTRICULAR ASSIST DEVICE (RVAD) ON (B)(6) 2024, THE BLEED FROM THE LUNGS BECAME UNCONTROLLED AND THE CIRCULATORY WAS DISRUPTED. IT WAS STATED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS INTRAOPERATIVE HEMORRHAGE. THE CAUSE OF BLEEDING WAS RELATED TO ANY RE-OPERATIVE PROCEDURE AND WAS POSSIBLY RELATED TO THE PUMP. THE DEVICE FUNCTIONED AS EXPECTED. THE PUMP WAS EXPLANTED.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED LOW FLOW ALARMS AT NIGHT AFTER IMPLANTATION AND THE FREQUENCY INCREASED. THE CAUSE OF THE LOW FLOW ALARMS WAS THOUGHT TO BE DUE TO THE EARLY POSTOPERATIVE PERIOD FOLLOWING IMPLANT. THE SITE PLANNED TO OBSERVE AS TROUBLESHOOTING FOR THE ALARMS. THERE WERE NO PATIENT SYMPTOMS DURING THE TIME OF THE LOW FLOWS. THE PATIENT REQUIRED A LOW FLOW SOFTWARE ADJUSTMENT FOR PERIOPERATIVE STABLE MANAGEMENT (QOL), WHICH WAS PERFORMED THE SAME DAY. THE LOW FLOWS PERSISTED. WHEN THE PATIENT WAS POSTOPERATIVE WEEK 1 THEY REQUIRED A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) FOR RIGHT HEART FAILURE. IT WAS NOTED THAT ALTHOUGH THE RIGHT HEART FAILURE HAD BEEN PROGRESSING IT WAS UNEXPECTED FOR RVAD IMPLANTATION. ALTHOUGH A CAUSAL RELATIONSHIP WITH THE DEVICE COULD NOT BE RULED OUT, IT WAS UNRELATED. THE PATIENT HAD EXPERIENCED LEFT VENTRICULAR COLLAPSE WITHOUT RVAD SUPPORT. DURING SURGERY BLEEDING BECAME UNCONTROLLABLE, LEADING TO CIRCULATORY COLLAPSE AND DEATH. THE PATIENT WAS EXPLANTED DUE TO POST OPERATION SURGICAL INCIDENT. THE CAUSE OF THE BLEEDING WAS UNDER INVESTIGATION AT THE TIME. LOG FILES WERE SENT FOR ANALYSIS TO CONFIRM HOW THE PUMP HAD BEEN OPERATING BETWEEN 14:00 AND 18:00 ON 25DEC2024. THE DATA RECORDED ON THE PERIODIC LOG FILE INDICATED THAT THERE WAS AN ACTIVE LVAD_OFF ALARM FLAG ASSOCIATED WITH AN (F69) INTENTIONAL PUMP STOP CONTROLLER FAULT FLAG ON 25DEC2024 16:45:42.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED MASSIVE BLEEDING ON (B)(6) 2024 IN THE UPPER GASTROINTESTINAL (GI) TRACT. THEY RECEIVED A BLOOD TRANSFUSION ON (B)(6) 2024, AND THE APPROXIMATE NUMBER OF UNITS OF RED CELL CONCENTRATE (RCC) TRANSFUSED PER DAY WAS 16 UNITS OR MORE. THE PATIENT'S PROTHROMBIN TIME (INR) ON (B)(6) 2024 WAS 2.0 IU 2024/12/10, ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT) WAS 46 SECONDS, AND THE PLATELET COUNT (PLT) WAS 11.4 ×104/L. THEY RECEIVED WARFARIN AND DRUG TREATMENT FOR THE BLEEDING. THE PATIENT HAD A HISTORY OF LESIONS OR DISEASE AT THE BLEEDING SITE, WHICH HASTENED THE DEVELOPMENT OF THE BLEEDING IN THE STOMACH. IT WAS DETERMINED THAT THE EVENT WAS "PROBABLY RELATED" TO THE DEVICE.Event Description: IT WAS LATER REPORTED THAT THE MASSIVE BLEEDING OF THE THORACIC CAVITY WAS THE SAME BLEEDING EVENT THAT LED TO THE PATIENT'S DEATH. IT WAS CONFIRMED THAT THE ACTIVATED CLOTTING TIME (ACT) WAS "OFF THE CHARTS" DURING THE SURGERY, WHICH SUGGESTED THAT DISSEMINATED INTRAVASCULAR COAGULATION OCCURRED FOR SOME REASON, WHICH EXHAUSTED THE COAGULATION FACTORS AND RESULTED IN THE MASSIVE BLEEDING EVENT. IT WAS FURTHER CLARIFIED THAT THE RIGHT HEART FAILURE OCCURRED ON (B)(6) 2024. THE RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS NEEDED ONE WEEK AFTER IMPLANT SURGERY, BUT THE ITEM SELECTION FOR RIGHT HEART FAILURE WAS NOT POSSIBLE ONE WEEK AFTER SURGERY. THE CAUSAL RELATIONSHIP TO THE DEVICE WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE RELATED ISSUES THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE. A DIRECT CORRELATION BETWEEN (B)(6) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED EVENT OF LOW FLOW ALARMS WAS CONFIRMED BY AN ONSITE ABBOTT REPRESENTATIVE. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE MODULAR CABLE WAS RETURNED ATTACHED TO THE PUMP CABLE. APPROXIMATELY 8¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP OUTLET PORT WITH THE BEND RELIEF PROPERLY ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES DID NOT REVEAL ANY ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE SUBMITTED CONTROLLER PERIODIC LOG FILE AND THE LVAD EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. A IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, BLEEDING, AND RIGHT HEART FAILURE. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7, ¿ALARMS AND TROUBLESHOOTING, EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. SECTION 7 ¿ALARMS AND TROUBLESHOOTING¿ PROVIDES ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE IN A SUB-SECTION ENTITLED "WHAT NOT TO DO: DRIVELINE AND CABLES". SECTION 8 ¿EQUIPMENT STORAGE AND CARE¿ STATES: "AS NEEDED, CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, USE WARM WATER AND MILD DISH SOAP." THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. C, ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. SECTION 4 OF THE PATIENT HANDBOOK ¿LIVING WITH THE HEARTMATE III¿ CONTAINS INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE. SECTION 4 ALSO INSTRUCTS THE USER TO ¿KEEP YOUR DRIVELINE CLEAN. WIPE OFF ANY DIRT OR GRIME. IF THE DRIVELINE GETS DIRTY, USE A TOWEL WITH MILD DISH SOAP AND WARM WATER TO GENTLY CLEAN IT. NEVER SUBMERGE THE DRIVELINE OR OTHER SYSTEM COMPONENTS IN WATER OR LIQUID.¿ FOLLOWING SOFTWARE UPGRADES, THE HEARTMATE 3 LVAS POCKET GUIDES TO ALARMS FOR PATIENTS, REV. A, AND CLINICIANS, REV. A, EXPLAIN THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.0 LPM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4 - PATIENT WEIGHT UPDATED. B5 NARRATIVE INFORMATION. H6 - ADVERSE EVENT PROBLEM CODES UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) , DID NOT REVEAL ANY DEVICE RELATED ISSUES THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE. A DIRECT CORRELATION BETWEEN (B)(6) AND THE REPORTED BLEEDING AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE REPORTED EVENT OF LOW FLOW ALARMS WAS CONFIRMED BY AN ONSITE ABBOTT REPRESENTATIVE. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE MODULAR CABLE WAS RETURNED ATTACHED TO THE PUMP CABLE. APPROXIMATELY 8¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP OUTLET PORT WITH THE BEND RELIEF PROPERLY ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE RETURNED PUMP, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO ADHERED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FLOW OR FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE SUBMITTED CONTROLLER PERIODIC LOG FILE AND THE LVAD EVENT AND PERIODIC LOG FILES RETRIEVED FROM THE RETURNED DEVICE APPEARED TO CAPTURE THE DEVICE FUNCTIONING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH AND BLEEDING. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. SECTION 7 ¿ALARMS AND TROUBLESHOOTING¿ PROVIDES ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE IN A SUB-SECTION ENTITLED "WHAT NOT TO DO: DRIVELINE AND CABLES". SECTION 8 ¿EQUIPMENT STORAGE AND CARE¿ STATES: "AS NEEDED, CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, USE WARM WATER AND MILD DISH SOAP." THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. SECTION 4 OF THE PATIENT HANDBOOK ¿LIVING WITH THE HEARTMATE III¿ CONTAINS INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE. SECTION 4 ALSO INSTRUCTS THE USER TO ¿KEEP YOUR DRIVELINE CLEAN. WIPE OFF ANY DIRT OR GRIME. IF THE DRIVELINE GETS DIRTY, USE A TOWEL WITH MILD DISH SOAP AND WARM WATER TO GENTLY CLEAN IT. NEVER SUBMERGE THE DRIVELINE OR OTHER SYSTEM COMPONENTS IN WATER OR LIQUID.¿ FOLLOWING SOFTWARE UPGRADES, THE HEARTMATE 3 LVAS POCKET GUIDES TO ALARMS FOR PATIENTS, REV. A, AND CLINICIANS, REV. A, EXPLAIN THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.0 LPM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152382&pc=BZD | 2518422-2025-002416 | 2024-01-17 05:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152535&pc=DSQ | 2916596-2025-00001 | 2024-12-26 05:00 | Death | THORATEC CORPORATION | 2025-01-14 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A CEREBROVASCULAR ACCIDENT (CVA). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152540&pc=HIG | 1222780-2025-00020 | 2024-12-13 05:00 | Death | HOLOGIC, INC | 2025-01-14 05:00 | HIG | FLUENT PRO FLUID MANAGEMENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K240886 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON DECEMBER (B)(6) 2024. INITIALLY THE PHYSICIAN RESECTED A FIBROID WITH THE FLUENT PRO. THE PHYSICIAN WAS ABLE TO RESECT THE FIBROID AND ALMOST HIT THE 2500 DEFICIT LIMIT. MYOSURE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PHYSICIAN PERFORMED A NOVASURE PROCEDURE WITH A SUCCESSFUL ABLATION. THE PATIENT WAS DISCHARGED AT THE TIME AND APPEARED FINE. 2 DAYS LATER THE PHYSICIAN WAS EXAMINING THE PATHOLOGY SAMPLES AND NOTED THAT THE PATIENT HAD DECEASED DUE TO A CARDIAC ARREST ON (B)(6) 2024. NO OTHER INFORMATION IS AVAILABLE, MULTIPLE ATTEMPTS AT OBTAINING INFORMATION WERE NOT RESPONDED. Manufacturer Narrative: THE CONSOLE WAS RETURNED FOR INVESTIGATION AN INFLOW SPEED FAULT ON THE FLUENT PRO CAN ONLY BE DISPLAYED DURING STARTUP OR DURING PRIMING. TO TEST FOR THIS FAILURE, THE CONSOLE WAS POWERED ON, SET UP WITH THE DISPOSABLES, AND WAS PRIMED. HERE, IT WAS OBSERVED THAT THE CONSOLE WORKED AS INTENDED AND DID NOT DISPLAY AN INFLOW SPEED FAULT. HENCE, THE FAILURE MODE REPORTED BY THE CUSTOMER WAS NOT REPLICATED DURING INVESTIGATION. A VISUAL INSPECTION WAS CONDUCTED AND THE WHEELS ON THE SYSTEM WERE DAMAGED. IT IS DETERMINED THAT THE DAMAGE TO THE WHEELS DOES NOT RELATE TO THE REPORTED ISSUE. TO FURTHER INVESTIGATE THE REPORTED FAILURE, THE BACK COVER WAS REMOVED TO VISUALLY INSPECT THE CABLES INVOLVED WITH THE INFLOW MOTOR. HERE, IT WAS OBSERVED THAT THERE WERE NO SIGNS OF DAMAGES ON THE INFLOW MOTOR AND ITS CABLES. ADDITIONALLY, THE SYSTEM LOG WAS REVIEWED FOR THE DATE OF THE PROCEDURE AND IT WAS OBSERVED THAT AN INFLOW MOTOR SPEED FAULT WAS PRESENT, AS SHOWN IN THE FILE ATTACHED. AN INSPECTION ONTO THE INFLOW MOTOR SHAFT WAS PERFORMED TO DETERMINE IF THERE IS EVIDENCE OF EXCESSIVE WEAR OR LOOSE SHAFT BEARINGS IN THE FLUID DECK. THERE WERE NO SIGNS OF DAMAGES OR ISSUES ON THE FLUID DECK, SPECIFICALLY THE INFLOW DECK. FOR THE REASON THAT THE REPORTED ISSUE CANNOT BE REPRODUCED, THE CAUSE OF THE REPORTED FAILURE COULD NOT BE FOUND. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. WE ARE CURRENTLY UNABLE TO ESTABLISH A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. Manufacturer Narrative: HOLOGIC IS SUBMITTING THIS REPORT IN ACCORDANCE WITH 21. CFR PART 803. THE SUBMISSION IS BASED UPON THE CURRENTLY AVAILABLE INFORMATION. THE SUBMISSION OF THIS REPORT DOES NOT REPRESENT A CONFIRMATION OF DEVICE MALFUNCTION INVOLVING THE HOLOGIC PRODUCTS IN THE EVENT DESCRIBED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED LOT/SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152604&pc=FPA | 3012307300-2025-00702 | 2024-12-14 05:00 | Death | SMITHS MEDICAL ASD, INC. | 2025-01-14 05:00 | FPA | CADD DISPOSABLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K111275 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPIRED WHILE THE PUMP WAS IN USE. PER REPORTER, THE PATIENT WAS FOUND AT HOME AND THE CAUSE OF DEATH IS UNKNOWN. THE PUMP HAD BEEN INFUSING AT A CONTINUOUS RATE OF 5 ML/HOUR WITHOUT ANY REPORTED ISSUES. Manufacturer Narrative: E1: FACILITY NAME (B)(6). H3: NEITHER SAMPLES NOR PICTURES WERE RECEIVED FOR THE ANALYSIS OF THIS COMPLAINT. NO DEVICE HISTORY RECORD WAS REVIEWED SINCE LOT NUMBER WAS NOT PROVIDED. WITHOUT THE RETURN OF THE USED SAMPLES A COMPREHENSIVE FAILURE INVESTIGATION CANNOT BE PERFORMED, AND A PROBABLE CAUSE CANNOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152969&pc=FCG | 3004904811-2025-00012 | 2024-10-01 04:00 | Death | COVIDIEN LP BARRX | 2025-01-14 05:00 | FCG | BNX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Pancreatitis | K103668 |  | 1 | Event Description: IT WAS REPORTED THAT THEY HAD A CASE OF SEVERE BLEEDING AND PANCREATITIS AFTER A FINE NEEDLE BIOPSY AND FINE NEEDLE ASPIRATION PROCEDURES FOR CYST ASPIRATION, CYST ASPIRATION WITH SOLID COMPONENTS, AND SOLID MASS TISSUE SAMPLE, AND ONE PATIENT DIED AS A RESULT OF THE COMPLICATION USING FINE NEEDLE ASPIRATION NEEDLE.Event Description: IT WAS REPORTED THAT THEY HAD A CASE OF SEVERE BLEEDING AND PANCREATITIS AFTER A FINE NEEDLE BIOPSY AND FINE NEEDLE ASPIRATION PROCEDURES FOR CYST ASPIRATION, CYST ASPIRATION WITH SOLID COMPONENTS, AND SOLID MASS TISSUE SAMPLE, AND ONE PATIENT DIED AS A RESULT OF THE COMPLICATION USING FINE NEEDLE ASPIRATION NEEDLE. PATIENT WAS ADMITTED THE DAY AFTER BEACON FNA PROCEDURE DUE TO FULMINANT ACUTE PANCREATITIS. PATIENT DIED THE FOLLOWING DAY AFTER ADMITTED, WITH MULTI-ORGAN FAILURE.Event Description: IT WAS REPORTED THAT A PATIENT DIED AFTER A FINE NEEDLE ASPIRATION PROCEDURE. PATIENT WAS ADMITTED THE DAY AFTER BEACON FNA 22G 1 PASS PROCEDURE CYST EMPTYING DUE TO FULMINANT ACUTE PANCREATITIS. PATIENT DIED THE FOLLOWING DAY AFTER ADMITTED, WITH MULTI-ORGAN FAILURE. PATIENT FINAL DIAGNOSIS ATYPIC SIMPLE CYST (CYTOHISTOLOGIC AND LAB NEGATIVE FOR CANCER) OF PANCREAS/MESENTERY. IMMEDIATE MANAGEMENT FOR FULMINANT ACUTE PANCREATITIS AND COMORBIDITY OF HYPERTENSION (IN TARGET WITH TREATMENT), MEDITERREAN ANEMIA TRAIT (NORMAL HGB LEVELS), NO CONCOMITANT TREATMENTS. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: DSC-22-01, DSC-22-01 SHARKCORE BIO SYS 22G SS NDL (LOT#B000003228). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFO B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: A2(AGE AT EVENT), A3A, B2, B3, B5, G3 D10 CONCOMITANT PRODUCT: UNK-BEACON, UNKNOWN BEACON (LOT#UNKNOWN). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21153758&pc=PJS | 3004209178-2025-00763 | 2024-07-12 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-14 05:00 | PJS | PERCEPT | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Awareness during Anaesthesia; Muscle Weakness; Shaking/Tremors; No Clinical Signs, Symptoms or Conditions | P960009 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE FAMILY MEMBER DID NOT KNOW IF THE DEATH WAS RELATED TO THE DEVICE OR THERAPY BUT SUSPECTS THE ANESTHETIC HAD SOMETHING TO DO WITH THEIR DYING. THE PATIENT NEVER USED THE DEVICE DUE TO AN INABILITY TO RETURN TO THE NEUROLOGIST. THE DEVICE NEVER FUNCTIONED FOR THE PATIENT. THE PATIENT'S CONDITION DETERIORATED RAPIDLY AFTER A BATTERY REPLACEMENT AND WIRE REPAIR ON (B)(6) 2024. A MANUFACTURER REPRESENTATIVE (REP) VISITED THE PATIENT¿S HOME TWICE TO ATTEMPT DEVICE TUNING RIGHT AFTER THE PATIENT REPLACED THE BATTERY IN (B)(6), BUT THE EFFORTS WERE UNSUCCESSFUL. THE FAMILY REPRESENTATIVE SAID THE SURGERY WAS DEEMED A TOTAL FAILURE. THE PATIENT WAS TOO WEAK TO HAVE IT DONE. THEY WENT TO PHYSICAL THERAPY AND COULDN'T DO IT. AFTER THE PATIENT CAME HOME, THEY WENT TO HOSPICE, AND THE DEVICE COULD NEVER BE USED, AND THEIR HANDS NEVER STOPPED SHAKING.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE MANUFACTURER REPRESENTATIVE AND THE SURGEON'S OFFICE INDICATING THAT THE PATIENT'S WOUND WAS CHECKED AFTER SURGERY, AND THE PATIENT WAS REFERRED BACK TO THEIR NEUROLOGIST FOR NEUROSTIMULATOR ADJUSTMENT. NEITHER THE REPRESENTATIVE NOR THE SURGEON'S OFFICE WERE AWARE OF ANY OTHER COMPLICATIONS. THE REPRESENTATIVE WILL FOLLOW UP WITH THE NEUROLOGIST FOR ADDITIONAL DETAILS. (B)(6) 2025 (B)(6) (HCP): NO NEW INFORMATION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE MANUFACTURER'S REPRESENTATIVE, WHO STATED THAT THEY SPOKE WITH THE NEUROLOGIST, WHO WAS ALSO UNAWARE THAT THE PATIENT HAD PASSED AWAY. THE REPRESENTATIVE ALSO NOTED THAT THE PATIENT HAD NOT VISITED THE NEUROLOGIST'S OFFICE SINCE THEIR BATTERY CHANGE PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: NEU_UNKNOWN, SERIAL# UNKNOWN, PRODUCT TYPE UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID NEU_UNKNOWN, SERIAL# UNKNOWN, PRODUCT TYPE: UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: NEU_UNKNOWN, SERIAL#: UNKNOWN AND PRODUCT TYPE: UNKNOWN. SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: NEU_UNKNOWN AND SERIAL/LOT #: UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154055&pc=MSX | 3009156722-2025-00001 | 2024-12-16 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-01-14 05:00 | MSX | TM80 | Device Alarm System | Insufficient Information | K220058 |  | 1 | Event Description: NO ALLEGED MALFUNCTION WAS REPORTED BY THE CUSTOMER. IT WAS REPORTED THAT ON (B)(6) 2024, AT 16:54 (4:54 PM), A PATIENT WAS PLACED ON THE TM80 TELEMETRY MONITOR CONNECTED TO A BENEVISION DMS, HAD VENTRICULAR TACHYCARDIA (V-TACH)/VENTRICULAR FIBRILLATION (V-FIB) TO ASYSTOLE EPISODES, CONTINUED UNTIL 17:56 (5:56 PM). THE ALARM WAS ALLEGEDLY SILENCED AT THE WORKSTATION. CODE WAS CALLED AND CHEST COMPRESSIONS STARTED. PATIENT EXPIRED.Event Description: NO ALLEGED MALFUNCTION WAS REPORTED BY THE CUSTOMER. IT WAS REPORTED THAT ON (B)(6) 2024, AT 16:54 (4:54 PM), A PATIENT WAS PLACED ON THE TM80 TELEMETRY MONITOR CONNECTED TO A BENEVISION DMS, HAD VENTRICULAR TACHYCARDIA (V-TACH)/VENTRICULAR FIBRILLATION (V-FIB) TO ASYSTOLE EPISODES, CONTINUED UNTIL 17:56 (5:56 PM). THE ALARM WAS ALLEGEDLY SILENCED AT THE WORKSTATION. CODE WAS CALLED AND CHEST COMPRESSIONS STARTED. PATIENT EXPIRED. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND PATIENT DATABASE FOR INVESTIGATION. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND PATIENT DATABASE FOR INVESTIGATION. MINDRAY EVALUATED THE DATA AND CONFIRMED THAT THE TM80 USED WITH THE BENEVISION DMS PERFORMED PER SPECIFICATIONS. MINDRAY'S TM80 TELEMETRY MONITOR AND THE BENEVISION DMS SERVER GENERATED THE CORRESPONDING ALARMS NORMALLY. HOWEVER, BECAUSE THE USER HAD THE WORKSTATION IN LOCAL STANDBY MODE FOR THE TM80 THAT WAS ACTIVELY MONITORING THE PATIENT, THE CORRESPONDING ALARMS WERE NOT DISPLAYED AT THE WORKSTATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154188&pc=OZD | 1220648-2025-25571 | 2024-09-20 04:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED HEMOLYSIS WITH A "PROBABLE" RELATIONSHIP TO THE IMPELLA DEVICE USED AND WOULD EVENTUALLY EXPIRE. THE PATIENT WAS ON MEDICAL FLOOR WAITING FOR SURGICAL AORTIC VALVE REPLACEMENT PLANNED AND BEGAN TO DECOMPENSATE. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT AND ON DOBUTAMINE, INTUBATED FOR AIRWAY PROTECTION, AND THEN TO THE CARDIOVASCULAR LAB FOR AN IMPELLA CP DEVICE PLACEMENT. PRIOR TO THE IMPELLA IMPLANT PROCEDURE, THE PATIENT HAD A LEFT VENTRICULAR END-DIASTOLIC PRESSURE (LVEDP) OF 40 AND 20% EJECTION FRACTION. A BIVENTRICULAR AORTIC VALVE REPLACEMENT WAS PERFORMED FOLLOWED BY THE IMPELLA CP DEVICE IMPLANT VIA THE LEFT FEMORAL ARTERY. THE IMPELLA IMPLANT WAS COMPLETED WITHOUT COMPLICATIONS, POSITION VERIFIED, AND THE PATIENT'S HEMODYNAMICS IMMEDIATELY RESPONDED. THE LVEDP WAS NOW 20. THE NEXT DAY THE NURSE REPORTED URINE OUTPUT CHANGED TO A TEA COLOR. THE IMPELLA CP PUMP PERFORMANCE LEVEL WAS REDUCED TO P-7. AORTIC VALVE REPLACEMENT SURGERY WAS PLANNED FOR THE NEXT MONDAY (APPROXIMATELY THREE DAYS LATER). THE FOLLOWING DAY, THE PATIENT WAS WEANING DOWN AT PERFORMANCE LEVEL P-6 WITH GOOD OUTPUTS AND CARDIAC INDEX. URINE OUTPUT AND COLOR IMPROVED. HOWEVER, TWO DAYS LATER, THE PATIENT DECOMPENSATED UPON PREPARATION FOR THE OPERATING ROOM AND EXPIRED. IT WAS REPORTED REGARDING THE HEMOLYSIS, THE PATIENT HAD NOT RECOVERED/THE EVENT HAD NOT RESOLVED. Manufacturer Narrative: THE INVESTIGATION FOR THE HEMOLYSIS HAS BEEN COMPLETED. DATA LOGS SHOWED PUMP RAN WITHOUT ISSUE DURING SUPPORT. THERE WERE SUCTION ALARMS TRIGGERED DURING THE CASE BUT RESOLVED QUICKLY. PRODUCT WAS NOT RETURNED FOR REVIEW. THE ROOT CAUSE OF HEMOLYSIS WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154193&pc=OZD | 1220648-2025-25572 | 2024-09-03 04:00 | Death | ABIOMED, INC. | 2025-01-14 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED VENTRICULAR FIBRILLATION WITH A POSSIBLE RELATIONSHIP TO THE IMPELLA DEVICE USED: PATIENT WENT INTO VENTRICULAR FIBRILLATION AT 13:04 ON (B)(6) 2024. PATIENT WAS SHOCKED AND RETURNED TO NORMAL SINUS RHYTHM (NSR). PATIENT WENT INTO VENTRICULAR FIBRILLATION AT 13:58 ON (B)(6) 2024. PATIENT WAS SHOCKED WITH RETURN TO NSR. AT 14:01 PATIENT WENT BACK INTO VENTRICULAR FIBRILLATION AND WAS SHOCKED AGAIN WITH RETURN TO NSR. THE PATIENT DECOMPENSATED AND EXPIRED. IMPELLA SUPPORT WAS FOR 14 HOURS. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED VENTRICULAR FIBRILLATION/VENTRICULAR TACHYCARDIA (VF/VT) HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE VF/VT COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154276&pc=MSX | 2221819-2025-00001 | 2024-12-16 05:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-01-14 05:00 | MSX | TM80 | Device Alarm System | Insufficient Information |  |  | 1 | Event Description: NO ALLEGED MALFUNCTION WAS REPORTED BY THE CUSTOMER. IT WAS REPORTED THAT ON (B)(6) 2024, AT 16:54 (4:54 PM), A PATIENT WAS PLACED ON THE TM80 TELEMETRY MONITOR CONNECTED TO A BENEVISION DMS, HAD VENTRICULAR TACHYCARDIA (V-TACH)/VENTRICULAR FIBRILLATION (V-FIB) TO ASYSTOLE EPISODES, CONTINUED UNTIL 17:56 (5:56 PM). THE ALARM WAS ALLEGEDLY SILENCED AT THE WORKSTATION. CODE WAS CALLED AND CHEST COMPRESSIONS STARTED. PATIENT EXPIRED.Event Description: NO ALLEGED MALFUNCTION WAS REPORTED BY THE CUSTOMER. IT WAS REPORTED THAT ON (B)(6) 2024, AT 16:54 (4:54 PM), A PATIENT WAS PLACED ON THE TM80 TELEMETRY MONITOR CONNECTED TO A BENEVISION DMS, HAD VENTRICULAR TACHYCARDIA (V-TACH)/VENTRICULAR FIBRILLATION (V-FIB) TO ASYSTOLE EPISODES, CONTINUED UNTIL 17:56 (5:56 PM). THE ALARM WAS ALLEGEDLY SILENCED AT THE WORKSTATION. CODE WAS CALLED AND CHEST COMPRESSIONS STARTED. PATIENT EXPIRED. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND PATIENT DATABASE FOR INVESTIGATION. MINDRAY EVALUATED THE DATA AND CONFIRMED THAT THE TM80 USED WITH THE BENEVISION DMS PERFORMED PER SPECIFICATIONS. MINDRAY'S TM80 TELEMETRY MONITOR AND THE BENEVISION DMS SERVER GENERATED THE CORRESPONDING ALARMS NORMALLY. HOWEVER, BECAUSE THE USER HAD THE WORKSTATION IN LOCAL STANDBY MODE FOR THE TM80 THAT WAS ACTIVELY MONITORING THE PATIENT, THE CORRESPONDING ALARMS WERE NOT DISPLAYED AT THE WORKSTATION. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND PATIENT DATABASE FOR INVESTIGATION. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154293&pc=BZD | 2518422-2025-100610 | 2025-01-09 05:00 | Death | RESPIRONICS, INC. | 2025-01-14 05:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. IT IS UNKNOWN IF THE PATIENT WAS USING THE DEVICE AT THE TIME OF DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154485&pc=FNL | 1824206-2025-00041 | 2024-12-17 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-14 05:00 | FNL | VERSACARE MED SURG BED | Improper or Incorrect Procedure or Method | Fall; Intracranial Hemorrhage | NA |  | 1 | Event Description: IT WAS REPORTED A PATIENT FELL WHEN THEY EXITED A VERSACARE BED. AT 5:42AM THE PATIENT EXITED THE BED FOR AN UNKNOWN CAUSE AND THE BED ALARM WAS TRIGGERED. THE NURSE DEACTIVATED THE ALARM AT THIS TIME. APPROXIMATELY TWO HOURS LATER, THE PATIENT AGAIN EXITED THE BED; HOWEVER, THIS TIME THE BED EXIT ALARM DID NOT TRIGGER. APPARENTLY, THE NURSE ¿FORGOT¿ TO REACTIVATE THE BED EXIT ALARM, AND THE PATIENT FELL SUSTAINING A SUBARACHNOID HEMORRHAGE. ¿A FEW HOURS AFTER THE FALL¿ THE PATIENT DIED. PER THE INSTRUCTIONS FOR USE SETTING THE BED'S EXIT ALARM: THE BED EXIT ALARM SYSTEM CONTROL IS LOCATED ON THE FLIP-UP CONTROL POD ON THE OUTSIDE OF THE HEAD-END. 1.) ZERO THE BED EXIT FEATURE, 2.) MAKE SURE THE BED IS PLUGGED INTO AC POWER; 3.) MAKE SURE THE PATIENT IS CENTERED ON THE BED AND ALIGNED WITH THE HIP LOCATOR, 4.) PRESS THE ENABLE CONTROL UNTIL THE INDICATOR ILLUMINATES. 5.) PRESS THE DESIRED MODE CONTROL. WHEN THE SYSTEM BEEPS ONE TIME AND THE INDICATOR STAYS ON SOLID, THE SYSTEM IS ARMED. ADDITIONALLY, THE IFU STATES THAT "IF THE SYSTEM DOES NOT ARM, THE SYSTEM WILL BEEP RAPIDLY FOR A FEW SECONDS AND THE SELECTED MODE INDICATOR WILL FLASH. THIS MEANS: THE PATIENT'S WEIGHT IS OUTSIDE OF THESE RANGES: A MODEL¿70 LB. TO 500 LB. (31.8 KG TO 227 KG) B THROUGH J MODELS¿50 LB TO 500 LB (23 KG TO 227 KG), THE PATIENT IS NOT IN THE CORRECT POSITION, OR THE SYSTEM HAS MALFUNCTIONED. THE BED EXIT ALARM SYSTEM DOES NOT OPERATE ON BATTERY BACK-UP. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154703&pc=NAY | 2955842-2025-00397 | 2024-12-05 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-14 05:00 | NAY | DAVINCI XI | Output Problem | Thromboembolism | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS SCHEDULED FOR A NON-ROBOTIC PARTIAL CYSTECTOMY PROCEDURE THAT HAD TO BE POSTPONED DUE TO THE SITE DA VINCI SYSTEM HAVING AN UNRECOVERABLE ERBE INTEGRATED ELECTROSURGICAL UNIT (IESU) ISSUE THAT WAS ENCOUNTERED ON THE DAY OF, BUT PRIOR TO, THIS PATIENT¿S PROCEDURE. THE SAME SURGEON HAD TWO ROBOTIC AND 2 NON-ROBOTIC PROCEDURES SCHEDULED. TWO OF THE PROCEDURES (ONE ROBOTIC, ONE NON-ROBOTIC) WERE COMPLETED THAT DAY; THE OTHER TWO HAD TO BE RESCHEDULED FOR A LATER DATE. THE PATIENT FOR THIS REPORT EXPIRED ONE WEEK LATER. THIS PATIENT¿S PROCEDURE WAS ORIGINALLY SCHEDULED FOR (B)(6) 2024 AS NON-ROBOTIC AND WAS RE-SCHEDULED FOR (B)(6) 2024 WHEN THE PATIENT UNDERWENT THE NON-ROBOTIC PARTIAL CYSTECTOMY SURGERY. THE PATIENT SUBSEQUENTLY EXPIRED ON (B)(6) 2024, 2 DAYS AFTER THE CYSTECTOMY WAS PERFORMED. THE PATIENT HAD A HISTORY OF ATRIAL FIBRILLATION AND WAS ON UNSPECIFIED ANTICOAGULATION MEDICATIONS, WHICH WERE HELD TWICE FOR 3 DAYS EACH TIME IN ANTICIPATION OF THE INITIAL AND RESCHEDULED SURGERIES. THE SURGEON SUSPECTS THE PATIENT SUFFERED A PROBABLE THROMBOEMBOLIC ACCIDENT EVENT AS THE PATIENT¿S ANTICOAGULANT TREATMENT WAS INTERRUPTED, THEN RESUMED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: ALTHOUGH THE SURGICAL PROCEDURE WAS NON-ROBOTIC, THE SURGEON REPORTED THAT THE PATIENT SUFFERED A PROBABLE THROMBOEMBOLIC ACCIDENT EVENT RELATED TO THE PATIENT¿S ANTICOAGULANT TREATMENT THAT WAS INTERRUPTED, THEN RESUMED FOLLOWING THE EARLIER DA VINCI IESU ERBE EVENT. A DA VINCI INTUITIVE TECHNICAL SUPPORT ENGINEER WAS CONTACTED REGARDING THE SYSTEM ISSUE THAT OCCURRED PRIOR TO THIS PATIENT'S RESCHEDULED PROCEDURE. AN ERROR CODE POINTING TO THE ERBE IESU WAS NOTED. THERE WAS NO MONOPOLAR ENERGY DELIVERY AVAILABLE. THE SITE DID NOT HAVE AN AVAILABLE BACKUP GENERATOR. THE GENERATOR WAS REPLACED LATER IN THE DAY. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED WHILE WAITING FOR A SURGICAL PROCEDURE. THE PATIENT HAD BEEN ON ANTICOAGULANTS BUT THESE MEDICINES HAD BEEN HELD IN PREPARATION FOR SURGERY. NO OPERATION HAD BEEN PERFORMED AND NO ALTERNATIVE IMMEDIATE OPTIONS FOR THEIR SURGERY WERE PROVIDED. THE CAUSE OF DEATH WAS NOT PROVIDED BUT SPECULATED TO MAY HAVE BEEN DUE TO HOLDING THEIR ANTICOAGULATION MEDICINE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: THIS PATIENT'S SURGICAL PROCEDURE WAS A SCHEDULED NON-ROBOTIC SURGERY; THE REPORTED PATIENT EVENT WAS ALLEGED TO BE INDIRECTLY RELATED BY THE SURGEON WHO REPORTED THAT THE PATIENT SUFFERED A PROBABLE THROMBOEMBOLIC ACCIDENT EVENT RELATED TO INTERRUPTION OF THE THE PATIENT¿S ANTICOAGULANT TREATMENT, THEN RESUMED FOLLOWING THE SURGEON'S EARLIER DA VINCI IESU ERBE EVENT, LEADING TO POSTPONEMENT OF THEIR SURGERY. BASED ON THE INVESTIGATION, THERE IS NO INDICATION THAT THE DEVICE DIRECTLY CAUSED THE REPORTED ADVERSE EVENT; THE CASCADE LED TO THE SURGEON'S SUBSEQUENT PROCEDURE(S) BEING POSTPONED BY THE SURGEON. PLEASE SEE SECTION H10 RELATED MEDWATCH REPORT FOR THE IESU ERBE DEVICE USED DURING A DIFFERENT PATIENT PROCEDURE. AN INTUITIVE SURGICAL, INC. (ISI) FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER SITE TO INVESTIGATE THE SURGEON'S UNRELATED EARLIER ROBOTIC CASE WHERE THE ERBE GENERATOR WAS REPLACED AS A RESULT OF MONOPOLAR POWER ACTIVATION PROBLEMS AND C-35 ERRORS NOTED DURING TESTING. ISI DID RECEIVE THE IESU ERBE GENERATOR FOR FAILURE ANALYSIS. DURING THE POWER-ON TEST, IESU CAUTERY ACTIVATION, AND VISUAL INSPECTION, THE ERROR WAS FOUND, CONFIRMING THE FAULT OCCURRED IN THE FIELD. UPON VISUAL INSPECTION, NO ISSUES WERE FOUND THAT WOULD BE RELATED TO THE REPORTED EVENT. THE COMPLAINT WAS CONFIRMED BASED ON THE FIELD EVALUATION AND FAILURE ANALYSIS, WHICH INDICATES THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE CUSTOMER REPORTED ISSUE. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THE FOLLOWING POSSIBLY RELATED SYSTEM ERRORS: C-34 - ENERGY ACTIVATION HALTED BY AN INSTRUMENT OR INSTRUMENT CABLE CONNECTED TO THE ERBE WHILE USING THE COAG FUNCTION. THE PROBABLE CAUSE OF THE ERRORS IS ATTRIBUTED TO A FAULTY ELECTRICAL COMPONENT INSIDE THE IESU. ONE ERROR INDICATES A LOWER VOLTAGE THAN EXPECTED DURING ENERGY ACTIVATION. THE OTHER ERROR INDICATES A HIGHER CURRENT MEASUREMENT VALUE THAN EXPECTED WHILE THE UNIT IS IN AN OFF STATE. ADDITIONAL B, C, D IMDRF ANNEX CODES ADDED REFLECT THE ISSUE THAT OCCURRED WITH THE IESU ERBE DURING A DIFFERENT SURGERY, RESULTING IN THIS PATIENT'S NON-ROBOTIC SURGERY BEING POSTPONED. C20 - NO FINDINGS AVAILABLE AND D15 - CAUSE NOT ESTABLISHED REFER TO THE IMPACT FOR THE PATIENT IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21134614&pc=BZD | 2518422-2025-001920 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21134808&pc=BZD | 2518422-2025-001954 | 2023-02-13 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21134921&pc=BZD | 2518422-2025-001951 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Hypersensitivity/Allergic reaction; Nausea; Pulmonary Dysfunction; Vomiting; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, SKIN IRRITATION, HYPERSENSITIVITY, NAUSEA, VOMITING AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21135518&pc=BZD | 2518422-2025-001995 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21135578&pc=BZD | 2518422-2025-001992 | 2024-06-13 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded; Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY DUE TO SINUS CANCER. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: PREVIOUSLY THIS REPORT WAS REPORTED AS NON-UNO DUE TO ALLEGATION ON REPAIRED DEVICE. NOW THE DEVICE INFORMATION HAS BEEN UPDATED AND IT WAS DETERMINED THAT THE REPORT SHOULD BE REPORTED AS FINAL UNO REPORT. ALL MANDATORY SECTION HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136064&pc=BZD | 2518422-2025-002099 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LIVER DISEASE, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136231&pc=BZD | 2518422-2025-001925 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Inflammation; Nausea; Vomiting; Dizziness; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136389&pc=NVZ | 3008973940-2025-00341 | 2024-12-20 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-13 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH AN IMPLANTABLE PULSE GENERATOR (IPG) PASSED AWAY, AND THERE WAS AN ALLEGATION THAT THE DEVICE WAS RELATED. THERE WAS AN INQUIRY REGARDING IRREGULAR HEART ARRHYTHMIAS AT THE TIME OF THE PATIENT'S PASSING. THE CALLER SOUGHT INFORMATION ABOUT ANY ACTIVITY THE DEVICE SHOWED AT THE TIME OF THE PATIENT'S DEATH AND WAS ADVISED TO CONTACT THE HEART SPECIALIST MANAGING THE PATIENT'S CARE FOR DETAILED DEVICE REPORTS.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS RESPIRATORY FAILURE AND AIRWAY COMPROMISE AND THE DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IPG. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136531&pc=PZS | 3011468686-2025-00002 | 2024-04-11 04:00 | Death | MC3 INC. | 2025-01-13 05:00 | PZS | MC3 | Adverse Event Without Identified Device or Use Problem | Shock; Brain Injury; Multiple Organ Failure; Cerebral Edema | K180151 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SINGLE-SITE CANNULATION USING AN ALTERNATIVE VENOUS ACCESS SITES FOR DUAL-LUMEN EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION. THE TIME FRAME OF THIS STUDY WAS JANUARY 2011 AND APRIL 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: MC3 CRESCENT DUAL-LUMEN CANNULAE (24 FR AND 28 FR). DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSES OF DEATH WERE FAILURE TO RECOVER AND WITHDRAWAL OF THERAPY, MULTI- ORGAN FAILURE, PATIENT WILL, ANOXIC BRAIN INJURY, UNCLEAR CEREBRAL OEDEMA AND HAEMORRHAGIC SHOCK AFTER A LIVER BIOPSY. AMONG PATIENT ADVERSE EVENTS INCLUDED: BLEEDING AT THE CANNULATION SITE THAT EVENTUALLY REQUIRED BLOOD TRANSFUSION AND WAS MANAGED BY ADAPTING THE HEPARIN DOSAGE. ONE PATIENT SUFFERED FROM RECURRING ¿SUCK DOWN¿ EPISODES OF THE ECMO CANNULA, WHICH WAS MANAGED WITH CONTINUOUS VOLUME ADMINISTRATION. THE CANNULA WAS FIRST REPOSITIONED AND ULTIMATELY CHANGED TO A FEMORAL-JUGULAR CONFIGURATION. UNFORTUNATELY, THIS DID NOT REDUCE THE FREQUENCY OF ¿SUCK DOWN¿ EPISODES AND THE PATIENT HAD TO FURTHER BE MANAGED WITH ADDITIONAL VOLUME ADMINISTRATION. PRODUCT MALFUNCTIONS INCLUDED: CHANGE OF ECMO CONFIGURATION AND REPOSITIONING OF THE CANNULAE. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 70128 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70124 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70128 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70124 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70128 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA. G2: CITATION: AUTHORS: ARMIN-KAI SCHOEBERL, DAWID STAUDACHER, MITSUAKI KAWASHIMA. ALTERNATIVE VENOUS ACCESS SITES FOR DUAL-LUMEN EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION. INTERDISCIPLINARY CARDIOVASCULAR AND THORACIC SURGERY 38(4) 2024. 10.1093/ICVTS/IVAE060 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136568&pc=PZS | 3011468686-2025-00003 | 2024-04-11 04:00 | Death | MC3 INC. | 2025-01-13 05:00 | PZS | MC3 | Adverse Event Without Identified Device or Use Problem | Shock; Brain Injury; Multiple Organ Failure; Cerebral Edema | K180151 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SINGLE-SITE CANNULATION USING AN ALTERNATIVE VENOUS ACCESS SITES FOR DUAL-LUMEN EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION. THE TIME FRAME OF THIS STUDY WAS JANUARY 2011 AND APRIL 2022. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: MC3 CRESCENT DUAL-LUMEN CANNULAE (24 FR AND 28 FR). DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSES OF DEATH WERE FAILURE TO RECOVER AND WITHDRAWAL OF THERAPY, MULTI- ORGAN FAILURE, PATIENT WILL, ANOXIC BRAIN INJURY, UNCLEAR CEREBRAL OEDEMA AND HAEMORRHAGIC SHOCK AFTER A LIVER BIOPSY. AMONG PATIENT ADVERSE EVENTS INCLUDED: BLEEDING AT THE CANNULATION SITE THAT EVENTUALLY REQUIRED BLOOD TRANSFUSION AND WAS MANAGED BY ADAPTING THE HEPARIN DOSAGE. ONE PATIENT SUFFERED FROM RECURRING ¿SUCK DOWN¿ EPISODES OF THE ECMO CANNULA, WHICH WAS MANAGED WITH CONTINUOUS VOLUME ADMINISTRATION. THE CANNULA WAS FIRST REPOSITIONED AND ULTIMATELY CHANGED TO A FEMORAL-JUGULAR CONFIGURATION. UNFORTUNATELY, THIS DID NOT REDUCE THE FREQUENCY OF ¿SUCK DOWN¿ EPISODES AND THE PATIENT HAD TO FURTHER BE MANAGED WITH ADDITIONAL VOLUME ADMINISTRATION. PRODUCT MALFUNCTIONS INCLUDED: CHANGE OF ECMO CONFIGURATION AND REPOSITIONING OF THE CANNULAE. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 70124 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70124 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70128 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70124 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; PRODUCT ID 70128 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA. G2: CITATION: AUTHORS: ARMIN-KAI SCHOEBERL, DAWID STAUDACHER, MITSUAKI KAWASHIMA. ALTERNATIVE VENOUS ACCESS SITES FOR DUAL-LUMEN EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION. INTERDISCIPLINARY CARDIOVASCULAR AND THORACIC SURGERY 38(4) 2024. 10.1093/ICVTS/IVAE060 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136899&pc=OZD | 1220648-2025-25511 | 2024-12-26 05:00 | Death | ABIOMED, INC. | 2025-01-13 05:00 | OZD | IMPELLA CP | Positioning Failure | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, A PATIENT HAD MULTIPLE RUNS OF VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION REQUIRING MULTIPLE SHOCKS FROM PATIENT IMPLANTABLE CARDIOVERTER DEFIBRILLATOR. THE DOCTOR BELIEVED THE IMPELLA WAS TOO DEEP AND POTENTIALLY CAUSING ARRHYTHMIAS. THE IMPELLA CP WAS PULLED BACK AND BELIEVED TO BE IN A BETTER POSITION AWAY FROM HEART STRUCTURES. THERE WAS NO HARM TO THE PATIENT AND THE PATIENT REMAINS ON SUPPORT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EVENTUALLY EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED POSITIONING ISSUE AND VENTRICULAR TACHYCARDIA (VT) HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE POSITIONING ISSUE AND VT COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ ADDED-B2 SELECTED DEATH AND ADDED DATE OF DEATH, B5 PATIENT OUTCOME OF DEATH, D6B-DATE OF EXPLANT, H1-UPDATED TO DEATH, H6 IMPACT CODE 1802.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21137047&pc=MNS | 2518422-2024-074811 | 2023-11-29 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | MNS | BIPAP AUTOSV | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY AND HIS MOTHER WOULD LIKE TO BE TAKEN OFF ALL MAILING LISTS AND CORRESPONDENCE REGARDING HIS CASE WITH THE CLASS ACTION AND RECALL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21137458&pc=BZD | 2518422-2025-002212 | 2024-02-08 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: "THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE CUSTOMER CALLED IN BECAUSE SHE WOULD WANT TO KNOW THE EXACT NAME OF THE RECALLED DEVICE. THE PATIENT HAS PASSED AWAY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21137785&pc=OZD | 1220648-2025-25520 | 2023-06-21 04:00 | Death | ABIOMED, INC. | 2025-01-13 05:00 | OZD | IMPELLA CP | Malposition of Device | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ST ELEVATION MYOCARDIAL INFARCTION WAS ADMITTED TO THE HOSPITAL WHERE IT WAS DISCOVERED SHE HAD A TOTAL OCCLUSION OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY (LAD). THE PATIENT'S LAD WAS OPENED, THE PATIENT REMAINED STABLE. PRIOR TO LEAVING THE CATHETERIZATION LAB, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AND FIBRILLATION WHERE SHE WAS RESCUED BY HER AUTOMATIC IMPLANTABLE CARDIOVERTER DEFIBRILLATOR. THE PATIENT WAS TRANSFERRED TO THE ICU WHERE SHE CONTINUED TO DECLINE AND DECOMPENSATE OVERNIGHT. THE FAMILY WISHED TO ESCALATE CARE OF THE PATIENT, AND THE PHYSICIAN DECIDED TO PLACE THE IMPELLA TO WEAN THE INOTROPE SUPPORT AS A LAST EFFORT, AND AN IMPELLA CP DEVICE WAS IMPLANTED FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED UPON CONNECTING THE IMPELLA CP TO THE AUTOMATED IMPELLA CONTROLLER (AIC) WITH THE SERIAL NUMBER (B)(6), THERE WAS A MESSAGE FOR "OPTICAL SENSOR NOT SUPPORTED". THE PUMP WAS DISCONNECTED AND RECONNECTED SEVERAL TIMES, HOWEVER THIS DID NOT RESOLVE THE ISSUE. THE AIC WAS SWAPPED WITH A BACKUP UNIT, AND THE PUMP WAS SUCCESSFULLY CONNECTED. IT WAS NOTED AFTER IMPELLA SUPPORTED WAS INITIATED, THE PATIENT WAS IN GOOD AND STABLE CONDITION AND WAS BROUGHT BACK TO THE ICU. IN THE ICU, THE AIC ALARMED FOR PLACEMENT SIGNAL LOW. STAT ECHOCARDIOGRAPHY WAS ORDERED FOR THE PATIENT, AND IT WAS OBSERVED THE IMPELLA CP WAS TOO DEEP WITHIN THE VENTRICLE. THE PHYSICIAN WAS ABLE TO SUCCESSFULLY PULL THE IMPELLA BACK PER BEST PRACTICES, AND "GREAT" WAVEFORMS ON THE AIC WERE NOTED. AS REPORTED ON (B)(6) 2023, THE PATIENT CONTINUED TO DECLINE DESPITE MAXIMAL MECHANICAL AND MEDICAL SUPPORT WITH WORSENING LIVER AND RENAL FUNCTION. THERE WAS A NEED FOR ESCALATION OF VASOPRESSOR AND INOTROPIC SUPPORT ON THE MORNING OF (B)(6) 2023 WITH POOR PROGNOSIS. IT WAS NOTED THAT THE IMPELLA WAS PLACED AND FUNCTIONING WELL, HOWEVER, DUE TO MULTIPLE METABOLIC DERANGEMENTS THE PATIENT CONTINUED TO DECLINED. THE COLLECTIVE DECISION WAS MADE BY THE CARDIOLOGY AND PULMONARY TEAMS IN CONJUNCTION WITH THE FAMILY TO TRANSITION THE PATIENT TO DO NOT RESUSCITATE COMFORT CARE. THEREFORE, THE DECISION WAS MADE TO WITHDRAW MECHANICAL CIRCULATORY SUPPORT AND THE IMPELLA WAS REMOVED AT THE BEDSIDE BY DR. (B)(6). PATIENT TOLERATED THE REMOVAL WELL AND HEMOSTASIS ACHIEVED. THE PATIENT LATER EXPIRED. ALTHOUGH THE PATIENT'S PRESENTING CONDITION MAY BE MORE LIKELY HAVE IMPACTED THE EVENT THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR IN THE PATIENT'S DEMISE AFTER A MALPOSITION EVENT OF THE IMPELLA WHICH MAY HAVE CONTRIBUTED TO LOWER THAN EXPECTED SUPPORT. THE MALPOSITION OF THE IMPELLA MAY HAVE BEEN RESULTANT OF THE OPTICAL SENSOR ISSUE REPORTED ON THE FIRST AIC THAT WAS IN USE. Manufacturer Narrative: THIS IS ONE OF TWO MEDICAL DEVICE REPORTS ASSOCIATED WITH THE EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2024-25212 FOR THE AUTOMATED IMPELLA CONTROLLER WITH SN (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS: ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ CORRECT IMPELLA CATHETER POSITION: ¿FOR OPTIMAL POSITIONING OF THE IMPELLA CATHETER, THE INLET AREA OF THE CATHETER SHOULD BE 3.5 CM BELOW THE AORTIC VALVE ANNULUS AND WELL AWAY FROM PAPILLARY MUSCLE AND SUBANNULAR STRUCTURES. THE OUTLET AREA SHOULD BE WELL ABOVE THE AORTIC VALVE.¿ THIS REPORT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF HISTORICAL RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21137863&pc=BZD | 2518422-2025-002181 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Pancreatitis; Unspecified Gastrointestinal Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, PANCREATIC AND ESOPHAGEAL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21137921&pc=BZD | 2518422-2025-002239 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND KIDNEY CANCER RENAL CELL LEFT KIDNEY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21137939&pc=NGV | 2135147-2025-00269 | 2024-10-21 04:00 | Death | ABBOTT MEDICAL | 2025-01-13 05:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device; No Apparent Adverse Event | Angina; Stroke/CVA; Fever; Low Blood Pressure/ Hypotension; Sepsis; Obstruction/Occlusion; No Clinical Signs, Symptoms or Conditions | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT AN UNKNOWN AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT ON (B)(6) 2024. THE OCCLUDER WAS IMPLANTED. POST-IMPLANT THE OCCLUDER EMBOLIZED TO THE MITRAL VALVE. ATTEMPTS WERE MADE TO RETRIEVE THE OCCLUDER. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT EXPERIENCED MULTIPLE STROKES IN BOTH HEMISPHERES OF THE BRAIN THAT LED TO BLINDNESS, THE INABILITY TO MOVE, AND GENERAL UNAWARENESS. THE PATIENT WAS ADMITTED TO A LONG-TERM ACUTE CARE HOSPITAL (LTACH). ONE WEEK AFTER BEING DISCHARGED FROM THE LTACH THE PATIENT DEVELOPED A 104 FAHRENHEIT TEMPERATURE DUE TO SEPSIS. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL WHERE IT WAS OBSERVED THAT THE PATIENT'S BREATHING WAS LABORED WITH LOW OXYGEN. THE PATIENT'S BLOOD PRESSURE WAS ALSO DROPPING. THE PATIENT WAS STABILIZED AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU). THE PATIENT WAS MOVED TO COMFORT CARE AND PASSED AWAY 12 HOURS LATER ON (B)(6) 2024.Event Description: CORRECTION: PLEASE RETRACT THIS REPORT 2135147-2025-00269, THE SAME ISSUE WAS PREVIOUSLY REPORTED AS 2135147-2024-05656.Event Description: CORRECTION: PLEASE RETRACT THIS REPORT 2135147-2025-00269, THE SAME ISSUE WAS PREVIOUSLY REPORTED AS 2135147-2024-05656. Manufacturer Narrative: H6 REMOVED HEALTH EFFECT - CLINICAL CODE 1770 STROKE/CVA, H6 REMOVED HEALTH EFFECT - CLINICAL CODE 2422 OBSTRUCTION/OCCLUSION, H6 REMOVED HEALTH EFFECT - CLINICAL CODE 1858 FEVER, H6 REMOVED HEALTH EFFECT - CLINICAL CODE 2067 SEPSIS, H6 REMOVED HEALTH EFFECT - CLINICAL CODE 1710 ANGINA, H6 REMOVED HEALTH EFFECT - CLINICAL CODE 1914 LOW BLOOD PRESSURE/HYPOTENSION, H6 REMOVED HEALTH EFFECT - IMPACT CODE 1802 DEATH, H6 REMOVED HEALTH EFFECT - IMPACT CODE 4641 UNEXPECTED MEDICAL INTERVENTION, H6 REMOVED HEALTH EFFECT - IMPACT CODE HOSPITALIZATION OR PROLONGED HOSPITALIZATION, H6 REMOVED MEDICAL DEVICE PROBLEM CODE 1395 MIGRATION OR EXPULSION OF DEVICE. CORRECTION: PLEASE RETRACT THIS REPORT 2135147-2025-00269, THE SAME ISSUE WAS PREVIOUSLY REPORTED AS 2135147-2024-05656. Manufacturer Narrative: CORRECTION: PLEASE RETRACT THIS REPORT 2135147-2025-00269, THE SAME ISSUE WAS PREVIOUSLY REPORTED AS 2135147-2024-05656. Manufacturer Narrative: B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138037&pc=NPT | 2025587-2025-00319 | 2024-12-10 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-13 05:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Apnea; Dyspnea; Renal Failure; Discomfort; Confusion/ Disorientation; Lethargy; Decreased Sensitivity; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED DURING A TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT (TAVR) PROCEDURE, PRIOR TO INSERTING THE DELIVERY CATHETER SYSTEM (DCS), THE PATIENT DEVELOPED A COMPLETE HEART BLOCK WITH NO ESCAPE RHYTHM. TRANSVENOUS CARDIAC PACING (TVP) WAS PLACED THROUGH AN ACCESS AT THE LEFT FEMORAL VEIN AND LEFT IN PLACE. VASCULAR ACCESS FOR THE DCS WAS ACHIEVED VIA THE RIGHT COMMON FEMORAL ARTERY. THE DCS AND LOADED VALVE WERE INSERTED INTO THE PATIENT AND NAVIGATED TO THE HEART. THE VALVE WAS IMPLANTED IN THE AORTIC POSITION. UPON EVALUATION, ATRIO-VENTRICULAR CONDUCTION RECOVERED. FOLLOWING THE TAVR, THE PATIENT WAS LETHARGIC FROM THE ANESTHESIA. THE PATIENT'S DAUGHTER AND MEDICAL POWER OF ATTORNEY (MPOA) REFUSED TO SIGN CONSENT FOR PERMANENT PACEMAKER IMPLANT AND REQUESTED FOR THE PATIENT TO MAKE THE DECISION. THE MPOA CHOSE TO WAIT 24-HOURS POST-ANESTHESIA FOR THE PATIENT TO MAKE A DECISION. OVER THE FOLLOWING 24-HOURS, THE PATIENT WAS TO BE MONITORED AND THEN RE-EVALUATION FOR PACEMAKER VERSUS A 14-DAY PORTABLE ELECTROCARDIOGRAM CARDIAC (HOLTER) MONITOR IN LIGHT OF RECOVERY OF ATRIAL VENTRICULAR CONDUCTION. THE PATIENT DIED TEN DAYS FOLLOWING THE TAVR DUE TO MULTIPLE ORGAN DYSFUNCTION SYNDROME. A PERMANENT PACEMAKER WAS NOT IMPLANTED. NO FURTHER DETAILS WERE PROVIDED RELATED TO THE PATIENT'S DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PRIOR TO THE IMPLANT OF THE VALVE A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. ONE DAY FOLLOWING THE IMPLANT OF THE VALVE, A PERMANENT PACEMAKER WAS IMPLANTED.Event Description: IT WAS REPORTED DURING A TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT (TAVR) PROCEDURE, PRIOR TO INSERTING THE DELIVERY CATHETER SYSTEM (DCS), THE PATIENT DEVELOPED A COMPLETE HEART BLOCK WITH NO ESCAPE RHYTHM. TRANSVENOUS CARDIAC PACING (TVP) WAS PLACED THROUGH AN ACCESS AT THE LEFT FEMORAL VEIN AND LEFT IN PLACE. VASCULAR ACCESS FOR THE DCS WAS ACHIEVED VIA THE RIGHT COMMON FEMORAL ARTERY. THE DCS AND LOADED VALVE WERE INSERTED INTO THE PATIENT AND NAVIGATED TO THE HEART. THE VALVE WAS IMPLANTED IN THE AORTIC POSITION. UPON EVALUATION, ATRIO-VENTRICULAR CONDUCTION RECOVERED. FOLLOWING THE TAVR, THE PATIENT WAS LETHARGIC FROM THE ANESTHESIA. THE PATIENT'S DAUGHTER AND MEDICAL POWER OF ATTORNEY (MPOA) REFUSED TO SIGN CONSENT FOR PERMANENT PACEMAKER IMPLANT AND REQUESTED FOR THE PATIENT TO MAKE THE DECISION. THE MPOA CHOSE TO WAIT 24-HOURS POST-ANESTHESIA FOR THE PATIENT TO MAKE A DECISION. OVER THE FOLLOWING 24-HOURS, THE PATIENT WAS TO BE MONITORED AND THEN RE-EVALUATION FOR PACEMAKER VERSUS A 14-DAY PORTABLE ELECTROCARDIOGRAM CARDIAC (HOLTER) MONITOR IN LIGHT OF RECOVERY OF ATRIAL VENTRICULAR CONDUCTION. THE PATIENT DIED TEN DAYS FOLLOWING THE TAVR DUE TO MULTIPLE ORGAN DYSFUNCTION SYNDROME. A PERMANENT PACEMAKER WAS NOT IMPLANTED. NO FURTHER DETAILS WERE PROVIDED RELATED TO THE PATIENT'S DEATH. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT UPON REMOVAL OF THE CATHETER AT THE LEFT FEMORAL ARTERIAL LINE, A LARGE HEMATOMA BEGAN TO FORM ABOVE THE REMOVAL SITE. MANUAL PRESSURE WAS APPLIED AND HELD FOR 30 MINUTES; THE SITE BECAME SOFT. A PRESSURE DRESSING AND GAUZE WERE APPLIED TO THE AREA AND THEN A SANDBAG WAS PLACED FOR TWO HOURS. THE HEMATOMA RESOLVED THE SAME DAY. OF NOTE, THE ARTERIAL ACCESS FOR INSERTION OF THE DCS WAS ACHIEVED VIA RIGHT COMMON FEMORAL ARTERY. ADDITIONAL INFORMATION WAS RECEIVED THAT PRIOR TO THE IMPLANT OF THE VALVE A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. ONE DAY FOLLOWING THE IMPLANT OF THE VALVE, A PERMANENT PACEMAKER WAS IMPLANTED. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY A PERMANENT PACEMAKER WAS NOT IMPLANTED ON THE FIRST POST-OPERATIVE DAY AS PREVIOUSLY NOTED. INSTEAD, A TEMPORARY PACEMAKER WAS REPLACED.Event Description: IT WAS REPORTED DURING A TRANSCATHETER AORTIC BIOPROSTHETIC VALVE REPLACEMENT (TAVR) PROCEDURE, PRIOR TO INSERTING THE DELIVERY CATHETER SYSTEM (DCS), THE PATIENT DEVELOPED A COMPLETE HEART BLOCK WITH NO ESCAPE RHYTHM. TRANSVENOUS CARDIAC PACING (TVP) WAS PLACED THROUGH AN ACCESS AT THE LEFT FEMORAL VEIN AND LEFT IN PLACE. VASCULAR ACCESS FOR THE DCS WAS ACHIEVED VIA THE RIGHT COMMON FEMORAL ARTERY. THE DCS AND LOADED VALVE WERE INSERTED INTO THE PATIENT AND NAVIGATED TO THE HEART. THE VALVE WAS IMPLANTED IN THE AORTIC POSITION. UPON EVALUATION, ATRIO-VENTRICULAR CONDUCTION RECOVERED. FOLLOWING THE TAVR, THE PATIENT WAS LETHARGIC FROM THE ANESTHESIA. THE PATIENT'S DAUGHTER AND MEDICAL POWER OF ATTORNEY (MPOA) REFUSED TO SIGN CONSENT FOR PERMANENT PACEMAKER IMPLANT AND REQUESTED FOR THE PATIENT TO MAKE THE DECISION. THE MPOA CHOSE TO WAIT 24-HOURS POST-ANESTHESIA FOR THE PATIENT TO MAKE A DECISION. OVER THE FOLLOWING 24-HOURS, THE PATIENT WAS TO BE MONITORED AND THEN RE-EVALUATION FOR PACEMAKER VERSUS A 14-DAY PORTABLE ELECTROCARDIOGRAM CARDIAC (HOLTER) MONITOR IN LIGHT OF RECOVERY OF ATRIAL VENTRICULAR CONDUCTION. THE PATIENT DIED TEN DAYS FOLLOWING THE TAVR DUE TO MULTIPLE ORGAN DYSFUNCTION SYNDROME. A PERMANENT PACEMAKER WAS NOT IMPLANTED. NO FURTHER DETAILS WERE PROVIDED RELATED TO THE PATIENT'S DEATH. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT UPON REMOVAL OF THE CATHETER AT THE LEFT FEMORAL ARTERIAL LINE, A LARGE HEMATOMA BEGAN TO FORM ABOVE THE REMOVAL SITE. MANUAL PRESSURE WAS APPLIED AND HELD FOR 30 MINUTES; THE SITE BECAME SOFT. A PRESSURE DRESSING AND GAUZE WERE APPLIED TO THE AREA AND THEN A SANDBAG WAS PLACED FOR TWO HOURS. THE HEMATOMA RESOLVED THE SAME DAY. OF NOTE, THE ARTERIAL ACCESS FOR INSERTION OF THE DCS WAS ACHIEVED VIA RIGHT COMMON FEMORAL ARTERY. ADDITIONAL INFORMATION WAS RECEIVED THAT PRIOR TO THE IMPLANT OF THE VALVE A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. ONE DAY FOLLOWING THE IMPLANT OF THE VALVE, A PERMANENT PACEMAKER WAS IMPLANTED. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY A PERMANENT PACEMAKER WAS NOT IMPLANTED ON THE FIRST POST-OPERATIVE DAY AS PREVIOUSLY NOTED. INSTEAD, A TEMPORARY PACEMAKER WAS REPLACED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE CHB DEVELOPED AFTER THE PRE-IMPLANT BALLOON DILATION, PRIOR TO THE INSERTION OF DCS. THE CHB WAS RELATED TO THE UNDERLYING INTRINSIC BIFASCICULAR BLOCK PRIOR TO THE PROCEDURE AND PROBABLY DUE TO THE PRE-IMPLANT BALLOON VALVULOPLASTY. THE PATIENT HAD A HISTORY OF RENAL TRANSPLANT, WAS TAKING CHRONIC IMMUNOSUPPRESSIVE THERAPY, ALONG WITH A HOST OF OTHER CHRONIC AND SEVERE MEDICAL COMORBIDITIES. THE PATIENT ORIGINALLY PRESENTED TO THE HOSPITAL FOR SYMPTOMATIC AORTIC STENOSIS AND A TAVR PROCEDURE. THE PATIENT HAD A COMPLICATED POST-OPERATIVE COURSE; WAS NOTED TO BE INCREASINGLY DELIRIOUS AND PROGRESSED TO A STATE OF TERMINAL DELIRIUM OVER THE COURSE OF APPROXIMATELY 1 WEEK. THROUGH THAT SAME TIME, THE PATIENT WAS NOTED TO BE IN RENAL FAILURE, SUSPECTED ACUTE TUBULAR NECROSIS (ATN), AND INCREASINGLY SHORT OF BREATH. THE PHYSICIAN DIAGNOSED FAILURE TO THRIVE; THE PATIENT WAS REMAINED UNCOMFORTABLE, DELIRIOUS, AND WITHOUT A SIGNIFICANT HOPE FOR ANY GOOD OUTCOMES DESPITE MOST AGGRESSIVE EFFORTS. FOR THIS REASON, THE PATIENT'S FAMILY WAS GIVEN THE OPTIONS FOR TRANSITIONING HIM TO COMFORT MEASURES ONLY WITH THE THOUGHT OF EVENTUAL LY LEADING INTO HOSPICE CARE. COMFORT MEASURES WERE EMPLOYED NINE DAYS AFTER THE TAVR AND THE PATIENT WAS TRANSFERRED FOR MORE ATTENTIVE SYMPTOMATIC CARE. IN THE MORNING OF THE NEXT DAY, THE PATIENT WAS EVALUATED AND WAS FOUND TO BE VERY COMFORTABLE AND RESTING. IN THE AFTERNOON OF THE SAME DAY, A NURSE ASSESSED THE PATIENT AND DETERMINED THE PATIENT WAS APNEIC AND UNRESPONSIVE. THE PATIENT WAS EXAMINED BY A PHYSICIAN AND PRONOUNCED DECEASED. PER THE PHYSICIAN, THE PATIENT'S MORTALITY WAS NOT RELATED TO THE DCS OR THE VALVE. THE MOST LIKELY CAUSE OF MORTALITY WAS DUE TO THE TAVR PROCEDURE AND THE MULTIPLE, UNDERLYING COMORBIDITIES. Manufacturer Narrative: UPDATED DATA: B5. A FIFTH PARAGRAPH WAS ADDED. B7. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B4. A FOURTH PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A1 - PT. IDENTIFIER A2 - DATE OF BIRTH B5 D - SUSPECT MEDICAL DEVICE INFORMATION H6 - ANNEX F. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138141&pc=DSP | 3010532612-2025-00049 | 2024-12-16 05:00 | Death | ARROW INTERNATIONAL LLC | 2025-01-13 05:00 | DSP | ARROW ULTRA 8 IAB: 8FR 30CC | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Respiratory Acidosis | K000729 |  | 1 | Event Description: IT WAS REPORTED "SICU ENDORSED BUSTED BALLOON AROUND 3:00 AM AND WAS INSTRUCTED NOT TO TURN ON THE MACHINE 7:00 AM ON-CALL DUTY PERFUSIONIST ARRIVED TO DISCONNECT THE MACHINE AND DOCTORS DECIDED TO PERFORM A REINSERTION AS PATIENT IS BALLOON DEPENDENT. PERFUSIONIST OBSERVED BLOOD AT HELIUM TUBING AND ENCOURAGED REMOVAL OF BURST CATHETER. SICU NURSES NOTED THAT THE BALLOON ONLY TURN OFF WHEN BLOOD CAN ALREADY BE SEEN FOR THE FIRST TIME, INCIDENT OCCURRED ON (B)(6) 2024 AT PHILIPPINE HEART CENTER SICU BED-SIDE". ADDITIONAL INFORMATION STATES THAT THE DEVICE WAS REMOVED IN ITS ENTIRETY AND REPLACED. THE SECOND IAB CATHETER WAS INSERTED INTO A DIFFERENT INSERTION SITE. AT THE TIME OF THIS REPORT, THE CUSTOMER HAS NOT RETURNED OUR REQUESTS FOR ADDITIONAL INFORMATION.Event Description: IT WAS REPORTED "SICU ENDORSED BUSTED BALLOON AROUND 3:00 AM AND WAS INSTRUCTED NOT TO TURN ON THE MACHINE 7:00 AM ON-CALL DUTY PERFUSIONIST ARRIVED TO DISCONNECT THE MACHINE AND DOCTORS DECIDED TO PERFORM A REINSERTION AS PATIENT IS BALLOON DEPENDENT. PERFUSIONIST OBSERVED BLOOD AT HELIUM TUBING AND ENCOURAGED REMOVAL OF BURSTED CATHETER. SICU NURSES NOTED THAT THE BALLOON ONLY TURN OFF WHEN BLOOD CAN ALREADY BE SEEN FOR THE FIRST TIME, INCIDENT OCCURRED ON (B)(6) 2024 AT (B)(6) CENTER BED-SIDE". ADDITIONAL INFORMATION STATES THAT THE DEVICE WAS REMOVED IN ITS ENTIRETY AND REPLACED. THE SECOND IAB CATHETER WAS INSERTED INTO A DIFFERENT INSERTION SITE. AT THE TIME OF THIS REPORT, THE CUSTOMER HAS NOT RETURNED OUR REQUESTS FOR ADDITIONAL INFORMATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: QN#(B)(4). THE REPORTED COMPLAINT FOR "BLOOD AT HELIUM TUBING" IS NOT ABLE TO BE CONFIRMED. THE PRODUCT WAS NOT RETURNED FOR INVESTIGATION. THE ROOT CAUSE OF THE COMPLAINT IS UNDETERMINED. NO FURTHER ACTION REQUIRED AT THIS TIME. THE REPORTED COMPLAINT WILL BE MONITORED FOR ANY DEVELOPING TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138550&pc=BZD | 2518422-2025-002218 | 2024-06-13 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO BIPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138551&pc=BZD | 2518422-2025-100474 | 2025-01-09 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138715&pc=BZD | 2518422-2025-002248 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Respiratory Failure; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, RESPIRATORY FAILURE WHILE USING THE DEVICE. THE PATIENT DIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138904&pc=BZD | 2518422-2025-100483 | 2025-01-09 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE RENTAL DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.Event Description: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO CORRECT THE PREVIOUS REPORT AND PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. AS A RESULT OF A THIRD-PARTY INSPECTION, NO FAULTS WERE FOUND, AND THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO CORRECT THE PREVIOUS REPORT AND PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. AS A RESULT OF A THIRD-PARTY INSPECTION, NO FAULTS WERE FOUND, AND THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21139103&pc=DSQ | 2916596-2025-00018 | 2024-12-14 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEMORRHAGIC STROKE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: E1: REPORTER PHONE NUMBER WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH AND BLEEDING. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21139125&pc=DSQ | 2916596-2025-00200 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT IN THE RESEARCH ARTICLE ¿SEVERE PERIPHERAL ARTERIAL DISEASE IS ASSOCIATED WITH INCREASED LIMB COMPLICATIONS AND STROKE RATE POST LVAD IMPLANTATION, BUT NOT ASSOCIATED WITH INCREASED MORTALITY¿, THAT THE HEARTMATE 3 MAY BE ASSOCIATED WITH VASCULAR COMPLICATION, STROKE, RENAL FAILURE AND DEATH. THE PURPOSE OF THIS STUDY WAS TO DETERMINE IF SEVERE PERIPHERAL ARTERIAL DISEASE (PAD) WAS ASSOCIATED WITH AN INCREASED RATE OF COMPLICATIONS AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. 198 PATIENTS WHO UNDERWENT HEARTMATE 3 IMPLANTATION FROM 2015 TO JULY 2023 WERE INCLUDED IN THE STUDY. PREOPERATIVE ARTERIAL BRACHIAL INDEX AND TOE PRESSURES DETERMINED ON VASCULAR ULTRASOUND WERE COLLECTED. SEVERE PAD WAS DEFINED AS TOE PRESSURE < 50 MMHG OR EITHER INFLOW OR OUTFLOW DISEASE. 37 PATIENTS (19%) DID NOT HAVE PREOPERATIVE ULTRASONOGRAPHIC EVALUATION AND WERE COUNTED IN THE NON-SEVERE PAD CATEGORY FOR THE ANALYSIS. THE PRIMARY OUTCOME WAS VASCULAR COMPLICATION (DRY GANGRENE OR CRITICAL LIMB ISCHEMIA). SECONDARY OUTCOMES INCLUDED STROKE, BLEEDING, RENAL FAILURE, DEATH DURING HOSPITALIZATION AND DEATH AT ONE YEAR. OF THE 198 PATIENTS STUDIED, 22 (11%) HAD SEVERE PAD. PATIENTS WITH SEVERE PAD WERE FOUR TIMES MORE LIKELY TO DEVELOP A VASCULAR COMPLICATION (18% VS 4%, P< 0.001) AND TWICE AS LIKELY TO DEVELOP A STROKE (4% VS 2%, P < 0.001). THERE WERE NO DIFFERENCES IN OTHER IN-HOSPITAL COMPLICATIONS OR MORTALITY AT ONE YEAR (84% VS 85%, P = 0.86). ADDITIONAL COMPLICATIONS INCLUDED BLEEDING, RENAL FAILURE AND DEATH DURING THE IMPLANT HOSPITALIZATION. WHILE STROKE AND VASCULAR COMPLICATIONS ARE HIGHER THAN PATIENTS WITHOUT SEVERE PAD, IT MAY BE AN ACCEPTABLE RISK FOR SOME PATIENTS AND CENTERS, ESPECIALLY WITH THE INTENT OF DESTINATION THERAPY. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. WILLIAMS, M., TASOUDIS, P., SUKHAVASI, A., BUCKERIDGE, S., MOSTERTZ, W., BYKU, M., & MERLO, A. (2024). SEVERE PERIPHERAL ARTERIAL DISEASE IS ASSOCIATED WITH INCREASED LIMB COMPLICATIONS AND STROKE RATE POST LVAD IMPLANTATION, BUT NOT ASSOCIATED WITH INCREASED MORTALITY. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S286. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1231. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WAS NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21139406&pc=DSQ | 2916596-2025-00022 | 2023-08-02 04:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PERSISTENT EXIT SITE INFECTION AND BACTEREMIA. THE PATIENT WAS ADMITTED TO HOSPICE AND THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF PER FAMILY WISHES ON (B)(6) 2024. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED.Event Description: THE PATIENT WAS ADMITTED TO HOSPICE AND THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF PER FAMILY WISHES ON (B)(6) 2023. THE PATIENT PASSED AWAY ON (B)(6) 2023. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION B5 CORRECTION. MANUFACTURER¿S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21139465&pc=DSQ | 3007042319-2025-00056 | 2025-01-06 05:00 | Death | HEARTWARE, INC. | 2025-01-13 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Stopped | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY AND CANCER, WHO HAD A VENTRICULAR ASSIST DEVICE IMPLANTED, WAS IN A TERMINAL CONDITION LEADING TO DEATH. THE FAMILY DID NOT ACCEPT THE EXPLANTATION OF THE VENTRICULAR ASSIST DEVICE AFTER THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY AND CANCER, WHO HAD A VENTRICULAR ASSIST DEVICE (VAD) IMPLANTED, WAS IN A TERMINAL CONDITION LEADING TO DEATH, AND A VAD STOP. THE FAMILY DID NOT ACCEPT THE EXPLANTATION OF THE VENTRICULAR ASSIST DEVICE AFTER THE PATIENT'S DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION TO THE EVENT DESCRIPTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS IN A TERMINAL CONDITION LEADING TO DEATH AND A VAD STOP; HOWEVER, THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. LOG FILE ANALYSIS COULD NOT BE PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED VAD STOP EVENT COULD NOT BE CONFIRMED. APPLICABLE RISK DOCUMENTATION AND EXPERIENCE WITH EVENTS OF SIMILAR CIRCUMSTANCES WERE CONSIDERED; A POSSIBLE ROOT CAUSE OF THE REPORTED VAD STOP EVENT CAN BE ATTRIBUTED BUT NOT LIMITED TO A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER, A CONTROLLER FAILURE, A PUMP FAILURE, OR THE VAD BEING STOPPED FOLLOWING THE PATIENT'S DEATH. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21139845&pc=BZD | 2518422-2025-100489 | 2024-12-17 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Chronic Obstructive Pulmonary Disease (COPD) | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY DUE TO CHRONIC LUNG DISEASE. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21140154&pc=DSQ | 2916596-2025-00162 | 2022-12-31 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT IN THE RESEARCH ARTICLE TITLED ¿CYSTATIN C - BASED ESTIMATED GLOMERULAR FILTRATION RATE: A PREDICTIVE METRIC FOR LONG-TERM HEARTMATE 3 SURVIVAL¿, THE HEARTMATE 3 (HM3) MAY BE RELATED TO DEATH. ALL HM3 PATIENTS IMPLANTED FROM FEB2016 TO DEC2022 WERE STUDIED IN THE SINGLE-CENTER STUDY. THE STUDY AIMED TO IDENTIFY PRE-IMPLANT PREDICTORS OF LONG-TERM MORTALITY IN A COHORT OF PATIENTS WITH AVAILABLE CYSC BECAUSE CYSTATIN C (CYSC)- BASED EGFR (EGFRCYSC) WAS INDEPENDENT FROM MUSCLE MASS AND NUTRITIONAL STATUS AND MAY PROVIDE A MORE ACCURATE MEASURE OF KIDNEY FUNCTION. USING A STEPWISE SELECTION PROCESS, POTENTIAL PRE-IMPLANT PREDICTORS WERE IDENTIFIED FROM A SET OF 32 VARIABLES, INCLUDING SERUM CREATININE (SCR)- BASED ESTIMATED GLOMERULAR FILTRATION RATE (EGFR) AND EGFRCYSC. AMONG 192 HEARTMATE 3 PATIENTS, DEATH OCCURRED IN 39 (20%) PATIENTS OVER A MEDIAN FOLLOW-UP OF 2 YEARS. EGFRCYSC EMERGED AS A STRONG INDICATOR OF MORTALITY IN THE COHORT OF HM3 PATIENTS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. COLUMBIA UNIVERSITY IRVING MEDICAL CENTER. MONDELLINI, G., PINSINO, A., HYNDS, M., VINOGRADSKY, A., LADANYI, A., SAYER, G., KAKU, Y., URIEL, N., TAKEDA, K., HUSAIN, S., MOHAN, S., COLOMBO, P., & YUZEFPOLSKAYA, M. (2024). CYSTATIN C- BASED ESTIMATED GLOMERULAR FILTRATION RATE: A PREDICTIVE METRIC FOR LONG-TERM HEARTMATE 3 SURVIVAL. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S171. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21140236&pc=CBK | MW5164583 | 2022-12-11 05:00 | Death | PHILIPS / RESPIRONICS, INC. | 2025-01-13 05:00 | CBK | TRILOGY VENTILATOR 100/200 |  | Cardiac Arrest |  |  | 1 | Event Description: A 14 YR OLD DISABLED CHILD USING TRILOGY VENTILATOR HAD A PLUGGED TRACH FROM THICK LUNG SECRETIONS IN THE VEHICLE (SO ON BATTERY) THE VENT DID NOT ALARM TO ALERT OF THE PLUG AND AS SUCH THE PATIENT CODED. SHE WAS BROUGHT BACK AND THEN PROMPTLY UPON PUTTING HER BACK ON THE VENT, THE SAME ISSUE HAPPENED BECAUSE WE HAD NO WAY TO KNOW IT WAS A PLUG WITH THE VENT NOT ALARMING. SHE WAS AGAIN BROUGHT BACK AND THE DAMAGE SUSTAINED DURING RESUSCITATION EFFORTS EVENTUALLY ENDED IN DEATH. SHE WOULD STILL BE HERE IF THE VENT ALARMED FOR THE PLUG AND WE COULD HAVE ATTEMPTED SUCTIONING. WE HAD NO IDEA THAT SUCTIONING WAS NEEDED BECAUSE THE VENT DID "WAS FINE" IF REQUESTED WE CAN GET THIS INFORMATION DIRECT FROM THE HOSPITAL OR SIGN A RELEASE FORM FOR YOU TO GATHER RECORDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21141568&pc=NRY | 2029214-2025-00117 | 2023-08-12 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-13 05:00 | NRY | UNKNOWN SOLITAIRE | Adverse Event Without Identified Device or Use Problem | Nervous System Injury |  |  | 1 | Event Description: BHOGAL, S., KALLUR, A., MERDLER, I., BEN-DOR, I., DEVINENI, A., HASHIM, H. D., BERNARDO, N. L., ROGERS, T., WERMERS, J. P., SATLER, L. F., GARCIA-GARCIA, H. M., & WAKSMAN, R. (2023). ASPIRATION THROMBECTOMY WITH AND WITHOUT CANGRELOR DURING PERCUTANEOUS CORONARY INTERVENTION. THE AMERICAN JOURNAL OF CARDIOLOGY, 209, 89¿91. HTTPS://DOI.ORG/10.1016/J.AMJCARD.2023.08.070. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A SINGLE-CENTER STUDY OF PATIENTS WHO UNDERWENT THROMBECTOMY PROCEDURES FROM NOVEMBER 2013 TO DECEMBER 2021. A TOTAL OF 588 PATIENTS WITH ASPIRATION THROMBECTOMY WERE IDENTIFIED, OF WHOM 202 RECEIVED CANGRELOR AND 386 DID NOT (CONTROL GROUP). MULTIPLE MANUFACTURERS' DEVICES WERE USED IN THE PROCEDURES, INCLUDING THE MEDTRONIC NEUROVASCULAR DEVICE, SOLITAIRE STENT RETRIEVER. IT WAS NOT SPECIFIED WHICH DEVICES WERE USED IN EACH CASE. NO DEVICE MALFUNCTION WAS REPORTED IN THE ARTICLE. THERE WAS NO DIFFERENCE IN IN-HOSPITAL ALL-CAUSE MORTALITY BETWEEN THE GROUPS (THOSE TREATED WITH CANGRELOR 1.5% VS CONTROL GROUP 1.1%, APPROXIMATELY 7 PATIENTS TOTAL). NO PATIENT CAUSE OF DEATH OR RELATIONSHIP TO ANY DEVICE OR PROCEDURE WAS SPECIFIED. ADDITIONAL ADVERSE EVENTS REPORTED IN THE ARTICLE INCLUDED: HEMATOMA RATE WAS 2.6% IN THE CONTROL GROUP (APPROXIMATELY 10 PATIENTS). 0 PATIENTS IN THE CANGRELOR GROUP EXPERIENCED HEMATOMA. HOWEVER, THE RATE OF BLOOD TRANSFUSION WAS 4% IN THE CANGRELOR GROUP VS 1.8% IN THE CONTROL GROUP, OR APPROXIMATELY 15 PATIENTS TOTAL. ISCHEMIC ENDPOINTS, INCLUDING IN-STENT THROMBOSIS, RATES WERE 0.5% VS 1.3% RESPECTIVELY, APPROXIMATELY 6 PATIENTS TOTAL. STROKE RATES WERE 3.5% VS 2.8%, RESPECTIVELY, APPROXIMATELY 18 PATIENTS TOTAL. Manufacturer Narrative: THIS REPORT IS SUBMITTED FOR THE PATIENT DEATHS NOTED IN THE LITERATURE ARTICLE. OTHER ADVERSE EVENTS/SERIOUS INJURIES REPORTED IN THE ARTICLE WILL BE REPORTED SEPARATELY. G3. PMA/501K NUMBER IS UNKNOWN AS THE MODEL INFORMATION FOR THE DEVICES ARE NOT AVAILABLE IN THE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21141847&pc=NEY | 3006451981-2025-00007 | 2024-01-31 05:00 | Death | COVIDIEN MEDICAL PRODUCTS | 2025-01-13 05:00 | NEY | UNKNOWN EMPRINT ANTENNA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ASSESSED 394 CASES OF LIVER MICROWAVE ABLATION (MWA) TO TREAT COLORECTAL LIVER METASTASIS BETWEEN 2006 AND 2022. CONCOMITANT PROCEDURES INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, AND COLECTOMY WERE PERFORMED IN 63% OF CASES. MICROWAVE ABLATIONS (MWA) WAS PERFORMED USING THE COMPETITOR SYSTEMS. DEATH OCCURRED IN ONE PATIENT AT HOME DUE TO UNKNOWN REASONS ON POST-OP DAY 20.Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ASSESSED 394 CASES OF LIVER MICROWAVE ABLATION (MWA) TO TREAT COLORECTAL LIVER METASTASIS BETWEEN 2006 AND 2022. CONCOMITANT PROCEDURES INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, AND COLECTOMY WERE PERFORMED IN 63% OF CASES. MICROWAVE ABLATIONS (MWA) WAS PERFORMED USING THE EVIDENT, EMPRINT, EMPRINT SX, EMPRINT HP, OR COMPETITOR SYSTEMS. DEATH OCCURRED IN ONE PATIENT AT HOME DUE TO UNKNOWN REASONS ON POST-OP DAY 20. ADVERSE EVENT WAS NOT RELATED TO THE DEVICE BUT RELATED TO THE PROCEDURE. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN) WELLS, ALEXANDRA, ET.AL., 2024, OUTCOMES AFTER SURGICAL MICROWAVE ABLATION FOR THE .TREATMENT OF COLORECTAL LIVER METASTASIS. AMERICAN COLLEGE OF SURGEONS. ISSN 1879-1190/22, VOL. 239, NO. 3, SEPTEMBER 2024 HTTPS://DOI.ORG/10.1097/XCS.0000000000001097 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H6 (RFR) ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21142147&pc=NEY | 3004962788-2025-00012 | 2024-01-31 05:00 | Death | COVIDIEN LP - SUPERDIMENSION INC | 2025-01-13 05:00 | NEY | UNKNOWN EMPRINT GENERATOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ASSESSED 394 CASES OF LIVER MICROWAVE ABLATION (MWA) TO TREAT COLORECTAL LIVER METASTASIS BETWEEN 2006 AND 2022. CONCOMITANT PROCEDURES INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, AND COLECTOMY WERE PERFORMED IN 63% OF CASES. MICROWAVE ABLATIONS (MWA) WAS PERFORMED USING THE COMPETITOR SYSTEMS. DEATH OCCURRED IN ONE PATIENT AT HOME DUE TO UNKNOWN REASONS ON POST-OP DAY 20.Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ASSESSED 394 CASES OF LIVER MICROWAVE ABLATION (MWA) TO TREAT COLORECTAL LIVER METASTASIS BETWEEN 2006 AND 2022. CONCOMITANT PROCEDURES INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, AND COLECTOMY WERE PERFORMED IN 63% OF CASES. MICROWAVE ABLATIONS (MWA) WAS PERFORMED USING THE EVIDENT, EMPRINT, EMPRINT SX, EMPRINT HP, OR COMPETITOR SYSTEMS. DEATH OCCURRED IN ONE PATIENT AT HOME DUE TO UNKNOWN REASONS ON POST-OP DAY 20. ADVERSE EVENT WAS NOT RELATED TO THE DEVICE BUT RELATED TO THE PROCEDURE. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN) WELLS, ALEXANDRA, ET.AL., 2024, OUTCOMES AFTER SURGICAL MICROWAVE ABLATION FOR THE TREATMENT OF COLORECTAL LIVER METASTASIS. AMERICAN COLLEGE OF SURGEONS. ISSN 1879-1190/22, VOL. 239, NO. 3, SEPTEMBER 2024 HTTPS://DOI.ORG/10.1097/XCS.0000000000001097. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H6 (RFR). ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21142151&pc=NEY | 3004962788-2025-00011 | 2024-01-31 05:00 | Death | COVIDIEN LP - SUPERDIMENSION INC | 2025-01-13 05:00 | NEY | UNKNOWN EMPRINT GENERATOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ASSESSED 394 CASES OF LIVER MICROWAVE ABLATION (MWA) TO TREAT COLORECTAL LIVER METASTASIS BETWEEN 2006 AND 2022. CONCOMITANT PROCEDURES INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, AND COLECTOMY WERE PERFORMED IN 63% OF CASES. MICROWAVE ABLATIONS (MWA) WAS PERFORMED USING THE COMPETITOR SYSTEMS. DEATH OCCURRED IN ONE PATIENT AT HOME DUE TO UNKNOWN REASONS ON POST-OP DAY 20.Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ASSESSED 394 CASES OF LIVER MICROWAVE ABLATION (MWA) TO TREAT COLORECTAL LIVER METASTASIS BETWEEN 2006 AND 2022. CONCOMITANT PROCEDURES INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, AND COLECTOMY WERE PERFORMED IN 63% OF CASES. MICROWAVE ABLATIONS (MWA) WAS PERFORMED USING THE EVIDENT, EMPRINT, EMPRINT SX, EMPRINT HP, OR COMPETITOR SYSTEMS. DEATH OCCURRED IN ONE PATIENT AT HOME DUE TO UNKNOWN REASONS ON POST-OP DAY 20. ADVERSE EVENT WAS NOT RELATED TO THE DEVICE BUT RELATED TO THE PROCEDURE. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK EMPRINT ANT, UNKNOWN EMPRINT ANTENNA (LOT#UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN); UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR (SN:UNKNOWN) WELLS, ALEXANDRA, ET.AL., 2024, OUTCOMES AFTER SURGICAL MICROWAVE ABLATION FOR THE TREATMENT OF COLORECTAL LIVER METASTASIS. AMERICAN COLLEGE OF SURGEONS. ISSN 1879-1190/22, VOL. 239, NO. 3, SEPTEMBER 2024 HTTPS://DOI.ORG/10.1097/XCS.0000000000001097. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H6 (RFR) ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143218&pc=LWS | 9614453-2025-00184 | 2024-12-12 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-01-13 05:00 | LWS | EVERA MRI XT DR SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis; Septic Shock | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BACTERIAL POCKET INFECTION REQUIRING ANTIBIOTICS. ADDITIONALLY, THE PATIENT UNDERWENT AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) POCKET DEBRIDEMENT SURGERY. THE ICD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED SEPSIS LEADING TO SEPTIC SHOCK. THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 1125 VAD OUTFLOW GRAFT, 1103 VAD PUMP SYSTEM, IMPLANTED ON (B)(6)2018. 5076-45 LEAD, IMPLANTED ON (B)(6)2008. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143226&pc=BZD | 2518422-2025-002267 | 2024-06-17 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY. THE PATIENT WENT INTO THE HOSPITAL WITH KIDNEY FAILURE AND NEVER USED THE DEVICE. THE PATIENT'S SON STATED THAT HE WAS USING THE DEVICE. THE SON HAS BEEN INSTRUCTED TO RETURN THE RECALLED DEVICE WHEN HE RECEIVES THE REPLACEMENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY. THE PATIENT WENT INTO THE HOSPITAL WITH KIDNEY FAILURE AND NEVER USED THE DEVICE. THE PATIENT'S SON STATED THAT HE WAS USING THE DEVICE. THE SON HAS BEEN INSTRUCTED TO RETURN THE RECALLED DEVICE WHEN HE RECEIVES THE REPLACEMENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143251&pc=BZD | 2518422-2025-002227 | 2022-01-18 05:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | PHILIPS CPAP DEVICE | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A PHILIPS CPAP DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143548&pc=DSQ | 2916596-2025-00150 | 2024-12-23 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Device Difficult to Setup or Prepare; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Bacterial Infection; Skin Infection; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS ADMITTED. THE PATIENT HAD AN ERODED MODULAR CABLE AND UNFORTUNATELY DID NOT RETURN TO HAVE IT REPAIRED. THE PATIENT HAD BEEN HAVING INTERMITTENT LOW FLOW ALARMS. THE OPPOSITE PORTION OF THEIR MODULAR CABLE THAT WAS ERODED AND COULD NOT BE REPLACED; THE SIDE THAT WAS CONNECTED TO THEIR DRIVELINE, AND PUMP. THE PATIENT PLANNED TO BE INPATIENT AT LEAST ANOTHER WEEK OR TWO. CARDIOTHORACIC SURGERY AND INFECTIOUS DISEASE TEAMS WERE CONSULTED FOR A PLAN OF CARE FOR A POCKET IN THEIR ABDOMEN NEAR THE DRIVELINE. LOG FILES CAPTURED LOW FLOW EVENTS ON (B)(6) 2024. THERE WERE ALSO SEVERAL LOW FLOW FLAGS. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL LOG FILES CAPTURED LOW FLOW EVENTS BEGINNING AT 7:42 PM ON (B)(6) 2025. THESE CONTINUED UNTIL 8:55 PM WITH FLOWS FLUCTUATING BETWEEN 2.4-2.5 LITER PER MINUTE (LPM). FURTHER LOW FLOW ALARMS OCCURRED STARTING AT 10:51 PM AND CONTINUED UNTIL THE DRIVELINE WAS DISCONNECTED AT 12:20 AM. FLOW VALUES DURING THIS PERIOD PROGRESSIVELY DROPPED FROM A BASELINE OF 3.5-4 LPM DOWN TO 0 LPM.Event Description: IT WAS FURTHER REPORTED THAT PATIENT ORIGINALLY PRESENTED A COMMUNICATION FAULT BACK IN (B)(6) OF 2024, WAS SCHEDULED FOR A DRIVELINE CLEANING, BUT LEFT AGAINST MEDICAL ADVICE BEFORE THE CLEANING COULD BE PERFORMED. DRIVELINE CLEANING WAS PERFORMED ON (B)(6) 2025 WITH A MODULAR CABLE EXCHANGE. THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL INFORMATION TO MANUFACTURER REPORT NUMBER 2916596-2024-05243. Manufacturer Narrative: B2 - OUTCOMES ATTRIBUTED TO ADVERSE: CORRECTED. B2 - DATE OF DEATH: CORRECTED. THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL INFORMATION, AND THE INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MFR# 2916596-2024-05243. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143700&pc=DSQ | 2916596-2025-00058 | 2024-12-05 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade; Fungal Infection; Epistaxis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS RE-IMPLANTED WITH A HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) PUMP AND A TEMPORARY CENTRIMAG RIGHT VENTRICULAR ASSIST DEVICE (RVAD) ON (B)(6) 2024. THEY UNDERWENT A REOPENING AND HEMOSTASIS ON (B)(6) 2024 AND WAS THEN EXTUBATED ON (B)(6) 2024. THE RVAD WAS SUCCESSFULLY WEANED AND EXPLANTED ON (B)(6) 2024. IT WAS REPORTED THE PATIENT WAS REINTUBATED ON (B)(6) 2024 DUE TO A MASSIVE EPISTAXIS AND ORAL BLEEDING, WHICH WERE CAUSED BY SIGNIFICANT COAGULOPATHY. THE PATIENT THEN WENT INTO TAMPONADE WITH CARDIOVASCULAR COLLAPSE ON THE NIGHT OF (B)(6) 2024, WHEN THEY UNDERWENT THEIR 3RD RE-OPENING. FOLLOWING THE OPERATION, THEY REMAINED HYPOTENSIVE WITH PERSISTENT LOW FLOW, SEVERE ACIDOSIS, AND WERE COAGULOPATHIC DESPITE TREATMENT WITH ELECTROLYTE REPLACEMENT, MAXIMAL INOTROPIC SUPPORT, AND TRANSFUSION OF BLOOD PRODUCTS. ON (B)(6) 2024 THE PATIENT HAD NO SIGN OF NEUROLOGICAL RECOVERY WITH PERSISTENT HYPOTENSION AND ACIDOSIS ON BACKGROUND OF BACTEREMIA AND FUNGEMIA. THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: A2 - PATIENT AGE: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING BLEEDING, PERICARDIAL FLUID COLLECTION, INFECTION, AND NEUROLOGICAL EVENTS (NOT STROKE RELATED). THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POST IMPLANT COMPLICATION. ALSO, (UNDER "ANTICOAGULATION") OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN (INCLUDING INR RANGE) FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK ALSO PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS SYSTEM ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION H6: DUE TO SYSTEM LIMITATIONS THE FOLLOWING COULD NOT BE INCLUDED: E2321: LOW BLOOD PRESSURE/HYPOTENSION, E0302: COAGULATION DISORDER, E1202: ELECTROLYTE IMBALANCE, NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143789&pc=MGB | 2024168-2025-00518 | 2018-03-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-01-13 05:00 | MGB | PERCLOSE PROGLIDE¿ | Off-Label Use; Insufficient Information; Patient Device Interaction Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THIS IS A RANDOMIZED CONTROLLED TRIAL OF PATIENTS UNDERGOING TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURES COMPARING VESSEL CLOSURE DEVICE (VCD) METHODS OF DUAL PRE-PLACED PERCLOSE PROGLIDE (PP) TECHNIQUE VERSUS ONE PERCLOSE PROGLIDE SUTURE AND AN ANGIO-SEAL (AS+PP) COMBINATION IN COMMON FEMORAL ARTERIES. STUDY RESULTS REVEALED VDC USE INCLUDED VASCULAR COMPLICATION, BLEEDING, DISSECTION, PERFORATION, ISCHEMIA, STENOSIS, NO HEMOSTASIS, AND DEATH. INTERVENTION WAS NOTED TO BE MANUAL COMPRESSION, USE OF ANOTHER DEVICE, BALLOONING, SELF-EXPANDING STENTS, COVERED STENTS, SURGERY AND UNSPECIFIED INTERVENTION. PATIENTS WITH THE AS+PP METHOD OF CLOSURE EXPERIENCED LOWER CHANCE OF SUTURE TEAR, POOR HEMOSTASIS, INCOMPLETE VASCULAR WALL APPOSITION, CALCIFICATION INTERFERENCE, UNPLANNED VASCULAR INTERVENTION, USE OF ADDITIONAL VCD USE, BLEEDING COMPLICATIONS, AND VASCULAR COMPLICATIONS. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE PATIENT EFFECT OF DEATH IS LISTED IN THE ELECTRONIC PERCLOSE PROGLIDE INSTRUCTIONS FOR USE (EIFU), AS A POTENTIAL ADVERSE EVENT OF USE OF THE DEVICE. BASED ON THE ARTICLE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND/OR SURGICAL INTERVENTION WAS LIKELY RELATED TO CASE CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. B2: DATE OF DEATH ESTIMATED B3: DATE OF EVENT ESTIMATED D4: PRIMARY UDI NUMBER - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. H6: MEDICAL DEVICE PROBLEM CODE 1494 - INDICATION FOR USE. LITERATURE ATTACHMENT: ARTICLE TITLED; "DUAL PROGLIDE VS PROGLIDE AND ANGIO-SEAL FOR FEMORAL ACCESS HEMOSTASIS AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT: A RANDOMISED COMPARATIVE TRIAL."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143802&pc=DSQ | 2916596-2025-00225 | 2020-12-02 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE EVENT WAS NOT DEVICE RELATED Manufacturer Narrative: D4: DEVICE LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: D4, EXPIRATION DATE AND PRIMARY UDI NUMBER, UPDATED H4, DEVICE MANUFACTURE DATE, UPDATED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. HEARTMATE II LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) LISTS ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143809&pc=BTI | 2518422-2025-002294 | 2023-07-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BTI | ESSENCE RENTAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K042655 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH AN ESSENCE RENTAL DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS MISSING AND UNAVAILABLE FOR RETURN TO THE MANUFACTURER FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143810&pc=BZD | 2518422-2025-002298 | 2023-08-15 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A PHILIPS CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS MISSING AND UNAVAILABLE FOR RETURN TO THE MANUFACTURER FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143851&pc=DSQ | 2916596-2025-00214 | 2024-12-30 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO OLD AGE ON (B)(6) 2024. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR PROVIDED. IT WAS UNKNOWN IF THE PUMP WOULD BE EXPLANTED OR RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143921&pc=BZD | 2518422-2025-002268 | 2024-07-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-13 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 01/15/2025 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21143972&pc=LWS | 9614453-2025-00188 | 2024-12-27 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-01-13 05:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Device Sensing Problem | Sepsis; No Clinical Signs, Symptoms or Conditions | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT AN EPISODE WAS DETECTED IN THE SUPRA VENTRICULAR TACHYCARDIA (SVT) ZONE THAT WAS SUSPECTED TO BE A MONITORED VENTRICULAR TACHYCARDIA (VT) EPISODE. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THREE DAYS AFTER THE ARRHYTHMIA EPISODE, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS REPROGRAMMED WITH ALL VENTRICULAR ARRHYTHMIA THERAPIES OFF AND PATIENT ALERTS OFF AS THE PATIENT WAS MOVED TO COMFORT CARE. THE PATIENT PASSED AWAY SHORTLY AFTER THE REPROGRAMMING, AND THE CAUSE OF DEATH WAS PROVIDED AS SEPSIS AND LIVER FAILURE. THE CAUSE OR SOURCE OF THE SEPSIS WAS NOT PROVIDED, HOWEVER ADDITIONAL INFORMATION RECEIVED STATED THE SEPSIS WAS NOT RELATED TO THE ICD SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145328&pc=LPB | 2029046-2025-00162 | 2024-12-18 05:00 | Death | BIOSENSE WEBSTER INC | 2025-01-13 05:00 | LPB | THERMOCOOL® SMART TOUCH® SF BI-DIRECTIONAL NAVIGATION CATHETER | Adverse Event Without Identified Device or Use Problem | Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION WITH A THERMOCOOL® SMART TOUCH® SF BI-DIRECTIONAL NAVIGATION CATHETER AND THE PATIENT EXPERIENCED AN ESOPHAGEAL FISTULA. THE PATIENT DIED. DURING THE CASE, AN ESOPHAGEAL FISTULA WAS NOTICED. IT WAS UNKNOWN WHAT CAUSED THE INJURY BUT IT IS BELIEVED THAT IT MAY HAVE BEEN DUE TO OTHER MEDICAL CONDITIONS. THE PATIENT IS DECEASED. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THE DEVICE HAS BEEN REPORTED AS DISCARDED, THEREFORE NO PRODUCT INVESTIGATION CAN BE PERFORMED, AND THE CUSTOMER COMPLAINT CANNOT BE CONFIRMED. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145503&pc=DSQ | 2916596-2025-00108 | 2010-12-28 05:00 | Death | THORATEC CORPORATION | 2025-01-13 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT, US, EP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE DEATH WAS NOT DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS FURTHER COMMUNICATED THAT THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. IT WAS COMMUNICATED THAT NO FURTHER INFORMATION WAS AVAILABLE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A4: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145530&pc=MGB | 2024168-2025-00536 | 2015-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-01-13 05:00 | MGB | PERCLOSE PROGLIDE¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THE AIM OF THIS META-ANALYSIS STUDY WAS TO EVALUATE THE IMPACT OF DIFFERENT LARGE-BORE ARTERIAL ACCESS CLOSURE DEVICES ON CLINICAL OUTCOMES AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT. MULTIPLE VASCULAR CLOSURE DEVICES WERE DISCUSSED, INCLUDING PROSTAR XL, PERCLOSE PROGLIDE, AND MANTA. ALTHOUGH THERE WERE SOME COMPLICATIONS (IN-HOSPITAL DEATH, MINOR AND MAJOR VASCULAR COMPLICATIONS, AND MAJOR OR LIFE-THREATENING BLEEDINGS), MANTA OUTPERFORMED PROGLIDE AND PROSTAR XL AS DESCRIBED IN THIS ARTICLE. PROSTAR XL DEVICE PERFORMED POORLY IN ALL EXPLORED ENDPOINTS. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE, TITLED "LARGE-BORE ARTERIAL ACCESS CLOSURE AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT: A SYSTEMATIC REVIEW AND NETWORK META-ANALYSIS" Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF THE PRODUCTION RECORD AND COMPLAINT HISTORY OF THE REPORTED LOTS COULD NOT BE CONDUCTED BECAUSE THE PART AND LOT NUMBERS WERE NOT PROVIDED. BASED ON THE CASE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE PATIENT EFFECTS OF DEATH IS LISTED IN THE ELECTRONIC PROGLIDE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF THE DEVICE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. ATTACHMENT ARTICLE, TITLED "LARGE-BORE ARTERIAL ACCESS CLOSURE AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT: A SYSTEMATIC REVIEW AND NETWORK META-ANALYSIS". B2- DEATH DATE WILL BE ESTIMATED AS (B)(6) 2015, THE EARLIEST ESTIMATED DATE OF EVENT. B3- THE DATE OF EVENT WAS ESTIMATED AS 1/1/2015 AS THE ARTICLE STATED THAT ONLY COHORTS ENROLLED AFTER 2015 WERE INCLUDED. D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND/OR MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145538&pc=MGB | 2024168-2025-00538 | 2015-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-01-13 05:00 | MGB | PROSTAR | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P960043 |  | 1 | Event Description: THE AIM OF THIS META-ANALYSIS STUDY WAS TO EVALUATE THE IMPACT OF DIFFERENT LARGE-BORE ARTERIAL ACCESS CLOSURE DEVICES ON CLINICAL OUTCOMES AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT. MULTIPLE VASCULAR CLOSURE DEVICES WERE DISCUSSED, INCLUDING PROSTAR XL, PERCLOSE PROGLIDE, AND MANTA. ALTHOUGH THERE WERE SOME COMPLICATIONS (IN-HOSPITAL DEATH, MINOR AND MAJOR VASCULAR COMPLICATIONS, AND MAJOR OR LIFE-THREATENING BLEEDINGS), MANTA OUTPERFORMED PROGLIDE AND PROSTAR XL AS DESCRIBED IN THIS ARTICLE. PROSTAR XL DEVICE PERFORMED POORLY IN ALL EXPLORED ENDPOINTS. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE, TITLED "LARGE-BORE ARTERIAL ACCESS CLOSURE AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT: A SYSTEMATIC REVIEW AND NETWORK META-ANALYSIS" Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF THE PRODUCTION RECORD AND COMPLAINT HISTORY OF THE REPORTED LOT COULD NOT BE CONDUCTED BECAUSE THE PART AND LOT NUMBERS WERE NOT PROVIDED. BASED ON THE CASE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. ATTACHMENT ARTICLE, TITLED "LARGE-BORE ARTERIAL ACCESS CLOSURE AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT: A SYSTEMATIC REVIEW AND NETWORK META-ANALYSIS". B2- DEATH DATE WILL BE ESTIMATED AS (B)(6) 2015, THE EARLIEST ESTIMATED DATE OF EVENT. B3- THE DATE OF EVENT WAS ESTIMATED AS (B)(6) 2015 AS THE ARTICLE STATED THAT ONLY COHORTS ENROLLED AFTER (B)(6) 2015 WERE INCLUDED. D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND/OR MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21148527&pc=LPB | MW5164608 | 2022-01-31 05:00 | Death | BIOSENSE WEBSTER, INC. | 2025-01-13 05:00 | LPB | THERMOCOOL SMARTTOUCH |  | Insufficient Information |  |  | 1 | Event Description: OMISSION OF THERMOCOOL SMARTTOUCH DEVICE LOT, UDI NUMBERS IN THE MEDICAL RECORDS. UNABLE TO DETERMINE SPECIFIC DEVICE USED ON PT WHO DIED FROM DEFINITIVE COMPLICATIONS FROM A CARDIAC ABLATION PROCEDURE. THE DEVICE COULD HAVE BEEN PART OF SEVERAL FDA RECALLS - Z-0169-2023, Z-0054-2018-20056-2018, Z-0056-2018, Z-0057-2018, Z-0058-2018, Z-0139-2015. VIOLATION OF 21 CFR PART 801 - FDA. FDA - UNIQUE DEVICE IDENTIFICATION (UDI) SYSTEM FINAL RULE. EMR - PAGE 39 - UNDER ID# (B)(6). ON PAGE 11 - UNDER ID# (B)(6). REF REPORT: MW5164609.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21148553&pc=LPB | MW5164609 | 2022-01-31 05:00 | Death | BIOSENSE WEBSTER, INC. | 2025-01-13 05:00 | LPB | THERMOCOOL SMARTTOUCH |  | Insufficient Information |  |  | 1 | Event Description: OMISSION OF THERMOCOOL SMARTTOUCH DEVICE LOT, UDI NUMBERS IN THE MEDICAL RECORDS. UNABLE TO DETERMINE SPECIFIC DEVICE USED ON PT WHO DIED FROM DEFINITIVE COMPLICATIONS FROM A CARDIAC ABLATION PROCEDURE. THE DEVICE COULD HAVE BEEN PART OF SEVERAL FDA RECALLS - Z-0169-2023, Z-0054-2018-20056-2018, Z-0056-2018, Z-0057-2018, Z-0058-2018, Z-0139-2015. VIOLATION OF 21 CFR PART 801 - FDA. FDA - UNIQUE DEVICE IDENTIFICATION (UDI) SYSTEM FINAL RULE. EMR - PAGE 39 - UNDER ID# (B)(6). ON PAGE 11 - UNDER ID# (B)(6). REF REPORT: MW5164608.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133917&pc=BZD | 2518422-2025-000744 | 2023-11-06 05:00 | Death | RESPIRONICS, INC. | 2025-01-12 05:00 | BZD | REMSTAR AUTO A-FLEX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR AUTO A-FLEX. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133922&pc=BZD | 2518422-2025-000745 | 2023-08-03 04:00 | Death | RESPIRONICS, INC. | 2025-01-12 05:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A PHILIPS CPAP DEVICE. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING PATIENT WAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H WAS UPDATED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133934&pc=BZD | 2518422-2025-001748 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-12 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21134052&pc=DSQ | 2916596-2025-00155 | 2024-12-22 05:00 | Death | THORATEC CORPORATION | 2025-01-12 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE MORNING OF (B)(6) 2024, PATIENT'S FLOW SUDDENLY INCREASED, BUT THE POWER, WHICH SUGGESTED ARRANGING FURTHER EXAMINATION TO RULE OUT ANY POSSIBILITY OF PUMP THROMBUS OR AORTIC INSUFFICIENCY (AI) SITUATION. THE FLOW INCREASED HIGHER ON (B)(6) 2024, EVEN OVER 10 LITERS PER MINUTE (LPM). THE POSSIBILITY OF AL WAS RULED OUT THROUGH TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) AND CHEST ECHOCARDIOGRAPHY. SYSTEM CONTROLLER WAS SWITCHED TO THE PATIENT'S BACKUP SYSTEM CONTROLLER, BUT THE DATA DID NOT DIFFER AS A RESULT. THE PATIENT AND THE DEVICE WERE CONTINUALLY MONITORED IN THE HOSPITAL. THE SPEED WAS TRIED TO DECREASE FROM 4800 TO 4400 REVOLUTIONS PER MINUTE (RPM) UNDER ECHOCARDIOGRAPHY AND FOUND OUT THAT THE SIZE OF THE LEFT VENTRICLE BECAME LARGER. THE LOG FILES DISPLAYED LOW SPEED ADVISOR(IES) ON (B)(6) 2024 AT 12:04 PM WHERE THE PUMP SPEED 4200 RPM WAS MOMENTARILY SET LOWER THAN THE LOW SET LIMIT 4600 RPM. THE PUMP SPEED WAS CHANGED TO 4800 RPM ON (B)(6) 2024 AT 12:05 PM. THE EVENT LOG DISPLAYED INCREASES IN FLOW RANGING FROM 3.5 LPM TO 11.8 LPM. PATIENT'S FAMILY DECIDED TO REMOVE ALL THE LIFE SUPPORT SYSTEMS, WHICH WAS DUE TO UNDERLYING CONDITION. DUE TO THE POOR CONDITION, THE CHEST WAS NOT OPENED, SO THROMBUS WAS UNABLE TO BE CONFIRMED. THERE WAS NO CHANGE OF ANTICOAGULATION, POST OPERATION HEPARIN USE. THE PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE ON (B)(6) 2024. THE EVENT WAS NOT CONSIDERED TO BE DEVICE RELATED SINCE THE OUTCOME WAS FROM THE PROGRESSION OF THE PATIENT'S OWN DISEASE. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. SECTION B7: OTHER RELEVANT HISTORY CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE, AND MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: THE SUSPECTED PUMP THROMBUS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION AS NO IMAGES WERE SUBMITTED FOR ANALYSIS AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION; HOWEVER, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED PUMP POWER AND FLOW ELEVATIONS THAT APPEARED CONSISTENT WITH MATERIAL SUCH AS A THROMBUS BEING INGESTED INTO THE PUMP AND CONTACTING THE ROTOR. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILES CAPTURED ELEVATIONS IN MOTOR POWER AND ESTIMATED FLOW FROM (B)(6) 2024, WITH PUMP PARAMETERS REACHING UP TO VALUES OF 5.3 WATTS AND 12.3 LITERS PER MINUTE. BASED ON COMPLAINT HISTORY AND SIMILARLY REPORTED EVENTS, THE INFORMATION CAPTURED WITHIN THE SUBMITTED LOG FILES IS CONSISTENT WITH MATERIAL, SUCH AS A THROMBUS, BEING INGESTED INTO THE PUMP AND CONTACTING THE ROTOR. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. SECTION E1: REPORTER ADDRESS LINE 1: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21134469&pc=OZP | 2032227-2025-107941 | 2024-11-19 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-12 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE CAUSE OF THE DEATH WAS DUE TO SMALL TUMOR IN LIVER. THE CUSTOMER PASSED AWAY IN HOSPITAL. THE LAST KNOWN PRODUCT (S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-381A, MMT-7040B, MMT-332A. THE PUMP WAS NOT WORN AT THE TIME OF PASSING. THE CUSTOMER WAS NOT USING A SENSOR AT THE TIME OF EVENT. THE CUSTOMER WAS USING THE INSULIN PUMP WITHIN 48 HOURS AND IT WAS UNKNOWN IF CUSTOMER USING THE AUTO-MODE/SMARTGUARD FEATURE AT THE TIME OF EVENT. MMT-1884 WAS REQUESTED, AND THE DEVICE WAS RETURNED FOR ANALYSIS. MMT-7040B WAS REQUESTED AND THE DEVICE WAS RETURNED FOR ANALYSIS. MMT-332A WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-381A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. THE PUMP PASSED THE FUNCTIONAL TESTS, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST AND DAT AT 0.08715 INCHES. SUCCESSFULLY DOWNLOADED HISTORY FILES AND TRACES USING THUMP. SUCCESSFULLY UPLOADED PUMP TO CARELINK. PLEASE SEE BELOW FOR THE DATE RANGE LISTED IN THE FORMATTED HISTORY FILE. THE FORMATTED HISTORY FILE LISTS DATA FROM (B)(6) 2024 TO (B)(6) 2024. THERE WAS NO DATA AVAILABLE TO VERIFY UNEXPECTED ALARM(S)/SUSPENDS AND BOLUS/BASAL DELIVERY FOR THE SS EVENT DATE OF 26-NOV-2024. IN FURTHER FULL REVIEW OF THE PUMP HISTORY/TRACES ON THE CUSTOMER'S EVENT DATE OF 19-NOV-2024, THERE IS NO UNEXPECTED ALARMS/SUSPENDS. UNABLE TO VERIFY DAILY TOTAL OF BASAL/BOLUS AND ALL INSULIN DELIVERED ON THE CUSTOMER'S EVENT DATE OF 19-NOV-2024 LISTED ON SMARTSOLVE DUE TO INSUFFICIENT DATA IN THE TRACE/HISTORY FILES. THE PUMP WAS PROGRAMMED WITH MULTIPLE BOLUS DELIVERIES AND ALL BOLUS DELIVERED PROPERLY THEIR INDICATED AMOUNTS (AT QUICK BOLUS SPEED) AND WERE PROPERLY RECORDED IN THE DAILY HISTORY. NO BOLUS DELIVERY ANOMALY OR HISTORY ANOMALY NOTED. PLEASE SEE BELOW FOR PUMP ERROR(S)/ALARM(S) NOTED 1 WEEK PRIOR TO THE SS EVENT DATE OF 26-NOV-2024 AND CUSTOMER'S EVENT DATE OF 19-NOV-2024 IN THE FORMATTED HISTORY FILE. LOW BATTERY ALERT WAS FOUND ON: (B)(6) 2024 10:12:00.000 AND (B)(6) 2024 21:45:00.000. INSERT BATTERY ALARM WAS FOUND ON: (B)(6) 2024 15:45:48.000, (B)(6) 2024 13:33:48.000, (B)(6) 2024 07:29:09.000 AND (B)(6) 2024 07:39:00.000. REPLACE BATTERY ALERT WAS FOUND ON: (B)(6) 2024 07:16:00.000 AND (B)(6) 2024 07:26:00.000. PUMP ERROR 23 ALARM WAS FOUND ON: (B)(6) 2024 07:40:04.000. POWER LOSS ALARM WAS FOUND ON: (B)(6) 2024 07:43:02.000 AND (B)(6) 2024 07:43:09.000. POWER MANAGEMENT GRAPH WAS SUCCESSFULLY GENERATED. THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE (UL VLITH) AND LOADED VOLTAGE (LOADED VLITH) WERE WITHIN SPEC RANGE. INSERT BATTERY ALARM WAS EXPECTED SINCE THE BATTERY WAS REMOVED FROM THE PUMP. EARLIEST POWER DATA AVAILABLE PER THE POWER MANAGEMENT TOOL/DETAIL TRACE FILE IS ON (B)(6) 2024 AT 2:09:56 AM. THERE WAS NO POWER DATA AVAILABLE FOR THE DATE OF (B)(6) 2024. UNABLE TO CHECK POWER DATA FOR LOW BATTERY ALERT AT (B)(6) 2024 10:12:00.000. UPON CHECKING ON THE POWER DATA/DETAIL TRACE FILE, LOW BATTERY ALERT AT (B)(6) 2024 21:45:00.000 AND REPLACE BATTERY ALERT WERE EXPECTED SINCE THE BATTERY IN THE PUMP IS LOW ON POWER OR DOES NOT HAVE ENOUGH POWER. THE CUSTOMER HAD USED A LOW/NO POWER/DEPLETED BATTERY. PUMP ERROR 23 ALARM AND POWER LOSS ALARM WERE EXPECTED SINCE THE PUMP RESETS AFTER THE BATTERY WAS REMOVED FOR MORE THAN 10 MINUTES. PUMP WAS MONITORED WITHOUT THE TEST ENERGIZER BATTERY INSIDE THE BATTERY COMPARTMENT AND REINSERTS IT BACK INTO THE BATTERY COMPARTMENT FOR LESS THAN 10 MINUTES AND NO UNEXPECTED PUMP ERROR 23 ALARM NOTED. NO UNEXPECTED LOW BATTERY ALERT, REPLACE BATTERY ALERT, PUMP ERROR 23 ALARM AND POWER LOSS ALARM NOTED DURING TESTING. SGCALIBRATIONERROR (776) WAS FOUND ON: (B)(6) 2024 21:47:20.000 LOSTSENSOR1ALERT (780) WAS FOUND ON: (B)(6) /2024 11:03:00.000 AND (B)(6) 2024 11:13:00.000. LOSTSENSOR2ALERT (781) WAS FOUND ON: (B)(6) 2024 11:30:00.000 AND (B)(6) 2024 11:40:00.000. SENSOREXPIREDALERT (794) WAS FOUND ON: (B)(6) 2024 20:17:32.000. THE PUMP WAS PROGRAMMED WITH A TEST GUARDIAN LINK (3) TRANSMITTER AND A GLUCOSE SENSOR SIMULATOR. THE PUMP CONNECTED SUCCESSFULLY TO THE TRANSMITTER AND DISPLAYED ¿TRANSMITTER CONNECTION SUCCESSFUL¿. THE PUMP COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR AND DISPLAYED THE CALIBRATE YOUR SENSOR ALARM PROPERLY AFTER COMPLETION OF THE WARM UP. THE PUMP CALIBRATED AND DISPLAYED THE PROGRAMMED VALUE OF 240 MG/DL PROPERLY ON THE DISPLAY GRAPH. NO SG CALIBRATION ERROR, LOST SENSOR ALERT, SENSOR EXPIRED ALERT OR UNEXPECTED SENSOR ERRORS OR ANOMALIES WERE NOTED DURING TESTING. TEST P-CAP AND RESERVOIR LOCKED PROPERLY INTO RESERVOIR COMPARTMENT DURING TESTING. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: A SCRATCHED CASE. THE PUMP WAS RECEIVED WITH THE ORIGINAL BATTERY CAP. NO CRACKED/DAMAGED BATTERY CAP NOTED DURING VISUAL INSPECTION. THE PUMP WAS RECEIVED WITHOUT A BATTERY INSTALLED. UNABLE TO VERIFY CUSTOMER'S DAILY TOTAL OF ALL INSULIN DELIVERED SURROUNDING THE DATE OF (B)(6) 2024 LISTED ON PUMP HISTORY AND THE DAYS PRIOR TO THAT DATE DUE TO INSUFFICIENT DATA IN THE TRACE/HISTORY FILES. UNABLE TO VERIFY CUSTOMER'S DAILY TOTAL OF ALL INSULIN DELIVERED SURROUNDING THE CUSTOMER'S EVENT DATE OF 19-NOV-2024 LISTED ON SMARTSOLVE AND THE DAYS PRIOR TO THE CUSTOMER'S EVENT DATE DUE TO INSUFFICIENT DATA IN THE TRACE/HISTORY FILES. THE PUMP PASSED ALL THE REQUIRED TESTING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21134535&pc=BZD | 2518422-2025-001914 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-12 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21132871&pc=DSQ | 2916596-2024-52977 | 2023-11-15 05:00 | Death | THORATEC CORPORATION | 2025-01-11 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. Manufacturer Narrative: SECTION A - PATIENT WEIGHT WAS UNKNOWN. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE DEVICE OPERATED AS EXPECTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A - PATIENT WEIGHT WAS UNKNOWN. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133048&pc=BZD | 2518422-2025-001865 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-11 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; THROAT/ESOPHAGEAL AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133104&pc=BZD | 2518422-2025-001871 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-11 05:00 | BZD | PHILIP CPAP DEVICE | Degraded | Cancer |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133271&pc=MIH | 2017233-2025-05713 | 2024-12-20 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-11 05:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Insufficient Information | Cardiac Arrest | P210032 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2024, A REINTERVENTION TOOK PLACE TO FIX A SUSPECTED ENDOLEAK ON AN EXISTING COOK GRAFT. A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS (TBE) BRANCH COMPONENT WAS ADVANCED TO THE LEVEL OF THE DESCENDING THORACIC AORTA, JUST BELOW THE AORTIC ARCH. THE FSA NOTED THE PHYSICIAN HAVING DIFFICULTY ADVANCING THE DEVICE THROUGH THE COOK GRAFT AND SUGGESTED PULLING THE DEVICE DOWN A BIT TO CHECK FOR WIRE WRAP. THE PHYSICIAN PULLED THE DEVICE DOWN, AND INSPECTED FOR WIRE WRAP, BUT FOUND NONE; BUT BEFORE READVANCING, THE PATIENT'S BLOOD PRESSURE DROPPED. THE PATIENT CODED. CPR WAS ADMINISTERED, BUT THE PATIENT DIED. A TRANSESOPHAGEAL ECHO (TEE) PROBE WAS INSERTED AND SHOWED A LOT OF BLOOD IN THE CHEST CAVITY. Manufacturer Narrative: D4 BASIC UDI UPDATED TO FULL NUMBER G3 DATE MANUFACTURER RECEIVED IS JANUARY 22, 2025, THE DATE OF THE IMAGING EVALUATION. H.6. INVESTIGATION FINDINGS CODE C20: THE DEVICE REMAINS IMPLANTED AND IS NOT AVAILABLE FOR ANALYSIS. H.6. INVESTIGATION FINDINGS CODE C19: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H.6. TYPE OF INVESTIGATION CODE B15: EVALUATION OF CT IMAGES FROM (B)(6) 2024 - 3D AND RECONSTRUCTION IMAGES SHOW LACK OF PROXIMAL NON-GORE DEVICE APPOSITION IN THE DISTAL THORACIC AORTIC ARCH. (BIRD-BEAKING OF THE PROXIMAL END OF THE NON-GORE DEVICE.) H.6. INVESTIGATION CONCLUSIONS CODE D15: THERE IS INSUFFICIENT INFORMATION AVAILABLE FOR GORE TO REASONABLY DRAW CONCLUSIONS RELATED TO ASPECTS OF THE EVENT, THEREFORE CONCLUSION CODE ¿D15: CAUSE NOT ESTABLISHED¿ IS BEING USED. INSUFFICIENT INFORMATION MAY INCLUDE LIMITED OR MISSING RELEVANT MEDICAL RECORDS, INVOLVEMENT OF MULTIPLE IMPLANTED DEVICES (INCLUDING NON-GORE DEVICES) IN THE FIELD OF TREATMENT, PATIENT NON-COMPLIANCE, AND/OR A GENERAL LACK OF AVAILABLE DETAIL OR SPECIFICITY RELATED TO AN ADVERSE EVENT AND/OR DEVICE. Manufacturer Narrative: H.6. INVESTIGATION FINDINGS CODE C21: INVESTIGATION PENDING PRODUCT HISTORY REVIEW AND IMAGING EVALUATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133275&pc=DSQ | 2916596-2025-00057 | 2024-12-21 05:00 | Death | THORATEC CORPORATION | 2025-01-11 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICES WERE WORKING AS INTENDED DURING THEIR TIME OF SUPPORT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS IN POOR HEALTH CONDITION PRIOR TO DEATH. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21133285&pc=DSQ | 2916596-2024-53028 | 2024-09-23 04:00 | Death | THORATEC CORPORATION | 2025-01-11 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE REPORTEDLY OPERATED AS EXPECTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: D4: DEVICE LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21262394&pc=NVZ | 3004209178-2025-01682 | 2025-01-09 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-29 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENTS REPRESENTATIVE THAT THE PATIENT HAD PASSED AWAY AND WANTED TO KNOW WHY THE IMPLANTABLE PULSE GENERATOR (IPG) FAILED. ADDITIONAL INFORMATION REGARDING THE CAUSE OF DEATH WAS REQUESTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263049&pc=MIH | 9612164-2025-00471 | 2019-02-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-29 05:00 | MIH | VALIANT NAVION | Leak/Splash | Insufficient Information | P100040 |  | 1 | Event Description: A VALIANT NAVION STENT GRAFT WAS IMPLANTED DURING AN UNKNOWN ENDOVASCULAR PROCEDURE. THE PATIENT HAD AN EXISTING VALIANT CAPTIVA AND ENDURANT IIS STENT GRAFT SYSTEM IMPLANTED SIX MONTHS EARLIER. IT WAS REPORTED THAT CORELAB REVIEW OF A CT TAKEN ON THE SAME DATE SHOWED AN IA AND IIIA ENDOLEAKS INVOLVING THE VALIANT NAVION. THERE WAS NO FRACTURE, STENT RING ENLARGEMENT OR STENT RING MIGRATION OBSERVED. THE MAXIMUM AORTIC DIAMETER WAS 96.8MM. IT WAS NOTED THAT THE VALIANT NAVION WAS FAILING TO SEAL WITH THE OTHER DEVICES. THE SITE COULD NOT CONFIRM WHICH DEVICE THE NAVION HAS LOST SEAL AGAINST. IT IS UNKNOWN IF THERE WERE INTERVENTIONS PERFORMED TO TREAT THE ENDOLEAKS. NO CAUSE WAS PROVIDED. IT WAS REPORTED THE PATIENT EXPIRED THE SAME YEAR. NO CAUSE OF DEATH WAS PROVIDED. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: PER CORELAB, IT WAS SAID FOR THE TYPE IIIA ENDOLEAK, IT'S LOSS OF SEAL BETWEEN THE NAVION AND ALL 3 OF THE DEVICES IN THE DISTAL END OF THE NAVION. IT WAS CONFIRMED THE IIIA ENDOLEAK IS BETWEEN THE VALIANT NAVION STENT GRAFT AND WHATEVER STENT GRAFT IS IMMEDIATELY DISTAL TO THE NAVION, WHICH IS UNKNOWN. THE SITE DOES NOT HAVE ANY INFORMATION ON THE IMPLANT ORDER OF THE DEVICES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263076&pc=MIH | 9612164-2025-00472 | 2019-02-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-29 05:00 | MIH | VALIANT CAPTIVIA - FF | Leak/Splash | Insufficient Information | P100040 |  | 1 | Event Description: A VALIANT NAVION STENT GRAFT WAS IMPLANTED DURING AN UNKNOWN ENDOVASCULAR PROCEDURE . THE PATIENT HAD AN EXISTING VALIANT CAPTIVA AND ENDURANT IIS STENT GRAFT SYSTEM IMPLANTED SIX MONTHS EARLIER. IT WAS REPORTED THAT CORELAB REVIEW OF A CT TAKEN ON THE SAME DATE SHOWED A IA AND IIIA ENDOLEAKS INVOLVING THE VALIANT NAVION. THERE WAS NO FRACTURE, STENT RING ENLARGEMENT OR STENT RING MIGRATION OBSERVED. THE MAXIMUM AORTIC DIAMETER WAS 96.8MM. IT WAS NOTED THAT THE VALIANT NAVION WAS FAILING TO SEAL WITH THE OTHER DEVICES. THE SITE COULD NOT CONFIRM WHICH DEVICE THE NAVION HAS LOST SEAL AGAINST. IT IS UNKNOWN IF THERE WERE INTERVENTIONS PERFORMED TO TREAT THE ENDOLEAKS. NO CAUSE WAS PROVIDED. IT WAS REPORTED THE PATIENT EXPIRED THE SAME YEAR. NO CAUSE OF DEATH WAS PROVIDED. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: PER CORELAB, IT WAS SAID FOR THE TYPE IIIA ENDOLEAK, IT'S LOSS OF SEAL BETWEEN THE NAVION AND ALL 3 OF THE DEVICES IN THE DISTAL END OF THE NAVION. IT WAS CONFIRMED THE IIIA ENDOLEAK IS BETWEEN THE VALIANT NAVION STENT GRAFT AND WHATEVER STENT GRAFT IS IMMEDIATELY DISTAL TO THE NAVION, WHICH IS UNKNOWN. THE SITE DOES NOT HAVE ANY INFORMATION ON THE IMPLANT ORDER OF THE DEVICES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263085&pc=MIH | 9612164-2025-00473 | 2019-02-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-29 05:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Leak/Splash | Insufficient Information | P100021 |  | 1 | Event Description: A VALIANT NAVION STENT GRAFT WAS IMPLANTED DURING AN UNKNOWN ENDOVASCULAR PROCEDURE. THE PATIENT HAD AN EXISTING VALIANT CAPTIVA AND ENDURANT IIS STENT GRAFT SYSTEM IMPLANTED SIX MONTHS EARLIER. IT WAS REPORTED THAT CORELAB REVIEW OF A CT TAKEN ON THE SAME DATE SHOWED A IA AND IIIA ENDOLEAKS INVOLVING THE VALIANT NAVION. THERE WAS NO FRACTURE, STENT RING ENLARGEMENT OR STENT RING MIGRATION OBSERVED. THE MAXIMUM AORTIC DIAMETER WAS 96.8MM. IT WAS NOTED THAT THE VALIANT NAVION WAS FAILING TO SEAL WITH THE OTHER DEVICES. THE SITE COULD NOT CONFIRM WHICH DEVICE THE NAVION HAS LOST SEAL AGAINST. IT IS UNKNOWN IF THERE WERE INTERVENTIONS PERFORMED TO TREAT THE ENDOLEAKS. NO CAUSE WAS PROVIDED. IT WAS REPORTED THE PATIENT EXPIRED THE SAME YEAR. NO CAUSE OF DEATH WAS PROVIDED. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED : PER CORELAB, IT WAS SAID FOR THE TYPE IIIA ENDOLEAK, IT'S LOSS OF SEAL BETWEEN THE NAVION AND ALL 3 OF THE DEVICES IN THE DISTAL END OF THE NAVION. IT WAS CONFIRMED THE IIIA ENDOLEAK IS BETWEEN THE VALIANT NAVION STENT GRAFT AND WHATEVER STENT GRAFT IS IMMEDIATELY DISTAL TO THE NAVION, WHICH IS UNKNOWN. THE SITE DOES NOT HAVE ANY INFORMATION ON THE IMPLANT ORDER OF THE DEVICES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263113&pc=MIH | 9612164-2025-00475 | 2019-02-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-29 05:00 | MIH | ENDURANT II ILIAC STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100021 |  | 1 | Event Description: A VALIANT NAVION STENT GRAFT WAS IMPLANTED DURING AN UNKNOWN ENDOVASCULAR PROCEDURE . THE PATIENT HAD AN EXISTING VALIANT CAPTIVA AND ENDURANT IIS STENT GRAFT SYSTEM IMPLANTED SIX MONTHS EARLIER. IT WAS REPORTED THAT CORELAB REVIEW OF A CT TAKEN ON THE SAME DATE SHOWED A IA AND IIIA ENDOLEAKS INVOLVING THE VALIANT NAVION. THERE WAS NO FRACTURE, STENT RING ENLARGEMENT OR STENT RING MIGRATION OBSERVED. THE MAXIMUM AORTIC DIAMETER WAS 96.8MM. IT WAS NOTED THAT THE VALIANT NAVION WAS FAILING TO SEAL WITH THE OTHER DEVICES. THE SITE COULD NOT CONFIRM WHICH DEVICE THE NAVION HAS LOST SEAL AGAINST. IT IS UNKNOWN IF THERE WERE INTERVENTIONS PERFORMED TO TREAT THE ENDOLEAKS. NO CAUSE WAS PROVIDED. IT WAS REPORTED THE PATIENT EXPIRED THE SAME YEAR. NO CAUSE OF DEATH WAS PROVIDED. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: PER CORELAB, IT WAS SAID FOR THE TYPE IIIA ENDOLEAK, IT'S LOSS OF SEAL BETWEEN THE NAVION AND ALL 3 OF THE DEVICES IN THE DISTAL END OF THE NAVION. IT WAS CONFIRMED THE IIIA ENDOLEAK IS BETWEEN THE VALIANT NAVION STENT GRAFT AND WHATEVER STENT GRAFT IS IMMEDIATELY DISTAL TO THE NAVION, WHICH IS UNKNOWN. THE SITE DOES NOT HAVE ANY INFORMATION ON THE IMPLANT ORDER OF THE DEVICES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263136&pc=BZD | 2518422-2025-101415 | 2024-10-21 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE PREVIOUSLY CALLED TO ASK FOR ALL COMMUNICATION TO STOP, BUT SHE IS STILL RECEIVING EMAILS AND LETTERS. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263334&pc=BZD | 2518422-2025-101419 | 2024-10-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR THE PATIENT TO BE REMOVED FROM THE MAILING LIST. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263444&pc=OZD | 1220648-2025-25868 | 2025-01-13 05:00 | Death | ABIOMED, INC. | 2025-01-29 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. HOWEVER, ON THE DAY OF IMPLANT, ECHOCARDIOGRAM SHOWED LEFT VENTRICLE THROMBOSIS AND PLAN WAS TO MONITOR AND MEDICALLY TREAT. IN THE FOLLOWING TWO DAYS THE PATIENT'S CONDITION DETERIORATED. CONTINUOUS RENAL REPLACEMENT THERAPY WAS STARTED, AND THE PATIENT WAS SHOCKED DUE TO SUPRAVENTRICULAR TACHYCARDIA. AND ATRIAL FIBRILLATION/RAPID VENTRICULAR RESPONSE. THE PATIENT WOULD EVENTUALLY EXPIRE. THE FAMILY DECIDED NOT TO ESCALATE CARE AND ULTIMATELY WITHDREW CARE. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. NO DEVICE WAS RECEIVED FOR EVALUATION. DEVICE HISTORY RECORD REVIEW CONFIRMED THAT THE PUMP PASSED ALL POST-STERILE INSPECTION CHECKS. DUE TO THE LIMITED CLINICAL INFORMATION AND LACK OF AVAILABLE DATA/PRODUCT THE ROOT CAUSE OF THIS INVESTIGATION IS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263572&pc=OZD | 1220648-2025-25869 | 2025-01-13 05:00 | Death | ABIOMED, INC. | 2025-01-29 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED ECTOPY AND LIMB ISCHEMIA WITH A DEMISED OUTCOME. THE PATIENT WAS A RECENT ADMISSION FOR ACUTE SYSTOLIC AND DIASTOLIC HEART FAILURE (BASELINE EJECTION FRACTION OF 25%) AND UNDERWENT 3-VESSEL CORONARY ARTERY BYPASS GRAFT SURGERY IN THE MORNING. THE PATIENT ARRIVED AT THE UNIT ON TWO INOTROPES/VASOPRESSORS AND THROUGHOUT THE DAY HAD INCREASED NEEDS FOR MEDICAL SUPPORT. CARDIOTHORACIC SURGERY CONSULTED WITH CARDIOLOGY FOR AN INTRA-AORTA BALLOON PUMP PLACEMENT, BUT CARDIOLOGY OPTED TO IMPLANT AN IMPELLA CP DEVICE WHICH WAS SUCCESSFULLY COMPLETED. AFTER START OF SUPPORT, SOME ECTOPY WAS EXPERIENCED AND WAS MOSTLY RESOLVED WITH MINOR POSITIONING ADJUSTMENTS. THE PATIENT HAD ANTI ARRHYTHMIC MEDICATIONS PRIOR TO IMPELLA DUE TO PREMATURE VENTRICULAR CONTRACTIONS POST-OPERATION. ADDITIONALLY, THE NEXT DAY, THE PATIENT WAS NOTED TO HAVE A POPLITEAL PULSE. HOWEVER, PULSES WERE ABSENT IN LEFT FOOT. THE PATIENT HAD A HISTORY OF RAYNAUDS AND THE TEAM THEREFORE WAS EVALUATING SUCH AS THE CAUSE BECAUSE THIS PRIMARILY HAS AFFECTED THE PATIENT'S LEFT FOOT IN THE PAST. ALSO, THERE WAS DISCUSSION REGARDING, IF NEEDED, PLACING A NEW IMPELLA CP DEVICE FROM AN AXILLARY APPROACH. HOWEVER, THE PATIENT WOULD EXPIRE THE NEXT DAY. THE FAMILY DECIDED TO WITHDRAW CARE. ALTHOUGH AN ALTERNATIVE CAUSE FOR THE LIMB ISCHEMIA WAS BEING CONSIDERED, THE CAUSE OF DEATH IS UNKNOWN, AND THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. THE PRODUCT WAS NOT RETURNED. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE ECTOPY AND LIMB ISCHEMIA WAS NOT DETERMINED. Manufacturer Narrative: G6 TYPE OF REPORT; 30 DAY WAS INADVERTENTLY NOT SELECTED MANUFACTURER DEVICE REPORT 1220648-2025-25869-1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263642&pc=BZD | 2518422-2025-101421 | 2024-10-22 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY THREE YEARS AGO. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS BEEN MISPLACED AND WILL NOT BE RETURNED FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263718&pc=NPT | 2025587-2025-00732 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-29 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Stroke/CVA | P130021 | E2014038 | 18 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE STROKE. THE AVERAGE TIME TO EVENT WAS SEVENTY-TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 62 TO 90 AND OVER. THERE WERE 6 MALES AND 12 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 18 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263937&pc=MHX | 9610816-2025-000073 | 2024-12-28 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-01-29 05:00 | MHX | INTELLIVUE MX450 PATIENT MONITOR | Defective Alarm | Cardiac Arrest | K130849 |  | 1 | Event Description: IT WAS REPORTED THE STAFF REPORTED NOT HEARING ALARMS FOR DESATURATION FROM THE STANDALONE MONITOR, AND THE NICU PATIENT PASSED AWAY. THE CUSTOMER REQUESTED ASSISTANCE WITH THE AUDIT LOGS. THE LOGS WERE INITIALLY REVIEWED BY WITH THE CUSTOMER, REVEALING THE CAUSE OF NO ALARM WAS THE ALARMS HAD BEEN PAUSED BY THE USER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. THIS INFORMATION INDICATED THE PATIENT WAS AN EMERGENCY 32-WEEK GESTATION C-SECTION DELIVERY WITH ALVEOLAR HEMORRHAGING AND CARDIAC ARREST. THE PATIENT ALMOST PASSED IN THE OR BUT STAFF WAS ABLE TO GET THE PATIENT TO THE NICU. THE PATIENT WAS MONITORED FOR ECG AND SPO2 AND STAFF INDICATED THEY DID NOT HEAR ALARMS FROM THIS STANDALONE MONITOR AND THE PATIENT SUBSEQUENTLY PASSING AWAY. AN INTERNAL INVESTIGATION WAS PERFORMED WITH THE MONITOR WITH ITS EXISTING LEADS AND CABLES USED DURING THE TIME OF THE EVENT. ALARMS WERE GENERATED NORMALLY WITH APPROPRIATE VOLUME SETTINGS; HOWEVER, IT WAS NOTED DURING LOG REVIEW THAT THE ALARMS HAD BEEN PAUSED BY USER, WITH THE PAUSE BEING THE DEFAULT OF 2 MINUTES. THE PAUSE TIME COULD NOT BE ADJUSTED AS THE CONFIGURED DEFAULT WAS 2 MINUTES. BASED ON THIS INFORMATION, THE DEVICE FUNCTIONED AS CONFIGURED; HOWEVER, USE ERROR AND ALARM MANAGEMENT MAY HAVE BEEN FACTORS. IF ADDITIONAL INFORMATION IS OBTAINED, PLEASE REQUEST RE-ASSESSMENT OF THE RECORD BY THE PMS CLINICAL EXPERT. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: PHILIPS RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT WAS AN EMERGENCY 32 WEEK DELIVERY THAT HAD ALVEOLI HEMORRHAGING AND NEONATE CARDIAC ARREST. THE PATIENT HAD TO BE DELIVERED PREMATURELY VIA C-SECTION DUE TO COMPLICATIONS. PATIENT ALMOST PASSED IN THE OR ROOM, BUT THEY WERE ABLE TO STABILIZE AND GET THE PATIENT TO THE NICU. THE FIELD SERVICE ENGINEER (FSE) MET WITH THE HOSPITAL RISK MANAGER TO EVALUATE THE DEVICE FROM THE EVENT. THE MONITOR SPEAKER WAS FUNCTIONING NORMALLY, AND THE SPEAKER SETTING WAS AT 7 AND RED ALARM WAS +1. THEY TESTED THE EXISTING CABLES AND LEADS THAT WAS USED DURING THE TIME OF THE EVENT. THE MONITORS SPO2 AND ECG ALARMED NORMALLY. RISK MANAGER ALSO WANTED TO KNOW WHAT THE "PAUSE ALARM" TIME LAPSE WAS AND/OR IF IT COULD BE CHANGED AFTER PRESSING THE "PAUSE ALARM" KEY. THE "PAUSE ALARM" WAS SET AT THE DEFAULT OF 2 MINUTES, AND THIS MONITOR CONFIG DIDN''T PROVIDE THE ABILITY TO CHOSE THE "PAUSE ALARM" TIME. IT WAS ALSO STATED THE DEVICE WAS NOT CONNECTED TO A PIC IX NETWORK. THIS DEVICE WAS CONNECTED TO A CAPSULE DEVICE. IT WAS FOUND THAT THE STAFF NEVER ADMITTED THE PATIENT INTO THE MONITOR AND NEVER DID AN "END CASE" AFTER THIS PATIENT OR THE PATIENTS PRIOR TO OR AFTER THIS EVENT. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL PRODUCT SPECIALIST REVIEWED THE LOG DATA. THE ALARMS SHOWED THAT THE STAFF ACKNOWLEDGED THE ALARMS. THE ALARM REVIEW HAS THE HISTORY OF THE ALARM NOTIFICATIONS WITH A MAXIMUM OF 300 ENTRIES. FIRST IN FIRST OUT STORAGE OF ALARMS. SO TIME IS OF ESSENCE, SINCE THE OLDEST ALARMS WILL BE OVERWRITTEN ONCE THE MAXIMUM IS REACHED. A REPORT CAN BE PRINTED IN BOTH PORTRAIT AND LANDSCAPE ORIENTATION. CAUSE: ALARMS WERE PAUSED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE PAUSE TIME COULD NOT BE ADJUSTED AS THE CONFIGURED DEFAULT WAS 2 MINUTES. BASED ON THIS INFORMATION, THE DEVICE FUNCTIONED AS CONFIGURED; HOWEVER, USE ERROR AND ALARM MANAGEMENT MAY HAVE BEEN FACTORS. THE CUSTOMER WAS PROVIDED WITH INFORMATION FROM REVIEW. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263940&pc=BZD | 2518422-2025-101425 | 2024-10-24 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DS2ADV AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND NO ISSUES WERE IDENTIFIED DURING THE ASSESSMENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263973&pc=MVK | 3008642652-2025-01032 | 2024-11-27 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-29 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 13:41:59 ON (B)(6) 2024. AT 12:56:26 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 190 BPM. AT 12:57:36, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 12:59:46, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 13:00:58, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 13:01:24, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 13:02:14, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AT 13:03:03, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 13:54:54 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263979&pc=DSQ | 2916596-2025-00335 | 2022-08-01 04:00 | Death | THORATEC CORPORATION | 2025-01-29 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "VALIDATION OF THE HEARTMATE 3 SURVIVAL RISK SCORE IN A LARGE LEFT VENTRICULAR ASSIST DEVICE CENTER" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH RIGHT VENTRICULAR FAILURE, INFECTION, BLEEDING, STROKE, PUMP THROMBOSIS, AND DEATH. A SINGLE-CENTER RETROSPECTIVE REVIEW OF ALL HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE RECIPIENTS BETWEEN SEPTEMBER 2017 AND AUGUST 2022. ONE- AND 2-YEAR KAPLAN¿MEIER SURVIVAL ANALYSES WERE CONDUCTED. A UNIVARIATE AND MULTIVARIABLE COX REGRESSION MODEL WAS USED TO IDENTIFY PREDICTORS. THE HEARTMATE 3 SURVIVAL RISK SCORE (HM3RS) WAS RECENTLY VALIDATED IN THE MULTICENTER STUDY OF MAGLEV TECHNOLOGY IN PATIENTS UNDERGOING MECHANICAL CIRCULATORY SUPPORT THERAPY WITH HEARTMATE 3 TO PREDICT PATIENT-SPECIFIC SURVIVAL IN HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE CANDIDATES. THE HEARTMATE 3 SURVIVAL RISK SCORE STRATIFIES INDIVIDUALS INTO TERTILES ACCORDING TO SURVIVAL PROBABILITY. THE PRIMARY END POINT OF THIS STUDY WAS TO VALIDATE THE HM3RS IN OUR COHORT. THE SECONDARY END POINTS WERE TO IDENTIFY PREDICTORS OF MORTALITY AND TO FIND OUR COHORT¿S OPTIMAL CUTOFF VALUE. A TOTAL OF 181 PATIENTS WERE INCLUDED IN THIS FINAL ANALYSIS. THE STUDY INCLUDED ALL PATIENTS WHO RECEIVED AN HM3 LVAD AND WERE NOT ENROLLED IN THE MOMENTUM-3 CLINICAL TRIAL. PATIENTS WHO UNDERWENT DEVICE EXCHANGE FROM AN HM3-TO-HM3 DURING THE STUDY PERIOD WERE ONLY FOLLOWED FROM THE TIME OF THE FIRST HM3 IMPLANTATION. PATIENTS WERE FOLLOWED FOR A PERIOD OF 2 YEARS OR UNTIL THEY REACHED A STUDY OUTCOME: DEATH, HEART TRANSPLANTATION, LVAD EXPLANT OR DECOMMISSIONING, OR LOSS TO FOLLOW-UP. THE MEDIAN AGE AT IMPLANT WAS 62.3 YEARS, 82.9% MALE, AND OF THE PATIENT POPULATION 22.1% WERE AFRICAN-AMERICAN, 3.3% WERE ASIAN, 21.5% WERE OTHER, AND THE MAJORITY WERE WHITE. THE MAJORITY OF PATIENTS WERE INTERAGENCY REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT (INTERMACS) PROFILE 1 (26.5%) OR 2 (42%). PRE-EXISTING CONDITIONS AND CO-MORBIDITIES INCLUDED: AN ISCHEMIC CAUSE OF HEART FAILURE (43.1%), HISTORY OF STROKE (0.6%), ATRIAL FIBRILLATION (13.8%), NYHA CLASS I (0.6%), CLASS II (1.7%), CLASS III (6.6%), CLASS IV (90.1%). THE INTENDED GOAL WAS DESTINATION THERAPY IN 77.9% OF PATIENTS. TRICUSPID VALVE REGURGITATION MODERATE/WORSE (39.8%). 5.5% (N = 10) OF PATIENTS WERE ON EXTRACORPOREAL MECHANICAL OXYGENATION BEFORE LVAD IMPLANTATION, 6.1% (N = 11) HAD A SERUM TOTAL BILIRUBIN GREATER THAN 2.5 MG/DL, 2.2% (N = 4) HAD A SERUM CREATININE LEVEL GREATER THAN 2.5 MG/DL, 13.2%(N ¼ 24) HAD PLATELET COUNTS BELOW 100,000/ML, AND 15.5%(N = 28) HAD ALBUMIN LEVELS LESS THAN 3 G/DL. THE MEAN HEARTMATE 3 SURVIVAL RISK SCORE FOR THE ENTIRE COHORT WAS 2.66 ± 0.66. TWO-YEAR SURVIVALS IN THE HIGH, AVERAGE, AND LOW SURVIVAL GROUPS WERE 93.5% ± 3.2%, 81.6% ± 7.4%, AND 82.0% ± 6.6%, RESPECTIVELY. AS A CONTINUOUS VARIABLE, THE UNADJUSTED HEARTMATE 3 SURVIVAL RISK SCORE WAS A SIGNIFICANT PREDICTOR OF MORTALITY (HAZARD RATIO, 2.20; 95% CI, 1.08-4.45; P = .029). THE AREAS UNDER THE CURVE WERE 0.70 AND 0.66 AT 1 AND 2 YEARS, RESPECTIVELY. WE WERE UNABLE TO DEMONSTRATE THE DISCRIMINATORY ABILITY OF THE HEARTMATE 3 SURVIVAL RISK SCORE USING THE ORIGINAL STRATIFICATION, BUT WE FOUND SIGNIFICANTLY INCREASED SURVIVAL IN THE HIGH SURVIVAL GROUP USING A BINARY CUTOFF (HAZARD RATIO, 4.8; 95% CI, 1.01-20.9; P = .038). THE HM3RS PROVIDES AN ESTIMATE OF THE INDIVIDUAL SURVIVAL PROBABILITY AT 2 YEARS AND CONSISTS OF 6 PARAMETERS: AGE (YEARS), HISTORY OF CORONARY ARTERY BYPASS GRAFTING (CABG) OR VALVULAR SURGERY, SERUM BLOOD UREA NITROGEN (MG/DL), SERUM SODIUM (MMOL/L), LEFT VENTRICULAR END-DIASTOLIC DIAMETER LESS THAN 5.5 CM, AND THE RATIO OF RIGHT ATRIAL PRESSURE TO PULMONARY CAPILLARY WEDGE PRESSURE GREATER THAN 0.6. ADVERSE EVENTS AT 2 YEARS AFTER HM3 LVAD IMPLANTATION INCLUDED: DEATH (17 PATIENTS), TRANSPLANT (32 PATIENTS), AND EXPLANT/DECOMMISSION (5 PATIENTS) DURING A MEDIAN OF 16.6 (8.36-27.8) MONTHS OF FOLLOW-UP. REASONS FOR EXPLANT/DECOMMISSION INCLUDED INFECTION (N = 2), MYOCARDIAL RECOVERY (N = 2), AND SELF-DISCONNECTION (N = 1). 3 PATIENTS UNDERWENT AN EXCHANGE. THE HM3RS STRATIFIED 68 PATIENTS INTO THE HIGH SURVIVAL GROUP, 58 PATIENTS INTO THE AVERAGE SURVIVAL GROUP, AND 55 PATIENTS INTO THE LOW SURVIVAL GROUP. THE UNADJUSTED HEARTMATE 3 SURVIVAL RISK SCORE WAS ASSOCIATED WITH POSTIMPLANT SURVIVAL IN PATIENTS OUTSIDE OF THE MULTICENTER STUDY OF MAGLEV TECHNOLOGY IN PATIENTS UNDERGOING MECHANICAL CIRCULATORY SUPPORT THERAPY WITH HEARTMATE 3 BUT DID NOT REMAIN AN INDEPENDENT PREDICTOR AFTER ADJUSTING FOR ISCHEMIC ETIOLOGY AND SEVERE DIABETES. THE HEARTMATE 3 SURVIVAL RISK SCORE WAS ABLE TO IDENTIFY PATIENTS AT HIGH SURVIVAL USING A BINARY CUTOFF, BUT WE WERE UNABLE TO DEMONSTRATE ITS DISCRIMINATORY ABILITY AMONG THE PREVIOUSLY PUBLISHED RISK TERTILES. (1) THE HM3RS WAS A SIGNIFICANT PREDICTOR OF MORTALITY IN THE UNIVARIATE (UNADJUSTED) REGRESSION ANALYSIS; (2) THE HM3RS WAS ABLE TO IDENTIFY PATIENTS AT HIGH RISK USING A BINARY CUTOFF, BUT NOT WHEN PATIENTS WERE SEPARATED INTO TERTILES; AND (3) AFTER ADJUSTING FOR ISCHEMIC ETIOLOGY AND SEVERE DIABETES IN A MULTIVARIABLE ANALYSIS, THE HM3RS IS NO LONGER AN INDEPENDENT PREDICTOR OF MORTALITY. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01AUG2022 AS DATA WAS COLLECTED BETWEEN SEPTEMBER 2017 AND AUGUST 2022. AUTHOR INFORMATION: MOELLER, C. M., RUBINSTEIN, G., OREN, D., VALLEDOR, A. F., LOTAN, D., RAIKHELKAR, J. K., CLERKIN, K. J., COLOMBO, P. C., LEAHY, N. E., FRIED, J. A., KAKU, Y., NAKA, Y., TAKEDA, K., YUZEFPOLSKAYA, M., TOPKARA, V. K., SAYER, G. T., & URIEL, N. (2024). VALIDATION OF THE HEARTMATE 3 SURVIVAL RISK SCORE IN A LARGE LEFT VENTRICULAR ASSIST DEVICE CENTER. THE JOURNAL OF THORACIC AND CARDIOVASCULAR SURGERY. HTTPS://DOI.ORG/10.1016/J.JTCVS.2024.03.009. DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE AND DEPARTMENT OF CARDIOTHORACIC SURGERY, COLUMBIA UNIVERSITY IRVING MEDICAL CENTER, NEW YORK, NY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264090&pc=BZD | 2518422-2025-101427 | 2024-10-24 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS BEING USED IN A NURSING HOME, AND A FAMILY MEMBER INDICATED THAT, HAD THEY STILL HAD POSSESSION OF IT, THEY WOULD HAVE RETURNED IT FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264109&pc=DSQ | 2916596-2025-00601 | 2023-05-01 04:00 | Death | THORATEC CORPORATION | 2025-01-29 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿DURABLE LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION IN JEHOVAH¿S WITNESS PATIENTS¿ THAT THE HEARTMATE 3 MAY BE RELATED TO DEATH, GI BLEEDING, BLEEDING, ARRYTHMIAS, AND NEUROLOGICAL EVENTS. LEFT VENTRICULAR ASSIST DEVICES ARE A CONVENTIONAL THERAPY FOR PATIENTS WITH END-STAGE HEART FAILURE, BUT ARE OFTEN NOT OFFERED TO JEHOVAH¿S WITNESS PATIENTS BECAUSE OF THEIR REFUSAL TO RECEIVE ALL BLOOD PRODUCTS. THE PURPOSE OF THIS STUDY WAS TO EXAMINE THE OUTCOMES OF A SINGLE CENTERS HIGH VOLUME EXPERIENCE OF JEHOVAH¿S WITNESS PATIENTS UNDERGOING LVAD IMPLANTATION. THE PATIENTS WHO UNDERWENT IMPLANTATION WAS BETWEEN JANUARY 2012 AND MAY 2023. BASELINE DEMOGRAPHICS, DIAGNOSTICS, ADVERSE EVENTS, TREATMENT MODALITIES, OUTCOMES AND SURVIVAL WERE ANALYZED. THIRTY-FIVE JEHOVAH¿S WITNESS PATIENTS UNDERWENT LVAD IMPLANTATION DURING THE STUDY AND THE AVERAGE PRE-OPERATIVE HEMOGLOBIN WAS 11.6 WITH AN AVERAGE HEMOGLOBIN OF 9.2 PRIOR TO DISCHARGE. PRE-OPERATIVELY 57% REQUIRED INOTROPES AND 37% REQUIRED ACUTE CIRCULATORY SUPPORT. THE MOST COMMON POST-OPERATIVE ADVERSE EVENT WAS ARRHYTHMIAS AT 63%. 29%EXPERIENCED A BLEEDING EVENT DURING THEIR IMPLANT ADMISSION, WITH 17% RELATED TO GASTROINTESTINAL BLEEDING. 11%OF OVERALL DEATHS WERE CAUSED BY A NEUROLOGICAL EVENT, WITH ONLY 6% RELATED TO A BLEEDING EVENT. THIS IS THE LARGEST COHORT OF JEHOVAH¿S WITNESS PATIENTS AT A SINGLE CENTER WHO HAVE UNDER GONE LVAD IMPLANTATION. OUR FINDINGS SUGGEST THAT JEHOVAH¿S WITNESS PATIENTS CAN SAFELY UNDER GO LVAD IMPLANTATION DESPITE THE REFUSAL OF BLOOD PRODUCTS. PATIENTS WITHIN THE JEHOVAH¿S WITNESS COMMUNITY SHOULD CONTINUE TO BE EVALUATED FOR LVAD THERAPY AND REFERRED TO AN EXPERIENCED, HIGH VOLUME BLOODLESS CENTER. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN JANUARY2012 AND MAY2023. CHICKERILLO, K., MEEHAN, K., CHINCO, A., LABUHN, C., MOORE, K., OKRAY, J., WRIGHT, N., HWANG, V., JAIN, S., & JEEVANANDAM, V. (2024). DURABLE LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION IN JEHOVAH¿S WITNESS PATIENTS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S280. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1218. (B)(6). THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING BLEEDING, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264177&pc=OZD | 1220648-2025-25873 | 2025-01-13 05:00 | Death | ABIOMED, INC. | 2025-01-29 05:00 | OZD | IMPELLA 5.5 | Detachment of Device or Device Component | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) FOLLOWED BY EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT SEVEN DAYS LATER. WHILE ON IMPELLA MCS, APPROXIMATELY 20 DAYS AFTER START OF SUPPORT, THE NURSED CALLED TO REPORT THAT THE DISTAL END OF THE STERILE SLEEVE HAD BECOME DETACHED AND WAS WRAPPED WITH TEGADERM. THE NURSED RAISED THE QUESTION REGARDING IF THERE IS ANY RISK OF INFECTION/CONTAMINATION WITH THE STERILE SLEEVE BEING SEPARATED AND WAS INFORMED TO CONSULT WITH THEIR PROVIDER AS ABIOMED HAS NO RECOMMENDATIONS PER THE INSTRUCTIONS FOR USE FOR A SEPARATED STERILE SLEEVE. FURTHER DISCUSSION NOTED THAT THE NURSE AND TEAM WERE ADVISED TO BE CONSCIENCE OF THE POTENTIAL RISK OF INFECTION WHEN/IF THE IMPELLA NEEDS REPOSITIONED. THE PATIENT WOULD EXPIRE FOUR DAYS LATER. THE FAMILY MADE THE DECISION TO WITHDRAW CARE. THE CAUSE OF DEATH WAS NOT PROVIDED, AND THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE USED AS A CONTRIBUTING FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF PRODUCT DAMAGE (STERILE SLEEVE DAMAGE) HAS BEEN COMPLETED. DUE TO LIMITED CLINICAL DETAILS AND NO PRODUCT RETURN, THE ROOT CAUSE OF THE STERILE SLEEVE DAMAGE COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25873 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264226&pc=DRA | 2135147-2025-00445 | 2025-01-07 05:00 | Death | ABBOTT MEDICAL | 2025-01-29 05:00 | DRA | MITRACLIP | Leak/Splash | Bradycardia; Cardiac Arrest; Non specific EKG/ECG Changes; Low Blood Pressure/ Hypotension; Myocardial Infarction; Tachycardia; Heart Failure/Congestive Heart Failure | K190167 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A SCHEDULED MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) GRADE 4. PT WAS ALLERGIC TO HEPARIN. ANGIOMAX WAS USED TO ANTICOAGULATE PATIENT AND ANTICOAGULANT CITRATE DEXTROSE SOLUTION USP (ACD) WAS USED TO PREP THE STEERABLE GUIDE CATHETER (SGC). THE SGC WAS INSERTED INTO THE PATIENT. ONCE THE SGC CROSSED THE SEPTUM AND THE DILATOR WAS REMOVED, AIR WAS NOTED IN THE SGC. THE SGC WAS TAKEN OFF THE STABILIZER AND HELD BELOW THE LEVEL OF THE PATIENT'S HEART AND ASPIRATION WAS PERFORMED UNTIL BLOOD FILLED THE HEMOSTASIS VALVE CHAMBER. PATIENT DEVELOPED ST ELEVATION IN ALL LEADS WITH BRADYCARDIA AND HYPOTENSION 70/40. ANESTHETIC MEDICATION WAS ADMINISTERED TO TREAT THE LOW BLOOD PRESSURE, BUT DURING THIS THE PATIENT'S RHYTHM SLOWED INTO ASYSTOLE. CARDIO-PULMONARY RESUSCITATION (CPR) WAS PERFORMED, AFTER WHICH THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA. CARDIOVERSION WAS THEN PERFORMED WHICH RECOVERED THE RHYTHM BUT IT WAS TACHYCARDIC GREATER THAN 130 BEATS PER MINUTE (BPM). BLOOD PRESSURE STILL REMAINED LOW IN 70/40 RANGE. ANESTHESIA CONTINUED TO WORK ON THE HYPOTENSION WHILE A CORONARY ANGIOGRAPHY WAS PERFORMED WHICH REVEALED CORONARIES WERE PATENT. ECHOCARDIOGRAM WAS ALSO PERFORMED WHICH SHOWED NO EFFUSION AND DECREASED VENTRICULAR FUNCTION FROM APPROXIMATELY 40% AT BASELINE TO 10-15%. BASAL SECTION OF THE HEART WAS CONTRACTING WELL BUT BOTH THE RIGHT AND LEFT VENTRICULAR FREE WALLS AND APEX CONTRACTILITY WAS DIMINISHED. NO AIR WAS OBSERVED IN THE ANATOMY. ANESTHESIA WAS ABLE TO RECOVER BLOOD PRESSURE SO THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU) FOR RECOVERY. LATER ON THE SAME NIGHT, THE PATIENT DIED. THE FAMILY DID NOT WANT EXTRAORDINARY MEASURES TO BE TAKEN. THE CAUSE OF THE PATIENT DETERIORATION DURING THE PROCEDURE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY DID NOT INDICATE A LOT-SPECIFIC QUALITY ISSUE. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED LEAK, DEATH, ST ELEVATION IN ALL LEADS, BRADYCARDIA, AND SUBSEQUENT HYPOTENSION. THE PATIENT EFFECTS OF MYOCARDIAL INFARCTION, CARDIA ARREST, TACHYCARDIA, AND HEART FAILURE APPEAR TO BE CASCADING EFFECTS. DEATH, BRADYCARDIA, HYPOTENSION, MYOCARDIAL INFARCTION, CARDIA ARREST, TACHYCARDIA, HEART FAILURE, AND NON-SPECIFIC EKG/ECG CHANGES ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED MEDICATION REQUIRED AND UNEXPECTED MEDICAL INTERVENTIONS WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES AS ANESTHETIC MEDICATION WAS ADMINISTERED TO TREAT THE LOW BLOOD PRESSURE, ADDITIONAL ASPIRATION WAS PROVIDED, CARDIO-PULMONARY RESUSCITATION (CPR) AND CORONARY ANGIOGRAPHY WERE PERFORMED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264324&pc=BZD | 2518422-2025-101431 | 2024-10-25 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE USER PASSED AWAY YEARS AGO. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS GIVEN TO A FRIEND AND IS NOT AVAILABLE TO BE RETURNED FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264529&pc=DSQ | 2916596-2025-00383 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-01-29 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT EXPIRED DUE TO AN AT HOME MECHANICAL FALL. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. IT WAS COMMUNICATED THERE WERE NO LOG FILES SINCE THE PATIENT IS DECEASED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264572&pc=OZD | 1220648-2025-25876 | 2025-01-16 05:00 | Death | ABIOMED, INC. | 2025-01-29 05:00 | OZD | IMPELLA 5.5 | Increase in Pressure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE US COMPLAINANT HAD A 5.5 PUMP SUPPORT PLACED TO SUPPORT A PATIENT ADMITTED IN AND ESCALATED FROM CP AND EXTRA CORPOREAL MEMBRANE OXYGENATION. THE PUMP HAD HIGH PURGE PRESSURES WHICH PROMPTED TROUBLESHOOTING. THE PURGE CASSETTE WAS REPLACED BUT THE HIGH PURGE PRESSURE ALARMS WERE NOT RESOLVED. THE TEAM USED TISSUE PLASMINOGEN ACTIVATOR IN THE PURGE AND THE PUMP REMAINED ON FOR SUPPORT AFTER 41 DAYS NOW AS CARE/COMFORT VERSUS TRANSPLANT BEING DISCUSSED. THE PATIENT EXPIRED WHEN THE PATIENT WAS MADE CARE/COMFORT AFTER DAY 42. Manufacturer Narrative: THE INVESTIGATION FOR THE HIGH PURGE PRESSURE HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR EVALUATION. DATA LOGS WERE REVIEWED AND LONG TERM LOG AT TIME OF ¿HIGH PURGE PRESSURE¿ ALARMS SHOWED A RISE IN PURGE PRESSURE OVER A PERIOD OF APPROXIMATELY 30 MINUTES PRIOR TO ALARMS. NO CHANGES IN MOTOR CURRENT WERE OBSERVED AT TIME OF ALARMS. FULL LONG TERM LOGS OF THE CASE SHOWED NO MOTOR CURRENT SPIKES PRIOR TO HIGH PURGE PRESSURE ALARMS. PURGE PRESSURE REMAINS HIGH FOR THE REMAINDER OF THE CASE. CLINICAL DETAILS NOTED THAT THE PATIENT WAS ON SUPPORT FOR 43 DAYS BEFORE "HIGH PURGE PRESSURE" ALARMS WERE TRIGGERED. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY DUE TO BIOMATERIAL SINCE THERE WAS NO MOTOR CURRENT SPIKE PRECEDING THE PURGE PRESSURE RISE. D.4 REVISED SERIAL NUMBER AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25876. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25876 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264580&pc=NGV | 2135147-2025-00448 | 2025-01-09 05:00 | Death | ABBOTT MEDICAL | 2025-01-29 05:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM AMPLATZER AMULET WAS IMPLANTED WITH NO REPORTED COMPLICATIONS. THE PATIENT'S CONDITION PRIOR TO IMPLANT WAS STABLE. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. IT IS UNKNOWN IF THE PATIENT'S DEATH WAS CLASSIFIED AS RELATED TO THE IMPLANTED AMULET OR THE PROCEDURE.Event Description: SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM AMPLATZER AMULET WAS IMPLANTED WITH NO REPORTED COMPLICATIONS. THERE WERE NO IMPLANT COMPLICATIONS OR POST-IMPLANT AND THE PATIENT WAS DISCHARGED FROM THE HOSPITAL. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. IT IS UNKNOWN IF THE PATIENT'S DEATH WAS CLASSIFIED AS RELATED TO THE IMPLANTED AMULET OR THE PROCEDURE. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE AMPLATZER AMULET. AN EVENT FOR PATIENT EFFECTS OF DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, A CAUSE FOR THE REPORTED DEATH COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE DUE TO PROCEDURE CONDITIONS/PATIENT CONDITIONS; HOWEVER, THIS CANNOT BE CONFIRMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FURTHERMORE, THIS COMPLAINT WAS REVIEWED BY MEDICAL AFFAIRS. THE REVIEWER STATED THAT,"NO IMAGING OF THE PROCEDURE OR SUBSEQUENT EVENTS WAS PROVIDED. WITHOUT SUCH IMAGING IT IS NOT POSSIBLE TO DETERMINE WHETHER THIS EVENT SHOULD BE ATTRIBUTED TO THE DEVICE OR TO UTILIZATION OUTSIDE OF THE AMULET IFU." CODES REMOVED: H6 HEALTH EFFECT-CLINICAL CODE: 4582. CODES ADDED: H6 HEALTH EFFECT-CLINICAL CODE: 4580.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264671&pc=PNJ | 9612164-2025-00492 | 2024-04-01 04:00 | Death | MEDTRONIC IRELAND | 2025-01-29 05:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Pericardial Effusion; Renal Impairment |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING PERICARDIAL EFFUSION REQUIRING INTERVENTION IN PATIENTS UNDERGOING LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS IN PATIENTS WITH AND WITHOUT THE DEVELOPMENT OF PERICARDIAL EFFUSION REQUIRING INTERVENTION. THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED PROCEDURE RELATED COMPLICATIONS WHICH INCLUDED PERICARDIAL EFFUSION WHICH REQUIRED PERCUTANEOUS DRAINAGE OR SURGICAL INTERVENTION, ATRIOVENTRICULAR (AV) FISTULA, PSEUDOANEURYSM, HEMATOMA, RETROPERITONEAL BLEEDING, VENOUS THROMBOEMBOLISM, AND ACUTE KIDNEY INJURY. THE STATUS OF THE LEADLESS IPGS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/80 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PERICARDIAL EFFUSION REQUIRING INTERVENTION IN PATIENTS UNDERGOING LEADLESS PACEMAKER IMPLANTATION: A REAL-WORLD ANALYSIS FROM THE NATIONAL INPATIENT SAMPLE DATABASE. HEART RHYTHM O2. 2024;5:217¿223. DOI: 10.1016/J.HROO.2024.02.004. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME MICRA PRODUCT ID MDT-TPS-DELSYS SERIAL: UNKNOWN D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264712&pc=BZD | 2518422-2025-101439 | 2024-12-02 05:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY LAST YEAR. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264713&pc=BZD | 2518422-2025-101435 | 2024-10-25 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO BIPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264863&pc=NIQ | 2024168-2025-01000 | 2025-01-06 05:00 | Death | ABBOTT VASCULAR INC. | 2025-01-29 05:00 | NIQ | XIENCE ALPINE | Adverse Event Without Identified Device or Use Problem | Tachycardia; Thrombosis/Thrombus | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025 THE PATIENT PRESENTED WITH ANGINA AND ANGIOGRAM REVEALED A TOTALLY OCCLUDED LESION IN THE RIGHT CORONARY ARTERY (RCA) WITH MILD CALCIFICATION AND HEAVY TORTUOSITY. THROMBUS ASPIRATION WAS PERFORMED AND THE 3X28MM XIENCE ALPINE STENT WAS IMPLANTED. HOWEVER, ON (B)(6) 2025 THE PATIENT HAD VENTRICULAR TACHYCARDIA AND WAS RETURNED TO THE CATH LAB. STENT THROMBOSIS WAS SUSPECTED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR 1 HOUR; HOWEVER, THE PATIENT EXPIRED PRIOR TO HAVING AN ANGIOGRAPHY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECT(S) OF DEATH AND THROMBOSIS ARE LISTED IN THE XIENCE ALPINE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265009&pc=DSQ | 2916596-2025-00405 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-01-29 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Headache; Hematoma; Visual Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025 DUE TO A DEVELOPING RESPIRATORY SYNCYTIAL VIRUS (RSV) INFECTION OF UNKNOWN SOURCE. THE PATIENT HAD TO USE THE RESTROOM AND BEGAN COMPLAINING OF VISUAL CHANGES AND HEADACHE, AND A CODE STROKE WAS CALLED. A COMPUTED TOMOGRAPHY (CT) SCAN WAS TAKEN OF THE PATIENT'S HEAD AND FOUND A LARGE SUBDURAL HEMATOMA OF MIXED DENSITY. THE CAUSE OF THE HEMATOMA WAS UNKNOWN, AND THE CUSTOMER PRESUMED THAT THE PATIENT FELL AT HOME FOR AN UNKNOWN REASON AND DID NOT INFORM ANYONE. NO ALARMS WERE IDENTIFIED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) WHERE THEY DECOMPENSATED. THE PATIENT WAS THEN TAKEN TO THE OPERATING ROOM FOR A DECOMPRESSION, WHICH THE PATIENT TOLERATED WELL. NEUROSURGERY SPOKE WITH THE PATIENT'S FAMILY THE NEXT MORNING, AND THE FAMILY CHOSE TO MOVE THE PATIENT TO COMFORT CARE. HOSPITAL CARE WAS WITHDRAWN, AND THE PATIENT PASSED AWAY FROM THE SUBDURAL HEMATOMA ON (B)(6) 2025. THE OUTCOME WAS STATED NOT TO BE LVAD DEVICE OR THERAPY RELATED, AND THE CUSTOMER STATED THE DEVICE FUNCTIONED AS INTENDED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND THE HEARTMATE 3 PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265094&pc=NAY | 2955842-2025-00965 | 2018-09-14 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-29 05:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH BREAST CANCER UNDERWENT A DA VINCI-ASSISTED BILATERAL NIPPLE-SPARING MASTECTOMY WITH IMMEDIATE RECONSTRUCTION IN (B)(6) 2018. THE PATIENT PASSED AWAY 2 YEARS POST-PROCEDURE. THE REPORTER DID NOT PROVIDE A CAUSE OF DEATH. THERE WAS NO REPORT THAT ANY DA VINCI DEVICE MALFUNCTION OR INTRAOPERATIVE ISSUES OCCURRED DURING THE PROCEDURE. REQUESTS FOR ADDITIONAL INFORMATION FROM THE OPERATING SURGEON HAVE BEEN SENT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE DEVICE LOG FOUND THAT THE FOLLOWING INSTRUMENTS WERE USED IN THE PROCEDURE: ENDOSCOPE, TWO MONOPOLAR CURVED SCISSORS, TWO FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. ALL INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES WITH THE EXCEPTION OF TWO OF THE INSTRUMENTS; ONE SINGLE-USE INSTRUMENT, AND ONE INSTRUMENT THAT HAD REACHED ITS FINAL LIFE USE LIMIT WITH THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. A PATIENT DEATH WAS REPORTED 2 YEARS AFTER A BILATERAL NIPPLE SPARING MASTECTOMY. NO OTHER INTRAOPERATIVE OR POST OPERATIVE DETAILS ARE PROVIDED. HOW THEY DIED OR THE EVENTS LEADING TO THEIR DEATH ARE NOT PROVIDED. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE DEVICE LOG FOUND THAT THE FOLLOWING INSTRUMENTS WERE USED IN THE PROCEDURE: ENDOSCOPE, TWO MONOPOLAR CURVED SCISSORS, TWO FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. ALL INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES EXCEPT FOR TWO OF THE INSTRUMENTS THAT WERE EXPIRED AFTER THIS PROCEDURE AND SINGLE-USE INSTRUMENT. SECTION H10 OF THIS REPORT (RELATED REPORT NUMBER) FOR THE RECEIVED VOLUNTARY MEDWATCH REPORT NUMBER (MW5164582). SECTION E: THE USER FACILITY WHERE THE SURGERY WAS PERFORMED IS DOCUMENTED. THE REPORTER'S NAME AND CORRESPONDING INFORMATION IS NOT AFFILIATED WITH THE USER FACILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265231&pc=DSQ | 3007042319-2025-00123 | 2024-11-15 05:00 | Death | HEARTWARE, INC. | 2025-01-29 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT SUFFERED A DEVASTATING HEAD BLEED AND SUBSEQUENTLY DIED. THE VENTRICULAR ASSIST DEVICE (HVAD) WAS TURNED OFF. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE BLEEDING AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265557&pc=MIH | 2017233-2025-05769 | 2024-12-06 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-29 05:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Patient Device Interaction Problem | Hemorrhage/Bleeding; Cardiogenic Shock; Respiratory Failure; Vascular Dissection; Ischemia Stroke | P210032 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE THROUGH THE PIVOTAL STUDY FOR THE GORE® ASCENDING STENT GRAFT IN THE TREATMENT OF LESIONS OF THE ASCENDING AORTA (ASG DEVICE): ON (B)(6) 2024, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF AN ASCENDING AORTIC ANEURYSM IN THE ASG22-02 ARISE II PIVOTAL STUDY USING A GORE® ASCENDING STENT GRAFT (ASG), GORE® TAG® THORACIC BRANCH ENDOPROSTHESES (TBE), AND A GORE® TAG® CONFORMABLE THORACIC STENT GRAFTS WITH ACTIVE CONTROL SYSTEM. THE MOST PROXIMAL LANDING ZONE WAS ZONE 0C AND THE MOST DISTAL LANDING ZONE WAS A PREVIOUSLY IMPLANTED ENDOVASCULAR DEVICE. IT WAS REPORTED THAT AN AORTIC DISSECTION OCCURRED DURING PLACEMENT OF THE TBE AORTIC COMPONENT. THE SUSPECTED CAUSE OF THE DISSECTION WAS THE RELEASE AND DEPLOYMENT OF THE AORTIC COMPONENT. THE ASG DEVICE WAS THEN INADVERTENTLY DEPLOYED IN THE FALSE LUMEN OF THE NEW DISSECTION. THE PHYSICIAN CONVERTED TO OPEN REPAIR TO TREAT THE DISSECTION AND INACCURATE ASG PLACEMENT. AN AORTIC ROOT REPLACEMENT WAS PERFORMED ALONG WITH REPLACEMENT OF THE ASCENDING AORTA. CORONARY ARTERY BYPASS GRAFTING WAS ALSO PERFORMED. THE PATIENT EXPERIENCED ACUTE BLOOD LOSS ANEMIA AND WAS PROVIDED INTRAOPERATIVE TRANSFUSION OF 10 UNITS OF PACKED RED BLOOD CELLS. THE PATIENT ALSO EXPERIENCED CARDIOGENIC SHOCK, WHICH WAS TREATED WITH VASOPRESSORS, AND ACUTE HYPOXIC RESPIRATORY FAILURE. THE PATIENT WAS PLACED ON A VENTILATOR. ON (B)(6) 2024, THE PATIENT EXPERIENCED A CEREBROVASCULAR ACCIDENT DUE TO EMBOLI. THE PATIENT WAS TREATED IN THE ICU. ON (B)(6) 2024, IT WAS REPORTED THAT MEDICAL CARE HAD BEEN WITHDRAWN AND COMFORT MEASURES WERE PUT IN PLACE. PATIENT DEATH WAS NOTED ON (B)(6) 2024. Manufacturer Narrative: D.10. CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: ASKED BUT UNAVAILABLE H.6. INVESTIGATION FINDINGS FOR ANALYSIS OF PRODUCTION RECORDS: CODE C21 UPDATED TO CODE C19 H.6. INVESTIGATION FINDINGS FOR ANALYSIS OF PRODUCTION RECORDS: CODE C19 - THE REVIEW OF THE MANUFACTURING PAPERWORK VERIFIED THAT THE LOTS INVOLVED IN THIS EVENT MET ALL PRE-RELEASE SPECIFICATIONS. H.6. INVESTIGATION CONCLUSIONS: CODE D16 UPDATED TO CODE D15. H.6. INVESTIGATION CONCLUSIONS: CODE D15 - THERE IS INSUFFICIENT INFORMATION AVAILABLE FOR GORE TO REASONABLY DRAW CONCLUSIONS RELATED TO ASPECTS OF THE EVENT, THEREFORE CONCLUSION CODE ¿D15: CAUSE NOT ESTABLISHED¿ IS BEING USED. INSUFFICIENT INFORMATION MAY INCLUDE LIMITED OR MISSING RELEVANT MEDICAL RECORDS, INVOLVEMENT OF MULTIPLE IMPLANTED DEVICES (INCLUDING NON-GORE DEVICES) IN THE FIELD OF TREATMENT, PATIENT NON-COMPLIANCE, AND/OR A GENERAL LACK OF AVAILABLE DETAIL OR SPECIFICITY RELATED TO AN ADVERSE EVENT AND/OR DEVICE. H.6. INVESTIGATION CONCLUSIONS: CODE D12 ADDED. ACCORDING TO THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE, POTENTIAL ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INCLUDE, BUT ARE NOT LIMITED TO, DISSECTION OF THE AORTIC VESSEL AND/OR SURROUNDING VASCULATURE, SURGICAL CONVERSION, BLEEDING (PROCEDURAL AND POST-TREATMENT), CARDIAC COMPLICATIONS (E.G., ARRHYTHMIA, MYOCARDIAL INFARCTION, CONGESTIVE HEART FAILURE, HYPOTENSION OR HYPERTENSION), NEUROLOGIC DAMAGE, LOCAL OR SYSTEMIC (E.G., STROKE) AND DEATH. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265633&pc=CAW | 2515872-2025-00003 | 2023-12-29 05:00 | Death | DEVILBISS HEALTHCARE | 2025-01-29 05:00 | CAW | DEVILBISS HEALTHCARE | Insufficient Information | Low Oxygen Saturation | K071397 |  | 1 | Event Description: DEVILBISS HEALTHCARE WAS NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY SERVICE OF A LEGAL COMPLAINT, WHICH ALLEGES THAT THE CONCENTRATOR WAS EMITTING NOISE AND ALARMING, BUT THAT THE DME PROVIDER DID NOT COME TO SERVICE, REPAIR OR REPLACE THE UNIT UNTIL SEVEN DAYS AFTER BEING NOTIFIED OF THE PROBLEM, WHICH ALLEGEDLY RESULTED IN THE PATIENT HAVING LOW OXYGEN SATURATION, DESPITE THE PATIENT ALSO DEMONSTRATING LOW SATURATION ON A DIFFERENT PORTABLE DEVICE. THE ALARMING DEVICE WAS REPLACED, AND IT WAS REPORTED THAT THE PATIENT PASSED AWAY THREE DAYS LATER. DEVILBISS HEALTHCARE WILL COOPERATE AND PARTICIPATE IN THE LITIGATION PROCESS TO GATHER FURTHER INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265761&pc=MCX | 3004742232-2025-00027 | 2025-01-11 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-01-29 05:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Vasoconstriction; Obstruction/Occlusion | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT THE MID AND PROXIMAL SEGMENTS OF A LEFT ANTERIOR DESCENDING (LAD) ARTERY. POST ORBITAL ATHERECTOMY, THE PHYSICIAN WENT TO PRE-DILATE THE LESION WITH A NON-COMPLIANT BALLOON, BUT PRIOR TO THE PRE-DILATION, THE PATIENT SUDDENLY WENT INTO CARDIAC ARREST. THE PATIENT WAS ATTEMPTED TO BE REVIVED, HOWEVER, THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS RECEIVED. IN THE PHYSICIAN'S OPINION, THE PRIMARY CAUSE OF DEATH WAS BRADYCARDIA AND SLOW/NO REFLOW; AND THE SECONDARY CAUSE OF DEATH WAS ABRUPT VESSEL CLOSURE. MEDICATION WAS ADMINISTERED TO IMPROVE SLOW FLOW AND ADDRESS THE BRADYCARDIA.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT THE MID AND PROXIMAL SEGMENTS OF A LEFT ANTERIOR DESCENDING (LAD) ARTERY. POST ORBITAL ATHERECTOMY, THE PHYSICIAN WENT TO PRE-DILATE THE LESION WITH A NON-COMPLIANT BALLOON, BUT PRIOR TO THE PRE-DILATION, THE PATIENT SUDDENLY WENT INTO CARDIAC ARREST. THE PATIENT WAS ATTEMPTED TO BE REVIVED, HOWEVER, THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS RECEIVED. IN THE PHYSICIAN'S OPINION, THE PRIMARY CAUSE OF DEATH WAS BRADYCARDIA AND SLOW/NO REFLOW; AND THE SECONDARY CAUSE OF DEATH WAS ABRUPT VESSEL CLOSURE. MEDICATION WAS ADMINISTERED TO IMPROVE SLOW FLOW AND ADDRESS THE BRADYCARDIA. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT ARRHYTHMIA, CARDIAC ARREST, OBSTRUCTION, DEATH, VASOCONSTRICTION AND RELATED MEDICAL INTERVENTION APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS LIKELY THAT THE REPORTED PATIENT ARRHYTHMIA, CARDIAC ARREST, OBSTRUCTION, DEATH, VASOCONSTRICTION AND RELATED MEDICAL INTERVENTION ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. PRODUCT PERFORMANCE ENGINEERING REVIEWED THE INCIDENT INFORMATION; HOWEVER, THE PRODUCT WAS NOT RETURNED TO ABBOTT VASCULAR FOR ANALYSIS. FACTORS THAT MAY CONTRIBUTE TO THE PATIENT ARRHYTHMIA, CARDIAC ARREST, OBSTRUCTION, DEATH, VASOCONSTRICTION AND RELATED MEDICAL INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO PATIENT ANATOMICAL MORPHOLOGY, PATIENT DISEASE STATE, DEVICE PLACEMENT OR USE TECHNIQUES EMPLOYED. DUE TO THE VERY LIMITED INFORMATION PROVIDED, IT CANNOT BE ESTABLISHED IF THE OAD CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE REPORTED PATIENT EFFECTS OF ARRHYTHMIA, CARDIAC ARREST, OBSTRUCTION, DEATH AND VASOCONSTRICTION ARE LISTED IN THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE INSTRUCTIONS FOR USE AS A KNOWN POTENTIAL PATIENT EFFECT ASSOCIATED WITH THE USE OF THE DEVICE. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266179&pc=NGV | 2124215-2025-04214 | 2025-01-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-29 05:00 | NGV | WATCHMAN FLX PRO | Device Dislodged or Dislocated | Hypoxia; Mitral Valve Insufficiency/ Regurgitation; No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION OCCURRED. THE DEVICE WAS EXPLANTED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. VENOUS FEMORAL ACCESS WAS OBTAINED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS USED FOR INTRAPROCEDURAL IMAGING. A DOUBLE CURVE WATCHMAN ACCESS SYSTEM (WAS) WAS INSERTED AND POSITIONED AT THE LAA. A CONTRAST INJECTION WAS PERFORMED AND MEASUREMENTS OF THE LAA TAKEN. A 35MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS ADVANCED AND THE CLOSURE DEVICE WAS DEPLOYED. RELEASE CRITERIA WAS DETERMINED TO BE MET, YET THE SHOULDER OF THE CLOSURE DEVICE WAS NOTED TO BE A THIRD TO BORDERLINE ONE HALF OF A SHOULDER AND THE PRESENCE OF A WATERMELON SEED DISTAL PUCK. AFTER THE PHYSICIAN PERFORMED A TUG TEST AND WAS CONTENT WITH THE STABILITY OF THE DEVICE, THE CLOSURE DEVICE WAS RELEASED AND IMPLANTED. THE CLOSURE DEVICE WAS OBSERVED WITH TEE AND THE PROCEDURE WAS CONCLUDED. DURING THE MIDDLE OF THE NIGHT WHEN THE PATIENT WAS STILL HOSPITALIZED, THE CLOSURE DEVICE WAS NOTED TO HAVE EMBOLIZED. THE PATIENT WAS TAKEN TO THE CARDIAC CATHETERIZATION LAB AND THE CLOSURE DEVICE WAS PERCUTANEOUSLY EXPLANTED BY RETROGRADE RETRIEVAL WITH A NON-BOSTON SCIENTIFIC (BSC) STEERABLE SHEATH AND GOOSENECK SNARE. THE PATIENT WAS THEN SENT TO THE CARDIOVASCULAR INTENSIVE CARE UNIT (CVICU) IN STABLE CONDITION TO BE MONITORED.Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION OCCURRED. THE DEVICE WAS EXPLANTED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. VENOUS FEMORAL ACCESS WAS OBTAINED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS USED FOR INTRAPROCEDURAL IMAGING. A DOUBLE CURVE WATCHMAN ACCESS SYSTEM (WAS) WAS INSERTED AND POSITIONED AT THE LAA. A CONTRAST INJECTION WAS PERFORMED AND MEASUREMENTS OF THE LAA TAKEN. A 35MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS ADVANCED AND THE CLOSURE DEVICE WAS DEPLOYED. RELEASE CRITERIA WAS DETERMINED TO BE MET, YET THE SHOULDER OF THE CLOSURE DEVICE WAS NOTED TO BE A THIRD TO BORDERLINE ONE HALF OF A SHOULDER AND THE PRESENCE OF A WATERMELON SEED DISTAL PUCK. AFTER THE PHYSICIAN PERFORMED A TUG TEST AND WAS CONTENT WITH THE STABILITY OF THE DEVICE, THE CLOSURE DEVICE WAS RELEASED AND IMPLANTED. THE CLOSURE DEVICE WAS OBSERVED WITH TEE AND THE PROCEDURE WAS CONCLUDED. DURING THE MIDDLE OF THE NIGHT WHEN THE PATIENT WAS STILL HOSPITALIZED, THE CLOSURE DEVICE WAS NOTED TO HAVE EMBOLIZED. THE PATIENT WAS TAKEN TO THE CARDIAC CATHETERIZATION LAB AND THE CLOSURE DEVICE WAS PERCUTANEOUSLY EXPLANTED BY RETROGRADE RETRIEVAL WITH A NON-BOSTON SCIENTIFIC (BSC) STEERABLE SHEATH AND GOOSENECK SNARE. THE PATIENT WAS THEN SENT TO THE CARDIOVASCULAR INTENSIVE CARE UNIT (CVICU) IN STABLE CONDITION TO BE MONITORED. IT WAS FURTHER REPORTED THE PATIENT HAD EXPERIENCED HYPOXIA, WHICH LED TO THE PHYSICIANS DISCOVERING THE EMBOLIZATION ON IMAGING. THE CLOSURE DEVICE HAD EMBOLIZED TO THE MIDDLE OF THE LEFT ATRIUM AND THE LEFT VENTRICLE. THE PERCUTANEOUS EXPLANTATION OF THE CLOSURE DEVICE PROCEDURE CAUSED THE PATIENT TO HAVE NEW MITRAL REGURGITATION (MR) DUE TO THE RETROGRADE APPROACH, WHICH RESULTED IN THE PATIENT REQUIRING A NON-BOSTON SCIENTIFIC MITRAL VALVE CLIP IMPLANT PROCEDURE. THE PATIENT THEN DIED AFTER THE MITRAL VALVE CLIP IMPLANT PROCEDURE OF AN UNKNOWN CAUSE OF DEATH. THE EVENTS LEADING UP TO THE DEATH ARE UNKNOWN.Event Description: IT WAS REPORTED THAT DEVICE EMBOLIZATION OCCURRED. THE DEVICE WAS EXPLANTED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. VENOUS FEMORAL ACCESS WAS OBTAINED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS USED FOR INTRAPROCEDURAL IMAGING. A DOUBLE CURVE WATCHMAN ACCESS SYSTEM (WAS) WAS INSERTED AND POSITIONED AT THE LAA. A CONTRAST INJECTION WAS PERFORMED AND MEASUREMENTS OF THE LAA TAKEN. A 35MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS ADVANCED AND THE CLOSURE DEVICE WAS DEPLOYED. RELEASE CRITERIA WAS DETERMINED TO BE MET, YET THE SHOULDER OF THE CLOSURE DEVICE WAS NOTED TO BE A THIRD TO BORDERLINE ONE HALF OF A SHOULDER AND THE PRESENCE OF A WATERMELON SEED DISTAL PUCK. AFTER THE PHYSICIAN PERFORMED A TUG TEST AND WAS CONTENT WITH THE STABILITY OF THE DEVICE, THE CLOSURE DEVICE WAS RELEASED AND IMPLANTED. THE CLOSURE DEVICE WAS OBSERVED WITH TEE AND THE PROCEDURE WAS CONCLUDED. DURING THE MIDDLE OF THE NIGHT WHEN THE PATIENT WAS STILL HOSPITALIZED, THE CLOSURE DEVICE WAS NOTED TO HAVE EMBOLIZED. THE PATIENT WAS TAKEN TO THE CARDIAC CATHETERIZATION LAB AND THE CLOSURE DEVICE WAS PERCUTANEOUSLY EXPLANTED BY RETROGRADE RETRIEVAL WITH A NON-BOSTON SCIENTIFIC (BSC) STEERABLE SHEATH AND GOOSENECK SNARE. THE PATIENT WAS THEN SENT TO THE CARDIOVASCULAR INTENSIVE CARE UNIT (CVICU) IN STABLE CONDITION TO BE MONITORED. IT WAS FURTHER REPORTED THE PATIENT HAD EXPERIENCED HYPOXIA, WHICH LED TO THE PHYSICIANS DISCOVERING THE EMBOLIZATION ON IMAGING. THE CLOSURE DEVICE HAD EMBOLIZED TO THE MIDDLE OF THE LEFT ATRIUM AND THE LEFT VENTRICLE. THE PERCUTANEOUS EXPLANTATION OF THE CLOSURE DEVICE PROCEDURE CAUSED THE PATIENT TO HAVE NEW MITRAL REGURGITATION (MR) DUE TO THE RETROGRADE APPROACH, WHICH RESULTED IN THE PATIENT REQUIRING A NON-BOSTON SCIENTIFIC MITRAL VALVE CLIP IMPLANT PROCEDURE. THE PATIENT THEN DIED AFTER THE MITRAL VALVE CLIP IMPLANT PROCEDURE OF AN UNKNOWN CAUSE OF DEATH. THE EVENTS LEADING UP TO THE DEATH ARE UNKNOWN. Manufacturer Narrative: B3: DATE OF DEATH LEFT BLANK AS IT IS UNKNOWN. H6: CODE ADDED: ANALYSIS OF DATA PROVIDED BY USER/THIRD PARTY. H6 EVALUATION CONCLUSION CODE UPDATED FROM CMC- NO PROBLEM DETECTED TO FAILURE TO FOLLOW INSTRUCTIONS. H6: CODE UPDATED FROM NO DEVICE PROBLEM FOUND TO OPERATIONAL PROBLEM IDENTIFIED. PROCEDURAL MEDIA WAS PROVIDED TO AID IN THE INVESTIGATION AND WAS REVIEWED BY A MEMBER OF THE BOSTON SCIENTIFIC MEDICAL SAFETY ORGANIZATION. 4 STILL IMAGES AND 10 SHORT VIDEO CLIPS WERE SUBMITTED FOR REVIEW. THE MEDIA SHOWED A CHICKEN WING LEFT ATRIAL APPENDAGE. THE WATCHMAN CLOSURE DEVICE MEASUREMENTS ON THE STILL IMAGES WERE SOMEWHAT UNDERESTIMATED AND IN THE 135-DEGREE VIEW, THE CLOSURE DEVICE WAS FORESHORTENED NOT ALLOWING CORRECT DEVICE MEASUREMENT OR ASSESSMENT OF DEVICE POSITION. IN THE 90-DEGREE VIEW, THE CLOSURE DEVICE POSITION WAS TOO PROXIMAL. THE MEDIA ALSO SHOWED A MID-CLOSURE DEVICE PINCHING THAT MIGHT DECREASE THE ANCHOR EFFECTIVENESS. THE TUG TEST BEFORE RELEASE SHOWED SOME WIRE BIAS THAT MIGHT DECREASE EFFECTIVENESS OF THE TEST. IN CONCLUSION, BASED ON THE MEDIA AVAILABLE, IT DOES NOT SEEM THAT RELEASE CRITERIA WERE PROPERLY ASSESSED OR MET BEFORE RELEASE AND THAT MIGHT HAVE CONTRIBUTED TO THE DEVICE EMBOLIZATION. NO MEDIA WAS SUBMITTED SHOWING THE PERCUTANEOUS REMOVAL OF THE CLOSURE DEVICE COMPLICATED BY MITRAL VALVE DAMAGE. Manufacturer Narrative: B2: UPDATED TO ADD DEATH; B3: DATE OF DEATH LEFT BLANK AS IT IS UNKNOWN; B5: INFORMATION ADDED; H1: UPDATED FROM SERIOUS INJURY TO DEATH; H6 PATIENT CODE UPDATED FROM NO CLINICAL SIGNS, SYMPTOMS OR CONDITIONS TO HYPOXIA AND MITRAL VALVE INSUFFICIENCY/REGURGITATION; H6 IMPACT CODE ADDED: DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266451&pc=QEW | 3020347218-2025-00005 | 2024-12-30 05:00 | Death | INARI MEDICAL, INC. | 2025-01-29 05:00 | QEW | CLOTTRIEVER XL CATHETER | Difficult to Remove; Use of Device Problem | Pulmonary Embolism | K242557 |  | 1 | Event Description: THE PATIENT WAS A 27-YEAR-OLD MALE WITH DEEP VEIN THROMBOSIS. THE MAJORITY OF THE PATIENT'S THROMBUS HAD BEEN REMOVED USING A TRIEVER20. IMAGING WAS PERFORMED WHICH REVEALED RESIDUAL CLOTS THE PHYSICIAN CHOSE TO ADDRESS WITH THE CLOTTRIEVER XL CATHETER (CTXL). NO INFERIOR VENA CAVA (IVC) FILTER OR PROTECTION WAS SEEN. THE CTXL WAS USED TO PERFORM A PASS IN THE COMMON FEMORAL VEIN. AFTER REMOVAL AND CLEANING THE BASKET, THE PATIENT'S O2 DROPPED TO THE LOW 90S. AFTER A COUPLE OF DEEP BREATHS, THE PATIENT'S O2 CAME BACK UP, AND THE PHYSICIAN PERFORMED ANOTHER PASS. AFTER THIS OCCURRED, THE PATIENT'S O2 DROPPED AGAIN BUT MORE SIGNIFICANTLY. A CODE WAS CALLED, AND THE PATIENT WAS FLIPPED ONTO HIS STRETCHER. MULTIPLE ROUNDS OF RESUSCITATION MEDICATIONS WERE ADMINISTERED TO THE PATIENT. RESUSCITATION EFFORTS (CPR) WERE CONTINUED FOR 45 MINUTES BEFORE THE CASE WAS CALLED AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE PROVIDED, AND THE LOT HISTORY RECORDS WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, OR DESIGN. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS THE RESULT OF INADVERTENT CLOT EMBOLISM. DISTAL EMBOLIZATION OF BLOOD CLOTS IS IDENTIFIED IN THE DEVICE LABELING AS POSSIBLE ADVERSE EVENTS/COMPLICATIONS. MANUFACTURER REFERENCE #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266497&pc=OZD | 1220648-2025-25883 | 2025-01-14 05:00 | Death | ABIOMED, INC. | 2025-01-29 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Vascular Dissection |  |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT COMPLICATIONS WERE ENCOUNTERED DURING UTILIZATION OF AN IMPELLA 5.5 PUMP FOR MECHANICAL CIRCULATORY SUPPORT. DELIVERY OF THE DEVICE PROVED COMPLICATED FOR WHICH MULTIPLE ATTEMPTS WERE REQUIRED TO ULTIMATELY PLACE THE PUMP. DURING THE COMPLICATED DELIVERY, A CONTRAST INJECTION SHOWED A POTENTIAL ISOLATED DISSECTION IN THE AXILLARY ARTERY FOR WHICH AN ALTERED APPROACH WAS REQUIRED. NO INTERVENTION WAS NECESSARY IN REGARDS TO THE DISSECTION. UPON COMPLETION OF SUPPORT, THE STAFF ATTEMPTED TO EXPLANT THE PUMP AT WHICH POINT THE PATIENT CODED. MULTIPLE CHEST COMPRESSIONS WERE PERFORMED IN AN ATTEMPT AT RESUSCITATION, BUT ULTIMATELY THE PATIENT EXPIRED DUE TO PERI-PROCEDURAL MYOCARDIAL INFARCTION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT AND VASCULAR DAMAGE - GREAT VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266725&pc=MKJ | 1220908-2025-00273 | 2024-12-27 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-29 05:00 | MKJ | X SERIES | Unable to Obtain Readings | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO MONITOR AN 81-YEAR-OLD MALE PATIENT, THE DEVICE DISPLAYED AN "ECG MONITORING FAILURE" MESSAGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED, HOWEVER THEY DO NOT BELIEVE IT WAS A RESULT OF THE REPORTED MALFUNCTION. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BEFORE THE UDI REQUIREMENTS. ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: THE DEVICE AND MULTIFUNCTION CABLE (MFC) WERE RETURNED TO ZOLL MEDICAL CORPORATION; THE CUSTOMER'S REPORT WAS DUPLICATED AND ATTRIBUTED TO THE RETURNED MFC CABLE WAS FOUND TO HAVE GASKET MATERIAL INSIDE THE MFC CABLE CONNECTOR CAUSING INTERFERENCE WITH DEVICE CONNECTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266850&pc=KDQ | 1417592-2025-00046 | 2024-10-01 04:00 | Death | MEDLINE INDUSTRIES, LP | 2025-01-29 05:00 | KDQ | OUTER CANISTER, SOFT LINER, 1500CC | Use of Device Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED BY THE USER FACILITY THAT DURING A MEDICAL EMERGENCY, SUCTION OF THE AIRWAY WAS REQUIRED IN ORDER TO INTUBATE THE PATIENT, HOWEVER THE WALLED SUCTION UNIT WAS NOT SET UP READY TO USE AND WHEN SET UP, IT DID NOT WORK CORRECTLY. Manufacturer Narrative: A NOTIFICATION WAS RECEIVED FROM MEDLINE FRANCE REGARDING AN ADVERSE EVENT ASSOCIATED WITH DYNDCLO1500. IT WAS REPORTED BY THE USER FACILITY THAT DURING A MEDICAL EMERGENCY, SUCTION OF THE AIRWAY WAS REQUIRED IN ORDER TO INTUBATE THE PATIENT, HOWEVER THE WALLED SUCTION UNIT WAS NOT SET UP READY TO USE AND WHEN SET UP, IT DID NOT WORK CORRECTLY. PER THE FACILITY IMMEDIATE ACTION WAS TAKEN TO USE THE CARDIAC ARREST TROLLEY TO MINIMIZE THE DELAY IN PROVIDING EMERGENCY SUCTION. ACCORDING TO THE FACILITY DURING AN INTERNAL INVESTIGATION AFTER THE EVENT, IT WAS FOUND THAT A 3 LITER LINER HAD BEEN PLACED INSIDE THE 1.5 LITER CANNISTER WHICH RESULTED IN THE SUCTION UNIT NOT WORKING CORRECTLY. THE CUSTOMER REPORTED THAT THE PATIENT HAS DIED, BUT THE DEATH IS NOT RELATED TO THE REPORTED EVENT. REPORTEDLY, USER ERROR WAS REPORTED TO HAVE CAUSED THE REPORTED DEVICE PROBLEM/ISSUE. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. NO SAMPLE FOR RETURN AND EVALUATION WAS INDICATED. DUE TO THE REPORTED INCIDENT AND PATIENT DEATH, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267082&pc=NIK | 3004209178-2025-01737 | 2024-09-25 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-29 05:00 | NIK | CLARIA MRI QUAD CRT-D SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED A BACTEREMIA INFECTION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER TWO MONTHS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THERE WAS ALSO A SMALL WOUND NOTED OVER THE CRT-D POCKET WHICH HAD BEEN CLOSED WITH ADHESIVE, BLOOD CULTURES SHOWED GRAM POSITIVE COCCI AND WERE POSITIVE FOR METHICILLIN-SENSITIVE STAPHYLOCOCCUS AUREUS AND THE PATIENT WAS TREATED WITH ANTIBIOTICS. AFTER PROLONGED HOSPITALIZATION THE PATIENT DEVELOPED SEPSIS AND DIED FOUR MONTHS AFTER THE SYSTEM WAS EXPLANTED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD INITIALLY PRESENTED TO THE HOSPITAL WITH EXACERBATION OF CONGESTIVE HEART FAILURE WHICH RESULTED IN THE PLACEMENT OF A LEFT VENTRICULAR ASSIST DEVICE (LVAD). FOLLOWING LVAD PLACEMENT THE PATIENT HAD A PROLONGED COMPLICATED HOSPITALIZATION WHICH INCLUDED THE PROGRESSION TO END STAGE RENAL DISEASE WITH DIALYSIS, MULTIPLE INFECTIONS, TRACHEOSTOMY, MULTIPLE CEREBROVASCULAR ACCIDENTS (CVA), ACALCULOUS CHOLECYSTITIS, A GASTROINTESTINAL BLEED, ASPIRATION AND PNEUMONIA. CULMINATING IN THE PATIENT BEING PLACED ON COMFORT CARE AND PASSING AWAY WITH CAUSE OF DEATH NOTED AS CARDIOGENIC SHOCK. Manufacturer Narrative: CORRECTION: H6 (EVAL CODE CONCLUSION: FDC/ANNEX D). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: (B)(4) LEAD IMPLANTED: (B)(6) 2014 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267108&pc=LWS | 2649622-2025-02994 | 2024-09-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-01-29 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED A BACTEREMIA INFECTION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER TWO MONTHS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THERE WAS ALSO A SMALL WOUND NOTED OVER THE CRT-D POCKET WHICH HAD BEEN CLOSED WITH ADHESIVE, BLOOD CULTURES SHOWED GRAM POSITIVE COCCI AND WERE POSITIVE FOR METHICILLIN-SENSITIVE STAPHYLOCOCCUS AUREUS AND THE PATIENT WAS TREATED WITH ANTIBIOTICS. AFTER PROLONGED HOSPITALIZATION THE PATIENT DEVELOPED SEPSIS AND DIED FOUR MONTHS AFTER THE SYSTEM WAS EXPLANTED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD INITIALLY PRESENTED TO THE HOSPITAL WITH EXACERBATION OF CONGESTIVE HEART FAILURE WHICH RESULTED IN THE PLACEMENT OF A LEFT VENTRICULAR ASSIST DEVICE (LVAD). FOLLOWING LVAD PLACEMENT THE PATIENT HAD A PROLONGED COMPLICATED HOSPITALIZATION WHICH INCLUDED THE PROGRESSION TO END STAGE RENAL DISEASE WITH DIALYSIS, MULTIPLE INFECTIONS, TRACHEOSTOMY, MULTIPLE CEREBROVASCULAR ACCIDENTS (CVA), ACALCULOUS CHOLECYSTITIS, A GASTROINTESTINAL BLEED, ASPIRATION AND PNEUMONIA. CULMINATING IN THE PATIENT BEING PLACED ON COMFORT CARE AND PASSING AWAY WITH CAUSE OF DEATH NOTED AS CARDIOGENIC SHOCK. Manufacturer Narrative: CORRECTION: H6 (EVAL CODE CONCLUSION: FDC/ANNEX D). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267125&pc=OJX | 2649622-2025-02996 | 2024-09-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-01-29 05:00 | OJX | ATTAIN PERFORMA MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P080006 |  | 1 | Event Description: IT WAS FURTHER REPORTED THERE WAS ALSO A SMALL WOUND NOTED OVER THE CRT-D POCKET WHICH HAD BEEN CLOSED WITH ADHESIVE, BLOOD CULTURES SHOWED GRAM POSITIVE COCCI AND WERE POSITIVE FOR METHICILLIN-SENSITIVE STAPHYLOCOCCUS AUREUS AND THE PATIENT WAS TREATED WITH ANTIBIOTICS. AFTER PROLONGED HOSPITALIZATION THE PATIENT DEVELOPED SEPSIS AND DIED FOUR MONTHS AFTER THE SYSTEM WAS EXPLANTED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD INITIALLY PRESENTED TO THE HOSPITAL WITH EXACERBATION OF CONGESTIVE HEART FAILURE WHICH RESULTED IN THE PLACEMENT OF A LEFT VENTRICULAR ASSIST DEVICE (LVAD). FOLLOWING LVAD PLACEMENT THE PATIENT HAD A PROLONGED COMPLICATED HOSPITALIZATION WHICH INCLUDED THE PROGRESSION TO END STAGE RENAL DISEASE WITH DIALYSIS, MULTIPLE INFECTIONS, TRACHEOSTOMY, MULTIPLE CEREBROVASCULAR ACCIDENTS (CVA), ACALCULOUS CHOLECYSTITIS, A GASTROINTESTINAL BLEED, ASPIRATION AND PNEUMONIA. CULMINATING IN THE PATIENT BEING PLACED ON COMFORT CARE AND PASSING AWAY WITH CAUSE OF DEATH NOTED AS CARDIOGENIC SHOCK.Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED A BACTEREMIA INFECTION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER TWO MONTHS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD; IMPLANTED: (B)(6) 2014. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H6 (EVAL CODE CONCLUSION: FDC/ANNEX D). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267136&pc=DTB | 2649622-2025-02997 | 2024-09-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-01-29 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P930039 |  | 1 | Event Description: IT WAS FURTHER REPORTED THERE WAS ALSO A SMALL WOUND NOTED OVER THE CRT-D POCKET WHICH HAD BEEN CLOSED WITH ADHESIVE, BLOOD CULTURES SHOWED GRAM POSITIVE COCCI AND WERE POSITIVE FOR METHICILLIN-SENSITIVE STAPHYLOCOCCUS AUREUS AND THE PATIENT WAS TREATED WITH ANTIBIOTICS. AFTER PROLONGED HOSPITALIZATION THE PATIENT DEVELOPED SEPSIS AND DIED FOUR MONTHS AFTER THE SYSTEM WAS EXPLANTED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD INITIALLY PRESENTED TO THE HOSPITAL WITH EXACERBATION OF CONGESTIVE HEART FAILURE WHICH RESULTED IN THE PLACEMENT OF A LEFT VENTRICULAR ASSIST DEVICE (LVAD). FOLLOWING LVAD PLACEMENT THE PATIENT HAD A PROLONGED COMPLICATED HOSPITALIZATION WHICH INCLUDED THE PROGRESSION TO END STAGE RENAL DISEASE WITH DIALYSIS, MULTIPLE INFECTIONS, TRACHEOSTOMY, MULTIPLE CEREBROVASCULAR ACCIDENTS (CVA), ACALCULOUS CHOLECYSTITIS, A GASTROINTESTINAL BLEED, ASPIRATION AND PNEUMONIA. CULMINATING IN THE PATIENT BEING PLACED ON COMFORT CARE AND PASSING AWAY WITH CAUSE OF DEATH NOTED AS CARDIOGENIC SHOCK.Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED A BACTEREMIA INFECTION. THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED AND REPLACED WITH A LEADLESS PACEMAKER TWO MONTHS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD IMPLANTED:(B)(6) 2014. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H6 (EVAL CODE CONCLUSION: FDC/ANNEX D). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267214&pc=DQY | 2124215-2025-04441 | 2025-01-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-29 05:00 | DQY | WATCHMAN FXD CURVE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Brain Injury |  |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A 20MM WATCHMAN FLX LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. SEDATION OF THE PATIENT PRIOR TO THE PROCEDURE AND THE MEDICATION MANAGEMENT WERE NOT SIMPLE. THE PATIENT NEEDED A HIGHER DOSE OF PROPOFOL AGENT AND ADDITIONAL MEDICATION FOR SEDATION, WHICH CAUSED A BREATHING DEPRESSION FOR A SHORT TIME AT THE BEGINNING OF THE PROCEDURE. AFTER MANAGING THE SEDATION, THE PROCEDURE WAS CONTINUED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT. BEFORE INSERTING THE WDS INTO THE WAS, THE PHYSICIAN NOTED THAT THERE WAS AN AIR EMBOLISM IN THE LEFT ATRIUM. THE PHYSICIAN USED A SYRINGE AND ASPIRATED THE AIR EMBOLISM OUT OF THE PATIENT. THE AIR WAS RESOLVED, AND THE PROCEDURE WAS CONTINUED. THE WDS WAS INSERTED INTO THE WAS AND THE CLOSURE DEVICE WAS THEN SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. POST PROCEDURE THE PATIENT DID NOT WAKE AND WAS TAKEN TO THE INTENSIVE CARE UNIT WHERE A COMPUTED TOMOGRAPHY SCAN OF THE PATIENT'S HEAD WAS PERFORMED. THE IMAGING SHOWED THAT THE PATIENT HAD HYPOXIC BRAIN DAMAGE. THE PATIENT WAS PLANNED TO BE TRANSFERRED TO ANOTHER FACILITY TO UNDERGO TREATMENT FOR THE BRAIN DAMAGE, BUT THE PATIENT DID NOT MAKE IT, AND THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267427&pc=CBK | 2518422-2025-101453 | 2024-12-27 05:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER WAS NOTIFIED OF A PATIENT DEATH WHILE THEY WERE ON VENTILATOR THERAPY USING A TRILOGY EVO VENTILATOR. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, IT WAS ALLEGED THAT THE PATIENT WITH PRIMARY DIAGNOSES OF CHRONIC RESPIRATORY FAILURE, COPD EXPIRED ON (B)(6) 2024 DURING CLINICAL AND THERAPEUTIC USE OF THE DEVICE. THE PATIENT WAS FOUND DECEASED WITH THE NON-INVASIVE INTERFACE ATTACHED AND THE MACHINE POWERED OFF. THERE HAS BEEN NO ALLEGATION OF DEVICE MALFUNCTION AT THIS TIME. IT IS UNKNOWN AT THIS TIME IF THE DEVICE ALARMED AT THE TIME OF THE EVENT. THE DEVICE SETTINGS AT THE TIME OF THE ALLEGED INCIDENT WERE AVAPS MODE VT500 8-10 RATE AUTO PS MIN 8 PS MAX 25. THE SECURE DATA CARD WILL NOT BE RETURNING WITH THE DEVICE BECAUSE THIS ITEM IS NOT APPLICABLE. THE CIRCUIT WILL NOT BE RETURNED WITH THE DEVICE. THE DEVICE HAS NOT BEEN RETRIEVED FOR FURTHER INVESTIGATION THUS FAR. Manufacturer Narrative: THE DEVICE IN QUESTION HAS BEEN RETRIEVED AND EVALUATED BY THE MANUFACTURER. DEVICE DRPT (DIAGNOSTIC REPORT) LOGS WERE INVESTIGATED FINDING NO MALFUNCTION OR FAILURE OF THE DEVICE TO PERFORM TO MANUFACTURER DECLARED SPECIFICATIONS. FURTHER ANALYSIS OF THE DEVICE INTERACTION TIME SIGNATURES FOUND THAT DURING THE DATE OF THE EVENT IN QUESTION (26-DEC-2024) THE DEVICE HAD BEEN IN STANDBY MODE AND WAS NOT PROVIDING THERAPY FROM 0730 (B)(6) 2024 UNTIL THERAPY WAS INITIATED AT 0905 (B)(6) 2024. THE DEVICE UNDERWENT FURTHER EVALUATION AND TESTING WITH THERAPY BEING PERFORMED FOR APPROXIMATELY 15.5 HOURS WITH NO FAILURE OR MALFUNCTIONS NOTED. BASED UPON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, NO CAUSAL RELATIONSHIP OF THE DEVICE TO THE PATIENT OUTCOME OF DEATH HAS BEEN NOTED. HOWEVER, CONTRIBUTION OF THE DEVICE TO THE PATIENT OUTCOME OF DEATH HAS BEEN CONFIRMED: FORESEEABLE USE ERROR - FAILURE TO TRANSITION THE DEVICE FROM STANDBY TO THERAPY STATUS-POST PATIENT CONNECTION TO THE DEVICE. DEVICE INHERENT SAFETY DESIGN PERMITS THE PATIENT TO BREATH SPONTANEOUSLY IN THE EVENT SUCH A USE ERROR OCCURS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267675&pc=DSQ | 2916596-2025-00705 | 2020-10-01 04:00 | Death | THORATEC CORPORATION | 2025-01-29 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PRE-IMPLANTATION LEFT VENTRICULAR END-DIASTOLIC DIAMETER (LVEDD) RATHER THAN BODY SURFACE AREA (BSA) SIGNIFICANTLY IMPACTS SURVIVAL AFTER DURABLE LEFT VENTRICULAR ASSIST DEVICE (D-LVAD) IMPLANTATION THROUGH THE RESEARCH ARTICLE "THE IMPACT OF BODY SURFACE AREA AND LEFT VENTRICULAR SIZE ON OUTCOMES FOLLOWING DURABLE LVAD IMPLANTATION". THIS SINGLE CENTER ANALYSIS EVALUATED 212 HEARTMATE 3 AND 101 HVAD PATIENTS WHO RECEIVED A CENTRIFUGAL D-LVAD BETWEEN AUG2015 AND OCT2020, WITH AVERAGE AGE OF 56.8 + 12.3 YEARS (72% MALE, 75% BLACK PATIENTS). PATIENTS WITH SMALLER PRE-IMPLANT LVEDD (52 OF 313 PATIENTS OR 17%) WERE FOUND TO HAVE SIGNIFICANTLY LOWER SURVIVAL AT 1-YEAR POST-IMPLANTATION COMPARED TO LARGER PRE-IMPLANT LVEDD IRRESPECTIVE OF BSA (HIGH LVEDD/HIGH BSA 85% AND HIGH LVEDD/LOW BSA 80% VS. LOW LVEDD/HIGH BSA 67% AND LOW LVEDD/LOW BSA 51%, P = 0.02).THE CORRELATION BETWEEN LVEDD AND BSA WAS MODEST (R = 0.21). Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01OCT2020 AS PATIENTS WERE IMPLANTED BETWEEN AUG2015 AND OCT2020. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. MEDSTAR HEART AND VASCULAR INSTITUTE, GEORGETOWN UNIVERSITY SCHOOL OF MEDICINE, WASHINGTON, DC; PIEDMONT HEART INSTITUTE - SAMSKY ADVANCED HEART FAILURE CENTER, ATLANTA, GA; MEDSTAR HEALTH RESEARCH INSTITUTE, WASHINGTON, DC; MEDSTAR HEART AND VASCULAR INSTITUTE, WASHINGTON, DC. ZVIMAN, A., RUBIN, G., MOLINA, E., RAO, S., CHOU, J., AFARI-ARMAH, N., KADAKKAL, A., VERA, M. P., KRISHNAN, M., GUPTA, R., RODRIGO, M., HOFMEYER, M., BALSARA, K., LAM, P., & SHEIKH, F. (2024). THE IMPACT OF BODY SURFACE AREA AND LEFT VENTRICULAR SIZE ON OUTCOMES FOLLOWING DURABLE LVAD IMPLANTATION. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S395. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1278. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: REPORT TYPE CORRECTED, SECTION B2: CORRECTED, SECTION D4: SERIAL NUMBER CORRECTED, SECTION E1: REPORTER NAME CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267744&pc=BZD | 2518422-2025-009217 | 2023-05-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21268072&pc=DSQ | 3007042319-2025-00124 | 2025-01-26 05:00 | Death | HEARTWARE, INC. | 2025-01-29 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Stopped | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF ISCHEMIC CARDIOMYOPATHY, WHO HAD A VENTRICULAR ASSIST DEVICE (VAD) IMPLANTED, EXPERIENCED A DEVICE STOP. THE PATIENT SUBSEQUENTLY PASSED AWAY. THE VAD TEAM COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR FURTHER TESTING OR ANALYSIS, FOLLOWING ESTABLISHED PROTOCOLS AND GUIDELINES. NO FURTHER INFORMATION WILL BE MADE AVAILABLE. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING A VAD STOP EVENT; HOWEVER, THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED VAD STOP EVENT COULD NOT BE CONFIRMED. APPLICABLE RISK DOCUMENTATION AND EXPERIENCE WITH EVENTS OF SIMILAR CIRCUMSTANCES WERE CONSIDERED; A POSSIBLE ROOT CAUSE OF THE REPORTED VAD STOP EVENT CAN BE ATTRIBUTED BUT NOT LIMITED TO A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER, A CONTROLLER FAILURE, OR A PUMP FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21268433&pc=DSQ | 3007042319-2025-00125 | 2024-12-13 05:00 | Death | HEARTWARE, INC. | 2025-01-29 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Obstruction of Flow; Pumping Problem | Stroke/CVA; Dysphasia; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CONGESTIVE HEART FAILURE PRESENTED TO THE HOSPITAL WITH ALTERED MENTAL STATUS, DISI NHIBITION, STEREOTYPICAL SPEECH, AND LOW FLOWS ON A LEFT VENTRICULAR ASSIST DEVICE (LVAD). DIAGNOSTIC TESTS REVEALED A LATE ACUTE INFARCT IN THE RIGHT TEMPORAL LOBE, AN OUTFLOW GRAFT THROMBUS, AND A PROXIMAL RIGHT INTERNAL CAROTID ARTERY FILLING DEFECT, WITH NO HE MORRHAGIC CONVERSION. LABORATORY RESULTS SHOWED LACTIC ACID DEHYDROGENASE AT 400, HAPTOGLOBIN AT 24, AND AN INTERNATIONAL NORMALISED RATIO OF 1.1, INDICATING SUBTHERAPEUTIC ANTICOAGULATION. THE PATIENT WAS NOT ENGAGING AND WAS REFUSING MEDICATIONS. FOLLOWING DISC USSIONS WITH THE PATIENT'S FAMILY, WHO NOTED A PROGRESSIVE DECLINE IN QUALITY OF LIFE AND INABILITY TO MEANINGFULLY ENGAGE WITH FAMILY, A DECISION WAS MADE TO TRANSITION TO PALLIATIVE CARE. LVAD SUPPORT WAS DISCONTINUED, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: ADDITIONAL PRODUCT: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ OUTFLOW GRAFT D4: MODEL #: 1125, CATALOG #: 1125, EXPIRATION DATE: 31-AUG-2022, SERIAL OR LOT#: (B)(6), D9: NO, H3: NO, H4: MFG DATE: 01-AUG-2017, H5: NO, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. ADDITIONAL PRODUCT: LOT# 16363498-0415 H3:NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE ASSOCIATED OUTFLOW GRAFT (LOT NUMBER 16363498-0415) WERE NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THE REPORTED THROMBUS AND LOW FLOW EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPERIENCED A STROKE AND WAS PLACED ON PALLIATIVE CARE. VAD SUPPORT WAS DISCONTINUED AND THE PATIENT SUBSEQUENTLY EXPIRED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INC CLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION, THROMBUS AND DEATH ARE KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTI-PLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21269772&pc=BZD | 2518422-2025-101454 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFICIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21269899&pc=LWP | 2124215-2025-04594 | 2024-12-12 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-29 05:00 | LWP | ACCOLADE MRI DR | Pacing Problem | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED BY THE FACILITY THAT FOLLOWING IMPLANT OF THE PACEMAKER THE PATIENT CODED AND EXPIRED. THERE WAS NOTED TO BE AN ABSENCE OF PACEMAKER SPIKES DURING THE CODE. THE PACEMAKER WAS PROGRAMMED TO DDDR MODE WITH A LOWER RATE OF 60 BEATS/MINUTE AND AN UPPER RATE OF 120 BEATS/MINUTE. IT WAS ALSO REPORTED THAT THE PATIENT HAD BEEN CRITICALLY ILL. NO OTHER INFORMATION (INCLUDING IDENTIFIABLE INFORMATION) WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270172&pc=BZD | 2518422-2025-101478 | 2025-01-22 05:00 | Death | RESPIRONICS, INC. | 2025-01-29 05:00 | BZD | DREAMSTATION AUTO BIPAP | Degraded | Liver Failure; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER AND KIDNEY SHUTDOWN AND DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270182&pc=OOB | 2124215-2025-04636 | 2023-10-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-29 05:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: JAPAN ALLIANCE REGISTRY IT WAS REPORTED THAT THE PATIENT DIED. A 4.00 MM X 20.00 MM AGENT DCB WAS USED IN THE SEVERELY CALCIFIED PROXIMAL PART OF THE LEFT CIRCUMFLEX ARTERY FOR TREATMENT OF SILENT ISCHEMIA. THE PATIENT WAS DISCHARGED 6 DAYS POST PROCEDURE. THE FOLLOWING MONTH, THE PATIENT WAS FOUND COLLAPSED AT HOME. EMERGENCY SERVICES WERE CONTACTED, BUT THE PATIENT HAD DIED. AN AUTOPSY WAS CONDUCTED, BUT THE OFFICIAL CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: A4 WEIGHT: 59.8.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270389&pc=NAY | 2955842-2025-01019 | 2024-12-27 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-29 05:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Bowel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT 3 DAYS AFTER AN UNEVENTFUL DA VINCI-ASSISTED PERFORATED BOWEL REPAIR, THE PATIENT REQUIRED A SECOND SURGERY AND LATER EXPIRED. THE SURGEON REPORTED THAT THE PATIENT PRESENTED TO THE HOSPITAL AND WAS DIAGNOSED WITH A BOWEL PERFORATION. THE INITIAL DA VINCI PROCEDURE WAS PERFORMED TO REPAIR THE PERFORATION; NO INFORMATION REGARDING THE ORIGIN OF THE PERFORATION OR THE PATIENT STATUS AT THE TIME WAS PROVIDED. ON POST-OPERATIVE DAY 3, THE PATIENT UNDERWENT A SECOND PROCEDURE TO RESOLVE A SECOND BOWEL PERFORATION THAT HAD NOT BEEN NOTED DURING THE ORIGINAL PROCEDURE. THE PATIENT DID NOT RECOVER FROM THE SECOND PROCEDURE AND EXPIRED ABOUT SIX DAYS LATER. THE SURGEON REPORTED THAT THERE WERE NO DA VINCI PRODUCT ISSUES. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: INFORMATION REGARDING THE PROCEDURE DATE WAS NOT PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. A SEARCH OF THE DATABASE FOR/NEAR THE ESTIMATED PROCEDURE DATE DID NOT MATCH INFORMATION REGARDING THE REPORTED ESTIMATED PROCEDURE DATE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BOWEL PERFORATION WHICH REQUIRED SURGERY. ALTHOUGH THIS WAS REPAIRED, A SECOND PERFORATION WAS FOUND ON A SUBSEQUENT OPERATION. THE DETAILS OF HOW THESE PERFORATIONS OCCURRED, IF THEY WERE BOTH PRESENT AT THE INDEX OPERATION OR THE SECOND PROCEDURE OCCURRED BEFORE, DURING OR AFTER THE FIRST PROCEDURE WAS NOT PROVIDED. NO OTHER INTRAOPERATIVE OR POST OPERATIVE DETAILS ARE AVAILABLE. THE DETAILS AND EVENTS LEADING TO THE PATIENT DEATH WAS NOT PROVIDED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTIONS B2 AND B3 WERE ESTIMATED BASED ON THE INFORMATION PROVIDED. THE ACTUAL EVENT DATE AND PATIENT DATE OF DEATH ARE UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM THE HOSPITAL RISK MANAGEMENT DEPARTMENT STATED, "[THE INFORMATION SUMMARY YOU PROVIDED] "IS NOT ACCURATE. THERE WAS NO INJURY TO A PATIENT INVOLVING DA VINCI EQUIPMENT." NO FURTHER INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270468&pc=LWS | 2124215-2025-04602 | 2025-01-02 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-29 05:00 | LWS | ENDOTAK RELIANCE G | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P910073/S043 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND LEAD SYSTEM EXPERIENCED AN INFECTION AND SEPSIS. AN EXTRACTION WAS PERFORMED AND THE CRT-D ALONG WITH THE LEFT VENTRICULAR (LV) AND RIGHT VENTRICULAR (RV) LEADS WERE EXPLANTED. HOWEVER, THE RIGHT ATRIAL (RA) LEAD BECAME FRACTURED DURING THE ATTEMPT AND WAS ABANDONED. THE PATIENT PASSED AWAY SOME TIME AFTER THE PROCEDURE, DUE TO THEIR ILLNESSES AND EFFECTS OF THE ANESTHESIA. NONE OF THE EXPLANTED PRODUCTS WERE GOING TO BE RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21249533&pc=BZD | 2518422-2025-008665 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250210&pc=BZD | 2518422-2025-008702 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250239&pc=OZP | 2032227-2025-116818 | 2024-12-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-28 05:00 | OZP | PUMP MMT-1880L MM770G BLE MG | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS UNKNOWN CAUSE OF DEATH. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1880L, MMT-332A, MMT-397A. TROUBLESHOOTING WAS NOT PERFORMED FOR DECEASED REPORTING. OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: NO CONTRIBUTING FACTORS WERE IDENTIFIED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER WOULD CONTINUE THE USE OF THE PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-1880L. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A.Event Description: THIS CASE IS NOT REPORTABLE BECAUSE AS PER THE CASE NOTES REPORTABILITY CRITERIA OF DEATH WAS NOT MENTIONED WITH RESPECT TO DEVICE ALLEGATION OR DIABETES. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT AND/OR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250273&pc=BZD | 2518422-2025-008722 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION CPAP | Degraded | Bronchitis | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CHRONIC BRONCHITIS AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250290&pc=PNJ | 9612164-2025-00435 | 2024-12-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-28 05:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN INFECTION. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED PURULENT DISCHARGE/PUS. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS EXPLANTED AND REPLACED BY A LEADLESS IPG. IT WAS FURTHER NOTED THAT THE INFECTION CONTINUED AND THE PATIENT NEVER RECOVERED AND LATER PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250615&pc=BZD | 2518422-2025-008774 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING COLON AND ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250662&pc=BZD | 2518422-2025-008669 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAM STATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, LUNG CANCER, BLADDER CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250700&pc=BZD | 2518422-2025-008775 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAM STATION CPAP | Degraded | Headache; Nausea; Vomiting; Dizziness; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250806&pc=OZP | 2032227-2025-117020 | 2024-12-06 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-28 05:00 | OZP | PUMP MMT-1884L MM780G BLE 3.0 MG/ML | Adverse Event Without Identified Device or Use Problem | Renal Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS NATURAL CAUSES. DIABETES AND END-STAGE RENAL DISEASE LISTED ON THE DEATH CERTIFICATE. OTHER RELEVANT HISTORY, THE PUMP WAS WORN AT THE TIME OF PASSING. THE CUSTOMER PASSED AWAY AT HOME. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884L, MMT-342G, MMT-442AG, MMT-442AK. TROUBLESHOOTING WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE DEVICE. PRODUCT RETURN REQUESTED FOR MMT-1884L. CUSTOMER DECLINED TO RETURN OR IS UNABLE TO RETURN. NO PRODUCT RETURN IS REQUIRED FOR MMT-342G. NO PRODUCT RETURN IS REQUIRED FOR MMT-442AG. NO PRODUCT RETURN IS REQUIRED FOR MMT-442AK. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251284&pc=BZD | 2518422-2025-101334 | 2024-08-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S SISTER HAS REQUESTED A RETURN LABEL FOR THE DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION, BUT THE EVALUATION IS NOT YET COMPLETE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251606&pc=MVK | 3008642652-2025-00980 | 2025-01-15 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-28 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 08:21:21 ON (B)(6) 2025. AT 13:31:00, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 200 BPM WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VT. MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT THE DEVICE WAS SHUT DOWN AT 13:32:05 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251714&pc=BZD | 2518422-2025-101337 | 2024-10-30 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE DEVICE IS NO LONGER AVAILABLE FOR RETURN, AS IT WAS RECYCLED FOLLOWING THE USER'S PASSING. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251821&pc=BZD | 2518422-2025-101348 | 2024-10-30 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT YET BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21252123&pc=BZD | 2518422-2025-101354 | 2024-10-31 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE USER'S WIFE HAS STATED THAT SHE IS UNSURE OF THE DEVICE'S CURRENT LOCATION AND IS UNABLE TO RETURN IT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21252223&pc=BZD | 2518422-2025-101356 | 2024-11-05 05:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21252587&pc=OZD | 1220648-2025-25849 | 2025-01-11 05:00 | Death | ABIOMED, INC. | 2025-01-28 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ESCALATED FROM IMPELLA CP TO IMPELLA 5.5 DUE TO LIMB ISCHEMIA. THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DESPITE THE IMPELLA CP BEING REMOVED, THE LEG CONTINUED TO BE ISCHEMIC DUE TO CLOTS THROUGHOUT THE LEG. THE PATIENT WAS TAKEN EMERGENTLY TO THE OPERATION ROOM FOR FEMORAL-POPLITEAL BYPASS. AFTER, THE PATIENT BECAME INCREASINGLY ACIDOTIC, WENT INTO RHABDOMYOLYSIS, AND CARE WAS WITHDRAWN, THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21252722&pc=NVY | 2649622-2025-02855 | 2024-10-11 04:00 | Death | MPRI | 2025-01-28 05:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD WAS REMOVED DUE TO INFECTION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED BY THE PATIENT'S SPOUSE, THAT THE PATIENT PASSED AWAY ONE WEEK AFTER THE LEAD WAS EXPLANTED DUE TO "SEPTIC HEART MONITOR". Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253190&pc=MHX | 1220063-2025-00008 | 2024-02-02 05:00 | Death | DRAEGER MEDICAL SYSTEMS, INC | 2025-01-28 05:00 | MHX | INFINITY DELTA | No Audible Prompt/Feedback | Asystole | K152407 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE DID NOT SOUND THE ASYSTOLE ALARM FAST ENOUGH. IT WAS FURTHER REPORTED THE PATIENT DIED BUT THE CONNECTION BETWEEN THE REPORTED EVENT AND THE PATIENT DEATH COULD NOT BE CLARIFIED YET. Manufacturer Narrative: THE LOCAL DRÄGER S&S ORGANIZATION HAS PERFORMED AN ON-SITE EVALUATION OF THE PATIENT MONITOR AND THE CENTRAL MONITORING STATION. THE PATIENT MONITOR HAD A MECHANICAL DAMAGE OF THE HOUSING AND WAS EQUIPPED WITH A NON-DRÄGER INTERNAL BATTERY WHICH WAS DEFECTIVE OR DID OTHERWISE NOT ALLOW OPERATION WITHOUT MAINS SUPPLY. BESIDES THESE NON-CONFORMITIES, MONITORING WAS FUNCTIONAL. THE LOGS FOR THE PARTICULAR MONITORING EPISODE HAD BEEN OVERWRITTEN ALREADY SINCE THE DEVICE REMAINED IN FURTHER USE. THE LOGS OF THE CENTRAL MONITORING STATION HAD BEEN ERASED AS WELL. IN GENERAL, THE REPORTED ASPECT OF DELAYED ALARMING CAN NEITHER BE CONFIRMED NOR DENIED DUE TO THE FACT THAT THE INFORMATION RELEVANT FOR AN ASSESSMENT WAS NOT AVAILABLE. THERE IS HOWEVER NO INDICATION FOR THE POTENTIAL PRESENCE OF A DEVICE MALFUNCTION. THE NON-CONFORMITIES OBSERVED WITH THE DEVICE SHOULD HAVE PREVENTED FROM PUTTING IT INTO USE, BUT THESE HAVE NO RELATION TO THE REPORTED EVENT; MONITORING WAS FUNCTIONAL WHEN THE DEVICE WAS TESTED IN FOLLOW-UP OF THE EVENT. ANOTHER ASPECT THAT SUBSTANTIATES THE POSTULATION THAT NO DEVICE-RELATED ERRORS HAVE LIKELY OCCURRED DURING THE COURSE OF EVENT IS THAT THE DEVICE WAS FURTHER USED AFTERWARDS - AN OPERATOR WOULD PROBABLY NOT DO THIS WHEN SIGNIFICANT CONCERNS AGAINST THE SAFETY AND PERFORMANCE OF THE DEVICE EXIST. A RELIABLE CONCLUSION IN REGARD TO THE EXACT ROOT CAUSE FOR THE REPORTED EVENT CANNOT BE MADE. A NUMBER OF CONTRIBUTING FACTORS MUST BE TAKEN INTO CONSIDERATION WHILE - ON THE OTHER HAND - A DIFFERENTIATION IS NOT POSSIBLE DUE TO LACK OF INFORMATION. FOR EXAMPLE, IF A SENSOR FOR A PARTICULAR MONITORING CHANNEL BECOMES DETACHED FROM THE PATIENT (ECG ELECTRODES, SPO2 SENSOR), THE DEVICE WILL POST A TECHNICAL ALARM. IT IS IN THE NATURE OF THINGS THAT LIFE-THREATENING CONDITIONS WHICH OCCUR IN SENSOR-OFF CONDITION CANNOT BE DETECTED ANYMORE VIA THIS CHANNEL; THE DEVICE WILL NOT POST PHYSIOLOGICAL ALARMS. IT WOULD ALSO BE PLAUSIBLE THAT A LOSS OF NETWORK CONNECTION HAS INTERRUPTED THE FORWARDING OF ALARMS TO THE CENTRAL STATION WHILE THE ALARM POSTED BY THE BEDSIDE MONITOR WAS RECOGNIZED WITH DELAY ONLY. CORRECTION - THE INITIAL REPORT WAS SENT UNDER MDR NUMBER 2510954-2024-00025 IN ERROR. THE CORRECT MDR FOR THIS SUBMISSION IS: 1220063-2025-00008. AFTER CONFIRMATION FROM THE BIOMED ABOUT THE EVENT THAT OCCURRED ON FEBRUARY 2ND. IT TURNS OUT THAT THE DOCTORS HAVE RULED THAT THESE WERE NOT UNDESIRABLE EVENTS, SO THERE IS NO REASON TO QUESTION THE MONITORS. THE CUSTOMER IS NO LONGER ALLEGING THAT A MALFUNCTION OCCURRED. NO FURTHER INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253406&pc=NGV | 2124215-2025-03773 | 2025-01-15 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Device Dislodged or Dislocated | Low Blood Pressure/ Hypotension; Heart Failure/Congestive Heart Failure | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A 35MM WATCHMAN FLX LAA CLOSURE DEVICE WAS IMPLANTED. APPROXIMATELY TWO (2) MONTHS POST IMPLANT, ON (B)(6) 2025, THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR THE SURGICAL REMOVAL OF THE 35MM WATCHMAN DEVICE AS IT HAD EMBOLIZED TO THE LEFT VENTRICLE. ON (B)(6) 2025, THE PATIENT DIED DUE TO COMPLICATIONS RELATED TO THE SURGICAL EXPLANTATION. IT WAS FURTHER REPORTED THAT THERE WERE NO SYMPTOMS REPORTED BY THE PATIENT OR FAMILY MEMBERS OTHER THAN SHORTNESS OF BREATH; HOWEVER, THE PATIENT WAS REPORTED TO HAVE HAD SHORTNESS OF BREATH SINCE BEFORE THE IMPLANTATION OF THE WATCHMAN DEVICE. THE MIGRATION OF THE CLOSURE DEVICE WAS DISCOVERED DURING A ROUTINE ECHOCARDIOGRAPHY. THE PHYSICIAN STATED THAT AFTER THE DEVICE REMOVAL SURGERY, WHICH ALSO INCLUDED A MITRAL VALVE REPAIR, THE PATIENT DEVELOPED HYPOTENSION AND PROGRESSED TO HEART FAILURE. THE PHYSICIAN ATTRIBUTED THE DEATH AND HEART FAILURE TO THE STRESS CAUSED BY THE OPEN-HEART SURGERY.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A 35MM WATCHMAN FLX LAA CLOSURE DEVICE WAS IMPLANTED. APPROXIMATELY TWO (2) MONTHS POST IMPLANT, ON (B)(6) 2025, THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR THE SURGICAL REMOVAL OF THE 35MM WATCHMAN DEVICE AS IT HAD EMBOLIZED TO THE LEFT VENTRICLE. ON (B)(6) 2025, THE PATIENT DIED DUE TO COMPLICATIONS RELATED TO THE SURGICAL EXPLANTATION. IT WAS FURTHER REPORTED THAT THERE WERE NO SYMPTOMS REPORTED BY THE PATIENT OR FAMILY MEMBERS OTHER THAN SHORTNESS OF BREATH; HOWEVER, THE PATIENT WAS REPORTED TO HAVE HAD SHORTNESS OF BREATH SINCE BEFORE THE IMPLANTATION OF THE WATCHMAN DEVICE. THE MIGRATION OF THE CLOSURE DEVICE WAS DISCOVERED DURING A ROUTINE ECHOCARDIOGRAPHY. THE PHYSICIAN STATED THAT AFTER THE DEVICE REMOVAL SURGERY, WHICH ALSO INCLUDED A MITRAL VALVE REPAIR, THE PATIENT DEVELOPED HYPOTENSION AND PROGRESSED TO HEART FAILURE. THE PHYSICIAN ATTRIBUTED THE DEATH AND HEART FAILURE TO THE STRESS CAUSED BY THE OPEN-HEART SURGERY. Manufacturer Narrative: MEDIA EVALUATED BY BSC MEDICAL SAFETY: PROCEDURAL ANGIOGRAPHIC MEDIA WAS PROVIDED TO ASSIST IN THE INVESTIGATION AND WAS REVIEWED BY BSC MEDICAL SAFETY. THE MEDIA REVIEW REPORT CONCLUDED: 7 PROCEDURAL FLUOROSCOPY VIDEO LOOPS SHOWING STATUS POST STERNOTOMY, INJECTION OF LEFT ATRIAL APPENDAGE (LAA) WHICH HAS AN ANTERIOR BEND AND A SECOND MAIN LOBE, DEPLOYMENT OF DEVICE WITH SOME COMPRESSION BASED ON SHAPE, TUG TEST WITHOUT SIGNIFICANT WIRE BIAS AND RELEASE OF DEVICE. POSITION AND SEAL ARE DIFFICULT TO ASSESS BASED ON FLUOROSCOPY ONLY, HENCE IT IS NOT POSSIBLE TO DETERMINE WHETHER POSITION, ANCHOR, SIZE AND SEAL (PASS) CRITERIA WERE MET. THE REMAINING LOOPS CONFIRM DEVICE EMBOLIZATION: A TRANSTHORACIC ECHOCARDIOGRAM (TTE) VIDEO LOOP SHOWS THE EMBOLIZED DEVICE IN LEFT VENTRICLE (LV) UNDER THE MITRAL VALVE, POSSIBLY ENTANGLED IN THE CHORDAE TENDINAE OF THE ANTERIOR LEAFLET AND ALSO A MODERATE MITRAL REGURGITATION. A VIDEO LOOP OF OPEN-HEART SURGERY SHOWS HOW DEVICE IS FREED FROM THE MITRAL VALVE AND REMOVED FROM THE LEFT VENTRICLE. IN CONCLUSION, THE MEDIA DOES NOT ALLOW TO DETERMINE WHAT CONTRIBUTED TO THE DEVICE EMBOLIZATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253497&pc=NRY | 2029214-2025-00226 | 2024-12-24 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-28 05:00 | NRY | SOLITAIRE FR | Break; Physical Resistance/Sticking | Thrombosis/Thrombus; No Clinical Signs, Symptoms or Conditions | K113455 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THE REACT CATHETER WAS FOUND KINKED. THE PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY TREATMENT FOR AN ACUTE ISCHEMIC STROKE. THE ACCESSED VESSEL WAS THE FEMORAL ARTERY WITH A DIAMETER OF 7.6 MM. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS SEVERE. IT WAS REPORTED THAT THE PATIENT HAD BASILAR ARTERY OCCLUSION, POOR VASCULAR DEVELOPMENT, AND ABDOMINAL BLOOD VESSELS WITH ABDOMINAL AORTIC ANEURYSM. AFTER PUNCTURE, A 0.035 LOACH GUIDEWIRE + 125CM ANGIOGRAPHY CATHETER WITH A 90CM LONG SHEATH OF LI KAI REACHED THE MIDDLE SECTION OF V1, AND WAS REPLACED WITH A TRESEND 3M GUIDEWIRE + REBAR 18 WITH REACT 71 AND DELIVERED TO THE V4 SECTION. THE MICRO-GUIDEWIRE AND MICRO-CATHETER FURTHER BROKE THROUGH THE BASILAR ARTERY APEX, AND REACT WAS DIFFICULT TO FOLLOW UP, SO THE WHOLE SYSTEM WAS WITHDRAWN AND THE LONG SHEATH DELIVERY WAS PERFORMED AGAIN. THE LONG SHEATH WAS DELIVERED TO THE UPPER END OF V1 FOR THE SECOND TIME, AND TRESEND 3M GUIDEWIRE + REBAR 18 WITH REACT 71 WAS DELIVERED AGAIN. THE GUIDEWIRE COULD NOT BREAK THROUGH THE BASILAR ARTERY OCCLUSION STAGE. AFTER REPEATED ATTEMPTS FAILED, REACT 71 WAS RAISED FOR ASPIRATION, AND A LITTLE THROMBUS WAS ASPIRATED. AFTER THE GUIDEWIRE WAS RAISED AGAIN AND COULD NOT BREAK THROUGH, THE WHOLE SYSTEM WAS WITHDRAWN, AND IT WAS FOUND THAT THE DISTAL END OF REACT WAS KINKED. THE SURGERY WAS OVER AND THE PATIENT WAS SENT TO THE ICU. IT WAS REPORTED THE CATHETER WAS KINKED IN THE DISTAL SECTION. THE REPORTED DEVICE AND ANY ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE CATHETER WAS FLUSHED AS INDICATED IN THE IFU. NO PATIENT SYMPTOMSOR FURTHER COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ANCILLARY DEVICES USED WERE: 8F FEMORAL ARTERY PUNCTURE SHEATH, LOCASTE 0.88 INCH 90CM GUIDE CATHETER, REBAR 18 0.021 IN 153 CM MICROCATHETER, BOSTON SCIENTIFIC TRESEND 0.014 INCH 300 CM GUIDEWIRE ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT IS RECOVERING WELL. THE PROCEDURE WAS COMPLETED AFTER REPLACING THE SOLITAIRE FR 6-30, IT WAS DELIVERED TO THE DISTAL END OF M1 THROUGH A MICROCATHETER AND THEN OPENED. USING THE STENT ANCHOR, THE INTERMEDIATE CATHETER WAS DELIVERED TO THE BEGINNING OF M1 AND THE SWIM TECHNIQUE WAS USED TO REMOVE THE THROMBUS SUCCESSFULLY. THERE WAS NO ADDITIONAL MEDICAL OR SURGICAL INTERVENTION. AFTER BEING TRANSFERRED TO THE ICU, THE PATIENT RECEIVED STANDARD TREATMENT AND POSTOPERATIVE RECOVERY AND IS CURRENTLY RECOVERING WELL. IT WAS NOTED THAT THE CAUSE OF THE CATHETER KINK/DIFFICULT NAVIGATION WAS SPECULATED THAT THE REASON MIGHT BE THAT WHEN THE STENT WAS DELIVERED THROUGH REBAR 18, THE GAP BETWEEN THE TAIL END OF THE MICROCATHETER AND THE STENT WAS TOO SMALL, RESULTING IN RESISTANCE DURING PUSHING AND CAUSING THE STENT TO BREAK.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT HAS DIED. THE SURGERY WAS COMPLETED AFTER THE MIDDLE CATHETER WAS REMOVED. SINCE THE GUIDEWIRE COULD NOT REACH THE DISTAL OCCLUSION AFTER REPEATED PROBING, NO SUBSEQUENT OPERATION WAS PERFORMED. THE THROMBUS WAS NOT REMOVED. NO ADDITIONAL MEDICAL INTERVENTION WERE DONE FOR THE PATIENT. AFTER THE OPERATION, THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU) ACCORDING TO NORMAL PROCEDURES AND WAS DISCHARGED THREE DAYS LATER.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS NOT THOUGHT TO BE RELATED TO MEDTRONIC DEVICE OR THERAPY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253543&pc=DQY | 2029214-2025-00227 | 2024-12-24 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-28 05:00 | DQY | REACT CATHETER | Difficult to Advance; Material Deformation; Physical Resistance/Sticking | Thrombosis/Thrombus; No Clinical Signs, Symptoms or Conditions | K182097 |  | 1 | Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT HAS DIED. THE SURGERY WAS COMPLETED AFTER THE MIDDLE CATHETER WAS REMOVED. SINCE THE GUIDEWIRE COULD NOT REACH THE DISTAL OCCLUSION AFTER REPEATED PROBING, NO SUBSEQUENT OPERATION WAS PERFORMED. THE THROMBUS WAS NOT REMOVED. NO ADDITIONAL MEDICAL INTERVENTION WERE DONE FOR THE PATIENT. AFTER THE OPERATION, THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU) ACCORDING TO NORMAL PROCEDURES AND WAS DISCHARGED THREE DAYS LATER.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS NOT THOUGHT TO BE RELATED TO MEDTRONIC DEVICE OR THERAPY.Event Description: MEDTRONIC RECEIVED A REPORT THAT THE REACT CATHETER WAS FOUND KINKED. THE PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY TREATMENT FOR AN ACUTE ISCHEMIC STROKE. THE ACCESSED VESSEL WAS THE FEMORAL ARTERY WITH A DIAMETER OF 7.6 MM. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS SEVERE. IT WAS REPORTED THAT THE PATIENT HAD BASILAR ARTERY OCCLUSION, POOR VASCULAR DEVELOPMENT, AND ABDOMINAL BLOOD VESSELS WITH ABDOMINAL AORTIC ANEURYSM. AFTER PUNCTURE, A 0.035 LOACH GUIDEWIRE + 125CM ANGIOGRAPHY CATHETER WITH A 90CM LONG SHEATH OF LI KAI REACHED THE MIDDLE SECTION OF V1, AND WAS REPLACED WITH A TRESEND 3M GUIDEWIRE + REBAR 18 WITH REACT 71 AND DELIVERED TO THE V4 SECTION. THE MICRO-GUIDEWIRE AND MICRO-CATHETER FURTHER BROKE THROUGH THE BASILAR ARTERY APEX, AND REACT WAS DIFFICULT TO FOLLOW UP, SO THE WHOLE SYSTEM WAS WITHDRAWN AND THE LONG SHEATH DELIVERY WAS PERFORMED AGAIN. THE LONG SHEATH WAS DELIVERED TO THE UPPER END OF V1 FOR THE SECOND TIME, AND TRESEND 3M GUIDEWIRE + REBAR 18 WITH REACT 71 WAS DELIVERED AGAIN. THE GUIDEWIRE COULD NOT BREAK THROUGH THE BASILAR ARTERY OCCLUSION STAGE. AFTER REPEATED ATTEMPTS FAILED, REACT 71 WAS RAISED FOR ASPIRATION, AND A LITTLE THROMBUS WAS ASPIRATED. AFTER THE GUIDEWIRE WAS RAISED AGAIN AND COULD NOT BREAK THROUGH, THE WHOLE SYSTEM WAS WITHDRAWN, AND IT WAS FOUND THAT THE DISTAL END OF REACT WAS KINKED. THE SURGERY WAS OVER AND THE PATIENT WAS SENT TO THE ICU. IT WAS REPORTED THE CATHETER WAS KINKED IN THE DISTAL SECTION. THE REPORTED DEVICE AND ANY ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE CATHETER WAS FLUSHED AS INDICATED IN THE IFU. NO PATIENT SYMPTOMSOR FURTHER COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ANCILLARY DEVICES USED WERE: 8F FEMORAL ARTERY PUNCTURE SHEATH, LOCASTE 0.88 INCH 90CM GUIDE CATHETER, REBAR 18 0.021 IN 153 CM MICROCATHETER, BOSTON SCIENTIFIC TRESEND 0.014 INCH 300 CM GUIDEWIRE ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT IS RECOVERING WELL. THE PROCEDURE WAS COMPLETED AFTER REPLACING THE SOLITAIRE FR 6-30, IT WAS DELIVERED TO THE DISTAL END OF M1 THROUGH A MICROCATHETER AND THEN OPENED. USING THE STENT ANCHOR, THE INTERMEDIATE CATHETER WAS DELIVERED TO THE BEGINNING OF M1 AND THE SWIM TECHNIQUE WAS USED TO REMOVE THE THROMBUS SUCCESSFULLY. THERE WAS NO ADDITIONAL MEDICAL OR SURGICAL INTERVENTION. AFTER BEING TRANSFERRED TO THE ICU, THE PATIENT RECEIVED STANDARD TREATMENT AND POSTOPERATIVE RECOVERY AND IS CURRENTLY RECOVERING WELL. IT WAS NOTED THAT THE CAUSE OF THE CATHETER KINK/DIFFICULT NAVIGATION WAS SPECULATED THAT THE REASON MIGHT BE THAT WHEN THE STENT WAS DELIVERED THROUGH REBAR 18, THE GAP BETWEEN THE TAIL END OF THE MICROCATHETER AND THE STENT WAS TOO SMALL, RESULTING IN RESISTANCE DURING PUSHING AND CAUSING THE STENT TO BREAK. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS: AS FOUND CONDITION: THE REACT-71 CATHETER WAS RETURNED FOR ANALYSIS WITHIN A SHIPPING BOX, A PLASTIC BIO-POUCH, AND A DISPENSER COIL. DAMAGE LOCATION DETAILS: NO DAMAGES WERE FOUND WITH THE REACT-71 CATHETER HUB. THE REACT-71 CATHETER BODY WAS FOUND STRETCHED FROM ~17.5CM TO ~15.5CM FROM THE CATHETER DISTAL TIP. THE REACT-71 CATHETER DISTAL MARKER/TIP WAS FOUND CRUSHED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253566&pc=KRA | 2029214-2025-00228 | 2024-12-24 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-28 05:00 | KRA | REBAR | Physical Resistance/Sticking | Thrombosis/Thrombus; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THE REACT CATHETER WAS FOUND KINKED. THE PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY TREATMENT FOR AN ACUTE ISCHEMIC STROKE. THE ACCESSED VESSEL WAS THE FEMORAL ARTERY WITH A DIAMETER OF 7.6 MM. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS SEVERE. IT WAS REPORTED THAT THE PATIENT HAD BASILAR ARTERY OCCLUSION, POOR VASCULAR DEVELOPMENT, AND ABDOMINAL BLOOD VESSELS WITH ABDOMINAL AORTIC ANEURYSM. AFTER PUNCTURE, A 0.035 LOACH GUIDEWIRE + 125CM ANGIOGRAPHY CATHETER WITH A 90CM LONG SHEATH OF LI KAI REACHED THE MIDDLE SECTION OF V1, AND WAS REPLACED WITH A TRESEND 3M GUIDEWIRE + REBAR 18 WITH REACT 71 AND DELIVERED TO THE V4 SECTION. THE MICRO-GUIDEWIRE AND MICRO-CATHETER FURTHER BROKE THROUGH THE BASILAR ARTERY APEX, AND REACT WAS DIFFICULT TO FOLLOW UP, SO THE WHOLE SYSTEM WAS WITHDRAWN AND THE LONG SHEATH DELIVERY WAS PERFORMED AGAIN. THE LONG SHEATH WAS DELIVERED TO THE UPPER END OF V1 FOR THE SECOND TIME, AND TRESEND 3M GUIDEWIRE + REBAR 18 WITH REACT 71 WAS DELIVERED AGAIN. THE GUIDEWIRE COULD NOT BREAK THROUGH THE BASILAR ARTERY OCCLUSION STAGE. AFTER REPEATED ATTEMPTS FAILED, REACT 71 WAS RAISED FOR ASPIRATION, AND A LITTLE THROMBUS WAS ASPIRATED. AFTER THE GUIDEWIRE WAS RAISED AGAIN AND COULD NOT BREAK THROUGH, THE WHOLE SYSTEM WAS WITHDRAWN, AND IT WAS FOUND THAT THE DISTAL END OF REACT WAS KINKED. THE SURGERY WAS OVER AND THE PATIENT WAS SENT TO THE ICU. IT WAS REPORTED THE CATHETER WAS KINKED IN THE DISTAL SECTION. THE REPORTED DEVICE AND ANY ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE CATHETER WAS FLUSHED AS INDICATED IN THE IFU. NO PATIENT SYMPTOMSOR FURTHER COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ANCILLARY DEVICES USED WERE: 8F FEMORAL ARTERY PUNCTURE SHEATH, LOCASTE 0.88 INCH 90CM GUIDE CATHETER, REBAR 18 0.021 IN 153 CM MICROCATHETER, BOSTON SCIENTIFIC TRESEND 0.014 INCH 300 CM GUIDEWIRE ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT IS RECOVERING WELL. THE PROCEDURE WAS COMPLETED AFTER REPLACING THE SOLITAIRE FR 6-30, IT WAS DELIVERED TO THE DISTAL END OF M1 THROUGH A MICROCATHETER AND THEN OPENED. USING THE STENT ANCHOR, THE INTERMEDIATE CATHETER WAS DELIVERED TO THE BEGINNING OF M1 AND THE SWIM TECHNIQUE WAS USED TO REMOVE THE THROMBUS SUCCESSFULLY. THERE WAS NO ADDITIONAL MEDICAL OR SURGICAL INTERVENTION. AFTER BEING TRANSFERRED TO THE ICU, THE PATIENT RECEIVED STANDARD TREATMENT AND POSTOPERATIVE RECOVERY AND IS CURRENTLY RECOVERING WELL. IT WAS NOTED THAT THE CAUSE OF THE CATHETER KINK/DIFFICULT NAVIGATION WAS SPECULATED THAT THE REASON MIGHT BE THAT WHEN THE STENT WAS DELIVERED THROUGH REBAR 18, THE GAP BETWEEN THE TAIL END OF THE MICROCATHETER AND THE STENT WAS TOO SMALL, RESULTING IN RESISTANCE DURING PUSHING AND CAUSING THE STENT TO BREAK.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT HAS DIED. THE SURGERY WAS COMPLETED AFTER THE MIDDLE CATHETER WAS REMOVED. SINCE THE GUIDEWIRE COULD NOT REACH THE DISTAL OCCLUSION AFTER REPEATED PROBING, NO SUBSEQUENT OPERATION WAS PERFORMED. THE THROMBUS WAS NOT REMOVED. NO ADDITIONAL MEDICAL INTERVENTION WERE DONE FOR THE PATIENT. AFTER THE OPERATION, THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU) ACCORDING TO NORMAL PROCEDURES AND WAS DISCHARGED THREE DAYS LATER.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS NOT THOUGHT TO BE RELATED TO MEDTRONIC DEVICE OR THERAPY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253764&pc=MVK | 3008642652-2025-01011 | 2024-11-29 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-28 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 07:14:50 ON (B)(6) 2024. AT 23:14:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF WITH RVR @ 150 BPM. AT 23:15:55, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. AF WITH RVR CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS AF WITH RVR @ 150 BPM. POST SHOCK RHYTHM WAS AF @ 110 BPM. AT 23:45:50, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SVT @ 150 BPM. AT 23:46:58, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. SVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SVT @ 150 BPM. POST SHOCK RHYTHM WAS SVT @ 150 BPM. AT 23:47:25, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. SVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SVT @ 150 BPM. POST SHOCK RHYTHM WAS SVT @ 160 BPM WITH MOTION ARTIFACT. AT 23:47:54, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. SVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SVT @ 160 BPM. POST SHOCK RHYTHM WAS SVT @ 150 BPM. AT 23:48:15, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. SVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SVT @ 160 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS AF @ 70 BPM WITH PVCS AND MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 23:52:08 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253839&pc=DSQ | 2916596-2025-00208 | 2024-07-12 04:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 FROM UNKNOWN CAUSES. THE DEATH WAS STATED TO NOT BE DEVICE OR THERAPY RELATED. THE CUSTOMER REPORTED THAT NO ADDITIONAL INFORMATION COULD BE PROVIDED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED OR WOULD BE PROVIDED. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED, AND IT WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D4: THE DEVICE LOT NUMBER WAS NOT PROVIDED, AND THE MANUFACTURE AND EXPIRATION DATES (H4) COULD NOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253859&pc=DSQ | 2916596-2025-00332 | 2024-12-10 05:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Anemia; Bacterial Infection; Sepsis; Post Traumatic Wound Infection; Thrombocytopenia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024 WITH A LOW HEMATOCRIT LEVEL AND ANEMIA, CAUSED BY MULTIPLE INFECTIONS, AND THE PATIENT WAS NOTED TO HAVE HAD A HISTORY OF IRON DEFICIENCY ANEMIA, AND NO BLEEDING WAS CONFIRMED. THE PATIENT'S WORSENING ANEMIA WAS ATTRIBUTED TO THE DEVELOPMENT OF SEPSIS, AND WAS TREATED WITH BLOOD PRODUCTS, AS WELL AS SEVERAL UNITS OF BLOOD THROUGHOUT ADMISSION. BLOOD CULTURES TAKEN ON (B)(6) 2024 WERE POSITIVE FOR CORYNEBACTERIUM STRIATUM (C. STRIATUM). THERE WAS ALSO A POSSIBLE DRIVELINE INFECTION ON THAT DATE, HOWEVER CULTURES WERE NEGATIVE. SACRAL WOUND CULTURES TAKEN ON (B)(6) 2024 WERE POSITIVE FOR KLEBSIELLA AND ESCHERICHIA COLI (E. COLI). LATER DRIVELINE CULTURES TAKEN WERE POSITIVE FOR C. STRIATUM AND THERE WAS SEROSANGUINOUS DRAINAGE FROM THE DRIVELINE. IT WAS SPECULATED THAT THE SACRAL WOUND COULD ALSO HAVE BEEN A SOURCE OF C. STRIATUM, HOWEVER IT DID NOT LOOK OVERLY INFECTED AT THE TIME. THE DATE OF ONSET FOR ALL INFECTIONS WAS DETERMINED TO BE (B)(6) 2024. THE INFECTIONS WERE TREATED INITIALLY WITH INTRAVENOUS (IV) VANCOMYCIN ON (B)(6) 2024, THEN THE PATIENT WAS SWITCHED TO ORAL LINEZOLID ON (B)(6) 2024 WHEN BLOOD CULTURES SENSITIVITIES CAME BACK AND THEN SWITCH AGAIN TO MINOCYCLINE ON (B)(6) 2024 DUE TO THROMBOCYTOPENIA. THE PATIENT WAS TRANSITIONED TO COMFORT CARE ON (B)(6) 2024 AND PASSED AWAY DUE TO SEPSIS ON (B)(6) 2025. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED, AND THE DEVICE WAS STATED TO HAVE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION B3: EVENT DATE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, SEPSIS, INFECTION, AND BLEEDING. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 OF THE IFU ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253896&pc=KDJ | 1416980-2025-00209 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-28 05:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS. THE CAUSE OF PERITONITIS WAS NOT REPORTED. IT WAS REPORTED THE PATIENT WAS HOSPITALIZED; HOWEVER, IT WAS NOT REPORTED IF THE PATIENT WAS HOSPITALIZED FOR THE PERITONITIS EVENT. TREATMENT FOR PERITONITIS WAS NOT REPORTED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED A LUNG AND BLOOD INFECTION. TREATMENT WAS NOT REPORTED. IT WAS REPORTED THAT THE PATIENT PASSED AWAY WHILE IN THE HOSPITAL. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF THE PERITONITIS WAS RESOLVED PRIOR TO DEATH. PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253899&pc=QZI | MW5165314 | 2024-11-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | QZI | FARAWAVE PFA ABLATION CATHETER |  | Insufficient Information |  |  | 1 | Event Description: A FARAWAVE PFA ABLATION CATHETER AND FARADRIVE SHEATH USED IN AN AFIB ABLATION RESULTED IN A PATIENT DEATH WHICH HAS NOT BEEN REPORTED BY THE MANUFACTURER (BOSTON SCIENTIFIC). REFERENCE REPORT MW5165315.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21253907&pc=DRA | MW5165315 | 2024-11-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | DRA | FARADRIVE SHEATH |  | Insufficient Information |  |  | 1 | Event Description: A FARAWAVE PFA ABLATION CATHETER AND FARADRIVE SHEATH USED IN AN AFIB ABLATION RESULTED IN A PATIENT DEATH WHICH HAS NOT BEEN REPORTED BY THE MANUFACTURER (BOSTON SCIENTIFIC). REF REPORT: MW5165314.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21254016&pc=MUQ | 1063481-2025-00002 | 2023-10-01 04:00 | Death | ARTIVION, INC. ¿ KENNESAW | 2025-01-28 05:00 | MUQ | BIOGLUE SYRINGE 5-PACK, 10 ML, JAPAN | Adverse Event Without Identified Device or Use Problem | Hypersensitivity/Allergic reaction; Insufficient Information | P01003 |  | 1 | Event Description: ACCORDING TO THE INITIAL NOTIFICATION FROM CENTURY MEDICAL, INC. POST MARKET VIGILANCE TEAM ON (B)(6) 2025, "THE PATIENT DEVELOPED MEDIAL DEGENERATION AND ACUTE AORTIC DISSECTION IN (B)(6) 2023. SUBSEQUENTLY, A BENTALL PROCEDURE WAS PERFORMED IN (B)(6) 2024. THE PATIENT ULTIMATELY DIED DUE TO RUPTURE OF THE SINUS OF VALSALVA. THE DATE OF DEATH WAS NOT REPORTED. IT WAS REPORTED THAT AN ALLERGIC REACTION WAS OBSERVED IN THE AORTIC WALL. WE ATTEMPTED TO OBTAIN INFORMATION SUCH AS THE LOT NUMBER OF THE BIOGLUE USED, THE METHOD OF USE, AND THE PROCEDURES FOLLOWED, BUT WE WERE UNABLE TO GET ANY RESPONSES BEYOND WHAT IS STATED." Manufacturer Narrative: ACCORDING TO THE INITIAL NOTIFICATION FROM CENTURY MEDICAL, INC. POST MARKET VIGILANCE TEAM ON (B)(6)2025, "THE PATIENT DEVELOPED MEDIAL DEGENERATION AND ACUTE AORTIC DISSECTION IN (B)(6) 2023. SUBSEQUENTLY, A BENTALL PROCEDURE WAS PERFORMED IN (B)(6) 2024. THE PATIENT ULTIMATELY DIED DUE TO RUPTURE OF THE SINUS OF VALSALVA. THE DATE OF DEATH WAS NOT REPORTED. IT WAS REPORTED THAT AN ALLERGIC REACTION WAS OBSERVED IN THE AORTIC WALL. WE ATTEMPTED TO OBTAIN INFORMATION SUCH AS THE LOT NUMBER OF THE BIOGLUE USED, THE METHOD OF USE, AND THE PROCEDURES FOLLOWED, BUT WE WERE UNABLE TO GET ANY RESPONSES BEYOND WHAT IS STATED IN THIS FORM." THIS INVESTIGATION IS RELEGATED TO BG3510-5-J, UNKNOWN LOT NUMBER, WITH THE ALLEGATION OF ALLERGIC REACTION IN WHICH THE PATIENT ULTIMATELY DIED. POTENTIAL LOT NUMBERS FROM SAP 6 MONTH LOOKBACK FROM THE DATE OF PROCEDURE LISTED: BG000888, BG001035, BG000972, BG000985, BG001005, BG000948, BG000915 A SAMPLE EVALUATION WAS NOT PERFORMED AS NO PRODUCT WAS RETURNED TO THE MANUFACTURER. THE MANUFACTURING RECORDS FOR BG3510-5-J, LOTS BG000888, BG001035, BG000972, BG000985, BG001005, BG000948, AND BG000915 WERE REVIEWED AND IT WAS CONFIRMED THAT ALL RECORDS WERE CONTROLLED, AVAILABLE FOR REVIEW, AND MET ALL SPECIFICATIONS PER THE DEVICE MASTER RECORD. ALL LOTS PASSED FUNCTIONAL TESTING AND MET RELEASE SPECIFICATIONS. DURING THE INVESTIGATION NO NON-CONFORMANCES OR DEVIATIONS WERE FOUND TO BE RELATED TO THE COMPLAINT. THE AVAILABLE INFORMATION WAS REVIEWED. ACCORDING TO THE INITIAL NOTIFICATION FROM CENTURY MEDICAL, INC. POST MARKET VIGILANCE TEAM ON (B)(6)2025, "THE PATIENT DEVELOPED MEDIAL DEGENERATION AND ACUTE AORTIC DISSECTION IN (B)(6)2023. SUBSEQUENTLY, A BENTALL PROCEDURE WAS PERFORMED IN (B)(6) 2024. THE PATIENT ULTIMATELY DIED DUE TO RUPTURE OF THE SINUS OF VALSALVA. THE DATE OF DEATH WAS NOT REPORTED. IT WAS REPORTED THAT AN ALLERGIC REACTION WAS OBSERVED IN THE AORTIC WALL. WE ATTEMPTED TO OBTAIN INFORMATION SUCH AS THE LOT NUMBER OF THE BIOGLUE USED, THE METHOD OF USE, AND THE PROCEDURES FOLLOWED, BUT WE WERE UNABLE TO GET ANY RESPONSES BEYOND WHAT IS STATED IN THIS FORM." THIS INVESTIGATION IS RELEGATED TO BG3510-5-J, UNKNOWN LOT NUMBER, WITH THE ALLEGATION OF ALLERGIC REACTION IN WHICH THE PATIENT ULTIMATELY DIED. POTENTIAL LOT NUMBERS FROM SAP 6 MONTH LOOKBACK: BG000888, BG001035, BG000972, BG000985, BG001005, BG000948, BG000915 BASED ON THE INFORMATION PROVIDED, INSUFFICIENT EVIDENCE IS PROVIDED TO DETERMINE WHAT CAUSED THE ALLEGED ALLERGIC REACTION. HOWEVER, A POTENTIAL REACTION TO BIOGLUE CANNOT BE EXCLUDED. THE FOLLOWING INFORMATION WAS NOT PROVIDED AND IS UNKNOWN: THE CONDITION OF THE NATIVE TISSUE, AND ANY CO-MORBIDITIES THE PATIENT MAY HAVE HAD. THERE ARE NO SURGICAL REPORTS OR MEDICAL RECORDS AVAILABLE. WE HAVE NO WAY TO EVALUATE IF BIOGLUE WAS APPLIED PER THE IFU. THE FOLLOWING IS PROVIDED IN THE APPLICABLE BIOGLUE INSTRUCTIONS FOR USE: ¿BIOGLUE IS NOT FOR PATIENTS WITH KNOWN SENSITIVITY TO MATERIALS OF BOVINE ORIGIN." ADDITIONALLY, THE FOLLOWING WARNING IS PROVIDED WITHIN THE IFU: "EXERCISE CAUTION WITH REPEAT EXPOSURE OF BIOGLUE IN THE SAME PATIENT. HYPERSENSITIVITY REACTIONS ARE POSSIBLE UPON EXPOSURE TO BIOGLUE. SENSITIZATION HAS BEEN OBSERVED IN ANIMALS." THE INSTRUCTIONS FOR USE ALSO LISTS DEATH AS AN OBSERVED UNDESIRABLE SIDE EFFECT/ADVERSE EVENT. THE ROOT CAUSE OF THE OBSERVED EVENT IS UNKNOWN. HOWEVER, A POTENTIAL REACTION TO BIOGLUE CANNOT BE EXCLUDED. A REVIEW OF IFU SHOWS ADEQUATE PRECAUTIONS AND WARNINGS PROVIDED ABOUT THE POTENTIAL FOR ALLERGIC REACTION. ADDITIONALLY, THE INSTRUCTIONS FOR USE LISTS DEATH AS AN OBSERVED UNDESIRABLE SIDE EFFECT/ADVERSE EVENT. NO FURTHER ACTION IS REQUIRED. THIS EVENT WAS REVIEWED. NO SAMPLES WERE RETURNED FOR INVESTIGATION AND A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED OVER A RANGE OF SEVEN (7) LOTS FROM A SIX (6) MONTH PERIOD SINCE THE LOT NUMBER WAS UNKNOWN. THE ROOT CAUSE OF THE OBSERVED EVENT IS UNKNOWN. HOWEVER, A POTENTIAL REACTION TO BIOGLUE CANNOT BE EXCLUDED. A REVIEW OF IFU SHOWS ADEQUATE PRECAUTIONS AND WARNINGS PROVIDED ABOUT THE POTENTIAL FOR ALLERGIC REACTION. ADDITIONALLY, THE INSTRUCTIONS FOR USE LISTS DEATH AS AN OBSERVED UNDESIRABLE SIDE EFFECT/ADVERSE EVENT. THE REPORTED EVENTS WILL CONTINUE TO BE MONITORED FOR TRENDS. RISK HAS BEEN REDUCED AS LOW AS POSSIBLE AND OVERALL RESIDUAL RISK IS ACCEPTABLE. A DEFINITIVE ROOT CAUSE FOR THE REPORTED EVENT COULD NOT BE DETERMINED. THE MANUFACTURING RECORDS WERE REVIEWED, AND IT WAS CONFIRMED THAT ALL RECORDS WERE CONTROLLED, AVAILABLE FOR REVIEW, AND MET ALL SPECIFICATIONS PER THE DEVICE MASTER RECORD. ALL LOTS PASSED FUNCTIONAL TESTING AND MET RELEASE SPECIFICATIONS. A REVIEW OF IFU SHOWS ADEQUATE PRECAUTIONS AND WARNINGS PROVIDED ABOUT THE POTENTIAL FOR ALLERGIC REACTION. ADDITIONALLY, THE INSTRUCTIONS FOR USE LISTS DEATH AS AN OBSERVED UNDESIRABLE SIDE EFFECT/ADVERSE EVENT. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION, PREVIOUSLY CRYOLIFE/JOTEC; THEREFORE, A CAPA EVALUATION IS NOT WARRANTED AT THIS TIME. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION. Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21254429&pc=OZD | 1220648-2025-25857 | 2025-01-11 05:00 | Death | ABIOMED, INC. | 2025-01-28 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Unintended Application Program Shut Down | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: EU COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA SUPPORT WHEN THE AUTOMATED IMPELLA CONTROLLER (AIC) TURNED OFF. THE AIC WAS REPLACED AND IMPELLA SUPPORT CONTINUED. THE PATIENT WAS ON IMPELLA SUPPORT FOR 9.08 HOURS BEFORE EXPIRING. Manufacturer Narrative: THE INVESTIGATION FOR THE CONSOLE (AIC) SHUTDOWN ISSUE HAS BEEN COMPLETED. CONSOLE LOGS FROM THE REPORTED DAY OF THE EVENT ARE CONSISTENT WITH THE COMPLAINT, SHOWING AN UNEXPECTED SHUTDOWN FOLLOWED BY "UNEXPECTED SHUTDOWN" ALARM AFTER THE BOOT-UP. THERE WERE NO WARNING SIGNS OR PRECURSORS PRIOR TO THE UNEXPECTED SHUTDOWN, AND NO SOFTWARE PROCESSES SHOWED SIGNS OF FAILURE. AFTER THE SHUTDOWN, THE CONSOLE¿S SOFTWARE AUTOMATICALLY RESTARTED, WITH NO INTERRUPTION IN PUMP OPERATION. THE CONSOLE WAS RETURNED FOR EVALUATION. DURING INSPECTION THE ISSUE COULD NOT BE REPRODUCED DURING TESTING WITH THE PUMP. NO UNEXPECTED SHUTDOWNS OR PERFORMANCE ISSUES WERE OBSERVED. A THOROUGH VISUAL INSPECTION OF THE INTERNAL COMPONENTS WAS CONDUCTED, BUT NO PHYSICAL PROBLEMS WERE FOUND. THERE WAS NO INTERRUPTION IN THE SUPPLY WERE DETECTED BETWEEN THE PBM AND THE 4-IN-1 ASSEMBLY. THE ROOT CAUSE OF THE INTERMITTENT, UNEXPECTED CONTROLLER SHUTDOWN COULD NOT BE DETERMINED, AS THERE WAS NO CLEAR INDICATION OF THE CAUSE, AND THE ISSUE COULD NOT BE REPRODUCED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE IMPELLA DEVICE HAS RECEIVED FROM THE CUSTOMER, UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255056&pc=MVK | 3008642652-2025-01019 | 2024-09-21 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-28 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON AN UNKNOWN DATE. THE PATIENT RECEIVED NINE INAPPROPRIATE TREATMENTS ON THEIR LAST WEAR DATE. THE DEVICE WAS STARTED UP AT 11:03:24 ON 9/20/2024. AT 05:01:46 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 05:02:22, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 05:08:08, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 05:09:37, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:10:27, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 05:12:17, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:13:29, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:14:06, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 05:15:13, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:15:56, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:16:26, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:17:17, THE PATIENT RECEIVED THE EIGHTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 05:17:49, THE PATIENT RECEIVED THE NINTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 05:17:55 ON (B)(6) 2024. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255236&pc=DQY | 2124215-2025-01346 | 2024-12-30 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | DQY | WATCHMAN TRUSTEER ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN GAINED ACCESS TO THE PATIENT AND THEN ATTEMPTED TO PERFORM THE TRANSSEPTAL PUNCTURE. THE FIRST ATTEMPT WAS UNSUCCESSFUL, BUT A SECOND ATTEMPT WAS MADE, AND THE PHYSICIAN SUCCESSFULLY COMPLETED THE TRANSSEPTAL PUNCTURE. HOWEVER, IT WAS BELIEVED THE BACK WALL OF THE HEART WAS PUNCTURED AT THIS TIME, BUT NOT SEEN. THE PHYSICIAN THEN EXCHANGED THE TRANSEPTAL SHEATH FOR THE WAS. THE WAS WENT ACROSS THE SEPTUM AND THE UNKNOWN PIGTAIL CATHETER WENT IN. AT THIS TIME PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE WAS NOTED. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS AND AUTO TRANSFUSED THE BLOOD BACK INTO THE PATIENT. IN ADDITION, THE PATIENT RECEIVED TWELVE (12) UNITS OF BLOOD, EIGHT (8) UNITS OF PLASMA AND ONE (1) UNIT OF PLATELETS. THE PATIENT WAS TAKEN TO THE OPERATING ROOM TO BE TREATED AND ON THE WAY THERE THE PATIENT WENT INTO CARDIAC ARREST. TWO HOURS POST PROCEDURE THE PATIENT WAS UNABLE TO BE RESUSCITATED, AND THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255386&pc=OZD | 1220648-2025-25858 | 2025-01-16 05:00 | Death | ABIOMED, INC. | 2025-01-28 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION WAS ESCALATED TO AN IMPELLA 5.5 FROM AN IMPELLA CP AND DURING IMPELLA 5.5 SUPPORT THE PATIENT EXPERIENCED AN ISCHEMIC CEREBRAL VASCULAR ACCIDENT (CVA) WITH A DEMISED OUTCOME. THE PATIENT HAD CHEST PAIN FOR ONE WEEK, WAS TAKEN TO THE EMERGENCY DEPARTMENT, AND DIAGNOSED WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE PATIENT WAS TRANSPORTED FOR AN EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI). ONE DRUG-ELUTING STENT (DES) WAS PLACED IN THE RIGHT CORONARY ARTERY (RCA), AND THE PATIENT DECOMPENSATED LEADING TO CARDIAC ARREST. IN BETWEEN RESUSCITATION ATTEMPTS, THE DECISION WAS MADE TO IMPLANT THE IMPELLA CP. ONCE ON MCS, THE PATIENT STABILIZED. THE PCI CONTINUED WITH A SECOND DES TO THE RCA AND TWO DES TO THE LEFT CIRCUMFLEX ARTERY. THE IMPELLA CP SITE WAS DRESSED WITH MINIMAL OOZING AND NO HEMATOMA PRESENT. THE PATIENT CONTINUED ON IMPELLA SUPPORT AND WAS PREPPED FOR AIREVAC TO AN OUTSIDE HOSPITAL. A COUPLE OF HOURS LATER, HOWEVER, A "LARGE" HEMATOMA AROUND THE IMPELLA ACCESS SITE WAS NOTED. HEPARIN RATE WAS REDUCED TO MANAGE THE CONDITION, AND THE SITE WAS MONITORED. FOLLOWING, APPROXIMATELY FIVE DAYS LATER, THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5. THE PATIENT WAS CONTINUED ON IMPELLA AND EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORTS. THREE DAYS AFTER START OF IMPELLA 5.5 SUPPORT, A COMPUTED TOMOGRAPHY SCAN WAS COMPLETED AND REVEALED AN ISCHEMIC CVA. AN ANGIOGRAM CONFIRMED A RIGHT COMMON CAROTID OCCLUSION OF UNKNOWN AGE. THE PATIENT EXPERIENCED A NEURO STATUS DECOMPENSATION AS A RESULT OF THE CVA AND CARE WAS WITHDRAWN. THE PATIENT SUBSEQUENTLY PASSED AWAY. A CAUSE FOR THE CVA IS UNKNOWN; THE NURSE REPORTED THE HOSPITAL WILL PERFORM AN AUTOPSY TO HELP DETERMINE THE CVA CAUSE. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CEREBROVASCULAR ACCIDENT (CVA) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS THIS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) : ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED FROM THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25858 AND WERE ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE PATIENT IMPLANT EXPERIENCE. REFER TO INITIAL MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 08-JAN-2025 AND IDENTIFIER ENDING IN (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255808&pc=DSQ | 2916596-2025-00262 | 2024-12-29 05:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A SUDDEN DROP IN FLOW WITH INCREASING PULSATILITY INDEX (PI). LOW FLOW ALARMS OCCURRED WITH BEEPING AND FLASHING OF THE RED HEART SYMBOL. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3, LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. A SPECIFIC CAUSE FOR THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES AND THE LOG FILES RETRIEVED FROM THE RETURNED DEVICE CONFIRMED THE REPORTED LOW FLOW EVENTS; HOWEVER, A SPECIFIC CAUSE FOR THE LOW FLOW EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE FULLY INTACT. THE MODULAR CABLE WAS RETURNED CONNECTED TO THE PUMP CABLE. APPROXIMATELY 3.0¿ OF THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF PROPERLY ENGAGED TO THE GRAFT HARDWARE. AN OUTFLOW GRAFT CLIP WAS ALSO RETURNED ATTACHED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE PUMP BODY, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED RELEVANT EVENTS ON (B)(6)2024 FROM 18:49:49 ¿ 19:35:04, PER THE TIMESTAMPS. THE MAJORITY OF THE FILE CONSISTED OF LOW FLOW EVENTS ASSOCIATED WITH SUSTAINED LOW FLOW ALARMS. A SPECIFIC CAUSE FOR THE LOW FLOW EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. TOWARD THE END OF THE FILE, THE SYSTEM WAS DISCONNECTED FROM EXTERNAL POWER FOLLOWED BY THE DISCONNECTION OF THE DRIVELINE RESULTING IN A PUMP STOP. THE CONTROLLER WAS THEN SHUT DOWN. THESE EVENTS APPEAR CONSISTENT WITH THE REPORTED EVENT OF THE PATIENT¿S EXPIRATION. NO OTHER NOTABLE EVENTS OR ALARMS REGARDING PUMP FUNCTION WERE CAPTURED, AND THE SYSTEM APPEARED TO BE OPERATING AS INTENDED PRIOR TO THE DISCONNECTION OF THE DRIVELINE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT LOG FILE RETRIEVED FROM THE RETURNED DEVICE CAPTURED LOW FLOW EVENTS ON (B)(6)2024, CONSISTENT WITH THE EVENTS OBSERVED IN THE SUBMITTED CONTROLLER EVENT LOG FILE. NO OTHER NOTABLE EVENTS WERE OBSERVED IN THE LVAD LOG FILES, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. G, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING FLOW, AND EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, PROVIDES INFORMATION ABOUT THE PUMP FLOW DISPLAY AND THE LOW FLOW HAZARD ALARM CONDITION. THIS SECTION STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LITERS PER MINUTE (LPM) AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", CAUTIONS THE USER THAT SHARP BENDS, TWISTS, OR KINKS IN THE DRIVELINE MAY MAKE IT MORE SUSCEPTIBLE TO WEAR AND FATIGUE OVER TIME. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT", CAUTIONS THE USER TO AVOID PULLING ON OR MOVING THE DRIVELINE. THIS SECTION FURTHER CAUTIONS: "DO NOT TWIST, KINK, OR SHARPLY BEND THE DRIVELINE, SYSTEM CONTROLLER POWER CABLES, THE POWER MODULE PATIENT CABLE, OR THE MOBILE POWER UNIT PATIENT CABLE, WHICH MAY CAUSE DAMAGE TO THE WIRES INSIDE, EVEN IF EXTERNAL DAMAGE IS NOT VISIBLE. DAMAGE TO THE DRIVELINE OR CABLES COULD CAUSE THE LEFT VENTRICULAR ASSIST DEVICE TO STOP. IF THE DRIVELINE OR CABLES BECOME TWISTED, KINKED, OR BENT, CAREFULLY UNRAVEL AND STRAIGHTEN.¿ SECTION 6 OF THE IFU (UNDER "CARING FOR THE DRIVELINE") INSTRUCTS THE USER TO CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS. SECTION 4 OF THE PATIENT HANDBOOK ALSO CONTAINS INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE. THIS SECTION INSTRUCTS THE USER: "CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE (CUTS, HOLES, TEARS). CALL YOUR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED)." AND "KEEP YOUR DRIVELINE CLEAN. WIPE OFF ANY DIRT OR GRIME. IF THE DRIVELINE GETS DIRTY, USE A TOWEL WITH MILD DISH SOAP AND WARM WATER TO GENTLY CLEAN IT. NEVER SUBMERGE THE DRIVELINE OR OTHER SYSTEM COMPONENTS IN WATER OR LIQUID." SECTION 7 OF THE IFU AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", PROVIDE ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE IN A SUB-SECTION ENTITLED "WHAT NOT TO DO: DRIVELINE AND CABLES". THESE SECTIONS ALSO OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THESE DOCUMENTS INSTRUCT THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. FINALLY, SECTION 8 OF THE IFU "EQUIPMENT STORAGE AND CARE" (UNDER "CLEANING THE DRIVELINE") EXPLAINS THAT AS NEEDED, THE USER CAN CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, WARM WATER AND MILD DISH SOAP SHOULD BE USED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4 PATIENT WEIGHT NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255965&pc=DSQ | 2916596-2025-00314 | 2024-11-05 05:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. ESCHERICHIA COLI WAS FOUND IN THE PERITONEAL FLUID. THE PATIENT REPORTEDLY HAD A RAPID DECOMPENSATION AND THE TEAM WITHDREW CARE. THE VENTRICULAR ASSIST DEVICE FUNCTIONED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP WAS NOT RETURNED. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255997&pc=DSQ | 2916596-2025-00629 | 2025-01-02 05:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Anemia; Arrhythmia; Bruise/Contusion; Fall; Pleural Effusion; Renal Impairment; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT FELL AT HOME AND WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025. THEY HAD EXTENSIVE BRUISING ON THEIR TRUNK, LOW HEMOGLOBIN AND POOR KIDNEY FUNCTION. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT FOR VENTRICULAR TACHYCARDIA OVERNIGHT FROM (B)(6) 2025 TO (B)(6) 2025. NURSE DOCUMENTATION SHOWED INCREASED POWER NUMBERS, BUT PERIODIC EVENT LOGS DID NOT REFLECT THIS. THE SITE ASKED FOR ASSESSMENT OF POWER TO SEE IF THE DOCUMENTATION PROVIDED WAS TRANSPOSITION OF THE PULSATILITY INDEX AND POWER VALUES. THE LOG FILE CAPTURED ROUTINE EVENTS AND THERE WERE NO NOTABLE ALARM CONDITIONS OR PARAMETER CHANGES. THE EVENT LOG DATED FROM (B)(6) 2025 THROUGH (B)(6) 2025, AND THE POWER WAS 3.2 THROUGH 3.6 MILLIWATTS.Event Description: IT WAS LATER REPORTED THAT THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT ON (B)(6) 2025 FOR ANEMIA, RENAL FAILURE AND LACTIC ACIDOSIS. ON (B)(6) 2025 THE PATIENT WAS TREATED FOR ANXIOUSNESS AND ALCOHOL WITHDRAWALS. THE PATIENT PROCEEDED TO HAVE VENTRICULAR FIBRILLATION ON (B)(6) 2025 WITH DEFIBRILLATION FROM IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR. ON (B)(6) 2025, THE PATIENT HAD A RUN OF VENTRICULAR TACHYCARDIA. THE PATIENT'S MENTAL STATUS WORSENED ON (B)(6) 2025 WITH A DECLINE TO CONFUSION AND REQUIRED INTUBATION TO HELP MAINTAIN THEIR AIRWAY ON (B)(6) 2025, AS WELL AS PRESSOR SUPPORT. LOG FILES WERE SUBMITTED FOR REVIEW ON 17JAN2025 FOR QUESTIONABLE FLUID COLLECTION AROUND THE VENTRICULAR ASSIST DEVICE (VAD). THERE APPEARED TO BE NO UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY AND THE PUMP LOG FILE APPEARED TO BE NORMAL AS WELL. COMPUTED TOMOGRAPHY (CT) OF THE PATIENT'S CHEST WAS PERFORMED ON (B)(6) 2025 AND FOUND MULTIFOCAL CAVITARY LESIONS WITH DECREASING WALL THICKNESS. THE CT IMPRESSIONS ALSO FOUND ASSOCIATED SURROUNDING PATCHY GROUNDGLASS WITH DECREASED DENSITY COMPARED TO PRIOR, PERSISTENT LEFT LOWER LOBE CONSOLIDATION AND SMALL LEFT PLEURAL EFFUSION. NEW MODERATE RIGHT LOWER LOBE SUPERIOR SEGMENT CONSOLIDATION WAS ALSO FOUND WITH TRACE RIGHT MAINSTEM BRONCHUS SECRETIONS. CT ALSO SHOWED POSTSURGICAL CHANGES FROM THE LEFT VENTRICULAR ASSIST DEVICE WITH STABLE SMALL RIGHT ANTERIOR MEDIASTINAL FLUID COLLECTION. BLOOD CULTURES UNCOVERED THE PRESENCE OF METHICILLIN RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) DESPITE ANTIBIOTIC TREATMENT AFTER MRSA WAS INDICATED IN A SPUTUM CULTURE ON (B)(6) 2025. PATIENT WAS TREATED WITH CEFTAROLINE, DAPTOMYCIN, AND LINEZOLID. NO PROCEDURE WAS CONDUCTED TO COLLECT AND CULTURE THE FLUID AROUND THE VAD. THE PATIENT WOULD CONTINUE TO BE MONITORED WITH REPEAT CT AND ANTIBIOTIC THERAPY. THE PATIENT WAS TRANSITIONED TO COMFORT CARE ON (B)(6) 2025. PRESSOR SUPPORT WAS INCREASED ON (B)(6) 2025, YET THE PATIENT ULTIMATELY WITHDREW FROM SUPPORT THE SAME DAY.Event Description: THE CAUSE OF THE FALL WAS LATER REPORTED TO NOT HAVE BEEN DEVICE OR THERAPY RELATED AS THE PATIENT INADVERTENTLY TRIPPED AND FELL. THE PATIENT'S RENAL DYSFUNCTION WAS ALSO SAID TO HAVE NOT BEEN DEVICE OR THERAPY RELATED AND HAD AN ONSET DATE OF (B)(6) 2025 FOR ACUTE KIDNEY INJURY DURING EMERGENCY DEPARTMENT ADMISSION. DIAGNOSTIC TESTING SHOWED PATIENTS CREATINE TO BE AT 7.5, METABOLIC ACIDOSIS (LACTATE 10.5), AND HYPERKALEMIA K AT 7.1. PATIENT CONTINUED TO REPORT OVERALL SORENESS ESPECIALLY AT SIGHT OF HEMATOMA OF RIGHT FLANK AND ABDOMEN. THE ARRHYTHMIA IN THIS EVENT WAS THOUGHT TO BE NOT DEVICE OR THERAPY RELATED DUE TO PREEXISTING ATRIAL FIBRILLATION AND VENTRICULAR TACHYCARDIA PRIOR TO LVAD IMPLANT. IT WAS ALSO REPORTED THAT THE PATIENT HAD AN ICD IMPLANTED PRIOR TO IMPLANT ON (B)(6) 2014 AND AN AV JUNCTION ABLATION POST IMPLANT FOR THE ARRYTHMIAS. DEATH WAS ALSO CONSIDERED TO NOT BE DEVICE OR THERAPY RELATED AND IT WAS SAID THE DEVICE WAS OPERATING AS INTENDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES COULD NOT CONFIRM THE REPORT OF ELEVATIONS IN PUMP POWER. REVIEW OF THE SUBMITTED LOG FILES SHOWED NORMAL OPERATION OF THE SYSTEM, WITH NO ATYPICAL ELEVATIONS IN POWER. NO NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, INFECTION, CARDIAC ARRHYTHMIA, RENAL DYSFUNCTION, BLEEDING, AND THROMBOEMBOLISM. SECTION 1 ALSO ADDRESSES ALL PUMP PARAMETERS INCLUDING PUMP POWER, SPEED, FLOW, AND PULSATILITY INDEX (PI). IN REFERENCE TO POWER, THIS SECTION STATES THAT PUMP POWER IS A DIRECT MEASUREMENT OF PUMP MOTOR VOLTAGE AND CURRENT. THIS SECTION FURTHER STATES THAT CHANGES IN PUMP SPEED, FLOW, OR PHYSIOLOGICAL DEMAND CAN AFFECT PUMP POWER. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION, ARRHYTHMIA, AND THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. SECTION 6 (UNDER ¿ANTICOAGULATION¿) ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL H6 HEATH IMPACT CODES: 1754 - BRUISE/CONTUSION. 2553 - CONFUSION/ DISORIENTATION. 2328 - ANXIETY. 2011 - PNEUMONIA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256436&pc=DSQ | 2916596-2025-00278 | 2025-01-02 05:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE PATIENT EXPERIENCED CARDIOGENIC SHOCK AND RECEIVED RESUSCITATION. IT WAS UNKNOWN WHETHER THE DEATH WAS CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B2 - OUTCOMES ATTRIBUTED TO ADVERSE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256484&pc=NGV | 2124215-2025-03694 | 2025-01-20 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | NGV | WATCHMAN FLX PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED. AT AN UNKNOWN TIME POST PROCEDURE, THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO NOT FEELING WELL. IT WAS DISCOVERED THAT THE CLOSURE DEVICE HAD EMBOLIZED OUT OF THE LAA. THE PATIENT WAS TRANSFERRED TO ANOTHER FACILITY TO HAVE THE CLOSURE DEVICE RETRIEVED AND REMOVED. AT AN UNKNOWN TIME, THE PATIENT DIED. Manufacturer Narrative: B2 DATE OF DEATH - AWARE DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - AWARE DATE USED AS EVENT DATE IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256495&pc=NGV | 2124215-2025-03804 | 2024-07-07 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: CHAMPION-AF STUDY. IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 24MM WATCHMAN FLX CLOSURE DEVICE WAS IMPLANTED WITH COMPLETE DEVICE SEAL NOTED. THE PATIENT WAS DISCHARGED ON RIVAROXABAN AND ASPIRIN. 929 DAYS POST INDEX PROCEDURE, IT WAS NOTED THE PATIENT DID NOT RETURN CALLS FOR THE PLANNED THREE (3) YEARS VISIT. AN OBITUARY SITE WAS CHECKED WHERE IT WAS FOUND THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS UNKNOWN. AT THE TIME OF THE EVENT, THE PATIENT WAS ON ASPIRIN. THERE ARE NO FURTHER DETAILS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256499&pc=DQY | 2124215-2025-03893 | 2025-01-23 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | DQY | WATCHMAN FXD CURVE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Bradycardia; Low Blood Pressure/ Hypotension; Hypoxia; Ventricular Fibrillation; Asystole | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS DONE IMMEDIATELY PRIOR TO PLACING THE PATIENT UNDER TWILIGHT CONSCIOUS SEDATION. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS PERFORMED WHILE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS INSERTED WITH VERSACROSS CONNECT (VCC) TO PERFORM TRANSSEPTAL PUNCTURE (TSP). TSP WENT SMOOTHLY AND NO SUSPICION OF EVENT FROM TSP. THE INTRAARTERIAL SEPTUM (IAS) WAS FLOSSED WITH THE WAS AND VCC AND THE WAS AND ICE CATHETER WAS POSITIONED ACROSS THE IAS INTO THE LEFT ATRIUM (LA). DURING A FIVE (5) MINUTE PERIOD, THE WAS AND ICE REMAINED IN THE LA UNATTENDED AS THE PHYSICIAN NEEDED TO LEAVE THE ROOM BRIEFLY, AND WHEN THE PHYSICIAN RETURNED THE WAS NOT ABLE TO BE VISUALIZED ON IMAGING. THE QRS INTERVAL BEGAN WIDENING, BRADYCARDIA WAS NOTED AND THEN ASYSTOLE. AS STAFF WAS PREPARING TO PERFORM CARDIOPULMONARY RESUSCITATION (CPR), VENTRICULAR FIBRILLATION WAS OBSERVED. THE STAFF AND PHYSICIAN CODED THE PATIENT FOR TWO (2) HOURS USING A TEMPORARY PACEMAKER, INTRA-AORTIC BALLOON PUMP (IABP), DIAGNOSTIC CARDIAC CATHETERIZATION (CC), AND DEFIBRILLATION. TEE WAS REINSERTED AND THERE WAS NO EVIDENCE OF PERICARDIAL EFFUSION. THE TEMPORARY PACEMAKER WAS FIRST UNABLE TO CAPTURE BETWEEN COMPRESSION CYCLES AND ULTIMATELY ONE LEAD WAS PLACED AND CAPTURED SUCCESSFULLY. THE DIAGNOSTIC CC REVEALED NO SIGN OF CORONARY DISEASE. ALTHOUGH IABP WAS SUCCESSFULLY PLACED, THE PATIENT MAINTAINED ONLY A WEAK BLOOD PRESSURE AND POOR OXYGEN SATURATION MAINTENANCE AT 70 OR BELOW. DUE TO THE LONG PERIOD OF TIME OF LOW OXYGEN SATURATION, THE PHYSICIAN DECIDED TO STOP RESUSCITATING THE PATIENT AND THE PATIENT DIED. IT WAS NOTED THAT THERE WAS NOT ANY HEAVY SNORING, NO GASPS OF AIR TAKEN BY THE PATIENT, AND NO CONTRAST INJECTIONS PERFORMED WITHIN THE LA. THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256574&pc=DSQ | 2916596-2025-00631 | 2024-12-18 05:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. IT WAS UNKNOWN IF AUTOPSY WAS PERFORMED. IT WAS UNKNOWN IF THE DEVICE EXPLANTED AND WOULD BE RETURNED FOR EVALUATION. Manufacturer Narrative: A4 PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION A3: PATIENT GENDER WAS NOT PROVIDED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256664&pc=DSQ | 2916596-2025-00364 | 2024-07-24 04:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2024 DUE TO RENAL FAILURE. THE PATIENT REFUSED HEMODIALYSIS AND MOVED TO HOSPICE CARE. THE PATIENT ALSO EXHIBITED RIGHT VENTRICULAR FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) TEAM WITHDREW CARE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256691&pc=DSQ | 2916596-2025-00379 | 2024-10-03 04:00 | Death | THORATEC CORPORATION | 2025-01-28 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2024 DUE TO RENAL FAILURE. THE PATIENT REFUSED HEMODIALYSIS AND MOVED TO HOSPICE CARE. THE PATIENT ALSO EXHIBITED RIGHT VENTRICULAR FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) TEAM WITHDREW CARE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B2: DATE OF DEATH CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256900&pc=LWS | 3004209178-2025-01668 | 2024-12-27 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-28 05:00 | LWS | VISIA AF MRI XT VR SURESCAN | Insufficient Information | Chest Pain; Dizziness | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT THAT THEY HAD BEEN FEELING LIGHTHEADED, CONGESTED, AND EXPERIENCED TIGHTNESS IN THEIR UPPER CHEST. THE PATIENT ALSO REPORTED HEARING TONES FROM THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). IT WAS FURTHER REPORTED VIA FOLLOW-UP WITH THE CLINIC THAT THE PATIENT PASSED AWAY ONE WEEK LATER. NO FURTHER INFORMATION WAS RECEIVED. Manufacturer Narrative: D10 CONTINUED: 977006 ,PAIN STIM IPG & 977C165, PAIN STIM LEAD, IMPLANTED (B)(6) 2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256929&pc=OZD | 1220648-2025-25860 | 2025-01-10 05:00 | Death | ABIOMED, INC. | 2025-01-28 05:00 | OZD | IMPELLA 5.5 | Failure to Advance; Patient-Device Incompatibility |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT HAD AN ATTEMPTED IMPELLA 5.5 IMPLANT DUE TO BEING ESCALATED FROM AN IMPELLA CP. THE PATIENT HAD A TORTUROUS ANATOMY AND THE CATHETER KINKED. IT WAS A DIFFICULT INSERTION, UNABLE TO IMPLANT FIRST 5.5 AND CASE ABORTED. PHYSICIAN ALTERED THE TECHNIQUE AND NEW 5.5 IMPLANTED. THERE WAS A PERIOD OF TIME (FAILED IMPLANT) WHEN PATIENT WAS NOT ON SUPPORT AND PATIENT PLACED BACK ON EXTRACORPOREAL MEMBRANE OXYGENATION AS THE PATIENT HAD SIGNIFICANT HEMODYNAMIC CHANGE. THE PATIENT WOULD EXPIRE 6 DAYS LATER. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE PRODUCT WAS RECEIVED FOR INVESTIGATION AND THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE IMPLANT OF PUMP WAS ABORTED DUE TO PATIENT'S ANATOMY. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25860 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257031&pc=LWP | 2017865-2025-11765 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-28 05:00 | LWP | ZENEX MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11766. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11767. IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257041&pc=NVN | 2017865-2025-11766 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-28 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11765. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11767. IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257051&pc=OZD | 1220648-2025-25864 | 2025-01-14 05:00 | Death | ABIOMED, INC. | 2025-01-28 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Key or Button Unresponsive/not Working | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT IN THE EU HAD A CONSOLE WITH GREY BUTTONS GETTING STUCK. THE CONSOLE WAS UTILIZED FOR THE 5.5 PUMP SUPPORT UNTIL THE PATIENT EXPIRED AFTER 24 HOURS. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - KBD/ROTARY KNOB ISSUES HAS BEEN COMPLETED. UNABLE TO CONFIRM THE REPORTED FAILURE WITH THE LOG EVIDENCE. DURING THE DEVICE ANALYSIS WAS UNABLE TO REPRODUCE THE REPORTED ISSUE. THE ROOT CAUSE OF THE GREY BUTTON GETTING STUCK WAS NOT DETERMINED DUE TO THE INABILITY TO REPRODUCE. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25864.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257052&pc=OZD | 1220648-2025-25865 | 2025-01-09 05:00 | Death | ABIOMED, INC. | 2025-01-28 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT IN JAPAN REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT HAD A CARDIAC PERFORATION REQUIRING THORACOTOMY. THE IMPELLA CAUSED/CONTRIBUTED TO THE CARDIAC PERFORATION. ADDITIONALLY, IT IS BELIEVED THAT THE PATIENT EXPIRED FROM HEMORRHAGIC SHOCK DUE TO LIVER RUPTURE. LIVER RUPTURE IS NOT IMPELLA RELATED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION ISSUE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25865 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257077&pc=NVN | 2017865-2025-11767 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-28 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11765. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11766. IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257263&pc=DTB | 2124215-2025-04215 | 2025-01-02 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | DTB | FINELINE II EZ STEROX | Fracture | Unspecified Infection; Sepsis | P960004 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND LEAD SYSTEM EXPERIENCED AN INFECTION AND SEPSIS. AN EXTRACTION WAS PERFORMED AND THE CRT-D ALONG WITH THE LEFT VENTRICULAR (LV) AND RIGHT VENTRICULAR (RV) LEADS WERE EXPLANTED. HOWEVER, THE RIGHT ATRIAL (RA) LEAD BECAME FRACTURED DURING THE ATTEMPT AND WAS ABANDONED. THE PATIENT PASSED AWAY SOME TIME AFTER THE PROCEDURE, DUE TO THEIR ILLNESSES AND EFFECTS OF THE ANESTHESIA. NONE OF THE EXPLANTED PRODUCTS WERE GOING TO BE RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257295&pc=NIK | 2124215-2025-04218 | 2025-01-02 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | NIK | MOMENTUM X4 CRT-D | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND LEAD SYSTEM EXPERIENCED AN INFECTION AND SEPSIS. AN EXTRACTION WAS PERFORMED AND THE CRT-D ALONG WITH THE LEFT VENTRICULAR (LV) AND RIGHT VENTRICULAR (RV) LEADS WERE EXPLANTED. HOWEVER, THE RIGHT ATRIAL (RA) LEAD BECAME FRACTURED DURING THE ATTEMPT AND WAS ABANDONED. THE PATIENT PASSED AWAY SOME TIME AFTER THE PROCEDURE, DUE TO THEIR ILLNESSES AND EFFECTS OF THE ANESTHESIA. NONE OF THE EXPLANTED PRODUCTS WERE GOING TO BE RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257411&pc=LWP | 2124215-2025-04217 | 2025-01-02 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-28 05:00 | LWP | ACUITY X4 SPIRAL L | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND LEAD SYSTEM EXPERIENCED AN INFECTION AND SEPSIS. AN EXTRACTION WAS PERFORMED AND THE CRT-D ALONG WITH THE LEFT VENTRICULAR (LV) AND RIGHT VENTRICULAR (RV) LEADS WERE EXPLANTED. HOWEVER, THE RIGHT ATRIAL (RA) LEAD BECAME FRACTURED DURING THE ATTEMPT AND WAS ABANDONED. THE PATIENT PASSED AWAY SOME TIME AFTER THE PROCEDURE, DUE TO THEIR ILLNESSES AND EFFECTS OF THE ANESTHESIA. NONE OF THE EXPLANTED PRODUCTS WERE GOING TO BE RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257534&pc=MNS | 2518422-2025-009002 | 2024-01-30 05:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | MNS | DREAMSTATION BIPAP AUTOSV | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K090539 |  | 1 | Event Description: A DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD DECEASED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257540&pc=BZD | 2518422-2025-101392 | 2024-10-29 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION DUE TO A USER OF DREAMSTATION AUTO CPAP. WIFE CALLED IN STATING PATIENT PASSED AWAY LAST 2022. SHE SAID THAT THE DEVICE HAS BEEN RETURNED AND NO LONGER HAS THE UNIT OR TRACKING DETAILS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21259963&pc=NPT | 2015691-2025-00535 | 2025-01-02 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-28 05:00 | NPT | US S3 UR 23MM | Device-Device Incompatibility | Organ Dehiscence | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT AN OFF-LABEL OPEN HEART TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION. THE PATIENT ALSO UNDERWENT AN OFF-LABEL OPEN HEART TMVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA IN THE NATIVE MITRAL POSITION THE SAME DAY. PER THE FCS, IT APPEARS THAT DURING THE DEPLOYMENT OF THE AORTIC VALVE, AN INTERACTION OCCURRED WITH THE MITRAL SAPIEN VALVE, RESULTING IN THE DEHISCENCE OF THE GRAFT AFFIXED TO THE LEFT ATRIUM. EFFORTS WERE MADE TO REMOVE THE AORTIC SAPIEN AND TO IMPLANT A SURGICAL AORTIC VALVE; HOWEVER, THE SURGICAL TEAM WAS UNABLE TO CONTROL THE ASSOCIATED BLEEDING AND THE PATIENT UNFORTUNATELY PASSED AWAY. AS PER FOLLOW-UP WITH THE FCS, THERE WAS NO ALLEGATION OF AN EW DEVICE MALFUNCTION OF THE AORTIC OR MITRAL VALVE THAT CAUSED OR CONTRIBUTED TO THE EVENTS AND DEATH. THERE WAS NO ALLEGATION ANY OF THE EW DEVICES WERE DEFICIENT IN SOME WAY DURING THE AORTIC OR MITRAL PROCEDURE. THERE WERE NO ISSUES WITH THE INITIAL PLACEMENT OF BOTH VALVES. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260006&pc=LYJ | 1644487-2025-00081 | 2023-06-01 04:00 | Death | CYBERONICS - HOUSTON | 2025-01-28 05:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD THE VNS FOR A YEAR AND A HALF BUT STILL PASSED AWAY DUE TO SUDEP. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260246&pc=LPB | 2029046-2025-00285 | 2024-12-31 05:00 | Death | BIOSENSE WEBSTER INC | 2025-01-28 05:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA ABLATION AND THE PATIENT EXPERIENCED CARDIAC ARREST / PULSELESS ELECTRICAL ACTIVITY (PEA) THAT REQUIRED IMPELLA PLACEMENT AND ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION). THE PATIENT EXPIRED. IT WAS REPORTED THAT THE PROCEDURE WAS AN EPICARDIAL ABLATION. DURING THE CASE, AFTER COMPLETING SOME LESIONS, AND DURING POST-ABLATION MAPPING AFTER VENTRICULAR TACHYCARDIA (VT) WAS INDUCED, THE PATIENT'S PRESSURE WAS LOW, AND THEIR EJECTION FRACTION (EF) WAS ALSO LOW. THE PATIENT WAS CARDIOVERTED TO TRY TO SHOCK THE PATIENT OUT OF IT, HOWEVER, THE PATIENT DID NOT RECOVER. THE PATIENT HAD NO EJECTION FRACTION AND WAS SHOWING SOME PULSELESS ELECTRICAL ACTIVITY. SOME OF IT CAME BACK AND THERE WAS A LITTLE BIT OF CONTRACTING. HOWEVER, AT THAT TIME, THEY KNEW THE CASE WAS OVER FOR THE MAPPING PORTION. THEY ATTEMPTED TO PUT IN AN EMERGENCY IMPELLA AND AT THAT POINT THE MAPPING WAS DONE. THE IMPELLA DEVICE DID NOT WORK, AND THE PATIENT WAS PUT ON ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION). THE PATIENT WAS TAKEN OFF ECMO AND THE PATIENT EXPIRED. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS THAT IT WAS PATIENT CONDITION RELATED. BASED ON THE INFORMATION PROVIDED AND THE LIMITED DETAILS ABOUT THE PROCEDURE, THIS EVENT IS BEING CONSERVATIVELY REPORTED AGAINST BIOSENSE WEBSTER INC. (BWI) DEVICE. IT SEEMS THAT THE MAIN CAUSE OF DEATH IS RELATED TO THE PATIENT¿S UNDERLYING CONDITION, AS THE PHYSICIAN'S OPINION ON THE CAUSE OF ADVERSE EVENT WAS PATIENT CONDITION. NO DEVICE MALFUNCTION WITH BWI PRODUCTS WERE NOTED DURING THE PROCEDURE. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260305&pc=NPT | 2015691-2025-00536 | 2025-01-02 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-28 05:00 | NPT | US S3 UR 23MM | Device-Device Incompatibility | Hemorrhage/Bleeding; Organ Dehiscence | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT AN OFF-LABEL OPEN HEART TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION. THE PATIENT ALSO UNDERWENT AN OFF-LABEL OPEN HEART TMVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA IN THE NATIVE MITRAL POSITION THE SAME DAY. PER THE FCS, IT APPEARS THAT DURING THE DEPLOYMENT OF THE AORTIC VALVE, AN INTERACTION OCCURRED WITH THE MITRAL SAPIEN VALVE, RESULTING IN THE DEHISCENCE OF THE GRAFT AFFIXED TO THE LEFT ATRIUM. EFFORTS WERE MADE TO REMOVE THE AORTIC SAPIEN AND TO IMPLANT A SURGICAL AORTIC VALVE; HOWEVER, THE SURGICAL TEAM WAS UNABLE TO CONTROL THE ASSOCIATED BLEEDING AND THE PATIENT UNFORTUNATELY PASSED AWAY. AS PER FOLLOW-UP WITH THE FCS, THERE WAS NO ALLEGATION OF AN EW DEVICE MALFUNCTION OF THE AORTIC OR MITRAL VALVE THAT CAUSED OR CONTRIBUTED TO THE EVENTS AND DEATH. THERE WAS NO ALLEGATION ANY OF THE EW DEVICES WERE DEFICIENT IN SOME WAY DURING THE AORTIC OR MITRAL PROCEDURE. THERE WERE NO ISSUES WITH THE INITIAL PLACEMENT OF BOTH VALVES. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSION AND H11 HAS BEEN UPDATED. CORRECTION TO H6 CLINICAL CODE. THE COMPLAINT FOR INTERVENTIONAL DEVICE INTERACTS WITH PRE-EXISTING IMPLANTABLE DEVICE WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF MEDICAL RECORDS AND IMAGERY. A REVIEW OF THE DHR DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF THE IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. IT SHOULD BE NOTED THAT THE THV WAS IMPLANTED USING OPEN-HEART PROCEDURE TECHNIQUE. THEREFORE, THIS WAS AN OFF-LABEL OPERATION. AS REPORTED, "A PATIENT UNDERWENT AN OFF-LABEL OPEN HEART TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION. THE PATIENT ALSO UNDERWENT AN OFF-LABEL OPEN HEART TMVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA IN THE NATIVE MITRAL POSITION THE SAME DAY. PER THE FCS, IT APPEARS THAT DURING THE DEPLOYMENT OF THE AORTIC VALVE, AN INTERACTION OCCURRED WITH THE MITRAL SAPIEN VALVE, RESULTING IN THE DEHISCENCE OF THE GRAFT AFFIXED TO THE LEFT ATRIUM. EFFORTS WERE MADE TO REMOVE THE AORTIC SAPIEN AND TO IMPLANT A SURGICAL AORTIC VALVE; HOWEVER, THE SURGICAL TEAM WAS UNABLE TO CONTROL THE ASSOCIATED BLEEDING AND THE PATIENT UNFORTUNATELY PASSED AWAY." IN THIS CASE, THE PROCEDURE WAS PERFORMED AS AN OPEN-HEART PROCEDURE RATHER THAN VIA TRANSCATHETER. THIS CAN INCREASE THE LIKELIHOOD FOR UNINTENTIONAL INTERACTION BETWEEN THE SYSTEM AND ANATOMY, WHICH CAN RESULT IN INTERACTION WITH PRE-EXISTING DEVICES AS REPORTED. AS SUCH AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (OFF-LABEL OPERATION) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260366&pc=NPT | 2025587-2025-00700 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 31 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE INFECTIONS. THE AVERAGE TIME TO EVENT WAS TWO HUNDRED AND FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 52 TO 90 AND OVER. THERE WERE 22 MALES AND 9 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 31 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260426&pc=NPT | 2025587-2025-00703 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS INFECTION. THE TIME TO EVENT WAS THREE HUNDRED FORTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS (B)(6). THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260431&pc=NPT | 2025587-2025-00701 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE INFECTIONS. THE AVERAGE TIME TO EVENT WAS SEVENTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 64 TO 90 AND OVER. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260434&pc=NPT | 2025587-2025-00705 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS RENAL. THE TIME TO EVENT WAS ONE HUNDRED THIRTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 75. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260435&pc=NPT | 2025587-2025-00704 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 8 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE RENAL. THE AVERAGE TIME TO EVENT WAS NINETY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 70 TO 90 AND OVER. THERE WERE 6 MALES AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 8 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260437&pc=NPT | 2025587-2025-00706 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS RENAL. THE TIME TO EVENT WAS SEVENTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 85. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260449&pc=NPT | 2025587-2025-00710 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 41 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED SIXTY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 60 TO 90 AND OVER. THERE WERE 19 MALES AND 22 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 41 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260452&pc=NPT | 2025587-2025-00711 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 4 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS TWENTY-FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 76 TO 90 AND OVER. THERE WERE 2 MALES AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260455&pc=NPT | 2025587-2025-00713 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS NINETY-SIX DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 85 TO 90 AND OVER. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260456&pc=NPT | 2025587-2025-00712 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS PULMONARY. THE TIME TO EVENT WAS THREE HUNDRED FIFTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 70. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260458&pc=NPT | 2025587-2025-00708 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS NEUROLOGIC. THE TIME TO EVENT WAS THREE HUNDRED FIFTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 80. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260459&pc=NPT | 2025587-2025-00707 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 11 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE NEUROLOGIC. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FIFTY-SEVEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 79 TO 90 AND OVER. THERE WERE 6 MALES AND 5 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 11 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260461&pc=NPT | 2025587-2025-00709 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE NEUROLOGIC. THE AVERAGE TIME TO EVENT WAS THREE HUNDRED THIRTY-TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 86 TO 88. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260481&pc=NPT | 2025587-2025-00719 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 4 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE LISTED AS UNKNOWN. THE AVERAGE TIME TO EVENT WAS TWO HUNDRED EIGHTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 75 TO 90 AND OVER. THERE WERE 3 MALES AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260482&pc=NPT | 2025587-2025-00718 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO PLUS VALVE | Patient Device Interaction Problem | Cardiac Arrest | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS DEATH DUE TO OTHER CARDIAC REASON. THE TIME TO EVENT WAS EIGHTY DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 61. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260491&pc=NPT | 2025587-2025-00715 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 86 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED TWENTY-EIGHT DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 24 TO (B)(6) AND OVER. THERE WERE 36 MALES AND 50 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 86 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260494&pc=NPT | 2025587-2025-00717 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS DEATH DUE TO OTHER CARDIAC REASON. THE TIME TO EVENT WAS FIFTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS (B)(6) OR OVER. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260496&pc=NPT | 2025587-2025-00716 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 1 | Event Description: THE TYPES OF DEATH BEING REPORTED IS SUDDEN CARDIAC DEATH. THE TIME TO EVENT WAS NINETEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 74. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260498&pc=NPT | 2025587-2025-00720 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 9 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE TRAUMA. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FORTY-FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 69 TO 90 AND OVER. THERE WERE 7 MALES AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 9 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260499&pc=NPT | 2025587-2025-00721 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Myocardial Infarction | P130021 | E2014038 | 6 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE MYOCARDIAL INFARCTIONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED TWENTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 47 TO (B)(6). THERE WERE 4 MALES AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 6 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260517&pc=NPT | 2025587-2025-00722 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 4 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE CARDIOVASCULAR PROCEDURE RELATED. THE AVERAGE TIME TO EVENT WAS EIGHTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 71 TO 83. THERE WERE 4 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260522&pc=NPT | 2025587-2025-00723 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS CARDIOVASCULAR PROCEDURE RELATED. THE TIME TO EVENT WAS ONE DAY FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 77. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260655&pc=NPT | 2025587-2025-00724 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Hemorrhage/Bleeding | P130021 | E2014038 | 12 | Event Description: THE TYPES OF DEATHS BEING REPORTED WERE HEMORRHAGIC AND CARDIOVASCULAR HEMORRHAGES. THE AVERAGE TIME TO EVENT WAS SEVENTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 63 TO 90 AND OVER. THERE WERE 6 MALES AND 6 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 12 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260737&pc=NPT | 2025587-2025-00726 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS HEART FAILURE. THE TIME TO EVENT WAS TWENTY-SIX DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 79. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260747&pc=NPT | 2025587-2025-00727 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS HEART FAILURE. THE TIME TO EVENT WAS TWO HUNDRED SIXTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 79. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260748&pc=NPT | 2025587-2025-00725 | 2024-11-06 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-28 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 34 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE HEART FAILURE. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED SIXTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 63 TO (B)(6) AND OVER. THERE WERE 18 MALES AND 16 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q4 2024 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 34 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21260993&pc=DQY | 3007284313-2025-03725 | 2025-01-10 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-28 05:00 | DQY | GORE® TRI-LOBE BALLOON CATHETER | Patient Device Interaction Problem | Vascular Dissection | K081799 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT OF A THORACIC AORTIC ARCH ANEURYSM USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM AND A NON-GORE DEVICE (ZENITH ALPHA EXTENSION). WHEN A 22FR GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS ADVANCED, REPORTEDLY THERE WAS RESISTANCE. TGMR313120J WAS IMPLANTED FIRST, THEN TGM343420J WAS IMPLANTED IN ZONE 2. HOWEVER, TGM343420J MOVED DISTALLY AND IMPLANTED IN ZONE 3. THEN, THE NON-GORE DEVICE WAS IMPLANTED PROXIMALLY. WHEN A TOUCH-UP USING A GORE® TRI-LOBE BALLOON CATHETER WAS PERFORMED TO THE NON-GORE DEVICE, THE BLOOD PRESSURE DROPPED, THE PATIENT LOST CONSCIOUSNESS AND THE PATIENT WAS INTO CARDIAC ARREST. A DRAINAGE AND A CARDIAC MASSAGE WERE ATTEMPTED, BUT THE PATIENT DIED. THE PHYSICIAN STATED THAT A RETROGRADE THORACIC AORTIC DISSECTION OCCURRED DUE TO THE TOUCH-UP TO THE NON-GORE DEVICE USING THE GORE® TRI-LOBE BALLOON CATHETER AND IT CAUSED A CARDIAC TAMPONADE. REPORTEDLY, THE GORE® TRI-LOBE BALLOON CATHETER WAS INFLATED WITH THE BALLOON WAS LOCATED INSIDE THE STENT GRAFT. Manufacturer Narrative: H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TRI-LOBE BALLOON CATHETER INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: TRAUMA TO THE VESSEL WALL, INCLUDING SPASM, DISSECTION, PERFORATION OR RUPTURE, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21261234&pc=BZD | 2518422-2025-009038 | 2023-10-13 04:00 | Death | RESPIRONICS, INC. | 2025-01-28 05:00 | BZD | PHILIPS CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 05/26/2025 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21238630&pc=BZD | 2518422-2025-008187 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Myocardial Infarction; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY,DEATH,KIDNEY DAMAGE, HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21238709&pc=BZD | 2518422-2025-008190 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Nausea; Vomiting; Dizziness; Cardiovascular Insufficiency; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA /VOMITING, ASTHMA (NEW OR WORSENING), REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21238885&pc=BZD | 2518422-2025-008231 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER, DEATH PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239027&pc=MHX | 1218950-2025-000056 | 2024-12-26 05:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-01-27 05:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm; No Audible Alarm | Cardiac Arrest | K183387 |  | 1 | Event Description: IT WAS REPORTED THE CENTRAL STATION DID NOT GENERATE AN ALARM WHEN THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY. THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHETHER ALARMS WERE TRIGGERED AT THE TIME OF THE EVENT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: PATIENT INVOLVEMENT WAS REPORTED, THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. THE DEVICE, INCLUDING THE ASSOCIATED BEDSIDE DEVICE(S), WERE TESTED AND DETERMINED TO BE FUNCTIONING AS DESIGNED. IT WAS ALSO INDICATED DURING TESTING THAT THE PIC IX SPEAKER WAS PUSHED TO THE BACK OF THE SHELF. THE PROVIDED LOGS WERE REVIEWED BY A PHILIPS PRODUCT SUPPORT ENGINEER AND CLINICAL SPECIALIST. IT WAS DETERMINED THAT THE DEVICE DID ALARM FOR ASYSTOLE DURING THE INDICATED TIMEFRAMES. THE DEVICE WAS TESTED AND DETERMINED TO BE FUNCTIONING AS DESIGNED. REVIEW OF THE LOGS ALSO INDICATED THE DEVICE WAS ALARMING AS EXPECTED. IT WAS CONFIRMED THAT THE DEVICE WAS TESTED AND IS STILL IN CLINICAL USE. REPORTING INSTITUTION PHONE#: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239086&pc=BZD | 2518422-2025-008285 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Inflammation; Nausea; Vomiting; Dizziness; Chronic Obstructive Pulmonary Disease (COPD); Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, AND/OR HEADACHE, NAUSEA / VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: BOX H WAS UPDATED FROM SERIOUS INJURY TO DEATH. Manufacturer Narrative: BOX B AND BOX H WAS CORRECTED AND UPDATED. Manufacturer Narrative: THE LEGACY COMPLAINT ID WAS UPDATED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239269&pc=BZD | 2518422-2025-008359 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239325&pc=BZD | 2518422-2025-008354 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239356&pc=DTZ | 8010762-2025-0000043 | 2025-01-26 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-01-27 05:00 | DTZ | TUBING SET | Gas Output Problem | Low Oxygen Saturation | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT A HLS SET WAS REPLACED AND AFTER THE REPLACEMENT (A FEW HOURS LATER) THE CUSTOMER NOTICED A DECREASE IN OXYGENATION VALUES. THE PATIENT OXYGENATION WAS IN THE LOW 80S. THE CUSTOMER CHECKED THE GAS SOURCE TO ENSURE PROPER CONNECTION AND THEY INCREASED THE SWEEP TEMPORARILY TO REMOVE ANY POTENTIAL CONDENSATION. THE POST OXYGENATION (PAO2) DECREASED FURTHER AND THE CUSTOMER EXCHANGED THE HLS SET AGAIN. AS THE PATIENT SUFFERED A LOW OXYGENATION, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4). Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER. Manufacturer Narrative: THE AFFECTED PRODUCT WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-04-28 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. ALL PERFORMANCE TESTS RELATED TO THE REPORTED FAILURE WERE FUNCTIONING AS PER SPECIFICATION. FURTHER, DURING VISUAL INSPECTION THIRD PARTY ARTICLES, CONNECTED TO THE MODULE, WERE FOUND. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT A HLS SET WAS REPLACED AND AFTER THE REPLACEMENT (A FEW HOURS LATER) THE CUSTOMER NOTICED A DECREASE IN OXYGENATION VALUES. THE PATIENT OXYGENATION WAS IN THE LOW 80`S. THE CUSTOMER CHECKED THE GAS SOURCE TO ENSURE PROPER CONNECTION AND THEY INCREASED THE SWEEP TEMPORARILY TO REMOVE ANY POTENTIAL CONDENSATION. THE POST OXYGENATION (PAO2) DECREASED FURTHER AND THE CUSTOMER EXCHANGED THE HLS SET AGAIN. NEW INFORMATION RECEIVED ON 2025-01-27 THAT THE PATIENT IS 33 YEARS OLD, WEIGHT 149.9KG AND A MALE. NO VISIBLE CLOTS WERE DETECTED. AFTER THE SECOND EXCHANGE TO THE THIRD SET THE PATIENT IMPROVED SLIGHTLY BUT THAT DID NOT LAST LONG. THE PATIENT HAD HIGH LEVELS OF TRIGLYCERIDES AND CONTINUED TO HAVE LOW OXYGENATION AND PASSED AWAY LATER IN THE AFTERNOON. ON 2025-02-03 THE INFORMATION WAS RECEIVED THAT THE FIRST AND THIRD USED SET WERE DISCARDED AND THE LOT NUMBERS ARE UNKNOWN. THE SECOND USED SET WAS AVAILABLE FOR INVESTIGATION. AS THE PATIENT SUFFERED A LOW OXYGENATION AND PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT (SECOND SET) WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-04-28 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. ALL PERFORMANCE TESTS RELATED TO THE REPORTED FAILURE WERE FUNCTIONING AS PER SPECIFICATION. FURTHER, DURING VISUAL INSPECTION THIRD PARTY ARTICLES, CONNECTED TO THE MODULE, WERE FOUND. FURTHER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-05-22 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT IN SCOPE INVOLVES A 33-YEAR-OLD MALE PATIENT ON VV-ECLS USING A CARDIOHELP SYSTEM WITH AN HLS SET EXPERIENCED A DECLINE IN OXYGENATION FOLLOWING A CIRCUIT EXCHANGE ON DAY SIX OF SUPPORT. WITHIN TEN HOURS OF INITIATING THE NEW CIRCUIT, POST-OXYGENATOR PAO2 DROPPED FROM 227 TO 210 MMHG (FIO2 1.0, SWEEP 6 L/MIN, BLOOD FLOW 5.8 - 6.28 L/MIN), AND THE PATIENT'S SPO2 FELL INTO THE LOW 80S. A SUBSEQUENT CIRCUIT REPLACEMENT DID NOT RESTORE ADEQUATE OXYGENATION, AND THE PATIENT SUBSEQUENTLY EXPIRED. INVESTIGATION OF THE RETURNED HLS SET - THE SECOND OF THREE CIRCUITS - REVEALED DAMAGE OF A GAS FILTER PART AND A CRACKED THIRD-PARTY STOPCOCK. THESE WERE ATTRIBUTED TO POST-USE HANDLING, WITH NO CLINICAL LEAKAGE REPORTED. INTERNAL INSPECTION SHOWED NO CLOTS OR MECHANICAL OBSTRUCTIONS AS DESCRIBED IN THE INVESTIGATION REPORT, AND THE DEVICE PASSED ALL LEAK AND FUNCTION TESTS AFTER REMOVING THE DAMAGED EXTERNAL STOPCOCK PART. ALTHOUGH, THE OBSERVED CRACKING OF THE THIRD-PARTY STOPCOCK MAY BE ASSOCIATED WITH ENVIRONMENTAL STRESS CRACKING OF POLYCARBONATE, AS DESCRIBED IN THE HANTEL TECHNOLOGIES WHITE PAPER ON MATERIAL SELECTION. THE REPORT HIGHLIGHTS THAT PC COMPONENTS USED IN MEDICAL DEVICES CAN FAIL WHEN EXPOSED TO LIPOPHILIC AGENTS, WHICH ARE COMMONLY USED IN CRITICAL CARE AND INTERVENTIONAL PROCEDURES. SUCH EXPOSURE - ESPECIALLY UNDER HOOP STRESS FROM HAND-TIGHTENING CONNECTIONS - CAN RESULT IN CRAZING AND CRACKING, EVEN IN THE ABSENCE OF VISIBLE DEFECTS UNDER NORMAL CONDITIONS. SPECIFICALLY, THE PAPER DEMONSTRATED THAT PC LUERS AND STOPCOCKS SUBJECTED TO LIPID-BASED AGENTS SUCH AS LIPIODOL OR SIMULATED CONTRAST MEDIA (1PERCENTAGE POPPY SEED OIL IN ETHYL STEARATE) SHOWED MICROCRACKS AND, EVENTUALLY, FULL STRUCTURAL FAILURE AFTER MECHANICAL STRESS. THE COMBINATION OF CHEMICAL EXPOSURE AND MECHANICAL FORCE AMPLIFIED THE HOOP STRESS BEYOND THE MATERIAL'S RESISTANCE THRESHOLD, LEADING TO ESC, A POTENTIAL MECHANISM CONSISTENT WITH THE CRACKED STOPCOCK OBSERVED IN THIS CASE. GAS PERMEABILITY WAS CONFIRMED DURING THE PRODUCT INVESTIGATION, INDICATING INTACT CO2 REMOVAL AND O2 DELIVERY FUNCTIONS. CLINICALLY, NO THROMBI WERE NOTED, ROUTINE FLUSHING WAS PERFORMED, AND TRANSMEMBRANE PRESSURE REMAINED WITHIN NORMAL LIMITS. THE DEVICE HAD BEEN PRIMED NINE DAYS PRIOR TO USE. HOWEVER, ACCORDING TO THE INSTRUCTIONS FOR USE, THE SET SHOULD ONLY BE PRIMED ¿[¿] SHORTLY BEFORE USE [¿]¿. NONETHELESS, THIS DEVIATION IS NOT EXPECTED TO HAVE IMPACTED THE INCIDENT DESCRIBED. ALTHOUGH THE TECHNICAL INVESTIGATION CONCLUDED THAT THERE WERE NO INDICATIONS OF POTENTIAL IMPAIRMENT TO OXYGENATION PERFORMANCE, SEVERAL CLINICAL AND CONTEXTUAL FACTORS MERIT CONSIDERATION AS POSSIBLE CONTRIBUTORS TO THE OBSERVED OXYGENATION DECLINE: 1) THE INVESTIGATION IDENTIFIED A CRACK IN THE THIRD-PARTY THREE-WAY STOPCOCK CONNECTED TO THE SAMPLE LINE AT THE QUICK-VENT LUER. WHILE DEEMED LIKELY TO HAVE OCCURRED POST-USE, POLYCARBONATE COMPONENTS LIKE THIS ARE KNOWN TO BE SUSCEPTIBLE TO ENVIRONMENTAL STRESS CRACKING (ESC), PARTICULARLY WHEN EXPOSED TO LIPOPHILIC AGENTS AND MECHANICAL STRESS. THE USE OF NON-ORIGINAL PARTS INTRODUCES MATERIAL-RELATED RISKS THAT CANNOT BE FULLY EXCLUDED AS CONTRIBUTING FACTORS 2) ANOTHER POTENTIAL CONTRIBUTOR, THOUGH CONSIDERED UNLIKELY, RELATES TO THE GAS SUPPLY SYSTEM. A LACK OF GAS PRESSURE FROM THE WALL SOURCE OR IMPROPER CALIBRATION OF THE SECHRIST GAS BLENDER COULD, HYPOTHETICALLY, IMPAIR OXYGEN TRANSFER. HOWEVER, SUCH ISSUES WOULD TYPICALLY AFFECT MULTIPLE PATIENTS USING THE SAME SYSTEM INFRASTRUCTURE. IN THIS CASE, NO ABNORMAL GAS SUPPLY CONDITIONS WERE NOTED, AND NO OTHER COMPLAINTS WERE REPORTED FROM THE FACILITY, MAKING THIS SCENARIO LESS PROBABLE. 3) THE CUSTOMER REPORTED OBSERVING "WHITE MATERIAL" DURING BLOOD SAMPLING, WHICH MAY BE RELATED TO THE PATIENT'S EXCEPTIONALLY HIGH TRIGLYCERIDE LEVEL (987 MG/DL) IN COMBINATION WITH THE USE OF PROPOFOL FOR SEDATION. BOTH HYPERTRIGLYCERIDEMIA AND PROLONGED PROPOFOL INFUSION ARE RECOGNIZED RISK FACTORS FOR MEMBRANE LUNG DYSFUNCTION AND IMPAIRED GAS EXCHANGE IN ECMO THERAPY. THIS ASSOCIATION IS SUPPORTED BY SEVERAL PUBLISHED REPORTS AND CASE STUDIES HIGHLIGHTING MECHANICAL FAILURE OF ECMO CIRCUITS DUE TO HYPERTRIGLYCERIDEMIA, PARTICULARLY IN THE SETTING OF CONTINUOUS PROPOFOL ADMINISTRATION WHILE THE DEVICE TESTING CONFIRMED THAT THE HLS SET MET PERFORMANCE SPECIFICATIONS AND NO TECHNICAL FAILURE WAS DETECTED, THE CLINICAL CONTEXT - ALONG WITH THE PUBLICATION EVIDENCE PRESENTED HERE - SUGGESTS THAT PATIENT-SPECIFIC METABOLIC FACTORS, INCLUDING HYPERTRIGLYCERIDEMIA AND THE USE OF PROPOFOL FOR SEDATION, MAY HAVE PLAYED A ROLE IN THE OBSERVED REDUCTION IN OXYGENATION. NOTABLY, THE NEED FOR MULTIPLE HLS SET EXCHANGES DUE TO ELEVATED TRIGLYCERIDE LEVELS FURTHER SUPPORTS THE LIKELIHOOD OF UNDERLYING METABOLIC STRESS. THESE CONDITIONS APPEAR TO BE THE MORE PLAUSIBLE CONTRIBUTORS TO THE PATIENT¿S CLINICAL DETERIORATION RATHER THAN A DEFICIENCY IN THE DEVICE ITSELF. IN CONCLUSION, FROM A TECHNICAL STANDPOINT, THE RETURNED HLS SET DID NOT DEMONSTRATE ANY INHERENT DEVICE MALFUNCTION THAT COULD EXPLAIN THE REPORTED OXYGENATION IMPAIRMENT. ALL RELEVANT FUNCTIONAL TESTS, AFTER REMOVAL OF THIRD-PARTY COMPONENTS, WERE PASSED SUCCESSFULLY. THE BROKEN GAS FILTER FITTING AND CRACKED STOPCOCK, WHILE NON-CONFORMING, WERE NOT LINKED TO GAS EXCHANGE PERFORMANCE FAILURE. CLINICALLY, HOWEVER, THE PRESENCE OF HIGH TRIGLYCERIDES AND LIPID-BASED SEDATIVE (PROPOFOL) USE MAY HAVE CONTRIBUTED TO COMPROMISED MEMBRANE PERFORMANCE AND REDUCTION OF AGGREGATE MEMBRANE AREA VIA LIPID FOULING - A PHENOMENON NOT DIRECTLY DETECTABLE THROUGH STANDARD POST-USE TECHNICAL TESTING. THE PATIENT¿S SEVERE ILLNESS, INSTABILITY, AND COMPLEX CLINICAL PICTURE (INCLUDING FLU A AND MULTIPLE CIRCUIT EXCHANGES WITHIN A SHORT TIMEFRAME) LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. LAST, NO DIMINUTION IN PRODUCT PERFORMANCE WAS DETECTED IN INTERNAL TESTING OR SUSPECTED OUTSIDE OF THE REVELATION OF SEVERE HYPERLIPIDEMIA (987 MG/DL)." ACCORDING TO THE INSTRUCTIONS FOR USE OF THE HLS SET, IN CHAPTER "SAFETY INSTRUCTIONS FOR THE HLS SET ADVANCED", IT IS STATED THAT " AN INSUFFICIENT SUPPLY OF POWER AND OXYGEN CAN LEAD TO INADEQUATE PATIENT SUPPORT. ENSURE THERE IS A SUFFICIENT SUPPLY OF MEDICAL OXYGEN. ENSURE THAT THE BATTERIES OF THE DRIVE ARE FULLY CHARGED", AS WELL AS "USE A MEDICAL GAS SUPPLY WITH DRY AIR AND OXYGEN. DO NOT USE AN AIR HUMIDIFIER IN THE GAS SUPPLY". IN CHAPTER "SAFETY INSTRUCTIONS FOR THE OXYGENATOR", IT IS INDICATED THAT "EXCEEDING THE PERMISSIBLE MAXIMUM VALUES DAMAGES THE OXYGENATOR. THIS CAN LEAD TO EMBOLISMS AND INADEQUATE PATIENT SUPPORT. OBSERVE THE PERMISSIBLE MAXIMUM VALUES FOR BLOOD FLOW AND GAS FLOW AS WELL AS PRESSURE ON THE BLOOD SIDE AND ON THE GAS SIDE", AS WELL "FLUSH THE OXYGENATOR REGULARLY TO MAINTAIN OR INCREASE THE GAS EXCHANGE CAPACITY. IF USED FOR LONGER THAN 6 HOURS: FLUSH THE OXYGENATOR AT LEAST ONCE A DAY. CARRY OUT FLUSHING ONLY WHEN THE PATIENT'S BLOOD GAS LEVELS PERMIT. MONITOR THE PATIENT'S BLOOD GAS PARAMETERS CAREFULLY DURING THE FLUSHING PROCESS. DISCONTINUE FLUSHING THE GAS FIBERS IF SIGNS OF HYPOCAPNIA OCCUR". LASTLY, IN THE SAME CHAPTER IT IS SAID THAT "THE OXYGENATOR PERFORMANCE IS RESTRICTED IF THE OXYGENATOR OPERATING POSITION IS INCORRECT OR IF THE GAS FLOW IS OBSTRUCTED. THIS CAN LEAD TO INADEQUATE PATIENT SUPPORT OR AIR EMBOLISMS IN THE PATIENT. ONLY OPERATE THE OXYGENATOR WITH THE GAS INLET AT THE TOP. DO NOT OCCLUDE OR BLOCK THE GAS OUTLET". THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT (SECOND SET) WERE REVIEWED ON 2025-05-26. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "LOW OXYGENATION" COULD NOT BE LINKED TO A DEVICE DEFECT. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: NEW INFORMATION RECEIVED ON 2025-01-27 THAT THE PATIENT IS 33 YEARS OLD, WEIGHT 149.9KG AND A MALE. NO VISIBLE CLOTS WERE DETECTED. AFTER THE SECOND EXCHANGE THE PATIENT IMPROVED SLIGHTLY BUT THAT DID NOT LAST LONG. THE PATIENT HAD HIGH LEVELS OF TRIGLYCERIDES AND CONTINUED TO HAVE LOW OXYGENATION AND EXPIRED LATER IN THE AFTERNOON. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239496&pc=BZD | 2518422-2025-008405 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239924&pc=BZD | 2518422-2025-008420 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD); Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER AND COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21239929&pc=MNS | 2518422-2025-101237 | 2025-01-23 05:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAD BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240000&pc=BZD | 2518422-2025-008425 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240094&pc=DTZ | 3013876692-2025-00005 | 2025-01-26 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-01-27 05:00 | DTZ | TUBING SET | Gas Output Problem | Low Oxygen Saturation | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT AN HLS SET WAS REPLACED AND AFTER THE REPLACEMENT (A FEW HOURS LATER) THE CUSTOMER NOTICED A DECREASE IN OXYGENATION VALUES. THE PATIENT OXYGENATION WAS IN THE LOW 80S. THE CUSTOMER CHECKED THE GAS SOURCE TO ENSURE PROPER CONNECTION AND THEY INCREASED THE SWEEP TEMPORARILY TO REMOVE ANY POTENTIAL CONDENSATION. THE POST OXYGENATION (PAO2) DECREASED FURTHER AND THE CUSTOMER EXCHANGED THE HLS SET AGAIN. AS THE PATIENT SUFFERED A LOW OXYGENATION, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: THE AFFECTED PRODUCT WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-04-28 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. ALL PERFORMANCE TESTS RELATED TO THE REPORTED FAILURE WERE FUNCTIONING AS PER SPECIFICATION. FURTHER, DURING VISUAL INSPECTION THIRD PARTY ARTICLES, CONNECTED TO THE MODULE, WERE FOUND. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT A HLS SET WAS REPLACED AND AFTER THE REPLACEMENT (A FEW HOURS LATER) THE CUSTOMER NOTICED A DECREASE IN OXYGENATION VALUES. THE PATIENT OXYGENATION WAS IN THE LOW 80`S. THE CUSTOMER CHECKED THE GAS SOURCE TO ENSURE PROPER CONNECTION AND THEY INCREASED THE SWEEP TEMPORARILY TO REMOVE ANY POTENTIAL CONDENSATION. THE POST OXYGENATION (PAO2) DECREASED FURTHER AND THE CUSTOMER EXCHANGED THE HLS SET AGAIN. NEW INFORMATION RECEIVED ON 2025-01-27 THAT THE PATIENT IS 33 YEARS OLD, WEIGHT 149.9KG AND A MALE. NO VISIBLE CLOTS WERE DETECTED. AFTER THE SECOND EXCHANGE TO THE THIRD SET THE PATIENT IMPROVED SLIGHTLY BUT THAT DID NOT LAST LONG. THE PATIENT HAD HIGH LEVELS OF TRIGLYCERIDES AND CONTINUED TO HAVE LOW OXYGENATION AND PASSED AWAY LATER IN THE AFTERNOON. ON 2025-02-03 THE INFORMATION WAS RECEIVED THAT THE FIRST AND THIRD USED SET WERE DISCARDED AND THE LOT NUMBERS ARE UNKNOWN. THE SECOND USED SET WAS AVAILABLE FOR INVESTIGATION. AS THE PATIENT SUFFERED A LOW OXYGENATION AND PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT (SECOND SET) WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-04-28 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. ALL PERFORMANCE TESTS RELATED TO THE REPORTED FAILURE WERE FUNCTIONING AS PER SPECIFICATION. FURTHER, DURING VISUAL INSPECTION THIRD PARTY ARTICLES, CONNECTED TO THE MODULE, WERE FOUND. FURTHER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-05-22 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT IN SCOPE INVOLVES A 33-YEAR-OLD MALE PATIENT ON VV-ECLS USING A CARDIOHELP SYSTEM WITH AN HLS SET EXPERIENCED A DECLINE IN OXYGENATION FOLLOWING A CIRCUIT EXCHANGE ON DAY SIX OF SUPPORT. WITHIN TEN HOURS OF INITIATING THE NEW CIRCUIT, POST-OXYGENATOR PAO2 DROPPED FROM 227 TO 210 MMHG (FIO2 1.0, SWEEP 6 L/MIN, BLOOD FLOW 5.8 - 6.28 L/MIN), AND THE PATIENT'S SPO2 FELL INTO THE LOW 80S. A SUBSEQUENT CIRCUIT REPLACEMENT DID NOT RESTORE ADEQUATE OXYGENATION, AND THE PATIENT SUBSEQUENTLY EXPIRED. INVESTIGATION OF THE RETURNED HLS SET - THE SECOND OF THREE CIRCUITS - REVEALED DAMAGE OF A GAS FILTER PART AND A CRACKED THIRD-PARTY STOPCOCK. THESE WERE ATTRIBUTED TO POST-USE HANDLING, WITH NO CLINICAL LEAKAGE REPORTED. INTERNAL INSPECTION SHOWED NO CLOTS OR MECHANICAL OBSTRUCTIONS AS DESCRIBED IN THE INVESTIGATION REPORT, AND THE DEVICE PASSED ALL LEAK AND FUNCTION TESTS AFTER REMOVING THE DAMAGED EXTERNAL STOPCOCK PART. ALTHOUGH, THE OBSERVED CRACKING OF THE THIRD-PARTY STOPCOCK MAY BE ASSOCIATED WITH ENVIRONMENTAL STRESS CRACKING OF POLYCARBONATE, AS DESCRIBED IN THE HANTEL TECHNOLOGIES WHITE PAPER ON MATERIAL SELECTION. THE REPORT HIGHLIGHTS THAT PC COMPONENTS USED IN MEDICAL DEVICES CAN FAIL WHEN EXPOSED TO LIPOPHILIC AGENTS, WHICH ARE COMMONLY USED IN CRITICAL CARE AND INTERVENTIONAL PROCEDURES. SUCH EXPOSURE - ESPECIALLY UNDER HOOP STRESS FROM HAND-TIGHTENING CONNECTIONS - CAN RESULT IN CRAZING AND CRACKING, EVEN IN THE ABSENCE OF VISIBLE DEFECTS UNDER NORMAL CONDITIONS. SPECIFICALLY, THE PAPER DEMONSTRATED THAT PC LUERS AND STOPCOCKS SUBJECTED TO LIPID-BASED AGENTS SUCH AS LIPIODOL OR SIMULATED CONTRAST MEDIA (1PERCENTAGE POPPY SEED OIL IN ETHYL STEARATE) SHOWED MICROCRACKS AND, EVENTUALLY, FULL STRUCTURAL FAILURE AFTER MECHANICAL STRESS. THE COMBINATION OF CHEMICAL EXPOSURE AND MECHANICAL FORCE AMPLIFIED THE HOOP STRESS BEYOND THE MATERIAL'S RESISTANCE THRESHOLD, LEADING TO ESC, A POTENTIAL MECHANISM CONSISTENT WITH THE CRACKED STOPCOCK OBSERVED IN THIS CASE. GAS PERMEABILITY WAS CONFIRMED DURING THE PRODUCT INVESTIGATION, INDICATING INTACT CO2 REMOVAL AND O2 DELIVERY FUNCTIONS. CLINICALLY, NO THROMBI WERE NOTED, ROUTINE FLUSHING WAS PERFORMED, AND TRANSMEMBRANE PRESSURE REMAINED WITHIN NORMAL LIMITS. THE DEVICE HAD BEEN PRIMED NINE DAYS PRIOR TO USE. HOWEVER, ACCORDING TO THE INSTRUCTIONS FOR USE, THE SET SHOULD ONLY BE PRIMED ¿[¿] SHORTLY BEFORE USE [¿]¿. NONETHELESS, THIS DEVIATION IS NOT EXPECTED TO HAVE IMPACTED THE INCIDENT DESCRIBED. ALTHOUGH THE TECHNICAL INVESTIGATION CONCLUDED THAT THERE WERE NO INDICATIONS OF POTENTIAL IMPAIRMENT TO OXYGENATION PERFORMANCE, SEVERAL CLINICAL AND CONTEXTUAL FACTORS MERIT CONSIDERATION AS POSSIBLE CONTRIBUTORS TO THE OBSERVED OXYGENATION DECLINE: 1) THE INVESTIGATION IDENTIFIED A CRACK IN THE THIRD-PARTY THREE-WAY STOPCOCK CONNECTED TO THE SAMPLE LINE AT THE QUICK-VENT LUER. WHILE DEEMED LIKELY TO HAVE OCCURRED POST-USE, POLYCARBONATE COMPONENTS LIKE THIS ARE KNOWN TO BE SUSCEPTIBLE TO ENVIRONMENTAL STRESS CRACKING (ESC), PARTICULARLY WHEN EXPOSED TO LIPOPHILIC AGENTS AND MECHANICAL STRESS. THE USE OF NON-ORIGINAL PARTS INTRODUCES MATERIAL-RELATED RISKS THAT CANNOT BE FULLY EXCLUDED AS CONTRIBUTING FACTORS 2) ANOTHER POTENTIAL CONTRIBUTOR, THOUGH CONSIDERED UNLIKELY, RELATES TO THE GAS SUPPLY SYSTEM. A LACK OF GAS PRESSURE FROM THE WALL SOURCE OR IMPROPER CALIBRATION OF THE SECHRIST GAS BLENDER COULD, HYPOTHETICALLY, IMPAIR OXYGEN TRANSFER. HOWEVER, SUCH ISSUES WOULD TYPICALLY AFFECT MULTIPLE PATIENTS USING THE SAME SYSTEM INFRASTRUCTURE. IN THIS CASE, NO ABNORMAL GAS SUPPLY CONDITIONS WERE NOTED, AND NO OTHER COMPLAINTS WERE REPORTED FROM THE FACILITY, MAKING THIS SCENARIO LESS PROBABLE. 3) THE CUSTOMER REPORTED OBSERVING "WHITE MATERIAL" DURING BLOOD SAMPLING, WHICH MAY BE RELATED TO THE PATIENT'S EXCEPTIONALLY HIGH TRIGLYCERIDE LEVEL (987 MG/DL) IN COMBINATION WITH THE USE OF PROPOFOL FOR SEDATION. BOTH HYPERTRIGLYCERIDEMIA AND PROLONGED PROPOFOL INFUSION ARE RECOGNIZED RISK FACTORS FOR MEMBRANE LUNG DYSFUNCTION AND IMPAIRED GAS EXCHANGE IN ECMO THERAPY. THIS ASSOCIATION IS SUPPORTED BY SEVERAL PUBLISHED REPORTS AND CASE STUDIES HIGHLIGHTING MECHANICAL FAILURE OF ECMO CIRCUITS DUE TO HYPERTRIGLYCERIDEMIA, PARTICULARLY IN THE SETTING OF CONTINUOUS PROPOFOL ADMINISTRATION WHILE THE DEVICE TESTING CONFIRMED THAT THE HLS SET MET PERFORMANCE SPECIFICATIONS AND NO TECHNICAL FAILURE WAS DETECTED, THE CLINICAL CONTEXT - ALONG WITH THE PUBLICATION EVIDENCE PRESENTED HERE - SUGGESTS THAT PATIENT-SPECIFIC METABOLIC FACTORS, INCLUDING HYPERTRIGLYCERIDEMIA AND THE USE OF PROPOFOL FOR SEDATION, MAY HAVE PLAYED A ROLE IN THE OBSERVED REDUCTION IN OXYGENATION. NOTABLY, THE NEED FOR MULTIPLE HLS SET EXCHANGES DUE TO ELEVATED TRIGLYCERIDE LEVELS FURTHER SUPPORTS THE LIKELIHOOD OF UNDERLYING METABOLIC STRESS. THESE CONDITIONS APPEAR TO BE THE MORE PLAUSIBLE CONTRIBUTORS TO THE PATIENT¿S CLINICAL DETERIORATION RATHER THAN A DEFICIENCY IN THE DEVICE ITSELF. IN CONCLUSION, FROM A TECHNICAL STANDPOINT, THE RETURNED HLS SET DID NOT DEMONSTRATE ANY INHERENT DEVICE MALFUNCTION THAT COULD EXPLAIN THE REPORTED OXYGENATION IMPAIRMENT. ALL RELEVANT FUNCTIONAL TESTS, AFTER REMOVAL OF THIRD-PARTY COMPONENTS, WERE PASSED SUCCESSFULLY. THE BROKEN GAS FILTER FITTING AND CRACKED STOPCOCK, WHILE NON-CONFORMING, WERE NOT LINKED TO GAS EXCHANGE PERFORMANCE FAILURE. CLINICALLY, HOWEVER, THE PRESENCE OF HIGH TRIGLYCERIDES AND LIPID-BASED SEDATIVE (PROPOFOL) USE MAY HAVE CONTRIBUTED TO COMPROMISED MEMBRANE PERFORMANCE AND REDUCTION OF AGGREGATE MEMBRANE AREA VIA LIPID FOULING - A PHENOMENON NOT DIRECTLY DETECTABLE THROUGH STANDARD POST-USE TECHNICAL TESTING. THE PATIENT¿S SEVERE ILLNESS, INSTABILITY, AND COMPLEX CLINICAL PICTURE (INCLUDING FLU A AND MULTIPLE CIRCUIT EXCHANGES WITHIN A SHORT TIMEFRAME) LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. LAST, NO DIMINUTION IN PRODUCT PERFORMANCE WAS DETECTED IN INTERNAL TESTING OR SUSPECTED OUTSIDE OF THE REVELATION OF SEVERE HYPERLIPIDEMIA (987 MG/DL)." ACCORDING TO THE INSTRUCTIONS FOR USE OF THE HLS SET, IN CHAPTER "SAFETY INSTRUCTIONS FOR THE HLS SET ADVANCED", IT IS STATED THAT " AN INSUFFICIENT SUPPLY OF POWER AND OXYGEN CAN LEAD TO INADEQUATE PATIENT SUPPORT. ENSURE THERE IS A SUFFICIENT SUPPLY OF MEDICAL OXYGEN. ENSURE THAT THE BATTERIES OF THE DRIVE ARE FULLY CHARGED", AS WELL AS "USE A MEDICAL GAS SUPPLY WITH DRY AIR AND OXYGEN. DO NOT USE AN AIR HUMIDIFIER IN THE GAS SUPPLY". IN CHAPTER "SAFETY INSTRUCTIONS FOR THE OXYGENATOR", IT IS INDICATED THAT "EXCEEDING THE PERMISSIBLE MAXIMUM VALUES DAMAGES THE OXYGENATOR. THIS CAN LEAD TO EMBOLISMS AND INADEQUATE PATIENT SUPPORT. OBSERVE THE PERMISSIBLE MAXIMUM VALUES FOR BLOOD FLOW AND GAS FLOW AS WELL AS PRESSURE ON THE BLOOD SIDE AND ON THE GAS SIDE", AS WELL "FLUSH THE OXYGENATOR REGULARLY TO MAINTAIN OR INCREASE THE GAS EXCHANGE CAPACITY. IF USED FOR LONGER THAN 6 HOURS: FLUSH THE OXYGENATOR AT LEAST ONCE A DAY. CARRY OUT FLUSHING ONLY WHEN THE PATIENT'S BLOOD GAS LEVELS PERMIT. MONITOR THE PATIENT'S BLOOD GAS PARAMETERS CAREFULLY DURING THE FLUSHING PROCESS. DISCONTINUE FLUSHING THE GAS FIBERS IF SIGNS OF HYPOCAPNIA OCCUR". LASTLY, IN THE SAME CHAPTER IT IS SAID THAT "THE OXYGENATOR PERFORMANCE IS RESTRICTED IF THE OXYGENATOR OPERATING POSITION IS INCORRECT OR IF THE GAS FLOW IS OBSTRUCTED. THIS CAN LEAD TO INADEQUATE PATIENT SUPPORT OR AIR EMBOLISMS IN THE PATIENT. ONLY OPERATE THE OXYGENATOR WITH THE GAS INLET AT THE TOP. DO NOT OCCLUDE OR BLOCK THE GAS OUTLET". THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT (SECOND SET) WERE REVIEWED ON 2025-05-26. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "LOW OXYGENATION" COULD NOT BE LINKED TO A DEVICE DEFECT. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: NEW INFORMATION RECEIVED ON 2025-01-27 THAT THE PATIENT IS 33 YEARS OLD, WEIGHT 149.9KG AND A MALE. NO VISIBLE CLOTS WERE DETECTED. AFTER THE SECOND EXCHANGE THE PATIENT IMPROVED SLIGHTLY BUT THAT DID NOT LAST LONG. THE PATIENT HAD HIGH LEVELS OF TRIGLYCERIDES AND CONTINUED TO HAVE LOW OXYGENATION AND EXPIRED LATER IN THE AFTERNOON. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240111&pc=BZD | 2518422-2025-008430 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-27 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, SKIN IRRITATION, REDUCED CARDIOPULMONARY RESERVE AND 2 SURGERIES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240384&pc=NPT | 2015691-2025-00478 | 2015-01-01 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-27 05:00 | NPT | SAPIEN VALVE UNKNOWN | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED IN MEDICAL ARTICLE "LACTATE DEHYDROGENASE LEVELS AFTER AORTIC VALVE REPLACEMENT: WHAT DO THEY TELL US?", A PATIENT UNDERWENT A TRANSAPICAL TAVI PROCEDURE WITH AN UNKNOWN SAPIEN VALVE IMPLANTED IN AORTIC POSITION. THE PATIENT HAD AN ANNULAR RUPTURE RESULTING IN A VENTRICULAR SEPTAL DEFECT AND RUPTURE OF THE PAPILLARY MUSCLE WITH A CONSECUTIVE FLAIL OF THE MITRAL ANTERIOR LEAFLET RESULTING IN SEVERE INSUFFICIENCY. AFTER CARDIOPULMONARY BYPASS STABILIZATION AND EXTUBATION, THE PATIENT DECIDED NOT TO UNDERGO ANOTHER INTERVENTION AND PASSED AWAY ON POSTOPERATIVE DAY 2. Manufacturer Narrative: THIS IS ONE OF THREE REPORTS BEING SUBMITTED FOR THIS CASE. ARTICLE CITATION: RINGS L, MAVROVA-RISTESKA L, HAEUSSLER A ET AL. LACTATE DEHYDROGENASE LEVELS AFTER AORTIC VALVE REPLACEMENT: WHAT DO THEY TELL US?. THE THORACIC AND CARDIOVASCULAR SURGEON 2024. DOI: 10.1055/A-2454-9020. IN THIS CASE, THE EXACT VALVE MODEL NUMBER IS NOT AVAILABLE. THEREFORE, SECTIONS D1 AND D4 OF THIS REPORT WILL REFLECT AN UNKNOWN EDWARDS TRANSCATHETER HEART VALVE. THE POSSIBLE PMA NUMBERS ASSOCIATED WITH AN EDWARDS SAPIEN TRANSCATHETER HEART VALVE ARE: P110021- EDWARDS SAPIEN TRANSCATHETER HEART VALVE; P130009 - EDWARDS SAPIEN XT TRANSCATHETER HEART VALVE; P140031- EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE AND EDWARDS SAPIEN 3 ULTRA HEART VALVE. THE DATES OF THE EVENTS ARE UNKNOWN; HOWEVER, THE ARTICLE PROVIDED THE FOLLOWING DATE RANGE ON PAGE 3: "(FROM JAN 2015 UNTIL DEC 2018)." THEREFORE, 01JAN2015 WAS USED AS THE OCCURRENCE DATE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS ARE INCONCLUSIVE AS SPECIFIC PATIENT AND PROCEDURAL FACTORS WERE NOT PROVIDED; HOWEVER, THE EVENT COULD BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240571&pc=OZD | 1220648-2025-25796 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-01-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED FROM ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR IMPELLA REGISTRY REGARDING A PATIENT WHO ENDURED CARDIOGENIC SHOCK WITH AN UNKNOWN/UNDETERMINED RELATIONSHIP TO THE DEVICE USED. THE PATIENT CONTINUED TO DECOMPENSATE AFTER IMPELLA PLACEMENT, DUE TO CRITICAL ILLNESS AND UNLIKELIHOOD OF RECOVERY, THE FAMILY DECIDED ON COMFORT CARE MEASURES. THE PATIENT WOULD EXPIRE APPROXIMATELY 24 HOURS AFTER START OF THIS SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25796. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240604&pc=OZD | 1220648-2025-25798 | 2025-01-07 05:00 | Death | ABIOMED, INC. | 2025-01-27 05:00 | OZD | IMPELLA CP | Fluid/Blood Leak | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND TWO DAYS AFTER START OF SUPPORT, A LEAKAGE BETWEEN IMPELLA FILTER AND PRESSURE RESERVOIR WAS OBSERVED. PROTOCOL FOR PURGE FILTER BYPASS WAS REQUESTED. REPORTEDLY THE FOLLOWING DAY MEDICAL STAFF DECIDED TO END THERAPY DUE TO THE PATIENT'S "POOR" GENERAL CONDITION. THE RELATIONSHIP BETWEEN THE PATIENT'S OUTCOME AND THE IMPELLA DEVICE USED IS UNKNOWN. NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF PURGE LEAK ¿ PUMP HAS BEEN COMPLETED. NO PRODUCT RETURNED AND NO DATA LOGS. CLINICAL DETAILS REPORTED A LEAKAGE BETWEEN THE FILTER AND PRESSURE RESERVOIR. BYPASS WAS PERFORMED. THE CAUSE OF THE PURGE LEAK - PUMP WAS MOST LIKELY A LEAK FROM THE SIDEARM BUT THE EXACT LOCATION OF THE LEAK WAS UNKNOWN. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: CLEANING. ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿ A.1 REVISED PATIENT IDENTIFIER AS IT WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25798. D.3 REVISED US ZIP CODE AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25798. E.1 REVISED LAST NAME AS IT WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25798. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25798 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.4 REVISED DEVICE MANUFACTURE DATE AS IT WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25798. H.10 INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25798.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240679&pc=OZD | 1220648-2025-25803 | 2025-01-08 05:00 | Death | ABIOMED, INC. | 2025-01-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DEVELOPED LIMB ISCHEMIA THE FOLLOWING DAY. THE PATIENT WOULD LATER EXPIRE. A PATIENT WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION CODED, RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED, AND WAS TRANSPORTED TO THE CATHETERIZATION LAB INTUBATED AND UNRESPONSIVE. THE PATIENT WAS PREPPED ON THE TABLE. RADIAL ACCESS WAS OBTAINED FOR CARDIAC CATHETERIZATION. THE PATIENT CODED AGAIN, AND CARDIOPULMONARY RESUSCITATION COMMENCED WITH ROSC ATTAINED AND MAINTAINED. REPEAT ARTERIAL BLOOD GASES AND ACTIVATED CLOTTING TIME (ACT) WERE COMPLETED EVERY FEW MINUTES AS ORDERED. AFTER ROSC, CORONARY ANGIOGRAM RESUMED WITH A TIGER CATHETER VIA RIGHT RADIAL ACCESS. PRIOR TO THE PERCUTANEOUS CORONARY INTERVENTION (PCI), THE IMPELLA CP DEVICE WAS PREPPED FOR CARDIAC SUPPORT. THE LEFT FEMORAL ACCESS OBTAINED FOR IMPELLA ACCESS, INSERTION WAS SUCCESSFUL AND CARDIAC FLOW WAS MAINTAINED AT 3.2 -3.3 FLOW THROUGHOUT THE PCI TO THE LEFT CIRCUMFLEX, MID LEFT ANTERIOR DESCENDING AND DIAGONAL BRANCH ARTERIES. ACT WAS MAINTAINED AT THERAPEUTIC RANGE BETWEEN 230 -250 RANGE. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT WITH STABLE VITAL SIGNS AND THE ACCESS SITE WAS CLEAN DRY AND INTACT (NO HEMATOMA OR BLEEDING). THE FOLLOWING DAY, HOWEVER, THE PATIENT EXPERIENCED LIMB ISCHEMIA AND A PLAN WAS MADE TO EXPLANT THE PUMP. THE PATIENT REMAINED INTUBATED AND SEDATED AND PRESSOR DEMANDS REMAINED UNCHANGED. THE NEXT DAY, THE PATIENT WAS NOTED BE IN MULTIORGAN FAILURE. GOALS WERE DISCUSSED WITH THE FAMILY AND THE PATIENT WAS MADE DO NOT RESUSCITATE/INTUBATE AND WOULD EXPIRE. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240866&pc=NPT | 2025587-2025-00629 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-27 05:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Fluid/Blood Leak; Perivalvular Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem; Central Regurgitation | Cardiac Arrest; Low Blood Pressure/ Hypotension; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-26 VALVE, INITIAL VALE DEPLOYMENT APPEARED TO BE AT A DEPTH OF APPROXIMATELY 3 MM ON THE NON-CORONARY CUSP. FOLLOWING THE IMPLANT OF THIS FIRST VALVE, HOWEVER, A DEPTH OF ABOUT -1 MM AND AORTIC INSUFFICIENCY WAS NOTED. IT WAS SUSPECTED THAT, UPON WITHDRAWAL OF THE PIGTAIL CATHETER, THE CATHETER PULLED TO VALVE UP TO THIS NEW DEPTH. A NEW VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM AND INSERTED INTO THE PATIENT. THERE WAS DIFFICULTY PASSING THE SECOND DCS THROUGH THE FIRST IMPLANTED VALVE AS THE DCS CAUGHT THE INNER AORTIC CURVE TAB. THE DCS WAS SUBSEQUENTLY PULLED INTO THE ASCENDING AORTA, RESULTING IN THE LOSS OF WIRE PLACEMENT. THE PATIENT'S BLOOD PRESSURE DROPPED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED WITH THE UNDERLYING RHYTHM BEING ASYSTOLE. LEFT FEMORAL ARTERY BIG SHEATH ACCESS WAS ESTABLISHED AND THE DCS WAS SUCCESSFULLY ADVANCED THROUGH THE OUTFLOW OF THE FIRST VALVE. THE SECOND VALVE WAS DEPLOYED TOO DEEP AT A DEPTH GREATER THAN 14 MM SO THE VALVE WAS PARTIALLY RECAPTURED. DURING CPR, THE VALVE WAS RE-DEPLOYED AT APPROXIMATELY 6 MM, BUT IT WAS NOTED THAT THE SECOND VALVE MOVED TO THE SAME LEVEL AS THE FIRST IMPLANTED VALVE. THE PROCEDURE WAS DELAYED BY AT LEAST ONE HOUR. DESPITE EFFORTS, THE PATIENT EXPERIENCED A POSSIBLE AORTIC DISSECTION, WHICH WAS CONSIDERED A POSSIBLE CAUSE OF DEATH. CORONARY FLOW WAS CONFIRMED TWICE DURING THE PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC REGURGITATION OBSERVED, WAS SEVERE PARAVALVULAR LEAK (PVL) AND SEVERE CENTRAL AORTIC REGURGITATION. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER, AND THE VALVE DISLODGED AORTIC. A MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE SECOND VALVE COULD HAVE BEEN A CONTRIBUTING FACTOR TO THE DEATH, AND BOTH DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS. ADDITIONALLY, THE PATIENT¿S UNDERLYING MEDICAL HISTORY DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-26 VALVE, INITIAL VALE DEPLOYMENT APPEARED TO BE AT A DEPTH OF APPROXIMATELY 3 MM ON THE NON-CORONARY CUSP. FOLLOWING THE IMPLANT OF THIS FIRST VALVE, HOWEVER, A DEPTH OF ABOUT -1 MM AND AORTIC INSUFFICIENCY WAS NOTED. IT WAS SUSPECTED THAT, UPON WITHDRAWAL OF THE PIGTAIL CATHETER, THE CATHETER PULLED TO VALVE UP TO THIS NEW DEPTH. A NEW VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM AND INSERTED INTO THE PATIENT. THERE WAS DIFFICULTY PASSING THE SECOND DCS THROUGH THE FIRST IMPLANTED VALVE AS THE DCS CAUGHT THE INNER AORTIC CURVE TAB. THE DCS WAS SUBSEQUENTLY PULLED INTO THE ASCENDING AORTA, RESULTING IN THE LOSS OF WIRE PLACEMENT. THE PATIENT'S BLOOD PRESSURE DROPPED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED WITH THE UNDERLYING RHYTHM BEING ASYSTOLE. LEFT FEMORAL ARTERY BIG SHEATH ACCESS WAS ESTABLISHED AND THE DCS WAS SUCCESSFULLY ADVANCED THROUGH THE OUTFLOW OF THE FIRST VALVE. THE SECOND VALVE WAS DEPLOYED TOO DEEP AT A DEPTH GREATER THAN 14 MM SO THE VALVE WAS PARTIALLY RECAPTURED. DURING CPR, THE VALVE WAS RE-DEPLOYED AT APPROXIMATELY 6 MM, BUT IT WAS NOTED THAT THE SECOND VALVE MOVED TO THE SAME LEVEL AS THE FIRST IMPLANTED VALVE. THE PROCEDURE WAS DELAYED BY AT LEAST ONE HOUR. DESPITE EFFORTS, THE PATIENT EXPERIENCED A POSSIBLE AORTIC DISSECTION, WHICH WAS CONSIDERED A POSSIBLE CAUSE OF DEATH. CORONARY FLOW WAS CONFIRMED TWICE DURING THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC REGURGITATION OBSERVED, WAS SEVERE PARAVALVULAR LEAK (PVL) AND SEVERE CENTRAL AORTIC REGURGITATION. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER, AND THE VALVE DISLODGED AORTIC. A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE SECOND VALVE COULD HAVE BEEN A CONTRIBUTING FACTOR TO THE DEATH, AND BOTH DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS. ADDITIONALLY, THE PATIENT¿S UNDERLYING MEDICAL HISTORY DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THERE WAS LIKELIHOOD OF THE MEDTRONIC GUIDEWIRE BEING REPLACED BY A NON-MEDTRONIC GUIDEWIRE DURING THE CASE. PER THE PHYSICIAN, THE MEDTRONIC GUIDEWIRE DID NOT CONTRIBUTE TO THE EVENT. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VOLUTFX-26, SERIAL/LOT #: (B)(6), UBD: 07-FEB-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {GWBC30}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. UPDATED DATA: B5. H6. H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240879&pc=NPT | 2025587-2025-00630 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-27 05:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Fluid/Blood Leak; Perivalvular Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem; Central Regurgitation | Cardiac Arrest; Low Blood Pressure/ Hypotension; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-26 VALVE, INITIAL VALE DEPLOYMENT APPEARED TO BE AT A DEPTH OF APPROXIMATELY 3 MM ON THE NON-CORONARY CUSP. FOLLOWING THE IMPLANT OF THIS FIRST VALVE, HOWEVER, A DEPTH OF ABOUT -1 MM AND AORTIC INSUFFICIENCY WAS NOTED. IT WAS SUSPECTED THAT, UPON WITHDRAWAL OF THE PIGTAIL CATHETER, THE CATHETER PULLED TO VALVE UP TO THIS NEW DEPTH. A NEW VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM AND INSERTED INTO THE PATIENT. THERE WAS DIFFICULTY PASSING THE SECOND DCS THROUGH THE FIRST IMPLANTED VALVE AS THE DCS CAUGHT THE INNER AORTIC CURVE TAB. THE DCS WAS SUBSEQUENTLY PULLED INTO THE ASCENDING AORTA, RESULTING IN THE LOSS OF WIRE PLACEMENT. THE PATIENT'S BLOOD PRESSURE DROPPED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED WITH THE UNDERLYING RHYTHM BEING ASYSTOLE. LEFT FEMORAL ARTERY BIG SHEATH ACCESS WAS ESTABLISHED AND THE DCS WAS SUCCESSFULLY ADVANCED THROUGH THE OUTFLOW OF THE FIRST VALVE. THE SECOND VALVE WAS DEPLOYED TOO DEEP AT A DEPTH GREATER THAN 14 MM SO THE VALVE WAS PARTIALLY RECAPTURED. DURING CPR, THE VALVE WAS RE-DEPLOYED AT APPROXIMATELY 6 MM, BUT IT WAS NOTED THAT THE SECOND VALVE MOVED TO THE SAME LEVEL AS THE FIRST IMPLANTED VALVE. THE PROCEDURE WAS DELAYED BY AT LEAST ONE HOUR. DESPITE EFFORTS, THE PATIENT EXPERIENCED A POSSIBLE AORTIC DISSECTION, WHICH WAS CONSIDERED A POSSIBLE CAUSE OF DEATH. CORONARY FLOW WAS CONFIRMED TWICE DURING THE PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC REGURGITATION OBSERVED, WAS SEVERE PARAVALVULAR LEAK (PVL) AND SEVERE CENTRAL AORTIC REGURGITATION. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER, AND THE VALVE DISLODGED AORTIC. A MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE SECOND VALVE COULD HAVE BEEN A CONTRIBUTING FACTOR TO THE DEATH, AND BOTH DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS. ADDITIONALLY, THE PATIENT¿S UNDERLYING MEDICAL HISTORY DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-26 VALVE, INITIAL VALE DEPLOYMENT APPEARED TO BE AT A DEPTH OF APPROXIMATELY 3 MM ON THE NON-CORONARY CUSP. FOLLOWING THE IMPLANT OF THIS FIRST VALVE, HOWEVER, A DEPTH OF ABOUT -1 MM AND AORTIC INSUFFICIENCY WAS NOTED. IT WAS SUSPECTED THAT, UPON WITHDRAWAL OF THE PIGTAIL CATHETER, THE CATHETER PULLED TO VALVE UP TO THIS NEW DEPTH. A NEW VALVE WAS LOADED INTO A NEW DELIVERY CATHETER SYSTEM AND INSERTED INTO THE PATIENT. THERE WAS DIFFICULTY PASSING THE SECOND DCS THROUGH THE FIRST IMPLANTED VALVE AS THE DCS CAUGHT THE INNER AORTIC CURVE TAB. THE DCS WAS SUBSEQUENTLY PULLED INTO THE ASCENDING AORTA, RESULTING IN THE LOSS OF WIRE PLACEMENT. THE PATIENT'S BLOOD PRESSURE DROPPED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED WITH THE UNDERLYING RHYTHM BEING ASYSTOLE. LEFT FEMORAL ARTERY BIG SHEATH ACCESS WAS ESTABLISHED AND THE DCS WAS SUCCESSFULLY ADVANCED THROUGH THE OUTFLOW OF THE FIRST VALVE. THE SECOND VALVE WAS DEPLOYED TOO DEEP AT A DEPTH GREATER THAN 14 MM SO THE VALVE WAS PARTIALLY RECAPTURED. DURING CPR, THE VALVE WAS RE-DEPLOYED AT APPROXIMATELY 6 MM, BUT IT WAS NOTED THAT THE SECOND VALVE MOVED TO THE SAME LEVEL AS THE FIRST IMPLANTED VALVE. THE PROCEDURE WAS DELAYED BY AT LEAST ONE HOUR. DESPITE EFFORTS, THE PATIENT EXPERIENCED A POSSIBLE AORTIC DISSECTION, WHICH WAS CONSIDERED A POSSIBLE CAUSE OF DEATH. CORONARY FLOW WAS CONFIRMED TWICE DURING THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC REGURGITATION OBSERVED, WAS SEVERE PARAVALVULAR LEAK (PVL) AND SEVERE CENTRAL AORTIC REGURGITATION. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER, AND THE VALVE DISLODGED AORTIC. A MEDTRONIC (CONFIDA) GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE SECOND VALVE COULD HAVE BEEN A CONTRIBUTING FACTOR TO THE DEATH, AND BOTH DCS COULD BE EXCLUDED AS CONTRIBUTING FACTORS. ADDITIONALLY, THE PATIENT¿S UNDERLYING MEDICAL HISTORY DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THERE WAS LIKELIHOOD OF THE MEDTRONIC GUIDEWIRE BEING REPLACED BY A NON-MEDTRONIC GUIDEWIRE DURING THE CASE. PER THE PHYSICIAN, THE MEDTRONIC GUIDEWIRE DID NOT CONTRIBUTE TO THE EVENT. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VOLUTFX-26, SERIAL/LOT #: (B)(6), UBD: 07-FEB-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: H6. PATIENT CODE/ANNEX E E062101 AND DEVICE CODES/ANNEX A A050405 AND A050406 WERE INADVERTENTLY ADDED TO THE PREVIOUS SUPPLEMENTAL REPORT. THOSE CODES WERE OMITTED FROM THIS REPORT AS THEY ARE ASSOCIATED WITH REGULATORY REPORT NUMBER 2025587-2025-00629. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {GWBC30}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} UPDATED DATA: B5. H6. H11. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21240963&pc=DSQ | 2916596-2025-00463 | 2024-08-01 04:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: EVENT SUMMARY: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE, "SURVIVAL IN END-STAGE HF PATIENTS INELIGIBLE FOR CARDIAC TRANSPLANT CAN EXCEED 10 YEARS WITH LVAD 'DESTINATION THERAPY'" THAT DESTINATION THERAPY (DT) PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD) ACHIEVED HIGH RATES OF SURVIVAL AND QUALITY OF LIFE WHEN COMPARED TO PATIENTS WITHOUT AN LVAD. THERE WERE 9 PATIENTS EXAMINED WITH VARYING DEVICES, THE HEARTMATE II, HEARTMATE 3, AND HEARTWARE LVADS. THE MEDIAN AGE AT IMPLANT WAS 71 YEARS. OF THE 9 PATIENTS SELECTED, 5 HAD NON-ISCHEMIC CARDIOMYOPATHY WITH A MEAN LEFT VENTRICULAR (LV) EJECTION FRACTION OF 16%. OF THE 9 TOTAL PATIENTS THERE WERE 4 WHO HAD SEVERE RIGHT VENTRICULAR DYSFUNCTION AS CLASSIFIED BY ECHOCARDIOGRAM CRITERIA. THE MEDIAN GFR WAS 40 ML/MIN/1.73M2. THE MEDIAN STAY IN THE INTENSIVE CARE UNIT (ICU) WAS 2 DAYS WITH 24 DAYS SPENT IN THE HOSPITAL OVERALL. SURVIVAL RATES WERE 100% AT 5 YEARS, 88% AT 7 YEARS, AND 4 PATIENTS LIVED BEYOND 10 YEARS WITH THE LVAD IN-SITU. THERE WERE 5 PATIENTS THAT REQUIRED ADMISSION DUE TO AN LVAD-RELATED INFECTION, 6 PATIENTS DUE TO GASTRO-INTESTINAL BLEEDING, AND 2 PATIENTS DUE TO STROKE. THE PATIENTS WERE CARED FOR THROUGH FREQUENT CLINICAL REVIEWS, CLOSE ANTICOAGULATION MANAGEMENT, AND PARTICIPATION IN LONG-TERM EXERCISE PROGRAM. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. EVENT DATE WAS ENTERED AS THE DATE OF PUBLICATION 01AUG2024, SINCE THE DATE OF DATA COLLECTION WAS NOT PROVIDED. MCLEAN, J., BAUMWOL, J., SHAH, A., SILBERT, B., HAYES, H., & DELLA-BOSCA, F. (2024). SURVIVAL IN END-STAGE HF PATIENTS INELIGIBLE FOR CARDIAC TRANSPLANT CAN EXCEED 10 YEARS WITH LVAD ¿DESTINATION THERAPY.¿ HEART LUNG AND CIRCULATION, 33, S287. HTTPS://DOI.ORG/10.1016/J.HLC.2024.06.350. REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21241056&pc=NKM | 2135147-2025-00402 | 2021-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-01-27 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Incomplete Coaptation | Cardiac Tamponade; Cardiogenic Shock; Pericardial Effusion; Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation | P100009 |  | 1 | Event Description: THE ARTICLE "COMPARATIVE ANALYSIS OF CENTRAL AND NONCENTRAL DEGENERATIVE MITRAL REGURGITATION (DMR) TREATED WITH TRANSCATHETER MITRAL VALVE EDGE-TO-EDGE REPAIR (TEER)" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE PROCEDURAL OUTCOMES AND PROGNOSIS BETWEEN NONCENTRAL DMR PATIENTS, OUTSIDE CLINICAL TRIAL ANATOMICAL CRITERIA, AND CENTRAL DMR PATIENTS UNDERGOING M-TEER. DEVICES MENTIONED INCLUDE MITRACLIP G3. THE ARTICLE CONCLUDED THAT NONCENTRAL DMR PATIENTS UNDERGOING M-TEER ACHIEVE SIMILAR PROCEDURAL SUCCESS RATES WITHOUT INCREASED RISK OF COMPLICATIONS COMPARED TO CENTRAL DMR PATIENTS. [(B)(6).] THE TIME FRAME OF THE STUDY WAS JANUARY 2021 TO FEBRUARY 2024. A TOTAL OF 136 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 69 AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED DEGENERATIVE MITRAL REGURGITATION, HYPERTENSION, CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, HEART FAILURE, DIABETES, CHRONIC LUNG DISEASE, RENAL INSUFFICIENCY, HYPERTROPHIC CARDIOMYOPATHY, PRIOR MALIGNANT TUMOR SURGERY, PRIOR OPEN-HEART SURGERY, PRIOR SURGICAL VALVE REPAIR, PRIOR TRANSCATHETER VALVE REPAIR, AORTIC VALVE REGURGITATION, TRICUSPID VALVE REGURGITATION, MITRAL ANNULAR CALCIFICATION, PULMONARY ARTERY HYPERTENSION, LEAFLET FLAIL, LEAFLET PROLAPSE. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), PERICARDIAL EFFUSION, DEATH, SURGICAL INTERVENTION, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, HEART FAILURE, RECURRENT MR, HOSPITALIZATION Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP G3 WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DEGENERATIVE MITRAL REGURGITATION, HYPERTENSION, CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, HEART FAILURE, DIABETES, CHRONIC LUNG DISEASE, RENAL INSUFFICIENCY, HYPERTROPHIC CARDIOMYOPATHY, PRIOR MALIGNANT TUMOR SURGERY, PRIOR OPEN HEART SURGERY, PRIOR SURGICAL VALVE REPAIR, PRIOR TRANSCATHETER VALVE REPAIR, AORTIC VALVE REGURGITATION, TRICUSPID VALVE REGURGITATION, MITRAL ANNULAR CALCIFICATION, PULMONARY ARTERY HYPERTENSION, LEAFLET FLAIL, LEAFLET PROLAPSE. COMPLICATIONS REPORTED INCLUDED SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), PERICARDIAL EFFUSION, DEATH, SURGICAL INTERVENTION, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, HEART FAILURE, RECURRENT MR, HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: COMPARATIVE ANALYSIS OF CENTRAL AND NONCENTRAL DEGENERATIVE MITRAL REGURGITATION (DMR) TREATED WITH TRANSCATHETER MITRAL VALVE EDGE-TO-EDGE REPAIR (TEER).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21241264&pc=OZD | 1220648-2025-25819 | 2025-01-10 05:00 | Death | ABIOMED, INC. | 2025-01-27 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINTS TEAM RECEIVED AN IMPACT STUDY THAT FOUND A POSSIBLE RELATIONSHIP BETWEEN THE IMPELLA DEVICE TO THE PATIENT'S BLEEDING. THE PATIENT HAD A BLOOD COAGULATION DISORDER THAT WAS POSSIBLY CAUSED BY IMPELLA THERAPY. THE PATIENT NEEDED A TRANSFUSION FOR THE BLOOD LOSS. THE PATIENT SURVIVED THE EVENT.Event Description: THE COMPLAINT TEAM RECEIVED NEW INFORMATION FROM AN IMPACT STUDY THAT REPORTED THE PATIENT WENT INTO CARDIAC ARREST, WENT ASYSTOLE AND EXPIRED. THE STUDY FOUND THERE WAS A REMOTE RELATIONSHIP BETWEEN THE IMPELLA AND THE PATIENT'S DEATH AND A POSSIBLE RELATIONSHIP BETWEEN THE IMPELLA PROCEDURE AND THE PATIENT'S DEATH. Manufacturer Narrative: THE INVESTIGATION OF VT/VF HAS BEEN COMPLETED. THE ROOT CAUSE OF THE VT/VF CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEED COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: B.2 DEATH AND DATE OF DEATH WERE ADDED. B.5 NEW INFORMATION WAS RECEIVED. H.1 REVISED DUE TO NEW INFORMATION RECEIVED. H.6 HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE WERE ADDDED DUE TO NEW INFORMATION RECEIVED. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE (IFUS) WERE REPORTED ON THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25819 INCORRECTLY. NO IFUS ARE NEEDED TO BE REPORTED FOR THIS REPORT IN ACCORDANCE WITH UPDATED PROCEDURES. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242058&pc=OZP | 2032227-2025-116445 | 2018-03-25 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-27 05:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Break; Insufficient Flow or Under Infusion | Hyperglycemia; Diabetic Ketoacidosis | P150001 |  | 1 | Event Description: MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY ON JANUARY 10, 2025, MEDTRONIC RECEIVED NOTICE OF A DEVICE/PRODUCT LEGAL HOLD NOTIFICATION. THE REPORTER ALLEGES THAT THE USE OF ONE OR MORE MEDTRONIC MINI MED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING, POSSIBLE UNDER DELIVERY ANOMALY CAUSED THE CLAIMANT TO DECEASED IN APARTMENT IN BED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, ARE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). IT WAS ALSO REPORTED FOR THE CUSTOMER EXPERIENCED HYPERGLYCEMIA, DIABETIC KETOACIDOSIS. ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY WHEN AND IF THE AFFECTED SERIAL NUMBER WAS IDENTIFIED. THE EVENT INVOLVED PRODUCT(S) MMT-1780KPK, UNOMEDICAL, AND MMT-332A. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND THE AUTO MODE/SMART GUARD FEATURE AT THE TIME OF THE EVENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242160&pc=MVK | 3008642652-2025-00950 | 2024-12-28 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 12:15:53 ON (B)(6) 2024. AT 16:53:00, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 16:53:35, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 18:54:43 ON (B)(6) 2024. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242184&pc=MVK | 3008642652-2025-00951 | 2024-12-09 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 13:09:35 ON (B)(6) 2024. AT 22:59:24 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 22:59:58, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PAC¿S AND PVC¿S. THE ELECTRODE BELT WAS DISCONNECTED AT 00:22:15 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242189&pc=MVK | 3008642652-2025-00952 | 2024-10-24 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 21:13:30 ON (B)(6) 2024. AT 11:49:40 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT WITH CPR/ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VT. CPR/ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 11:50:51, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT TIME OF TREATMENT WAS VT @ 150 BPM WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 10 BPM WITH CPR/ ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 12:44:46 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242197&pc=MVK | 3008642652-2025-00953 | 2024-10-30 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 18:09:56 ON (B)(6) 2024. AT 19:05:08 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. CPR/MOTION ARTIFACT AND RB USE PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE WAS SHUT DOWN AT 19:05:44 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242209&pc=MVK | 3008642652-2025-00955 | 2025-01-11 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 21:42:43 ON (B)(6)2025. AT 14:29:53 ON (B)(6)2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 90 BPM WITH PVC¿S. THE RHYTHM THEN DEGRADED TO VF WITH CPR/MOTION ARTIFACT. AT 14:30:28, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 14:34:13, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 250 BPM/VF. AT 14:34:44, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM. AT 14:36:59, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 270 BPM/VF. AT 14:37:31, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 14:39:41, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 90 BPM WITH PAC¿S AND PVC¿S. THE ELECTRODE BELT WAS DISCONNECTED AT 14:40:11 ON (B)(6)2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242213&pc=MVK | 3008642652-2025-00957 | 2024-12-24 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 22:07:12 ON (B)(6) 2024. AT 04:40:16 ON (B)(6) 2024, AND ARRHYTHMIA WAS DETECTED. ECG SHOWS ECG SHOWS VF VARYING RATES/AMPLITUDES. THE LIFEVEST PROPERLY DETECTED VF. VARYING RATES/AMPLITUDES AND RESPONSE BUTTON USE DELAYED AND PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE WAS SHUT DOWN AT 04:41:13 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242218&pc=MVK | 3008642652-2025-00958 | 2024-09-28 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-27 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 04:20:56 ON (B)(6) 2024. AT 05:09:37, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 05:11:12, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF THE NON-LIFEVEST DEFIBRILLATION WAS VF WITH CPR/MOTION ARTIFACT. PER REVIEW OF THE HOLTER DATA, POST SHOCK RHYTHM WAS CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 05:15:28 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242758&pc=NPT | 2025587-2025-00643 | 2024-12-23 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-27 05:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A HEART CATHETER PROCEDURE WHICH OCCURRED APPROXIMATELY 8 MONTHS AND 20 DAYS FOLLOWING THE IMPLANT OF AN EVOLUTFX-29 VALVE, THE GUIDE CATHETER THAT HAD ENGAGED THE LEFT CORONARY OSTIA BECAME CAUGHT ON THE TOP PORTION OF THE EVOLUT. THE PHYSICIAN ATTEMPTED MULTIPLE METHODS TO DISENGAGE THE GUIDE CATHETER BUT WAS UNSUCCESSFUL. PER THE PHYSICIAN, OTHER UNSPECIFIED COMPLICATIONS OCCURRED DURING THE PROCEDURE WHICH INCREASED THE DIFFICULTY OF DISENGAGING THE GUIDE CATHETER. FOLLOWING THE PROCEDURE, THE PATIENT EXPERIENCED CARDIAC ARREST IN THE INTENSIVE CARE UNIT AND SUBSEQUENTLY DIED.Event Description: IT WAS REPORTED THAT DURING A HEART CATHETER PROCEDURE WHICH OCCURRED APPROXIMATELY 8 MONTHS AND 20 DAYS FOLLOWING THE IMPLANT OF AN EVOLUTFX-29 VALVE, THE GUIDE CATHETER THAT HAD ENGAGED THE LEFT CORONARY OSTIA BECAME CAUGHT ON THE TOP PORTION OF THE EVOLUT. THE PHYSICIAN ATTEMPTED MULTIPLE METHODS TO DISENGAGE THE GUIDE CATHETER BUT WAS UNSUCCESSFUL. PER THE PHYSICIAN, OTHER UNSPECIFIED COMPLICATIONS OCCURRED DURING THE PROCEDURE WHICH INCREASED THE DIFFICULTY OF DISENGAGING THE GUIDE CATHETER. FOLLOWING THE PROCEDURE, THE PATIENT EXPERIENCED CARDIAC ARREST IN THE INTENSIVE CARE UNIT AND SUBSEQUENTLY DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY OF CORONARY DISEASE WAS A CONTR IBUTING FACTOR TO THE DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, B7. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242844&pc=NIQ | 2124215-2025-03418 | 2025-01-13 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-27 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY. IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2023, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE DISTAL LEFT CIRCUMFLEX ARTERY (LCX) WITH 90% STENOSIS AND WAS 14 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.50 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 16 MM SYNERGY STENT SYSTEM. THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. POST-DILATATION WAS NOT PERFORMED. THE SECOND TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO MID LAD WITH 90% STENOSIS AND WAS 30 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.75 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.75 MM X 32 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. IT WAS NOTED THAT THE SUBJECT WAS ALSO IMPLANTED WITH AN ADDITIONAL DRUG ELUTING STENT. THREE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2025, THE SUBJECT DIED. AT THE TIME OF THE EVENT THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. NO ACTION WAS TAKEN TO TREAT THE EVENT, AND THE PRIMARY CAUSE OF DEATH IS UNKNOWN. IT IS ALSO UNKNOWN IF AN AUTOPSY WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242855&pc=NIQ | 2124215-2025-03428 | 2025-01-13 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-27 05:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: SYNERGY CHINA REGISTRY: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2023, THE SUBJECT WAS REFERRED FOR CARDIAC CATHETERIZATION. THE FIRST TARGET LESION WAS LOCATED IN THE DISTAL LEFT CIRCUMFLEX ARTERY (LCX) WITH 90% STENOSIS AND WAS 14 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.50 MM. THE FIRST TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.50 MM X 16 MM SYNERGY STENT SYSTEM. THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. POST-DILATATION WAS NOT PERFORMED. THE SECOND TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD) EXTENDING UP TO MID LAD WITH 90% STENOSIS AND WAS 30 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 2.75 MM. THE SECOND TARGET LESION WAS TREATED WITH PRE-DILATATION AND PLACEMENT OF 2.75 MM X 32 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. IT WAS NOTED THAT THE SUBJECT WAS ALSO IMPLANTED WITH AN ADDITIONAL DRUG ELUTING STENT. THREE DAYS LATER, THE SUBJECT WAS DISCHARGED ON ASPIRIN AND CLOPIDOGREL. IN (B)(6) 2025, THE SUBJECT DIED. AT THE TIME OF THE EVENT THE SUBJECT WAS ON ASPIRIN AND CLOPIDOGREL. NO ACTION WAS TAKEN TO TREAT THE EVENT, AND THE PRIMARY CAUSE OF DEATH IS UNKNOWN. IT IS ALSO UNKNOWN IF AN AUTOPSY WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242872&pc=FKX | 0008030665-2025-00217 | 2025-01-11 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-27 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pneumonia | K181108 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP DOCUMENTATION RECEIVED FROM THE PATIENT¿S HOME PROGRAM FACILITY ADMINISTRATOR (FA) CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2025 WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. PER THE FA, THE PATIENT¿S CAUSE OF DEATH IS UNCONFIRMED DUE TO THE PENDING AUTOPSY RESULTS, HOWEVER IT WAS LIKELY A CARDIAC ARREST. ALTHOUGH THE SPECIFICS ARE LARGELY UNKNOWN, IT APPEARS THE PATIENT COMPLETED THEIR LAST CCPD TREATMENT ON (B)(6) 2025 WITHOUT REPORTED ISSUE. PER THE FA¿S RESPONSE, THERE WAS NO ALLEGATION THE SERIOUS ADVERSE EVENTS WERE RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION.Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP DOCUMENTATION RECEIVED FROM THE PATIENT¿S HOME PROGRAM FACILITY ADMINISTRATOR (FA) CONFIRMED THE PATIENT EXPIRED ON (B)(6)2025 WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. PER THE FA, THE PATIENT¿S CAUSE OF DEATH IS UNCONFIRMED DUE TO THE PENDING AUTOPSY RESULTS, HOWEVER IT WAS LIKELY A CARDIAC ARREST. ALTHOUGH THE SPECIFICS ARE LARGELY UNKNOWN, IT APPEARS THE PATIENT COMPLETED THEIR LAST CCPD TREATMENT ON (B)(6)2025 WITHOUT REPORTED ISSUE. PER THE FA¿S RESPONSE, THERE WAS NO ALLEGATION THE SERIOUS ADVERSE EVENTS WERE RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN PD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF A PROBABLE CARDIAC ARREST AND DEATH, AS THE PATIENT WAS CONNECTED TO THE DEVICE WHEN SHE EXPIRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN (AUTOPSY RESULTS AND DEATH CERTIFICATE UNAVAILABLE); THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON FOLLOW-UP DOCUMENTATION FROM THE HOME PROGRAMS¿ FA, THERE WAS NO ALLEGATION OR INDICATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) WAS DEFICIENT OR MALFUNCTIONED IN ANY WAY. CARDIAC DISEASE IS THE LEADING CAUSE OF DEATH IN DIALYSIS PATIENTS, AND MORTALITY RATES ARE UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION PRESENTLY AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN PATIENT¿S CARDIAC ARREST AND DEATH. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS CONNECTED TO THE DEVICE WHEN SHE EXPIRED, THE PENDING AUTOPSY REPORT, AND NO CONFIRMED CAUSE OF DEATH; THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE LIBERTY SELECT CYCLER FROM HAVING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS EXPERIENCED BY THE PATIENT. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE CYCLER WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT ANY FAILURES OR PROBLEMS. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. SYSTEM AIR LEAK, TEACH PUMP, AND VALVE ACTUATION TESTS WERE PERFORMED AN PASSED. AN INTERNAL VISUAL OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242891&pc=MIH | 2247858-2025-00018 | 2024-03-26 04:00 | Death | BOLTON MEDICAL, INC. | 2025-01-27 05:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P190015 |  | 1 | Event Description: "SUBJECT (B)(6) WAS CONTACTED ON (B)(6) 2024 TO SCHEDULE THEIR 2-YEAR FOLLOW-UP, AND THE SITE WAS TOLD BY THE SUBJECT'S SPOUSE THAT THE SUBJECT HAD PASSED AWAY AT HOME ON (B)(6) 2024. THE CAUSE OF DEATH WAS PREVIOUSLY UNKNOWN DUE TO SIGNIFICANT SITE DIFFICULTIES OBTAINING ANY SOURCE. THE DEATH CERTIFICATE WAS RECENTLY OBTAINED CONFIRMING THE DATE OF DEATH AND CAUSE OF DEATH STATES ABDOMINAL AORTIC ANEURYSM. THE DEVICE RELATEDNESS IS PENDING ASSESSMENT BY THE PRINCIPAL INVESTIGATOR. FURTHER, THE SITE IS STILL ATTEMPTING TO OBTAIN POSSIBLE HOSPITALIZATION SOURCE FROM AN OUTSIDE INSTITUTION WHICH OCCURRED SHORTLY PRIOR TO THE SUBJECT'S DEATH TO FURTHER REVIEW AGAINST THE DEATH CERTIFICATE." PATIENT OUTCOME: "SUBJECT DEATH AS NOTED ABOVE." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH FOUR DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00015, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00016, DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00017, AND DEVICE 4 IS BEING REPORTED UNDER MDR-2247858-2025-00018. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242892&pc=MIH | 2247858-2025-00015 | 2024-03-26 04:00 | Death | BOLTON MEDICAL, INC. | 2025-01-27 05:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P190015 |  | 1 | Event Description: "SUBJECT (B)(6) WAS CONTACTED ON (B)(6)2024 TO SCHEDULE THEIR 2-YEAR FOLLOW-UP, AND THE SITE WAS TOLD BY THE SUBJECT'S SPOUSE THAT THE SUBJECT HAD PASSED AWAY AT HOME ON (B)(6)2024. THE CAUSE OF DEATH WAS PREVIOUSLY UNKNOWN DUE TO SIGNIFICANT SITE DIFFICULTIES OBTAINING ANY SOURCE. THE DEATH CERTIFICATE WAS RECENTLY OBTAINED CONFIRMING THE DATE OF DEATH AND CAUSE OF DEATH STATES ABDOMINAL AORTIC ANEURYSM. THE DEVICE RELATEDNESS IS PENDING ASSESSMENT BY THE PRINCIPAL INVESTIGATOR. FURTHER, THE SITE IS STILL ATTEMPTING TO OBTAIN POSSIBLE HOSPITALIZATION SOURCE FROM AN OUTSIDE INSTITUTION WHICH OCCURRED SHORTLY PRIOR TO THE SUBJECT'S DEATH TO FURTHER REVIEW AGAINST THE DEATH CERTIFICATE." PATIENT OUTCOME: "SUBJECT DEATH AS NOTED ABOVE." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH FOUR DEVICES. DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00016, DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00017, AND DEVICE 4 IS BEING REPORTED UNDER MDR-2247858-2025-00018. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242893&pc=MIH | 2247858-2025-00016 | 2024-03-26 04:00 | Death | BOLTON MEDICAL, INC. | 2025-01-27 05:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P190015 |  | 1 | Event Description: "SUBJECT TPA-004-002 WAS CONTACTED ON 27MAR2024 TO SCHEDULE THEIR 2-YEAR FOLLOW-UP, AND THE SITE WAS TOLD BY THE SUBJECT'S SPOUSE THAT THE SUBJECT HAD PASSED AWAY AT HOME ON (B)(6) 2024. THE CAUSE OF DEATH WAS PREVIOUSLY UNKNOWN DUE TO SIGNIFICANT SITE DIFFICULTIES OBTAINING ANY SOURCE. THE DEATH CERTIFICATE WAS RECENTLY OBTAINED CONFIRMING THE DATE OF DEATH AND CAUSE OF DEATH STATES ABDOMINAL AORTIC ANEURYSM. THE DEVICE RELATEDNESS IS PENDING ASSESSMENT BY THE PRINCIPAL INVESTIGATOR. FURTHER, THE SITE IS STILL ATTEMPTING TO OBTAIN POSSIBLE HOSPITALIZATION SOURCE FROM AN OUTSIDE INSTITUTION WHICH OCCURRED SHORTLY PRIOR TO THE SUBJECT'S DEATH TO FURTHER REVIEW AGAINST THE DEATH CERTIFICATE." PATIENT OUTCOME: "SUBJECT DEATH AS NOTED ABOVE." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH FOUR DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00015, DEVICE 2 IS BEING REPORTED UNDER, DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00017, AND DEVICE 4 IS BEING REPORTED UNDER MDR-2247858-2025-00018. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242894&pc=MIH | 2247858-2025-00017 | 2024-03-26 04:00 | Death | BOLTON MEDICAL, INC. | 2025-01-27 05:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P190015 |  | 1 | Event Description: "SUBJECT (B)(6) WAS CONTACTED ON (B)(6) 2024 TO SCHEDULE THEIR 2-YEAR FOLLOW-UP, AND THE SITE WAS TOLD BY THE SUBJECT'S SPOUSE THAT THE SUBJECT HAD PASSED AWAY AT HOME ON (B)(6) 2024. THE CAUSE OF DEATH WAS PREVIOUSLY UNKNOWN DUE TO SIGNIFICANT SITE DIFFICULTIES OBTAINING ANY SOURCE. THE DEATH CERTIFICATE WAS RECENTLY OBTAINED CONFIRMING THE DATE OF DEATH AND CAUSE OF DEATH STATES ABDOMINAL AORTIC ANEURYSM. THE DEVICE RELATEDNESS IS PENDING ASSESSMENT BY THE PRINCIPAL INVESTIGATOR. FURTHER, THE SITE IS STILL ATTEMPTING TO OBTAIN POSSIBLE HOSPITALIZATION SOURCE FROM AN OUTSIDE INSTITUTION WHICH OCCURRED SHORTLY PRIOR TO THE SUBJECT'S DEATH TO FURTHER REVIEW AGAINST THE DEATH CERTIFICATE." PATIENT OUTCOME: "SUBJECT DEATH AS NOTED ABOVE." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH FOUR DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00015, DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00016, DEVICE 3 IS BEING REPORTED UNDER MDR-2247858-2025-00017, AND DEVICE 4 IS BEING REPORTED UNDER MDR-2247858-2025-00018. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242901&pc=PNJ | 9612164-2025-00429 | 2024-02-01 05:00 | Death | MEDTRONIC IRELAND | 2025-01-27 05:00 | PNJ | MICRA | Appropriate Term/Code Not Available | Unspecified Infection; Pneumothorax; Pericardial Effusion; Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE ASSOCIATION BETWEEN FRAILTY AND IN-HOSPITAL OUTCOMES IN PATIENTS UNDERGOING LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN BUT DEFINED AS IN-HOSPITAL MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED UNKNOWN VASCULAR, PERICARDIAL, OR DEVICE COMPLICATIONS, INFECTIONS, OR PNEUMOTHORAX. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/76 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ASSOCIATION BETWEEN FRAILTY AND IN-HOSPITAL OUTCOMES IN PATIENTS UNDERGOING LEADLESS PACEMAKER IMPLANTATION: A NATIONWIDE ANALYSIS. HEART RHYTHM O2. 2024; 5:85¿94. DOI: 10.1016/J.HROO.2023.12.007. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME MICRA PRODUCT ID MDT-TPS-DELSYS, SERIAL: UNKNOWN, D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21243192&pc=DSQ | 2916596-2025-00293 | 2024-08-09 04:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY (B)(6) 2024 AFTER A HEMORRHAGIC CEREBRAL VASCULAR ACCIDENT (CVA). CARE WAS WITHDRAWN AFTER NEUROSURGERY DECLINED INTERVENTION. THE DEVICE FUNCTIONED AS INTENDED. Manufacturer Narrative: MANUFACTURER'S CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21243884&pc=NPT | 2015691-2025-00483 | 2025-01-07 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-27 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Material Integrity Problem | Unspecified Vascular Problem | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 26 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION, THE VALVE WAS UNABLE TO PASS THROUGH THE ESHEATH+ AND A STRUT BENT WHILE TRYING TO PUSH THE VALVE THROUGH IT. A SECOND VALVE WAS ALSO USED IN THE SAME MANNER AND ALSO DAMAGED. THE DELIVERY SYSTEM WAS UNABLE TO ADVANCE THE FIRST TIME OUT OF THE PARTIAL EXPANDING PORTION AND THE SECOND TIME IT WAS STUCK JUST DISTAL TO THE PARTIAL EXPANDING PORTION. IT'S NOTED THERE WAS A TEAR AT THE PARTIALLY EXPANDING TO FULLY EXPANDING TRANSITION POINT OF THE FIRST ESHEATH+. THE SECOND ESHEATH+ WAS DAMAGED AFTER THE CASE IN THE PARTIALLY EXPANDING PORTION DUE TO TRYING TO GET THE VALVE OUT TO EXAMINE. THE PATIENT HAD RAPID DECOMPENSATION WITH NO EFFUSION NOTED IN THE CHEST. THE PHYSICIAN BELIEVES CAUSE WOULD POSSIBLY BE FROM RETROPERITONEAL BLEED FROM THE ILIAC ARTERY, HOWEVER NO ANGIO WAS TAKEN DURING CPR AND LITTLE TO NO FLOW MADE IT INTO THE ILLIACS TO EXAMINE IF/ WHERE THE PATIENT WAS BLEEDING FROM. ADDITIONAL INFORMATION RECEIVED NOTED IT'S STILL AN UNKNOWN CAUSE OF DEATH AND AN AUTOPSY IS BEING DONE, THOUGH THE RESULTS HAVE NOT YET BEEN MADE AVAILABLE. AFTER REVIEWING THE CASE WITH THE CT SURGEON, IT WAS NOTED IN THE ANESTHESIA CHART THAT THE PATIENT'S PRESSURE WAS MAPPING IN THE 30-40 FOR A TOTAL OF 8 MINUTES BEFORE ANY ACLS OR ACTION WAS TAKEN. IT'S UNKNOWN IF THE PATIENT HAD A BLEED SOMEWHERE OR IF THEY JUST DECOMPENSATED AND COULD NOT RECOVER. THE PERCEIVED ROOT CAUSE OF THE BENT STRUTS WAS THAT THE STRUTS CAUSE ON THE ILIAC STENTS. THE PERCEIVED ROOT CAUSE OF THE RESISTANCE WAS EITHER THE STENTS IN THE ILIAC ARTERY OR CALCIUM BURDEN. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED FROM A PRODUCT EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED: UPDATE TO B4, G3, G6, H2, H3, H6, H11. THE PRODUCT WAS RETURNED; THEREFORE, A PRODUCT EVALUATION INVESTIGATION WAS COMPLETED. DUE TO THE NATURE OF THE COMPLAINT, NO FUNCTIONAL OR DIMENSIONAL TESTING WAS PERFORMED. THE RETURNED DEVICE WAS VISUALLY EXAMINED FOR ANY ABNORMALITIES AND THE FOLLOWING WAS OBSERVED: CRIMPED VALVE: ONE (1) STRUT BENT OUTWARD AT THE INFLOW SIDE. POST-EXPANDED VALVE: BENT STRUT WAS CORRECTED. LEAFLETS WRINKLED AND DEHYDRATED DUE TO STORAGE CONDITION (PROLONGED CRIMPING) DURING THE RETURN HANDLING PROCESS. FRAME DISTORTED/CANTED. THE PROVIDED IMAGERY WAS REVIEWED, AND THE FOLLOWING WAS NOTED: ACCESS SIDE (LEFT OR RIGHT) WAS NOT SPECIFIED. PATIENT'S ACCESS VESSELS HAD PRESENCE OF CALCIFICATION, TORTUOSITY, AND UNDERSIZED VESSEL VESSELS. PER THE IFU, MINIMUM LUMINAL DIAMETER (MLD) OF 5.5MM IS REQUIRED FOR USE OF THE 14F ESHEATH+. AS THE RETURNED DEVICE MEETS SPECIFICATION WITH NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT HAS CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. WHILE THE IFU/TRAINING MANUALS SPECIFIC TO THIS COMPLAINT EVENT ARE NOT LISTED IN THE TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, REVIEW OF THE INSTRUCTIONS REVEALS SIMILAR CONTENT AS DOCUMENTED IN THE TECHNICAL SUMMARY. THEREFORE, THE REVIEW OF THE INSTRUCTIONS/MANUALS WITHIN THE TECHNICAL SUMMARY REMAIN EQUIVALENT TO THE INSTRUCTIONS/MANUALS FOR THIS COMPLAINT EVENT. THE CONTENT ADEQUATELY PROVIDES WARNINGS AND INSTRUCTIONS FOR USE REGARDING THE REPORTED COMPLAINT EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, IFU WARNINGS AND CAUTIONS, AND PHYSICIAN TRAINING. NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION.' THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR FRAME DAMAGE WAS CONFIRMED BASED ON EVALUATION OF THE RETURNED DEVICE. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY, UNDERSIZED VESSEL) AND/OR PROCEDURAL FACTORS (HIGH PUSH FORCE) LIKELY CONTRIBUTED TO THE EVENT AS REVIEW OF PROVIDED 3MENSIO IMAGERY REVEALED PRESENCE OF CALCIFICATION, TORTUOSITY, AND UNDERSIZED VESSEL DIAMETERS IN THE ACCESS VASCULATURE. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SCRATCHES OBSERVED ON THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF CALCIFIED NODULES WITHIN THE ACCESS VESSEL. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. EXCESSIVE DEVICE MANIPULATION OR HIGH PUSH FORCE CAN LEAD TO THE VALVE STRUTS INTERACTING WITH THE SHEATH SHAFT AND RESULT IN THE STRUT DAMAGE AT THE VALVE INFLOW SIDE. THE PRODUCT RISK ASSESSMENT (PRA) IS CAPTURED IN THE TECHNICAL SUMMARY, WHICH APPLIES TO THIS COMPLAINT EVENT. THE RISKS OUTLINED IN THE PRA REMAIN THE SAME. THE PRA DOCUMENTS THE DIFFICULTY ADVANCING THE DELIVERY SYSTEM THROUGH THE SHEATH, RESULTING IN DIFFICULTY OR THE INABILITY TO INTRODUCE DELIVERY SYSTEM/LOADER AND ADVANCE THROUGH SHEATH, PROLONGING PROCEDURE, WHICH DID OCCUR DURING THIS EVENT. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. AS THE COMPLAINT DID NOT EXCEED THE PRA RE-ESCALATION THRESHOLD, NO FURTHER ACTION IS REQUIRED. A CAPA/SCAR IS NOT REQUIRED BECAUSE AN EDWARDS/SUPPLIER DEFECT WHICH COULD HAVE RESULTED IN THE COMPLAINT WAS NOT CONFIRMED. PRIOR CORRECTIVE ACTION, DETAILED IN THE TECHNICAL SUMMARY, CAPTURES PRIOR HIGH PUSH FORCE INVESTIGATION AND CORRECTIVE/PREVENTATIVE ACTION. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. NO FURTHER ESCALATION IS NEEDED AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21243892&pc=NPT | 2015691-2025-00484 | 2025-01-07 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-27 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Material Integrity Problem | Unspecified Vascular Problem | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 26 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION, THE VALVE WAS UNABLE TO PASS THROUGH THE ESHEATH+ AND A STRUT BENT WHILE TRYING TO PUSH THE VALVE THROUGH IT. A SECOND VALVE WAS ALSO USED IN THE SAME MANNER AND ALSO DAMAGED. THE DELIVERY SYSTEM WAS UNABLE TO ADVANCE THE FIRST TIME OUT OF THE PARTIAL EXPANDING PORTION AND THE SECOND TIME IT WAS STUCK JUST DISTAL TO THE PARTIAL EXPANDING PORTION. IT'S NOTED THERE WAS A TEAR AT THE PARTIALLY EXPANDING TO FULLY EXPANDING TRANSITION POINT OF THE FIRST ESHEATH+. THE SECOND ESHEATH+ WAS DAMAGED AFTER THE CASE IN THE PARTIALLY EXPANDING PORTION DUE TO TRYING TO GET THE VALVE OUT TO EXAMINE. THE PATIENT HAD RAPID DECOMPENSATION WITH NO EFFUSION NOTED IN THE CHEST. THE PHYSICIAN BELIEVES CAUSE WOULD POSSIBLY BE FROM RETROPERITONEAL BLEED FROM THE ILIAC ARTERY, HOWEVER NO ANGIO WAS TAKEN DURING CPR AND LITTLE TO NO FLOW MADE IT INTO THE ILLIACS TO EXAMINE IF/ WHERE THE PATIENT WAS BLEEDING FROM. ADDITIONAL INFORMATION RECEIVED NOTED IT'S STILL AN UNKNOWN CAUSE OF DEATH AND AN AUTOPSY IS BEING DONE, THOUGH THE RESULTS HAVE NOT YET BEEN MADE AVAILABLE. AFTER REVIEWING THE CASE WITH THE CT SURGEON, IT WAS NOTED IN THE ANESTHESIA CHART THAT THE PATIENT'S PRESSURE WAS MAPPING IN THE 30-40 FOR A TOTAL OF 8 MINUTES BEFORE ANY ACLS OR ACTION WAS TAKEN. IT'S UNKNOWN IF THE PATIENT HAD A BLEED SOMEWHERE OR IF THEY JUST DECOMPENSATED AND COULD NOT RECOVER. THE PERCEIVED ROOT CAUSE OF THE BENT STRUTS WAS THAT THE STRUTS CAUSE ON THE ILIAC STENTS. THE PERCEIVED ROOT CAUSE OF THE RESISTANCE WAS EITHER THE STENTS IN THE ILIAC ARTERY OR CALCIUM BURDEN. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED FROM A PRODUCT EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED: UPDATE TO B4, G3, G6, H2, H3, H6, H11. THE PRODUCT WAS RETURNED; THEREFORE, A PRODUCT EVALUATION INVESTIGATION WAS COMPLETED. DUE TO THE NATURE OF THE COMPLAINT, NO FUNCTIONAL OR DIMENSIONAL TESTING WAS PERFORMED. SAMPLE #1: THE RETURNED DEVICE WAS VISUALLY EXAMINED FOR ANY ABNORMALITIES AND THE FOLLOWING WAS OBSERVED: CRIMPED VALVE: TWO (2) BENT VALVE STRUTS AT THE INFLOW SIDE, EXPOSED THROUGH PUNCTURED SHEATH/STRAIN RELIEF AT APPROXIMATELY 1.5" FROM COM-NUT. POST-EXPANDED VALVE: BENT STRUTS WERE CORRECTED. LEAFLETS WRINKLED AND DEHYDRATED DUE TO STORAGE CONDITION (PROLONGED CRIMPING) DURING THE RETURN HANDLING PROCESS. FRAME IS DISTORTED. THE PROVIDED IMAGERY WAS REVIEWED, AND THE FOLLOWING WAS NOTED: ACCESS SIDE (LEFT OR RIGHT) WAS NOT SPECIFIED. PATIENT'S ACCESS VESSELS HAD PRESENCE OF CALCIFICATION, TORTUOSITY, AND UNDERSIZED VESSEL VESSELS. PER THE IFU, MINIMUM LUMINAL DIAMETER (MLD) OF 5.5MM IS REQUIRED FOR USE OF THE 14F ESHEATH+. AS THE RETURNED DEVICE MEETS SPECIFICATION WITH NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT HAS CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. WHILE THE IFU/TRAINING MANUALS SPECIFIC TO THIS COMPLAINT EVENT ARE NOT LISTED IN THE TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, REVIEW OF THE INSTRUCTIONS REVEALS SIMILAR CONTENT AS DOCUMENTED IN THE TECHNICAL SUMMARY. THEREFORE, THE REVIEW OF THE INSTRUCTIONS/MANUALS WITHIN THE TECHNICAL SUMMARY REMAIN EQUIVALENT TO THE INSTRUCTIONS/MANUALS FOR THIS COMPLAINT EVENT. THE CONTENT ADEQUATELY PROVIDES WARNINGS AND INSTRUCTIONS FOR USE REGARDING THE REPORTED COMPLAINT EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, IFU WARNINGS AND CAUTIONS, AND PHYSICIAN TRAINING. NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION.' THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR FRAME DAMAGE WAS CONFIRMED BASED ON EVALUATION OF THE RETURNED DEVICE. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY, UNDERSIZED VESSEL) AND/OR PROCEDURAL FACTORS (HIGH PUSH FORCE) LIKELY CONTRIBUTED TO THE EVENT AS REVIEW OF PROVIDED 3MENSIO IMAGERY REVEALED PRESENCE OF CALCIFICATION, TORTUOSITY, AND UNDERSIZED VESSEL DIAMETERS IN THE ACCESS VASCULATURE. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SCRATCHES OBSERVED ON THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF CALCIFIED NODULES WITHIN THE ACCESS VESSEL. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. EXCESSIVE DEVICE MANIPULATION OR HIGH PUSH FORCE CAN LEAD TO THE VALVE STRUTS INTERACTING WITH THE SHEATH SHAFT AND RESULT IN THE STRUT DAMAGE AT THE VALVE INFLOW SIDE. THE PRODUCT RISK ASSESSMENT (PRA) IS CAPTURED IN THE TECHNICAL SUMMARY, WHICH APPLIES TO THIS COMPLAINT EVENT. THE RISKS OUTLINED IN THE PRA REMAIN THE SAME. THE PRA DOCUMENTS THE DIFFICULTY ADVANCING THE DELIVERY SYSTEM THROUGH THE SHEATH, RESULTING IN DIFFICULTY OR THE INABILITY TO INTRODUCE DELIVERY SYSTEM/LOADER AND ADVANCE THROUGH SHEATH, PROLONGING PROCEDURE, WHICH DID OCCUR DURING THIS EVENT. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. AS THE COMPLAINT DID NOT EXCEED THE PRA RE-ESCALATION THRESHOLD, NO FURTHER ACTION IS REQUIRED. A CAPA/SCAR IS NOT REQUIRED BECAUSE AN EDWARDS/SUPPLIER DEFECT WHICH COULD HAVE RESULTED IN THE COMPLAINT WAS NOT CONFIRMED. PRIOR CORRECTIVE ACTION, DETAILED IN THE TECHNICAL SUMMARY, CAPTURES PRIOR HIGH PUSH FORCE INVESTIGATION AND CORRECTIVE/PREVENTATIVE ACTION. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. NO FURTHER ESCALATION IS NEEDED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21244574&pc=DQX | 2126666-2025-00004 | 2024-12-06 05:00 | Death | LAKE REGION MEDICAL | 2025-01-27 05:00 | DQX | SAFARI2¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | K151244 |  | 1 | Event Description: EVENT DESCRIPTION: WHEN THE SAFARI WAS PRESSED AGAINST A PATIENT WITH A SMALL LEFT VENTRICLE, WIRE PERFORATION OCCURRED. PHYSICIAN COMMENTS: PATIENT OUTCOME: NO PATIENT INJURY. INTRODUCER SHEATH USED: NOT AVAILABLE AT THIS MOMENT. GUIDEWIRE USED: YES - UNKNOWN - SAFARI 2 GUIDE. CATHETER USED: NOT AVAILABLE AT THIS MOMENT. BALLOON CATHETER USED: NOT AVAILABLE AT THIS MOMENT. STENT USED: NONE. INFLATION DEVICE USED: NOT AVAILABLE AT THIS MOMENT. IMAGING CATHETER: UNKNOWN. ADDITIONAL INFORMATION 17JAN2025: QUESTION: PLEASE DESCRIBE THE PATIENT'S CURRENT CONDITION. ANSWER: DECEASED ABOUT 1 MONTH AFTER SURGERY (DUE TO ANOTHER COMPLICATION WHILE RECUPERATING THAT WAS NOT DIRECTLY RELATED TO THIS EVENT). QUESTION: PLEASE PROVIDE DETAILS OF THE CIRCUMSTANCES LEADING UP TO THE PERFORATION. ANSWER: AFTER TAVI VALVE IMPLANTATION, WHEN THE TAVI DEVICE IS REMOVED, THE SAFARI WIRE ADVANCES TOWARD THE APEX OF THE HEART, EXPOSING THE WINDING OF THE TIP OF THE SAFARI WIRE TO THE OUTSIDE OF THE HEART. QUESTION: WHAT IS THE LOCATION OF THE PERFORATION? ANSWER: SEVERAL CENTIMETERS HEADWARD FROM THE APEX OF THE HEART. QUESTION: PLEASE DESCRIBE IN DETAIL WHAT KIND OF PROCEDURE WAS PERFORMED OR PLANNED. (DRAINAGE, HEMOSTASIS OF THE PERFORATION BY THORACOTOMY, ETC.) DRAINAGE WAS PERFORMED, HEMOSTASIS OF PERFORATED AREA BY OPEN CHEST, ETC.). ANSWER: HEMOSTASIS BY OPEN CHEST. QUESTION: WHAT IS THE OPINION OF THE DOCTOR IN CHARGE REGARDING THE CAUSE OF THE PERFORATION? ANSWER: THE SAFARI WIRE WAS HELD IN PLACE TO PREVENT IT FROM BEING PULLED OUT WHEN THE TAVI DEVICE WAS REMOVED, BUT THE SAFARI WIRE WENT TOO DEEP. THE PATIENT WAS A VERY ELDERLY WOMAN WITH A SMALL BUILD AND A VERY SMALL LEFT VENTRICLE, SO THE RISK WAS PROBABLY HIGH. QUESTION: WAS THERE ANY PROBLEM WITH THE APPEARANCE OR FUNCTIONALITY OF THE PRODUCT IN QUESTION? ANSWER: NOTHING IN PARTICULAR. QUESTION: SHAPING? ANSWER: NONE. PROCEDURE NAME: TAVI BY LEFT CAROTID APPROACH. TREATMENT SITE: A FEW CENTIMETERS HEADWARD FROM THE APEX. THE DEVICE IN QUESTION WILL NOT BE RETURNED. Manufacturer Narrative: PRODUCT WAS NOT RECEIVED FOR EVALUATION; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU) WARNINGS SECTION: MONITOR WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. NEVER ADVANCE THE GUIDEWIRE AGAINST RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED IN THE VENTRICLE. THE CURVE OF THE SAFARI2 GUIDEWIRE SHOULD BE CONSTRAINED WITH IN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1. ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2. CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3. AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4. INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5. CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6. REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7. ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8. CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9. MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10. TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT CLINICAL AND/OR PROCEDURAL FACTORS HAVE CONTRIBUTED TO THE EVENT. CARDIAC PERFORATION, ADDITIONAL SURGICAL PROCEDURE, AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE SAFARI2 INSTRUCTIONS FOR USE (IFU). DESPITE ATTEMPTS TO GATHER FURTHER DETAILS, NO FURTHER INFORMATION HAS BEEN RECEIVED TO DATE. PER THE DISTRIBUTOR, NO ADDITIONAL INFORMATION IS AVAILABLE REGARDING THIS EVENT. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED. THE SECTIONS LEFT BLANK OR UNKNOWN ON THIS FORM COULD NOT BE COMPLETED DUE TO MISSING INFORMATION. DATE OF DEATH WAS ESTIMATED BASED ON THE REPORT OF THE PATIENT EXPIRED ABOUT ONE MONTH AFTER SURGERY. ALTHOUGH ANNEX G INSTRUCTIONS STATE THAT WITH STAND-ALONE DEVICES ANNEX G TERM SHOULD NOT BE USED, CODE 4755 WAS SELECTED DUE TO COMPONENT CODE (G) SHOWING UP ON THE MISSING DATA REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245277&pc=NPT | 2025587-2025-00649 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-27 05:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A TRANSCATHETER AORTIC VALVE REPLACEMENT USING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-26 IN A PATIENT WITH A MEAN ANNULAR DIAMETER OF 21.2 MM, PERIMETER OF 66.7 MM, AND A SMALL SINUS OF VALSALVA, A TRANSFEMORAL APPROACH AND 14 FR INTRODUCER SHEATH WERE USED. THE VALVE WAS IMPLANTED IN THE PATIENT. AFTER THE PROCEDURE WAS COMPLETED, AND WHILE WAITING FOR A BED, ANESTHESIA USED A DEFIBRILLATOR DEVICE ON THE PATIENT. CARDIOPULMONARY RESUSCITATION WAS INITIATED, BUT THE PATIENT SUBSEQUENTLY PASSED AWAY. CONTRIBUTING FACTORS INCLUDED THE PATIENT'S LARGE BODY MASS INDEX AND ANATOMY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN IT WAS UNKNOWN IF THE VALVE CAUSED OR CONTRIBUTE TO THE EVENT BUT THE PATIENT'S LARGE BODY MASS INDEX AND OTHER COMORBIDITIES MAY HAVE BEEN A CONTRIBUTING FACTOR.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRANSFEMORAL APPROACH AND 14 FRENCH INTRODUCER SHEATH WERE NOT USED AS AN ALTERNATIVE ACCESS WAS USED. Manufacturer Narrative: IMAGE REVIEW: ONE MEDIA FILE WAS PROVIDED FOR REVIEW. THE MEDIA FILE PROVIDED SHOWS EVIDENCE OF AN ALTERNATIVE ACCESS APPROACH, AND A VALVE DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE. IMAGES OF THE FULLY RELEASED VALVE WERE NOT PROVIDED THEREFORE IT IS NOT POSSIBLE TO VALIDATE CAUSE OF DEATH. THE PATIENT¿S EXECUTIVE SUMMARY WAS PROVIDED FOR ANATOMICAL REVIEW. THE PATIENT¿S ANNULUS PERIMETER MEASURED 66.7MM WITH A PERIMETER DERIVED DIAMETER OF 21.2MM SUGGESTING A 26MM EVOLUT. SINUSES OF VALSALVA HEIGHTS MEASURED WITHIN RECOMMENDED RANGES; HOWEVER, SINUSES OF VALSALVA DIAMETERS MEASURED LESS THAN THE RECOMMENDED 27MM FOR THE 26MM EVOLUT. UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}; PRODUCT ID {L-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245716&pc=NPT | 2025587-2025-00651 | 2024-01-21 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-27 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 MONTHS AND 24 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT WENT INTO CARDIOPULMONARY ARREST AND WAS TRANSPORTED BY AMBULANCE TO THE HOSPITAL. THE PATIENT DIED WITHOUT RETURN OF SPONTANEOUS CIRCULATION. FOLLOWING ARRIVAL AT THE FACILITY, COMPUTED TOMOGRAPHY SCAN REVEALED NO APPARENT CAUSE OF DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245737&pc=DSQ | 2916596-2025-00611 | 2022-12-01 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Shock; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿PERFORMANCE OF RISK SCORES IN PREDICTING MAJOR BLEEDING IN LEFT VENTRICULAR ASSIST DEVICE RECIPIENTS: A COMPARATIVE EXTERNAL VALIDATION¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH BLEEDING, THROMBOSIS, AND DEATH. THIS SINGLE-CENTER RETROSPECTIVE STUDY EVALUATED 10 PATIENTS WHO WERE IMPLANTED WITH HM3 BETWEEN NOV2010 AND DEC2022. OF NOTE, 94 PATIENTS WITH A HEARTWARE WERE ALSO INCLUDED IN THE ANALYSIS RESULTS, PATIENT BACKGROUND INFORMATION, AND OUTCOMES. DATA WAS COLLECTED FROM ELECTRONIC HEALTH RECORDS. THE STUDY AIMED TO COMPARE THE PREDICTIVE ACCURACY OF CURRENT RISK SCORES FOR BLEEDING EVENTS IN LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS. THE MEDIAN FOLLOW-UP AMONG THE 104 PATIENTS WAS 1916 DAYS. THE AVERAGE AGE AT IMPLANTATION WAS 64 YEARS OLD. THERE WERE 21 FEMALES IN THE STUDY AND 83 MALE PATIENTS. 24 PATIENTS EXPERIENCED A GASTROINTESTINAL (GI) BLEED. 15 PATIENTS EXPERIENCED A MEDIASTINAL BLEED, 11 PATIENTS EXPERIENCED AN INTRACRANIAL BLEED, 5 EXPERIENCED A BLEED IN THE PLEURAL SPACE, 5 EXPERIENCED A NASAL BLEED, 5 EXPERIENCED A SUBCUTANEOUS OR INTRAMUSCULAR BLEED, 3 EXPERIENCED A BLEED AT THE DRIVELINE EXIT SITE, 2 EXPERIENCED A RETROPERITONEAL BLEED, 1 EXPERIENCED AN INTRAABDOMINAL BLEED, 1 EXPERIENCED A GENITOURINARY BLEED, 1 EXPERIENCED AN INTRAORAL BLEED. 1 PATIENT EXPERIENCED A HEMORRHAGIC SHOCK LEADING TO DEATH. IN TOTAL 23 PATIENTS PASSED AWAY. 11 PATIENTS PASSED AWAY AS A RESULT OF BLEEDING EVENTS, AND 1 OF THOSE EVENTS WAS POST THROMBOLYSIS FOR PUMP THROMBOSIS. OUTCOMES WERE NOT BROKEN DOWN BY LVAD TYPE. THE STUDY FOUND THAT THE RISK SCORES CURRENTLY USED TO ASSESS LVAD RECIPIENTS, IN THEIR EXISTING FORMS, FAILED TO ACCURATELY PREDICT BLEEDING EVENTS IN HIGH-RISK POPULATIONS. THE STUDY HIGHLIGHTED THE NEED FOR A MORE ACCURATE RISK ASSESSMENT TO BE ABLE TO RELIABLY IDENTIFY LVAD PATIENTS AT HIGH RISK OF BLEEDING EVENTS.Event Description: ONE PATIENT ALSO EXPERIENCED AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR POCKET BLEED. ADDITIONALLY, 17 PATIENTS DIED OR EXPERIENCED A MAJOR BLEEDING EVENT BEFORE PHENPROCOUMON INITIATION. OUTCOMES WERE NOT BROKEN DOWN BY LVAD TYPE. THESE PATIENTS RECEIVED UNFRACTIONATED HEPARIN, EITHER AS MONOTHERAPY OR WITH ANTIPLATELET THERAPY. IT WAS NOTED THAT 4 OF THE MAJOR BLEEDING EVENTS OCCURRED AFTER INTENSIFIED ANTICOAGULANT THERAPY. FINALLY, 4 OF THE MAJOR BLEEDING EVENTS WERE RELATED TO ALTEPLASE ADMINISTRATION IN THE CONTEXT OF PUMP THROMBOSIS. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN, DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01DEC2022 AS DATA WAS COLLECTED BETWEEN NOV2010 AND DEC2022. AUTHOR INFORMATION: VAN DER HORST, S. F. B., DE JONG, Y., VAN REIN, N., JUKEMA, J. W., PALMEN, M., JANSSEN, E., BONNEVILLE, E. F., KLOK, F. A., HUISMAN, M. V., TOPS, L. F., & DEN EXTER, P. L. (2024). PERFORMANCE OF RISK SCORES IN PREDICTING MAJOR BLEEDING IN LEFT VENTRICULAR ASSIST DEVICE RECIPIENTS: A COMPARATIVE EXTERNAL VALIDATION. RESEARCH AND PRACTICE IN THROMBOSIS AND HAEMOSTASIS, 8(4), 102437. HTTPS://DOI.ORG/10.1016/J.RPTH.2024.102437 LEIDEN UNIVERSITY MEDICAL CENTER, LEIDEN, THE NETHERLANDS. D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND THE HEARTMATE 3 PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, DEVICE THROMBOSIS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245886&pc=DSQ | 2916596-2025-00538 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Aortic Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH LOW FLOWS. THE EVENT LOG FILE CAPTURED LOW FLOW ALARMS ON (B)(6) 2025 THAT BECAME MORE FREQUENT ON (B)(6) 2025 AT 1257 TO 16XX. THE ESTIMATED FLOWS WERE AVERAGING FROM 0.1 TO 2.4 LPM AND THE PULSATILITY INDEX (PI) WAS AVERAGING FROM 1.9 TO 6.6 DURING THIS EVENTS. THERE WERE NO UNUSUAL ROTOR FAULTS OR ANY OTHER ABNORMAL PUMP FAULTS THAT WERE SEEN DURING THESE EVENTS. IT WAS LATER REPORTED THAT THE PATIENT PASSED AWAY. IT WAS NOTED THAT THE PUMP WAS NOT FUNCTIONING AS INTENDED. HOWEVER, THE EXACT CAUSE WAS UNCLEAR. THE OUTCOME HAPPENED VERY QUICKLY AND IT WAS UNEXPECTED. THE PATIENT WAS TALKING ONE MINUTE STATING TO HAVE COLD LIKE SYMPTOMS AND STARTED TO HAVE THE LOW FLOW ALARMS. IN AN HOUR, THE PATIENT WAS IN THE EMERGENCY ROOM WITH THE TEAM CALLING TO REPORT THE PATIENT LOOKED VERY BAD AND NEEDED TO BE FLOWN OVER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT ARRIVED TO THE EMERGENCY ROOM BY FLIGHT CARE AND WAS OBTUNDED ON ARRIVAL LEADING TO IMMEDIATE INTUBATION. THE BLOOD PRESSURE MEASURED WITH DOPPLER WAS IN THE 40S-60S, IMPROVED SLIGHTLY WITH LEVOPHED (NOREPINEPHRINE BITARTRATE) UP TO 2 MCG/KG/MIN AND ADDITION OF DOBUTAMINE. STAT ECHOCARDIOGRAM WAS OBTAINED DEMONSTRATING SEVERELY ENLARGED LEFT AND RIGHT VENTRICLE WITH THE AORTIC VALVE OPENING WITH EVERY BEAT. FLOWS COULD NOT BE ASSESSED ACROSS INFLOW. HOWEVER, THE LEFT VENTRICULAR ASSIST DEVICE WITH OVER 3 HOURS OF EXTREMELY LOW FLOW FROM 0.5-0.7, INITIALLY 1.0. THE PATIENT WAS BLUE ON EXAM, VERY COLD, AND POOR PULSES. THE LABS WERE EVALUATED AND REVIEWED AND BICARB WAS ALSO GIVEN. A DECISION WAS FINALLY MADE TO MAKE THE PATIENT COMFORTABLE AND TURN OF THE LVAD. IT WAS UNKNOWN IF THE DEATH WAS CONSIDERED TO BE DEVICE RELATED. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. SECTION B3: DATE OF DEATH CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE AND MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), CONFIRMED EXTRINSIC OBSTRUCTION OF THE OUTFLOW GRAFT; HOWEVER, A SPECIFIC CAUSE FOR THE OBSERVED OUTFLOW GRAFT OBSTRUCTION AND THE DURATION FOR WHICH IT WAS PRESENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE INTACT. THE PUMP AND OUTFLOW GRAFT WERE ENCAPSULATED IN TISSUE. THE MODULAR CABLE WAS RETURNED, PROPERLY CONNECTED AT THE INLINE CONNECTOR. THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE BEND RELIEF PROPERLY ENGAGED AT THE GRAFT HARDWARE. AN OUTFLOW GRAFT CLIP WAS ALSO RETURNED, PROPERLY ATTACHED. THE APICAL CUFF WAS RETURNED, PROPERLY ENGAGED WITH THE CUFF LOCK. UPON REMOVAL OF THE BEND RELIEF, A CREASE IN THE GRAFT MATERIAL WAS OBSERVED. AN ACCUMULATION OF TISSUE WAS OBSERVED IN THE SPACE BETWEEN THE GRAFT AND THE BEND RELIEF, FILLING INTO THE CREASED AREA, SUGGESTING THAT THE DEFORMATION WAS PRESENT FOR AN UNDETERMINED AMOUNT OF TIME WHILE (B)(6) WAS SUPPORTING THE PATIENT. THESE FINDINGS ARE CONSISTENT WITH EXTRINSIC OUTFLOW GRAFT OBSTRUCTION DURING SUPPORT AND WOULD HAVE CONTRIBUTED TO THE REPORTED DECREASE IN FLOW AND PATIENT SYMPTOMS. VISUAL EXAMINATION OF THE PUMP¿S REMAINING BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF ADHERED OR DEVELOPED DEPOSITIONS OR THROMBUS FORMATIONS. THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. SECTION 5, UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5 "SURGICAL PROCEDURES" CAUTIONS THE USER THAT SHARP BENDS, TWISTS, OR KINKS IN THE DRIVELINE MAY MAKE IT MORE SUSCEPTIBLE TO WEAR AND FATIGUE OVER TIME. SECTION 6 "PATIENT CARE AND MANAGEMENT" CAUTIONS THE USER TO AVOID PULLING ON OR MOVING THE DRIVELINE AND TO NOT TWIST, KINK, OR SHARPLY BEND ANY CABLES, AS IT MAY CAUSE DAMAGE TO THE WIRES INSIDE THAT MAY NOT BE OUTWARDLY VISIBLE AND THAT COULD CAUSE THE PUMP TO STOP. IF KINKING, TWISTING, OR BENDING DOES OCCUR, CAREFULLY UNRAVEL AND STRAIGHTEN. SECTION 6 INSTRUCTS THE USER TO CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE. SECTION 8 "EQUIPMENT STORAGE AND CARE" STATES, "AS NEEDED, CLEAN EXTERIOR SURFACES OF THE DRIVELINE CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, USE WARM WATER AND MILD DISH SOAP." SECTION 5, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH.". THE HEARTMATE 3 LVAS PATIENT HANDBOOK IS ALSO CURRENTLY AVAILABLE. SECTION 4 "LIVING WITH THE HEARTMATE 3" (UNDER "CARING FOR THE DRIVELINE") INSTRUCTS THE USER: "CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE (CUTS, HOLES, TEARS). CALL YOUR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED)" AND "KEEP YOUR DRIVELINE CLEAN. WIPE OFF ANY DIRT OR GRIME. IF THE DRIVELINE GETS DIRTY, USE A TOWEL WITH MILD DISH SOAP AND WARM WATER TO GENTLY CLEAN IT. NEVER SUBMERGE THE DRIVELINE IN WATER OR LIQUID." IN ADDITION, SECTION 5 "ALARMS AND TROUBLESHOOTING¿ CAUTIONS THE USER NOT TO TWIST, KINK, OR SHARPLY BEND THE DRIVELINE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245958&pc=DSQ | 2916596-2025-00408 | 2025-01-11 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Renal Failure; Hemorrhagic Stroke; Thrombosis/Thrombus; Drug Resistant Bacterial Infection | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO WITHDRAWAL OF CARE DUE TO A STROKE. THE PATIENT HAD BEEN OFF ALL ANTICOAGULATION FOR SIGNIFICANT PERIOD OF TIME PRECEDING HOSPITALIZATION DUE TO SIGNIFICANT BLEEDING HISTORY AND DESPITE THIS, REQUIRED NUMEROUS TRANSFUSIONS OFTEN MORE THAN WEEKLY WITH DIALYSIS SESSIONS TO MAINTAIN HEMOGLOBIN LEVELS GREATER THAN 8. A COMPUTED TOMOGRAPHY (CT) SCAN IDENTIFIED NUMEROUS SMALL STROKES, THOUGH WHETHER THEY WERE ISCHEMIC THAT TRANSFORMED INTO HEMORRHAGIC IS UNKNOWN. THE PATIENT BECAME NEUROLOGICALLY AND COGNITIVELY LESS INTACT AS A RESULT AND WAS TRANSITIONED TO COMFORT MEASURES. IT WAS NOTED BY THE SITE THAT THEY THOUGHT IT QUITE LIKELY THAT BEING OFF ANTICOAGULATION AND PATIENT'S UNDERLYING INFECTION HISTORY CONTRIBUTED TO A PROTHROMBOTIC STATE RESULTING IN CLOT THAT COULD HAVE SHOWERED EMBOLI TO THE BRAIN FROM THE LEFT VENTRICULAR ASSIST DEVICE. Manufacturer Narrative: D4 - EXPIRATION DATE: CORRECTION. H5 - DEVICE MANUFACTURE DATE: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE ACCOUNT COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, UNDER "ANTICOAGULATION", CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245965&pc=DSQ | 2916596-2025-00346 | 2025-01-09 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OFFICIAL CAUSE OF DEATH WAS NOT CONFIRMED, ALTHOUGH THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED BY THE HOSPITAL CLINICIANS. THE DEVICE OPERATED AS EXPECTED AND DESIGNED WITHOUT ANY ALARMS, NO TECHNICAL ISSUES WERE HEARD FROM THE CLINICIANS. THE PATIENT HAD GOOD DRUG COMPLIANCE AND WAS GOOD AT HANDLING DEVICES. THE CLINICIANS DID NOT SPECIFY THE ROOT CAUSE OF DEATH, BUT THE PATIENT PASSED AWAY AT HOME AND THE DEVICE DID NOT SHOW ANY ISSUES.Event Description: IT WAS REPORTED ON (B)(6)2025 THAT THE PATIENT HAD PASSED AWAY AND LATER NOTED THAT THE DEATH HAD OCCURRED ON (B)(6)2025. Manufacturer Narrative: CORRECTION B5: EVENT SUMMARY DETAILS WERE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21245983&pc=DSQ | 2916596-2025-00292 | 2025-01-09 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO OUTFLOW THROMBOSIS. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED AND THERE WERE NO RECENT CHANGES TO PUMP PARAMETERS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED TO CONFIRM THE OUTFLOW THROMBOSIS, BUT THE IMAGES WERE NOT AVAILABLE FOR EVALUATION. THE PATIENT HAD HEART FAILURE SYMPTOMS INCLUDING DYSPNEA AND SUDDEN DECREASE IN FLOW. A THROMBOLYSIS WAS PERFORMED BEFORE THE PATIENT PASSED AWAY. THE DEATH WAS NOT CONSIDERED TO BE DIRECTLY RELATED TO THE DEVICE, BUT THE OUTFLOW THROMBOSIS COULD NOT BE EXPLAINED. THE DEVICE REPORTEDLY OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORTED OUTFLOW THROMBOSIS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION AS THE DEVICE IS NOT AVAILABLE FOR INVESTIGATION AND NO IMAGES WERE SUBMITTED FOR REVIEW. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW EVENTS WHICH THE ACCOUNT ATTRIBUTED TO THE SUSPECTED OUTFLOW THROMBOSIS. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER LOG FILE CAPTURED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. CONTINUOUS LOW FLOW ALARMS WERE CAPTURED ON (B)(6) 2025. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS AND DEATH WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. ADDITIONALLY, IT ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU ALSO DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, STATES TO USE THE PUMP STOP BUTTON TO TURN OFF THE PUMP. TO USE THIS FEATURE, PRESS AND HOLD THE PUMP STOP BUTTON WHILE THE PUMP STOP COUNTDOWN COUNTS DOWN FROM 10 TO 1, THEN A STOPPING PUMP MESSAGE WILL BE DISPLAYED. THIS SECTION FURTHER STATES THAT THE PUMP STOPS WITHIN THE FIRST FEW SECONDS OF HOLDING DOWN THE PUMP STOP BUTTON. HOWEVER, IF THE BUTTON IS RELEASED BEFORE THE COUNTDOWN IS COMPLETE, THE PUMP RESUMES AT THE PREVIOUSLY SET MODE AND SPEED. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS THROMBOEMBOLISM AS A LATE POSTIMPLANT COMPLICATION. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDE INFORMATION ON SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21246337&pc=NGV | 2124215-2025-03203 | 2025-01-06 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-27 05:00 | NGV | WATCHMAN FLX | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: CHAMPION-AF STUDY: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED WITH COMPLETE DEVICE SEAL NOTED. THE PATIENT WAS DISCHARGED ON APIXABAN AND CLOPIDOGREL. 312 DAYS POST INDEX PROCEDURE, THE PATIENT DIED. AT THE TIME OF THE EVENT, THE PATIENT WAS ON ASPIRIN AND CLOPIDOGREL. THE CAUSE OF DEATH REMAINS UNKNOWN AS THE NOTICE OF THE DEATH WAS FOUND ON THE OBITUARY SITE, AND NO FURTHER DETAILS WERE MADE AVAILABLE AS NO CALLS TO FAMILY WILL BE MADE PER THE FACILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21246375&pc=DSY | 2017233-2025-05760 | 2024-12-29 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-01-27 05:00 | DSY | GORE-TEX® STRETCH VASCULAR GRAFT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K880167 |  | 1 | Event Description: REPORTEDLY: UNEXPECTED DEATH, POSSIBLY RELATED TO AV (ARTERIOVENOUS) GRAFT HEMORRHAGE AT ACCESS SITE. PATIENT WAS FOUND UNRESPONSIVE IN ROOM, HEMORRHAGE EVIDENT. DNR (DO-NOT-RESUSCITATE ORDER). 2200-A:-OTHER USED TO CAPTURE HEMORRHAGE. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. MANUFACTURING RECORDS WERE REVIEWED, AND THE DEVICE MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: UPDATED H6 HEALTH EFFECT IMPACT CODE TO, F02 - DEATH. ADDED RELATED REPORT NUMBER TO H10/11.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21247626&pc=DSQ | 2916596-2025-00626 | 2024-12-19 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Tachycardia; Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THEIR SYSTEM CONTROLLER WAS RETURNED TO CLINIC. LOG FILES WERE SENT FOR EVALUATION. THE LOG FILE CONTAINED DATA FROM (B)(6) 2024 18:30 - (B)(6) 2024 0:15 BASED ON THE DATE/TIME STAMP DATA RECORDED. THE LOG FILE INDICATED THAT THE PATIENT HAD NOT CONNECTED TO POWER MODULE/MPU POWER DURING THIS HISTORY. THIS INDICATES THE PATIENT HAD BEEN SLEEPING ON BATTERY POWER. THERE WERE 109 LOW BATTERY ADVISORY EVENTS, AND 6 LOW BATTERY HAZARD EVENTS RECORDED OVER THE ENTIRE RECORDED HISTORY. THE FOLLOWING ARE THE SEQUENCE OF EVENTS FROM (B)(6) 2024 THAT WE ARE ABLE TO DETERMINE BASED ON THE LOG FILE DATA. ON (B)(6) 2024 AT 8:56 THE PATIENT PERFORMED BATTERY EXCHANGE TO FULLY CHARGED SET OF BATTERIES. LATER THAT DAY AT 21:21 A LOW POWER ADVISORY ALARM WAS ACTIVATED, THERE WAS ABOUT 12 HOURS AND 24 MINUTES OF RUNTIME ON BATTERY POWER. AT 22:39 THAT A LOW POWER HAZARD ALARM WAS ACTIVATED AND THERE WAS ABOUT A 1 HOUR, 18 MINUTES OF RUNTIME IN LOW POWER ADVISORY ALARM STATE. ON (B)(6) 2024 AT 0:15 A NO EXTERNAL POWER ALARM ACTIVATED, AND THE SYSTEM CONTINUED OPERATING ON THE EMERGENCY BACKUP BATTERY. THE HEARTMATE II PUMP REMAINED OPERATING ON EMERGENCY BACKUP BATTERY POWER FOR AN UNKNOWN AMOUNT OF TIME. ONCE FULLY DEPLETED, THE DATE/TIME STAMP WAS NO LONGER MAINTAINED AND WAS RESET TO DEFAULT. PUMP OFF TIME WAS UNKNOWN AND THE DATE/TIME BECAME UNKNOWN. THE WHITE POWER LEAD OF THE SYSTEM CONTROLLER WAS CONNECTED TO POWER MODULE POWER, LIKELY FOR LOG FILE DOWNLOAD. DATE/TIME CAPTURED AS 1JAN2001 1:00.Event Description: IT WAS LATER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN AS WELL AS IF DEATH WAS CONSIDERED DEVICE RELATED AND OPERATING AS INTENDED AT THE TIME OF DEATH.Event Description: IT WAS LATER REPORTED THAT ON (B)(6) 2025 THE ARRHYTHMIA DEPARTMENT RECEIVED A CONCERNING TRANSMISSION CLOSE TO 1AM. CLINIC TRIED REACHING OUT TO PATIENT AND PATIENTS' FAMILY BUT WAS UNSUCCESSFUL. EMERGENCY SERVICES (911) WAS THEN CALLED AND ASKED TO PERFORM A WELLNESS CHECK AT PATIENT'S ADDRESS. AT 12:22 911 CALLED BACK TO REPORTED THAT PATIENT WAS DECEASED. THE DEPUTY ON THE SCENE SAID THAT " THERE WAS NO NOISE COMING FROM THE VENTRICULAR ASSIST DEVICE, THE PATIENT WAS SITTING UP ON THE SOFA WITH A BAG IN THEIR HAND" AND THAT "THE BAG HAD SOMETHING IN IT CONNECTED TO THEIR STOMACH." THE DEPUTY PROCEEDED TO SAY THAT THEY WERE NOT ALLOWED TO TOUCH THE BODY UNTIL THE CORONER ARRIVED ON THE SCENE AND PROCESSED THE SCENE. AT 12:48 THE PATIENTS RELATIVE CALLED BACK TO REPORT THAT THE PATIENT PASSED AWAY AND THAT THEY WERE IN ROUTE TO PATIENT'S HOME. THE PATIENT HAD AN ICD IMPLANTED FOR ARRHYTHMIA PRIOR TO IMPLANT. THE PATIENT HAD MONOMORPHIC VENTRICULAR TACHYCARDIA THAT WAS DID NOT EXIST PRIOR TO IMPLANT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILE CONFIRMED A PUMP STOP DUE TO BATTERY DEPLETION; HOWEVER, A SPECIFIC CAUSE FOR THIS EVENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A REVIEW OF THE SUBMITTED LOG FILE FOUND THAT THE PUMP STOPPED ON (B)(6) 2024 DUE TO EXTERNAL 14-VOLT BATTERY DEPLETION FOLLOWED BY CONTROLLER BACKUP BATTERY DEPLETION. THE SYSTEM OTHERWISE APPEARED TO BE OPERATING AS INTENDED AT THE SET SPEED, AND THERE WERE NO OTHER NOTABLE ALARMS ACTIVE IN THE LOG FILE. IT WAS REPORTED THAT HEARTMATE II LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS, INCLUDING DEATH AND CARDIAC ARRHYTHMIA. SECTION 7 ¿ALARMS AND TROUBLESHOOTING¿ CONTAINS INFORMATION REGARDING ALL VISUAL/AUDIO ALARMS, AND WHAT ACTION(S) SHOULD BE PERFORMED WHEN THEY OCCUR. THE HEARTMATE II PATIENT HANDBOOK, REV. C, IS CURRENTLY AVAILABLE. SECTION 5 ¿ALARMS AND TROUBLESHOOTING¿ CONTAINS INFORMATION REGARDING SYSTEM CONTROLLER ALARMS AND THE PROPER ACTIONS ASSOCIATED WITH THEM. SECTION 8 "HANDLING EMERGENCIES" LISTS EXAMPLES OF EMERGENCIES AND WHAT ACTIONS TO TAKE, INCLUDING WHEN THE PUMP HAS STOPPED. THE PATIENT HANDBOOK INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE PATIENT THINKS THAT, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5 NARRATIVE INFORMATION NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21248170&pc=OZD | 1220648-2025-25829 | 2025-01-10 05:00 | Death | ABIOMED, INC. | 2025-01-27 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, BLEEDING WAS NOTED AT THE IMPELLA INSERTION SITE. BLEEDING SUCCESSFULLY MANAGED BY MANUAL PRESSURE AND APPLYING A FEM-STOP. HEPARIN WAS OFF DUE TO BLEEDING AT IMPELLA INSERTION SITE. LARGE HEMATOMA REMAINED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE PRODUCT WAS NOT RETURNED. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. B.2 REVISED AS REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25829. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25829 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21249104&pc=DSQ | 2916596-2025-00628 | 2024-12-31 05:00 | Death | THORATEC CORPORATION | 2025-01-27 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THE CAUSE WAS NOT PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND WAS DUE TO PRE-EXISTING MEDICAL CONDITIONS. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION WERE SENT TO THE CUSTOMER; HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED AT THIS TIME. IT WAS REPORTED THAT NO AUTOPSY WAS PERFORMED, AND HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU), REVISION D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21237074&pc=BZD | 2518422-2025-008092 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-26 05:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM-RECRT | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LUNG DISEASE DEATH AND OTHER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H: TYPE OF REPORT WAS INCORRECT IN PREVIOUS REPORT, NOW CORRECTED IN THIS REPORT. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX D : PRODUCT BRAND NAME, MODEL NUMBER, CATALOG ITEM IDENTIFIER, SERIAL NUMBER MISSED TO CAPTURE IN PREVIOUS REPORT, NOW CORRECTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21237234&pc=BZD | 2518422-2025-008127 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-26 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21238167&pc=DSQ | 2916596-2025-00417 | 2022-09-01 04:00 | Death | THORATEC CORPORATION | 2025-01-26 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿LONG-TERM IMPLANTABLE VENTRICULAR ASSIST DEVICE SUPPORT IN CHILDREN¿ THAT LONG-TERM OUTPATIENT MANAGEMENT USING IMPLANTABLE CONTINUOUS-FLOW VENTRICULAR ASSIST DEVICES (IC-VADS) IS FEASIBLE IN PEDIATRIC PATIENT WITH FAVORABLE OUTCOMES. THIS SINGLE-CENTER RETROSPECTIVE STUDY REVIEWED 100 CONSECUTIVE PATIENTS UNDERGOING IC-VAD SUPPORT BETWEEN MAY2008 AND SEP2022, 16 OF WHICH WERE IMPLANTED WITH A HEARTMATE 3. OF NOTE, 67 PATIENTS WITH A HEARTWARE AND 17 PATIENTS WITH A HEARTMATE II WERE ALSO INCLUDED IN THE ANALYSIS RESULTS. THE MEDIAN AGE, WEIGHT, AND BODY SURFACE AREA AT IMPLANTATION FOR ALL IC-VADS WERE 14.1 YEARS, 54.8 KG, AND 1.6 M², RESPECTIVELY. CARDIOMYOPATHY (58%) WAS THE MOST COMMON ETIOLOGY, FOLLOWED BY CONGENITAL HEART DISEASE (37%, INCLUDING 13 SINGLE VENTRICLE). AT 6 MONTHS OF IC-VAD SUPPORT, 94 ENCOUNTERS ACHIEVED POSITIVE OUTCOMES: ONGOING SUPPORT (59), TRANSPLANT (33), AND CARDIAC RECOVERY (2). EIGHTY-TWO ENCOUNTERS (82%) RESULTED IN HOME DISCHARGE WITH ONGOING VAD SUPPORT, INCLUDING 38 (46%, OUT OF 82) REQUIRING READMISSION AND 7 (9%, OUT OF 82) RESULTING IN DEATH. THERE WAS A CLINICALLY SIGNIFICANT DECREASE IN MORBIDITY RATES BEFORE VERSUS AFTER HOME DISCHARGE: BLEEDING (1.55 VS 0.06), INFECTION (0.84 VS 0.37), AND STROKE (0.84 VS 0.15 EVENT PER PATIENT-YEAR). OVERALL, 86 ENCOUNTERS (86%) REACHED POSITIVE END POINTS AT THE LATEST FOLLOW-UP (64 TRANSPLANT, 15 ONGOING SUPPORT, AND 7 RECOVERY). INFECTION (29%; 4 OF 14) WAS THE MOST COMMON CAUSE OF NEGATIVE OUTCOMES, FOLLOWED BY CEREBROVASCULAR ACCIDENT (21%; 3), AND UNRESOLVED FRAILTY (21%; 3). OTHER OUTCOMES SUCH AS RIGHT VENTRICULAR FAILURE, CARDIAC ARREST, NEUROLOGICAL DYSFUNCTION, PUMP THROMBOSIS, HEMOLYSIS, DEVICE MALFUNCTION, RESPIRATORY FAILURE, AND RENAL FAILURE WERE ALSO REPORTED. THE ESTIMATED OVERALL SURVIVAL AT 1, 2, AND 5 YEARS WAS 90%, 86%, AND 77%, RESPECTIVELY. OUTCOMES WERE NOT BROKEN DOWN BY IC-VAD TYPE.Event Description: OTHER OUTCOMES SUCH AS RIGHT VENTRICULAR FAILURE (12), CARDIAC ARREST (1), NEUROLOGICAL DYSFUNCTION (29), PUMP THROMBOSIS (16), HEMOLYSIS (4), DEVICE MALFUNCTION (4), RESPIRATORY FAILURE (12), AND RENAL FAILURE (3) WERE ALSO REPORTED. OUTCOMES WERE NOT BROKEN DOWN BY IC-VAD TYPE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. CONGENITAL HEART SURGERY, TEXAS CHILDREN'S HOSPITAL, BAYLOR COLLEGE OF MEDICINE, HOUSTON, TEX. CHO, J., TUNUGUNTLA, H. P., TUME, S. C., SPINNER, J. A., BOCCHINI, C. E., TERUYA, J., HEINLE, J. S., HICKEY, E. J., & ADACHI, I. (2023). LONG-TERM IMPLANTABLE VENTRICULAR ASSIST DEVICE SUPPORT IN CHILDREN. JOURNAL OF THORACIC AND CARDIOVASCULAR SURGERY, 167(4), 1417-1426.E1. HTTPS://DOI.ORG/10.1016/J.JTCVS.2023.10.048. Manufacturer Narrative: SECTION E: REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION B, IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236244&pc=MHX | 1218950-2025-000058 | 2024-10-24 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-01-25 05:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Cardiac Arrest | K183387 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT FELL OFF THE STRETCHER IN THE EMERGENCY DEPARTMENT AND THERE WAS NO ALARM AT THE CENTRAL STATION. IT REMAINS UNKNOWN WHAT TYPE OF ALARM WAS EXPECTED. THE MONITORS WERE TESTED WITH VARIOUS SETTINGS, REVEALING THE MONITORS WERE ALARMING AS DESIGNED FOR BOTH THE BEDSIDE MONITOR AND CENTRAL STATION. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THE PATIENT PASSED AWAY. AT THE TIME OF THE EVENT, THE PATIENT'S ECG WAS BEING MONITORED. THE PATIENT WENT INTO CARDIAC ARREST, FELL OFF THE STRETCHER, AND THE LEADS WERE DISCONNECTED. THE ECG LEADS OFF ALARM WAS GENERATED AT THE CENTRAL STATION, WHICH WAS ACKNOWLEDGED BY THE MONITOR TECH MULTIPLE TIMES. TWENTY-SIX MINUTES LATER THE NURSE WENT TO THE ROOM AND DISCOVERED THE PATIENT. SECURITY FOOTAGE WAS REVIEWED AND MULTIPLE MEETINGS WITH HOSPITAL STAFF DETERMINED THE MONITOR TECH LEFT THE MONITOR UNATTENDED AND WAS DEALING WITH ALARM FATIGUE. IN ADDITION, IT WAS DETERMINED THERE WAS NO ISSUE WITH ANY PHILIPS DEVICE. THE EMERGENCY DEPARTMENT MANAGER INDICATED THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT DEATH. Manufacturer Narrative: THE RSE TALKED TO THE CUSTOMER WHO STATED THE HOSPITAL REVIEWED SECURITY CAMERA FOOTAGE AND DETERMINED THAT IT WAS A STAFF ISSUE AND PHILIPS EQUIPMENT WAS NOT THE ISSUE. THE SECURITY FOOTAGE WAS REVIEWED AND MULTIPLE MEETINGS WITH THE HOSPITAL STAFF AND IT WAS DETERMINED THE MONITOR TECH LEFT THE MONITOR UNATTENDED AND WAS DEALING WITH ALARM FATIGUE. THE ECG LEADS OFF ALARM WAS GENERATED AT THE CENTRAL STATION, WHICH WAS ACKNOWLEDGED BY THE MONITOR TECH MULTIPLE TIMES. TWENTY-SIX MINUTES LATER THE NURSE WENT TO THE ROOM AND DISCOVERED THE PATIENT. BIOMED TESTED THE X3 AND STATED THAT ALL ALARMS WORKED AS INTENDED. THE HOSPITAL DECIDED TO PUT THE EQUIPMENT BACK IN USE. A CLINICAL SPECIALIST REVIEWED THE LOGS AND DETERMINED THAT THE LOGS DID NOT SHOW ANY ISSUE WITH PHILIPS EQUIPMENT. THE PHILIPS FIELD SUPPORT ENGINEER (FSE) PULLED THE CONFIGURATION FILES AND LOGS AND UPLOADED THEM. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) EVALUATED THE DEVICE LOGS AND CONFIRMED THAT THE DEVICE WAS WORKING AS DESIGNED. THE CUSTOMER WAS UNABLE TO PROVIDE A MORE SPECIFIC TIMESTAMP OF THE EVENT WHICH CAN BE MATCHED TO THE AUDIT LOGS. ACCORDING TO THE INFORMATION PROVIDED, THE PIC IX SYSTEM OPERATED AS DESIGNED AND GENERATED MULTIPLE ALARMS AND INOPS (INOPERATIVE) MESSAGES AROUND THE TIME OF THE EVENT. ADDITIONALLY, NO ECG OFF ALARM WILL BE GENERATED DUE TO ECG LEAD FALLBACK OPERATING MODE. IN CASE ECG LEAD FALLBACK MODE IS CONFIGURED ON, ANOTHER AVAILABLE LEAD BECOMES THE PRIMARY LEAD IF LEADS OFF INOP ERROR MESSAGES ARE GENERATED FOR MORE THAN 10 SECONDS FOR THE LEAD WHICH WAS THE ORIGINAL PRIMARY LEAD. WHEN THE LEADS OFF CONDITION IS CORRECTED, THE LEADS ARE AUTOMATICALLY SWITCHED BACK. THE EMERGENCY DEPARTMENT MANAGER OF THE HOSPITAL INDICATED THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT DEATH. BASED ON THIS INFORMATION, A DEVICE MALFUNCTION DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT; HOWEVER, ALARM FATIGUE WITH SUBSEQUENT USE ERROR MAY HAVE BEEN A FACTOR. THE REPORTED PROBLEM COULD NOT BE CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236355&pc=MVK | 3008642652-2025-00895 | 2024-12-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 20:26:45 ON (B)(6) 2024. AT 01:43:33 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 01:44:31, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 07:51:12 ON 12/17/2024. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236356&pc=MVK | 3008642652-2025-00896 | 2024-12-11 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND TWELVE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 06:36:58 ON (B)(6) 2024. AT 06:31:58 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 100 BPM WITH CPR/ELECTRODE LEAD FALL OFF. AT 06:32:35, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 100 BPM WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:33:07, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:33:34, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:35:51, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/ELECTRODE LEAD FALL OFF. AT 06:36:20, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:37:54, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:39:35, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/ELECTRODE LEAD FALL OFF. AT 06:40:41, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:41:56, THE PATIENT RECEIVED THE EIGHTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 06:43:15, THE PATIENT RECEIVED THE NINTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM WITH CPR/ELECTRODE LEAD FALL OFF. AT 06:45:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS INTERMITTENT CARDIAC ACTIVITY AND CPR/ELECTRODE LEAD FALL OFF. AT 06:46:12, THE PATIENT RECEIVED THE TENTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS INTERMITTENT CARDIAC ACTIVITY AND CPR/ELECTRODE LEAD FALL OFF. AT 06:46:47, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS INTERMITTENT CARDIAC ACTIVITY AND CPR/ELECTRODE LEAD FALL OFF. AT 06:48:10, THE PATIENT RECEIVED THE ELEVENTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS INTERMITTENT CARDIAC ACTIVITY AND CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS INTERMITTENT CARDIAC ACTIVITY AND CPR/ELECTRODE LEAD FALL OFF. AT 06:48:43, THE PATIENT RECEIVED THE TWELFTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. THE DEVICE WAS SHUT DOWN AT 07:49:14 ON 12/11/2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236357&pc=MVK | 3008642652-2025-00899 | 2024-10-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 19:12:03 ON (B)(6) 2024. AT 19:55:56, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ATRIAL FLUTTER @ 60 BPM WITH DEGRADING TO ASYSTOLE WITH FLUTTER WAVES AND CPR/MOTION ARTIFACT. AT 19:56:49, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 20:57:14 ON 10/26/2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236359&pc=MVK | 3008642652-2025-00897 | 2024-11-01 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 07:38:00 ON 10/31/2024. AT 03:22:02 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING AMPLITUDES. THE DEVICE PROPERLY DETECTED VF. VARYING AMPLITUDES PREVENTED THE LIFEVEST FROM TREATING THE PATIENT THE DEVICE WAS SHUT DOWN AT 03:45:44 ON 11/1/2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236360&pc=MVK | 3008642652-2025-00901 | 2024-10-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON AN UNKNOWN DATE. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT ON THE LAST WEAR DATE OF THE LIFEVEST. THE DEVICE WAS STARTED UP AT 08:16:57 ON (B)(6) 2024. AT 00:31:48 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING RATES/AMPLITUDES. AT 00:33:11, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH VARYING RATES/AMPLITUDES. POST SHOCK RHYTHM WAS ASYSTOLE WITH PACER SPIKES. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 05:06:55 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236361&pc=MVK | 3008642652-2025-00900 | 2024-10-15 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 17:59:22 ON (B)(6) 2024. AT 01:49:47 ON(B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM WITH TALL T WAVES AND CPR/MOTION ARTIFACT. AT 01:50:33, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS RHYTHM @ 90 BPM WITH TALL T WAVES AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 140 BPM WITH CPR/MOTION ARTIFACT. AT 01:51:31, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 50 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 02:26:31 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236363&pc=MVK | 3008642652-2025-00903 | 2024-11-10 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 05:32:35 ON 11/9/2024. AT 16:30:34 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 16:31:09, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 16:31:36, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 16:32:04, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 16:32:31, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 16:32:58, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 17:48:24 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236364&pc=MVK | 3008642652-2025-00902 | 2024-11-03 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-25 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 07:05:41 ON (B)(6) 2024. AT 14:57:28 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 14:58:39, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 15:12:45 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226813&pc=BZD | 2518422-2025-007619 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Adult Respiratory Distress Syndrome; Pneumonia; Chronic Obstructive Pulmonary Disease (COPD) | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION DEATH, ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS), CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND PNEUMONIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226852&pc=BZD | 2518422-2025-007335 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226908&pc=BZD | 2518422-2025-007655 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Cancer; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FOR DEVICE INFORMATION ON 31.03.2025 IN BOX D, PRODUCT BRAND NAME, MODEL NUMBER, CATALOG NUMBER, SERIAL NUMBER, UDI NUMBER HAVE BEEN UPDATED OR CORRECTED IN THIS REPORT. IN BOX H, MANUFACTURE DATE IS UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226987&pc=BZD | 2518422-2025-007322 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227167&pc=BZD | 2518422-2025-007701 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227255&pc=BZD | 2518422-2025-007688 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Stroke/CVA; Pneumonia; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DEATH, OTHER, MINI STROKES & PNEUMONIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227276&pc=BZD | 2518422-2025-007721 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAM STATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227406&pc=BZD | 2518422-2025-007739 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer; Heart Failure/Congestive Heart Failure; Respiratory Arrest | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, ASTHMA (NEW OR WORSENING), LUNG CANCER, PULMONARY ARREST AND CONGESTIVE HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227568&pc=BZD | 2518422-2025-007760 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227708&pc=BZD | 2518422-2025-007808 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Vomiting; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21227763&pc=BZD | 2518422-2025-007818 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228088&pc=BZD | 2518422-2025-007859 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228334&pc=BZD | 2518422-2025-007863 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD); Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX B: DESCRIBE EVENT OR PROBLEM WAS CORRECTED. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. BOX D: EXPIRATION DATE WAS CORRECTED. BOX H: PATIENT OUTCOME CODE GRID WAS CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228515&pc=MVK | 3008642652-2025-00863 | 2024-12-16 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-24 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 10:54:07 ON (B)(6)2024. AT 03:39:34 ON (B)(6)2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 100 BPM WITH PVCS/NSVT TRANSITIONING TO VT @ 250 BPM. AT 03:40:11, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 280 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM WITH PVCS/NSVT. AT 03:40:44, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 100 BPM WITH PVCS/NSVT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM WITH PVCS AND MOTION ARTIFACT. AT 03:42:34, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 110 BPM WITH PVCS/NSVT. AT 03:46:02, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BEAT TRANSITIONING BACK TO VT @ 180 BPM. AT 03:46:35, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS NSVT. AT 03:47:08, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 280 BPM. POST SHOCK RHYTHM WAS VT @ 260 BPM. AT 03:47:30, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 260 BPM. AT 03:48:04, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA/BRADYCARDIA FROM 100-50 BPM WITH PVCS/NSVT THAT DEGRADES TO VT/VF WITH CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 05:14:32 ON (B)(6)2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228686&pc=BZD | 2518422-2025-101162 | 2025-01-22 05:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION AUTO CPAP | Contamination; Device Emits Odor; Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED FOR INVESTIGATION. THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE DEVICE TO BE CONTAMINATED, THE BLOWER BOX WAS VERY DIRTY, AND THE DEVICE HAD AN ODOR. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED SECTION H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228726&pc=BZD | 2518422-2025-007892 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | DREAMSTATION AUTO BIPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE TWO ERROR CODE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228764&pc=OZD | 1220648-2025-25761 | 2025-01-06 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hypoxia; Vascular Dissection | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS BEING IMPLANTED WITH AN IMPELLA CP VIA SINGLE ACCESS WHEN THERE WAS A VESSEL DISSECTION. THE PATIENT WOULD LATER EXPIRE. UPON ARRIVAL TO CATHETERIZATION LAB THE PATIENT REQUIRED CARDIOPULMONARY RESUSCITATION (CPR) AND ACHIEVED RETURN OF SPONTANEOUS CIRCULATION (ROSC). FIRST, IMPELLA PREPPED, INSERTED AND STARTED AS RECOMMENDED, AND SINGLE ACCESS ATTEMPTED FOR ANGIOGRAM, BUT RIGHT FEMORAL ARTERY ENDED UP HAVING SMALL DISSECTION IN FEMORAL ARTERY SO LEFT FEMORAL ARTERY WAS ACCESSED USING A COMPANION SHEATH. ANGIOGRAM REVEALED CULPRIT LESION OF LEFT CIRCUMFLEX (LCX) ARTERY WITH LEFT ANTERIOR DESCENDING (LAD) ARTERY TOTAL OCCLUSION. A TOTAL OF TWO STENTS WAS DEPLOYED IN LCX AND TWO STENTS IN LAD. DURING PERCUTANEOUS CORONARY INTERVENTION, THE PATIENT REMAINED WITH A MEAN ARTERIAL PRESSURE OVER 80. HOWEVER, DUE TO HYPOXIA, THE CARDIOTHORACIC SURGERY CONSULTED FOR ECMO, BUT SURGEONS REFUSED. THE PATIENT CONTINUED TO DECOMPENSATE AND REQUIRED 12 SHOCKS AND CPR. ROSC ACHIEVED AND ADDITIONAL DRIPS ADDED. TEAM DECIDED TO FURTHER STABILIZE PATIENT ON THE UNIT. A SWAN CATHETER WAS SWAPPED TO A DIALYSIS CATHETER, AND THE PATIENT TRANSFERRED TO UNIT. HOWEVER, THE PATIENT WOULD LATER EXPIRE THE MORNING AFTER A TOTAL OF SEVEN HOURS ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE VASCULAR DAMAGE HAS BEEN COMPLETED. PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE RUN. AS PER THE GIVEN CLINICAL, DURING THE PROCEDURE, A SMALL DISSECTION (TEAR) OCCURRED IN THE RIGHT FEMORAL ARTERY (RFA) WHILE ATTEMPTING SINGLE ACCESS FOR THE ANGIOGRAM. DUE TO THIS COMPLICATION, THE TEAM SWITCHED TO THE LEFT FEMORAL ARTERY (LFA), USING A COMPANION SHEATH FOR ACCESS TO CONTINUE THE PROCEDURE. THE CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228899&pc=NIM | 2024168-2025-00830 | 2016-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-01-24 05:00 | NIM | ACCULINK | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; High Blood Pressure/ Hypertension; Nausea; Transient Ischemic Attack; Coma | P040012 |  | 1 | Event Description: THE PROCEDURE WAS TO TREAT AN OCCLUSION OF THE RIGHT INTERNAL CAROTID ARTERY (ICA). AFTER PRE-DILATION USING A 2.50X30MM RYUJIN BALLOON A 6MM SPIDER FILTER WAS DEPLOYED IN THE HIGH CERVICAL ICA. THE STENOSIS WAS PRE-DILATED AGAIN WITH A 5.50X30MM VIATRAC BALLOON AND THEN A 7-10X40MM ACCULINK STENTED WAS IMPLANTED. AFTER THE PROCEDURE, THE INTRACRANIAL CIRCULATION OF THE LEFT ICA WAS GREATLY IMPROVED; HOWEVER, IRRITABILITY, IRRESPONSIVENESS, AND PROJECTILE VOMITING WERE OBSERVED 10 MINUTES AFTER THE OPERATION, AND THE BLOOD PRESSURE INCREASED TO 186/98 MMHG. EMERGENCY CT REVEALED LEFT BASAL GANGLIA AND INTRAVENTRICULAR HEMORRHAGE AND EXTENSIVE SUBARACHNOID HEMORRHAGE. MEDICATION WAS IMMEDIATELY ADMINISTERED. SHORTLY THEREAFTER, THE PATIENT DEVELOPED A COMA ACCOMPANIED BY ANISOCORIA AND A DULLED PUPILLARY LIGHT REFLEX. THEREFORE, THE DIAGNOSIS OF CEREBRAL HERNIA WAS CONSIDERED. THE PATIENT¿S BLOOD PRESSURE GRADUALLY DECREASED AFTER EMERGENCY TRACHEAL INTUBATION, TREPANATION, AND DRAINAGE OF THE HEMATOMA. A CATHETER WAS PLACED INTO THE LEFT CAROTID ARTERY TO IDENTIFY THE SOURCE OF THE HEMORRHAGE, AND DIGITAL SUBTRACTION ANGIOGRAPHY (DSA) SHOWED SPECKLED EXTRAVASATION OF CONTRAST AGENT AT THE SMALL BRANCHES OF THE MIDDLE CEREBRAL ARTERY (MCA). THE PATIENT DIED 40 HOURS AFTER CAS. NO ADDITIONAL INFORMATION WAS PROVIDED. DETAILS ARE LISTED IN THE ARTICLE TITLED, ¿FATAL INTRACRANIAL HEMORRHAGE AFTER CAROTID ARTERY STENTING: THREE CASE REPORTS AND A LITERATURE REVIEW". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A PRODUCTION RECORD REVIEW AND A REVIEW OF THE COMPLAINT HANDLING DATABASE WAS NOT PERFORMED BECAUSE THE PART AND LOT NUMBER WAS NOT REPORTED. THE REPORTED PATIENT EFFECTS OF HYPOTENSION, HEMORRHAGE AND DEATH IS LISTED IN THE RX ACCULINK CAROTID STENT SYSTEM INSTRUCTIONS FOR USE AS KNOWN PATIENT EFFECTS ASSOCIATED WITH THE USE OF A STENT IN CAROTID ARTERIES. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. A CONCLUSIVE CAUSE FOR THE REPORTED COMA, HIGH BLOOD PRESSURE/ HYPERTENSION, INTRACRANIAL HEMORRHAGE, TRANSIENT ISCHEMIC ATTACK AND NAUSEA AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE TREATMENTS APPEAR TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. BASED ON THE REPORTED INFORMATION AND RESULTS OF THE COMPLAINT INVESTIGATION THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. LITERATURE ATTACHMENT: ARTICLE TITLE ¿FATAL INTRACRANIAL HEMORRHAGE AFTER CAROTID ARTERY STENTING: THREE CASE REPORTS AND A LITERATURE REVIEW¿ B2: DATE OF DEATH ESTIMATED AS 1/03/2016 B5: DATE OF PROCEDURE ESTIMATED AS (B)(6) 2016 D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: DATE OF IMPLANT ESTIMATED AS (B)(6) 2016.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21228990&pc=DSQ | 2916596-2025-00221 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY WITH AN UNKNOWN CAUSE OF DEATH. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: SECTION A - ALL KNOWN PATIENT INFORMATION WAS PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS NOTED THAT THE PUMP WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE ACCOUNT DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21229333&pc=OZD | 1220648-2025-25764 | 2025-01-07 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND SAME DAY OF IMPLANT HAD HEAVY BLEEDING AT THE ACCESS SITE REQUIRING TWO UNITS RED BLOOD CELLS. STATUS OF THE PATIENT AS A RESULT OF THE EVENT IS UNKNOWN. NO FURTHER DETAILS WERE PROVIDED. THE PATIENT CONTINUES ON IMPELLA MCS.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS THE DEVICE REMAINS IMPLANTED, SUPPORTING THE PATIENT AT THE TIME OF THIS MDR WRITING, AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25764 TO REFLECT THE PATIENT OUTCOME; DEATH, DATE OF DEATH WAS UNKNOWN. B5 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25764 TO REFLECT THE DETAILS OF THE PATIENT OUTCOME. D3 MANUFACTURE ZIP CODE WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25764. D6B WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25764. H1 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25764 TO REFLECT THE DETAILS OF THE PATIENT OUTCOME. H6 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25764 TO REFLECT THE DETAILS OF THE PATIENT OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21229794&pc=CAW | 3004672275-2025-00009 | 2024-12-16 05:00 | Death | INOGEN, INC. | 2025-01-24 05:00 | CAW | INOGEN ONE G3 PORTABLE OXYGEN CONCENTRATOR | Failure to Run on Battery | Hypoxia | K032818 |  | 1 | Event Description: ACCORDING TO THE PATIENT'S WIFE, HER HUSBAND DID HAVE TO GO TO HOSPITAL AND DID EXPERIENCE ANY ADVERSE HEALTH CONSEQUENCES FROM EVENT. HE WAS TAKEN TO THE HOSPITAL VIA AMBULANCE. IT WAS REPORTED THAT UPON ARRIVAL THE PATIENT WAS NON-RESPONSIVE. SHE REPORTED A BATTERY CONNECTION ISSUE, AND HER HUSBAND HAD LACK OF O2. THERE WERE NO REPORTED AUDIBLE OR VISUAL ALARMS AT THE TIME OF THE EVENT. THERE WAS ALTERNATIVE OXYGEN SUPPLY FOR USE DURING THE DEVICE ISSUE. THE PATIENT EXPERIENCED A MASSIVE HEART ATTACK AND DIED. Manufacturer Narrative: A RETURN OF THE DEVICE HAS BEEN INITIATED. ONCE THE DEVICE IS RECEIVED AN INVESTIGATION WILL BE CONDUCTED AND A FOLLOW UP WILL BE SUBMITTED IF REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21230344&pc=CAT | 3012316249-2025-00006 | 2025-01-04 05:00 | Death | Unknown Manufacturer | 2025-01-24 05:00 | CAT | UNKNOWN | Use of Device Problem | Partial thickness (Second Degree) Burn |  |  | 1 | Event Description: PATIENT SUFFERED FLASH BURN TO FACE WHILE SMOKING AND USING OXYGEN CONCENTRATOR. PATIENT SUSTAINED 2ND DEGREE BURNS TO FACE. PATIENT WAS TRANSFERRED TO TRAUMA CENTER AND WAS PLACED ON A VENTILATOR. PATIENT WAS EXTUBATED SAME DAY AND SUBSEQUENTLY DIED. THERE IS NO DATA PROVIDED TO SUPPORT THAT THE DEVICE(S) IN QUESTION MALFUNCTIONED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21230698&pc=OZD | 1220648-2025-25769 | 2024-11-25 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: COMPLAINANT IN JAPAN REPORTED THE PATIENT WAS ON IMPELLA CP SMART ASSIST SUPPORT AND DURING SUPPORT, THE PATIENT HAD SYSTEMIC BLEEDING. THE PATIENT'S CHEST WAS REOPENED AND HEMOSTASIS WAS ACHIEVED. THE CAUSE OF THE BLEEDING WAS NOT FROM THE IMPELLA, HOWEVER THE DOCTOR STATED THAT THE IMPELLA EXACERBATED THE BLEEDING. THE PATIENT HAD A REPLACEMENT PUMP AND EXPIRED ON SUPPORT OF THE REPLACEMENT PUMP. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE BLEEDING WAS OCCURRING PRIOR TO THE IMPELLA INSERTION. AFTER PUMP INSERTION, THE BLEEDING WAS POSSIBLY EXACERBATED AND HEMOSTASIS WAS ACHIEVED BY REOPENING THE CHEST. NO OTHER INFORMATION REGARDING THE VASCULAR DAMAGE WAS PROVIDED. THEREFORE, THE ROOT CAUSE OF THE VASCULAR DAMAGE WAS UNABLE TO BE DETERMINED AS INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND THE RELATION TO IMPELLA IS UNKNOWN. B.2 REVISED AS REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25769. H.6 CODE 4641 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25769 AND A NEW CODE WAS ADDED. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE (IFUS) WERE REPORTED ON THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25769 INCORRECTLY. NO IFUS ARE NEEDED TO BE REPORTED FOR THIS REPORT IN ACCORDANCE WITH UPDATED PROCEDURES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21230865&pc=OZD | 1220648-2025-25773 | 2025-01-09 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT DEVELOPED BLEEDING FROM THE IMPELLA ACCESS SITE. THE PATIENT GROIN SITE PREVIOUSLY STABLE HAD BEGUN OOZING DESPITE PROPER PATIENT POSITIONING AND DRESSING TECHNIQUES. THE PATIENT WAS SWITCHED TO BIVAL AND IS OFF HEPARIN. THE PATIENT REQUIRED TRANSFUSION OF PACKED RED BLOOD CELLS AND PLATELETS. IT WAS FURTHER REPORTED THAT THE IMPELLA WAS EXPLANTED FOR PATIENT TO GO ON COMFORT MEASURES AND THE PATIENT EXPIRED AFTER WITHDRAWAL. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS UNABLE TO BE DETERMINED AS INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21230887&pc=BZD | 2518422-2025-007997 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-24 05:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM-RECRT | Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF CANCER AND DEATH . THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231085&pc=DSQ | 3007042319-2025-00106 | 2025-01-15 05:00 | Death | HEARTWARE, INC. | 2025-01-24 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | High Blood Pressure/ Hypertension; Renal Failure; Thrombosis/Thrombus; Lactate Dehydrogenase Increased | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) EXPERIENCED SUSPECTED VAD THROMBUS AND HIGH-POWER ALARMS. THE PATIENT INITIALLY CONTACTED THE DEVICE COORDINATORS DUE TO THESE ALARMS. ELEVATED LACTATE DEHYDROGENASE (LDH) LEVELS WERE NOTED, AND THE PATIENT WAS PLACED ON AN ANTICOAGULANT DRIP WITH PLANS FOR THROMBOLYTIC THERAPY IF TOLERATED. DESPITE MULTIPLE ROUNDS OF THROMBOLYTICS, THE PATIENT REQUIRED ADDITIONAL SUPPORT FOR BLOOD PRESSURE, AND KIDNEY FAILURE ENSUED. THE DECISION WAS MADE TO DISCONTINUE CARE AND TURN OFF THE VENTRICULAR ASSIST DEVICE, AFTER WHICH THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE LOG FILE ANALYSIS REVEALED SEVERAL INCREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS BEGINNING ON (B)(6) 2025, LEADING TO PARAMETERS ABOVE NORMAL OPERATING RANGE, AND ONE HUNDRED AND SEVEN (107) HIGH WATT ALARMS LOGGED SINCE (B)(6) 2025. AS A RESULT, THE REPORTED HIGH POWER EVENT WAS CONFIRMED. THE REPORTED THROMBUS COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS TREATED WITH AN ANTICOAGULANT DRIP AND MULTIPLE ROUNDS OF THROMBOLYTICS. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER EXPERIENCING THROMBUS, HYPERTENSION AND KIDNEY FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE AVAILABLE INFORMATION AND HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF THE HIGH POWER EVENT CAN BE ATTRIBUTED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS, HYPERTENSION, RENAL DYSFUNCTION, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: DDBC3D4 ICD IMPLANT DATE: (B)(6)2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231123&pc=GCJ | 3005075853-2025-00629 | 2024-12-26 05:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-01-24 05:00 | GCJ | 5MM XCEL STABILITY SLEEVE | Crack | Insufficient Information | K032676 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A LAPAROSCOPIC APPENDECTOMY, A TROCAR WAS INSERT INTO THE PATIENT, AS INSERTING THE TROCAR HIT THE ILIAC WALL AT THIS TIME THE TROCAR FELT LIKE " IT BENT LIKE A STRAW" STATED BY FACILITY. THEY TRIED TO CONVERT TO OPEN. CASE RESULTED IN PATIENT DEATH. Manufacturer Narrative: (B)(4). DATE SENT: 2/24/2025. D4: BATCH # UNK. THE DEVICE UPON WHICH THIS MEDWATCH IS BASED HAS BEEN RECEIVED, HOWEVER, THE EVALUATION IS NOT YET COMPLETE. ANY FURTHER INFORMATION DERIVED FROM THE EVALUATION WILL BE SUBMITTED IN A SUPPLEMENTAL 3500A FORM. Manufacturer Narrative: (B)(4). DATE SENT: 5/2/2025. D4: BATCH # UNK. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. VISUAL INSPECTION WAS CONDUCTED ON THE RETURNED DEVICE. VISUAL ANALYSIS OF THE RETURNED SAMPLE DETERMINED THAT THE CB5LT DEVICE WAS RECEIVED WITH THE SLEEVE BROKEN. ADDITIONAL MATERIALS ANALYSIS WAS PERFORMED ON THE DEVICE. SEM (SCANNING ELECTRON MICROSCOPY) IMAGING OF THE BROKEN SURFACE INDICATES THAT THE FRACTURE IS DUCTILE IN NATURE. ONE POSSIBLE CAUSE FOR THE DAMAGE FOUND MAY BE DUE TO EXCESSIVE FORCE BEING APPLIED TO THE DEVICE. PLEASE NOTE ONCE PARTIAL ENTRY HAS BEEN ACCOMPLISHED, VERY LITTLE PRESSURE MAY BE REQUIRED TO COMPLETE ENTRY. EXCESSIVE PRESSURE COULD CAUSE INJURY TO INTRA-ABDOMINAL OR INTRA-THORACIC STRUCTURES. FOR ADDITIONAL INFORMATION PLEASE REFER TO THE INSTRUCTIONS FOR USE FOR ADDITIONAL INFORMATION REGARDING PROPER DEVICE USAGE. IN ADDITION, AN ELEMENTAL ANALYSIS IDENTIFIED SILICONE LUBRICANT THAT IS PRESENT ON THE TROCAR AND GLYCEROL TRIOLEATE AS SUGGESTED CHEMICAL MATCHES FOR THE SPECTRUM. GLYCEROL TRIOLEATE IS LIKELY SURGICAL RESIDUE. NO EVIDENCE OF CONTAMINANTS WERE FOUND IN THE COMPLAINT SAMPLE. THE MANUFACTURING RECORDS COULDN'T BE REVIEWED AS THE BATCH NUMBER IS UNKNOWN. Manufacturer Narrative: (B)(4). DATE SENT: 1/24/2025. D4: BATCH # UNK. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: "WAS THIS THE FIRST TROCAR THAT WAS INSERTED IN THE PROCEDURE? YES. THIS WAS THE FIRST TROCAR INSERTED DURING THE PROCEDURE. -WAS THE ABDOMEN INSUFFLATED FIRST? NO. THE ABDOMEN WAS NOT INSUFFLATED FIRST. -WHAT OBTURATOR WAS USED WITH THIS TROCAR? IT WAS A STANDARD OBTURATOR WITH THE OPTIMAL CUTTING TROCAR (HAS A CLEAR TIP). -WILL THE OBTURATOR BE RETURNED FOR ANALYSIS? NO. -DID THE TROCAR BREAK DURING INSERTION OR DURING USE WITHIN THE SURGICAL PROCEDURE? THE TROCAR BROKE DURING THE PROCEDURE; IT WAS LOCKED IN PLACE, BUT DID NOT BREAK AT THAT TIME. THE SURGEON TOOK THE CAMERA OUT WITH THE OBTURATOR. IN AN ATTEMPT TO WITHDRAW THE TROCAR AND INSUFFLATE, TOO MUCH TORQUE WAS APPLIED ON THE TROCAR, SNAPPING IT IN THE MID SHAFT DUE TO THE PATIENT¿S BODY HABITUS. -IF DURING INSERTION, WAS THE OBTURATOR LOCKED IN PLACE DURING THE INSERTION? WHAT TYPE OF OBTURATOR? -IF DURING USE, WHAT INSTRUMENTS MAY HAVE BEEN WITHIN THE TROCAR AT THE TIME AND WERE THE INSTRUMENTS BEING TORQUED TO ONE DIRECTION? THERE WERE NO INSTRUMENTS IN THE TROCAR AT THE TIME IT BROKE. THE CAMERA HAD ALREADY BEEN REMOVED." ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231250&pc=DSQ | 3007042319-2025-00107 | 2024-11-12 05:00 | Death | HEARTWARE, INC. | 2025-01-24 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO AGE-RELATED REASONS, AND IT WAS NOTED THAT THE VENTRICULAR ASSIST DEVICE WAS NOT INVOLVED IN THE CAUSE OF DEATH. THE DEVICE WAS IMPLANTED AND IS NOW OUT OF SERVICE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231272&pc=OBJ | 2124215-2025-03471 | 2024-12-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-24 05:00 | OBJ | OPTICROSS HD | Entrapment of Device; Detachment of Device or Device Component; Material Twisted/Bent | Cardiac Arrest; Foreign Body In Patient; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT A CATHETER ISSUE LED TO THE PATIENT'S DEATH. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE TIP OF THE OPTICROSS HD CATHETER SNAPPED WHILE IN THE RIGHT CORONARY ARTERY. FOLLOWING THIS DETACHMENT, THE PATIENT HAD DIED.Event Description: IT WAS REPORTED THAT A CATHETER ISSUE LED TO THE PATIENT'S DEATH. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE TIP OF THE OPTICROSS HD CATHETER SNAPPED WHILE IN THE RIGHT CORONARY ARTERY. FOLLOWING THIS DETACHMENT, THE PATIENT HAD DIED. IT WAS LATER REPORTED THAT THIS EVENT WAS ALREADY REPORTED UNDER ANOTHER EVENT. TRACKWISE NUMBER (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231300&pc=OBJ | 2124215-2025-03470 | 2024-12-17 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-24 05:00 | OBJ | OPTICROSS HD | Detachment of Device or Device Component | Cardiac Arrest; Foreign Body In Patient; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT A CATHETER ISSUE LED TO THE PATIENT'S DEATH. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE TIP OF THE OPTICROSS HD CATHETER SNAPPED WHILE IN THE RIGHT CORONARY ARTERY. FOLLOWING THIS DETACHMENT, THE PATIENT HAD DIED.Event Description: IT WAS REPORTED THAT A CATHETER ISSUE LED TO THE PATIENT'S DEATH. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. DURING THE PROCEDURE, THE TIP OF THE OPTICROSS HD CATHETER SNAPPED WHILE IN THE RIGHT CORONARY ARTERY. FOLLOWING THIS DETACHMENT, THE PATIENT HAD DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231383&pc=DSQ | 2916596-2025-00147 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Septic Shock; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. THE SEPSIS WAS ASSUMED TO BE DUE TO AN INFECTION, THOUGH A SOURCE WASN¿T IDENTIFIED. IT WAS LIKELY DEVELOPING AT TIME OF VENTRICULAR ASSIST DEVICE (VAD) IMPLANT. THE INFECTION WAS DUE TO THE DETERIORATION OF CARDIAC FUNCTION AND THE NEED FOR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT WAS NOT IN MULTI SYSTEM ORGAN FAILURE (MSOF) PRIOR TO VAD IMPLANT THOUGH DEVELOPED IT AFTERWARDS. THE SEPSIS WAS THOUGHT TO MOST LIKELY PRE-DATE THE VAD IMPLANT (OR BE CAUSED BY REMOVAL OF ECMO CANNULA DURING THE OPERATION), HOWEVER, WAS NOT RECOGNIZED ON THE DAY. THE PATIENT'S DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. DEVICE WAS WORKING AS EXPECTED AND WOULD NOT BE RETURNED. NO LOG FILES WERE AVAILABLE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C AND PATIENT HANDBOOK, REV. B ARE CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING SEPSIS, INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), MULTIPLE TYPES OF ORGAN FAILURE, AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION. THE IFU ALSO PROVIDES SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231413&pc=MOZ | MW5165189 | 2024-05-01 04:00 | Death | FERRING PHARMACEUTICALS INC. | 2025-01-24 05:00 | MOZ | EUFLEXXA |  | Insufficient Information |  |  | 1 | Event Description: EVENT: PT DECEASED. FREQ: INJECT 1 SYRINGE INTO EACH KNEE WEEKLY FOR 3 WEEKS. PRESCRIBER INFO: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231590&pc=LOZ | 3003761017-2025-00083 | 2024-12-26 05:00 | Death | SYNCARDIA SYSTEMS, LLC | 2025-01-24 05:00 | LOZ | SYNCARDIA 70CC TAH | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P030011 |  | 1 | Event Description: PER HOSPITAL STAFF, REPORTED DEATH OF PATIENT FROM MULTI-SYSTEM ORGAN FAILURE AFTER 742 DAYS ON SUPPORT. THE OUTCOME FORM INDICATES THERE WAS NO AUTOPSY, THE DEVICE WAS NOT EXPLANTED, AND THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. Manufacturer Narrative: PATIENT DEATH DUE TO MULTI-SYSTEM ORGAN FAILURE. DEVICE CONFIRMED BY HOSPITAL STAFF TO NOT HAVE CAUSED OR CONTRIBUTED TO PATIENT DEATH. DEVICE WAS NOT EXPLANTED AND NO AUTOPSY WAS PERFORMED. SYNCARDIA HAS COMPLETED ITS EVALUATION AND IS CLOSING THIS FILE. IF NEW OR ADDITIONAL INFORMATION IS RECEIVED IN THE FUTURE, SYNCARDIA WILL FILE A FOLLOW-UP MDR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231594&pc=PYX | 1220648-2025-25777 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | PYX | IMPELLA RP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA RP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND BEGAN BLEEDING FROM AN UNKNOWN SOURCE. THE PATIENT WOULD EVENTUALLY EXPIRE. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT AFTER HAVING A VENTRICULAR FIBRILLATION ARREST AT HOME. IN EMERGENCY DEPARTMENT THE PATIENT CODED MULTIPLE TIMES, INTUBATED, AND WAS TRANSFERRED TO CATHETERIZATION LAB FOR AN ANGIOGRAM. AFTER 10% EJECTION FRACTION ON LEFT VENTRICULAR GRAM, THE RIGHT CORONARY ARTERY (RCA) DISEASE FROM ANGIOGRAM, AND CARDIOGENIC SHOCK WITH UNSTABLE HEMODYNAMIC STATUS, THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP PRIOR TO PERCUTANEOUS CORONARY INTERVENTION (PCI). AFTER THE IMPELLA CP INSERTED, THE PCI WAS PERFORMED ON RCA, THE PATIENT LOST PULSATILITY AND CONTINUED TO DECOMPENSATE HEMODYNAMICALLY. A RIGHT HEART CATHETERIZATION WAS COMPLETED AND PULMONARY ARTERY PULSATILITY INDEX REVEALED AS 0.3. THE DECISION WAS MADE TO ESCALATE TO AN ADDITIONAL IMPELLA RP DEVICE (WITH HEPARIN IN THE PURGE SYSTEM). THE PATIENT'S PULSATILITY CAME BACK AS SOON AS THE IMPELLA RP WAS STARTED, AND THE MEAN ARTERIAL PRESSURE WAS MAINTAINED ABOVE 55. THE IMPELLA WAS SECURED, THE SHEATH WAS SECURED AND DRESSED, AND THE PATIENT TRANSFERRED TO UNIT ON BIPELLA SUPPORT. FOLLOWING THE IMPLANT, PULSES IN THE RIGHT LEG (IMPELLA CP SIDE) BECAME WEAK AND THE RIGHT FOOT APPEARED MOTTLED. THE LEG BEGAN TO SWELL, AND A FASCIOTOMY WAS SUBSEQUENTLY COMPLETED TO RESTORE PULSE TO THE LEG. FOUR DAYS LATER, PULSE WAS ONCE AGAIN ABSENT IN THE PATIENT'S LEG (IMPELLA CP SIDE). THE PATIENT HAD BEEN GIVEN TEN UNITS OF BLOOD AS A RESULT OF THE CONDITION AND INTERVENTION. COINCIDING WITH THE ISCHEMIA, THE PATIENT ALSO DEVELOPED HEMOLYSIS WITH FOUR UNITS OF BLEED BEING GIVEN TO HELP MANAGE THE HEMOLYSIS. THE SUPPORT LEVEL WAS LOWERED TO REDUCE SUCTION THAT MAY HAVE BEEN CONTRIBUTING TO THE HEMOLYSIS. AS A RESULT OF THE PRIOR INTERVENTIONS, THE PATIENT EXPERIENCED BLEEDING COMPLICATIONS FROM THE FASCIOTOMY SITES. MULTIPLE BLOOD PRODUCTS WERE GIVEN TO COUNTERACT THE BLEED AND HEPARIN LEVELS WERE REDUCED. THE PATIENT LATER PASSED AWAY AFTER SEVEN DAYS OF SUPPORT. IT IS UNKNOWN IF THE COMPLICATIONS CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE EVENT DATE REFLECTS THE SAME AS THE ASSOCIATED MEDICAL DEVICE REPORT ACCORDINGLY (BEST ESTIMATE) GIVEN TOTALLY OF THE INFORMATION FOR THE EVENT. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. NOTE: AS HEPARIN WAS IN THE IMPELLA RP PURGE, THE IMPACT/EXACERBATION TO THE BLEEDING IS UNKNOWN. THEREFORE, THIS IS ONE OF TWO MEDICAL DEVICES ASSOCIATED WITH THIS EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-25587. USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER: ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿ SECTION: WARNINGS AND CAUTIONS: ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND. ¿BE SURE THAT THE STOPCOCK ON THE REPOSITIONING SHEATH IS ALWAYS KEPT IN THE CLOSED POSITION. SIGNIFICANT BLEED BACK CAN RESULT IF THE STOPCOCK IS OPEN.¿ ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE FEMORAL VEIN. CLOSE AND DRESS THE WOUND.¿ Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. G1 MANUFACTURING SITE NAME AND ADDRESS WAS INADVERTENTLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25777.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231613&pc=DSQ | 2916596-2025-00582 | 2024-11-19 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Renal Failure; Sepsis; Unspecified Mental, Emotional or Behavioural Problem; Pulmonary Hypertension; Unspecified Respiratory Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS ADMITTED AND SUFFERING FROM SEPSIS. THE PATIENT'S FAMILY DECIDED TO TRANSITION THE PATIENT TO PALLIATIVE CARE. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE DEATH WAS NOT CONSIDERED TO BE RELATED TO THE FUNCTIONING OF THE HEARTMATE 3. LOG FILES CONTAINED A LOSS OF EXTERNAL POWER EVENT AS WELL OTHER ROUTINE LOW VOLTAGE EVENTS. THE LOSS OF EXTERNAL POWER EVENT APPEARED TO HAVE BEEN A RESULT OF A SIMULTANEOUS SYSTEM CONTROLLER LEAD DISCONNECT FROM POWER. THE LOG FILES ALSO REVEALED THAT AT 23:12:14, THE BUS AND RELATIVE STATE OF CHARGE (RSOC) VOLTAGE ON THE BLACK POWER CABLE CONNECTED TO THE MOBILE POWER UNIT (MPU) BEGAN TO DECREASE CONSISTENT WITH DISRUPTION OF POWER TO THE MPU. VOLTAGES CONTINUE TO DECREASE UNTIL 23:29:23 ACTIVATING LOW VOLTAGE AND POWER CABLE DISCONNECT ALARMS. AT 23:31:29, VOLTAGE ON THE MPU REVERTED TO EXPECTED VALUES. IT WAS REPORTED UNKNOWN WHETHER THE MPU HAD A V-LOCK CONNECTOR ON THE AC POWER CORD. IT WAS UNKNOWN WHETHER THE AC POWER CORD CAME LOOSE AT THE WALL OUTLET OR FROM THE BACK OF THE UNIT. IT WAS ALSO UNKNOWN WHETHER THERE WERE ANY ENVIRONMENTAL CIRCUMSTANCES THAT WOULD HAVE AFFECTED AC POWER AT THE TIME OF THE EVENT. THE MPU WAS IN SERVICE AT THE TIME. THE LOSS OF EXTERNAL POWER WAS BELIEVED TO BE RELATED TO ERROR IN TRANSPORTATION AND TRANSFERRING FROM THE SKILLED NURSING FACILITY TO THE AMBULANCE AND THEN TO EMERGENCY DEPARTMENT. THE EMERGENCY ROOM STAFF STATED THE PATIENT WAS RECEIVED WITH ONE POWER LEAD ATTACHED TO A 14 VOLT BATTERY AND ONE POWER LEAD ATTACHED TO THE MPU WHICH WAS NOT PLUGGED INTO WALL POWER.Event Description: ADDITIONAL INFORMATION STATED THAT ONLY INFECTION SOURCE WAS THE CONFIRMED CORONAVIRUS DISEASE (COVID) 19 INFECTION (IMAGING SUGGESTIVE OF PNEUMONIA VS ACUTE LUNG INJURY). BLOOD CULTURES WERE NEGATIVE. THE PATIENT EXPERIENCED HYPOTENSION, RESPIRATORY DISTRESS, ALTERED MENTAL STATUS, AND ACUTE RENAL FAILURE. THEY WERE TREATED FOR BACTERIAL PNEUMONIA IN ADDITION BROAD COVERAGE ZOSYN, FLUCONAZOLE, AZITHROMYCIN, DEXAMETHASONE 6MG INTRAVENOUS (IV) DAILY, VANCOMYCIN 500MG ORAL 4 TIMES A DAY WITH IV FLAGYL 500 MG THREE TIMES A DAY AND VANCOMYCIN ENEMA. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED IMAGES CONFIRMED THE REPORTED PUMP CABLE DAMAGE; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE ACCOUNT¿S SUBMITTED IMAGES REVEALED SUPERFICIAL DAMAGE TO THE OUTER JACKET OF THE PUMP CABLE, EXPOSING THE UNDERLYING BIONATE LAYER AT ONE LOCATION. IT WAS ALSO CONFIRMED THAT THE BEND RELIEF WAS MISSING ABOVE THE MODULAR CONNECTION. THE DAMAGE APPEARED TO BE COSMETIC ONLY. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2024. NO NOTABLE EVENTS WERE CAPTURED, AND THE PUMP APPEARED TO OPERATE AS INTENDED AT THE FIXED SPEED. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS ADVERSE EVENTS, INCLUDING SEPSIS, INFECTION, RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE IFU AND PATIENT HANDBOOK PROVIDE INFORMATION REGARDING HOW TO CLEAN AND CARE FOR THE DRIVELINE AND INSTRUCT THE USER TO INSPECT THEIR DRIVELINE FOR SIGNS OF DAMAGE. THE IFU AND PATIENT HANDBOOK FURTHER INSTRUCT THE USE TO CLEAN EXTERIOR SURFACES OF THE DRIVELINE AND CABLES WITH A DAMP, LINT-FREE CLOTH. IF MORE AGGRESSIVE CLEANING IS NEEDED, USE WARM WATER AND MILD DISH SOAP. SEVERAL SECTIONS OF THE HEARTMATE 3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE PATIENT HANDBOOK CAUTIONS THE USERS TO CALL THEIR HOSPITAL CONTACTS IF THEY THINK, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232054&pc=DSQ | 2916596-2025-00288 | 2024-10-31 04:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE. THE PATIENT HAD MODERATE RIGHT HEART FUNCTION FROM IMPLANTATION BUT DETERIORATED. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: THE DEVICE OPERATED AS EXPECTED. THE HEART FAILURE EXISTED PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) MODERATELY. A DEVICE ISSUE DID NOT CAUSE OR CONTRIBUTE TO RIGHT HEART FAILURE. THERE WAS NO INFORMATION PROVIDED ON PATIENT SYMPTOMS AND TREATMENT. Manufacturer Narrative: SECTION A3 - PATIENT GENDER: CORRECTED. SECTION B7 - OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: CORRECTED. SECTION D5 - OPERATOR OF DEVICE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AS WELL AS RECORDS IN THE MANUFACTURING EXECUTION SYSTEM AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS RIGHT HEART FAILURE AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS RIGHT HEART FAILURE AS A POTENTIAL RISK/ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232406&pc=DSQ | 2916596-2025-00228 | 2024-01-18 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Coagulation Disorder; Depression; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE THAT THE PATIENT PASSED AWAY FROM SEPTIC MULTI-ORGAN FAILURE CAUSED BY A DRIVELINE INFECTION. AFTER INITIAL IMPLANTATION OF THE HEARTMATE 3, THE PATIENT COULD ONLY ATTAIN AN INTERNATIONAL NORMALIZED RATIO (INR) OF 1.5 (INSTEAD OF THE TARGET 2.0 TO 3.0). THE PATIENT CHANGED FROM THE ANTI-COAGULATION THERAPY TO JANTOVEN (WARFARIN) COMBINED WITH PLAVIX (CLOPIDOGREL) 75MG ONCE A DAY. IN ADDITION, THE THERAPY WAS OPTIMIZED ACCORDING TO THE COAGULATION FINDINGS, AND VITAMIN K SUPPLEMENTS WERE USED TO STABILIZE THE PATIENT¿S CONDITION. BECAUSE OF THE POORLY CONTROLLED COAGULATION, THE PATIENT REQUIRED MULTIPLE LONG-TERM HOSPITAL STAYS WITH INTRAVENOUS HEPARIN BRIDGING. THE PATIENT ALSO EXPERIENCED DEPRESSIVE EPISODES NECESSITATING PSYCHIATRIC TREATMENT. THE THERAPY WAS SWITCHED TO RIVAROXABAN WITH CLOPIDOGREL. THE PATIENT REMAINED OFF WARFARIN FOR 4 YEARS UNTIL THEY DIED OF SEPTIC MULTI-ORGAN FAILURE CAUSED BY A DRIVELINE INFECTION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DEPARTMENT OF CARDIOTHORACIC SURGERY, HELIOS CLINIC SIEGBURG, SIEGBURG, GERMANY DEPARTMENT OF CARDIOTHORACIC SURGERY, UNIVERSITY HOSPITAL ZURICH CLINIC OF CARDIOVASCULAR SURGERY, ZURICH, SWITZERLAND DEPARTMENT OF CARDIOTHORACIC SURGERY AND TRANSPLANT SURGERY, MEDICAL HIGH SCHOOL HANNOVER, HANNOVER, GERMANY AUTHOR CITATION: DHARUMAN Y, ET AL. LONG-TERM DISCONTINUATION OF WARFARIN IN A PATIENT WITH HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE WITHOUT THROMBOEMBOLIC EVENTS. J INT MED RES. 2024 JUN;52(6):3000605241258474. DOI: 10.1177/03000605241258474. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE. SECTION E1: CUSTOMER CONTACT INFORMATION WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED DRIVELINE INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENT OF A DEPRESSIVE EPISODE, AS WELL AS THE PATIENT OUTCOME, COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ""INTRODUCTION"", LISTS DRIVELINE INFECTION, SEPSIS, AND MULTIPLE ORGAN FAILURES/DYSFUNCTIONS, AND PSYCHIATRIC EPISODE AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232469&pc=DSQ | 2916596-2025-00368 | 2024-06-29 04:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO LIVER FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) TEAM WITHDREW CARE. THE VAD WAS FUNCTIONING AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS HEPATIC DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232629&pc=QFG | 3013756811-2025-11146 | 2024-10-11 04:00 | Death | TANDEM DIABETES CARE | 2025-01-24 05:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: A MALFUNCTION ALARM OCCURRED (PREVIOUSLY SUBMITTED VIA MDR(B)(6)) ON 10/10/2024 AT APPROXIMATELY AT 12:08 PM AND THE CUSTOMER MANUALLY RESUMED INSULIN DELIVERY AT 5:42 PM. ON(B)(6)2024 AT 3:36 AM, THE CUSTOMER MANUALLY SUSPENDED INSULIN DELIVERIES. ON THE SAME DAY, THE CUSTOMER WAS FOUND UNRESPONSIVE WITH A BLOOD GLUCOSE LEVEL OF OVER 500 MG/DL AND HIGH KETONE LEVELS AND WAS TRANSPORTED TO THE EMERGENCY ROOM (ER). THE CUSTOMER WAS HOSPITALIZED AND WAS TREATED WITH LONG-LASTING MANUAL INSULIN INJECTIONS ONCE DIABETIC KETOACIDOSIS (DKA) SYMPTOMS HAD ALREADY BEGAN. ON (B)(6)2024, THE CUSTOMER PASSED AWAY. PER THE CONTACT, THE CAUSE OF DEATH WAS DKA THAT LED TO ORGAN FAILURE. THE CUSTOMER WAS NOT WEARING THE PUMP AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. H3 OTHER TEXT : DEVICE NOT RETURNED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233033&pc=DSQ | 2916596-2025-00338 | 2024-07-22 04:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THEY WERE FOUND DEAD, AND THE CAUSE OF DEATH WAS UNCLEAR. THE PUMP WORKED AS INTENDED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED NO ADDITIONAL INFORMATION ABOUT THE REPORTED PATIENT OUTCOME WAS SHARED BY THE CUSTOMER. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233036&pc=DSQ | 2916596-2025-00317 | 2024-01-09 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2025. TRAUMATIC SUBARACHNOID HEMORRHAGE WAS REPORTED.Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION A2: PATIENT AGE CORRECTED. SECTION B3: DATE OF DEATH AND DATE OF EVENT CORRECTED. SECTION B5: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED PATIENT OUTCOME; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233576&pc=MWI | 3005053560-2024-00008 | 2024-12-26 05:00 | Death | IRADIMED CORPORATION | 2025-01-24 05:00 | MWI | 3880 MRI PATIENT MONITORING SYSTEM | Adverse Event Without Identified Device or Use Problem | Encephalopathy | K180903 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED AN EVENT INVOLVING AN IRADIMED PATIENT MONITOR. ACCORDING TO THE CUSTOMER, AN ICU PATIENT WAS TRANSFERRED TO THE MRI, AND THE MONITOR INTERMITTENTLY LOST SIGNALS (ECG, NIBP) DURING THE MRI PROCEDURE. THEY ATTEMPTED TROUBLESHOOTING AND SWITCHED THE NIBP CUFF TO THE OTHER ARM BUT CONTINUED TO EXPERIENCE INTERMITTENT SIGNALS. THE CUSTOMER ALSO REPORTED THAT THE DEVICE SCREEN WENT COMPLETELY BLANK FOR APPROXIMATELY 5 SECONDS. THE PATIENT WAS 87 YEARS OLD AND HAD PREEXISTING CONDITIONS (ENCEPHALOPATHY). Manufacturer Narrative: THE REPORTED EVENT INVOLVED THE PATIENT MONITOR S/N (B)(6). FOLLOWING THE INITIAL REPORT, IRADIMED CONTACTED THE BIOMEDICAL ENGINEER, WHO INDICATED THAT THE DEVICE HAD BEEN REMOVED FROM SERVICE FOR INVESTIGATION. THE CUSTOMER STATED THAT THEY TROUBLESHOT THE DEVICE, CONDUCTED FUNCTIONAL TESTS, BUT WERE UNABLE TO REPLICATE THE FAILURE MODE. BASED ON THE BIOMEDICAL ENGINEERING ASSESSMENT, THE DEVICE WAS OPERATING AS EXPECTED. THE DEVICE WAS RETURNED TO IRADIMED FOR EVALUATION AND A REVIEW OF OPERATIONAL DETAILS. IT WAS EVALUATED ON 01/16/2025, DURING WHICH A PHYSICAL INSPECTION AND FUNCTIONAL TESTING WERE PERFORMED. NO MALFUNCTIONS WERE IDENTIFIED. IRADIMED WAS UNABLE TO DUPLICATE THE REPORTED FAILURE MODE, AS THE EQUIPMENT OPERATED AS DESIGNED AND EXPECTED. THE AVAILABLE INFORMATION DOES NOT REASONABLY SUGGEST THAT THE DEVICE (3880) MALFUNCTIONED, AS NO EVIDENCE OF A MALFUNCTION WAS FOUND. THE CUSTOMER CONFIRMED THAT THE ICU PATIENT HAD A PREEXISTING CONDITION, "ALCOHOL-INDUCED ENCEPHALOPATHY," WHICH MIGHT HAVE CONTRIBUTED TO THE OUTCOME. THEREFORE, NO FURTHER INFORMATION IS AVAILABLE AT THIS TIME. IF THE USER PROVIDES ADDITIONAL DETAILS DURING THE INVESTIGATION, AN MDR SUPPLEMENT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233796&pc=DSQ | 2916596-2025-00429 | 2022-08-01 04:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿DAYS ALIVE OUT OF THE HOSPITAL AFTER PLACEMENT OF A CONTEMPORARY LEFT VENTRICULAR ASSIST DEVICE AND HEART TRANSPLANTATION¿ THAT HEARTMATE 3 (HM3) IMPLANT WAS ASSOCIATED WITH OUTCOMES SUCH AS HEART TRANSPLANT, PUMP DECOMMISSION, OR DEATH. THE HOSPITALIZATION BURDEN OF LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS COMPARED TO HEART TRANSPLANTATION (HT) WAS UNKNOWN SO A SINGLE CENTER, RETROSPECTIVE COHORT STUDY WAS CONDUCTED. ALL CONSECUTIVE PATIENTS GREATER THAN 18 YEARS OLD AND RECEIVING HM3 OR HT FROM AUGUST 2017 TO AUGUST 2022 WERE INCLUDED EXCEPT FOR THOSE WHO DID NOT SURVIVE THEIR INDEX HOSPITALIZATION. TO REDUCE CENSORING BIAS, PATIENTS BRIDGED TO HEART TRANSPLANTATION (BTT) OR DECOMMISSIONED LESS THAN 1 YEAR AFTER HM 3 PLACEMENT WERE ALSO EXCLUDED. DAYS ALIVE OUT OF THE HOSPITAL (DAOH) AT 1-YEAR WERE COLLECTED AFTER DISCHARGE AND COMPARED BETWEEN PATIENTS THAT RECEIVED HM 3 AND HT. OUTPATIENT ENCOUNTERS FOR INVASIVE PROCEDURES WERE CONSIDERED AS TIME SPENT IN THE HOSPITAL. OVERALL, 255 PATIENTS (AGE: 57, IQR: 48-64 YEARS, 27% FEMALE) COMPRISED THE STUDY COHORT. THESE PATIENTS SPENT 351, IQR: 335-358 DAOH AT 1 YEAR AFTER INDEX DISCHARGE. FIFTEEN PATIENTS CROSSED OVER TO HT AFTER 1 YEAR. A TOTAL OF 78 PATIENTS UNDERWENT HM3 IMPLANTATION AND 192 RECEIVED HT. DAOH AT 1 YEAR WERE HIGHER FOR PATIENTS WITH HM3 (357, IQR: 329-364 DAYS) COMPARED TO HT (350, IQR: 336-356 DAYS, P=0.0016). FINDINGS REMAINED SIMILAR AFTER EXCLUDING CROSSOVERS, HM3: 357, IQR: 329-364 DAYS VERSUS HT: 350, IQR: 337-356 DAYS, P < 0.01. LENGTH PER ENCOUNTER WAS SIMILAR FOR HM3 (3, IQR: 1-8 DAYS) VS. HT (1.5, IQR: 1-3 DAYS, P=0.16). OVER 95% OF THE DAYS WERE SPENT OUT OF THE HOSPITAL AFTER EITHER HM 3 OR HT DURING THE FIRST YEAR FOLLOWING DISCHARGE FROM THE INDEX HOSPITALIZATION. PATIENTS RECEIVING HM3 SPENT MORE DAYS OUT OF THE HOSPITAL IN COMPARISON TO HT AT 1 YEAR. Manufacturer Narrative: A1-A6: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. B3: DATE OF EVENT HAS BEEN ENTERED AS 01AUG2022 AS THE IMPLANT DATE OCCURRED BETWEEN (B)(6) 2022 B5: DEVICE WAS IMPLANTED AT TIME OF EVENT. D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. RAMOS, J. O., FLETCHER, R., ISLAM, R., MEHTA, A., FERRELL, B., DIEZ-LOPEZ, C., ROCHLANI, Y., MADAN, S., MURTHY, S., SHIN, J., VUKELIC, S., FOREST, S., SIMS, D., PATEL, S., JORDE, U., GOLDSTEIN, D., & SAEED, O. (2024). DAYS ALIVE OUT OF THE HOSPITAL AFTER PLACEMENT OF A CONTEMPORARY LEFT VENTRICULAR ASSIST DEVICE AND HEART TRANSPLANTATION. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S231-S232. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1109. MONTEFIORE MEDICAL CENTER, BRONX, NY; ALBERT EINSTEIN COLLEGE OF MEDICINE, ALBERT EINSTEIN COLLEGE OF MEDICINE, BRONX, NY; JACOBI MEDICAL CENTER/ ALBERT EINSTEIN COLLEGE OF MEDICINE, BRONX, NY; BELLVITGE UNIVERSITY HOSPITAL, BARCELONA, SPAIN. THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF WI 90979616). NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233798&pc=NPT | 2015691-2025-00470 | 2024-12-30 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-24 05:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Calcified | Dyspnea | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST, APPROXIMATELY 7 YEARS AND 3 MONTHS AFTER A TRANSFEMORAL TAVR PROCEDURE WITH A 26 MM SAPIEN 3 VALVE, AN ECHOCARDIOGRAM DEMONSTRATED SEVERE STENOSIS WITH A VALVE AREA INDEX OF 1.1CM2, MODERATE TO SEVERE AORTIC REGURGITATION, MODERATELY CALCIFIED LEAFLETS, AND A MEAN GRADIENT OF 50MMHG. THE PERCEIVED ROOT CAUSE OF THE STENOSIS AND REGURGITATION WAS SEVERE PROSTHETIC VALVE DEGENERATION. THE PATIENT PRESENTED WITH SHORTNESS OF BREATH WITH MINIMAL EXERTION AND INTERMITTENT BILATERAL LOWER EXTREMITY EDEMA. THE PATIENT WAS BEING WORKED UP FOR A VALVE IN VALVE PROCEDURE BUT PASSED AWAY PRIOR TO SCHEDULED TESTING. DESPITE MULTIPLE INVESTIGATIONAL ATTEMPTS, ADDITIONAL DETAILS REGARDING THE PATIENT'S DEATH (INCLUDING CAUSE OF DEATH AND DATE OF EVENT) REMAIN UNKNOWN DUE TO THE LIMITED INFORMATION PROVIDED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER A TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, CALCIFIC DEGENERATION IS A COMMON CAUSE OF BIOPROSTHETIC HEART VALVE FAILURES. MANY FACTORS CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION. THESE INCLUDE PATIENT FACTORS (AGE, DISEASE STATE, PHARMACOLOGICAL INTERVENTION, ETC.), MECHANICAL STRESS RELATED TO THE VALVE'S HEMODYNAMIC PERFORMANCE, AND GLUTARALDEHYDE FIXATION OF TISSUE. OF THESE, THE FIXATION PROCESS IS A RELATIVELY MINOR CONTRIBUTOR TO CALCIFICATION FOR EDWARDS' TISSUE VALVES DUE TO SUBSEQUENT ANTI-CALCIFICATION TREATMENTS DURING MANUFACTURING. THOUGH NUMEROUS STUDIES HAVE BEEN CONDUCTED ON PREVENTIVE CALCIFICATION STRATEGIES IN BIOPROSTHETIC HEART VALVES, THE CAUSES OF CALCIFICATION ARE NOT FULLY UNDERSTOOD AND THERE ARE STILL NO MECHANISMS OR MEDICAL THERAPIES WHICH FULLY PREVENT BIOPROSTHESES FROM CALCIFYING. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (HYPERTENSION AND CHRONIC KIDNEY DISEASE) MAY HAVE CONTRIBUTED TO THE CALCIFICATION. DUE TO LIMITED INFORMATION AVAILABLE, THE CAUSE OF DEATH AND RELATIONSHIP TO THE DEVICE IS UNKNOWN. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233828&pc=DSQ | 2916596-2025-00218 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED AT NIGHT ON (B)(6) 2025 DUE TO AN ACUTE CIRCULATORY SHOCK SITUATION. THE LOG FILES WERE RETRIEVED AND SHOWED A SUDDEN DECLINE IN FLOW. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED, AND NO OBVIOUS OUTFLOW GRAFT OCCLUSION WAS DETECTED. THE PATIENT WAS INTUBATED IN THE INTENSIVE CARE UNIT (ICU) WHERE THEY PASSED AWAY DUE TO HEMODYNAMIC PROBLEMS FOLLOWING THE CIRCULATORY SHOCK. THE PUMP WAS EXPLANTED, AND AN AUTOPSY WAS TO BE PERFORMED BUT RESULTS WOULD NOT BE PROVIDED. ADDITIONAL INFORMATION WAS PROVIDED THAT THE REASON FOR THE SUDDEN DECREASE IN FLOW WAS AN OUTFLOW GRAFT OBSTRUCTION (OFG). NO IMAGES WERE AVAILABLE. THE PATIENT PASSED AWAY ON
(B)(6) 2025. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: ANALYSIS OF THE SUBMITTED LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS. ALTHOUGH A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED, THE ACCOUNT REPORTED THEY WERE DUE TO AN OUTFLOW GRAFT OBSTRUCTION. THE REPORTED OUTFLOW GRAFT OBSTRUCTION COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6)2024 THROUGH (B)(6)2025. ESTIMATED FLOW RANGED FROM 3.7 - 5.1 LITERS PER MINUTE (LPM) FROM THE BEGINNING OF THE FILE THROUGH 03:56:45 ON (B)(6)2025. MULTIPLE LOW FLOW HAZARD ALARMS WERE CAPTURED BETWEEN 07:27:53 THROUGH 08:08:15 ON (B)(6)2025. ESTIMATED FLOW RANGED FROM 2.2 ¿ 2.4 LPM DURING THESE EVENTS. FOLLOWING THIS TIMESTAMP, A PARTIAL RECOVERY IN ESTIMATED FLOW WAS CAPTURED, RANGING FROM 2.8-3.2 LPM. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED. THE PATIENT REMAINS ONGOING ON HM3 LVAS, SERIAL NUMBER (B)(6). NO FURTHER RELATED EVENTS HAVE BEEN REPORTED AT THIS TIME. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. G, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", OUTLINES CONSIDERATIONS FOR PUMP PLACEMENT AND PROVIDES INSTRUCTIONS REGARDING THE PREPARATION, INSTALLATION, AND ORIENTATION OF THE SEALED OUTFLOW GRAFT. THIS SECTION UNDER "ATTACHING THE SEALED OUTFLOW GRAFT TO THE PUMP," INSTRUCTS THE USER TO VERIFY THAT THE GRAFT IS NOT TWISTED OR KINKED BY CHECKING THE POSITION OF THE BLACK LINE ON THE GRAFT ABOVE AND BELOW THE BEND RELIEF AND ENSURING THAT THE LINE IS STRAIGHT. SECTION 5 OF THE IFU, UNDER "PREIMPLANT PROCEDURES" AND "IMPLANT PROCEDURES" CAUTIONS THE USER: "THE SEALED OUTFLOW GRAFT MUST NOT BE KINKED OR POSITIONED WHERE IT COULD ABRADE AGAINST A PUMP COMPONENT OR BODY STRUCTURE" AND "STRETCH THE SEALED OUTFLOW GRAFT COMPLETELY PRIOR TO MEASURING AND CUTTING THE GRAFT TO THE APPROPRIATE LENGTH." SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21234171&pc=NVN | 2017865-2025-11489 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-24 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11482. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11488. FOLLOWING THE ELECTRONICS PERFORMANCE INDICATOR (EPI), AN ALERT WAS RECEIVED BY THE CLINICIAN AND AWAITING DEVICE EXPLANT. THE PATIENT PASSED AWAY FROM PNEUMONIA PRIOR TO ANY INTERVENTION TAKING PLACE FOR THE EPI. IT WAS NOT STATED IF PHYSICIAN BELIEVED THE INFECTION WAS RELATED TO ANY ABBOTT PRODUCTS/ PROCEDURES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) WAS REVIEWED; NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND ASSOCIATED WITH THIS SERIAL NUMBER. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21234370&pc=LWP | 2017865-2025-11482 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-24 05:00 | LWP | ASSURITY MRI | Premature Discharge of Battery; Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P140033 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11488. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11489. FOLLOWING THE ELECTRONICS PERFORMANCE INDICATOR (EPI), AN ALERT WAS RECEIVED BY THE CLINICIAN AND AWAITING DEVICE EXPLANT. THE PATIENT PASSED AWAY FROM PNEUMONIA PRIOR TO ANY INTERVENTION TAKING PLACE FOR THE EPI. IT WAS NOT STATED IF PHYSICIAN BELIEVED THE INFECTION WAS RELATED TO ANY ABBOTT PRODUCTS/ PROCEDURES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) WAS REVIEWED; NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND ASSOCIATED WITH THIS SERIAL NUMBER. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21234570&pc=NVN | 2017865-2025-11488 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-24 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11482. RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-11489. FOLLOWING THE ELECTRONICS PERFORMANCE INDICATOR (EPI), AN ALERT WAS RECEIVED BY THE CLINICIAN AND AWAITING DEVICE EXPLANT. THE PATIENT PASSED AWAY FROM PNEUMONIA PRIOR TO ANY INTERVENTION TAKING PLACE FOR THE EPI. IT WAS NOT STATED IF PHYSICIAN BELIEVED THE INFECTION WAS RELATED TO ANY ABBOTT PRODUCTS/ PROCEDURES. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) WAS REVIEWED; NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND ASSOCIATED WITH THIS SERIAL NUMBER. FURTHER INFORMATION WAS REQUESTED, BUT NOT YET AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21234849&pc=OZD | 1220648-2025-25785 | 2024-02-23 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP IN A 65 YEAR OLD MALE PATIENT FOR MECHANICAL CIRCULATORY SUPPORT FOR MYOCARDIAL INFARCTION. IT WAS REPORTED THAT UPON EXPLANTING THE PUMP, THE PATIENT HAD A HEMATOMA EVACUATION AND WASHOUT ON THEIR RIGHT GROIN. THE SITE WAS THEN SURGICALLY CLOSED. HOWEVER, FURTHER DETAILS STATES THAT DUE TO THE PATIENT'S RAPID DETERIORATION, THE PATIENT CONTINUED TO CODE AND A NEW PUMP WAS PLACED. THE PATIENT EVENTUALLY EXPIRED WHILE ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE CAUSE OF THE ACCESS SITE BLEEDING WAS NOT DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND NOT ENOUGH CLINICAL INFORMATION WAS PROVIDED. E.1 REVISED US PHONE NUMBER AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25785. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25785 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.1 REVISED REPORTING CONTACT FACILITY NAME AND REPORTING CONTACT FAX NUMBER AS THEY WERE INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25785. H.2 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25785 AS IT IS WAS AN INITIAL MANUFACTURER DEVICE REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21235590&pc=OZD | 1220648-2025-25786 | 2024-12-25 05:00 | Death | ABIOMED, INC. | 2025-01-24 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5, ALONG WITH ECMO AND PROTEK DUO, IN A 72 YEAR OLD FEMALE PATIENT FOR A BRIDGE TO RECOVERY FOR CARDIOMYOPATHY. THE TEAM OBSERVED A LEFT VENTRICLE (LV) CLOT AND WERE UNABLE TO INCREASE IMPELLA FLOWS. HOWEVER THE PUMP REMAINED ON FOR SUPPORT FOR OVER 32 DAYS. CARE WAS EVENTUALLY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED THROMBUS HAS BEEN COMPLETED. NEITHER THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THE ROOT CAUSE OF THROMBUS COULD NOT BE DETERMINED AS THE DEVICE WAS NOT RETURNED NOR SUFFICIENT CLINICAL DETAILS. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21235668&pc=DSQ | 2916596-2025-00528 | 2017-08-01 04:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "EFFECT OF PREOPERATIVE MITRAL REGURGITATION ON LVAD OUTCOMES IN PATIENTS WITH ELEVATED PULMONARY VASCULAR RESISTANCE" THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH GASTROINTESTINAL BLEEDING, STROKE, THE NEED FOR RIGHT VENTRICULAR ASSIST DEVICES (RVAD), AND DEATH. ADVERSE EVENT INCLUDED: THE DEATH OF 103 PATIENTS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE, AND DIVISION OF CARDIOTHORACIC TRANSPLANTATION AND CIRCULATORY SUPPORT, DEPARTMENT OF SURGERY, BAYLOR COLLEGE OF MEDICINE; DEPARTMENT OF CARDIOLOGY, THE TEXAS HEART INSTITUTE. KHERALLAH, R. Y., LAMBA, H. K., CIVITELLO, A. B., NAIR, A. P., SIMPSON, L., SHAFII, A. E., LOOR, G., GEORGE, J. K., DELGADO, R. M., LIAO, K. K., STAINBACK, R. F., FRAZIER, O. H., & KONERU, S. (2024). EFFECT OF PREOPERATIVE MITRAL REGURGITATION ON LVAD OUTCOMES IN PATIENTS WITH ELEVATED PULMONARY VASCULAR RESISTANCE. CARDIOVASCULAR DRUGS AND THERAPY. HTTPS://DOI.ORG/10.1007/S10557-024-07581-1. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT DEATHS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED "EFFECT OF PREOPERATIVE MITRAL REGURGITATION ON LVAD OUTCOMES IN PATIENTS WITH ELEVATED PULMONARY VASCULAR RESISTANCE" REPORTED THAT IN PATIENTS WITH END-STAGE HEART FAILURE WHO UNDERGO LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION, HIGHER PULMONARY VASCULAR RESISTANCE (PVR) IS ASSOCIATED WITH HIGHER RIGHT HEART FAILURE RATES AND INELIGIBILITY FOR HEART TRANSPLANT. CONCOMITANT MITRAL REGURGITATION (MR) COULD POTENTIALLY WORSEN PULMONARY HEMODYNAMICS AND LEAD TO WORSE OUTCOMES; HOWEVER, ITS EFFECTS IN THIS PATIENT POPULATION HAVE NOT BEEN SPECIFICALLY EXAMINED. USING AN INSTITUTIONAL DATABASE SPANNING NOVEMBER 2003 TO AUGUST 2017, THE STUDY RETROSPECTIVELY IDENTIFIED PATIENTS WITH ELEVATED PVR WHO UNDERWENT LVAD IMPLANTATION. OF 644 LVAD RECIPIENTS, 232 OF THESE PATIENTS MET THE CRITERIA FOR INCLUSION IN THE ANALYSIS OF WHICH TWO HAD HEARTMATE 3. THE ARTICLE NOTED ADVERSE EVENTS AND CLINICAL OUTCOMES WHICH INCLUDED DEATH (103 PATIENTS), RIGHT HEART FAILURE (59 PATIENTS), GASTROINTESTINAL BLEEDING (55 PATIENTS), CEREBROVASCULAR ACCIDENT (64 PATIENTS), HEART TRANSPLANT (53 PATIENTS), LVAD EXCHANGE (33 PATIENTS), AND LVAD EXPLANT (8 PATIENTS). ADDITIONALLY, 15 PATIENTS WERE LOST TO FOLLOW-UP. THE ARTICLE CONCLUDED THAT AMONG PATIENTS WITH ELEVATED PREOPERATIVE PVR, THOSE WITH MODERATE/SEVERE MR HAD BETTER OVERALL SURVIVAL AND HIGH TRANSPLANT RATES THAN THOSE WITH MILD/NO MR. IN PATIENTS WITH PREOPERATIVE ELEVATED PVR, MR SEVERITY MAY BE A PROGNOSTIC SIGN THAT CAN INFORM PATIENT SELECTION FOR END-STAGE HEART FAILURE THERAPY. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21235783&pc=NPT | 2015691-2025-00473 | 2020-01-29 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-01-24 05:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Perivalvular Leak; Malposition of Device |  | P140031 |  | 1 | Event Description: AS REPORTED BY IMPLANT PATIENT REGISTRY, AND CLARIFIED VIA MEDICAL RECORD, THE PATIENT PASSED AWAY 6 DAYS POST-IMPLANTATION OF A 29MM SAPIEN 3 (S3) VALVE IN THE RIGHT ATRIOVENTRICULAR VALVE POSITION. PER THE MEDICAL RECORD, THE PATIENT HAD BEEN HOSPITALIZED 3 WEEKS PRIOR TO IMPLANT OF S3 VALVE AND ON POST-OPERATIVE DAY 2, SHE WAS NOTED TO BE SIGNIFICANTLY HYPOTENSIVE WITH A NARROW PULSE PRESSURE AND IN A JUNCTIONAL TACHYCARDIA. HER CENTRAL VENOUS PRESSURE WAS IN THE 20'S. DWINDLE EPINEPHRINE WAS GIVEN AND THE PATIENT THEN DECOMPENSATED REQUIRING CARDIOPULMONARY RESUSCITATION. THE CHEST WAS THEN OPENED, AND CLOTS WERE REMOVED FROM THE CAVITY. SHE STABILIZED BUT THEN ARRESTED AGAIN AND WAS THEN CANNULATED ONTO VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). AN ECHO WAS PERFORMED AFTER CANNULATION AND THERE WAS A PERIVALVULAR LEAK IN THE S3 VALVE. ON POD 4, A REPEAT ECHO THEN SHOWED THE S3 VALVE TO BE MOBILE AND POORLY ANCHORED. THE PATIENT IMMEDIATELY WENT TO THE HYBRID LAB AND HAD THE VALVE BALLOONED. SHE TOLERATED THE PROCEDURE WELL BUT DID HAVE AN EPISODE OF VENTRICULAR FIBRILLATION/VENTRICULAR TACHYCARDIA WITH AIR NOTED IN THE ECMO CANNULAS. SHE RECEIVED CARDIAC MASSAGE AND DEFIBRILLATION. SHE STABILIZED AND WAS BROUGHT BACK TO THE CARDIOVASCULAR ICU. ON POD 5, HER SEDATION WAS LIGHTENED, THE PATIENT DID NOT WAKE UP AS EXPECTED AND WAS FOUND TO HAVE A MASSIVE LEFT HEMISPHERIC STROKE. SHE WAS THEN DECANNULATED FROM ECMO, THAT SHE TOLERATED REASONABLY WELL WITH MINIMAL INCREASE IN INOTROPIC SUPPORT AND SATURATIONS IN THE 80'S. GIVEN THE GRAVE PROGNOSIS OF INTRACRANIAL BLEED, THE FAMILY ELECTED TO MOVE TO COMFORT CARE. HER EPINEPHRINE WAS TURNED OFF, SHE WAS EXTUBATED AND SHE PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SENT AS ADMINISTRATIVE REVIEW SHOWED THERE WAS MISSING INFORMATION IN THE ADDITIONAL MANUFACTURER NARRATIVE. B4, G3, G6, AND H2 HAVE BEEN UPDATED. H11 HAS BEEN CORRECTED. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. PER THE INSTRUCTIONS FOR USE (IFU), HYPOTENSION AND CARDIOGENIC SHOCK ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PATIENTS UNDERGOING THE TVR PROCEDURE CAN BE NON-OPERATIVE OR HIGH RISK, HAVE COMPLEX MEDICAL HISTORIES AND HAVE MULTIPLE CO-MORBIDITIES. PATIENT RISK FACTORS FOR HYPOTENSION INCLUDE LOW EJECTION FRACTION, CORONARY ARTERY DISEASE, CONGESTIVE HEART FAILURE, ARTERIOSCLEROSIS, HYPOVOLEMIA, AND ANEMIA. ADDITIONALLY, THESE PATIENTS ARE ROUTINELY ADMINISTERED MULTIPLE VASOACTIVE DRUGS DURING THE PROCEDURE AND ARE INTENTIONALLY MADE HYPOTENSIVE, UTILIZING RAPID VENTRICULAR PACING, TO FACILITATE ACCURATE VALVE DEPLOYMENT. AS A RESULT OF THESE FACTORS, INTRA-OPERATIVE HYPOTENSION IS NOT UNCOMMON AND IS TREATED WITH STANDARD THERAPIES, INCLUDING VASOACTIVE DRUGS. IT IS ALSO NOT UNCOMMON TO INITIATE BRIEF CHEST COMPRESSIONS OR CARDIAC MASSAGE TO FACILITATE THE DISTRIBUTION OF THESE VASOACTIVE DRUGS. IN SOME CASES, THESE STANDARD MANEUVERS ARE NOT ADEQUATE, AND INITIATION OF CARDIOPULMONARY BYPASS, INSERTION OF INTRA-AORTIC BALLOON PUMP, AND/OR CONVERSION TO OPEN SURGERY IS REQUIRED. PERIODS OF PROLONGED HYPOTENSION CAN RESULT IN CARDIOGENIC SHOCK. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. AS A PRECAUTION, THE TRAINING MANUALS INSTRUCT THE OPERATOR TO MINIMIZE THE NUMBER AND DURATION OF RAPID BURST PACING EPISODES, AND ALLOW SUFFICIENT HEMODYNAMIC RECOVERY BEFORE INITIATING ANOTHER EPISODE OF RAPID PACING. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MALPOSITION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO MALPOSITION, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT LANDING ZONE CALCIFICATION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE VALVE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF PARAVALVULAR LEAKAGE, MALPOSITION OF THE THV AND HEMODYNAMIC INSTABILITY ENDING IN THE DEATH OF THE PATIENT IS UNKNOWN BUT MAY BE RELATED TO ONE OR MORE OF THE ABOVE-MENTIONED FACTORS (I.E. CONGENITAL HEART DEFECTS; SEE PATIENT HISTORY IN B7), AND/OR MECHANISMS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE TRICUSPID POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO TRICUSPID PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), HYPOTENSION AND CARDIOGENIC SHOCK ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PATIENTS UNDERGOING THE TVR PROCEDURE CAN BE NON-OPERATIVE OR HIGH RISK, HAVE COMPLEX MEDICAL HISTORIES AND HAVE MULTIPLE CO-MORBIDITIES. PATIENT RISK FACTORS FOR HYPOTENSION INCLUDE LOW EJECTION FRACTION, CORONARY ARTERY DISEASE, CONGESTIVE HEART FAILURE, ARTERIOSCLEROSIS, HYPOVOLEMIA, AND ANEMIA. ADDITIONALLY, THESE PATIENTS ARE ROUTINELY ADMINISTERED MULTIPLE VASOACTIVE DRUGS DURING THE PROCEDURE AND ARE INTENTIONALLY MADE HYPOTENSIVE, UTILIZING RAPID VENTRICULAR PACING, TO FACILITATE ACCURATE VALVE DEPLOYMENT. AS A RESULT OF THESE FACTORS, INTRA-OPERATIVE HYPOTENSION IS NOT UNCOMMON AND IS TREATED WITH STANDARD THERAPIES, INCLUDING VASOACTIVE DRUGS. IT IS ALSO NOT UNCOMMON TO INITIATE BRIEF CHEST COMPRESSIONS OR CARDIAC MASSAGE TO FACILITATE THE DISTRIBUTION OF THESE VASOACTIVE DRUGS. IN SOME CASES, THESE STANDARD MANEUVERS ARE NOT ADEQUATE, AND INITIATION OF CARDIOPULMONARY BYPASS, INSERTION OF INTRA-AORTIC BALLOON PUMP, AND/OR CONVERSION TO OPEN SURGERY IS REQUIRED. PERIODS OF PROLONGED HYPOTENSION CAN RESULT IN CARDIOGENIC SHOCK. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. AS A PRECAUTION, THE TRAINING MANUALS INSTRUCT THE OPERATOR TO MINIMIZE THE NUMBER AND DURATION OF RAPID BURST PACING EPISODES, AND ALLOW SUFFICIENT HEMODYNAMIC RECOVERY BEFORE INITIATING ANOTHER EPISODE OF RAPID PACING. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MALPOSITION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO MALPOSITION, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT LANDING ZONE CALCIFICATION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE VALVE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF PARAVALVULAR LEAKAGE, MALPOSITION OF THE THV AND HEMODYNAMIC INSTABILITY ENDING IN THE DEATH OF THE PATIENT IS UNKNOWN BUT MAY BE RELATED TO ONE OR MORE OF THE ABOVE-MENTIONED FACTORS (I.E. CONGENITAL HEART DEFECTS; SEE PATIENT HISTORY IN B7), AND/OR MECHANISMS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE TRICUSPID POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO TRICUSPID PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21235896&pc=DSQ | 2916596-2025-00321 | 2024-07-25 04:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Skin Infection; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE CUSTOMER REPORTED THAT NO ADDITIONAL INFORMATION COULD BE PROVIDED. THIS WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WOULD NOT BE RETURNED.Event Description: THE CAUSE OF DEATH WAS BEING SEPTIC DUE TO RECURRENT INFECTIONS AND MULTIPLE HOSPITALIZATIONS WITHOUT RESOLUTION. THE FAMILY THEN DECIDED TO TRANSITION TO COMFORT CARE. THE DEVICE OPERATED AS EXPECTED AND WAS NOT DEVICE OR THERAPY RELATED.Event Description: THE SOURCE OF THE INFECTIONS WAS A CHRONIC DRIVELINE EXIT SITE INFECTION. THE ONSET OF THESE INFECTIONS WAS UNABLE TO BE DETERMINED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21235909&pc=DSQ | 2916596-2025-00310 | 2025-01-07 05:00 | Death | THORATEC CORPORATION | 2025-01-24 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypoxia; Renal Failure; Tachycardia; Cardiogenic Shock; Respiratory Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK FROM RIGHT VENTRICULAR FAILURE. IT WAS STATED THAT THE DEATH WAS NOT DEVICE OR THERAPY RELATED AND WAS DUE TO THE CARDIOGENIC SHOCK FROM RIGHT VENTRICULAR FAILURE. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT RETURN. THE RIGHT HEART FAILURE EXISTED PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD). A DEVICE RELATED ISSUE DID NOT CAUSE/CONTRIBUTE TO RIGHT HEART FAILURE. THE PATIENT HAD SYMPTOMS OF HYPOXIC RESPIRATORY FAILURE, FLUID OVERLOAD WITH PULMONARY EDEMA AND ACUTE CHRONIC KIDNEY FAILURE, AND SUBSEQUENT VENTRICULAR TACHYCARDIA (VT)/ VENTRICULAR FIBRILLATION (VF) ARREST. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOLLOWING ANEMIA DUE TO UTERINE BLEEDING. THE PATIENT RECEIVED CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) INITIATION ON (B)(6) 2025, INTUBATION DUE TO WORSENING HEMODYNAMICS, HYPOXIA, ENCEPHALOPATHY, REFRACTORY VT/VF ARREST WITH ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) PER AMERICAN HEART ASSOCIATION (AHA) PROTOCOLS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED AND WOULD NOT BE PROVIDED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, CARDIAC ARRHYTHMIA, BLEEDING, AND DEATH. SECTION 6, "PATIENT CARE AND MANAGEMENT," LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. UNDER ¿ANTICOAGULATION¿, THIS SECTION ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214645&pc=BZD | 2518422-2025-006849 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Pulmonary Dysfunction; Vomiting; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214723&pc=BZD | 2518422-2025-006876 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE, TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214871&pc=BZD | 2518422-2025-006889 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING OF LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214939&pc=BZD | 2518422-2025-006922 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215012&pc=BZD | 2518422-2025-006912 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND MULTIPLE MYELOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. GENERAL INFORMATION TAB WAS UPDATED. BOX A, B, E, WAS UPDATED. BOX G, H WAS UPDATED. SECTION H6 PATIENT OUTCOME CODE, HEALTH IMPACT CODE WAS UPDATED. ADDITIONAL NARRATIVE WAS UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215017&pc=BZD | 2518422-2025-006920 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD) | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND COPD. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215100&pc=BZD | 2518422-2025-006958 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Myocardial Infarction; Cancer; Heart Failure/Congestive Heart Failure; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH, HEART ATTACK AND HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215103&pc=BZD | 2518422-2025-006991 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER CANCER, BLADDER CANCER, STOMACH CANCER, KIDNEY CANCER, DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215748&pc=BZD | 2518422-2025-007108 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH CANCER AND ESOPHAGEAL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215751&pc=BZD | 2518422-2025-007122 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215756&pc=BZD | 2518422-2025-007061 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, LUNG CANCER AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215800&pc=BZD | 2518422-2025-006957 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING OF CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215809&pc=BZD | 2518422-2025-007130 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215818&pc=BZD | 2518422-2025-007075 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pneumonia; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER DEATH, PANCREATIC CANCER AND PNEUMONIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21215908&pc=BZD | 2518422-2025-007181 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216056&pc=BZD | 2518422-2025-007115 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION RESPIRATORY TRACT IRRITATION LUNG DISEASE REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216058&pc=BZD | 2518422-2025-007183 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM-RECRT | Adverse Event Without Identified Device or Use Problem | Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION AUTO CPAP WITH AN ALLEGATION OF DEATH AND KIDNEY DISEASE/TOXICITY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216201&pc=BZD | 2518422-2025-007109 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216256&pc=BZD | 2518422-2025-006933 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216287&pc=BZD | 2518422-2025-005990 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND NSC LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216288&pc=BZD | 2518422-2025-007143 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS AND/OR HEADACHE, LUNG DISEASE, CANCER, DEATH AND CHRONIC MYELOID LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216353&pc=BZD | 2518422-2025-006861 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND BLADDER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216641&pc=BZD | 2518422-2025-007340 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216698&pc=BZD | 2518422-2025-101081 | 2025-01-21 05:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216797&pc=OZD | 1220648-2025-25720 | 2025-01-04 05:00 | Death | ABIOMED, INC. | 2025-01-23 05:00 | OZD | IMPELLA CP | Pumping Stopped | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND APPROXIMATELY 11 DAYS AFTER START OF THE SUPPORT THE PUMP STOPPED. SUBSEQUENT CONVERSATION WITH THE PHYSICIAN NOTED THE PUMP STOPPED DUE TO HIGH MOTOR CURRENT, AND THERE WAS NO ATTEMPT PERFORMED TO RESTART THE PUMP. THE AUTOMATED IMPELLA CONTROLLER WAS SHUT OFF, AND THE PUMP WAS PULLED BACK IN THE AORTA. THE PATIENT PASSED AWAY ONE DAY LATER REFERRED TO HIS CARDIAC DISEASE. HOWEVER, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: PATIENT-SPECIFIC INFORMATION IS UNAVAILABLE AT THE TIME OF THIS MDR WRITING; THEREFORE, RESPECTIVE FIELDS REFLECT "UNKNOWN" OR LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION FOR THE PUMP STOP HAS BEEN COMPLETED. THE RETURNED PRODUCT WAS INSPECTED AND AN ABNORMALLY LARGE PURGE GAP WAS SEEN AND THE BEARING WAS EXPOSED. THE DATA LOGS CONFIRMED THAT THE PUMP WAS RUN FOR 290 HOURS (APPROXIMATELY 12 DAYS) LONGER THAN THE DESIGN LIFE OF THE IMPELLA WHICH IS 8 DAYS. THE ROOT CAUSE OF THE PUMP STOP WAS DETERMINED TO BE DUE TO BEARING WEAR. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ D.9 DEVICE RETURN DATE/INFORMATION WAS ADDED. B.2 REVISED AS THERE WAS AN INCORRECT SELECTION ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25720. E.1 REVISED AS FACILITY NAME WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25720. G.1 REVISED MANUFACTURING SITE NAME AND ADDRESS SECTION WAS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25720. H.10 REVISED INSTRUCTIONS FOR USE AS THEY WERE REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25720.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216863&pc=BZD | 2518422-2025-007317 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX B: DATE OF DEATH WAS CORRECTED. BOX H: TYPE OF REPORTED COMPLAINT WAS CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21216997&pc=BZD | 2518422-2025-007381 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, KIDNEY DISEASE/TOXICITY AND LIVER DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217186&pc=DSQ | 2916596-2025-00563 | 2014-07-30 04:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Unexpected Shutdown | Ischemia; Tachycardia; Embolism/Embolus | P060040 |  | 1 | Event Description: THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE. THE PERFUSIONIST REPORTED THAT THE PUMP'S PERIODIC SPEED HAD DROPPED. ELEVATED POWER AND WIDE PULSE PRESSURE WERE OBSERVED. MINIMAL SYSTOLIC, BUT NO DIASTOLIC FLOW WAS OBSERVED ON ECHO. IT WAS ALSO REPORTED THAT AFTER EVERY BEAT, THE AORTIC VALVE WAS NOTED TO HAVE OPENED. TWO DAYS LATER THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR) BECAUSE OF A RED HEART ALARM, INDICATING THAT THE PUMP HAD STOPPED WORKING. THE PATIENT ARRIVED IN THE OR WITH THE PUMP ON, BUT WITH NO FLOW. THE PATIENT RECEIVED A PUMP EXCHANGE. THE PATIENT EXPIRED ON (B)(6) 2014. ACCORDING TO THE VAD COORDINATOR, THE PATIENT'S EXPIRATION WAS NOT DEVICE RELATED AND THE PUMP FUNCTIONED AS INTENDED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE INTERMACS REGISTRY STATING: TEE CONFIRMED THROMBUS WITH HER AORTIC ROOT SUBSEQUENT EMBOLISM TO BOWEL AND LOWER EXTREMITIES. MIN SYSTOLIC AND NO DIASTOLIC FLOW THRU PUMP. PUMP OFF EVENT ON (B)(6) 2015 AT 2AM. ADDITIONAL INFORMATION ALSO INCLUDED INDICATED THAT THE PATIENT EXPIRED ON (B)(6) 2014 DUE TO WITHDRAWAL OF SUPPORT, SPECIFY. Manufacturer Narrative: SECTION H10: CORRECTION. Manufacturer Narrative: A USER FACILITY MEDWATCH REPORT WAS RECEIVED FROM THE INTERMACS REGISTRY. THE USER FACILITY NUMBER WAS NOT PROVIDED. THE INTERMACS IDENTIFIER IS LISTED IN H10. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. H10: INTERMACS NYFL-100256(102244).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217204&pc=DSQ | 2916596-2025-00529 | 2014-10-06 04:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Mechanical Problem | Stroke/CVA; Hemolysis; Hematuria; Thrombosis/Thrombus | P060040 |  | 1 | Event Description: THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD). IT WAS REPORTED THAT THE PATIENT PRESENTED WITH ELEVATED LACTATE DEHYDROGENASE (LDH) LEVELS, HEMATURIA, AND POWER ELEVATIONS. BIVALIRUDIN WAS INITIATED AND THE PATIENT'S NUMBERS NORMALIZED; HOWEVER, WHEN BIVALIRUDIN WAS DISCONTINUED, LDH BEGAN TO RISE AND HEMATURIA OCCURRED. IT WAS REPORTED THAT THE SURGICAL PLACEMENT OF THE ORIGINAL PUMP WAS DIFFICULT. THE PUMP WAS SUBSEQUENTLY EXCHANGED AND THE HOSPITAL REPORTED THAT THEY PLUGGED THE OLD SITE USING A FELT PLUG AND RE-CORED THE HEART. A THROMBUS WAS SEEN ON THE OUTFLOW STATOR BLADES. ADDITIONAL INFORMATION PROVIDED TO THE MANUFACTURER INDICATED THAT THE HEART WAS RE-CORED BECAUSE THE HOSPITAL FELT THAT THE INFLOW WAS MALPOSITIONED AND WAS POTENTIALLY COMPROMISING FLOW. A DEVICE TRACKING FORM WAS RECEIVED INDICATING THAT THE PATIENT EXPIRED ON (B)(6) 2014 APPROXIMATELY 1 MONTH AFTER THE PUMP EXCHANGE. THE CAUSE OF DEATH WAS NOTED TO BE A MIDDLE CEREBRAL ARTERY (MCA) CEREBROVASCULAR ACCIDENT (CVA) DUE TO VAD THROMBUS. IT WAS REPORTED THAT ANTICOAGULATION LIKELY CONTRIBUTED TO THE STROKE. THE SECOND PUMP WAS REPORTED TO BE FUNCTIONING AS INTENDED AND NO DEVICE ISSUES WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE INTERMACS REGISTRY STATING: POWER SPIKES, DARK URINE LDH ELEVATED ADDITIONAL INFORMATION ALSO INCLUDED INDICATED THAT THE PATIENT EXPIRED ON (B)(6) 2014 DUE TO WITHDRAWAL OF SUPPORT. PATIENT HAD A CVA. DEVICE FUNCTION NORMAL: YES. Manufacturer Narrative: SECTION H10: CORRECTION. Manufacturer Narrative: G8: THIS REPORT IS BEING SUBMITTED AS A CORRECTION/FOLLOW UP TO MFR. #2916596-2014-01974. H10: CORRECTED TO INCLUDE INTERMACS MDUM-9795 (114992).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217229&pc=DSQ | 2916596-2025-00530 | 2015-08-31 04:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE II®, SYSTEM CONTROLLER | Device Difficult to Setup or Prepare | Insufficient Information | P060040 |  | 1 | Event Description: THE PATIENT WAS IMPLANTED WITH A LEFT VENTRICULAR ASSIST DEVICE (LVAD) ON (B)(6) 2013. EMERGENCY DEPARTMENT PERSONNEL FROM AN OUTLYING HOSPITAL NOTIFIED THE IMPLANTING CENTER VAD COORDINATOR THAT THE PATIENT WAS IN FULL CARDIAC ARREST. THEY REPORTED THAT ON THE EVENING OF (B)(6) 2015, A BACKUP BATTERY FAULT ALARM SOUNDED. THE PATIENT LIVES 3 HOURS AWAY FROM THE IMPLANTING CENTER. THE PATIENT AND THE CAREGIVER ATTEMPTED TO EXCHANGE THE SYSTEM CONTROLLER, AS THE MESSAGE DISPLAY INDICATED BACKUP BATTERY FAULT - REPLACE CONTROLLER; HOWEVER, THEY WERE UNABLE TO "MAKE THE CONNECTION". ONCE THE BACKUP SYSTEM CONTROLLER WAS ENGAGED, THE SAME ALARM AND MESSAGE FOR BACKUP BATTERY FAULT - REPLACE CONTROLLER OCCURRED WITH THE BACKUP SYSTEM CONTROLLER. IT WAS ALSO REPORTED THAT THIS PATIENT¿S AORTIC VALVE HAD BEEN SEWN CLOSED. THE PATIENT EXPIRED ON (B)(6) 2015. A USER FACILITY MEDWATCH REPORT WAS RECEIVED ON 10/08/2015 THAT STATES: "PATIENT AND WIFE WERE IN BED WHEN THE CONTROLLER BEGAN TO ALARM WITH A ¿BACKUP BATTERY FAULT¿ ALARM. WIFE SAID THE ALARM MESSAGE STATED TO REPLACE THE CONTROLLER. THEY ATTEMPTED TO CHANGE OUT THE CONTROLLER BUT WERE NOT ABLE TO COMPLETE THE EXCHANGE. WIFE STATED SHE WENT TO CALL THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATORS AND THEN HEARD PATIENT FALL TO THE GROUND. PATIENT WAS UNRESPONSIVE WHEN SHE GOT BACK TO PATIENT AND SHE CALLED EMERGENCY MEDICAL SERVICES. IT WAS UNCLEAR AT THE TIME OF THE CALL AS TO WHAT WAS PREVENTING A SUCCESSFUL CONTROLLER EXCHANGE. THE MANUFACTURER ISSUED AN URGENT MEDICAL DEVICE CORRECTION NOTIFICATION ON SEPTEMBER 14, 2015. DEVICE USAGE PROBLEM: DEVICE WAS HARD TO USE." Manufacturer Narrative: SECTION H10: CORRECTION. Manufacturer Narrative: G8: THIS REPORT IS BEING SUBMITTED AS A CORRECTION/FOLLOW UP TO MFR. #2916596-2015-01676. H10: CORRECTED TO INCLUDE (B)(4). NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217604&pc=BZD | 2518422-2025-007406 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217606&pc=BZD | 2518422-2025-007417 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Bronchitis; Dyspnea; Pleural Effusion; Pneumonia; Fibrosis; Cancer; Heart Failure/Congestive Heart Failure; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, DEATH, EAR INFECTION/INFLAMMATION, SHORTNESS OF BREATH, CHRONIC SINUS INFECTIONS, HEART FAILURE, CHRONIC PNEUMONIA, BRONCHITIS, PLEURAL EFFUSION, PULMONARY FIBROSIS AND REACTIVE AIRWAY DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217812&pc=MVK | 3008642652-2025-00833 | 2024-12-31 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-23 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing; Failure to Convert Rhythm; Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SEVEN APPROPRIATE TREATMENTS AND TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:09:00 ON (B)(6) 2024. AT 17:04:31 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 90 BPM DEGRADING TO VT @ 220 BPM. AT 17:09:05, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH PAC¿S AND PVC¿S. AT 17:09:32, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS NINE SECONDS OF ASYSTOLE TRANSITIONING TO SINUS BRADYCARDIA @ 25 BPM. AT 17:10:41, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 30 BPM. AT 17:11:51, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 20 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM. AT 17:13:19, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 260 BPM WITH CPR/MOTION ARTIFACT. AT 17:14:51, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH CPR/MOTION ARTIFACT. AT 17:16:10, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 17:16:41, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 17:17:06, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 250 BPM. AT 17:17:37, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 60 BPM DEGRADING TO VT @ 220 BPM/VF. AT 17:18:12, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 190 BPM/VF. AT 17:18:48, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. THE ELECTRODE BELT WAS DISCONNECTED AT 18:10:35 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217823&pc=MVK | 3008642652-2025-00834 | 2024-10-12 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-23 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 12:23:03 ON (B)(6) 2024. AT 20:07:33, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 150 BPM WITH CPR/ELECTRODE LEAD FALL OFF. AT 20:08:43, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 20:21:34 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217829&pc=MVK | 3008642652-2025-00835 | 2024-12-06 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-23 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 12:11:15 ON (B)(6) 2024. AT 16:04:59, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 240 BPM WITH CPR/MOTION ARTIFACT. AT 16:05:34, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS VF WITH MOTION ARTIFACT. AT 16:06:04, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE TRANSITIONING TO SINUS BRADYCARDIA @ 50 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 16:20:36, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. AT 16:21:55, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 16:22:22 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217837&pc=MVK | 3008642652-2025-00836 | 2024-12-02 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-23 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 11:36:09 ON (B)(6) 2024. AT 08:31:54 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF @ 120 BPM TRANSITIONING TO VT @ 200 BPM WITH MOTION ARTIFACT. AT 08:32:33, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 180 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 180 BPM WITH MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 09:46:40 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218001&pc=NVZ | 3008973940-2025-00637 | 2024-12-08 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-23 05:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Cardiac Arrest; Tachycardia; Syncope/Fainting | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH AN IMPLANTABLE PULSE GENERATOR (IPG) COLLAPSED ON THE FLOOR. CHEST COMPRESSIONS WERE PERFORMED BY THE CALLER UNTIL EMERGENCY MEDICAL TECHNICIANS ARRIVED. DURING INTERROGATION AT THE HOSPITAL EMERGENCY ROOM, A GLITCH WAS IDENTIFIED IN THE IPG. THE PATIENT ULTIMATELY PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS CARDIAC ARREST FROM VENTRICULAR TACHYCARDIA (VT) RECORDED ON THE IPG. THE DEATH WAS NOT BELIEVED TO BE LINKED TO THE IPG PERFORMANCE AND THE DEVICE WAS NOTED TO BE FUNCTIONING NORMALLY DURING HOSPITALIZATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218038&pc=OZD | 1220648-2025-25727 | 2025-01-08 05:00 | Death | ABIOMED, INC. | 2025-01-23 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding; Renal Failure | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD ACCESS SITE BLEEDING. PRESSURE WAS HELD AND ANGLE MATCHING WAS PERFORMED. HEPARIN WAS WITHHELD UNTIL LABS WERE RAN. ADDITIONALLY, WHILE ON SUPPORT, THE PATIENT HAD HEMOLYSIS. DIALYSIS WAS STARTED AND THE PATIENT BECAME HYPOTENSIVE SO CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED. THE PATIENT DID HAVE ACUTE KIDNEY INJURY PRIOR TO IMPELLA SUPPORT, HOWEVER IT IS UNKNOWN IF THE IMPELLA CONTRIBUTED TO THE PATIENT'S KIDNEY ISSUES. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED AFTER 68.47 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE INVESTIGATION FOR THE HEMOLYSIS, ACCESS SITE BLEED, RENAL FAILURE HAS BEEN COMPLETED. PUMP WAS NOT RETURNED FOR INVESTIGATION. AS PER THE CLINICAL INFORMATION PROVIDED, HEMATOMA AND ACCESS SITE WAS BLEEDING ON DAY 2 AND WAS ADDRESSED WITH MANUAL DRESSING AND ANGLE MATCHING. NO OTHER INFORMATION WAS PROVIDED ABOUT ACTS AND HEPARIN WAS WITHHELD. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. DATA LOGS WERE INVESTIGATED AND AROUND THE TIME OF REPORTED FAILURE THERE WERE NO POSITIONING ISSUES, SUCTION ALARMS, MOTOR CURRENT ISSUES, OR PLACEMENT SIGNAL TRENDS. CLINICAL MENTIONED THAT THE PUMP WAS SEATED DEEP AND REPOSITIONED AT 4.1CM MARK BUT DATA LOGS DON'T SUPPORT THAT. THEREFORE, THE ROOT CAUSE OF THE HEMOLYSIS ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED. AS PER THE CLINICAL INFORMATION PROVIDED, THE PATIENT HAS AN ACUTE KIDNEY INJURY PRIOR TO IMPELLA SUPPORT. ON DAY 3 OF IMPELLA SUPPORT WHEN PATIENT WAS GOING THROUGH DIALYSIS, THE PATIENT WAS SWITCHED TO CRRT DUE TO PATIENT BEING HYPOTENSIVE. IT IS UNCLEAR IF THE IMPELLA PUMP CONTRIBUTED TO THE RENAL ISSUES. THEREFORE, THE ROOT CAUSE OF THE RENAL ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): LIMB¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ IMPELLA CP WITH SMARTASSIST: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218081&pc=PYX | 1220648-2025-25728 | 2024-12-27 05:00 | Death | ABIOMED, INC. | 2025-01-23 05:00 | PYX | IMPELLA RP FLEX | Increase in Pressure | Cardiac Arrest; Cardiovascular Insufficiency | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DEVELOPED HIGH PURGE PRESSURE. THE PATIENT HAD A CENTRAL RIGHT VENTRICULAR ASSIST DEVICE (RVAD) FOR FOUR DAYS, AND WAS TAKEN FOR RVAD DECANNULATION TWO DAYS PRIOR TO THE IMPELLA RP FLEX IMPLANT. DURING THIS PROCEDURE LARGE AMOUNT OF CLOTS IN CHEST WAS NOTED, DECANNULATION WAS ABORTED AND THE CHEST WAS LEFT OPEN. THE RVAD DECANNULATION AND IMPELLA RP FLEX IMPLANT WAS COMPLETED TWO DAYS LATER. TWO DAYS AFTER START OF IMPELLA MCS, IT WAS NOTED PATIENT WAS RESTING ON THE IMPELLA RP FLEX SUPPORT WHILE ASSESSED FOR RIGHT VENTRICULAR RECOVERY POST DURABLE LEFT VENTRICULAR ASSIST DEVICE PROCEDURE. NOW 16 DAYS AFTER START OF IMPELLA MCS, HIGH PURGE PRESSURE AND PURGE SYSTEM BLOCKED ALARMS WERE ENCOUNTERED. RESOLUTION AND ACTION TAKEN ARE UNCLEAR. APPROXIMATELY ELEVEN DAYS LATER THE PATIENT EXPIRED. CARE WAS WITHDRAWN; HOWEVER, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE EVENT WITH IMPELLA RP FLEX AS AN ASSOCIATED FACTOR WITH THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR ANALYSIS. THE LOGS SHOWED NO MOTOR CURRENT SPIKE BEFORE HIGH PURGE PRESSURE ALARMS. THE LOGS SHOWED A GRADUAL RISE IN PURGE PRESSURE BEFORE ALARMS WERE TRIGGERED. THERE WAS A MOTOR CURRENT SPIKE RIGHT BEFORE HIGH PURGE PRESSURE WAS RESOLVED LIKELY AS A RESULT OF BIOMATERIAL BEING PUSHED OFF THE IMPELLER. IT IS UNKNOWN IF TISSUE PLASMINOGEN WAS ADDED TO THE PURGE. THE ROOT CAUSE OF HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL E.1 FACILITY NAME AND US PHONE NUMBER WERE INADVERTENTLY OMITTED FROM INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25728 AND WERE ADDED. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25728 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4614 AND 1762 WERE REPORTED INCORRECTLY ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25728. A NEW CODE WAS ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218136&pc=FGE | 2183870-2025-00040 | 2024-12-19 05:00 | Death | COVIDIEN | 2025-01-23 05:00 | FGE | PROTEGE GPS SELF-EXPANDING STENT SYSTEM 035 STANDARD | Fracture; Activation, Positioning or Separation Problem; Detachment of Device or Device Component | Unspecified Tissue Injury; No Clinical Signs, Symptoms or Conditions | K032580 |  | 1 | Event Description: IT WAS REPORTED WHILE ATTEMPTING TO DEPLOY THE PROTEGE GPS STENT, THE CATHETER BECAME LODGED WITH THE STENT PARTIALLY DEPLOYED. WHILE TRYING TO COMPLETELY DEPLOY AND DETACH FROM THE STENT, THE CATHETER FRACTURED LEAVING THE DISTAL END OF THE CATHETER LODGED IN THE PATIENT. ATTEMPTS TO REMOVE THE RETAINED PORTION OF THE CATHETER FROM THE PATIENT WERE NOT SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: WHEN THE DETACHMENT OCCURRED, THE DISTAL END OF STENT WAS PARTIALLY ACROSS THE TARGET LESION. THE FRACTURE OCCURRED APPROX. 5 MM FROM DISTAL TIP OF CATHETER. THERE WAS PATIENT INJURY, UNCLEAR OF THE CAUSE OF THE INJURY AND THE PATIENT EXPIRED WITHIN ABOUT 24 HOURS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW 31 PROCEDURAL IMAGES WERE RETURNED THE FIRST 16 IMAGES SHOW INSERTION OF A GUIDE CATHETER AND GUIDEWIRE TO THE RIGHT SIDE OF THE CEREBRAL VENOUS SINUS. IMAGES 17 18 SHOW A STENT BEING DEPLOYED IN THE CEREBRAL VENOUS SINUS. IMAGE 19 20 SHOW THE STENT WITH PARTIAL BLOOD FLOW. IMAGES 21 22 SHOWS A PARTIALLY DEPLOYED STENT AND THE TIP OF THE CATHETER WITH SNARE. IMAGE 23 SHOWS MULTIPLE ATTEMPTS TO SNARE THE TIP OF THE CATHETER WITH THE PARTIALLY DEPLOYED STENT ALONG WITH NUMEROUS ATTEMPTS TO SNAG THE STENT WITH MANIPULATION OF THE GUIDEWIRE. IMAGES 24 25 SHOW THE GUIDEWIRE THROUGH THE STENT. IMAGE 26 SHOWS THE STENT PULLED PROXIMALLY IN THE VEIN. IMAGES 27 29 SHOW THE REPOSITIONED STENT AND CATHETER TIP. IMAGE 30 SHOWS A SNARE ATTEMPT OF THE STENT AND DETACHED TIP. IMAGE 31 SHOWS A SNARE ATTEMPT AND THE STENT FULLY DEPLOYED WITH THE DETACHED TIP STILL ATTACHED TO THE PROXIMAL END. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218226&pc=MNS | 2518422-2025-101102 | 2025-01-21 05:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION THE DEVICE WAS OPENED AND PUT BACK TOGETHER. AN ALARM #93 PRESENTED AFTER REPLACING THE BLOWER. THE DEVICE WAS OPENED BACK UP AND THE BLOWER WAS RESEATED. THE DEVICE IS OPERATIONAL. FILTERS, TUBING AND MANUAL REPLACED. DEVICE TESTED AND ALL TESTING PASSED. THE DEVICE PUT BACK INTO LOAN STOCK. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218372&pc=DSQ | 3007042319-2025-00095 | 2024-10-12 04:00 | Death | HEARTWARE, INC. | 2025-01-23 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Use of Device Problem; Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CONGESTIVE HEART FAILURE WAS FOUND DECEASED BY FAMILY MEMBERS WITH BOTH BATTERIES DISCONNECTED FROM THE CONTROLLER OF A VENTRICULAR ASSIST DEVICE (VAD). THIS DISCONNECTION CAUSED A VAD STOP. THERE WERE NO ALLEGATIONS AGAINST THE VAD OR ITS ACCESSORIES, AND THE FAMILY, ALONG WITH THE VENTRICULAR ASSIST DEVICE TEAM, DID NOT SUSPECT A MALFUNCTION O F THE VAD OR PERIPHERAL EQUIPMENT. NO TREATMENT OR INTERVENTION WAS ADMINISTERED AS THE PATIENT WAS FOUND DECEASED. NEITHER THE VAD NOR THE CONTROLLER OR BATTERIES WERE SENT BACK FOR ANALYSIS. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: 31-OCT-2018 / SERIAL OR LOT#: (B)(6), UDI #: (B)(4). D9: NO. H3: NO. H4: MFG DATE: 31-OCT-2017. H5: NO. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: 31-OCT-2018 / SERIAL OR LOT#: (B)(6), UDI #: (B)(4), D9: NO, H3: NO, H4: MFG DATE: 31-OCT-2017, H5: NO, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION. REVIEW OF LOG FILES COULD NOT BE PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED. A POSSIBLE ROOT CAUSE OF THE REPORTED VAD STOPPED EVENT CAN BE ATTRIBUTED, BUT NOT LIMITED, TO THE REPORTED DISCONNECTION OF BOTH POWER SOURCES FROM THE CONTROLLER AS DESCRIBED IN THE EVENT DETAILS. BASED ON THE LIMITED AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF STROKE. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCTS: CONTROLLER 2.0 (B)(6) H3: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218573&pc=DSQ | 2916596-2025-00527 | 2024-12-09 05:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia; Necrosis; Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO SEPTIC SHOCK FROM BACTEREMIA. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO SEPTIC SHOCK FROM BACTEREMIA AND GANGRENE. THE DEVICE OPERATED AS EXPECTED.Event Description: LEFT LOWER EXTREMITY ISCHEMIA WAS NOTED ON 09DEC2024 WHICH DEVELOPED INTO DRY GANGRENE ON 16DEC2024. SEVERE SEPTIC SHOCK WAS NOTED ON 06JAN2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, A SPECIFIC CAUSE FOR THE REPORTED INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS SEPSIS, LOCALIZED INFECTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218714&pc=LJS | 3006260740-2025-00089 | 2025-01-01 05:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-01-23 05:00 | LJS | GROSHONG NXT CLEARVUE 4F SINGLE-LUMEN PICC BASIC TRAY (WITH MICROINTRODUCER) | Fluid/Blood Leak | Low Blood Pressure/ Hypotension | K034020 |  | 1 | Event Description: IT WAS REPORTED, THE UNIT WAS USING 7655405 WHEN IT WAS SUDDENLY NOTICED THAT THE CRITICALLY ILL PATIENT'S BLOOD PRESSURE DROPPED. NOREPINEPHRINE DID NOT WORK, AND LATER A CATHETER WAS FOUND TO BE CRACKED AND THE SHEETS WERE WET. THE EXTENSION TUBE IN THE BODY IS CRACKED. RESUSCITATION OF THE PATIENT WAS STARTED IMMEDIATELY, BUT ULTIMATELY FAILED. DURATION OF CATHETERIZATION: ABOUT TEN DAYS TO ONE MONTH, IN ADDITION, MAINTENANCE EVERY 5-7 DAYS. ADDITIONAL INFORMATION PROVIDED: IT WAS REPORTED, THE CLIENT DESCRIBED TO A BD EMPLOYEE THAT THE PATIENT DIED BECAUSE THE RESUSCITATION FAILED AFTER BLOOD PRESSURE DROPPED. THE TIME BETWEEN THE PATIENT'S DROP IN BLOOD PRESSURE AND THE DISCOVERY OF THE LEAK IS UNKNOWN. THE PATIENT'S BLOOD PRESSURE WAS STABLE UNTIL IT DROPPED. A CVC WAS TO BE PLACED DURING RESUSCITATION, BUT THE DOCTOR STATED THE PATIENT COULD NOT BE SAVED AND IT WAS TOO LATE TO GET A CVC. THERE ARE NO OTHER SUSPICIOUS FACTORS THAT COULD CAUSE A CATHETER RUPTURE. THE PICC HAD BEEN IN PLACE FOR ABOUT 1 MONTH. NO OTHER INFORMATION WAS PROVIDED.Event Description: CAUSE OF DEATH WAS REQUESTED BUT NOT PROVIDED. OFFICIAL MEDICAL RECORDS AND DOCTOR'S ORDERS ARE NOT AVAILABLE. CLARIFICATION ON WHETHER THE PATIENT WAS ON A NOREPINEPHRINE DRIP OR IF THEY RECEIVED NOREPINEPHRINE AS PART OF RESUSCITATION EFFORTS WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE MANUFACTURER HAS RECEIVED THE SAMPLE AND WILL BE EVALUATED. A SUPPLEMENTAL WILL BE SUBMITTED WITH EVALUATION RESULTS. Manufacturer Narrative: H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE INITIAL COMPLAINT WAS SUBMITTED WITH (B)(6) 2024 AS THE DATE OF AWARENESS. AFTER FURTHER REVIEW, IT WAS DETERMINED THE DATE OF AWARENESS FOR THIS EVENT IS 07 JANUARY 2025. THE COMPLAINT OF A DAMAGED CATHETER WAS CONFIRMED, AND THE CAUSE APPEARED TO BE RELATED TO PRODUCT USE CONDITIONS. THE PRODUCT RETURNED FOR EVALUATION WAS ONE 4FR SINGLE LUMEN GROSHONG NXT CLEARVUE CATHETER. USAGE RESIDUES WERE OBSERVED THROUGHOUT THE SAMPLE. THE CATHETER WAS ASSEMBLED WITH A TWO-PIECE CONNECTOR. THE CATHETER WAS CUT APPROXIMATELY 3CM FROM THE ASSEMBLED CONNECTOR. AN S-SHAPED SPLIT WAS OBSERVED IN THE EXTENSION TUBING. AN ATTEMPT TO INFUSE WATER THROUGH THE SAMPLE USING A 12ML SYRINGE REVEALED THE SAMPLE TO BE PATENT TO INFUSION; HOWEVER, LEAKAGE WAS OBSERVED AT THE SPLIT SITE. THE CATHETER WAS PATENT DISTAL OF THE SPLIT. TACTILE INSPECTION OF THE SAMPLE REVEALED TENSILE WEAKNESS AT THE SPLIT SITE. MICROSCOPIC INSPECTION OF THE SPLIT REVEALED SHARPLY DEFINED, GRANULAR FRACTURE SURFACES. THE COARSENESS OF THE GRANULARITY VARIED ALONG THE LENGTH OF THE SPLIT, WITH THE FINEST TEXTURES OCCURRING AT THE ENDS OF THE FRACTURE. MULTIPLE SUPERFICIAL STRESS FRACTURES WERE OBSERVED IN THE VICINITY OF THE SPLIT. THE SPLIT FEATURES, STRESS FRACTURES AND TENSILE WEAKNESS WERE CONSISTENT WITH CATHETER DAMAGE CAUSED BY OVER-PRESSURIZATION OF THE CATHETER. THIS MAY OCCUR IF THE CATHETER IS FLUSHED AGAINST AN OCCLUSION OR A KINK. IT CAN ALSO OCCUR IF FLUSHING IS PERFORMED USING A SYRINGE SMALLER THAN 10ML. THE ENTIRE CATHETER A REVIEW OF SIMILAR COMPLAINTS BASED ON THE PRODUCT BATCH NUMBER WAS PERFORMED AND ONE ADDITIONAL COMPLAINT EXHIBITING SIMILAR CHARACTERISTICS WAS FOUND. A REVIEW OF THE MANUFACTURE QUALITY AND INSPECTION RECORDS FOR THE IMPLICATED PRODUCT BATCH INDICATED NO ISSUES POTENTIALLY RELATED TO PRODUCT MANUFACTURE, ASSEMBLY, AND PACKAGING. MANUFACTURING CONTROLS ARE IN PLACE TO MITIGATE THE OCCURRENCE OF THIS TYPE OF FAILURE. THOSE INSPECTION PROCEDURES ARE DOCUMENTED IN (B)(4). THE OBSERVED CATHETER DAMAGE WAS LIKELY CAUSED BY OVER-PRESSURIZATION OF THE CATHETER DURING PRODUCT USE. IT IS CURRENTLY UNKNOWN TO WHAT EXTENT CATHETER DAMAGE CONTRIBUTED TO THE REPORTED DROP IN THE PATIENT'S BLOOD PRESSURE AND ULTIMATELY TO DECLINING PATIENT CONDITIONS AND THE FAILURE OF RESUSCITATION EFFORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218727&pc=MNS | 2518422-2025-101109 | 2025-01-21 05:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE TOTAL BLOWER HOURS WERE 5834.4 WITH NO ERRORS OR REPORTED FAULTS. THE DEVICE PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218869&pc=GDW | 1219930-2025-00414 | 2024-12-26 05:00 | Death | COVIDIEN LP LLC NORTH HAVEN | 2025-01-23 05:00 | GDW | GIA | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Fistula; Sepsis; Peritonitis; Organ Dehiscence | K111825 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE OF DEATH CURRENTLY UNKNOWN.Event Description: IT WAS REPORTED THAT DURING A HEMICOLECTOMY WITH LATEROLATERAL ANASTOMOSIS PROCEDURE, THE DEVICE PERFORMED AS USUAL. POSTOPERATIVELY, THE STAPLE LINE DID NOT HOLD OR OPENED UP (DEHISCENCE). THE PATIENT EXPERIENCED EXTENDED HOSPITALIZATION FOR 8 DAYS AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. PLANS TO RE-OPERATE ON THE PATIENT WERE MADE, BUT IT WAS NOT POSSIBLE. THE PATIENT OUTCOME WAS DEATH, WITH THE CAUSE AS FISTULA IN THE 6DPO, WITH REAPPROACH FOR FECAL PERITONITIS, PROGRESSING WITH SEPTICEMIA AND DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218882&pc=MNS | 2518422-2025-101113 | 2023-09-01 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | MNS | REP DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION ST30 DEVICE PASSED AWAY IN (B)(6) 2023, AND THE PATIENT'S SON WANTED TO RETURN THE REPLACEMENT DEVICE HE RECEIVED FOR HIS FATHER AS IT IS NOT NEEDED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21219479&pc=DSQ | 2916596-2025-00513 | 2024-09-30 04:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21219623&pc=NPT | 2025587-2025-00569 | 2024-12-31 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-23 05:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 YEARS AND 5 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNSPECIFIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21219624&pc=MKJ | 2112020-2025-00044 | 2025-01-13 05:00 | Death | CARDIAC SCIENCE CORPORATION | 2025-01-23 05:00 | MKJ | G3 POWERHEART AED, BRITISH | Failure to Power Up | Cardiac Arrest | P160033 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 31-YEAR-OLD MALE PATIENT, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS VERIFIED DURING EVALUATION. A REVIEW OF THE DEVICE LOG SHOWS THE DEVICE WAS NOTIFYING A THE USER TO A DEPLETED BATTERY A MONTH PRIOR TO THE DEVICE BEING DEPLOYED ON A PATIENT THAT WAS NOT ADDRESSED. THE G3 OPERATOR AND SERVICE MANUAL (70-00914-01) STATES UNDER DAILY MAINTENANCE, "CHECK THE STATUS INDICATOR TO ENSURE THAT IT IS GREEN. WHEN THE INDICATOR IS GREEN, THE AED IS READY FOR A RESCUE." IT IS SUGGESTED THE USER CHECK THE NVI DAILY AND PRIOR TO CLINICAL USE. THE DEVICE WAS SCRAPPED AND THE CUSTOMER WAS SENT A REPLACEMENT DEVICE. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21219627&pc=FKX | 0008030665-2025-00189 | 2025-01-02 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-23 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue | K222318 |  | 1 | Event Description: ON 13/JAN/2025, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. DURING FOLLOW-UP, THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT WAS HOSPITALIZED ON (B)(6) 2025 DUE TO SHORTNESS OF BREATH AND WEAKNESS. ALTHOUGH MUCH OF THE PATIENT¿S HOSPITAL COURSE IS UNKNOWN (RECORD CLOSED), IT APPEARS THE PATIENT UNDERWENT A CARDIAC CATHETERIZATION, WHICH REVEALED THE NEED FOR THE PATIENT TO UNDERGO CORONARY ARTERY BYPASS SURGERY. THE PATIENT WAS TOLD BY THE CARDIOLOGIST AND SURGEON THAT SHE WAS NOT A CANDIDATE FOR SURGERY DUE TO HER MULTIPLE COMORBIDITIES (SPECIFICS NOT PROVIDED) AND INSTEAD ELECTED TO ENTER HOSPICE CARE ON (B)(6) 2025. THE PDRN CONFIRMED THE PATIENT PASSED AWAY PEACEFULLY AT HOME WITH FAMILY PRESENT ON (B)(6) 2025. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED ON THE MORNING OF (B)(6) 2025 (PRIOR TO DISCHARGE) WITHOUT ANY REPORTED ISSUES. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME FROM THE PATIENT¿S RESIDENCE, AND NO AUTOPSY WAS PERFORMED. THE HPM REPORTED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS WITHDRAWAL FROM DIALYSIS/UREMIA (CODE 104). ADDITIONALLY, THE PDRN REPORTED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP DOES NOT EXIST BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH, AS THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 9 DAYS PRIOR TO THE PATIENT¿S EXPIRATION. THE PDRN REPORTED THAT THE PATIENT WITHDREW FROM CCPD THERAPY ON (B)(6) 2025 AND ENTERED HOSPICE CARE LATER THE SAME DAY. THE PATIENT EXPIRED ON (B)(6) 2025 AS A RESULT FROM WITHDRAWING FROM CARE. HOSPICE CARE IS DEFINED AS A MEDICAL INTERVENTION, WITH THE EXPECTED OUTCOME BEING THE PATIENT WILL EXPIRE AS A RESULT. PER THE HPM, THE PATIENT¿S DEATH WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. ADDITIONALLY, THERE WAS NO REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21220257&pc=DSQ | 2916596-2025-00485 | 2025-01-15 05:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223184&pc=MVK | 3008642652-2025-00861 | 2024-11-04 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-23 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 21:36:46 ON (B)(6) 2024. AT 21:35:29 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 40 BPM WITH MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL ON THE SS CHANNEL. AT 21:36:46, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS BRADYCARDIA @ 40 BPM WITH MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL ON THE SS CHANNEL. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL ON THE SS CHANNEL. THE ELECTRODE BELT WAS DISCONNECTED AT 21:42:10 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223188&pc=MVK | 3008642652-2025-00862 | 2024-12-11 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-23 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 15:43:31 ON (B)(6) 2024. AT 22:11:48, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VF. MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE WAS SHUT DOWN AT 22:13:41 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223349&pc=BZD | 2518422-2025-007510 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, KIDNEY DISEASE/TOXICITY AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223553&pc=DSQ | 2916596-2025-00554 | 2024-09-10 04:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Renal Failure; Respiratory Failure; Heart Failure/Congestive Heart Failure; Gastrointestinal Hemorrhage; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT'S CREATININE ROSE TO 2.5 MILLIGRAMS PER DECILITER (MG/DL) AND THEY REQUIRED 1 UNIT OF PACKED RED BLOOD CELLS (PRBC) FOR AN UNKNOWN REASON. THE FOLLOWING DAY, A CHEST TUBE WAS PLACED FOR POSSIBLE HEMATOMA. ON (B)(6) 2024 PATIENT WAS IN AND OUT OF ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. THEY WERE REINTUBATED ON (B)(6) 2024 WITH BRONCHOSCOPY FOR LEFT LUNG MUCUS PLUG. ON (B)(6) 2024, THE PATIENT WAS SUCCESSFULLY EXTUBATED AND CONTINUOUS RENAL REPLACEMENT THERAPY WAS INITIATED. GASTROINTESTINAL (GI) BLEED SUSPECTED ON (B)(6) 2024 DUE TO DARK STOOLS AND 1 UNIT OF BLOOD WAS GIVEN. ON (B)(6) 2024 THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION STORM. 12 IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SHOCKS AND LIDOCAINE, MAGNESIUM, AND CALCIUM WERE ADMINISTERED. THE PATIENT REQUIRED REINTUBATION. TRANSESOPHAGEAL ECHOCARDIOGRAM REVEALED NO SUCTION AND A DILATED RIGHT VENTRICLE WITH MODERATE TO SEVERE DYSFUNCTION. THE RIGHT VENTRICULAR ASSIST DEVICE (RVAD) FLOWS WERE INCREASED. LEFT VENTRICULAR ASSIST DEVICE WAS DECREASED FROM 5500 TO 5400. VENTRICLE PACING AND SENSING (VVI) WAS INCREASED TO 100 BPM. HEMOGLOBIN AND HEMATOCRIT DECREASED TO 6.8 GRAMS PER DECILITER (G/DL) AND 22%. PLATELET COUNT WAS 57,000. TWO UNITS OF PRBC WERE ADMINISTERED. ON (B)(6) 2024 THE PATIENT WAS EXTUBATED, AND 2 UNITS OF PLATELETS WERE TRANSFUSED. THE OXYGENATOR WAS CLAMPED FOR TRIAL AND REMOVED ON (B)(6) 2024. GI BLEEDING CONTINUED (B)(6) 2024 AND THE PATIENT RECEIVED 4 ADDITIONAL UNITS OF BLOOD. ESOPHAGOGASTRODUODENOSCOPY (EGD) PERFORMED ON (B)(6) 2024 WAS NEGATIVE. (B)(6) 2024 FLEXIBLE SIGMOIDOSCOPY SHOWED MULTIPLE LARGE ULCERS IN DISTAL RECTUM. THE PATIENT REQUIRED FURTHER TRANSFUSIONS. ON (B)(6) 2024, THE PATIENT EXPERIENCED AN ACUTE RECURRENT GASTROINTESTINAL BLEED WITH BRIGHT RED BLOOD AND LOW FLOW ALARMS. THE PATIENT EXPERIENCED ALTERED MENTAL STATUS AND WAS REINTUBATED. FIVE UNITS OF PRBC WERE TRANSFUSED. THE FLEXIBLE SIGMOIDOSCOPY SHOWED ARTERIAL BLEEDING ABOVE DENTATE LINE WHICH RESOLVED WITH EPINEPHRINE INJECTION AND 3 CLIPS. THE PATIENT'S DRIVELINE HAD YELLOW DRAINAGE SO THEY WERE STARTED ON PIPERACILLIN AND TAZOBACTAM (ZOSYN) WITH NO KNOWN CULTURE OBTAINED. THE PATIENT HAD ALSO DEVELOPED DELIRIUM. THE PATIENT'S PROGNOSIS WAS POOR AND SO THE FAMILY ULTIMATELY DECIDED TO WITHDRAWAL CARE. THE PATIENT PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: H6 - DELIRIUM E0207. HEMATOMA E0505. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED EVENTS, AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. THE REPORTED LOW FLOW HAZARD ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE SUBMITTED FOR REVIEW; HOWEVER, THE ACCOUNT REPORTED THAT THE ALARMS OCCURRED DUE TO THE PATIENT¿S PROFUSE GASTROINTESTINAL (GI) BLEEDING. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, ARRYTHMIA, BLEEDING, INFECTION, NEUROLOGICAL EVENTS (NOT STROKE RELATED), AND DEATH. SECTION 1 OF THE IFU ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA, INFECTION, AND NEUROLOGIC DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. THIS SECTION, UNDER ¿ANTICOAGULATION¿, PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. SECTION 6 OF THE IFU, UNDER "RIGHT HEART FAILURE", DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. SECTION 5 OF THE HEARTMATE 3 LVAS PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223796&pc=NGV | 2124215-2025-03098 | 2024-12-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-23 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 20MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. POST PROCEDURE THE PATIENTS BLOOD PRESSURE WAS LOW FOLLOWING EXTUBATION. A TRANSTHORACIC ECHOCARDIOGRAM WAS PERFORMED AND RULED OUT A PERICARDIAL EFFUSION. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL WITH NO TREATMENT OR INTERVENTION. LATER THAT EVENING THE PATIENT BEGAN TO HAVE TROUBLE BREATHING AND THE PATIENT DIED. THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21224267&pc=OZD | 1220648-2025-25749 | 2025-01-07 05:00 | Death | ABIOMED, INC. | 2025-01-23 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: TH COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, PERSISTENT SUCTION AND LOW FLOW PRESENT DESPITE ADEQUATE PRELOAD AND POSITION. THE PUMP WAS REPOSITIONED, LVEDP 20S, AND GOOD RIGHT VENTRICLE FUNCTION WAS EVALUATED BEFORE REMOVING DEVICE AND REPLACING WITH NEW IMPELLA. THE PATIENT NEVER SHOWED ANY IMPROVEMENT IN NEURO STATUS. BRAIN FLOW STUDY COMPLETED AND INDICATIVE OF BRAIN DEATH. THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION FOR THE LOW PUMP FLOW HAS BEEN COMPLETED. DATA LOGS WERE REVIEWED FINDING NO (MOTOR CURRENT) MC SPIKE OBSERVED. NO POSITIONING ISSUE ALARM. NO PLACEMENT SIGNAL OCCURRED. THE CAUSE OF THE LOW PUMP FLOW ISSUE CANNOT BE DETERMINED SINCE THE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND LACK OF CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21224378&pc=BZD | 2518422-2025-007584 | 2024-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-01-23 05:00 | BZD | REMSTAR AUTO | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE TWO ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21224811&pc=PYX | 1220648-2025-25759 | 2025-01-12 05:00 | Death | ABIOMED, INC. | 2025-01-23 05:00 | PYX | IMPELLA RP | Patient-Device Incompatibility | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA RP INSERTION, THE RP WAS ADVANCED OVER WIRE BUT UNABLE TO PASS THE PULMONIC VALVE DESPITE MULTIPLE ATTEMPTS. IMPELLA WIRE WAS REMOVED AND EXCHANGED BECAUSE THE 0.027 WIRE WAS SLIGHTLY KINKED/BENT. SHEATH FLUSHED AGAIN AND IMPELLA RP ADVANCED OVER WIRE AND SUCCESSFUL ADVANCED INTO PULMONARY ARTERY WITH PIGTAIL DEFLECTING TO RIGHT PULMONARY ARTERY. IMPELLA STARTED. IT WAS FURTHER REPORTED THAT THE PATIENT CONTINUED TO DETERIORATE AND WAS MADE A DO-NOT-RESUSCITATE. THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY USE ISSUE SINCE NON - ABIOMED PRODUCT WAS USED WHICH IS NOT RECOMMENDED PER INSTRUCTIONS FOR USE. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25759 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4582 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25759 AND A NEW CODE WAS ADDED. CODE 4629 WAS INADVERTENTLY OMITTED FROM INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25759 AND WAS ADDED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA & RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21225932&pc=DSQ | 2916596-2025-00436 | 2024-06-13 04:00 | Death | THORATEC CORPORATION | 2025-01-23 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT, US, EP | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Liver Damage/Dysfunction; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Valvular Insufficiency/ Regurgitation; Lactate Dehydrogenase Increased | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS SEEN IN CLINIC AND NOTED TO HAVE SOME ISOLATED EVENTS OF ELEVATED FLOWS AND POWER. PATIENT HAD NOT NOTICED ANY CHANGES TO THEIR PARAMETERS AND HAD NOT HAD ANY ALARMS OR DARK/DISCOLORED URINE. LABS DRAWN IN THE CLINIC WERE STILL PENDING. THE PATIENT DID NOT TAKE COUMADIN DUE TO PREVIOUS GASTROINTESTINAL BLEEDS. THE PATIENT DID TAKE ACETYLSALICYLIC ACID 81 MILLIGRAMS DAILY. LOG FILES WERE SENT FOR REVIEW, AND THEY CAPTURED SEVERAL ELEVATED POWERS AND FLOWS OVER THE COURSE OF THE LOG FILE WITH POWER NOTED AS HIGH AS 9.1W. THESE WERE ALL ASSOCIATED WITH PULSATILITY INDEX (PI) EVENTS. THE ELEVATED PUMP POWERS WERE ALL TRANSIENT AND NOT SUSTAINED. THE CAUSE OF THE ELEVATED PUMP POWERS WAS NOT IDENTIFIED AT THE TIME. THE PATIENT DID NOT HAVE ANY ASSOCIATED SYMPTOMS WITH THE ELEVATED FLOWS AND POWER. AS TREATMENT THE PATIENT WAS RE-STARTED ON COUMADIN WITH AN INTERNATIONAL NORMALIZED RATIO (INR) GOAL RANGE OF 1.8-2.3. AS ON (B)(6) 2025, THE PATIENT HAD NOT NOTICED ANY MORE FLOW OR POWER ELEVATIONS. SUPPLEMENTAL INFORMATION PROVIDED NEW LOG FILES, AND LAB WORK WAS PENDING. THE PATIENT WAS SEEN IN CLINIC ON (B)(6) 2025 FOR SURVEILLANCE OF PREVIOUSLY REPORTED INCREASE IN LACTATE DEHYDROGENASE (LDH). PATIENT DENIED ANY CHANGES IN HEART FAILURE SYMPTOMS (NO SHORTNESS OF BREATH OR VOLUME OVERLOAD SIGNS). LOG FILES WERE REVIEWED AND THEY CAPTURED TRANSIENT POWER AND FLOW ELEVATION ASSOCIATED WITH PI EVENTS. THE PI EVENTS WERE CONSIDERED ROUTINE AND DROPPED THE PUMP TO ITS LOW-SPEED LIMIT OF 9000 RPM. OTHERWISE, THE LOG FILE CAPTURED ROUTINE EVENTS, WITH NO NOTABLE ALARM CONDITIONS OR PARAMETER CHANGES.Event Description: FURTHER INFORMATION FROM (B)(6)2025 PROVIDED THAT THE PATIENT WAS ADMITTED AFTER GOING TO THE LOCAL EMERGENCY ROOM WITH CHEST PAIN. THE PATIENT REPORTED HAVING LOWER EXTREMITY WEAKNESS OVER THE PAST COUPLE OF DAYS. THERE WAS A CONCERN OF LEFT VENTRICULAR ASSIST DEVICE (LVAD) THROMBOSIS WITH ELEVATED LACTATE DEHYDROGENASE (LDH). THEY WERE RESTARTED ON WARFARIN RECENTLY AND INCREASED ASPIRIN TO 325 MG DAILY. LABS ON (B)(6)2025 SHOWED ELEVATED LDH, AND LIVER FUNCTION TESTS. NO ALARMS WERE SEEN ON THE INITIAL LVAD INTERROGATION, BUT THEY DID HAVE SOME ELEVATIONS IN FLOW AND POWER WHICH WAS NOT NEW. A BEDSIDE ECHO WAS BEING DONE. PATIENT WOULD ALSO GO FOR A COMPUTED TOMOGRAPHY (CT) SCAN THE MORNING OF (B)(6)2025. ADDITIONAL INFORMATION AFTER INQUIRING REGARDING THROMBUS PROVIDED THAT CHANGES TO THE PATIENT'S CONDITION OR ANTICOAGULATION STATUS INCLUDED COUMADIN BEING STARTED AND LOW MOLECULAR WEIGHT HEPARIN (LMWH) BEING STARTED. THE PATIENT EXPERIENCED INTERMITTENT POWER AND FLOW ELEVATIONS, NO CHANGES TO SPEED. AN ECHO WAS PERFORMED ON (B)(6)2025. THE LEFT VENTRICLE WAS SEVERELY DILATED, AND LEFT VENTRICULAR SYSTOLIC FUNCTION WAS SEVERELY REDUCED. THERE WAS A LVAD INFLOW CANNULA NOTED IN THE LEFT VENTRICLE. THE RIGHT VENTRICLE WAS ENLARGED, AND THE RIGHT VENTRICULAR SYSTOLIC FUNCTION WAS REDUCED. THERE WAS MODERATE MITRAL REGURGITATION AND MILD TRICUSPID REGURGITATION. RIGHT VENTRICULAR SYSTOLIC PRESSURE (RVSP) WAS ELEVATED, ESTIMATED TO BE 55-60 MMHG. THE AORTIC VALVE DID NOT APPEAR TO OPEN ON ANY BEATS, AS EXPECTED WITH AN LVAD. THERE WAS MODERATE CONTINUOUS AORTIC REGURGITATION. THE INFERIOR VENA CAVA WAS DILATED WITHOUT RESPIRATORY COLLAPSE, INDICATING A RIGHT ATRIAL PRESSURE OF APPROXIMATELY 15 MMHG. COMPARED TO THE PRIOR TRANSTHORACIC ECHOCARDIOGRAM (TTE) FROM (B)(6)2024, THE MITRAL REGURGITATION AND AORTIC REGURGITATION WERE SIMILAR. THE ESTIMATED RVSP WAS HIGHER LIKELY RELATED TO AN INCREASE IN THE ESTIMATED RIGHT ATRIAL PRESSURE (RAP). A CT EXAM WAS COMPLETED ON (B)(6)2025 WHEN THE PATIENT WAS ADMITTED FOR ACUTE DECOMPOSITION WITH PUMP CLOT. ON (B)(6)2025, THE PATIENT HAD INCREASED SHORTNESS OF BREATH, AND ON (B)(6)2025, THE PATIENT HAD CHEST PAIN AND SHORTNESS OF BREATH. ON (B)(6)2025, COUMADIN WAS RESTARTED AS TREATMENT, AND INCREASED FOLLOW UP. THE PATIENT DID NOT WANT TO BE ADMITTED TO THE HOSPITAL AT THIS TIME FOR ADDITIONAL TREATMENT OPTIONS. ON (B)(6)2025, THE PATIENT WAS ADMITTED TO THE HOSPITAL AND WAS NOT END OF LIFE. THROMBUS HAD NOT RESOLVED, AND THE PATIENT REMAINED END OF LIFE.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS PUMP THROMBOSIS. THE PATIENT WAS NOT AN EXCHANGE OR TRANSPLANT CANDIDATE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORTED DEVICE THROMBUS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED ELEVATIONS IN PUMP POWER AND FLOW; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHERMORE, A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), (B)(6), AND THE REPORTED RIGHT HEART FAILURE, REGURGITATION, ELEVATED LACTATE DEHYDROGENASE (LDH), AND ELEVATED LIVER FUNCTION TESTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILES COLLECTIVELY CONTAINED EVENTS FROM 11DEC2024 THROUGH 29JAN2025. SLIGHT, TRANSIENT ELEVATIONS IN PUMP POWER AND FLOW WERE CAPTURED ON (B)(6) 2024, AND(B)(6) 2025, NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED FOR THE DURATION OF THE FILE. THE ACCOUNT REPORTED THAT HEARTMATE II LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. C, AND THE HEARTMATE II LVAS PATIENT HANDBOOK, REV. C, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS, RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND HEMOLYSIS, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP SPEED, POWER, FLOW, AND PULSATILITY INDEX (PI). IN REFERENCE TO POWER, THIS SECTION EXPLAINS THAT PUMP POWER IS A DIRECT MEASUREMENT OF MOTOR VOLTAGE AND CURRENT; THEREFORE, CHANGES IN PUMP SPEED, FLOW, OR PHYSIOLOGICAL DEMAND CAN AFFECT PUMP POWER. THIS SECTION ALSO NOTES THAT DEVICE FLOW AND POWER GENERALLY RETAIN A LINEAR RELATIONSHIP AT A GIVEN SPEED; HOWEVER, WHILE POWER IS DIRECTLY MEASURED BY THE SYSTEM CONTROLLER, THE REPORTED FLOW IS ESTIMATED BASED ON POWER. ADDITIONALLY, ANY INCREASE IN POWER NOT RELATED TO INCREASED FLOW CAUSES ERRONEOUSLY HIGH FLOW READINGS. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, WARNS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. SECTION 6, "PATIENT CARE AND MANAGEMENT" (UNDER "PUMP PERFORMANCE MONITORING"), EXPLAINS THAT PUMP PERFORMANCE IS SENSITIVE TO CHANGES IN SYSTEMIC VASCULAR RESISTANCE AND LEFT VENTRICULAR FILLING. SECTION 6, UNDER "ANTICOAGULATION", CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. SECTION 6 ALSO DESCRIBES THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LEFT VENTRICULAR ASSIST SYSTEM DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORTED DEVICE THROMBUS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED ELEVATIONS IN PUMP POWER AND FLOW; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHERMORE, A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILES COLLECTIVELY CONTAINED EVENTS FROM (B)(6)2024 THROUGH (B)(6)2025. SLIGHT, TRANSIENT ELEVATIONS IN PUMP POWER AND FLOW WERE CAPTURED ON (B)(6)2024, (B)(6)2025. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED FOR THE DURATION OF THE FILE. THE PATIENT REMAINS ONGOING ON HEARTMATE II LVAS, SERIAL NUMBER (B)(6), AND NO FURTHER RELATED EVENTS HAVE BEEN REPORTED AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) REV. C, AND THE HEARTMATE II LVAS PATIENT HANDBOOK, REV. C, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEVICE THROMBOSIS, RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND HEMOLYSIS, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP SPEED, POWER, FLOW, AND PULSATILITY INDEX (PI). IN REFERENCE TO POWER, THIS SECTION EXPLAINS THAT PUMP POWER IS A DIRECT MEASUREMENT OF MOTOR VOLTAGE AND CURRENT; THEREFORE, CHANGES IN PUMP SPEED, FLOW, OR PHYSIOLOGICAL DEMAND CAN AFFECT PUMP POWER. THIS SECTION ALSO NOTES THAT DEVICE FLOW AND POWER GENERALLY RETAIN A LINEAR RELATIONSHIP AT A GIVEN SPEED; HOWEVER, WHILE POWER IS DIRECTLY MEASURED BY THE SYSTEM CONTROLLER, THE REPORTED FLOW IS ESTIMATED BASED ON POWER. ADDITIONALLY, ANY INCREASE IN POWER NOT RELATED TO INCREASED FLOW CAUSES ERRONEOUSLY HIGH FLOW READINGS. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, WARNS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. SECTION 6, "PATIENT CARE AND MANAGEMENT" (UNDER "PUMP PERFORMANCE MONITORING"), EXPLAINS THAT PUMP PERFORMANCE IS SENSITIVE TO CHANGES IN SYSTEMIC VASCULAR RESISTANCE AND LEFT VENTRICULAR FILLING. SECTION 6, UNDER "ANTICOAGULATION", CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. SECTION 6 ALSO DESCRIBES THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LEFT VENTRICULAR ASSIST SYSTEM DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21225950&pc=QZI | 2124215-2025-03287 | 2024-12-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-23 05:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A FARAWAVE PULSED FIELD ABLATION CATHETER, NO INTRAPROCEDURAL COMPLICATIONS OCCURRED. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL AND THE FOLLOWING DAY HAD A DECOMPENSATION. THE PATIENT WENT BACK TO THE HOSPITAL WHERE AN ECHOCARDIOGRAM WAS PERFORMED BUT PASSED AWAY.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A FARAWAVE PULSED FIELD ABLATION CATHETER, NO INTRAPROCEDURAL COMPLICATIONS OCCURRED. THE PATIENT WAS DISCHARGED FROM THE HOSPITAL AND THE FOLLOWING DAY HAD A DECOMPENSATION. THE PATIENT WENT BACK TO THE HOSPITAL WHERE A ECHOCARDIOGRAM WAS PERFORMED BUT PASSED AWAY. IT WAS FURTHER REPORTED THAT 48 LESIONS WERE DELIVERED DURING THE PROCEDURE. THE AUTOPSY WAS NOT PERFORMED DUE TO FAMILY RELIGIOUS BELIEFS AND PREFERENCES. THE PATIENT PASSED AWAY AT HOME. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: UPDATED FIELD B5 WITH ADDITIONAL INFORMATION RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226495&pc=OZP | 2032227-2025-115186 | 2024-11-13 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-23 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Diabetic Ketoacidosis; Confusion/ Disorientation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON UNKNOWN DATE OF DEATH. CAUSE OF DEATH WAS BRAIN IN FOGGED FARCT. DIABETES DIABETIC KETOACIDOSIS. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS ADMITTED TO A HOSPITAL. THE CUSTOMER¿S BLOOD GLUCOSE LEVEL WAS UNKNOWN. THE INSULIN PUMP WAS NOT WORN AT THE TIME OF DEATH. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-213A, MMT-332A, MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-213A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. MMT-1884 WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226497&pc=OZP | 2032227-2025-115187 | 2021-11-21 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-23 05:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Break; Material Integrity Problem | Insufficient Information | P150001 |  | 1 | Event Description: MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY. IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON 21 NOVEMBER. CAUSE OF DEATH WAS UNKNOWN. TROUBLESHOOTING WAS NOT PERFORMED IT WAS UNKNOWN WHETHER THE PUMP WAS USED WITHIN 48 HOURS AND WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE DEVICE. MMT-1780KL, MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. FRN-UNK-RSVR, UNOMED INF SETEvent Description: UPDATED SUMMARY: MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY. IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025, DATE OF DEATH. CAUSE OF DEATH WAS UNKNOWN CAUSE OF DEATH. TROUBLESHOOTING WAS NOT PERFORMED IT WAS UNKNOWN WHETHER THE PUMP WAS USED WITHIN 48 HOURS AND WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE DEVICE. MMT-1780KL, MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THIS IS 1 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN RECEIVED WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. THE INFORMATION HAS BEEN PROVIDED IN SECTION B5 WITH THIS REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21226500&pc=OZP | 2032227-2025-115188 | 2021-11-21 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-23 05:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Break; Material Integrity Problem | Insufficient Information | P150001 |  | 1 | Event Description: MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY. IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON 21 NOVEMBER. CAUSE OF DEATH WAS UNKNOWN. TROUBLESHOOTING WAS NOT PERFORMED IT WAS UNKNOWN WHETHER THE PUMP WAS USED WITHIN 48 HOURS AND WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE DEVICE. MMT-1780KL, MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK.Event Description: UPDATED SUMMARY: THE REFERENCED LEGAL MATTER WAS RECEIVED BY MEDTRONIC ON 03/05/2025 FROM THE ATTORNEY OF THE FAMILY. IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON 21 NOVEMBER 2021 (DATE OF DEATH). CAUSE OF DEATH WAS UNKNOWN CAUSE OF DEATH. TROUBLESHOOTING WAS NOT PERFORMED IT WAS UNKNOWN WHETHER THE PUMP WAS USED WITHIN 48 HOURS AND WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. IT WAS NOT KNOWN IF THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: IT WAS UNKNOWN WHETHER THE CUSTOMER WILL CONTINUE USING THE DEVICE. MMT-1780KL, MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THIS IS 2 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204943&pc=BZD | 2518422-2025-006103 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Dyspnea; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, DIFFICULTY BREATHING SHORTNESS OF BREATH AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205089&pc=BZD | 2518422-2025-006151 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG CANCER AND ASTHMA (NEW OR WORSENING). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205102&pc=BZD | 2518422-2025-006116 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, CANCER AND KIDNEY CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205349&pc=BZD | 2518422-2025-006226 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205503&pc=BZD | 2518422-2025-006286 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH (NO CAUSE OF DEATH LISTED). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205510&pc=BZD | 2518422-2025-006279 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, PANCREATIC, LIVER AND KIDNEY CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205562&pc=BZD | 2518422-2025-006289 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205571&pc=BZD | 2518422-2025-006283 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND NON-HODGKINS LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER AND NON-HODGKINS LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX B: DESCRIBED EVENT OR PROBLEM (B5 VERBIAGE) WAS CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205627&pc=BZD | 2518422-2025-006327 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Inflammation; Nausea; Vomiting; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205633&pc=BZD | 2518422-2025-006318 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205679&pc=BZD | 2518422-2025-006257 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded; Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX D: PRODUCT BRAND NAME WAS CORRECTED. BOX H6: DEVICE PROBLEM CODE GRID AND HEALTH IMPACT GRID WAS CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205685&pc=BZD | 2518422-2025-006024 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205687&pc=BZD | 2518422-2025-006312 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205696&pc=BZD | 2518422-2025-006319 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205827&pc=BZD | 2518422-2025-006290 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH LUNG PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205907&pc=BZD | 2518422-2025-006360 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LIVER, KIDNEY AND BONE CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21205961&pc=CBK | 2518422-2025-006395 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | CBK | TRILOGY 100 | Degraded | Headache; Nausea; Vomiting; Dizziness; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K083526 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: BOX D: SUSPECT MEDICAL DEVICE , 510K AND RECALL (Z) NUMBER HAS BEEN UPDATED/CORRECTED IN THIS REPORT. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206024&pc=BZD | 2518422-2025-006415 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206311&pc=BZD | 2518422-2025-006417 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ASTHMA (NEW OR WORSENING), LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206355&pc=BZD | 2518422-2025-006498 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Headache; Hypersensitivity/Allergic reaction; Pulmonary Dysfunction; Dizziness; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206470&pc=BZD | 2518422-2025-005662 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206474&pc=BZD | 2518422-2025-006521 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, LUNG DISEASE, DEATH, SEVERE EAR, NOSE, AND THROAT INFLAMMATION AND REACTIVE AIRWAY DISEASE(RAD). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206611&pc=MVK | 3008642652-2025-00748 | 2024-06-22 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-22 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 13:01:35 ON (B)(6) 2024. AT 09:37:11 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 09:38:29, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 10:44:52 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206693&pc=BZD | 2518422-2025-006615 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX B: DATE OF DATE WAS CORRECTED. BOX H6: PATIENT OUTCOME CODE GRID WAS CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206694&pc=PNJ | 9612164-2025-00374 | 2025-01-02 05:00 | Death | MEDTRONIC IRELAND | 2025-01-22 05:00 | PNJ | MICRA¿ AV2 | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT, DURING THE DEPLOYMENT AND RECAPTURE OF A LEADLESS IMPLANTABLE PULSE GENERATOR (MAV2 DEVICE) IN THE RIGHT VENTRICLE, SEVERAL ISSUES AROSE. THE INVASIVE ARTERIAL LINE BLOOD PRESSURE DROPPED TO 50/30, AND THERE WAS NO CAPTURE AT 2.5V DURING TESTING, DESPITE CONFIRMATION OF 2-3 TINES CAUGHT VIA A TUG-TEST. A PERFORATION OF THE RIGHT VENTRICULE OCCURRED, LEADING TO THE INITIATION OF PERICARDIOCENTESIS AND EXTRACORPOREAL MEMBRANE OXYGENATION, WHICH DID NOT RESOLVE THE SITUATION. A CODE WAS CALLED DUE TO THE PATIENT'S CONDITION, AND OPEN-HEART REPAIR SURGERY WAS REQUIRED FOR THE PERFORATED RIGHT VENTRICLE. THE PATIENT WAS DESCRIBED AS "VERY SICK" IN THE CARDIOTHORACIC INTENSIVE CARE UNIT POST-SURGERY AND PASSED AWAY SEVERAL DAYS AFTER THE EVENT. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206705&pc=PNJ | 9612164-2025-00375 | 2024-11-05 05:00 | Death | MEDTRONIC IRELAND | 2025-01-22 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Pulmonary Edema; Sepsis; Cardiogenic Shock; Renal Impairment | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT'S CHRONIC KIDNEY DISEASE WORSENED AFTER THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) WAS IMPLANTED. THE DEVICE REMAINS IN USE. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT PATIENT'S CHRONIC KIDNEY DISEASE DID NOT WORSEN DUE TO THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) IMPLANT, THAT IT WAS NOT RELATED TO THE LEADLESS IPG.Event Description: IT WAS REPORTED THAT THE PATIENTS CHRONIC KIDNEY DISEASE WORSENED AFTER THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) WAS IMPLANTED. THE DEVICE REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER NOTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE PATIENT WAS TREATED WITH MEDICATION. IT WAS ALSO NOTED THAT THE PATIENT EXPERIENCED CARDIOMEGALY WITH MILD PULMONARY EDEMA.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 22-JAN-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206738&pc=BZD | 2518422-2025-006609 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND INHALED THE FOAM INTO HER LUNGS WHEN IT BLEW OUT HER LEFT LUNG COLLAPSE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: IN THE PREVIOUS REPORT, SECTION B2 OUTCOMES ATTRIBUTED TO ADVERSE EVENT WAS CAPTURED INCORRECTLY AND IT WAS CORRECTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206748&pc=BZD | 2518422-2025-006613 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN THE PREVIOUS REPORT, SECTION B2 OUTCOMES ATTRIBUTED TO ADVERSE EVENT WAS CAPTURED INCORRECTLY AND IT WAS CORRECTED IN THIS REPORT. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206765&pc=BZD | 2518422-2025-006624 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206824&pc=BZD | 2518422-2025-006618 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Inflammation; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206825&pc=BZD | 2518422-2025-006593 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Chronic Obstructive Pulmonary Disease (COPD) | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND COPD. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21206857&pc=MVK | 3008642652-2025-00749 | 2024-11-06 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-22 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6)2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 22:45:17 ON (B)(6)2024. AT 23:33:35, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 23:37:55 ON (B)(6)2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207046&pc=LZG | 3004464228-2025-03102 | 2024-12-09 05:00 | Death | INSULET CORPORATION | 2025-01-22 05:00 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Insufficient Information | Coma | K211575 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE PATIENT HAS PASSED AWAY ON (B)(6) 2024. THE PATIENT REPORTEDLY DECEASED FROM A ¿DIABETIC COMA¿. THE PATIENT'S EX-HUSBAND REPORTED THE PATIENT'S DEATH TO THE PRESTATAIRE ON (B)(6) 2025, WHO THEN REPORTED IT TO INSULET. THE PATIENT'S EX-HUSBAND DIDN¿T PROVIDE ANY FURTHER INFORMATION TO THE PRESTATAIRE. ACCORDING TO THE PRESTATAIRE, ¿THE POD EXPIRATION OCCURRED ON (B)(6) 2024 AT 6:48 P.M. WITHOUT ANY REACTION FROM THE PATIENT¿. GRAPHS FROM MYDIABBY APP HAS BEEN PROVIDED TO INSULET BY THE PRESTATAIRE. THE GRAPH SHOWS THE PATIENT ADMINISTERED A BOLUS DOSE OF (DOSE AMOUNT UNKNOWN) ON (B)(6) 2024 IN THE EVENING AND NO MORE INSULIN INJECTIONS CAN BE OBSERVED ON MYDIABBY GRAPH FROM (B)(6) 2024. Manufacturer Narrative: THE PERSONAL DIABETES MANAGER (PDM) SERIAL NUMBER (B)(6), HAS BEEN REQUESTED TO THE PRESTATAIRE FOR INVESTIGATION. THE PRESTATAIRE INFORMED INSULET THAT THE PDM IS AVAILABLE, AND THE PRESTATAIRE WILL BE ABLE TO SEND IT TO INSULET BY ¿END OF JANUARY¿. THE PRESTATAIRE CONFIRMED TO INSULET THE POD IS NOT AVAILABLE FOR INVESTIGATION. THE POD LOT NUMBER HAS NOT BEEN REPORTED. ADDITIONAL FOLLOW-UPS HAVE BEEN MADE WITH THE PATIENT¿S ENDOCRINOLOGIST AND GENERAL PRACTITIONER. THE PATIENT'S ENDOCRINOLOGIST INFORMED INSULET THAT HER DEPARTMENT HAS BEEN MADE AWARE OF THE PATIENT¿S DEATH BY THE PATIENT'S EX-HUSBAND, BUT THE PATIENT'S EX-HUSBAND DIDN¿T PROVIDE ANY FURTHER INFORMATION TO THE ENDOCRINOLOGY DEPARTMENT, AND THEY WERE NOT CONTACTED BY ANYONE ELSE REGARDING THIS DEATH. THE GENERAL PRACTITIONER WAS NOT AWARE OF THE PATIENT¿S DEATH. THE PATIENT¿S ENDOCRINOLOGIST AND GENERAL PRACTITIONER DO NOT HAVE ANY ADDITIONAL INFORMATION TO PROVIDED REGARDING THE PATIENT OR THE PATIENT¿S DEATH. LOT RELEASE RECORDS FOR THE PDM WERE REVIEWED, AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA AT THE TIME OF RELEASE. Manufacturer Narrative: THE CASE DESCRIPTION STATES THE USER HAD HIGH BLOOD GLUCOSE LEVELS (BG) WHILE USING THE SYSTEM. INSPECTION OF THE LOG FILES FROM THE DEVICE FOUND NO ABNORMALITIES OR ISSUES WITH INSULIN DELIVERY. THE LOG FILES SHOW THAT ON (B)(6) 2024 AT 19:44 A 10 UNIT (U) BOLUS WAS COMPLETED BEFORE THE CURRENT POD WAS DEACTIVATED AND A NEW POD WAS ACTIVATED. THE POD WAS RUNNING THE USER'S SET "BASAL 1" BASAL PROGRAM DELIVERING A VARYING RATE THROUGH OUT THE DAY BETWEEN 0.6 U/HR TO 0.85 U/HR. AT 23:03 ON (B)(6) 2024 A 10U BOLUS WAS DELIVERED. AT 14:14 ON (B)(6) 2024 A MANUAL BLOOD GLUCOSE (BG) READING OF 237 MG/DL WAS ENTERED INTO THE PERSONAL DIABETES MANAGER (PDM) FOLLOWED BY ANOTHER 10U BOLUS. ON (B)(6) 2024 2 MORE BOLUSES WERE DELIVERED, OF 10U AND 15U, RESPECTIVELY. ON (B)(6) 2024 AT 12:56 A MANUAL BG ENTRY OF 277 MG/DL WAS ENTERED INTO THE PDM FOLLOWED BY A 9U BOLUS. A 10U BOLUS WAS ALSO DELIVERED ON (B)(6) 2024 AT 15:21. A FINAL MANUAL BG ENTRY OF 246 MG/DL WAS ENTERED AT 22:57 FOLLOWED BY A 9U BOLUS ON (B)(6) 2024. AT 03:49 ON (B)(6) 2024 THE POD CORRECTLY GENERATED AN EXPIRATION ALARM AFTER 80 HOURS OF USE. THIS IS A SAFETY FEATURE THAT DOES NOT INDICATE A DEVICE FAILURE. NO PODS WERE FOUND TO BE ACTIVATED AFTER THIS POD WAS DEACTIVATED AND THUS NO INSULIN DELIVERY. LOG FILES SHOW THAT ALL BOLUSES THAT WERE PROGRAMMED WERE FULLY DELIVERED AND NO ISSUES WITH INSULIN DELIVERY WERE FOUND. THE SYSTEM WAS FOUND TO ACT AS EXPECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207059&pc=OZD | 1220648-2025-25690 | 2025-01-03 05:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT; ANATOMICAL ISSUES PREVENTED IMPLANT OF AN IMPELLA 5.5. THE PATIENT WAS SENT TO THE UNIT IN CRITICAL CONDITION. APPROXIMATELY 17 DAYS INTO SUPPORT, THE PATIENT DEVELOPED LOW PURGE PRESSURE ALARMS WHICH WERE UNRESOLVED FOLLOWING TROUBLESHOOTING. DISCUSSIONS WITH THE CARDIOVASCULAR SURGEON SURROUNDING GOALS OF CARE WERE MADE. DUE TO NO CURRENT EXIT PLAN AND THE COMPLEXITIES OF THE PATIENT, THE DECISION WAS MADE TO INCREASE THE DEXTROSE IN THE PURGE TO D20 WITH 25MEQ BICARBONATE AND WAIT/WATCH UNTIL THE IMPELLA CP POTENTIALLY FAILS. THERE WAS AN AUDIBLE HISSING/BUZZING SOUND COMING FROM THE INSERTION SITE; HOWEVER, THERE WAS NO VISIBLE PURGE FLUID. THE CARDIOVASCULAR SURGEON AND CRITICAL CARE TEAM PUT A PLAN IN PLACE TO PULL THE IMPELLA CP BACK INTO THE AORTA SHOULD THE PUMP FAIL. EXPLANT SUPPLIES REMAINED AT BEDSIDE. TWO DAYS LATER IT WAS REPORTED THE IMPELLA CP DEVICE HAD BEEN EXPLANTED. CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. NO DEVICE WAS RECEIVED FOR EVALUATION. DEVICE HISTORY RECORD REVIEW CONFIRMED THAT THE PUMP PASSED ALL POST-STERILE INSPECTION CHECKS. DATA LOGS WERE REVIEWED. LOW PURGE PRESSURE ALARMS OCCURRED AT EVENING OF 1/2 AT P-7 UNRESOLVED POST TROUBLESHOOTING. NO EVIDENCE OBSERVED TO CONFIRM THE ROOT CAUSE FOR REPORTED ISSUE. THE PUMP LEFT ON SUPPORT UNTIL PATIENT EXPIRED. DUE TO THE LIMITED CLINICAL INFORMATION AND LACK OF AVAILABLE DATA/PRODUCT THE ROOT CAUSE OF THIS INVESTIGATION IS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS SAVED FOR RETURN. HOWEVER, IT WAS NOT RECEIVED FROM THE CUSTOMER AT THE TIME OF THIS MDR WRITING, AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207146&pc=BZD | 2518422-2025-006654 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Respiratory Failure; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021. Manufacturer Narrative: IN PREVIOUS REPORT BOX G: CONTACT OFFICE ENTITY WAS ADDED INCORRECTLY, IT WAS CORRECTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207167&pc=LZG | 3004464228-2025-03106 | 2025-01-07 05:00 | Death | INSULET CORPORATION | 2025-01-22 05:00 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Improper or Incorrect Procedure or Method; Insufficient Information | Hyperglycemia | K211575 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WHO IS NON-VERBAL WITH DOWN SYNDROME WAS FOUND DECEASED IN HIS ROOM AT A GROUP HOME. THE STAFF REPORTED CHANGING THE PATIENT'S POD 3 DAYS PRIOR TO HIS DEATH AND THAT HIS BLOOD SUGAR WAS HIGH FOR THE ENTIRE 3 DAYS FOLLOWING THE POD CHANGE DUE TO POSSIBLE INFUSION SITE FAILURE. THE PATIENT'S HCP (HEALTH CARE PROVIDER) REPORTED THE OMNIPOD SYSTEM WAS APPROPRIATELY PROVIDING ALERTS DUE TO HIGH BLOOD GLUCOSE LEVELS, AND THE STAFF FAILED TO ASSESS AND CONFIRM THE DEVICE STATUS OR THE PATIENT'S BLOOD GLUCOSE LEVELS. ADDITIONALLY, THE HCP REPORTED THE STAFF NEGLECTED TO PROVIDE INSULIN BOLUSES FOR MEALS WHICH RESULTED IN THE PATIENT RECEIVING CONSIDERABLY LESS TOTAL DAILY INSULIN. THE BACK-UP PLAN ORDERED BY THE HCP WAS NOT INITIATED. BLOOD GLUCOSE LEVELS WERE REPORTED TO BE HIGH (>400 MG/DL). Manufacturer Narrative: ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED HYPERGLYCEMIA AND DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207172&pc=OZD | 1220648-2025-25691 | 2024-10-09 04:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Sepsis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED HIT, HIGH PURGE PRESSURE AND SEPSIS. THE PATIENT WOULD EVENTUALLY EXPIRE. THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5, AND WAS ON A RESPIRATOR, VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AND DIALYSIS THERAPIES PRIOR. NINE DAYS INTO SUPPORT ON THE IMPELLA 5.5 PUMP, THE PHYSICIAN REPORTED A PURGE PRESSURE HIGH ALARM WITH PURGE FLOWS AT 1.6-2.6ML AND PURGE PRESSURE AT 1060MMHG. ATTENTION WAS DRAWN TO THE FACT THAT THE PUMP CAN STOP AT ANY TIME, AND THE RECOMMENDATION MADE TO CHANGE THE PUMP; DISCUSSION REGARDING LYSIS OPTION WAS MADE WITH SHARING OF THE LYSIS PROTOCOL AS WELL. THE FOLLOWING DAY, THE PHYSICIAN INDICATED THE PUMP WAS LEFT IN. THE PURGE FLOW WAS AT MAXIMUM 1ML, AND PURGE PRESSURE NOTED SOMETIMES FELL BELOW 1000MMHG AND THE NEXT DAY STATED THE PUMP WAS SUCCESSFULLY LYSED, THE PATIENT WAS AWAKE AND SUCCESSFULLY IN WEANING. HOWEVER, APPROXIMATELY THREE DAYS LATER, THE PATIENT WAS REPORTEDLY INTUBATED AND SEDATED AND AT THE BEGINNING SEPSIS WAITING FOR FURTHER DECISIONS. TWO DAYS LATER THERAPIES WERE STOPPED, AND CARE WAS WITHDRAWN DUE TO LACK OF RECOVERY AND POSSIBLE POSITIVE OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR ANALYSIS. NO DATA LOGS WERE RETURNED FOR ANALYSIS. THE ROOT CAUSE OF HIGH PURGE PRESSURE AND SEPSIS WAS NOT DETERMINED AS NO PRODUCT OR DATA LOGS WERE RETURNED FOR ANALYSIS. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ B.2 LIFE-THREATENING WAS INADVERTENTLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER AND SHOULD OF NOT BEEN. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G.1 REVISED MANUFACTURING SITE NAME AND ADDRESS AS IT WAS PREVIOUSLY ENTERED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED FROM THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207275&pc=QFG | 3004464228-2025-03110 | 2025-01-08 05:00 | Death | INSULET CORPORATION | 2025-01-22 05:00 | QFG | OMNIPOD 5 POD | Adverse Event Without Identified Device or Use Problem | Hyperglycemia | K203768 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD HIGH BLOOD GLUCOSE LEVELS THAT WERE FELT TO BE ELEVATED DUE TO THE FLU. ON THE WAY TO THE HOSPITAL, THE PATIENT'S HEART STOPPED RESULTING IN THE PATIENT PASSING AWAY. THE PATIENT WAS REPORTED TO ALSO HAVE ADDISON'S DISEASE, OSTEOPOROSIS, THYROID ISSUE AND HIGH BLOOD PRESSURE. THE PATIENT'S SISTER WAS NOT AWARE OF ANY DEVICE MALFUNCTIONS WITH THE OMNIPOD SYSTEM AT TIME OF DEATH. BLOOD GLUCOSE LEVELS WERE NOT PROVIDED. Manufacturer Narrative: ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED HYPERGLYCEMIA AND DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. CLOUD - LOCKED DOWN/SMARTPHONE. LOCKDOWN. CLOUD - OMNIPOD 5 SOFTWARE APP VERSION. 3.1.1. CLOUD - SMARTPHONE OPERATING SYSTEM. N5004L-AM-Q-MV01602-06-01.06. CLOUD - SMARTPHONE HARDWARE. N5004L. CLOUD - CGM SENSOR TYPE. G6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207290&pc=OZD | 1220648-2025-25698 | 2025-01-04 05:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA CP | Malposition of Device; Infusion or Flow Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND THE FOLLOWING DAY POSITIONING ALARMS TRIGGERED DESPITE PULSATILE MOTOR CURRENT. THE IMPELLA WAS REPOSITIONED AND PROPER PLACEMENT CONFIRMED UNDER ECHOCARDIOGRAM. PATIENT STILL REQUIRED CHEMICAL SUPPORT TO MAINTAIN HEMODYNAMICS, POSSIBLE WORK UP FOR IMPELLA 5.5 WAS INDICATED. THE FOLLOWING DAY, WAVEFORMS ON REMAINED UNCHANGED SINCE THE MORNING WHEN ECHOCARDIOGRAM CONFIRMED INLET 3.8CM FROM AORTIC VALVE. PLACEMENT MONITORING WAS DISABLED. ANOTHER ECHOCARDIOGRAM WAS PLANNED WITH REPOSITIONING. NEXT DAYS FOLLOWING, IT WAS NOTED DUE TO FLOW SATURATION, INACCURATE POSITION MONITORING, AND LACK OF IMPROVEMENT, A DECISION WAS MADE TO EXPLANT THE IMPELLA CP WITHOUT EXCHANGE OR ESCALATION IN CARE THE IMPELLA CP DEVICE WAS REMOVED ON DAY FOUR OF SUPPORT, AND THE PATIENT WAS MEDICALLY MANAGED. THE PATIENT EXPIRED THE SAME DAY AFTER EXPLANT, HOWEVER. Manufacturer Narrative: THE INVESTIGATION OF SATURATED FLOW AND OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. SATURATED FLOW: DATA LOGS WERE REVIEWED AND THE SATURATED FLOWS WERE CONFIRMED ON (B)(6) 2025. LEADING UP TO THE FIRST POSITION ALARM THERE IS A DROP IN MOTOR CURRENT, SPEED DEVIATION, AND SUCTION ALARM WHILE RUNNING AT A CONSTANT P-LEVEL. ADDITIONALLY, THE LEFT VENTRICLE PLACEMENT SIGNAL (PS) AMPLITUDE BECOMES VARIABLE, OFTEN SHRINKING DOWN TO MATCH THE AORTIC PS. HOWEVER, THERE IS NO MOTOR CURRENT SPIKE THAT WOULD BE EXPECTED DURING BIOMATERIAL INGESTION. FOLLOWING THIS EVENT, THE USER STARTS TO CHANGE P-LEVEL OFTEN, BOTH POSITION IN VENTRICLE AND AORTA ALARMS INTERMITTENTLY TRIGGER, AND PUMP FLOWS BECOME SATURATED DEPENDING ON THE P-LEVEL. LATER IN THE CASE, SPECIFICALLY WHILE RUNNING CONSTANT AT P-7 THE EVENING OF (B)(6) 2025, MOTOR CURRENT/PUMP FLOWS CAN BE SEEN TO TREND UP AND BECOME SATURATED. RETURNED PRODUCT WAS INVESTIGATED AND THERE WERE NO ABNORMALITIES. THERE WAS NO SIGN OF BIOMATERIAL OR BLOOD INGRESS INTO THE PURGE GAP OR EVEN THE MOTOR HOUSING WHEN IT WAS TORN DOWN. THE PUMP WAS RUN IN A BUCKET FOR 15 MINUTES AND FLOWS WERE WITHIN SPECIFICATION. SINCE DATA LOGS DID NOT SHOW CONCLUSIVE CAUSAL EVIDENCE AND THE ISSUE DID NOT REPRODUCE, THE ROOT CAUSE OF THE SATURATED FLOWS COULD NOT BE DETERMINED. OPTICAL SIGNAL ISSUE: DATA LOGS WERE REVIEWED AND THE POSITIONING ALARMS WERE CONFIRMED ON (B)(6)2025. LEADING UP TO THE FIRST POSITION ALARM THERE IS A DROP IN MOTOR CURRENT, SPEED DEVIATION, AND SUCTION ALARM WHILE RUNNING AT A CONSTANT P-LEVEL. ADDITIONALLY, THE LEFT VENTRICLE PLACEMENT SIGNAL AMPLITUDE BECOMES VARIABLE, OFTEN SHRINKING DOWN TO MATCH THE AORTIC PS. FOLLOWING THIS EVENT, THE USER STARTS TO CHANGE P-LEVEL OFTEN, AND BOTH POSITION IN VENTRICLE AND AORTA ALARMS INTERMITTENTLY TRIGGER WHILE PUMP FLOWS BECOME SATURATED AT TIMES. IN TERMS OF 5S PARAMETERS, GAIN, LIGHT, AND LAMP WERE ALL STABLE THROUGH THE CASE. THE SIGNAL-TO-NOISE RATIO AND CONTRAST WERE CONSISTENTLY LOW, AND THIS WOULD NOT CAUSE THE POSITIONING ALARMS, ONLY EXACERBATE THE ISSUE. RETURNED PRODUCT WAS INVESTIGATED AND THERE WERE NO ABNORMALITIES. THE PUMP PASSED LASER CONTINUITY AND THERE WAS NO SIGN OF BIOMATERIAL IN THE PURGE GAP OR ANYWHERE IN THE MOTOR HOUSING WHEN TORN DOWN. WHEN CONNECTED TO A CONSOLE, ALL 5S PARAMETERS AND PUMP FLOWS WERE WITHIN SPECIFICATION WHEN RUNNING THE PUMP IN A BUCKET OF WATER FOR 16 MINUTES. CLINICAL DETAILS REPORT THAT THE PUMPS POSITION WAS CONFIRMED SEVERAL TIMES WITH ECHOCARDIOGRAPHY. SINCE DATA LOGS DO NOT SHOW CONCLUSIVE CAUSAL EVIDENCE, AND THE ISSUE DID NOT REPRODUCE IN THE LAB, THE ROOT CAUSE OF THE OPTICAL SIGNAL ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207540&pc=BZD | 2518422-2025-006670 | 2023-05-23 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207653&pc=BZD | 2518422-2025-006678 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207688&pc=BZD | 2518422-2025-006153 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN B: OUTCOMES ATTRIBUTED TO AE AND DATE OF DEATH WAS CORRECTED. IN H: PATIENT OUTCOME CODE GRID WAS UPDATED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207695&pc=OZP | 2032227-2025-114428 | 2024-12-29 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-22 05:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Dementia; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS DEMENTIA, HEART FAILURE, SOME RELATED TO DIABETES. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-7040A, MMT-396A, MMT-1884, MMT-332A. TROUBLESHOOTING WAS PERFORMED AND PUMP WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. NO PRODUCT RETURN IS REQUIRED FOR MMT-396A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207760&pc=BZD | 2518422-2025-006694 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21208057&pc=NGV | 2135147-2025-00369 | 2009-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-01-22 05:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "SINGLE VS DUAL ANTIPLATELET THERAPY AFTER LEFT ATRIAL APPENDAGE CLOSURE: A PROPENSITY SCORE MATCHING ANALYSIS", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, SINGLE CENTER STUDY TO COMPARE BOTH ISCHEMIC AND BLEEDING OUTCOMES IN PATIENTS RECEIVING SINGLE ANTIPLATELET THERAPY (SAPT) OR DUAL ANTIPLATELET THERAPY (DAPT) AFTER SUCCESSFUL LEFT ATRIAL APPENDAGE CLOSURE (LAAC). DEVICES INCLUDED IN THE STUDY WERE AMPLATZER AMULET, AMPLATZER CARDIAC PLUG, WATCHMAN FLX, AND WATCHMAN 2.5. THE ARTICLE CONCLUDED IN THIS PROPENSITY SCORE MATCHING ANALYSIS OF A SINGLE-CENTER LAAC COHORT, ISCHEMIC AND BLEEDING OUTCOMES DID NOT DIFFER AT 1 YEAR FOR PATIENTS DISCHARGED WITH SAPT OR DAPT. THESE RESULTS HAVE TO BE CONFIRMED IN AN ADEQUATELY POWERED RANDOMIZED CLINICAL TRIAL. [THE PRIMARY AND CORRESPONDING AUTHOR WAS LORENZ RAEBER, DEPARTMENT OF CARDIOLOGY, BERN UNIVERSITY HOSPITAL, UNIVERSITY OF BERN, CH-3010, BERN, SWITZERLAND, WITH CORRESPONDING E-MAIL ADDRESS: LORENZ.RAEBER@INSEL.CH] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2009 TO AUGUST 2023. A TOTAL OF 384 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 290 RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 74.6 YEARS, AVERAGE WEIGHT WAS 79.0 KG, AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, PAROXYSMAL ATRIAL FIBRILLATION, CEREBROVASCULAR EVENT, CORONARY HEART DISEASE, MYOCARDIAL INFARCTION, VASCULAR DISEASE, CONGESTIVE HEART FAILURE, BLEEDING (GASTROINTESTINAL, INTRACRANIAL), HEMATOLOGIC-COAGULATION DISORDER, DIFFUSE INTRACRANIAL AMYLOID ANGIOPATHY, AND ANGIODYSPLASIA/GASTRODUODENAL ULCER. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE: SINGLE VS DUAL ANTIPLATELET THERAPY AFTER LEFT ATRIAL APPENDAGE CLOSURE: A PROPENSITY SCORE MATCHING ANALYSIS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER AMULET WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ARTERIAL HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, PAROXYSMAL ATRIAL FIBRILLATION, CEREBROVASCULAR EVENT, CORONARY HEART DISEASE, MYOCARDIAL INFARCTION, VASCULAR DISEASE, CONGESTIVE HEART FAILURE, BLEEDING (GASTROINTESTINAL, INTRACRANIAL), HEMATOLOGIC-COAGULATION DISORDER, DIFFUSE INTRACRANIAL AMYLOID ANGIOPATHY, AND ANGIODYSPLASIA/GASTRODUODENAL ULCER. COMPLICATIONS REPORTED INCLUDED DEATH DURING THE INDEX HOSPITALIZATION, STROKE (ISCHEMIC, HEMORRHAGIC), BLEEDING, THROMBUS; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLE: "SINGLE VS DUAL ANTIPLATELET THERAPY AFTER LEFT ATRIAL APPENDAGE CLOSURE: A PROPENSITY SCORE MATCHING ANALYSIS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21208117&pc=FKX | 0008030665-2025-00175 | 2025-01-11 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-22 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., DEATH CERTIFICATE, ESRD DEATH NOTIFICATION, CAUSALITY, TREATMENT RECORD, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL.Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6)2025 WHILE UNDERGOING CCPD THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., DEATH CERTIFICATE, ESRD DEATH NOTIFICATION, CAUSALITY, TREATMENT RECORD, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN HE EXPIRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENT(S) IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT A LACK OF ADDITIONAL CLINICAL INFORMATION PRECLUDED A MORE COMPREHENSIVE INVESTIGATION. HOWEVER, BASED ON THE INITIAL REPORTING, THERE WAS NO INDICATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) WAS DEFICIENT OR MALFUNCTIONED IN ANY WAY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENT(S). THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. HOWEVER, GIVEN THE LACK OF CLINICAL FOLLOW-UP, NO TREATMENT RECORD, NO CAUSALITY, AND NO MANUFACTURER INVESTIGATION OF THE SUSPECT DEVICE, THIS LIBERTY SELECT CYCLER CANNOT BE EXCLUDED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENT(S) EXPERIENCED BY THE PATIENT. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE CYCLER WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE. IT WAS NOTED THAT THE DOOR LOGO WAS MISSING. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION AND SYSTEM AIR LEAK TESTS WERE PERFORMED AND PASSED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21208281&pc=OZD | 1220648-2025-25703 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. WHILE ON SUPPORT, HEMOLYSIS IN URINE, LABS REPORTING FINE. ADDITIONALLY, NO PULSES WERE NOTED IN BILATERAL LOWER LEG EXTREMITIES. CARE WAS WITHDRAWN AND PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE LIMB ISCHEMIA AND HEMOLYSIS HAS BEEN COMPLETED. DATA LOGS WERE REVIEWED. THERE WAS A MOTOR CURRENT AND PLACEMENT SIGNAL SPIKE, AND SPEED DEVIATION FOLLOWED BY LOW PUMP FLOWS AND SUCTION DURING THE TIME OF THE REPORTED HEMOLYSIS. THE MOTOR CURRENT REMAINED VARIABLE THROUGH THE CASE. THE CAUSE OF THE HEMOLYSIS ISSUE MOST LIKELY WAS BIOMATERIAL INGESTION. THE CAUSE OF THE LOW PUMP FLOW MOST LIKELY WAS DUE TO BIOMATERIAL INGESTION. THE ROOT CAUSE OF THE LIMB ISCHEMIA COULD NOT BE DETERMINED WITHOUT ADDITIONAL INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21208300&pc=BZD | 2518422-2025-006714 | 2024-01-10 05:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DS2ADV AUTO CPAP | Insufficient Information | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND PCA ELECTRICAL DAMAGE. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE SERVICE CENTER AND ERROR CODE FOUND. ADDITIONALLY, FAILURES OBSERVED WERE REUSABLE POLLEN FILTER SERVICE, ISO PORT KIT, PLUS CONTAMINATION, BLOWER BOX KIT CONTAMINATION AND HEATER PLATE KIT PHYSICAL DAMAGE. THE CUSTOMER COMPLAINT WAS REPRODUCED, AND THE DEVICE WAS NOT REPAIRED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21208456&pc=OZD | 1220648-2025-25704 | 2025-01-03 05:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA CP | Patient-Device Incompatibility | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE INSERTION OF IMPELLA CP DEVICE TO A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS UNSUCCESSFUL. THE PATIENT PRESENTED IN PROFOUND CARDIOGENIC SHOCK ON MULTIPLE PRESSORS AND INOTROPES. THE IMPELLA SHEATH AND WIRE INSERTED (ACCESS SITE LOCATION AND SIDE UNKNOWN), AND THE PUMP WAS PREPPED. HOWEVER, THE PATIENT'S ANATOMY WOULD NOT ACCOMMODATE DELIVERY OF THE IMPELLA CP DEVICE. REPORTEDLY, THE PATIENT EXPIRED SHORTLY THEREAFTER. NO FURTHER DETAILS WERE MADE AVAILABLE. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE THE PUMP WAS UNABLE TO PASS THROUGH VASCULATURE. Manufacturer Narrative: SOME PATIENT-SPECIFIC INFORMATION IS UNAVAILABLE AT THE TIME OF THIS MDR WRITING AND THEREFORE, RESPECTIVE FIELDS WERE REFLECTED "UNKNOWN" OR LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21208900&pc=MFA | 3007284006-2025-00015 | 2024-12-26 05:00 | Death | SPECTRANETICS CORPORATION | 2025-01-22 05:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Great Vessel Perforation; Pericardial Effusion | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND A RIGHT ATRIAL (RA) LEAD DUE TO NON-FUNCTION. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH AND A SPECTRANETICS MEDIUM VISISHEATH DILATOR SHEATH ON THE RV LEAD, THE PATIENT'S BLOOD PRESSURE DROPPED, AND AN EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON, STERNOTOMY, AND BYPASS. AN INFERIOR VENA CAVA (IVC) PERFORATION WAS DISCOVERED AND REPAIRED. THE EXTRACTION CONTINUED WITH THE SAME GLIDELIGHT DEVICE AND VISISHEATH, AND THE RV AND RA LEAD WERE BOTH REMOVED, BUT NEW BLEEDING WAS DISCOVERED AT THE SVC/INNOMINATE JUNCTION. THE PERFORATION WAS REPAIRED; HOWEVER, FOUR DAYS LATER, THE DECISION WAS MADE TO WITHDRAW CARE DUE TO HYPOXIC ANOXIC BRAIN INJURY. THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE 14F GLIDELIGHT DEVICE IN USE WHEN THE PERFORATIONS OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A3B): PATIENT GENDER UNK. D4/H4): THE LOT NUMBER WAS NOT PROVIDED; THUS, THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, SERIAL NUMBER, EXPIRATION DATE, AND MANUFACTURE DATE. H3): THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6): PERFORATION OF VESSELS, GREAT VESSEL PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209295&pc=DSQ | 3007042319-2025-00092 | 2024-12-11 05:00 | Death | HEARTWARE, INC. | 2025-01-22 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT HAS EXPIRED DUE TO MULTI ORGAN SYSTEM FAILURE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING MULTI-ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, INCLUDING THE OCCURRENCE OF THE EVENT SEVERAL YEARS POST-IMPLANT AND THE STATEMENT BY THE PHYSICIAN THAT THE EVENT IS UNRELATED TO THE DEVICE, THE DEVICE MOST LIKELY DID NOT CAUSE OF CONTRIBUTE TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, MULTI-ORGAN FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209446&pc=BZD | 2518422-2025-006752 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER (LUNG AND PANCREATIC CANCER). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209713&pc=OZD | 1220648-2025-25708 | 2025-01-06 05:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA CP | Positioning Failure | Sepsis | P140003 |  | 1 | Event Description: A COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT. THE IMPELLA WAS DISLOCATED COMPLETELY IN AORTA. INITIALLY, THEY PULLED THE PUMP BACK COMPLETELY INTO THE AORTA AND LET IT RUN ON LOW P-LEVEL. SUDDENLY THE PATIENT BECAME VERY CRITICAL, HAD TO BE RESUSCITATED FOR 3 MINUTES. THE PUMP WAS REPOSITIONED WITH TRANS-ESOPHAGEAL ECHO AND IT WORKED WELL. POSITION WAS NOT PERFECT, BUT WELL ENOUGH TO LET IT RUN ON P6. ALSO, THE PATIENT IS SEPTIC. THE PATIENT EXPIRED WHILE ON SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF INFECTION/SEPSIS COULD NOT BE DETERMINED. THE ROOT CAUSE OF POSITIONING ISSUE WAS MOST LIKELY USE RELATED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209949&pc=DSQ | 2916596-2025-00382 | 2025-01-05 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209977&pc=DSQ | 2916596-2025-00389 | 2025-01-04 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEMORRHAGIC STROKE. THE DEATH WAS RELATED TO THE PATIENT'S HEALTH CONDITION. THE DEVICE WAS WORKING AS INTENDED DURING TIME OF SUPPORT. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210027&pc=QLG | 2954323-2025-02611 | 2024-12-02 05:00 | Death | ABBOTT DIABETES CARE INC | 2025-01-22 05:00 | QLG | FREESTYLE LIBRE 2 | Device Alarm System | Hypoglycemia | K193371 |  | 1 | Event Description: ABBOTT DIABETES CARE (ADC) RECEIVED A COMPLAINT FROM A RELATIVE THAT ON (B)(6) 2024, THE CUSTOMER PASSED AWAY. THE REPORTER DISCOVERED THE DECEDENT AT APPROXIMATELY 7:00 AM; HOWEVER, IT WAS INDICATED THAT THE CUSTOMER PASSED AWAY AT APPROXIMATELY AT 5:00 AM. THE RELATIVE ALLEGED THAT THE ALARM DID NOT ANNUNCIATE ON THE FREESTYLE LIBRE 2 DEVICE. ADDITIONALLY, IT WAS REPORTED THAT THE CUSTOMER WAS BLIND AND DEPENDED ON THE LOW BLOOD GLUCOSE (BG) ALARMS FOR TREATMENT DECISIONS. THE RELATIVE ALLEGED THAT THE CUSTOMER DIED FROM LOW BLOOD SUGAR; HOWEVER, THERE WAS NO OFFICIAL CAUSE OF DEATH REPORTED, AND NO DEATH CERTIFICATE WAS PROVIDED. IN ADDITION, IT WAS INDICATED THAT THE CUSTOMER EXPERIENCED KIDNEY FAILURE, HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED REGARDING THE TIMELINE OR IMPACT THIS HAD ON THE EVENT. ADC HAS REQUESTED THAT THE DEVICE BE RETURNED FOR A PHYSICAL INVESTIGATION. MULTIPLE FOLLOW UP ATTEMPTS HAVE BEEN MADE TO CONTACT THE REPORTER, AND THUS-FAR HAVE BEEN UNSUCCESSFUL. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT CANNOT BE REASONABLY CONCLUDED THAT THE FREESTYLE LIBRE 2 DEVICE HAS LED TO OR CONTRIBUTED TO THE DEATH. A FOLLOW UP WILL BE SENT IF FURTHER INFORMATION IS OBTAINED. Manufacturer Narrative: THE REPORTED PRODUCT IS NOT EXPECTED TO BE RETURNED AS THE CUSTOMER DECLINED TO RETURN THE DEVICE. AN INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND IT HAS BEEN DETERMINED THAT THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. THE DHRS (DEVICE HISTORY REVIEW) FOR THE FREESTYLE LIBRE READER WERE REVIEWED AND THE DHRS SHOWED THE FREESTYLE LIBRE READER PASSED ALL TESTS PRIOR TO RELEASE. THE DHRS FOR THE PRODUCT WERE REVIEWED AND THE DHRS SHOWED THE PRODUCT PASSED ALL TESTS PRIOR TO RELEASE. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. NO FURTHER INVESTIGATION IS PLANNED. IN THE EVENT THAT PRODUCT IS RECEIVED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210032&pc=NPT | 2025587-2025-00531 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-22 05:00 | NPT | EVOLUT PRO PLUS VALVE | Fluid/Blood Leak; Perivalvular Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem | Obstruction/Occlusion; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE DEPLOYMENT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON. FOLLOWING VALVE DEPLOYMENT, THE VALVE DISLODGED PROXIMALLY, LEADING TO SEVERE PARA VALVULAR LEAK. A SNARE WAS USED TO TEMPORARILY STABILIZE THE VALVE AT THE ASCENDING AORTA WHILE A SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED; HOWEVER, UPON ADVANCING THE DCS, THE DISLODGED VALVE DISPLACED FURTHER AND BECAME ANCHORED AT THE SINUS LEVEL, CAUSING AN OCCLUSION OF BOTH THE RIGHT AND LEFT CORONARY ARTERIES. MULTIPLE ATTEMPTS WERE MADE USING TWO SEPARATE SNARES AND BALLOONS TO REPOSITION AND RETRIEVE THE ANCHORED VALVE FROM THE SINUS OF VALSALVA, ALONG WITH PERFORMING CARDIOPULMONARY RESUSCITATION; HOWEVER, THE INTERVENTIONS WERE UNSUCCESSFUL. SUBSEQUENTLY, THE PATIENT DIED THE SAME DAY. PER THE PHYSICIAN, THE DCS CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH, AND THE CAUSE OF DEATH WAS THE VALVE OCCLUDING BOTH THE RIGHT AND THE LEFT CORONARY ARTERIES. Manufacturer Narrative: PRODUCT ANALYSIS: UPON RECEIPT OF THE SUSPECT COMPLAINT DEVICE AT MEDTRONIC'S QUALITY LABORATORY, VISUAL ANALYSIS SHOWED THE DELIVERY CATHETER SYSTEM WAS RECEIVED WITHOUT A VALVE LOADED WITHIN THE CAPSULE. THE DEVICE WAS RECEIVED WITH THE NOSECONE OVER-CAPTURED BY THE CAPSULE. THE CAPSULE APPEARED INTACT WITH NO EVIDENCE OF DAMAGE. THE HANDLE APPEARED INTACT. THE DEVICE WAS RETURNED WITH THE END CAP/SCREW GEAR SNAP FIT CONNECTED. THE HANDLE APPEARED TO RETRACT AND ADVANCE THE CAPSULE. THE TRIGGER MOVED TO FULLY ADVANCED AND RETRACTED POSITIONS AND LOCKED IN PLACE WHEN RELEASED. THE TIP-RETRIEVAL MECHANISM APPEARED INTACT. THE INNER MEMBER SHAFT AND SPINDLE HUB APPEARED INTACT WITH NO EVIDENCE OF DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: THREE MEDIA FILES WERE PROVIDED FOR REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS NOT PROVIDED FOR REVIEW THUS IT IS NOT POSSIBLE TO VALIDATE A GOOD LOAD. THE FIRST MEDIA FILE SHOWS A RELEASED VALVE THAT APPEARS TO BE UNDER EXPANDED AT AN UNDETERMINED DEPTH. OF NOTE, IF A VALVE IS UNDER-EXPANDED: REMOVE SYSTEM, PERFORM PRE-DILATATION, AND IMPLANT NEW VALVE WITH NEW DELIVERY SYSTEM. IN THIS CASE, DESPITE UNDER EXPANSION, THE VALVE WAS RELEASED. EVIDENCE SUGGESTS THAT DURING REMOVAL OF THE DELIVERY CATHETER SYSTEM THE NOSE CONE MIGHT HAVE CONTRIBUTED TO THE DISLODGMENT OF THE VALVE. MEDTRONIC RECOMMENDS CAUTION WHILE WITHDRAWING THE DELIVERY CATHETER SYSTEM TO MINIMIZE INTERACTION WITH THE EVOLUT FRAME. A SNARE WAS USED TO REPOSITION THE VALVE IN THE ASCENDING AORTA, AND A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED. IT WAS REPORTED THAT THE DISLODGED VALVE SHIFTED TO THE LEVEL OF THE SINUSES OF VALVE DURING ADVANCEMENT OF THE SECOND DELIVERY CATHETER SYSTEM CAUSING CORONARY OCCLUSION AND SUBSEQUENT DEATH. IMAGES OF THE SECOND VALVE IMPLANT ATTEMPT WERE NOT PROVIDED THUS IT IS NOT POSSIBLE TO VALIDATE CAUSE OF DEATH. HOWEVER, AS STATED IN THE INSTRUCT IONS FOR USE: PHYSICIANS SHOULD USE JUDGMENT WHEN CONSIDERING REPOSITIONING A FULLY DEPLOYED BIOPROSTHESIS (FOR EXAMPLE, USING A SNARE, BALLOON, AND/OR FORCEPS). REPOSITIONING THE BIOPROSTHESIS IS NOT RECOMMENDED, EXCEPT IN CASES WHERE IMMINENT SERIOUS HARM OR DEATH IS POSSIBLE (FOR EXAMPLE, CORONARY OCCLUSION). REPOSITIONING OF A DEPLOYED VALVE MAY CAUSE AORTIC ROOT DAMAGE, CORONARY ARTERY DAMAGE, MYOCARDIAL DAMAGE, VASCULAR COMPLICATIONS, PROSTHETIC VALVE DYSFUNCTION (INCLUDING DEVICE MALPOSITION), EMBOLIZATION, STROKE, AND/OR EMERGENT SURGERY. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: D-EVPROP23-29, SERIAL/LOT #: (B)(6), UDI#: (B)(4). OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: EVPROPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: L-EVPROP23-29 (LOT: 0012395511); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: D-EVPROP23-29 (LOT: 0012436660); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210059&pc=CBK | 2518422-2025-101055 | 2022-12-11 05:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | CBK | TRILOGY | Electrical /Electronic Property Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MWMW5164583) IN REFERENCE TO A 14-YEAR OLD DISABLED CHILD USING TRILOGY VENTILATOR HAD A PLUGGED TRACH FROM THICK LUNG SECRETIONS IN THE VEHICLE (SO ON BATTERY) THE VENT DID NOT ALARM TO ALERT OF THE PLUG AND AS SUCH THE PATIENT CODED. SHE WAS BROUGHT BACK AND THEN PROMPTLY UPON PUTTING HER BACK ON THE VENT, THE SAME ISSUE HAPPENED BECAUSE WE HAD NO WAY TO KNOW IT WAS A PLUG WITH THE VENT NOT ALARMING. SHE WAS AGAIN BROUGHT BACK AND THE DAMAGE SUSTAINED DURING RESUSCITATION EFFORTS EVENTUALLY ENDED IN DEATH. SHE WOULD STILL BE HERE IF THE VENT ALARMED FOR THE PLUG AND WE COULD HAVE ATTEMPTED SUCTIONING. WE HAD NO IDEA THAT SUCTIONING WAS NEEDED BECAUSE THE VENT DID "WAS FINE" IF REQUESTED WE CAN GET THIS INFORMATION DIRECT FROM THE HOSPITAL OR SIGN A RELEASE FORM FOR YOU TO GATHER RECORDS." THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH (MWMW5164583) IN REFERENCE TO A 14-YEAR OLD DISABLED CHILD USING TRILOGY VENTILATOR HAD A PLUGGED TRACH FROM THICK LUNG SECRETIONS IN THE VEHICLE (SO ON BATTERY) THE VENT DID NOT ALARM TO ALERT OF THE PLUG AND AS SUCH THE PATIENT CODED. SHE WAS BROUGHT BACK AND THEN PROMPTLY UPON PUTTING HER BACK ON THE VENT, THE SAME ISSUE HAPPENED BECAUSE WE HAD NO WAY TO KNOW IT WAS A PLUG WITH THE VENT NOT ALARMING. SHE WAS AGAIN BROUGHT BACK AND THE DAMAGE SUSTAINED DURING RESUSCITATION EFFORTS EVENTUALLY ENDED IN DEATH. SHE WOULD STILL BE HERE IF THE VENT ALARMED FOR THE PLUG AND WE COULD HAVE ATTEMPTED SUCTIONING. WE HAD NO IDEA THAT SUCTIONING WAS NEEDED BECAUSE THE VENT DID "WAS FINE" IF REQUESTED WE CAN GET THIS INFORMATION DIRECT FROM THE HOSPITAL OR SIGN A RELEASE FORM FOR YOU TO GATHER RECORDS." ADDITIONAL INFORMATION WAS OBTAINED FOR THIS RECORD. PER THE REPLIES FROM THE EMAIL RECEIVED, IT WAS STATED THAT THERE WAS NO ALARM SOUND AT THE TIME OF THIS EVENT. THEY STATED THAT "ALL ALARMS WERE OPERATIONAL WHEN DOING ALARM CHECKS BY PULLING THE TUBE OFF THE SIDE OF THE VENT". THE ALARM SOUND VOLUME WAS FULL. THE DEVICE WAS MAINTAINED BY THE HOSPITAL. THEY STATED THAT "NO ALARM SOUND, NO RED LINE" APPEARED ON THE SCREEN. THE PATIENT WENT TO A PARTY AND LEFT THE PARTY HAPPY. WHILE ON TRANSPORT, THE PATIENT COULD NOT SPEAK LOUD ENOUGH WHEN THE PHLEGM PLUG HAPPENED TO ALERT ANYONE AND BY THEN SHE WAS CHOKING AND TRYING TO FRANTICALLY KICK. THEY PULLED INTO AN AMBULANCE BAY NOTICING THAT SHE WAS TURNING BLUE. NO ALARMS WERE GOING OFF, SO THEY HAD NO REASON TO SUSPECT A VENT ISSUE. THE PATIENT CODED AND WAS RESUSCITATED USING AN AMBU BAG, THE VENT WAS RECONNECTED AND SOUNDED FINE BUT THEN THE PATIENT CODED AGAIN BECAUSE OF THE PLUG THAT WAS GOING UNNOTICED AND UNARMED. THE PATIENT WAS REVIVED AGAIN, BUT BECAUSE OF RECITATION EFFORTS ON HER MUSCULAR DYSTROPHY, HER BODY FAILED FOUR DAYS LATER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. AT THIS TIME, A FINAL REPORT WILL BE SUBMITTED BUT IF ANY ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210107&pc=NPT | 2025587-2025-00533 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-22 05:00 | NPT | EVOLUT PRO PLUS DCS | Failure to Advance; Device Dislodged or Dislocated; Patient Device Interaction Problem | Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE DEPLOYMENT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING AN 18 MM NON-MEDTRONIC BALLOON. FOLLOWING VALVE DEPLOYMENT, THE VALVE DISLODGED PROXIMALLY, LEADING TO SEVERE PARA VALVULAR LEAK. A SNARE WAS USED TO TEMPORARILY STABILIZE THE VALVE AT THE ASCENDING AORTA WHILE A SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED; HOWEVER, UPON ADVANCING THE DCS, THE DISLODGED VALVE DISPLACED FURTHER AND BECAME ANCHORED AT THE SINUS LEVEL, CAUSING AN OCCLUSION OF BOTH THE RIGHT AND LEFT CORONARY ARTERIES. MULTIPLE ATTEMPTS WERE MADE USING TWO SEPARATE SNARES AND BALLOONS TO REPOSITION AND RETRIEVE THE ANCHORED VALVE FROM THE SINUS OF VALSALVA, ALONG WITH PERFORMING CARDIOPULMONARY RESUSCITATION; HOWEVER, THE INTERVENTIONS WERE UNSUCCESSFUL. SUBSEQUENTLY, THE PATIENT DIED THE SAME DAY. PER THE PHYSICIAN, THE DCS CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH, AND THE CAUSE OF DEATH WAS THE VALVE OCCLUDING BOTH THE RIGHT AND THE LEFT CORONARY ARTERIES. Manufacturer Narrative: PRODUCT ANALYSIS: UPON RECEIPT OF THE SUSPECT COMPLAINT DEVICE AT MEDTRONIC'S QUALITY LABORATORY, VISUAL ANALYSIS SHOWED THE DELIVERY CATHETER SYSTEM WAS RECEIVED WITHOUT A VALVE LOADED WITHIN THE CAPSULE. THE DEVICE WAS RECEIVED WITH THE CAPSULE PARTIALLY OPENED. THERE WAS A SLIGHT BEND TO THE NOSE CONE TIP AND A SLIGHT BEND TO THE CAPSULE. DELAMINATION WAS OBSERVED OVER THE NITINOL REINFORCING FRAME ALONG THE CAPSULE. THERE WERE BENDS TO THE STABILITY SHAFT. THERE WAS SLIGHT TEARING AND STRESS MARKING TO THE INLINE SHEATH. THE HANDLE APPEARED INTACT. THE DEVICE WAS RETURNED WITH THE END CAP/SCREW GEAR SNAP FIT CONNECTED. THE TIP RETRACT COVERS WERE RETURNED DETACHED AND COULD BE REATTACHED WITHOUT ISSUE. THE HANDLE APPEARED TO RETRACT AND ADVANCE THE CAPSULE. THE TRIGGER MOVED TO FULLY ADVANCED AND RETRACTED POSITIONS AND LOCKED IN PLACE WHEN RELEASED. THE TIP-RETRIEVAL MECHANISM APPEARED INTACT. THERE WERE TWO KINKS TO THE PROXIMAL END OF THE INNER MEMBER SHAFT. THE SPINDLE HUB APPEARED INTACT WITH NO EVIDENCE OF DAMAGE. UPDATED DATA: D9. H3. H6. CORRECTED DATA: H6. A150204 REMOVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: THREE MEDIA FILES WERE PROVIDED FOR REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS NOT PROVIDED FOR REVIEW THUS IT IS NOT POSSIBLE TO VALIDATE A GOOD LOAD. THE FIRST MEDIA FILE SHOWS A RELEASED VALVE THAT APPEARS TO BE UNDER EXPANDED AT AN UNDETERMINED DEPTH. OF NOTE, IF A VALVE IS UNDER-EXPANDED: REMOVE SYSTEM, PERFORM PRE-DILATATION, AND IMPLANT NEW VALVE WITH NEW DELIVERY SYSTEM. IN THIS CASE, DESPITE UNDER EXPANSION, THE VALVE WAS RELEASED. EVIDENCE SUGGESTS THAT DURING REMOVAL OF THE DELIVERY CATHETER SYSTEM THE NOSE CONE MIGHT HAVE CONTRIBUTED TO THE DISLODGMENT OF THE VALVE. MEDTRONIC RECOMMENDS CAUTION WHILE WITHDRAWING THE DELIVERY CATHETER SYSTEM TO MINIMIZE INTERACTION WITH THE EVOLUT FRAME. A SNARE WAS USED TO REPOSITION THE VALVE IN THE ASCENDING AORTA, AND A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED. IT WAS REPORTED THAT THE DISLODGED VALVE SHIFTED TO THE LEVEL OF THE SINUSES OF VALVE DURING ADVANCEMENT OF THE SECOND DELIVERY CATHETER SYSTEM CAUSING CORONARY OCCLUSION AND SUBSEQUENT DEATH. IMAGES OF THE SECOND VALVE IMPLANT ATTEMPT WERE NOT PROVIDED THUS IT IS NOT POSSIBLE TO VALIDATE CAUSE OF DEATH. HOWEVER, AS STATED IN THE INSTRUCT IONS FOR USE: PHYSICIANS SHOULD USE JUDGMENT WHEN CONSIDERING REPOSITIONING A FULLY DEPLOYED BIOPROSTHESIS (FOR EXAMPLE, USING A SNARE, BALLOON, AND/OR FORCEPS). REPOSITIONING THE BIOPROSTHESIS IS NOT RECOMMENDED, EXCEPT IN CASES WHERE IMMINENT SERIOUS HARM OR DEATH IS POSSIBLE (FOR EXAMPLE, CORONARY OCCLUSION). REPOSITIONING OF A DEPLOYED VALVE MAY CAUSE AORTIC ROOT DAMAGE, CORONARY ARTERY DAMAGE, MYOCARDIAL DAMAGE, VASCULAR COMPLICATIONS, PROSTHETIC VALVE DYSFUNCTION (INCLUDING DEVICE MALPOSITION), EMBOLIZATION, STROKE, AND/OR EMERGENT SURGERY. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. CONTINUATION OF D10: PRODUCT ID: EVPROPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. PRODUCT ID: EVPROPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: L-EVPROP23-29 (LOT: 0012395511); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID: D-EVPROP23-29 (LOT: 0012436660); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210115&pc=NPT | 2025587-2025-00534 | 2025-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-22 05:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Chest Pain; Myocardial Infarction; Cardiac Tamponade; Vascular Dissection; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A PREVIOUS A SURGICAL AORTIC VALVE REPLACEMENT IN THE AORTIC POSITION UNDERWENT AN ATTEMPTED TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, THE TRANSCATHETER VALVE WAS DEPLOYED TO 80% WHEN THE PATIENT EXPERIENCED SEVERE CHEST PAIN AND ST ELEVATION. THE VALVE WAS RECAPTURED AND REMOVED, AND FURTHER INVESTIGATION WAS CONDUCTED. IMAGING, INCLUDING TRANSTHORACIC ECHOCARDIOGRAM AND TRANSESOPHAGEAL ECHOCARDIOGRAM, REVEALED A LACERATION OF THE LEFT VENTRICLE, WHICH RESULTED IN A PERICARDIAL EFFUSION THAT LED TO TAMPONADE. PER THE PHYSICIAN, THE LACERATION WAS CAUSED BY THE WORKING WIRE. DUE TO ADHESIONS FROM PREVIOUS BYPASS SURGERY, NO PERICARDIAL CENTESIS WAS PERFORMED. THE EFFUSION CLOTTED OFF, CAUSING THE RIGHT VENTRICLE TO COLLAPSE. CONTRIBUTING FACTORS INCLUDED PATIENT ANATOMY AND POSSIBLE FRAIL TISSUE IN THE LEFT VENTRICLE. ANESTHESIA WAS CALLED TO INTUBATE THE PATIENT TO STABILIZE THE AIRWAY. THE PROCEDURE WAS ULTIMATELY ABORTED, AND THE PATIENT DIED. Manufacturer Narrative: UPDATED DATA: D4. H4. CORRECTED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID T505U225, (SERIAL: (B)(6) ); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2013; PRODUCT ID EVFXPLUS-26, (SERIAL: (B)(6) ); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210288&pc=NPT | 2025587-2025-00536 | 2025-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-22 05:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVOLUT FX+ VALVE, SEVERE CONSTRAINT WAS OBSERVED IN THE CUSP OVERLAP VIEW DURING INITIAL DEPLOYMENT TO A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 3 MM ON THE LEFT CORONARY CUSP (LCC). THE DECISION WAS MADE TO RELEASE THE VALVE, HOWEVER, THE VALVE SPONTANEOUSLY DISLODGED TO A DEPTH OF -3 MM ON THE NCC AND -3 ON THE LCC AFTER FULL RELEASE. THE DISLODGED VALVE OCCLUDED THE CORONARY ARTERIES. THE OCCLUSION NECESSITATED THE INITIATION OF CARDIOPULMONARY RESUSCITATION (CPR). THE VALVE WAS SNARED AND PULLED INTO THE ASCENDING AORTA, TEMPORARILY STABILIZING BLOOD PRESSURE AND HEART RATE. A SECOND VALVE DEPLOYMENT CAUSED THE ORIGINAL EMBOLIZED VALVE TO OCCLUDE THE CORONARY ARTERIES AGAIN, LEADING TO A SECOND CPR ATTEMPT, BUT STABLE BLOOD PRESSURE AND RHYTHM COULD NOT BE REGAINED. THE SECOND DELIVERY CATHETER SYSTEM AND VALVE WERE REMOVED WITHOUT DEPLOYMENT. ACCESS WAS GAINED VIA THE RIGHT RADIAL ARTERY TO ATTEMPT SNARING THE EMBOLIZED VALVE FROM A DIFFERENT ANGLE, BUT THIS WAS UNSUCCESSFUL. A 22 MM NON-MEDTRONIC BALLOON WAS INSERTED TO ATTEMPT PULLING BACK THE WEDGED VALVE, BUT THIS WAS ALSO UNSUCCESSFUL. THE PATIENT ANATOMY, INCLUDING A SHORT ASCENDING AORTA AND SMALL SINOTUBULAR JUNCTION, ALONG WITH MILD CALCIUM AND FIBROTIC LEAFLETS, WERE IDENTIFIED AS CONTRIBUTING FACTORS. THE PATIENT ULTIMATELY DIED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVOLUT FX+ VALVE, SEVERE CONSTRAINT WAS OBSERVED IN THE CUSP OVERLAP VIEW DURING INITIAL DEPLOYMENT TO A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 3 MM ON THE LEFT CORONARY CUSP (LCC). THE DECISION WAS MADE TO RELEASE THE VALVE, HOWEVER, THE VALVE SPONTANEOUSLY DISLODGED TO A DEPTH OF -3 MM ON THE NCC AND -3 ON THE LCC AFTER FULL RELEASE. THE DISLODGED VALVE OCCLUDED THE CORONARY ARTERIES. THE OCCLUSION NECESSITATED THE INITIATION OF CARDIOPULMONARY RESUSCITATION (CPR). THE VALVE WAS SNARED AND PULLED INTO THE ASCENDING AORTA, TEMPORARILY STABILIZING BLOOD PRESSURE AND HEART RATE. A SECOND VALVE DEPLOYMENT CAUSED THE ORIGINAL EMBOLIZED VALVE TO OCCLUDE THE CORONARY ARTERIES AGAIN, LEADING TO A SECOND CPR ATTEMPT, BUT STABLE BLOOD PRESSURE AND RHYTHM COULD NOT BE REGAINED. THE SECOND DELIVERY CATHETER SYSTEM (DCS) AND VALVE WERE REMOVED WITHOUT DEPLOYMENT. ACCESS WAS GAINED VIA THE RIGHT RADIAL ARTERY TO ATTEMPT SNARING THE EMBOLIZED VALVE FROM A DIFFERENT ANGLE, BUT THIS WAS UNSUCCESSFUL. A 22 MM NON-MEDTRONIC BALLOON WAS INSERTED TO ATTEMPT PULLING BACK THE WEDGED VALVE, BUT THIS WAS ALSO UNSUCCESSFUL. THE PATIENT ANATOMY, INCLUDING A SHORT ASCENDING AORTA AND SMALL SINOTUBULAR JUNCTION, ALONG WITH MILD CALCIUM AND FIBROTIC LEAFLETS, WERE IDENTIFIED AS CONTRIBUTING FACTORS. THE PATIENT ULTIMATELY DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID-PACED DURING THE POST-IMPLANT BALLOON DILATION. PER THE PHYSICIAN, THE DCS CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH, HOWEVER, THE DEATH WAS CAUSED BY THE CORONARY OCCLUSION. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210304&pc=NPT | 2025587-2025-00537 | 2025-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-22 05:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Obstruction/Occlusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVOLUT FX+ VALVE, SEVERE CONSTRAINT WAS OBSERVED IN THE CUSP OVERLAP VIEW DURING INITIAL DEPLOYMENT TO A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 3 MM ON THE LEFT CORONARY CUSP (LCC). THE DECISION WAS MADE TO RELEASE THE VALVE, HOWEVER, THE VALVE SPONTANEOUSLY DISLODGED TO A DEPTH OF -3 MM ON THE NCC AND -3 ON THE LCC AFTER FULL RELEASE. THE DISLODGED VALVE OCCLUDED THE CORONARY ARTERIES. THE OCCLUSION NECESSITATED THE INITIATION OF CARDIOPULMONARY RESUSCITATION (CPR). THE VALVE WAS SNARED AND PULLED INTO THE ASCENDING AORTA, TEMPORARILY STABILIZING BLOOD PRESSURE AND HEART RATE. A SECOND VALVE DEPLOYMENT CAUSED THE ORIGINAL EMBOLIZED VALVE TO OCCLUDE THE CORONARY ARTERIES AGAIN, LEADING TO A SECOND CPR ATTEMPT, BUT STABLE BLOOD PRESSURE AND RHYTHM COULD NOT BE REGAINED. THE SECOND DELIVERY CATHETER SYSTEM AND VALVE WERE REMOVED WITHOUT DEPLOYMENT. ACCESS WAS GAINED VIA THE RIGHT RADIAL ARTERY TO ATTEMPT SNARING THE EMBOLIZED VALVE FROM A DIFFERENT ANGLE, BUT THIS WAS UNSUCCESSFUL. A 22 MM NON-MEDTRONIC BALLOON WAS INSERTED TO ATTEMPT PULLING BACK THE WEDGED VALVE, BUT THIS WAS ALSO UNSUCCESSFUL. THE PATIENT ANATOMY, INCLUDING A SHORT ASCENDING AORTA AND SMALL SINOTUBULAR JUNCTION, ALONG WITH MILD CALCIUM AND FIBROTIC LEAFLETS, WERE IDENTIFIED AS CONTRIBUTING FACTORS. THE PATIENT ULTIMATELY DIED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE EVOLUT FX+ VALVE, SEVERE CONSTRAINT WAS OBSERVED IN THE CUSP OVERLAP VIEW DURING INITIAL DEPLOYMENT TO A DEPTH OF 2 MM ON THE NON-CORONARY CUSP (NCC) AND 3 MM ON THE LEFT CORONARY CUSP (LCC). THE DECISION WAS MADE TO RELEASE THE VALVE, HOWEVER, THE VALVE SPONTANEOUSLY DISLODGED TO A DEPTH OF -3 MM ON THE NCC AND -3 ON THE LCC AFTER FULL RELEASE. THE DISLODGED VALVE OCCLUDED THE CORONARY ARTERIES. THE OCCLUSION NECESSITATED THE INITIATION OF CARDIOPULMONARY RESUSCITATION (CPR). THE VALVE WAS SNARED AND PULLED INTO THE ASCENDING AORTA, TEMPORARILY STABILIZING BLOOD PRESSURE AND HEART RATE. A SECOND VALVE DEPLOYMENT CAUSED THE ORIGINAL EMBOLIZED VALVE TO OCCLUDE THE CORONARY ARTERIES AGAIN, LEADING TO A SECOND CPR ATTEMPT, BUT STABLE BLOOD PRESSURE AND RHYTHM COULD NOT BE REGAINED. THE SECOND DELIVERY CATHETER SYSTEM (DCS) AND VALVE WERE REMOVED WITHOUT DEPLOYMENT. ACCESS WAS GAINED VIA THE RIGHT RADIAL ARTERY TO ATTEMPT SNARING THE EMBOLIZED VALVE FROM A DIFFERENT ANGLE, BUT THIS WAS UNSUCCESSFUL. A 22 MM NON-MEDTRONIC BALLOON WAS INSERTED TO ATTEMPT PULLING BACK THE WEDGED VALVE, BUT THIS WAS ALSO UNSUCCESSFUL. THE PATIENT ANATOMY, INCLUDING A SHORT ASCENDING AORTA AND SMALL SINOTUBULAR JUNCTION, ALONG WITH MILD CALCIUM AND FIBROTIC LEAFLETS, WERE IDENTIFIED AS CONTRIBUTING FACTORS. THE PATIENT ULTIMATELY DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RAPID-PACED DURING THE POST-IMPLANT BALLOON DILATION. PER THE PHYSICIAN, THE DCS CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH, HOWEVER, THE DEATH WAS CAUSED BY THE CORONARY OCCLUSION. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: EVFXPLUS-26 (J300621); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210415&pc=DSQ | 2916596-2025-00286 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia; Loss of consciousness | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY (B)(6) 2025 FROM VENTRICULAR TACHYCARDIA (VT) ARREST. THE PATIENT WAS FOUND UNRESPONSIVE BY THEIR WIFE. THE PROVIDER DID NOT BELIEVE THE DEATH WAS RELATED TO THE HEARTMATE 3. THE DEVICE WAS BELIEVED TO BE FUNCTIONING AS INTENDED. THE PATIENT WILL NOT RECEIVE AN AUTOPSY, AND THE DEVICE WOULD NOT BE RETURNED.Event Description: IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF VENTRICULAR TACHYCARDIA PRIOR TO IMPLANT. THE PATIENT ALSO HAD A MEDTRONIC IMPLANTABLE CARDIOVERTER-DEFIBRILLATER (ICD). Manufacturer Narrative: B5: ADDITIONAL INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) AND THE REPORTED EVENT COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT NO AUTOPSY WILL BE PERFORMED, AND THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION", LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING CARDIAC ARRHYTHMIA. ADDITIONALLY, SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210881&pc=FKX | 0008030665-2025-00182 | 2025-01-15 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-22 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding | K222318 |  | 1 | Event Description: A PERITONEAL REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO AN INTRACEREBRAL HEMORRHAGE (ICH). IT WAS UNCLEAR WHETHER THE PATIENT WAS CONNECTED TO THEIR LIBERTY SELECT CYCLER ON OR AROUND THE TIME OF DEATH AS A PATIENT CONTACT REPORTED TO THE OUTPATIENT CLINIC THAT THE CYCLER WAS UNPLUGGED WHEN HE WAS FOUND UNRESPONSIVE ON HIS BEDROOM FLOOR. NO FURTHER INFORMATION WAS PROVIDED BY FAMILY TO THE OUTPATIENT CLINIC CONCERNING THIS EVENT. IT WAS UNKNOWN WHETHER EMERGENCY SERVICES WERE ACTIVATED BUT IT WAS AFFIRMED THAT THE PATIENT WAS PRONOUNCED DECEASED IN HIS HOME WITH NO TRANSPORT TO HOSPITAL. IT WAS CONFIRMED THAT THE PATIENT¿S ICH AND SUBSEQUENT DEATH WERE UNRELATED TO CCPD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S).Event Description: A PERITONEAL REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO AN INTRACEREBRAL HEMORRHAGE (ICH). IT WAS UNCLEAR WHETHER THE PATIENT WAS CONNECTED TO THEIR LIBERTY SELECT CYCLER ON OR AROUND THE TIME OF DEATH AS A PATIENT CONTACT REPORTED TO THE OUTPATIENT CLINIC THAT THE CYCLER WAS UNPLUGGED WHEN HE WAS FOUND UNRESPONSIVE ON HIS BEDROOM FLOOR. NO FURTHER INFORMATION WAS PROVIDED BY FAMILY TO THE OUTPATIENT CLINIC CONCERNING THIS EVENT. IT WAS UNKNOWN WHETHER EMERGENCY SERVICES WERE ACTIVATED BUT IT WAS AFFIRMED THAT THE PATIENT WAS PRONOUNCED DECEASED IN HIS HOME WITH NO TRANSPORT TO HOSPITAL. IT WAS CONFIRMED THAT THE PATIENT¿S ICH AND SUBSEQUENT DEATH WERE UNRELATED TO CCPD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: IT IS UNKNOWN WHETHER A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CAN BE ATTRIBUTED TO AN ACUTE ICH WITH NO INDICATION THIS EVENT AND HIS SUBSEQUENT DEATH WERE DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. CORRECTION PROVIDED IN G4. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE: HEATER TRAY DAMAGED (PAINT). THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210984&pc=DSQ | 2916596-2025-00232 | 2024-12-16 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT ARRIVED DEAD AT THE IMPLANTING CENTER FROM AN OUTSIDE EMERGENCY ROOM. THE PATIENT WAS OPERATING AS INTENDED AND APPEARED TO HAVE DIED FROM ADVANCED STAGES OF HEART FAILURE. AN AUTOPSY WAS PERFORMED AND THE DEVICE WAS EXPLANTED. THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR WAS TO RETURN THE PUMP. THE SYSTEM CONTROLLER WOULD ALSO BE RETURNED. AN EXPLANT KIT WAS REQUESTED. THE PATIENT HAD ARRIVED TO THE OUTSIDE EMERGENCY DEPARTMENT WITH SHORTNESS OF BREATH REQUIRING INTUBATION. THE AUTOPSY CONFIRMED HEART FAILURE RELATED CAUSES. THIS WAS DUE TO THE PROGRESSION OF RIGHT HEART FAILURE. THIS WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD BE RETURNED.Event Description: THE RIGHT HEART FAILURE DID NOT EXIST PRE-VAD IMPLANT. A DEVICE RELATED ISSUE DID NOT CAUSE/CONTRIBUTE TO THE RIGHT HEART FAILURE. THE PATIENT HAD SYMPTOMS OF SHORTNESS OF BREATH. THE RIGHT HEART FAILURE WAS NOT TREATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION G2: REPORT SOURCE CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE 4641 - UNEXPECTED MEDICAL INTERVENTION ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6), WAS RETURNED ASSEMBLED WITH THE PUMP CABLE SEVERED APPROXIMATELY 4¿¿ FROM THE PUMP HEADER. THE REMAINDER OF THE PUMP CABLE (INCLUDING THE INLINE CONNECTOR) AND THE MODULAR CABLE WERE NOT RETURNED. THE OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT, WITH THE OUTFLOW GRAFT BEND RELIEF ENGAGED TO THE GRAFT ATTACHMENT HARDWARE. THE OUTFLOW GRAFT AND BEND RELIEF MATERIAL HAD BEEN SEVERED NEAR THEIR RESPECTIVE HARDWARE. THE OUTFLOW GRAFT CLIP WAS RETURNED PROPERLY SEATED OVER THE OUTFLOW GRAFT HARDWARE. THE MINI APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED. (B)(6) APPEARED TO HAVE BEEN EXPOSED TO A FIXATIVE AGENT PRIOR TO THE PUMP¿S RETURN FOR EVALUATION. UPON DISASSEMBLY, VISUAL EXAMINATION OF THE PUMP¿S BLOOD-CONTACTING SURFACES DID NOT REVEAL ANY EVIDENCE OF DEPOSITIONS OR THROMBUS FORMATIONS. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE LEFT VENTRICULAR ASSIST DEVICE EVENT AND PERIODIC LOG FILES WERE RETRIEVED FROM (B)(6) AND CAPTURED THE PUMP FUNCTIONING AS INTENDED. THE PUMP WAS CLEANED, REASSEMBLED, AND OPERATED ON A MOCK CIRCULATORY LOOP. THE DEVICE FUNCTIONED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211230&pc=DSQ | 2916596-2025-00475 | 2019-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿ASSESSMENT OF PLATELET RESPONSE TO ASPIRIN THERAPY AND HEMOCOMPATIBILITY-RELATED ADVERSE EVENTS IN HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE RECIPIENTS¿ THAT HEARTMATE 3 (HM3) IMPLANT WAS ASSOCIATED WITH OUTCOMES SUCH AS HEMOCOMPATIBILITY RELATED ADVERSE EVENTS (HRAES) INCLUDING THROMBOSIS, STROKE AND BLEEDING. PATIENTS RECEIVING LEFT VENTRICULAR ASSIST DEVICES (LVADS) TYPICALLY UNDERGO ANTICOAGULATION AND ASPIRIN (ASA) THERAPY. THE HM3 HAS SHOWN PROMISE IN REDUCING HEMOCOMPATIBILITY-RELATED ADVERSE EVENTS (HRAES) LIKE STROKE, THROMBOSIS AND BLEEDING. HOWEVER, IT REMAINS UNCERTAIN WHETHER ASA RESPONSIVENESS IMPACTS THE INCIDENCE OF HRAES AND IF ASA SENSITIVITY CHANGES DURING HM3 SUPPORT. IN THIS PROSPECTIVE SINGLE-CENTER COHORT STUDY, 32 HM3 LVAD PATIENTS (AGE: 58.9 ± 9.9 YEARS, FEMALE: 15.6%, BMI: 29.1 ± 5.4KG/M2) IMPLANTED BETWEEN 2019 AND 2022 WERE ENROLLED. SERIAL ASA PLATELET SENSITIVITY TESTING WAS CONDUCTED. RESISTANCE TO ASA ANTIPLATELET THERAPY WAS DEFINED AS ASA REACTIVITY UNITS (ARU) > 550. THE PRIMARY ENDPOINT WAS THE INCIDENCE OF HRAES (ASA RESPONDERS VS. NONRESPONDERS) FROM 6 MONTHS BEFORE BASELINE UNTIL THE NEXT OUTPATIENT FOLLOW UP VISIT (3 MONTHS ± 30 DAYS). SECONDARY ENDPOINTS WERE CHANGES IN ASA RESISTANCE OVER TIME. AT THE FIRST ARU ASSESSMENT (361 DAYS POST-HM3 IMPLANTATION), 13 (40.6%) PATIENTS WERE ASA RESISTANT, WITH 28.6% (P=0.22) DEMONSTRATING RESISTANCE DURING FOLLOW-UP. MEDIAN (IQR) ARU VALUES SIGNIFICANTLY (P=0.03) CHANGED OVER TIME: 535 (396-590) VS. 497 (412-575). ASA RESPONDERS AND NON-RESPONDERS RECEIVED COMPARABLE DAILY ASA DOSES PER-PROTOCOL (105 ± 23MG VS. 100 ± 0MG, P=0.82), WITHOUT SIGNIFICANT DIFFERENCES IN BASELINE DEMOGRAPHICS (AGE: 60.6 ± 8.2 VS. 56.4 ± 11.9 YEARS, P=0.24; FEMALE: 10.5% VS. 23.1%, P=0.37; BMI: 28.5 ± 6.4 VS. 30.1 ± 3.7KG/M2, P=0.36; ISCHEMIC HEART FAILURE: 55.6% VS. 45.5%, P=0.35) OR RISK FACTORS (ATRIAL FIBRILLATION: 21.1% VS. 15.4%, DIABETES: 21.1% VS. 30.8% AND STROKE HISTORY: 15.8% VS. 7.7%, P > 0.63). THE INCIDENCE OF HRAES WAS SIMILAR BETWEEN ASA RESPONDERS AND NON-RESPONDERS 6 MONTHS BEFORE THE INITIAL MEASUREMENT (15.8% VS. 0.0%, P=0.14) AND BETWEEN THE TWO ARU MEASUREMENTS (5.3% VS.7.7%, P=0.77). WHILE ASA RESISTANCE VARIES AMONG LVAD OUTPATIENTS AND OVER TIME, THE HM3 DEMONSTRATES EXCELLENT HEMOCOMPATIBILITY IN BOTH ASA RESPONDERS AND NON-RESPONDERS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: IT WAS DETERMINED THAT THIS EVENT DID NOT MEET THE REQUIREMENTS OF A COMPLAINT. NO DEATHS WERE REPORTED IN THE RESEARCH ARTICLE. THE HEARTMATE 3 LVAS IFU, REV. G IS CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE-RELATED ISSUES WERE REPORTED, SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A1-A6: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2019 AS THE IMPLANT DATES OCCURRED BETWEEN 2019-2022. B5: DEVICE WAS IMPLANTED AT TIME OF EVENT. SCHLOEGLHOFER, T., ASADI, H. A., ABART, T., SCHWARZ, C., SCHIMA, H., RIEBANDT, J., MARKO, C., WIEDEMANN, D., MESSNER, B., & ZIMPFER, D. (2024). ASSESSMENT OF PLATELET RESPONSE TO ASPIRIN THERAPY AND HEMOCOMPATIBILITY-RELATED ADVERSE EVENTS IN HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE RECIPIENTS. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 43(4), S282. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.02.1221. DEPARTMENT OF CARDIAC SURGERY, CENTER FOR MEDICAL PHYSICS AND BIOMEDICAL ENGINEERING, MEDICAL UNIVERSITY OF VIENNA, VIENNA, AUSTRIA; DEPARTMENT OF CARDIAC SURGERY, MEDICAL UNIVERSITY OF VIENNA, VIENNA, AUSTRIA; DEPARTMENT OF CARDIAC SURGERY, MEDICAL UNIVERSITY OF GRAZ AND MEDICAL UNIVERSITY OF VIENNA, GRAZ AND VIENNA, AUSTRIA. D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211325&pc=MVK | 3008642652-2025-00802 | 2024-11-24 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-22 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 10:01:36 ON (B)(6) 2024. AT 13:59:05, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING RATES/AMPLITUDES AND MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VF. VF WITH VARYING RATES/AMPLITUDES AND MOTION ARTIFACT PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 14:53:29 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211335&pc=DSQ | 2916596-2025-00197 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNCLEAR BUT IT DID NOT APPEAR TO BE DEVICE RELATED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. IT WAS COMMUNICATED ON (B)(6) 2025 THAT THE CAUSE OF DEATH WAS UNKNOWN. THE DEVICE WAS STATED TO HAVE OPERATED AS EXPECTED. Manufacturer Narrative: D4 - UDI: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211443&pc=DSQ | 2916596-2025-00229 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pneumonia; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PNEUMONIA AND SEPSIS.Event Description: IT WAS REPORTED THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND WAS STATED TO HAVE OPERATED AS EXPECTED. THERE WERE NO DETAILS PROVIDED ON CLINICAL SYMPTOMS/CULTURES IDENTIFIED. IT WAS NOTED THAT THERE WAS NO DRIVELINE OR DEVICE INFECTION IN PLACE AT TIME OF OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211647&pc=JOL | 1417592-2025-00031 | 2024-12-24 05:00 | Death | MEDLINE INDUSTRIES, LP | 2025-01-22 05:00 | JOL | YANKAUER,FLEXIBLE HANDLE,REGLR CAPACITY | Connection Problem | Aspiration/Inhalation |  |  | 1 | Event Description: ACCORDING TO THE FACILITY ON (B)(6) 2024 THERE WAS A DELAY IN SUCTIONING THE PATIENT AS THE YANKAUER WOULDN'T "PROPERLY CONNECT TO THE SUCTION TUBING". Manufacturer Narrative: ACCORDING TO THE FACILITY ON (B)(6) 2024 THERE WAS A DELAY IN SUCTIONING THE PATIENT AS THE YANKAUER WOULDN'T "PROPERLY CONNECT TO THE SUCTION TUBING". PER THE FACILITY THE DEVICE WAS BEING USED "IN PREPARATION FOR INTUBATION" AND "YANKAUER WAS CONNECTED TO SUCTION TUBING AND SUCTION TURNED ON THEN PLACED AT HEAD OF BED". PER THE FACILITY THE DEVICE WAS TESTED PRIOR TO USE AND THE SUCTION WAS "TURNED TO HIGHEST SETTING FOR MAX SUCTION". PER THE FACILITY THE "PATIENT ASPIRATED, REQUIRED BRONCHOSCOPY TO REMOVE ASPIRATED MATERIAL". PER THE FACILITY THE PATIENT HAS EXPIRED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IT HAS BEEN DETERMINED THAT THE REPORTED EVENT CAUSED OR CONTRIBUTED TO SERIOUS INJURY REQUIRING MEDICAL INTERVENTION, THEREFORE, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. MULTIPLE LOT NUMBERS WERE PROVIDED BY THE FACILITY AND WERE LISTED AS: 10001, 30001, 40001.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211955&pc=BZD | 2518422-2025-006779 | 2023-12-07 05:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. SISTER OF THE PATIENT CALLED IN STATING THAT SHE WILL RETURN THE DEVICE BECAUSE THE PATIENT PASSED AWAY. SHE WOULD LIKE TO RETURN THE OLD DEVICE. SHE THINKS THAT THE PATIENT DID NOT REGISTER THE OLD DEVICE AND SHE JUST WANTS TO GET A RETURN LABEL VIA EMAIL. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: IN THIS REPORT, SECTION H - EVALUATION METHOD, EVALUATION RESULTS AND CONCLUSION CODE HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21212142&pc=DSQ | 2916596-2025-00089 | 2024-12-23 05:00 | Death | THORATEC CORPORATION | 2025-01-22 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY AFTER A TIME IN HOSPICE. NO OTHER DETAILS OF THEIR DEATH WERE KNOWN. NO EQUIPMENT WAS RETURNING. THE VENTRICULAR ASSIST DEVICE WAS FUNCTIONING AS INTENDED. Manufacturer Narrative: SECTION H4 - DEVICE MANUFACTURE DATE: CORRECTED. SECTION H6 HEALTH EFFECT - IMPACT CODE: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21213648&pc=OZD | 1220648-2025-25719 | 2024-12-25 05:00 | Death | ABIOMED, INC. | 2025-01-22 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombocytopenia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DEVELOPED SHORTNESS OF BREATH AND SUBSEQUENTLY EXPERIENCED CARDIOPULMONARY ARREST WITH A DEMISED OUTCOME. THE PATIENT WAS ON INDEPENDENT SUPPORT BY THE IMPELLA 5.5 (RECENTLY WEANED OFF FROM EXTRACORPOREAL MEMBRANE OXYGENATION) AND HAD DEVELOPED SHORTNESS OF BREATH WHILE SITTING DURING REHABILITATION ON THE SEVENTH DAY OF IMPELLA MCS. AN ANGIOGRAPHY REVEALED A PULMONARY EMBOLISM IN THE MAIN PULMONARY ARTERY WHICH WAS DIFFICULT TO TREAT LEADING TO A DO NOT ATTEMPT RESUSCITATION. THE PATIENT WAS DIAGNOSED WITH HEPARIN INDUCED THROMBOCYTOPENIA (HIT) AT THE TIME OF THE ABOVE CONDITIONS AND IT WAS BELIEVED THAT THE HIT HAD CONTRIBUTED TO THE PULMONARY EMBOLISM. HEPARIN WAS IN THE PURGE LINE LEADING UP TO THE EVENT. THE PATIENT WOULD EXPIRE TWO DAYS LATER. REPORTEDLY, THE CAUSE OF DEATH WAS PULMONARY EMBOLISM, AND THE CAUSE OF HIT WAS THE USE OF HEPARIN. IT IS UNKNOWN OF THE HIT CONTRIBUTED TO THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED THROMBOCYTOPENIA AND THROMBUS HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE THROMBOCYTOPENIA AND THROMBUS COULD NOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ B.2 REVISED SELECTIONS AS AN INCORRECT SELECTION WAS PREVIOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25719 AND A NEW SELECTION WAS ADDED. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25719 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED FROM THE PREVIOUSLY SUBMITTED INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25719.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214103&pc=NVN | 2124215-2025-03064 | 2024-12-26 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-22 05:00 | NVN | INGEVITY+ | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, A CARDIAC TAMPONADE OCCURRED DURING LEAD INSERTION AND THE PATIENT EXPIRED ON THE OPERATING TABLE. THE ATTEMPTED LEAD WAS EXPECTED TO BE RETURNED FOR ANALYSIS.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, A CARDIAC TAMPONADE OCCURRED DURING LEAD INSERTION AND THE PATIENT EXPIRED ON THE OPERATING TABLE. THE ATTEMPTED LEAD WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE PRODUCT IS EXPECTED TO BE RETURNED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON RETURN AND COMPLETION OF ANALYSIS. UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, A THOROUGH EVALUATION OF THE LEAD WAS PERFORMED. TESTING WAS COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE AND INNER INSULATION INTEGRITY. MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS. MICROSCOPIC INSPECTIONS OF THE TERMINAL PIN ASSEMBLY, LEAD BODY, AND ELECTRODE TIP FOUND NO ANOMALIES. LABORATORY TESTING WAS UNABLE TO REPRODUCE THE REPORTED CLINICAL OBSERVATIONS AND DETAILED ANALYSIS DID NOT REVEAL ANY ABNORMALITIES. LABORATORY TESTING DID NOT IDENTIFY ANY LEAD CHARACTERISTICS THAT WOULD HAVE RESULTED IN THE CLINICAL OBSERVATION OF CARDIAC TAMPONADE. NO LEAD ISSUES WERE NOTED THAT WOULD HAVE MADE CARDIAC TAMPONADE MORE LIKELY THAN WHAT IS DESCRIBED IN THE INSTRUCTIONS FOR USE AS A KNOWN INHERENT RISK OF THE USE OF THIS PRODUCT. Manufacturer Narrative: THE PRODUCT IS EXPECTED TO BE RETURNED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON RETURN AND COMPLETION OF ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214420&pc=BZD | 2518422-2025-006805 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ASTHMA (NEW OR WORSENING), CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214452&pc=BZD | 2518422-2025-006815 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-22 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194700&pc=BZD | 2518422-2025-005374 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194707&pc=BZD | 2518422-2025-005386 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Inflammation; Pulmonary Dysfunction; Dizziness; Epistaxis; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, DEATH, NOSE BLEEDS, LUNG INFLAMMATION AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194744&pc=BZD | 2518422-2025-005402 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194758&pc=BZD | 2518422-2025-005395 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Inflammation; Dizziness; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194841&pc=BZD | 2518422-2025-005449 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Respiratory Failure; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, KIDNEY DISEASE/TOXICITY AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194850&pc=BZD | 2518422-2025-005396 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194881&pc=BZD | 2518422-2025-005493 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Inflammation; Dizziness; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194896&pc=BZD | 2518422-2025-005485 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Myocardial Infarction | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING HEART ATTACK AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194899&pc=BZD | 2518422-2025-005420 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Breast Cancer; Headache; Inflammation; Nausea; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Skin Inflammation/ Irritation; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER, OTHER (UNSPECIFIED EVENT), BREAST, MULTIPLE MYELOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194936&pc=BZD | 2518422-2025-005514 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, LUNG DISEASE, CANCER AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, INFLAMMATORY RESPONSE, LUNG DISEASE, CANCER AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX B: DATE OF DATE WAS CORRECTED. BOX B: DESCRIBED EVENT OR PROBLEM (B5 VERBIAGE) WAS CORRECTED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194941&pc=BZD | 2518422-2025-005535 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 | 317963345 | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN PREVIOUS INITIAL/FINAL REPORT, THE DATE OF DEATH WAS GIVEN WRONGLY, NOW IT WAS UPDATE CORRECTLY. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194982&pc=BZD | 2518422-2025-005452 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194997&pc=BZD | 2518422-2025-005541 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195009&pc=BZD | 2518422-2025-005570 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, CANCER, DEATH, CANCER, BLADDER AND PROSTATE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195039&pc=BZD | 2518422-2025-005565 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: IN PREVIOUS INITIAL/FINAL REPORT, THE DATE OF DEATH WAS GIVEN WRONGLY, NOW IT WAS UPDATED CORRECTLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195046&pc=BZD | 2518422-2025-005573 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Dyspnea | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND SHORTNESS OF BREATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195107&pc=BZD | 2518422-2025-005589 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195198&pc=BZD | 2518422-2025-004717 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195199&pc=BZD | 2518422-2025-005651 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF DEATH, LUNG DISEASE AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H: TYPE OF REPORTED COMPLAINT WAS CAPTURED SERIOUS INJURY IN PREVIOUS REPORT. TYPE OF REPORTED COMPLAINT WAS CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195201&pc=BZD | 2518422-2025-004784 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195317&pc=BZD | 2518422-2025-005673 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER.
THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195472&pc=BZD | 2518422-2025-005712 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195590&pc=BZD | 2518422-2025-005719 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Cancer; Lymphoma; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, LYMPHOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195650&pc=BZD | 2518422-2025-005782 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Inflammation; Dizziness; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195711&pc=BZD | 2518422-2025-005799 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195721&pc=BZD | 2518422-2025-005789 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Adult Respiratory Distress Syndrome | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND ARDS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195731&pc=BZD | 2518422-2025-005804 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Dizziness | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS AND HEADACHE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195784&pc=BZD | 2518422-2025-005525 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195792&pc=BZD | 2518422-2025-005255 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Headache; Hypersensitivity/Allergic reaction; Inflammation; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, REDUCED PULMONARY RESERVE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195797&pc=BZD | 2518422-2025-005784 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Heart Failure/Congestive Heart Failure | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195845&pc=BZD | 2518422-2025-100863 | 2025-01-14 05:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION HUMIDIFIER, GB-RENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE DREAMSTATION HUMIDIFIER DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.Event Description: THIS REPORT IS BEING SUBMITTED TO CORRECT THE DESCRIPTION OF THE EVENT OR PROBLEM PREVIOUSLY REPORTED. IT WAS REPORTED THAT THE DREAMSTATION HUMIDIFIER DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE IS NO EVIDENCE THAT THE DREAMSTATION HUMIDIFIER DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THIS COMPLAINT IS NOT REPORTABLE TO ANY REGULATORY AGENCIES. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED TO CORRECT THE DESCRIPTION OF THE EVENT OR PROBLEM PREVIOUSLY REPORTED. IT WAS REPORTED THAT THE DREAMSTATION HUMIDIFIER DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE IS NO EVIDENCE THAT THE DREAMSTATION HUMIDIFIER DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THIS COMPLAINT IS NOT REPORTABLE TO ANY REGULATORY AGENCIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195864&pc=BZD | 2518422-2025-005681 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAM STATION AUTO CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195867&pc=BZD | 2518422-2025-005836 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Gastrointestinal Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING STOMACH ISSUES, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195876&pc=BZD | 2518422-2025-005758 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH, THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195905&pc=KRH | 9612164-2025-00350 | 2024-12-23 05:00 | Death | MEDTRONIC MEXICO | 2025-01-21 05:00 | KRH | DURAN ANCORE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | K032810 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH A 27MM MITRAL DURAN ANCORE ANNULOPLASTY BAND DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE ANNULOPLASTY BAND COULD NOT BE EXCLUDED. IT WAS STATED THAT PRIOR TO DEATH, THE PATIENT WAS IN CARDIAC ARREST, AND ACTIVE CARDIOPULMONARY RESUSCITATION WAS BEING PERFORMED. UPON ARRIVAL TO THE FACILITY, THERE WAS NO HEART RATE DETECTED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195980&pc=BZD | 2518422-2025-005858 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21195982&pc=BZD | 2518422-2025-005856 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Hepatic or Biliary Problem; Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, APPENDIX, COLON, LIVER PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196100&pc=BZD | 2518422-2025-005894 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196120&pc=FTL | 1213643-2025-090027 |  | Death | DAVOL INC., SUB. C.R. BARD, INC. | 2025-01-21 05:00 | FTL | COMPOSIX MESH E/X | Defective Device; Patient Device Interaction Problem | Insufficient Information | K002684 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, COMPOSIX E/X ON (B)(6) 2008 AND COMPOSIX L/P ON (B)(6) 2011. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST COMPOSIX KUGEL HERNIA PATCH AND COMPOSIX E/X. ATTORNEY ALLEGES THAT THE PATIENT HAD SUBSEQUENT SURGICAL INTERVENTION DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING SUBSEQUENT SURGICAL INTERVENTION AND DEATH. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. THIS EMDR REPRESENTS BARD/DAVOL COMPOSIX MESH E/X (DEVICE #2). AN ADDITIONAL EMDR WAS SUBMITTED TO REPRESENT BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196121&pc=FTL | 1213643-2025-090026 |  | Death | DAVOL INC., SUB. C.R. BARD, INC. | 2025-01-21 05:00 | FTL | MESH ¿ COMPOSIX KUGEL | Defective Device; Patient Device Interaction Problem | Insufficient Information | K061314 |  | 1 | Event Description: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, COMPOSIX E/X ON (B)(6) 2008 AND COMPOSIX L/P ON (B)(6) 2011. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST COMPOSIX KUGEL HERNIA PATCH AND COMPOSIX E/X. ATTORNEY ALLEGES THAT THE PATIENT HAD SUBSEQUENT SURGICAL INTERVENTION DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. Manufacturer Narrative: NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING SUBSEQUENT SURGICAL INTERVENTION AND DEATH. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. THIS EMDR REPRESENTS BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1). AN ADDITIONAL EMDR WAS SUBMITTED TO REPRESENT BARD/DAVOL COMPOSIX MESH E/X (DEVICE #2). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196188&pc=BZD | 2518422-2025-005919 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196560&pc=OZD | 1220648-2025-25658 | 2024-08-26 04:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | OZD | IMPELLA 5.5 | Positioning Failure | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN AGE AND ETHNICITY) WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED POSITIONING ISSUES AND WOULD SUBSEQUENTLY EXPIRE. THE IMPELLA PUMP WAS SUCCESSFULLY IMPLANTED AND PATIENT ON THE UNIT WITH EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). APPROXIMATELY FOUR DAYS AFTER START OF SUPPORT, THE PHYSICIAN REPORTED A RECURRING ALARM ¿IMPELLA IN AORTA¿ AND "IMPELLA POSITION UNKNOWN." THE POSITION WAS NOTED "GOOD" AT 4.9CM IN THE VENTRICLE, PULSATILE CURVES. THE PATIENT WAS WITH HIS/HER OWN EJECTION FRACTION OF APPROXIMATELY 10-15% AND ON EXTRACORPOREAL MEMBRANE OXYGENATION THERAPY. SUPPORT CONTINUED UNTIL REPORTEDLY THE PATIENT EXPIRED APPROXIMATELY 19 DAYS LATER. THE DECISION WAS MADE TO WITHDRAW CARE. NO FURTHER INFORMATION REGARDING THE CAUSE OF DEATH WAS PROVIDED, AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RETURNED; THEREFORE, AN EVALUATION, ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. THE ROOT CAUSE OF POSITIONING ISSUE WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. Manufacturer Narrative: B.2 REQUIRED INTERVENTION WAS SELECTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25658. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25658 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196570&pc=OZD | 1220648-2025-25659 | 2024-11-21 05:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hematoma | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED AN ACCESS SITE COMPLICATION AND STROKE WITH A DEMISED OUTCOME. THE PATIENT PRESENTED WITH CONGESTIVE HEART FAILURE EXACERBATION. AN INTRA-AORTIC BALLOON PUMP WAS PLACED AND SHOWED WORSENING HEMODYNAMICS IN WHICH THE DECISION WAS MADE TO CANNULATE FOR VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). FOLLOWING, THE PATIENT WAS TRANSPORTED TO THE CARDIOVASCULAR OPERATING ROOM FOR AN AXILLARY IMPELLA 5.5 DEVICE PLACEMENT. THE IMPELLA IMPLANT PROCEDURE WAS SUCCESSFULLY COMPLETED, PATIENT EXTUBATED AND TRANSFERRED TO THE INTENSIVE CARE UNIT. APPROXIMATELY SEVEN DAYS AFTER START OF THE SUPPORT, AN ACCESS SITE HEMATOMA DEVELOPED. AN EVACUATION WAS PERFORMED TO THE HEMATOMA, AND MCS CONTINUED. FIVE DAYS LATER, THE PATIENT WAS SEVERELY DIAPHORETIC, POSSIBLE INFECTION FROM PRIOR ECMO GROIN SITE; HOWEVER, RESULTS WERE NEGATIVE. THE PATIENT HAD AN INCISION AND DRAINAGE WASHOUT TWO DAYS FOLLOWING, AND A COUPLE OF DAYS THEREAFTER, THE PATIENT EXPERIENCED LEFT SIDED WEAKNESS NOTED TO BE FROM PRESSING ON A NERVE DURING THE ECMO GROIN SITE WASHOUT PROCEDURE. THREE DAYS FOLLOWING THE LEFT-SIDED WEAKNESS, THE PATIENT HAD SLURRED SPEECH WITH MAINTAINED LEFT SIDED WEAKNESS. COMPUTED TOMOGRAPHY SHOWED THE PATIENT HAD A STROKE ON THE BRAIN STEM. THE PATIENT WAS INTUBATED TO MAINTAIN THE AIRWAY BUT WOULD LATER EXPIRE THIS SAME DAY. THE CAUSE OF DEATH WAS NOT PROVIDED; THEREFORE, UNKNOWN AND THE IMPELLA COULD NOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ACCESS SITE BLEED HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ACCESS SITE BLEED COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196699&pc=OZD | 1220648-2025-25663 | 2025-01-02 05:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | OZD | IMPELLA 5.5 | Patient-Device Incompatibility; Deformation Due to Compressive Stress | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT MULTIPLE COMPLICATIONS WERE ENCOUNTERED DURING IMPELLA 5.5 SUPPORT. DURING DELIVERY OF THE DEVICE, IT WAS NOTED THAT RESISTANCE WAS ENCOUNTERED AT THE INNOMINATE/SUBCLAVIAN JUNCTION RESULTING IN THE CATHETER KINKING. DESPITE THE NOTED KINK, THE DELIVERY WAS ULTIMATELY SUCCESSFUL AND SUPPORT WAS INITIATED. ROUGHLY THREE DAYS INTO SUPPORT, OOZING WAS NOTED AT THE IMPELLA ACCESS SITE. HEPARIN WAS STOPPED TO MANAGE THE CONDITION AND THE PATIENT REMAINS ON SUPPORT.Event Description: IT WAS REPORTED THAT THE PATIENT UNFORTUNATELY SUFFERED FROM SUSPECTED DEAD BOWEL AND WAS ON SUPPORT UNTIL CHILDREN COULD COME TO VISIT BEFORE EXPIRING. PATIENT EXPIRED ON SUPPORT 1/18 AT 1830. PER THE MEDICAL SAFETY OFFICER REVIEW THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME, DELAY IN NEEDED THERAPY. THE USER FOR SOME REASON UNKNOWN HAD DIFFICULTY PASSING THE DEVICE BEYOND THE INNOMINATE/SUBCLAVIAN INJUNCTION; THERE WAS BUCKLING. THERE IS NO ALLEGATION REGARDING MANUFACTURING/QUALITY ISSUE AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE FAILURE MODE: PRODUCT DAMAGE (CANNULA KINK) CHANGED TO FAILURE MODE: DELIVERY ISSUE SINCE THE CANNULA KINK OCCURRED DURING INSERTION. THE ROOT CAUSE OF DELIVERY ISSUE CANNOT BE DETERMINED SINCE THE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. THE ROOT CAUSE OF ACCESS SITE BLEEDING - MAJOR ISSUE CANNOT BE DETERMINED SINCE THE COMPLIANT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. B2 OUTCOME REVISED TO REFLECT DEATH WHICH INADVERTENTLY WAS NOT INCLUDED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25663. B5 REVISED TO REFLECT THE PATIENT OUTCOME AND MSO STATEMENT WHICH INADVERTENTLY WAS NOT INCLUDED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25663. D6B REVISED TO REFLECT THE EXPLANT DATE WHICH INADVERTENTLY WAS NOT INCLUDED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25663. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25663 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H1 REVISED TO DEATH WHICH INADVERTENTLY WAS NOT INCLUDED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25663. H6 REVISED TO REFLECT PATIENT OUTCOME OF DEATH WHICH INADVERTENTLY WAS NOT INCLUDED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25663. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196781&pc=OZD | 1220648-2025-25666 | 2025-01-02 05:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DEVELOPED LIMB ISCHEMIA. THE PATIENT WOULD EVENTUALLY EXPIRE. THE PATIENT WAS IN THE INTENSIVE CARE UNIT (ICU) FOR 12 HOURS AND HAD ISSUES WITH BLOOD PRESSURE AS THE VASOPRESSORS WERE INCREASING. THE DECISION WAS MADE TO PLACE IMPELLA CP DEVICE FOR ADDED SUPPORT. THE RIGHT FEMORAL ARTERY WAS ACCESSED USING MICRO-PUNCTURE AND ULTRASOUND. PERCLOSE X2 WERE PLACED. IT WAS NOTED 6500 UNITS HEPARIN WAS GIVEN. THE IMPELLA WAS SUCCESSFULLY IMPLANTED, DRESSING WAS PLACED, AND PATIENT WAS SENT TO ICU WITHOUT INCIDENT. HOWEVER, APPROXIMATELY 4 HOURS LATER, THE PATIENT'S LEG BECAME CYANOTIC, THE RIGHT VENTRICLE WAS STRUGGLING, AND THE IMPELLA COULD NOT RUN ABOVE P4 WITHOUT SUCTION. RIGHT VENTRICLE SUPPORT WAS DISCUSSED. HOWEVER, THE PATIENT WAS NOT EXPECTED TO SURVIVE THROUGHOUT THE NIGHT. THE PATIENT WOULD EVENTUALLY EXPIRE NEXT MORNING OVERNIGHT. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED LIMB ISCHEMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE LIMB ISCHEMIA COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196825&pc=OZD | 1220648-2025-25670 | 2025-01-06 05:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD A HEMORRHAGIC BLEED TO THE CEREBELLUM. THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR AN EXTERNAL VENTRICULAR DRAIN AND CRANIOTOMY. THE PATIENT EXPIRED AFTER 150.07 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25670. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196876&pc=OZD | 1220648-2025-25674 | 2025-01-07 05:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | OZD | IMPELLA CP | Increase in Pressure | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP DEVICE WAS IMPLANTED PREOPERATIVELY AND TWO DAYS INTO IMPELLA CP SUPPORT, THE SYSTEM TRIGGERED HIGH PURGE PRESSURE AND PURGE SYSTEM BLOCKED ALARMS. THE PURGE CASSETTE WAS REPLACED DURING TROUBLESHOOTING, BUT THE ISSUE PERSISTED. HEPARIN DOSE WAS INCREASED IN RESPONSE TO THE CONDITION AND ONGOING MONITORING HAD BEEN OUTLINED FOR SUPPORT. UPON FOLLOW UP, IT WAS VERIFIED THAT THE PURGE ISSUE RESOLVED WITH THE NOTED INTERVENTION. SUPPORT CONTINUED FOR TWO MORE DAYS UNTIL THE PATIENT SUFFERED A CARDIOPULMONARY ARREST WHILE WAITING FOR VENTRICULAR SEPTAL PERFORATION REPAIR AND PASSED AWAY. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. HOWEVER, THE IMPELLA IS UNLIKELY RELATED TO THE REPORTED DEATH. THE PATIENT WAS CONTINUED ON IMPELLA SUPPORT, AND THE DEVICE WAS NOT REMOVED BECAUSE OF THE PURGE PRESSURE ISSUE. THE PATIENT WAS SUBSEQUENTLY TAKEN TO THE OPERATING ROOM FOR THE CARDIAC SURGICAL PROCEDURE: VENTRICULAR SEPTAL PERFORATION REPAIR WHERE HE MET HIS DEMISE. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THERE WERE NO DATA LOGS RETURNED FOR THE COMPLAINT. THE CLINICAL DETAILS INDICATED THAT THERE WAS HIGH PURGE PRESSURE, TWO DAYS AFTER IMPLANT. THE CASSETTE WAS REPLACED BUT THAT DID NOT RESOLVE THE ISSUE. THE HEPARIN CONCENTRATION WAS INCREASED TO 50 (U/L) AND THE PURGE FLOW RATE IMPROVED FROM 1 MILLILITER (ML) TO 3ML AND THERE WAS NO PROBLEM WITH PUMP OPERATION. THE NEXT DAY, THE PATIENT WENT INTO CARDIOPULMONARY ARREST WHICH WAS NOT RELATED TO THE IMPELLA SO THE PUMP WAS EMERGENTLY EXPLANTED AS THE PATIENT EXPIRED. THE RETURNED PUMP WAS CUT AT THE CATHETER AND NO RED HANDLE WAS RETURNED TO CONFIRM THE SERIAL NUMBER OF THE PUMP. THERE WERE DEFORMATIONS ON THE BLADES OF THE IMPELLER BUT NONE WERE FRACTURED. REGARDING THE HIGH PURGE PRESSURE, THERE WAS NO KINKS SEEN IN THE RETURNED PORTION OF THE CATHETER. THERE WAS A SLIGHT FILM IN THE PURGE GAP BUT THAT WAS CLEARED AFTER THE PUMP WAS SOAKED. AS PART OF THE INVESTIGATION THE CATHETER WAS THEN CUT NEAR THE MOTOR HOUSING AND NO KINK IN THE PURGE LINE WAS SEEN OR SIGNS OF BLOOD INGRESS. A SYRINGE WITH FLUID WAS USED IN THE PURGE LINE AND FLUID EXPELLED VERY EASILY FROM THE PURGE GAP. DUE TO LACK OF SUFFICIENT PRODUCT RETURNED AS WELL AS DATA LOGS, THE ROOT CAUSE OF THE HIGH PURGE PRESSURE COULD NOT BE DETERMINED. D.9 DATE DEVICE RETURNED/INFORMATION WAS ADDED. E.1 ADDRESS LINE 2 INFORMATION WAS ADDED AS IT WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25674. E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25674 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4641 WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25674. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196959&pc=BZD | 2518422-2025-005973 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, METASTATIC LUNG CANCER AND BRAIN AND BONE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197451&pc=DXF | 2124215-2025-01600 | 2024-12-30 05:00 | Death | BAYLIS MEDICALE CIE INC | 2025-01-21 05:00 | DXF | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A WATCHMAN FLX PRO PROCEDURE, A VERSACROSS ACCESS SOLUTION KIT WAS SELECTED FOR USE. A TRACE EFFUSION WAS NOTED PRIOR TO PROCEDURE. THE PHYSICIAN OBTAINED RIGHT FEMORAL VEIN ACCESS AND AN 18GUAGE SHEATH WAS PLACED. THE VERSACROSS SHEATH WAS INSERTED. THE PHYSICIAN DROPPED BELOW THE SEPTUM ON FIRST ATTEMPT, SO IT WAS RE-ATTEMPTED TO RE-WIRE THE SUPERIOR VENA CAVA (SVC) TO COME BACK DOWN TO THE SEPTUM. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND THE VERSACROSS SHEATH WAS EXCHANGED FOR THE WATCHMAN TRUSTEER SHEATH OVER THE WIRE. A PIGTAIL CATHETER WENT IN OVER THE WIRE. AT THIS POINT, AN EFFUSION WAS NOTED. AUTO TRANSFUSION BEGAN WITH THAT PIGTAIL CATHETER WHILE A PERICARDIOCENTESIS WAS PERFORMED. AFTER THAT THERE WERE TWO PIGTAIL CATHETERS PLACED IN THE PERICARDIAL SPACE. 12 UNITS OF PACKED RED BLOOD CELLS (PRBCS), 8 UNITS OF FRESH FROZEN PLASMA (FFP) AND 1 UNIT OF PLATELETS WERE GIVEN AND THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR). THE PATIENT ARRESTED BEFORE GETTING TRANSFERRED THE OR TABLE, PATIENT CHEST WAS OPENED BUT WERE UNABLE TO RESUSCITATE. IT WAS DETERMINED THE CAUSE OF DEATH WAS DUE TO PERFORATION AT THE BACK WALL OF THE HEART LEADING TO PERICARDIAL EFFUSION, CARDIAC TAMPONADE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. IT WAS FURTHER CONFIRMED VIA GFE DATED 10JAN2025, THERE WAS NO IMAGING/VISUALIZATION ISSUES NOTED. THE TSP WAS WAS DIFFICULT BECAUSE OF THE REACHABILITY. THE PHYSICIAN BROUGHT THE SHEATH OUT OVER THE WIRE AND PUT MORE OF A BEND ON THE VERSACROSS AND WENT BACK IN TO THE SVC AND BACK DOWN. THE TSP WAS VERY INFERIOR AND POSTERIOR. THE PERFORATION COULD HAVE PARTIALLY CAUSED BY THE WIRE HOWEVER THE PHYSICIAN PUT THE TRUSTEER OVER THE WIRE AND INTO THE PERICARDIAL SPACE. THERE WAS NO MALFUNCTION WITH THE VERSACROSS SYSTEM. ACTIVATED CLOTTING TIME (ACT) GREATER THAN 400 WITH AN INR OF 2.5 ON CASE DAY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197538&pc=DSQ | 2916596-2025-00004 | 2024-12-25 05:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Hemorrhage/Bleeding; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER SURGERY, THE PATIENT HAD SHOWN TENDENCIES OF RIGHT HEART FAILURE AND WAS MONITORED WITH LOW FLOW ALARMS WITH SOFTWARE VERSION 2.0. THE PATIENT'S CONDITION CONTINUED TO WORSEN, AND LOW FLOW PERSISTED. DUE TO WORSENING RIGHT HEART FAILURE, SURGERY TO IMPLANT A RIGHT VENTRICULAR ASSIST DEVICE WAS PERFORMED ON (B)(6) 2024. DURING SURGERY BLEEDING BECAME UNCONTROLLABLE, LEADING TO CIRCULATORY COLLAPSE AND DEATH. THE PATIENT WAS EXPERIENCING LOW FLOW ALARMS AT THE TIME OF THE EVENT. THE PATIENT WAS EXPLANTED DUE TO THE POST OPERATION SURGICAL INCIDENT. THE CAUSE OF THE BLEEDING WAS UNDER INVESTIGATION. LOG FILES WERE SENT FOR ANALYSIS TO CONFIRM HOW THE PUMP HAD BEEN OPERATING BETWEEN 14:00 AND 18:00 ON (B)(6) 2024. THE DATA RECORDED ON THE PERIODIC LOG FILE INDICATED THAT THERE WAS AN ACTIVE LVAD_OFF ALARM FLAG ASSOCIATED WITH AN (F69) INTENTIONAL PUMP STOP CONTROLLER FAULT FLAG ON (B)(6) 2024 16:45:42. THE PATIENT PASSED AWAY ON (B)(6) 2024. Manufacturer Narrative: A1-A6: PATIENT INFORMATION IS UNKNOWN D4: PRODUCT INFORMATION, UPDATED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS EVENT WAS DETERMINED TO BE SUPPLEMENTAL AND THE INVESTIGATION FINDINGS WILL BE SUBMITTED UNDER MFR# 2916596-2025-00036. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197654&pc=LWS | 3008973940-2025-00580 | 2025-01-09 05:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-01-21 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Under-Sensing | Tachycardia; Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. POSSIBLE UNDERSENSING OF VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF) WAS NOTED. THE REMOTE TRANSMISSION SHOWED VT/VF THAT WAS NOT FALLING WITHIN THE DETECTION ZONES. THE DEVICE WAS WORKING WELL WITHIN PROGRAMMED PARAMETERS. THE CAUSE OF DEATH WAS INDICATED AS SUSPECTED CARDIAC EVENT.Event Description: IT WAS FURTHER REPORTED THAT A LEAD INTEGRITY ALERT (LIA) TRIGGERED FOR TWO OR MORE HIGH RATE-NON SUSTAINED EPISODES AND HIGH SENSING INTEGRITY COUNTER (SIC). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5.DESC EVT PROBLEM, DEVICE CODES (FDD/ANNEX A) MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197778&pc=BZD | 2518422-2025-005943 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21198154&pc=NKM | 2135147-2025-00356 | 2021-10-07 04:00 | Death | ABBOTT MEDICAL | 2025-01-21 05:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P100009 |  | 1 | Event Description: IT WAS REPORTED A XTR MITRACLIP WAS IMPLANTED ON (B)(6) 2021 TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. IT IS NOTED THAT THE PATIENT ALSO UNDERWENT AN IMPLANT OF A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) WITH A NON-ABBOTT VALVE THE SAME DAY. INFORMATION RECEIVED ON (B)(6) 2024, THAT THE PATIENT HAD PASSED AWAY. IT WAS REPORTED THAT THE POSSIBLE CAUSE DEATH WAS DEVICE INFECTION RELATED TO THE NON-ABBOTT AORTIC VALVE AND THE MITRACLIP. THE EXACT TYPE AND LOCATION OF THE INFECTION WAS UNKNOWN. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED DEATH AND INFECTION. DEATH AND INFECTION ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21198846&pc=BZD | 2518422-2025-100972 | 2024-11-13 05:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21198902&pc=NEY | 3004962788-2025-00015 | 2024-06-30 04:00 | Death | COVIDIEN LP - SUPERDIMENSION INC | 2025-01-21 05:00 | NEY | UNKNOWN EMPRINT GENERATOR | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY FROM 2007-2022 EVALUATED OUTCOMES OF MICROWAVE ABLATION (MWA) TO 397 LIVER TUMORS RESULTING FROM NEUROENDOCRINE TUMOR METASTASES IN 94 PATIENTS. MWA WAS COMPLETED USING EMPRINT, EVIDENT, OR COMPETITOR PRODUCTS. PROCEDURES WERE PERFORMED VIA LAPAROSCOPIC, OPEN, HAND-ASSISTED LAPAROSCOPIC, AND ROBOTIC TECHNIQUES. CONCOMITANT PROCEDURES WERE PERFORMED INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, WHIPPLE PROCEDURE, DISTAL PANCREATECTOMY, AND LIVER RESECTION IN 62% OF PATIENTS. ONE PATIENT DEATH WAS REPORTED IN A PATIENT WITH CIRRHOSIS WHO UNDERWENT A MWA WITHOUT CONCOMITANT PROCEDURE AND DIED ON POSTOPERATIVE DAY THREE. NO OTHER DETAILS WERE PROVIDED REGARDING THIS MORTALITY. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNK - EVIDENT GEN, UNKNOWN EVIDENT GENERATOR, (SN: UNKNOWN) UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR, (SN: UNKNOWN) UNK - EVIDENT GEN, UNKNOWN EVIDENT GENERATOR, (SN: UNKNOWN) SURGICAL MICROWAVE ABLATION OF 397 NEUROENDOCRINE LIVER METASTASES: A RETROSPECTIVE COHORT ANALYSIS OF 16 YEARS OF EXPERIENCE ALEXANDRA WELLS, VINCENT BUTANO, MICHAEL PHILLIPS, JOSHUA DAVIS, ERIN BAKER, JOHN MARTINIE, DAVID IANNITTI ACCEPTED: 30 JUNE 2024 HTTPS://DOI.ORG/10.1007/S00464-024-11021-4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21198920&pc=NEY | 3004962788-2025-00016 | 2024-06-30 04:00 | Death | COVIDIEN LP - SUPERDIMENSION INC | 2025-01-21 05:00 | NEY | UNKNOWN EVIDENT GENERATOR | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY FROM 2007-2022 EVALUATED OUTCOMES OF MICROWAVE ABLATION (MWA) TO 397 LIVER TUMORS RESULTING FROM NEUROENDOCRINE TUMOR METASTASES IN 94 PATIENTS. MWA WAS COMPLETED USING EMPRINT, EVIDENT, OR COMPETITOR PRODUCTS. PROCEDURES WERE PERFORMED VIA LAPAROSCOPIC, OPEN, HAND-ASSISTED LAPAROSCOPIC, AND ROBOTIC TECHNIQUES. CONCOMITANT PROCEDURES WERE PERFORMED INCLUDING HEPATECTOMY, CHOLECYSTECTOMY, WHIPPLE PROCEDURE, DISTAL PANCREATECTOMY, AND LIVER RESECTION IN 62% OF PATIENTS. ONE PATIENT DEATH WAS REPORTED IN A PATIENT WITH CIRRHOSIS WHO UNDERWENT A MWA WITHOUT CONCOMITANT PROCEDURE AND DIED ON POSTOPERATIVE DAY THREE. NO OTHER DETAILS WERE PROVIDED REGARDING THIS MORTALITY. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR, (SN: UNKNOWN) UNK - EMPRINT GEN, UNKNOWN EMPRINT GENERATOR, (SN: UNKNOWN) UNK - EVIDENT GEN, UNKNOWN EVIDENT GENERATOR, (SN: UNKNOWN) SURGICAL MICROWAVE ABLATION OF 397 NEUROENDOCRINE LIVER METASTASES: A RETROSPECTIVE COHORT ANALYSIS OF 16 YEARS OF EXPERIENCE ALEXANDRA WELLS, VINCENT BUTANO, MICHAEL PHILLIPS, JOSHUA DAVIS, ERIN BAKER, JOHN MARTINIE, DAVID IANNITTI ACCEPTED: 30 JUNE 2024 HTTPS://DOI.ORG/10.1007/S00464-024-11021-4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199054&pc=BZD | 2518422-2025-006023 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded |  | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING THYROID AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199152&pc=BZD | 2518422-2025-006029 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION CPAP | Degraded | Asthma; Nausea; Vomiting; Cancer; Unspecified Respiratory Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199635&pc=PYX | 1220648-2025-25679 | 2025-01-04 05:00 | Death | ABIOMED, INC. | 2025-01-21 05:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Hemolysis; Thrombosis/Thrombus | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPLANT, AN IMPELLA RP FLEX DEVICE EXPERIENCED A SIGNIFICANT DECREASE IN PUMP FLOW. IT WAS NOTED THAT HEPARIN WAS NOT BEING SUPPLIED DUE TO BLEEDING AT THE EXTRACORPOREAL MEMBRANE OXYGENATION SITE AND IT WAS BELIEVED A THROMBUS HAD ENTERED THE PUMP. HEPARIN WAS RESTARTED TO ADDRESS THE ISSUE AND SUPPORT CONTINUED. A COUPLE DAYS LATER, THE PATIENT DEVELOPED HEMOLYSIS. CARE WAS WITHDRAWN AND THE PATIENT PASSED AWAY. IT IS UNKNOWN IF THE COMPLICATIONS CONTRIBUTED TO THE PATIENT'S PASSING. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS, LOW PUMP FLOW, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW FAILURE MODE WAS CHANGED TO PLACEMENT SIGNAL ISSUE SINCE THE CHARACTERISTICS DETERMINE A FALSE LOW FLOW WHICH IS A CHARACTERISTIC OF SENSOR DRIFTING. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. DATA LOGS SHOW WHILE THE PUMP FLOW WAS DRIFTING DOWN, THE PLACEMENT SIGNAL WAS OBSERVED TO BE DRIFTING UPWARDS. THE PLACEMENT SIGNAL (PS) DRIFTED FROM 100MMHG TO 300MMHG IN THE SPAN OF 90 MINUTES AND THEN EVENTUALLY STARTS TO STABILIZE. MOTOR CURRENT, CVP VALUES AND OPTICAL SIGNAL WERE STABLE AT THE TIME. NO PSNR ALARMS WERE PRESENT AT THE TIME. THESE CHARACTERISTICS INDICATE A DRIFTING DP SENSOR LEADING TO FALSE FLOW VALUES. THEREFORE, THE ROOT CAUSE OF THE PLACEMENT SIGNAL ISSUES WAS DETERMINED TO BE MOST LIKELY DP SENSOR DRIFT. DATA LOGS WERE INVESTIGATED. AT THE TIME OF THE REPORTED FAILURE, THERE WERE NO MOTOR CURRENT SPIKES OR SUCTION ALARMS OBSERVED. PLACEMENT SIGNAL WAS TRENDING DOWN TO NORMAL FROM THE PREVIOUS RISE SINCE LAST 34 HOURS. PUMP FLOW WAS SEEN GETTING BETTER FROM 2L/MIN TO 2.5L/MIN AT P9. THE EXPECTED FLOW AT P9 IS (3.3L/MIN TO 3.7L/MIN). EVENTUALLY THE PS STABILIZES AND ACTS SIMILAR TO CVP AND MOTOR CURRENT. THE PATIENT HAD RECEIVED FEW UNITS OF BLOOD WHICH HELPED THE FLOW TO GET BETTER. THEREFORE, THE ROOT CAUSE OF THE HEMOLYSIS ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED". Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199721&pc=DSQ | 2916596-2025-00482 | 2024-12-28 05:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AND IT WAS SAID THE DEVICE WAS FUNCTIONING AS INTENDED. NO EQUIPMENT WILL BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199771&pc=DSQ | 3007042319-2025-00083 | 2024-11-04 05:00 | Death | HEARTWARE, INC. | 2025-01-21 05:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Electrical /Electronic Property Problem; Pumping Stopped; Battery Problem; Power Problem; Pumping Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF ISCHEMIC CARDIOMYOPATHY EXPERIENCED AN UNEXPLAINED DEATH, POSSIBLY DUE TO A LOSS OF POWER TO THE VENTRICULAR ASSIST DEVICE. THE PATIENT HAD SUFFERED A STROKE THE PREVIOUS YEAR, WHICH AFFECTED MOBILITY. THE DEVICE'S LOGFILES RECORDED MULTIPLE CRITICAL BATTERY ALARMS AND CONTROLLER FAULT ALARMS, ALONG WITH LOW FLOW ALARMS AND A VENTRICULAR ASSIST DEVICE STOP ALARM. IT WAS ALLEGED THAT THE PATIENT MAY HAVE SLEPT ON THE BATTERIES WITHOUT THE MAIN ADAPTOR PLUGGED IN, AND POSSIBLY ATTEMPTED TO GET OUT OF BED OR TO A CHAIR TO REACH THE POWER SUPPLY AND THE CONTROLLER EXHIBITED MULTIPLE LOSSES OF POWER. IT WAS ALSO REPORTED THAT THE CONTROLLER EXHIBITED UNEXPECTED LOSSES OF POWER IN SEVERAL MONTHS AGO. THE DEATH WAS REPORTED TO THE CORONER FOR INVESTIGATION. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿CONTROLLER 2.0 D4: MODEL #: 1420/ CATALOG #: 1420/ EXPIRATION DATE: 31-JUL-2023 / SERIAL OR LOT#: (B)(6) D9: NO H3: NO H4: MFG DATE: 11-JUL-2022 H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ BATTERY D4: MODEL #: 1650/ CATALOG #: 1650/ EXPIRATION DATE: 31-MAR-2023 / SERIAL OR LOT#: (B)(6) D9: NO H3: NO H4: MFG DATE: 10-MAR-2022 H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿BATTERY D4: MODEL #: 1650 / CATALOG #: 1650/ EXPIRATION DATE: 31-MAR-2024 / SERIAL OR LOT#: (B)(6) D9: NO H3: NO H4: MFG DATE: 10-MAR-2023 H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6), ONE (1) CONTROLLER (B)(6), AND TWO (2) BATTERIES (B)(6) WERE NOT RETURNED FOR EVALUATION AS THE VENTRICULAR ASSIST DEVICE (VAD)REMAINS IMPLANTED, AND THE CONTROLLER & (2) BATTERIES DEVICES LOCATION IS UNKNOWN. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED THREE (3) CONTROLLER POWER UP EVENTS WITH ASSOCIATED PUMP START EVENTS WERE LOGGED ON (B)(6) 2024 AT 12:59:03, (B)(6) 2024 AT 05:53:43, AND (B)(6) 2025 AT 22:20:22. THE CONTROLLER WAS WITHOUT POWER FOR TWENTY-SIX (26) SECONDS, EIGHTEEN (18) SECONDS, AND ELEVEN (11) SECONDS, RESPECTIVELY. THE CONTROLLER CAN ONLY STORE A MAXIMUM OF 30 DAYS OF DATA. AFTER REACHING THE LIMIT, THE CONTROLLER INITIATES A FIRST-IN FIRST-OUT METHOD WHEREBY THE OLDEST DATA POINT IS DELETED TO ALLOW THE NEWEST DATA POINT TO BE RECORDED. AS A RESULT, THE DATA LOG FILE COVERING THE FIRST LOSS OF POWER WAS NOT AVAILABLE FOR ANALYSIS. NO ANOMALIES WERE RECORDED LEADING UP TO THE LOSSES OF POWER ON (B)(6) 2024 AND (B)(6) 2025. REVIEW OF THE ALARM LOG FILE REVEALED FIFTY-THREE (53) CRITICAL BATTERY ALARMS INVOLVING BATTERIES (B)(6) WERE LOGGED ON (B)(6) 2025. THE CRITICAL BATTERY ALARMS WERE THE RESULT OF THE BATTERIES DEPLETING BELOW 10% RELATIVE STATE OF CHARGE (RSOC). ANALYSIS OF THE LOG FILES REVEALED THAT, AT THE TIME OF THE FIRST CRITICAL BATTERY ALARM, (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 18% RELATIVE STATE OF CHARGE (RSOC) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 9% RSOC. BOTH BATTERIES WERE ALLOWED TO CONTINUE DEPLETING, THUS TRIGGERING CONTROLLER FAULT ALARMS, WHICH WILL OCCUR IF A BATTERY IS APPROACHING 0% RSOC. TWENTY-FOUR (24) CONTROLLER FAULT ALARMS WERE LOGGED ON (B)(6) 2025. A REVIEW OF THE EVENT LOG FILE REVEALED TWO (2) SUBSEQUENT CONTROLLER POWER-UP EVENTS ON (B)(6) 2025 AT 09:20:21 AND 09:22:29. THE CONTROLLER WAS WITHOUT POWER FOR OVER AN HOUR DURING THE FIRST CONTROLLER LOSS OF POWER EVENT, AND THIRTY-SEVEN (37) SECONDS DURING THE SECOND CONTROLLER LOSS OF POWER EVENT. THE DATA POINT RECORDED BEFORE THE POWER-UP EVENTS REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2); THE BATTERY WAS ALLOWED TO DEPLETE COMPLETELY, RESULTING IN A LOSS OF POWER TO THE CONTROLLER. THE BATTERY LIKELY RECOVERED ENOUGH CHARGE TO START THE CONTROLLER AGAIN, BUT QUICKLY DEPLETED, CAUSING THE ADDITIONAL CONTROLLER POWER-UP EVENT. ADDITIONALLY, LOG FILES REVEALED TWO (2) LOW FLOW ALARMS WERE LOGGED ON (B)(6) 2025 AT 09:20:30 AND AT 09:22:38. AS A RESULT, THE REPORTED CONTROLLER LOSS OF POWER, CRITICAL BATTERY ALARMS, CONTROLLER FAULT ALARMS, AND LOW FLOW ALARM EVENTS WERE CONFIRMED. THE VAD STOP ALARM COULD NOT BE CONFIRMED; HOWEVER, IT IS LIKELY THAT THE LOSSES OF POWER WERE PERCEIVED AS THE REPORTED VAD STOP. A POSSIBLE ROOT CAUSE OF THE LOSS OF POWER EVENTS THAT OCCURRED ON (B)(6) 2024, AND (B)(6) 2025 CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. BASED ON THE AVAILABLE INFORMATION, THE MOST LIKELY ROOT CAUSE OF THE CRITICAL BATTERY ALARMS, CONTROLLER FAULT ALARMS, AND CONTROLLER LOSS OF POWER EVENTS THAT OCCURRED ON (B)(6) 2025 CAN BE ATTRIBUTED TO THE BATTERIES BEING ALLOWED TO DEPLETE COMPLETELY. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. OF NOTE, THE AVAILABLE INFORMATION INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, AND/OR POOR VAD FILLING. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF NEUROLOGICAL DYSFUNCTION. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - 1420-CONTROLLER 2.0 D4: SERIAL#: (B)(6) H3: YES D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - BATTERY D4: SERIAL#: (B)(6) H3: YES D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - BATTERY D4: SERIAL#: (B)(6) H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199983&pc=NPT | 2025587-2025-00498 | 2024-05-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-21 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT SUDDENLY DETERIORATED AFTER RETURNING HOME FROM A REGULAR HOSPITAL VISIT. THE PATIENT WAS URGENTLY TRANSPORTED BACK TO THE HOSPITAL AND WAS NOTED TO BE IN ASYSTOLE UPON ARRIVAL. SUBSEQUENTLY, THE PATIENT PASSED AWAY. A CLEAR CAUSE OF DEATH WAS UNABLE TO BE DETERMINED. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, IT IS UNKNOWN IF THE PATIENT'S CAUSE OF DEATH COULD BE CONTRIBUTED TO THE PATIENT'S UNDERLYING MEDICAL HISTORY OR NOT. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200038&pc=DSQ | 2916596-2025-00169 | 2024-12-31 05:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR (RV) FAILURE. IT WAS NOTED THAT THE PATIENT HAD RV DYSFUNCTION BEFORE LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. THE PATIENT RECEIVED MULTIPLE ROUNDS OF ALS, INITIATION OF HIGH DOSE INOTROPES, AND THERE WERE CONSIDERATIONS OF VENOVENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE REPORTEDLY OPERATED AS EXPECTED. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS FURTHER COMMUNICATED THAT THE PUMP WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200053&pc=MCX | 3004742232-2025-00018 | 2024-12-27 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-01-21 05:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Vasoconstriction; Obstruction/Occlusion | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A HEAVILY CALCIFIED, MILDLY TORTUOUS LEFT ANTERIOR DESCENDING (LAD) ARTERY. AFTER THE ORBITAL ATHERECTOMY TREATMENT, ST ELEVATION AND SLOW FLOW WERE OBSERVED. THE PATIENT BEGAN TO GASP FOR AIR AND THEIR BLOOD PRESSURE SUDDENLY DROPPED AFTER STARTING INOTROPES. THE PATIENT WAS INTUBATED AND HAD NO CAROTID PULSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR 35 MINUTES. AN INTRA-AORTIC BALLOON PUMP (IABP) AND TEMPORARY PACEMAKER IMPLANT WERE PLACED. THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A HEAVILY CALCIFIED, MILDLY TORTUOUS LEFT ANTERIOR DESCENDING (LAD) ARTERY. AFTER THE ORBITAL ATHERECTOMY TREATMENT, ST ELEVATION AND SLOW FLOW WERE OBSERVED. THE PATIENT BEGAN TO GASP FOR AIR AND THEIR BLOOD PRESSURE SUDDENLY DROPPED AFTER STARTING INOTROPES. THE PATIENT WAS INTUBATED AND HAD NO CAROTID PULSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR 35 MINUTES. AN INTRA-AORTIC BALLOON PUMP (IABP) AND TEMPORARY PACEMAKER IMPLANT WERE PLACED. THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE SLOW FLOW WAS PERSISTENT, AND WAS CAUSED BY A VESSEL SPASM. THE PATIENT EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT HARMS INCLUDING ARRHYTHMIA, OBSTRUCTION, CARDIAC ARREST, LOW BLOOD PRESSURE, VASOCONSTRICTION, DEATH, AND RELATED HOSPITALIZATION, AND MEDICAL INTERVENTION APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. DUE TO THE VERY LIMITED INFORMATION PROVIDED, THE MAIN CAUSE OF DEATH IS MOST LIKELY DUE TO PROCEDURAL CIRCUMSTANCES AND POSSIBLY THE PATIENT¿S DISEASE STATE ALTHOUGH THE USE OF THE ORBITAL ATHERECTOMY PROBABLY INDIRECTLY CONTRIBUTED TO THE OUTCOME OF DEATH. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. B2: CORRECTED TO REFLECT THE PATIENT OUTCOME. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200312&pc=GAM | 9612501-2025-00186 | 2024-09-22 04:00 | Death | DAVIS & GECK CARIBE LTD | 2025-01-21 05:00 | GAM | UNKNOWN - MAXON | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A RETROSPECTIVE ANALYSIS WAS COMPLETED DETAILING 300 CASES OF LAPAROSCOPIC COMMON BILE DUCT EXPLORATION (LCBDE) DUE TO CHOLEDOCHOLITHIASIS. A TOTAL OF 300 PATIENTS UNDERWENT LCBDE. THE FIRST 100 CASES WERE PERFORMED IN THE 1990S; THE REMAINING 200 CASES WERE PERFORMED BETWEEN 2017 AND 2021. BOTH DATABASES INCLUDED CONSECUTIVE UNSELECTED PATIENTS WITH A PROVEN OR SUSPECTED DIAGNOSIS OF CHOLEDOCHOLITHIASIS. IF A CHOLEDOCHOTOMY WAS NECESSARY, A PRIMARY CLOSURE WITH A RUNNING 5-0 MAXON SUTURE WAS PERFORMED. COMPLICATION IN THE ENTIRE COHORT INCLUDE ONE DEATH. NO DETAILS REGARDING INTERVENTIONS WERE DESCRIBED. NON-DEVICE RELATED COMPLICATIONS INCLUDE CONVERSION TO OPEN AND PANCREATITIS. AN ENDOCLOSE WAS ALSO USED BUT WAS NOT RELATED TO COMPLICATIONS. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: UNKNOWN - MAXON, UNKNOWN - MAXON, (LOT #UNKNOWN) REFERENCE: STRUCTURED LEARNING AND MENTORING: SHORTENING THE LEARNING CURVE IN LAPAROSCOPIC COMMON BILE DUCT EXPLORATION DAVID MARTINEZ-CECILIA, 2024, SURGICAL ENDOSCOPY HTTPS://DOI.ORG/10.1007/S00464-024-11304-W. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200510&pc=MKJ | 3030677-2025-000224 | 2025-01-14 05:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-01-21 05:00 | MKJ | HEARTSTART XL+ DEFIBRILLATOR/MONITOR | Defibrillation/Stimulation Problem | Burn(s) | K110825 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE HEARTSTART XL+ DEFIBRILLATOR/MONITOR INDICATING THE DEVICE BURNED THE PATIENT'S SKIN DURING USE. THE PATIENT HAD NO HEARTBEAT AND WAS UNCONSCIOUS DURING USE. THE DEPARTMENT TEACHER REPORTED THAT THE PATIENT¿S SKIN WAS BURNT AND BLACK. THE PATIENT DIED. ADDITIONAL DETAILS HAVE BEEN REQUESTED.Event Description: PHILIPS RECEIVED A COMPLAINT ON THE XL+ DEVICE INDICATING THAT THE DEVICE BURNED THE PATIENT'S SKIN DURING USE. THE PATIENT HAD NO HEARTBEAT AND WAS UNCONSCIOUS IN THE PROCESS OF THE BURNING OF THE PATIENTS SKIN. THE PATIENT DIED. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. Manufacturer Narrative: THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS REMOTE SERVICE ENGINEER (RSE) AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. MULTIPLE ATTEMPTS WERE MADE TO REQUEST INFORMATION REGARDING RESOLUTION OF THE REPORTED PROBLEM. NO RESPONSE WAS RECEIVED, SO NO INFORMATION IS AVAILABLE. THIS WILL BE DOCUMENTED AS A MALFUNCTION, THE CAUSE OF WHICH WAS NOT DETERMINED. THE DEVICE REMAINS AT THE CUSTOMER SITE AND NO FURTHER EVALUATION IS WARRANTED AT THIS TIME. BASED ON THE INFORMATION AVAILABLE THERE IS INSUFFICIENT INFORMATION TO CONFIRM THE REPORTED PROBLEM. BASED ON THE INFORMATION AVAILABLE AND RESULTS OF ADDITIONAL ANALYSIS, NO FURTHER ACTION IS NECESSARY AT THIS TIME. WE ARE UNABLE TO CONFIRM THE FINAL DISPOSITION OF THE DEVICE BECAUSE THE CUSTOMER DID NOT RESPOND TO REQUESTS FOR ADDITIONAL INFORMATION. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200512&pc=MGQ | MW5164997 |  | Death | SOLEVA PHARMA LLC. | 2025-01-21 05:00 | MGQ | MUGARD ORAL RINSE 240ML/BTL |  | Insufficient Information |  |  | 1 | Event Description: EVENT: THE PATIENT WAS MARKED DECEASED IN WALGREENS RETAIL SYSTEM WHICH FLOWED TO WALGREENS SPECIALTY SYSTEM AND WE ARE REPORTING PER CONTRACTUAL REQUIREMENTS AND STANDARD OPERATING PROCEDURE WITHIN ONE BUSINESS DAY; HOWEVER, NO CAUSE OF DEATH KNOWN, DATE OF DEATH UNKNOWN AND NO FURTHER INFORMATION AVAILABLE. FREQ: SWISH AND EXPEL OR SWALLOW 5-10 ML 4-6 TIMES DAILY AS DIRECTED FOR MUCOSITIS. MALIGNANT NEOPLASM OF LIVER, PRIMARY, UNSPECIFIED AS TO TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200514&pc=CAW | MW5164998 | 2024-08-09 04:00 | Death | INOGEN, INC. | 2025-01-21 05:00 | CAW | INOGEN G5 PORTABLE OXYGEN CONCENTRATOR |  | Distress; Respiratory Insufficiency |  |  | 1 | Event Description: I AM WRITING TO FORMALLY REPORT A DEFECTIVE MEDICAL DEVICE, THE LNOGEN GS PORTABLE OXYGEN CONCENTRATOR (POC), WHICH WAS PURCHASED FROM NATIONAL MEDICAL EQUIPMENT, INC. AND CAUSED SIGNIFICANT HARM TO MY HUSBAND'S HEALTH DURING ITS USE. THIS MATTER INVOLVES BOTH THE MALFUNCTION OF THE DEVICE AND THE COMPANY'S FAILURE TO ADEQUATELY ADDRESS THE SITUATION. BACKGROUND: ON (B)(6) 2024, I PURCHASED AN INOGEN GS POC FROM (B)(4) FOR (B)(6). THE DEVICE WAS REQUIRED TO FACILITATE THE SAFE AIR TRANSPORT OF MY HUSBAND, WHO HAD SUFFERED A CATASTROPHIC ACCIDENT AND WAS BEING TRANSFERRED FROM (B)(6) TO (B)(6) FOR CONTINUED MEDICAL CARE. THE DEVICE WAS REPRESENTED AS FAA-APPROVED AND CAPABLE OF FUNCTIONING FOR 13 HOURS, WHICH WAS CRITICAL FOR OUR 3-HOUR FLIGHT AND ANY POTENTIAL DELAYS. DEVICE FAILURE: SHORTLY AFTER TAKEOFF, THE INOGEN GS POC FAILED. THIS MALFUNCTION LEFT MY HUSBAND WITHOUT SUFFICIENT OXYGEN FOR OVER SEVEN HOURS DURING THE FLIGHT, RESULTING IN DELIRIUM AND A SIGNIFICANT WORSENING OF HIS ALREADY CRITICAL CONDITION. THE LACK OF OXYGEN DURING THIS PERIOD EXACERBATED HIS HEALTH PROBLEMS AND CAUSED IMMENSE DISTRESS. TRAGICALLY, HIS CONDITION DETERIORATED RAPIDLY FOLLOWING THIS INCIDENT, AND HE PASSED AWAY ON (B)(6) 2024. ISSUES WITH (B)(4): MISREPRESENTATION OF DEVICE CAPABILITIES: THE COMPANY ASSURED ME THE POC WOULD OPERATE FOR 13 HOURS, BUT IT FAILED WITHIN APPROXIMATELY 2.5 HOURS. FAILURE TO FILE (B)(6) CLAIM: DESPITE (B)(6) CONFIRMING THE DEVICE WOULD HAVE BEEN COVERED, THE COMPANY REQUIRED UPFRONT PAYMENT AND DID NOT PROCESS THE CLAIM. DEFECTIVE EQUIPMENT: THE DEVICE'S MALFUNCTION IN A LIFE-CRITICAL SITUATION RAISES QUESTIONS ABOUT ITS QUALITY AND SAFETY. INADEQUATE RESOLUTION: DESPITE RETURNING THE DEVICE PER THEIR INSTRUCTIONS, THE COMPANY INITIALLY REFUSED DELIVERY AND ONLY ACCEPTED IT AFTER A SECOND ATTEMPT. THEY OFFERED A NOMINAL REFUND OF (B)(6), FAR BELOW THE PURCHASE PRICE, AND FAILED TO PROVIDE ANY JUSTIFICATION OR ACCOUNTABILITY FOR THE DEFECTIVE PRODUCT. IMPACT: THE FAILURE OF THIS MEDICAL DEVICE HAD CATASTROPHIC CONSEQUENCES FOR MY HUSBAND'S HEALTH AND CONTRIBUTED TO HIS UNTIMELY PASSING. THIS SITUATION ALSO CAUSED EMOTIONAL AND FINANCIAL DISTRESS TO OUR FAMILY. ADDITIONALLY, THE COMPANY'S REFUSAL TO HONOR ITS COMMITMENTS HIGHLIGHTS BROADER CONCERNS ABOUT THEIR ETHICAL PRACTICES. REQUEST FOR FDA ACTION: I REQUEST THAT THE FDA INVESTIGATE THE FOLLOWING: DEVICE RELIABILITY: CONDUCT A THOROUGH REVIEW OF THE INOGEN GS POC TO DETERMINE WHETHER IT MEETS SAFETY AND PERFORMANCE STANDARDS. COMPANY PRACTICES: INVESTIGATE NATIONAL MEDICAL EQUIPMENT, INC. FOR POTENTIAL VIOLATIONS OF FDA REGULATIONS, INCLUDING THE DISTRIBUTION OF DEFECTIVE DEVICES AND MISREPRESENTATION OF PRODUCT CAPABILITIES. RECALL EVALUATION: ASSESS WHETHER THE INOGEN GS POC SHOULD BE SUBJECT TO A RECALL OR ADDITIONAL SCRUTINY TO PREVENT HARM TO OTHER PATIENTS. PLEASE FIND ATTACHED SUPPORTING DOCUMENTATION, INCLUDING: PROOF OF PURCHASE/SHIPPING AND RETURN DETAILS I APPRECIATE YOUR PROMPT ATTENTION TO THIS MATTER. ENSURING THE SAFETY AND RELIABILITY OF MEDICAL DEVICES IS VITAL TO PROTECTING PATIENTS AND THEIR FAMILIES FROM SIMILAR TRAGEDIES. THANK YOU FOR YOUR TIME AND CONSIDERATION. PLEASE DO NOT HESITATE TO CONTACT ME IF YOU REQUIRE ADDITIONAL INFORMATION OR DOCUMENTATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200600&pc=DSQ | 2916596-2025-00462 | 2024-10-27 04:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Bacterial Infection; Muscle Weakness; Renal Failure; Presyncope | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT (ED) FOR PRESYNCOPE AND GENERALIZED WEAKNESS. THE PATIENT HAD VENTRICULAR TACHYCARDIA ON ARRIVAL TO THE ED AND HAD PROFOUND RIGHT VENTRICULAR DILATION ON ECHOCARDIOGRAM. IT WAS NOTED THAT THE RIGHT VENTRICULAR FAILURE WAS A KNOWN CONDITION AND THE PATIENT WAS ON INOTROPES BEFORE ADMISSION. CHEST X-RAY SHOWED RIGHT LOWER LOBE CONSOLIDATION/EFFUSION AND LUNG ULTRASOUND WITH LARGE EFFUSION WITH COMPLEX FLUID. THE PATIENT WAS TREATED WITH CYANOKIT (HYDROXOCOBALAMIN), METHYLENE BLUE, VASOPRESSIN, LEVOPHED, EPINEPHRINE, MILRINONE, BICARBONATE, LIDOCAINE, AND FLOLAN (EPOPROSTENOL). THE PATIENT WAS ALSO INTUBATED AND DIRECT CURRENT CARDIOVERSION WAS PERFORMED WHICH WAS NOT SUCCESSFUL. CHEST TUBE WAS INSERTED AND CULTURES WERE POSITIVE FOR STENOTROPHOMONAS MALTOPHILIA. BLOOD CULTURES WERE NEGATIVE. KEPPRA (LEVETIRACETAM), VANCOMYCIN, AND CEFEPIME WERE STARTED. THE PATIENT HAD WORSENING CREATININE AND BECAME ANURIC. THE PATIENT WAS TRANSITIONED TO COMFORT CARE AND PASSED AWAY ON (B)(6) 2024.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: B5 - NARRATIVE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING ARRHYTHMIA, RIGHT HEART FAILURE, RESPIRATORY FAILURE, INFECTION, RENAL FAILURE, AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS ARRHYTHMIA, RIGHT HEART FAILURE, AND INFECTION AS POTENTIAL LATE POST-IMPLANT COMPLICATIONS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200618&pc=BZD | 2518422-2025-100998 | 2024-09-12 04:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200658&pc=DSQ | 2916596-2025-00160 | 2025-01-02 05:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: G6 - TYPE OF REPORT: WAS NOT FILLED OUT IN INITIAL REPORT; SHOULD HAVE BEEN POPULATED WITH '30-DAY' AND 'INITIAL.' MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE (IFU), REVISION D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200731&pc=DSQ | 2916596-2025-00168 | 2024-12-30 05:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN IF AN AUTOPSY WOULD BE PERFORMED OR PROVIDED. IT WAS UNKNOWN IF THE PUMP WOULD BE EXPLANTED OR RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4 PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200824&pc=NPT | 2025587-2025-00502 | 2024-10-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-21 05:00 | NPT | EVOLUT PRO PLUS DCS | Failure to Advance; Difficult to Advance; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Pneumothorax; Vascular Dissection; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE, DUE TO THE FUSION OF THE RIGHT CORONARY CUSP (RCC) AND NON-CORONARY CUSP (NCC), AND HORIZONTAL ORIENTATION OF THE AORTA, IT WAS DIFFICULT TO PASS THE SYSTEM ACROSS THE NATIVE VALVE. THE DELIVERY CATHETER SYSTEM (DCS) EVENTUALLY CROSSED AND THE VALVE WAS IMPLANTED. THE PATIENT EXPERIENCED A SIGNIFICANT DROP IN BLOOD PRESSURE DURING THE PROCEDURE DUE TO BOTH AN AORTIC ROOT DISSECTION AND PERICARDIAL EFFUSION THAT WAS ATTRIBUTED TO THE DCS. THE PATIENT WAS INITIALLY STABILIZED AND MAINTAINING BLOOD PRESSURE USING A PARACENTESIS KIT FOR FLUID DRAINAGE BUT LATER DEVELOPED A PNEUMOTHORAX. DESPITE EFFORTS, THE BLOOD PRESSURE WAS UNABLE TO BE RESTORED, RESULTING IN THE PATIENT'S DEATH THE FOLLOWING DAY. PER THE PHYSICIAN, THE DEATH MAY HAVE POTENTIALLY BEEN CAUSED BY A POSSIBLE INADVERTENT PUNCTURE OF THE LUNG DURING THE PARACENTESIS RESULTING IN A PNEUMOTHORAX. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: L-EVPROP23-29 (LOT: 0012281961); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: EVOLUTR-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: ON (B)(6) 2024. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONCLUSION: A MEDICAL SAFETY ASSESSMENT WAS COMPLETED AND DETERMINED THAT, BASED ON THE INFORMATION PROVIDED, THE EVENT OF AORTIC ROOT DISSECTION IS ASSESSED AS LIKELY RELATED TO THE DELIVERY CATHETER SYSTEM (DCS) WITH UNSUCCESSFUL ATTEMPTS TO CROSS THE NATIVE VALVE LIKELY RESULTING IN A PERICARDIAL EFFUSION, PERICARDIOCENTESIS, PNEUMOTHORAX, HEMODYNAMIC INSTABILITY AND SUBSEQUENT DEATH. THE PATIENT'S CHALLENGING ANATOMY WAS LIKELY A CONTRIBUTING FACTOR TO THE EVENT. DIFFICULTIES ADVANCING THE DCS THROUGH THE ACCESS VESSEL IS KNOWN TO BE RELATED TO FACTORS SUCH AS PATIENT ANATOMY AND PHYSICIAN TECHNIQUE, INCLUDING GUIDEWIRE AND INTRODUCER SHEATH SELECTION. IN THIS CASE, IT WAS REPORTED THAT IT WAS DIFFICULT TO PASS THE SYSTEM ACROSS THE NATIVE VALVE DUE TO THE FUSION OF THE RCC AND NCC, AND HORIZONTAL ORIENTATION OF THE AORTA. HYPOTENSION IS AN EFFECT THAT IS HIGHLY DEPENDENT ON THE PATIENT'S PRE-PROCEDURAL CONDITION AND CAN OCCUR DESPITE A NORMALLY-FUNCTIONING DEVICE OR MODEL IMPLANT PROCEDURE. IN THIS CASE, IT WAS REPORTED THAT THE HYPOTENSION WAS DUE TO BOTH AN AORTIC ROOT DISSECTION AND PERICARDIAL EFFUSION. HYPOTENSION, VASCULAR COMPLICATIONS AND CARDIOVASCULAR INJURIES, SUCH AS DISSECTION AND EFFUSION, ARE A KNOWN POTENTIAL ADVERSE PATIENT EFFECT PER THE DEVICE INSTRUCTIONS FOR USE (IFU), AND ARE TYPICALLY RELATED TO PATIENT FACTORS (ANATOMY, COMORBIDITIES, ETC.), AND/ OR PROCEDURAL EFFECTS (SHEATH USED, USER TECHNIQUE, PUNCTURE CUT LOCATION, ETC.). PER THE PHYSICIAN, THE DEATH MAY HAVE POTENTIALLY BEEN CAUSED BY A POSSIBLE INADVERTENT PUNCTURE OF THE LUNG DURING THE PARACENTESIS RESULTING IN A PNEUMOTHORAX. THERE WAS NO INFORMATION TO SUGGEST THAT A FAILURE TO MEET MANUFACTURING SPECIFICATIONS WAS RELATED TO THESE EVENTS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21201136&pc=NPT | 2025587-2025-00503 | 2025-01-13 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-21 05:00 | NPT | EVOLUT PRO PLUS VALVE | Perivalvular Leak; Insufficient Information; Patient Device Interaction Problem | Bradycardia; Low Blood Pressure/ Hypotension; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE TRANSCATHETER AORTIC VALVE EVOLUT PRO+ 26MM IMPLANT, THE PATIENT EXPERIENCED HYPOTENSION AND BRADYCARDIA. THESE CONDITIONS WERE TREATED WITH MEDICATION AND A TEMPORARY PACEMAKER. IMAGING REVEALED A SIGNIFICANT PARAVALVULAR LEAK (PVL), WHICH WAS ADDRESSED WITH THREE POST-IMPLANT BALLOON DILATIONS USING A SIZE 22MM X 40 MM NON-MEDTRONIC BALLOON, RESULTING IN A MILD PVL. THE PROCEDURE WAS COMPLETED, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE NEXT MORNING, THE PATIENT AGAIN EXPERIENCED HYPOTENSION AND BRADYCARDIA. CARDIOPULMONARY RESUSCITATION MANEUVERS WERE PERFORMED BUT WERE UNSUCCESSFUL, AND THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: 0012436657); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVPROP23-29 (LOT: 0012155070); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21203140&pc=DXT | 2134243-2025-00001 | 2024-11-27 05:00 | Death | ACIST MEDICAL SYSTEMS, INC. | 2025-01-21 05:00 | DXT | ACIST | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Air Embolism | K010390 |  | 1 | Event Description: DURING A PERCUTANEOUS CORONARY INTERVENTION USING THE ACIST CVI INJECTION SYSTEM, AIR WAS INJECTED INTO A PATIENT. THE AIR INJECTION OCCURRED AFTER INSERTION OF A CORONARY GUIDING CATHETER AND ENGAGEMENT INTO THE LEFT MAIN CORONARY ARTERY AND UPON INJECTION OF CONTRAST MEDIA. THE PATIENT EXPERIENCED CARDIOGENIC CRASH AND SHOCK AND EXPIRED THE SAME DAY. Manufacturer Narrative: THE ACIST ANGIOGRAPHIC INJECTION SYSTEM, MODEL CVI, SYSTEM SERIAL NUMBER (B)(6), WAS RECEIVED FOR EVALUATION ON FEBRUARY 4, 2025. THE INJECTION SYSTEM WAS FUNCTIONALLY TESTED AND MET THE PRE-ESTABLISHED SPECIFICATIONS. THERE WAS NO EVIDENCE OF DEVICE MALFUNCTION RELATED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE HAVE BEEN REVIEWED AND NO INADEQUACIES WERE IDENTIFIED REGARDING WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR USE OF THE DEVICE. PER THE ACIST CVI USER'S MANUAL, THE AIR COLUMN DETECT SENSOR IS DESIGNED TO AID THE USER IN THE DETECTION OF AIR COLUMNS IN THE INJECTION LINE, BUT IT IS NOT DESIGNED TO REPLACE THE VIGILANCE AND CARE REQUIRED OF THE OPERATOR IN VISUALLY INSPECTING FOR AIR AND CLEARING AIR FROM THE ENTIRE PATIENT KIT AND ANGIOGRAPHIC CATHETER. THE AIR COLUMN DETECT MECHANISM IS TO BE USED IN CONJUNCTION WITH AND TO COMPLEMENT THE USER'S OTHER PROCEDURES FOR PREVENTING AIR INJECTIONS. THIS REPORT IS CLOSED. Manufacturer Narrative: THE ACIST ANGIOGRAPHIC INJECTION SYSTEM, MODEL CVI, SYSTEM SERIAL NUMBER (B)(6), HAS NOT YET BEEN RETURNED TO ACIST. THE CONSUMABLE KITS USED DURING THE EVENT WERE DISCARDED BY THE USER FACILITY AND THE LOT NUMBERS ARE UNKNOWN. THE CINE-ANGIOGRAMS HAVE BEEN REQUESTED FOR EVALUATION BUT HAVE NOT YET BEEN RETURNED TO ACIST. A CLINICAL ASSESSMENT BY AN ACIST MEDICAL ADVISORY BOARD MEMBER WILL BE PERFORMED ON THE CINE-ANGIOGRAMS UPON RECEIPT. A FOLLOW-UP REPORT WILL BE SUBMITTED TO FDA UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21203141&pc=JOL | 1417592-2025-00033 | 2024-12-24 05:00 | Death | MEDLINE INDUSTRIES, LP | 2025-01-21 05:00 | JOL | TUBING, SUCTION, 1/4" X 10', STRAIGHT | Connection Problem | Aspiration/Inhalation |  |  | 1 | Event Description: ACCORDING TO THE FACILITY ON (B)(6) 2024 THERE WAS A DELAY IN SUCTIONING THE PATIENT AS THE YANKAUER WOULDN'T "PROPERLY CONNECT TO THE SUCTION TUBING". Manufacturer Narrative: ACCORDING TO THE FACILITY ON (B)(6) 2024 THERE WAS A DELAY IN SUCTIONING THE PATIENT AS THE YANKAUER WOULDN'T "PROPERLY CONNECT TO THE SUCTION TUBING". PER THE FACILITY THE DEVICE WAS BEING USED "IN PREPARATION FOR INTUBATION" AND "YANKAUER WAS CONNECTED TO SUCTION TUBING AND SUCTION TURNED ON THEN PLACED AT HEAD OF BED". PER THE FACILITY THE DEVICE WAS TESTED PRIOR TO USE AND THE SUCTION WAS "TURNED TO HIGHEST SETTING FOR MAX SUCTION". PER THE FACILITY THE "PATIENT ASPIRATED, REQUIRED BRONCHOSCOPY TO REMOVE ASPIRATED MATERIAL". PER THE FACILITY THE PATIENT HAS EXPIRED. THE DEVICE WAS NOT SAVED FOR EVALUATION. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IT HAS BEEN DETERMINED THAT THE REPORTED EVENT CAUSED OR CONTRIBUTED TO SERIOUS INJURY, THEREFORE, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. MULTIPLE LOT NUMBERS WERE PROVIDED BY THE FACILITY AND WERE LISTED AS: 26023120010 AND 48623120001. Manufacturer Narrative: UPDATE TO H6: INVESTIGATION FINDINGS UPDATE TO H6: INVESTIGATION CONCLUSIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21203282&pc=NPT | 2015691-2025-00422 | 2025-01-08 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-21 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Perivalvular Leak; Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. AFTER VALVE DEPLOYMENT, MODERATE PARAVALVULAR LEAK WAS NOTED WITH UNSTABLE PRESSURE. POST-DILATION (-2CC) WAS PERFORMED BUT, PERICARDIAL TAMPONADE OCCURRED. DURING TEE VERIFICATION, THE BLOOD SELECTION FROM THE PERICARDIUM WAS NOT SUCCESSFUL, INDICATING A PROBABLE RUPTURE OF THE AORTIC ROOT (LCC). AN EMERGENCY OPEN CHEST SURGERY CONVERSION WAS PERFORMED. THE RUPTURE WAS SUTURED AND STOPPED THE BLEEDING HOWEVER, THE PATIENT DEVELOPED CARDIAC ARREST AND EXPIRED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS INDICATE POST-DILATION WAS PERFORMED TO TREAT MODERATE PARAVALVULAR LEAK WAS NOTED WITH UNSTABLE PRESSURE BUT, PERICARDIAL TAMPONADE OCCURRED. TEE INDICATE PROBABLE RUPTURE OF THE AORTIC ROOT (LCC). OPEN SURGERY WAS PERFORMED HOWEVER, THE PATIENT DEVELOPED CARDIAC ARREST AND EXPIRED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21203592&pc=BZD | 2518422-2025-101011 | 2024-11-06 05:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY, AND THE DEVICE WAS RETURNED. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204004&pc=BZD | 2518422-2025-101017 | 2024-11-20 05:00 | Death | RESPIRONICS, INC. | 2025-01-21 05:00 | BZD | DREAMSTATION | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204268&pc=DSQ | 2916596-2025-00170 | 2024-12-19 05:00 | Death | THORATEC CORPORATION | 2025-01-21 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Renal Failure; Septic Shock; Multiple Organ Failure; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT EXPERIENCE SEPTIC SHOCK WITH E. COLI BACTEREMIA AND HISTORY OF E. COLI DRIVELINE INFECTION (MFR # 2916596-2024-02329). THEY PROGRESSED INTO MULTI-SYSTEM ORGAN FAILURE REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY AND THEIR FAMILY MADE THE DECISION TO WITHDRAW CARE. THE PATIENT PASSED AWAY. Manufacturer Narrative: SECTION H3: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ""INTRODUCTION"", LISTS INFECTION, SEPSIS, AND MULTIPLE ORGAN FAILURES/DYSFUNCTIONS (INCLUDING RENAL DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21281209&pc=OZP | 2032227-2025-118929 | 2025-01-10 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-01-31 05:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Fatigue; Hyperglycemia; Renal Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AT HOSPITAL. THE CAUSE OF DEATH WAS HYPERGLYCEMIA, CARDIAC ARREST, AND KIDNEY FAILURE. THE CUSTOMER WAS WEAK AND IN PAIN WITH BLOOD GLUCOSE LEVELS OF 500 MG/DL DURING HOSPITALIZATION. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-398A, MMT-1780KPK.TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER WAS NOT USING THE INSULIN PUMP SYSTEM AT THE TIME OF REPORTED EVENT. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-398A. PRODUCT RETURN REQUESTED FOR MMT-1780KPK. CUSTOMER DECLINED TO RETURN OR IS UNABLE TO RETURN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21281505&pc=DSQ | 2916596-2025-00315 | 2024-08-16 04:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED HYPERCAPNIC RESPIRATORY FAILURE IN THE HOSPITAL THAT REQUIRED CONTINUOUS BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) AND ICU LEVEL OF CARE. THE PATIENT WAS ALSO DIALYSIS DEPENDENT AT THE TIME. THE PATIENT WAS OBTUNDED AND UNABLE TO WEAN OFF BIPAP NOR PARTICIPATE IN THEIR OWN CARE AND RECOVERY. THE RECOMMENDATION FOR FEEDING TUBE PLACEMENT WAS DISCUSSED BY PROVIDERS WITH THE PATIENT'S FAMILY, AND AFTER LENGTHY DISCUSSION AROUND THE PATIENT'S QUALITY OF LIFE AND LIKELIHOOD FOR RECOVERY, THE FAMILY DECIDED IN CORROBORATION WITH THE PATIENT'S ADVANCED HEALTH CARE DIRECTIVE TO TRANSITION OF COMFORT FOCUSED TREATMENT. THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT RESPIRATORY FAILURE AND RENAL DYSFUNCTION DID NOT EXIST PRIOR TO THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE CAUSE WAS UNKNOWN. THE DEVICE WAS NOT THOUGHT TO HAVE CAUSED OR CONTRIBUTED TO THE RESPIRATORY FAILURE OR RENAL DYSFUNCTION. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (HM3 LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21281762&pc=BZD | 2518422-2025-101552 | 2024-12-31 05:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | BZD | PHILIP'S CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A PHILIPS CPAP DEVICE. THE PATIENT'S WIFE HAS REQUESTED THAT ALL COMMUNICATION CEASE, AS THE PATIENT PASSED AWAY THREE YEARS AGO. NO ADDITIONAL DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. FURTHERMORE, NO DEVICE INFORMATION WAS SUPPLIED, AS THE DEVICE HAS BEEN DISCARDED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21281851&pc=BZD | 2518422-2025-101555 | 2024-10-29 04:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S WIFE NO LONGER HAS THE DEVICE AS HER HUSBAND PASSED AWAY TWO YEARS AGO. NO ADDITIONAL DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21281997&pc=BZD | 2518422-2025-101557 | 2024-12-06 05:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S WIFE CONTACTED US TO REPORT THAT SHE HAD RECEIVED A LETTER REGARDING A REPLACEMENT DEVICE FOR HER HUSBAND, WHICH HE NEVER RECEIVED. SHE CLARIFIED THAT HER HUSBAND HAS PASSED AWAY AND NO LONGER REQUIRES A REPLACEMENT. NO FURTHER DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21282555&pc=OZD | 1220648-2025-25932 | 2025-01-14 05:00 | Death | ABIOMED, INC. | 2025-01-31 05:00 | OZD | IMPELLA CP | Fluid/Blood Leak | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE IMPELLA IMPLANT WAS NOTED AS UNEVENTFUL AND THE PATIENT WAS ADDITIONALLY PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION DUE TO A WEAK RIGHT VENTRICLE AND A WEAK LUNG. WHILE ON SUPPORT THE PATIENT EXPERIENCED ACCESS SITE BLEEDING AND HAD A PURGE FLUID LEAK. APPROXIMATELY, FOUR DAYS AFTER START OF THE REPORT IS WHEN THE PATIENT HAD AN ACCESS SITE BLEED AND WAS ADMINISTERED "AT LEAST" TWO UNITS OF RED BLOOD CELLS. HEMOSTASIS WAS NOT ACHIEVED, HOWEVER. ADDITIONAL ACTION TAKEN AS A RESULT OF THE EVENT IS UNKNOWN. HOWEVER, IT WAS NOTED THAT PURGE FLUID WAS LEAKING FROM THE RED IMPELLA PLUG (NOT AT THE FILTER AND NOT AT THE YELLOW LUERLOCK), AND PURGE VALUES WERE WITHIN THE NORMAL RANG. IT WAS EXPLAINED, HOWEVER, THAT IT WAS IMPOSSIBLE TO UNDERSTAND WHETHER AND HOW MUCH PURGE FLUID IS STILL FLOWING THROUGH THE MOTOR GAP, AND THE PUMP CAN STOP AT ANY TIME. PUMP REPLACEMENT AS SOON AS POSSIBLE WAS RECOMMENDED AND TO CLOSELY MONITOR THE MOTOR CURRENT UNTIL THEN. EXTENSIVE DISCUSSION ABOUT ESCALATION WAS MADE, AS THE PATIENT STILL HAD A RELATIVELY HIGH CATECHOLAMINE REQUIREMENT. THE DECISION ON ESCALATION WAS TO BE MADE THE FOLLOWING DAY. THE FOLLOWING DAY, HOWEVER, THE PATIENT WOULD EXPIRE. THE PATIENT WAS IN MULTI-ORGAN FAILURE (MOF) AND CARE WAS WITHDRAWN. THE PATIENT REQUIRED EXTRACORPOREAL LIFE SUPPORT ALONG WITH IMPELLA MCS AND WOULD PASS AWAY NINE DAYS AFTER START OF SUPPORT IT APPEARS BECAUSE OF MOF. NO FURTHER DETAILS REGARDING THE FINAL OUTCOME OF THE BLEEDING EVENT AND PURGE FLUID LEAK WERE PROVIDED. GIVEN THE IMPELLA-RELATED EVENTS, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS A POSSIBLE CONTRIBUTING FACTOR OF THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND PURGE LEAK ¿ PUMP HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INVESTIGATION. A DATA LOG WAS NOT PROVIDED FOR THE CASE RUN. NO PHYSICAL ABNORMALITIES WERE OBSERVED ON THE RETURNED PRODUCT. THE REPO SHEATH WAS FOUND SUTURED THROUGH THE CATHETER, AND THIS IS WHERE THE LEAK WAS REPRODUCED. THE CATHETER WAS FURTHER CUT DOWN, AND THE DAMAGED PURGE LINE WAS OBSERVED LEAKING. THE PURGE PRESSURE WAS WITHIN SPECIFICATIONS DURING TESTING DUE TO THE SLOW LEAK. THE LEAK INSIDE THE CATHETER COULD REACH THE RED HANDLE, AS MENTIONED IN CLINICAL DETAILS. THE CAUSE OF THE PURGE LEAK WAS DAMAGED PURGE TUBING AT THE 75 CENTIMETER MARK DUE TO THE APPLICATION OF INAPPROPRIATE SUTURES. THE CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS PATIENT CONDITION RELATED TO COAGULATION DISORDER. NO HEPARIN WAS GIVEN IN THE PURGE FLUID AT THE TIME OF BLEED. D.9 DATE DEVICE RETURNED/INFORMATION WAS ADDED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: CLEANING ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿ E.4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25932 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H.6 CODE 4641 WAS REPORTED INCORRECTLY ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25932. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED FROM THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25932 AND WERE ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21282573&pc=CBK | 8010042-2025-0000183 | 2025-01-14 05:00 | Death | MAQUET CRITICAL CARE AB | 2025-01-31 05:00 | CBK | SERVO-S BASE UNIT | Connection Problem; No Flow | Cardiac Arrest; Respiratory Arrest | K123149 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTILATOR WAS NOT CONNECTED TO O2 SUPPLY WHILE ON PATIENT. THE PATIENT DIED THE DAY AFTER THE EVENT. MANUFACTURER'S REF.#: (B)(4).Event Description: MANUFACTURER'S REF.#: (B)(4). Manufacturer Narrative: THE UDI INFORMATION IS NOT AVAILABLE SINCE THE DEVICE WAS MANUFACTURED BEFORE 2015. Manufacturer Narrative: THE VENTILATOR WAS INVESTIGATED BY OUR FIELD SERVICE ENGINEER. THE VENTILATOR PASSED ALL FUNCTIONAL AND SAFETY TESTS AND WAS CLEARED FOR CLINICAL USE. EVALUATION OF THE RECEIVED LOGS REVEALED THAT THE USER STARTED THE VENTILATION AT 12:55:41 (B)(6) 2025 AND AT 12:55:48 (B)(6) 2025 A MEDIUM PRIORITY ALARM FOR LOW O2 SUPPLY PRESSURE WAS ACTIVATED (AUDIBLE). PROBABLY THE HIGH PRIORITY LOW EXPIRATORY VOLUME ALARM WAS SHOWN IN THE ALARM/MESSAGE AREA AS THIS ALARM HAS A HIGHER PRIORITY THAN THE O2 SUPPLY PRESSURE ALARM. FROM (B)(6) 2025 12:55:48 TO 12:55:54 THE USER TRIED TO ADJUST THE SET O2 CONCENTRATION. AFTER A USER ADJUSTMENT OF THE O2 CONCENTRATION, THE O2 CONCENTRATION ALARMS ARE DELAYED BY APPROXIMATELY 50 S IN ORDER TO PREVENT NUISANCE ALARM DURING STABILIZATION OF THE O2 CONCENTRATION IN THE SYSTEM. THE HIGH PRIORITY LOW O2 CONCENTRATION ALARM IS ACTIVATED (AUDIBLE AND VISIBLE) AT 12:56:40 (B)(6) 2025 AND THE USER SILENCES THE ALARMS AFTER THREE SECONDS AT 12:56:43. THE TECHNICAL LOG DOES NOT CONTAIN ANY TECHNICAL ERROR CODES TO INDICATE THAT THERE WAS A VENTILATOR MALFUNCTION. THE VENTILATOR WAS PASSED PRE-USE CHECK SUCCESSFULLY PRIOR TO VENTILATION START AT DATE OF EVENT. ACCORDING TO USER FACILITY, THE O2 SUPPLY WAS NOT CONNECTED FROM THE BEGINNING. THE VENTILATOR DETECTED THE LACK OF O2 SUPPLY PRESSURE AND ALARMED FOR IT ACCORDING TO SPECIFICATION. THERE IS NO ALLEGATION FROM THE USER THAT A VENTILATOR MALFUNCTION IN SOME WAY CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. THE CONCLUSION IS THAT THE VENTILATOR HAS BEEN WORKING ACCORDING TO SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21283212&pc=NPT | 2025587-2025-00892 | 2024-11-28 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-31 05:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Renal Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 23 DAYS AFTER THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH A HISTORY OF DECLINING RENAL FUNCTION, THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL. AFTER THE TRANSFER, THE PATIENT'S RENAL FUNCTION BEGAN TO DECLINE FURTHER AND RENAL FAILURE WAS DIAGNOSED. HEMODIALYSIS WAS NOT INTRODUCED AT THE REQUEST OF THE PATIENT'S FAMILY. THE PATIENT PASSED AWAY FROM RENAL FAILURE APPROXIMATELY 2 WEEKS LATER. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT'S UNDERLYING MEDICAL HISTORY DID NOT CONTRIBUTED TO THE DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21283439&pc=OZD | 1220648-2025-25939 | 2025-01-19 05:00 | Death | ABIOMED, INC. | 2025-01-31 05:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Hemolysis; Gastrointestinal Hemorrhage; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED AN IMPELLA 5.5 WHILE ON PATIENT SUPPORT HAD HIGH PURGE PRESSURE AND PURGE SYSTEM BLOCKED ALARMS. NO KINKS WERE OBSERVED IN THE TUBING. IMPELLA 5.5 WAS DECREASED TO P-3 AND P-2 WHERE FLOWS WERE APPROPRIATE. TRANSESOPHAGEAL ECHOCARDIOGRAPHY WAS PERFORMED TO CONFIRM IMPELLA 5.5 OUTLET POSITION RELATED TO CENTRALLY CANNULATED AORTIC CANNULA. THE IMPELLA 5.5 MEASURED FIVE CENTIMETERS IN THE LEFT VENTRICLE (LV) AND WAS PROPERLY POSITIONED ABOVE THE AORTIC VALVE. HEPARIN PURGE WAS STARTED ALONG WITH SYSTEMIC HEPARIN DRIP. PURGE FLOW, PURGE PRESSURE, AND LV WAVEFORM WAS SPONTANEOUSLY RETURNED TO NORMAL ALONG WITH NORMAL FLOWS AT P-4. THE PATIENT REMAINS ON IMPELLA SUPPORT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT¿S LABS WERE HEMOLYZING AND THAT THE PATIENT DEVELOPED GASTROINTESTINAL (GI) BLEED. BIVAL WAS DISCONTINUED. THE NEXT DAY THE GI BLEED WAS STABILIZED. THE FAMILY DECIDED TO WITHDREW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: B.1 ADDED ADVERSE EVENT BASED ON THE ADDITIONAL INFORMATION RECEIVED. B.2 ADDED DEATH AND DATE OF DEATH DUE TO NEW INFORMATION RECEIVED. B.5 ADDITIONAL INFORMATION WAS RECEIVED. D.6B EXPLANT DATE WAS ADDED. H.1 REVISED TO DEATH BASED ON THE ADDITIONAL INFORMATION RECEIVED. H.6 ADDED ADDITIONAL CODES FOR THE NEW INFORMATION RECEIVED. H.10 ADDITIONAL INSTRUCTIONS FOR USE WERE ADDED BASED ON THE ADDITIONAL INFORMATION RECEIVED. THE INVESTIGATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR ANALYSIS. THE DATA LOGS CONFIRMED PURGE PRESSURE HIGH ALARMS. THE LOGS SHOWED A MOTOR CURRENT SPIKE WITH SPEED DEVIATION AND A DROP IN PURGE FLOW RIGHT BEFORE HIGH PURGE PRESSURE OCCURRENCE. IT WAS UNKNOWN IF TISSUE PLASMINOGEN ACTIVATOR WAS ADDED TO THE PURGE. THERE WAS HIGH PURGE PRESSURE FOR ABOUT TWO HOURS AFTER WHICH VALUES STARTED DECREASING TOWARDS NORMAL PURGE PRESSURE VALUES. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY MOTOR CURRENT SPIKE IN THE LOGS BEFORE HEMOLYSIS OCCURRENCE. THE ROOT CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL WAS NOT DETERMINED AS INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284137&pc=KDI | 0008030665-2025-00262 | 2025-01-15 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-31 05:00 | KDI | 2008T HD SYS. CDX W/BIBAG BLUE STAR | Adverse Event Without Identified Device or Use Problem | Loss of consciousness | K173972 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS UNKNOWN PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). ON (B)(6) 2025, THIS UNKNOWN PATIENT REQUESTED HD THERAPY BE PAUSED TO UTILIZE THE RESTROOM. WHILE RETURNING THE PATIENT¿S BLOOD, IT WAS REPORTED THE PATIENT BECAME UNRESPONSIVE. THE HD REGISTERED NURSE (HDRN) REPORTEDLY INITIATED CARDIOPULMONARY RESUSCITATION (CPR) MEASURES, HOWEVER THE PATIENT EXPIRED. THE 2008T HEMODIALYSIS SYSTEM WAS SEQUESTERED POST-EVENT, HOWEVER NO DOCUMENTATION REGARDING THE ANY POST-EVENT DEVICE TESTING WAS OBTAINED. IT SHOULD BE NOTED THAT PRIOR TO TREATMENT, THE 2008T HEMODIALYSIS SYSTEMS¿ CONDUCTIVITY WAS WITHIN ACCEPTABLE LIMITS, NO PRETREATMENT ALARMS WERE NOTED, AND THERE WERE NO DEFECTS ON EITHER THE BLOODLINES OR DIALYZER NOTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, DEATH CERTIFICATE, TREATMENT RECORDS, MACHINE RECORDS) WERE UNSUCCESSFUL.Event Description: IT WAS REPORTED TO FRESENIUS THAT THIS UNKNOWN PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). ON (B)(6) 2025, THIS UNKNOWN PATIENT REQUESTED HD THERAPY BE PAUSED TO UTILIZE THE RESTROOM. WHILE RETURNING THE PATIENT¿S BLOOD, IT WAS REPORTED THE PATIENT BECAME UNRESPONSIVE. THE HD REGISTERED NURSE (HDRN) REPORTEDLY INITIATED CARDIOPULMONARY RESUSCITATION (CPR) MEASURES, HOWEVER THE PATIENT EXPIRED. THE 2008T HEMODIALYSIS SYSTEM WAS SEQUESTERED POST-EVENT, HOWEVER NO DOCUMENTATION REGARDING THE ANY POST-EVENT DEVICE TESTING WAS OBTAINED. IT SHOULD BE NOTED THAT PRIOR TO TREATMENT, THE 2008T HEMODIALYSIS SYSTEMS¿ CONDUCTIVITY WAS WITHIN ACCEPTABLE LIMITS, NO PRETREATMENT ALARMS WERE NOTED, AND THERE WERE NO DEFECTS ON EITHER THE BLOODLINES OR DIALYZER NOTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, DEATH CERTIFICATE, TREATMENT RECORDS, MACHINE RECORDS) WERE UNSUCCESSFUL. Manufacturer Narrative: PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. ADDITIONALLY, NO ON-SITE EVALUATION WAS PERFORMED BY A FRESENIUS FIELD SERVICE TECHNICIAN (FST). A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HD THERAPY UTILIZING A 2008T HEMODIALYSIS SYSTEM AND THE SERIOUS ADVERSE EVENTS OF LOSS OF CONSCIOUSNESS, AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT EXPIRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE MEDICAL ASSESSMENT. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR APPROXIMATELY 45% OF ALL ESRD DEATHS, AND APPROXIMATELY 25% OF THOSE DEATHS ARE DUE TO SUDDEN CARDIAC DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284575&pc=DSQ | 2916596-2025-00456 | 2024-05-13 04:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Syncope/Fainting; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED TO AN OUTSIDE FACILITY WITH COMPLAINTS OF LOW FLOWS AND SYNCOPE. THE PATIENT WAS FOUND TO HAVE AN AORTIC ROOT THROMBUS ON ECHOCARDIOGRAM AND RIGHT VENTRICULAR FAILURE (RVF). THEY THEN DEVELOPED MULTI SYSTEM ORGAN FAILURE. AN IMPELLA RP WAS IMPLANTED AT THE OUTSIDE FACILITY AND THEN THE PATIENT WAS TRANSFERRED TO THE IMPLANTING CENTER. THE PATIENT WAS MAINTAINED ON A HEPARIN DRIP DUE TO THE THROMBUS. THE PATIENT ULTIMATELY PASSED AWAY DUE TO MULTI SYSTEM ORGAN FAILURE ON (B)(6) 2024. IT WAS REPORTED THAT THE HEARTMATE 3 FUNCTIONED AS INTENDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, THE LOW FLOW ALARMS COULD NOT BE CONFIRMED THROUGH THIS EVALUATION. A SPECIFIC CAUSE FOR THESE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PATIENT ULTIMATELY EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. IT WAS COMMUNICATED THAT THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE), THROMBOEMBOLISM, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 OF THE IFU, (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINE SYSTEM CONTROLLER ALARMS, INCLUDING THE LOW FLOW HAZARD ALARM, AND PROVIDE INFORMATION REGARDING HOW TO RESPOND TO AND TROUBLESHOOT THE ALARMS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284603&pc=OUT | 2029214-2025-00266 | 2025-01-21 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-01-31 05:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Rupture; Hemorrhagic Stroke | P100018 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A NEUROVASCULAR FLOW DIVERSION PROCEDURE USING THE PIPELINE FLEX SHIELD 3.50MM X 16MM FOR AN UNRUPTURED SACCULAR ANEURYSM LOCATED IN THE RIGHT POSTERIOR COMMUNICATING ARTERY. THE ANEURYSM HAD A MAXIMUM DIAMETER OF 5MM AND A NECK DIAMETER OF 3.5MM, WITH MINIMAL VESSEL TORTUOSITY. POST-PROCEDURE, THE PATIENT EXPERIENCED A RUPTURE AND SUBARACHNOID HEMORRHAGE, WHICH WAS IDENTIFIED AS THE CAUSE OF DEATH. THE ANGIOGRAPHIC RESULT POST-PROCEDURE WAS SATISFACTORY. THE PATIENT HAD A HISTORY OF SUBARACHNOID RUPTURE AND A FAMILY HISTORY OF RUPTURE. IT WAS NOTED THAT THE POST-PROCEDURE RUPTURE WAS NOT BELIEVED TO BE DIRECTLY RELATED TO THE DEVICE IMPLANTED, AND WAS CONSIDERED A COINCIDENCE DUE TO THE PATIENT'S MEDICAL HISTORY. THE DEVICE REMAINS IMPLANTED AND IN SERVICE.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT PASSED AWAY THE WEEK OF (B)(6) 2025. THE CAUSE OF DEATH WAS NOT DETERMINED BUT WAS THOUGHT TO BE RELATED TO A PARENCHYMAL STROKE/ HEMORRHAGE IN THE CONTRALATERAL HEMISPHERE. THERE IS NO KNOWLEDGE OF MEDICAL EV ENT/PROCEDURES LEADING UP TO THE PATIENT'S DEATH. Manufacturer Narrative: B2 DATE OF DEATH MONTH VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284627&pc=DSQ | 2916596-2025-00671 | 2025-01-25 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST ON (B)(6) 2025. WHETHER THE EVENT WAS DEVICE OR THERAPY RELATED WAS UNKNOWN. IT WAS UNKNOWN IF THE PUMP WAS EXPLANTED. WHETHER AN AUTOPSY WOULD BE PERFORMED OR IF THE PUMP WOULD BE RETURNED FOR EVALUATION WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284671&pc=DSQ | 2916596-2025-00385 | 2025-01-14 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Cardiovascular Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD CHOSEN COMFORT CARE AFTER DEVELOPING SEVERE PERIPHERAL ARTERY DISEASE (PAD) WITH DECREASED FUNCTION IN LEGS AND DECREASED QUALITY OF LIFE (QOL). THE PATIENT PASSED AWAY ON (B)(6) 2025.Event Description: IT WAS NOTED THAT THE PATIENT DID NOT HAVE PAD PRE-IMPLANT, BUT THE LVAD WAS NOT THOUGHT TO HAVE CONTRIBUTED TO OR WORSENED THE PAD. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: D4:DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284674&pc=DSQ | 2916596-2025-00412 | 2024-08-05 04:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Purulent Discharge; Erythema; Muscle Weakness; Sepsis; Discomfort; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH A KNOWN DRIVELINE INFECTION WAS FEELING WELL SINCE BEING DISCHARGE HOME ON (B)(6) 2025. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) DRESSING WAS STILL WITH FREQUENT SATURATION WITH WHITISH DISCHARGE. THE PATIENT DENIED ANY OTHER LVAD SPECIFIC ISSUES. ON (B)(6) 2024, THE PATIENT PRESENTED TO INFECTIOUS DISEASE FOR AN OFFICE VISIT WITH WORSENING LVAD DRAINAGE. THE PATIENT STARTED ON DOXYCYCLINE 100 MILLIGRAM (MG) TABLETS BY MOUTH EVERY 12 HOURS AND BACTRIM 800-160 MG TABLET WITH THE PLAN TO TAKE FOR AT LEAST 4 MONTHS. ON (B)(6) 2024 THE PATIENT STOPPED AMOXICILLIN AND CONTINUED BACTRIM AND DOXYCYCLINE. THE PATIENT WAS ADMITTED ON (B)(6) 2024 FOR WORSENING ERYTHEMA, DISCOMFORT, AND PURULENCE AT DRIVELINE SITE. THE PATIENT RECEIVED INTRAVENOUS ANTIBIOTICS AND FURTHER EVALUATION WITH COMPUTED TOMOGRAPHY (CT) OF THE ABDOMEN AND PELVIS WAS TO BE PERFORMED. ON (B)(6) 2024, THE PATIENT UNDERWENT IRRIGATION AND DEBRIDEMENT OF THE DRIVELINE WITH REPOSITIONING, AND PLACEMENT OF A WOUND VACUUM DRESSING. A CT SCAN OF THE ABDOMEN WAS PERFORMED ON (B)(6) 2024. THE PATIENT WAS DISCHARGED ON (B)(6) 2024 AND FELT VERY WEAK THAT WEEK. THEY WERE NOT ABLE TO WALK AND HAD DIFFICULTY STANDING. DISCHARGE TO REHABILITATION WAS RECOMMENDED BY PHYSICAL THERAPY. THE PATIENT REFUSED AND REQUESTED TRANSFER TO HOME WITH HOME SERVICES. PER INFECTIOUS DISEASE, BACTRIM WAS STOPPED (B)(6) 2024 AND THE PATIENT WAS ON LINEZOLID 600MG Q12H FROM (B)(6) 2024. THE PATIENT CONTINUED IMIPENEM 1000 MG IV TWICE DAILY (FOR A PLAN OF AT LEAST 6-8 WEEKS) AND DOXYCYCLINE 100 MG TWICE DAILY (INDEFINITELY). ON (B)(6) 2024, THE PATIENT WAS ADMITTED FOR ELEVATED INTERNATIONAL NORMALIZED RATIO (INR) 8. DURING THIS ADMISSION, INFECTIOUS DISEASE AND CARDIOTHORACIC SURGERY WERE CONSULTED FOR MANAGEMENT OF THE DRIVELINE INFECTION. THERE WAS NO FURTHER DEBRIDEMENT OPTIONS AS THE PATIENT HAD BEEN DEEMED A PROHIBITIVE RISK TO UNDERGO PUMP EXCHANGE. PATIENT REQUESTED TO BE NOT RESUSCITATE (DNR). ON (B)(6) 2024, THE PATIENT HAD LOW FLOW ALARMS. PUMP PARAMETERS WERE SPEED OF 5300 ROTATIONS PER MINUTE (RPM), FLOW OF 2.8 LITERS PER MINUTE (LPM), PULSATILITY INDEX (PI) OF 11.5, POWER OF 4.1 WATTS, AND MEAN ARTERIAL PRESSURE (MAP) OF 64. GIVEN THE LACK OF MEDICAL AND SURGICAL OPTIONS, THE PATIENT DECIDED TO BE TRANSITIONED TO HOSPICE AND REQUESTED THAT THEIR LVAD BE TURNED OFF. ON (B)(6) 2024, THE LVAD WAS TURNED OFF. THE PATIENT WAS NOTED TO BE PULSELESS AND APNEIC. THEIR PUPILS WERE FIXED AND DILATED, AND THE PATIENT WAS PRONOUNCED DECEASED. RELATED INFECTION MFR: 2916596-2024-04467.Event Description: IT WAS REPORTED THAT THE PATIENT WITH A KNOWN DRIVELINE INFECTION WAS FEELING WELL SINCE BEING DISCHARGE HOME ON (B)(6) 2024. THE CAUSE OF DEATH WAS REPORTED TO BE SEPSIS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B5: EVENT DETAILS CORRECTED. SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION AND SEPSIS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE SUBMITTED FOR REVIEW. ALTHOUGH A SPECIFIC CAUSE FOR THE REPORTED ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION, THE ACCOUNT STATED THE ALARMS WERE DUE TO HYPOTENSION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), SEPSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284872&pc=DSQ | 2916596-2025-00468 | 2025-01-14 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEART FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED HEART FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284944&pc=FKX | 0008030665-2025-00266 | 2025-01-11 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-01-31 05:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: ON (B)(6) 2025, A HOME CARE FACILITY ADMINISTRATOR AND PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PDRN, IT WAS REPORTED THIS PATIENT EXPIRED ON (B)(6) 2025 AT HOME DUE TO A CARDIAC ARREST, ATTRIBUTED TO MULTIMORBIDITY THAT INCLUDED A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE. IT WAS EXPLAINED THAT THE PATIENT HAD BEEN IN AND OUT OF THE HOSPITAL IN RECENT MONTHS FOR A DECLINE IN OVERALL HEALTH AND FOR REASONS UNRELATED TO PD THERAPY OR THE USE OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). THE PATIENT WAS FOUND UNRESPONSIVE BY FAMILY ON THE MORNING OF (B)(6) 2025 WHILE STILL CONNECTED TO HER CYCLER. EMERGENCY SERVICES WERE NOT CONTACTED, AND THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME WITHOUT TRANSPORT TO THE HOSPITAL. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO HER DEATH. THE CAUSE OF THIS PATIENT¿S CARDIAC ARREST AND SUBSEQUENT DEATH CAN BE ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE WITH MULTIMORBIDITY. THERE WAS NO INDICATION OF AN ASSOCIATED DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS CONFIRMED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285291&pc=NIK | 2017865-2025-12010 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-31 05:00 | NIK | NEUTRINO NXT HF | Failure to Deliver Shock/Stimulation; Under-Sensing | Arrhythmia; No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12011, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12012, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12014. IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM AFTER THEY COLLAPSED IN THE HOME. THE PATIENT EXPIRED THAT SAME DAY. A REVIEW OF MERLIN.NET FOUND ABORTED SHOCK AND UNDER-SENSING DURING A VENTRICULAR TACHYCARDIA (VT) EPISODE. HOWEVER, THE PHYSICIAN CONCLUDED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO THE PERFORMANCE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA) OR LEFT VENTRICULAR LEADS. Manufacturer Narrative: THE PROVIDED SESSION RECORDS WERE REVIEWED BY TECHNICAL SERVICES AND IT WAS OBSERVED THAT THE ARRHYTHMIA WAS BELOW PROGRAMMED THERAPY ZONE; HENCE, NO THERAPY WAS DELIVERED. THE DEVICE WAS PERFORMING PER PROGRAMMED SETTINGS. DHR REVIEWED AND FOUND TO BE COMPLETED. THE PRODUCT PASSED ALL QUALITY CONTROL TESTS PRIOR TO ITS DISTRIBUTION. Manufacturer Narrative: CORRECTION: H6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285297&pc=NVN | 2017865-2025-12011 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-31 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12010, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12012, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12014. IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM AFTER THEY COLLAPSED IN THE HOME. THE PATIENT EXPIRED THAT SAME DAY. A REVIEW OF MERLIN.NET FOUND ABORTED SHOCK AND UNDER-SENSING DURING A VENTRICULAR TACHYCARDIA (VT) EPISODE. HOWEVER, THE PHYSICIAN CONCLUDED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO THE PERFORMANCE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA) OR LEFT VENTRICULAR LEADS. Manufacturer Narrative: CORRECTION: UPON REVIEW, THE PULSE LEAD SHOULD NOT HAVE BEEN SUBMITTED AS A MEDICAL DEVICE REPORT (MDR) AS THE EVENT DID NOT INDICATE A MALFUNCTION CAUSED A SERIOUS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285313&pc=DSQ | 2916596-2025-00764 | 2006-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT, US, EP | Mechanical Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT IT WAS THAT A TOUGHER DRIVELINE WAS NEEDED. IT WAS NOTED THAT THERE WAS DEATHS IN THE HEARTMATE II TRIAL FROM CAR DOOR TRANSECTIONS.Event Description: IT WAS REPORTED THAT THE ISSUE OCCURRED MORE THAN ONCE (MFR # 2916596-2025-02228). Manufacturer Narrative: SECTION A - SPECIFIC PATIENT INFORMATION IS UNKNOWN. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B: DATE OF DEATH WAS ESTIMATED AS DATE OF EVENT. MANUFACTURER'S INVESTIGATION CONCLUSION: IT WAS REPORTED BY THE ACCOUNT THAT A TOUGHER DRIVELINE WAS NEEDED. IT WAS NOTED THAT IN A HEARTMATE II BRIDGE TO TRANSPLANT (BTT) TRIAL, THE STRENGTH OF THE HEARTMATE II DRIVELINE WAS TESTED VIA CAR DOOR TRANSECTIONS. EQUIPMENT WAS "WEAR" TESTED BY SLAMMING IN DOORS. IT WAS NOTED THAT THE TRIAL RESULTED IN DEATHS, AND THAT THE HEARTMATE II DRIVELINE WAS VERY DELICATE AND FLIMSY IN COMPARISON TO THAT OF HEARTMATE 3. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE II DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE (HM) II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE II PATIENT HANDBOOK DOCUMENT ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. THE PATIENT HANDBOOK INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THE PATIENT THINKS THAT, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. SECTIONS 6 AND 8 OF THE HMII IFU, AS WELL AS SECTIONS 4 AND 6 OF THE HMII PATIENT HANDBOOK, PROVIDE INFORMATION REGARDING HOW TO CARE FOR THE DRIVELINE. THE PATIENT HANDBOOK ALSO INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED). NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285565&pc=DSQ | 2916596-2025-00439 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285571&pc=NVY | 2017865-2025-12012 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-31 05:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Deliver Shock/Stimulation; Under-Sensing; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Insufficient Information | P950022 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12010, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12011, RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12014. IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM AFTER THEY COLLAPSED IN THE HOME. THE PATIENT EXPIRED THAT SAME DAY. A REVIEW OF MERLIN.NET FOUND ABORTED SHOCK AND UNDER-SENSING DURING A VENTRICULAR TACHYCARDIA (VT) EPISODE. HOWEVER, THE PHYSICIAN CONCLUDED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO THE PERFORMANCE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA) OR LEFT VENTRICULAR LEADS.Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM AFTER THEY COLLAPSED IN THE HOME. THE PATIENT EXPIRED THAT SAME DAY. A REVIEW OF MERLIN.NET FOUND ABORTED SHOCK AND UNDER-SENSING DURING A VENTRICULAR TACHYCARDIA (VT) EPISODE. HOWEVER, THE PHYSICIAN CONCLUDED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO THE PERFORMANCE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA) OR LEFT VENTRICULAR LEADS. Manufacturer Narrative: DHR REVIEWED AND FOUND TO BE COMPLETE. THE PRODUCT PASSED ALL QUALITY CONTROL TESTS PRIOR TO ITS DISTRIBUTION. Manufacturer Narrative: CORRECTION:H6 - MEDICAL DEVICE PROBLEM CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285573&pc=OZD | 1220648-2025-25943 | 2025-01-13 05:00 | Death | ABIOMED, INC. | 2025-01-31 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE AND EXPERIENCED ACCESS SITE BLEEDING WITH BLOOD ADMINISTRATION. CARE WAS EVENTUALLY WITHDRAWN, AND THE PATIENT MET THEIR DEMISE. PRIOR TO THE DAY OF STARTING MECHANICAL CIRCULATORY SUPPORT, THE PATIENT UNDERWENT A LEFT HEART CATHETERIZATION WHICH REVEALED LEFT ANTERIOR DESCENDING ARTERY (LAD) LESION WHICH WAS TREATED WITH ONE DRUG-ELUTING STENT. THE PATIENT WAS STABLE AND SENT TO THE UNIT FOR RECOVERY. HOWEVER, SUBSEQUENTLY, THE PATIENT DECOMPENSATED WITH DROP IN BLOOD PRESSURE REQUIRING THREE AMPS OF EPINEPHRINE. LEVOPHED AND DOPAMINE DRIPS WERE STARTED, AND THE PATIENT WAS TAKEN BACK TO THE CATHETERIZATION LAB. IMAGING REVEALED NON-OBSTRUCTIVE CORONARY ARTERY DISEASE WITH PATENT LAD STENT. RIGHT HEART CATHETERIZATION WAS COMPLETED WITH ELEVATED PRESSURES. THE PATIENT WAS IN CARDIOGENIC SHOCK STAGE E. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE. THE FIRST IMPELLA CP BEING PREPPED FOR SUPPORT WAS INADVERTENTLY TURNED ON OUTSIDE THE PATIENT'S BODY WITH THE EASY GUIDE LUMEN IN PLACE. A NEW IMPELLA CP DEVICE WAS OBTAINED, AND THE PATIENT WAS SUCCESSFULLY PLACED ON IMPELLA CP SUPPORT IN THE SETTING OF CARDIOGENIC SHOCK AFTER GOING INTO HEART BLOCK APPROXIMATELY THREE HOURS PRIOR. DUE TO HIGH MEAN ARTERIAL PRESSURE IN THE 140S, PERFORMANCE LEVEL WAS DECREASED TO P-7 AS DOPAMINE AND LEVOPHED WERE TITRATED. DOPAMINE WAS DISCONTINUED. SUCTION ALARMS BEGAN AND PERFORMANCE LEVEL WERE ADJUSTED TO RESOLVE. THE IMPELLA CP PLACEMENT WAS CONFIRMED AND VOLUME EVALUATION WAS CONTINUED. THE PATIENT WAS TRANSFERRED BACK TO THE UNIT IN CRITICAL CONDITION AND SOON AFTER ARRIVING, THE PATIENT'S BLOOD PRESSURE "SIGNIFICANTLY" DROPPED. ECHOCARDIOGRAM CONFIRMED THE IMPELLA CP DEPTH AND RIGHT HEART DYSFUNCTION. DOBUTAMINE WAS ORDERED, AND THE PATIENT REQUIRED AN AMP OF EPINEPHRINE WITH AN EPINEPHRINE DRIP THAT FOLLOWED. THE TEAM DISCUSSED THE PATIENT'S STATUS WITH FAMILY AND NO ESCALATION WAS EXPECTED AT SUCH SAID TIME. THE FOLLOWING DAY IT WAS NOTED THAT THE PATIENT HAD MULTIPLE SUCTION EVENTS OVERNIGHT REQUIRING 500 MILLILITERS OF NORMAL SALINE BOLUS. SUCTION EVENTS RESOLVED AFTER THE NORMAL SALINE BOLUS, AND THE IMPELLA CP SUPPORT LEVEL WAS WEANED FROM SUPPORT LEVEL P-5 TO P-8. EPIPHRINE WAS WEANED DOWN, AND IT WAS NOTED THAT THE SITE WAS CLEAR OF BLEEDING. THE PATIENT CONTINUED ON SUPPORT AND TWO DAYS LATER, HOWEVER, ACCESS SITE OOZING WAS NOW NOTED. BLOOD ADMINISTRATION WAS STARTED WITH ONE UNIT OF RED BLOOD CELLS AND ONE UNIT OF FRESH FROZEN PLASMA. THE IMPELLA CP'S DEPTH WAS ADJUSTED TO 3.66 CENTIMETERS FROM 4.23 CENTIMETERS WITH NO PLANS TO ESCALATE AT SUCH SAID TIME. ADDITIONALLY, IT WAS NOTED THERE WAS CONCERN FOR BLEEDING IN THE LUNGS BASED ON COMPUTED TOMOGRAPHY SCAN EARLIER IN THE DAY. HOWEVER, A BRONCHOSCOPY WAS PERFORMED SHOWING NO ACTIVE BLEEDING. HEPARIN REMAINED ON HOLD. THE PATIENT EXPERIENCED INTERMITTENT EPISODES OF NON PULSATILITY REQUIRING SOME TITRATION OF DRIPS. THE FOLLOWING DAY CARE WAS WITHDRAWN PER THE FAMILY'S WISHES. THE PATIENT WAS CRITICALLY ILL IN STAGE E (EXTREMIS) CARDIOGENIC SHOCK AND NEVER FULLY RECOVERED FROM THE CARDIAC EVENT. HOWEVER, THE IMPELLA CANNOT BE EXCLUDED AS A POSSIBLE CONTRIBUTING FACTOR TO THE OVERALL EVENT OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY ANTICOAGULATION MANAGEMENT SINCE BLEEDING WAS RESOLVED BY WITHHOLDING THE HEPARIN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS, ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THIS IS ONE OF TWO DEVICES RELATED TO THE PATIENT'S IMPLANT EXPERIENCE. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-25928 FOR THE INITIAL IMPELLA CP DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286227&pc=NIK | 2017865-2025-12014 | 2025-01-05 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-31 05:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information | P030054 |  | 1 | Event Description: RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12010. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12011. RELATED MANUFACTURER REPORT NUMBER: 2017865-2025-12012. IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM AFTER THEY COLLAPSED IN THE HOME. THE PATIENT EXPIRED THAT SAME DAY. A REVIEW OF MERLIN.NET FOUND ABORTED SHOCK AND UNDER-SENSING DURING A VENTRICULAR TACHYCARDIA (VT) EPISODE. HOWEVER, THE PHYSICIAN CONCLUDED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO THE PERFORMANCE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA) OR LEFT VENTRICULAR LEADS. Manufacturer Narrative: CORRECTION: UPON REVIEW, THE LEAD SHOULD NOT HAVE BEEN SUBMITTED AS A MEDICAL DEVICE REPORT (MDR) AS THE EVENT DID NOT INDICATE A MALFUNCTION CAUSED A SERIOUS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286251&pc=DSQ | 2916596-2025-00736 | 2024-12-20 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: IT WAS LATER REPORTED THAT PATIENT MADE THE DECISION TO STOP ALL MEDICAL TREATMENT AND DISCHARGE HOME ON HOSPICE. IT WAS SAID THAT DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THAT THE DEVICE WAS OPERATING AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A4- PATIENT WEIGHT NOT YET PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286253&pc=KDQ | 1417592-2025-00050 | 2024-10-01 04:00 | Death | MEDLINE INDUSTRIES LP | 2025-01-31 05:00 | KDQ | LINER,SEMI-RIGID,3000CC,50EA/CS | Use of Device Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED FROM MEDLINE FRANCE REGARDING AN ADVERSE EVENT ASSOCIATED WITH OR530. IT WAS REPORTED BY THE USER FACILITY THAT DURING A MEDICAL EMERGENCY, SUCTION OF THE AIRWAY WAS REQUIRED IN ORDER TO INTUBATE THE PATIENT, HOWEVER THE WALLED SUCTION UNIT WAS NOT SET UP READY TO USE AND WHEN SET UP, IT DID NOT WORK CORRECTLY. Manufacturer Narrative: A NOTIFICATION WAS RECEIVED FROM MEDLINE FRANCE REGARDING AN ADVERSE EVENT ASSOCIATED WITH OR530. IT WAS REPORTED BY THE USER FACILITY THAT DURING A MEDICAL EMERGENCY, SUCTION OF THE AIRWAY WAS REQUIRED IN ORDER TO INTUBATE THE PATIENT, HOWEVER THE WALLED SUCTION UNIT WAS NOT SET UP READY TO USE AND WHEN SET UP, IT DID NOT WORK CORRECTLY. PER THE FACILITY IMMEDIATE ACTION WAS TAKEN TO USE THE CARDIAC ARREST TROLLEY TO MINIMIZE THE DELAY IN PROVIDING EMERGENCY SUCTION. ACCORDING TO THE FACILITY DURING AN INTERNAL INVESTIGATION AFTER THE EVENT, IT WAS FOUND THAT A 3 LITER LINER HAD BEEN PLACED INSIDE THE 1.5 LITER CANNISTER WHICH RESULTED IN THE SUCTION UNIT NOT WORKING CORRECTLY. THE CUSTOMER REPORTED THAT THE PATIENT HAS DIED, BUT THE DEATH IS NOT RELATED TO THE REPORTED EVENT. REPORTEDLY, USE ERROR MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEVICE PROBLEM/ISSUE. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. NO SAMPLE FOR RETURN AND EVALUATION WAS INDICATED. DUE TO THE REPORTED INCIDENT AND PATIENT DEATH, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286317&pc=NVN | 2017865-2025-12019 | 2024-12-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-31 05:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER:2017865-2025-12018. IT WAS REPORTED THAT DURING IMPLANT PROCEDURE OF PACEMAKER AND RIGHT VENTRICULAR (RV) LEAD IMPLANT THE PATIENT DIED. IT WAS NOTED THAT PATIENT HAD OTHER COMORBIDITIES. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DECEASED. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: G2.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286325&pc=LWS | 2649622-2025-03206 | 2025-01-20 05:00 | Death | MPRI | 2025-01-31 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED SUSPECTED INTERMITTENT OVERSENSING NOTED ON RECORDED EPISODES. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DIED THREE DAYS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286341&pc=LPB | 2029046-2025-00334 | 2025-01-07 05:00 | Death | BIOSENSE WEBSTER INC | 2025-01-31 05:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Fistula; Pericarditis | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION, AND THE PATIENT EXPERIENCED ESOPHAGEAL FISTULA THAT REQUIRED SURGICAL INTERVENTION. AS A RESULT, THE PATIENT HAD A STROKE AND DIED. THE PATIENT HAD A RADIOFREQUENCY ABLATION OF ATRIAL FIBRILLATION ON (B)(6) 2024. OVER A MONTH LATER, ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED FOR "DRY PERICARDITIS TABLE". TWO DAYS INTO THE HOSPITALIZATION (B)(6) 2024), THE PATIENT HAD A CEREBROVASCULAR ACCIDENT. ON (B)(6) 2024, THE PATIENT HAD A THORACIC SCAN THAT SHOWED ATRIOESOPHAGEAL FISTULA. THE PATIENT HAD SURGICAL TREATMENT OF THE FISTULA HOWEVER THE PATIENT DIED A FEW WEEKS LATER. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. E1. INITIAL REPORTER PHONE: (B)(6). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 30-JAN-2025. IT WAS REPORTED THAT THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS THE PROCEDURE. THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS 338 SECONDS. THE CAUSE OF DEATH WAS STROKE SECONDARY TO ATRIOESOPHAGEAL FISTULA. DEATH AFTER DIAGNOSIS OF BRAIN DEATH. IN ADDITION, THE CONCOMITANT PRODUCT SECTION WAS UPDATED BASED ON THE INFORMATION RECEIVED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286344&pc=LPB | 2029046-2025-00335 | 2025-01-07 05:00 | Death | BIOSENSE WEBSTER INC | 2025-01-31 05:00 | LPB | NGEN GENERATOR | Adverse Event Without Identified Device or Use Problem | Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION, AND THE PATIENT EXPERIENCED ESOPHAGEAL FISTULA THAT REQUIRED SURGICAL INTERVENTION. AS A RESULT, THE PATIENT HAD A STROKE AND DIED. THE PATIENT HAD A RADIOFREQUENCY ABLATION OF ATRIAL FIBRILLATION ON (B)(6) 2024. OVER A MONTH LATER, ON (B)(6) 2024, THE PATIENT WAS HOSPITALIZED FOR "DRY PERICARDITIS TABLE". TWO DAYS INTO THE HOSPITALIZATION (B)(6) 2024), THE PATIENT HAD A CEREBROVASCULAR ACCIDENT. ON (B)(6) 2024, THE PATIENT HAD A THORACIC SCAN THAT SHOWED ATRIOESOPHAGEAL FISTULA. THE PATIENT HAD SURGICAL TREATMENT OF THE FISTULA HOWEVER THE PATIENT DIED A FEW WEEKS LATER. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 30-JAN-2025. IT WAS REPORTED THAT THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT IS THE PROCEDURE. THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS 338 SECONDS. THE CAUSE OF DEATH WAS STROKE SECONDARY TO ATRIOESOPHAGEAL FISTULA. DEATH AFTER DIAGNOSIS OF BRAIN DEATH. IN ADDITION, THE CONCOMITANT PRODUCT SECTION WAS UPDATED BASED ON THE INFORMATION RECEIVED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN CLARIFICATION ON WHICH SYSTEM WAS USED DURING THIS PROCEDURE OR IF SERVICE WAS NEEDED. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. SINCE THERE IS NO CLARIFICATION, NO PRODUCT INVESTIGATION CAN BE PERFORMED, AND THE CUSTOMER COMPLAINT CANNOT BE CONFIRMED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING CLARIFICATION OF THE SYSTEM USED IN THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. SINCE NO SERIAL NUMBER WAS PROVIDED, NO MANUFACTURER RECORD EVALUATION COULD BE PERFORMED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: NOTE: PER FDA REQUEST, MDR SUBMISSIONS FOR THE NGEN RF GENERATOR ARE TO BE REPORTED WITH PMA DETAILS OF THE CATHETER USED ALONG WITH THE NGEN RF GENERATOR. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. E1. INITIAL REPORTER PHONE: (B)(6). MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN CLARIFICATION ON WHICH SYSTEM WAS USED DURING THIS PROCEDURE. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. SINCE THERE IS NO CLARIFICATION, NO PRODUCT INVESTIGATION CAN BE PERFORMED, AND THE CUSTOMER COMPLAINT CANNOT BE CONFIRMED. SINCE NO SERIAL NUMBER WAS PROVIDED, NO MANUFACTURER RECORD EVALUATION COULD BE PERFORMED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286352&pc=DXY | 2017865-2025-12018 | 2024-12-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-01-31 05:00 | DXY | ENDURITY PACEMAKER, SR, MRI RESTRICTED | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: RELATED MANUFACTURER REFERENCE NUMBER: 2017865-2025-12019. IT WAS REPORTED THAT DURING IMPLANT PROCEDURE OF PACEMAKER AND RIGHT VENTRICULAR (RV) LEAD IMPLANT THE PATIENT DIED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, THE DEVICE WAS WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286401&pc=MNS | 2518422-2025-101420 | 2025-01-23 05:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286460&pc=BZD | 2518422-2025-010075 | 2024-06-13 04:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE CUSTOMER ALLEGES THE PATIENT HAS PASSED AWAY, DID NOT RECEIVE REPLACEMENT DEVICE. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER¿S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286621&pc=DYB | 2015691-2025-00652 | 2024-07-29 04:00 | Death | EDWARDS LIFESCIENCES | 2025-01-31 05:00 | DYB | EDWARDS ESHEATH+ INTRODUCER | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, DURING A TRANSFEMORAL TAVR PROCEDURE WITH A 23 MM SAPIEN 3 ULTRA RESILIA VALVE, THERE WAS DAMAGE FROM THE RIGHT COMMON ILIAC ARTERY (CIA) TO THE ABDOMINAL AORTA OBSERVED, AND A STANFORD TYPE A AORTIC DISSECTION. THERE WAS ALSO RETROPERITONEAL BLEEDING. AN EMERGENCY LAPAROTOMY ARTIFICIAL BLOOD VESSEL REPLACEMENT WAS PERFORMED. AFTER THE PROCEDURE, THE PATIENT'S HYPOTENSION PERSISTED, AND A NON-OCCLUSIVE MESENTERIC INFARCTION OCCURRED. THE PATIENT DIED 3 DAYS AFTER THE TAVR PROCEDURE. Manufacturer Narrative: HEALTH EFFECT - CLINICAL CODE: AORTIC DISSECTION. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE AORTIC DISSECTION/CARDIOVASCULAR INJURIES CANNOT BE DETERMINED, IT IS POSSIBLE THE MECHANISMS OF THE TAVR PROCEDURE DESCRIBED ABOVE OR PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286765&pc=MIH | 2247858-2025-00039 | 2024-12-13 05:00 | Death | BOLTON MEDICAL, INC. | 2025-01-31 05:00 | MIH | CUSTOM-MADE TREO ABDOMINAL STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Patient Device Interaction Problem | Cardiac Arrest; Hemorrhage/Bleeding; Ischemic Heart Disease; Diminished Pulse Pressure | P190015 |  | 1 | Event Description: (B)(6). THE PATIENT WOKE UP FROM THE PROCEDURE, AND LATER HE THEN DROPPED HIS BP, WHICH, UNFORTUNATELY, DIDN'T RECOVER. ADDITIONAL INFORMATION RECEIVED FROM THE SITE ON 01/28/2025: THIS PATIENT UNDERWENT A FENESTRATED ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR ON (B)(6) 2024. THE PROCEDURE WAS MORE COMPLICATED THAN USUAL DUE TO DIFFICULTY IN ACHIEVING STABLE CANNULATION OF THE VISCERAL VESSELS. THIS ADDED TO THE DURATION OF THE PROCEDURE. BLOOD LOSS THROUGH THE DELIVERY SYSTEM OF THE ENDOVASCULAR GRAFT WHILST SLOW, WAS CUMULATIVELY SIGNIFICANT AT AROUND 1500ML. THE ANEURYSM REPAIR WAS HOWEVER COMPLETED SATISFACTORILY. THE IMMEDIATE POST-OPERATIVE RECOVERY OF THE PATIENT WAS UNEVENTFUL BUT SEVERAL HOURS AFTER THE PROCEDURE THEY BECAME UNWELL AND, WHILST UNCONFIRMED, THE CLINICAL ASSESSMENT WAS THAT THIS WAS THE RESULT OF AN ACUTE CORONARY EVENT. THEY SUBSEQUENTLY DEVELOPED SIGNS AND SYMPTOMS OF SPINAL CORD ISCHAEMIA AND BECAME INCREASINGLY UNWELL. THE PATIENT DIED APPROXIMATELY 24 HOURS AFTER THE ANEURYSM REPAIR. FOLLOWING DISCUSSION WITH THE HOSPITAL'S (B)(6) MEDICAL EXAMINER, THE DEATH WAS REPORTED TO (B)(6) CORONER WHO HAS OPENED AN INQUEST, WHICH IS AWAITED."Event Description: "PAT 5111-B1 + PAT 5209-L1 (28-LC15W10012S1949, LOT # 2408190329) THE PATIENT WOKE UP FROM THE PROCEDURE, AND LATER HE THEN DROPPED HIS BP, WHICH, UNFORTUNATELY, DIDN'T RECOVER. ADDITIONAL INFORMATION RECEIVED FROM THE SITE ON 01/28/2025: "THIS PATIENT UNDERWENT A FENESTRATED ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR ON (B)(6) 2024. THE PROCEDURE WAS MORE COMPLICATED THAN USUAL DUE TO DIFFICULTY IN ACHIEVING STABLE CANNULATION OF THE VISCERAL VESSELS. THIS ADDED TO THE DURATION OF THE PROCEDURE. BLOOD LOSS THROUGH THE DELIVERY SYSTEM OF THE ENDOVASCULAR GRAFT WHILST SLOW, WAS CUMULATIVELY SIGNIFICANT AT AROUND 1500ML. THE ANEURYSM REPAIR WAS HOWEVER COMPLETED SATISFACTORILY. THE IMMEDIATE POST-OPERATIVE RECOVERY OF THE PATIENT WAS UNEVENTFUL BUT SEVERAL HOURS AFTER THE PROCEDURE THEY BECAME UNWELL AND, WHILST UNCONFIRMED, THE CLINICAL ASSESSMENT WAS THAT THIS WAS THE RESULT OF AN ACUTE CORONARY EVENT. THEY SUBSEQUENTLY DEVELOPED SIGNS AND SYMPTOMS OF SPINAL CORD ISCHAEMIA AND BECAME INCREASINGLY UNWELL. THE PATIENT DIED APPROXIMATELY 24 HOURS AFTER THE ANEURYSM REPAIR. FOLLOWING DISCUSSION WITH THE UNIVERSITY HOSPITALS OF LEICESTER NHS TRUST MEDICAL EXAMINER THE DEATH WAS REPORTED TO HM CORONER WHO HAS OPENED AN INQUEST, WHICH IS AWAITED." PATIENT OUTCOME: "PATIENT DEATH." Manufacturer Narrative: THE PATIENT DEATH WAS REPORTED TO (B)(6) ON 12/17/2024 WITH NO ADDITIONAL INFORMATION. EMAILS WERE SENT TO THE CASE REP ON 12/17/2024, 12/19/2024, 12/20/2024, 12/26/2024, 01/06/2025, AND 01/10/2025 TO TRY TO OBTAIN ADDITIONAL INFORMATION REGARDING THE PATIENT DEATH. AN EMAIL FROM THE CASE REP DATED 01/10/2025 STATED THAT THE SURGEON IS UNABLE TO PROVIDE ANY DETAILS UNTIL HE RECEIVES GUIDANCE FROM THE LEGAL DEPARTMENT AT HIS TRUST. ON 01/28/2025, PERMISSION WAS GIVEN FROM THE HEAD OF PRIVACY AND TRUST LEGAL DEPARTMENT TO SEND THE EVENT DETAILS. THEREFORE, THIS REPORT IS SUBMITTED AFTER THE 30-DAY TIME-FRAME. BOLTON MEDICAL IS VOLUNTARILY REPORTING AN EVENT RELATED TO A TREO CUSTOM-MADE DEVICE. THE CUSTOM-MADE TREO DEVICES ARE NOT MARKETED IN THE US; HOWEVER, THEY ARE SIMILAR TO THE TREO ABDOMINAL STENT GRAFT DELIVERY SYSTEM APPROVED FOR SALE IN THE US (P190015). THE EVENT OCCURRED IN THE UNITED KINGDOM. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: BOLTON MEDICAL IS VOLUNTARILY REPORTING AN EVENT RELATED TO A TREO CUSTOM-MADE DEVICE. THE CUSTOM MADE TREO DEVICES ARE NOT MARKETED IN THE US, HOWEVER THEY ARE SIMILAR TO THE TREO ABDOMINAL STENT GRAFT DELIVERY SYSTEM APPROVED FOR SALE IN THE US (P190015). THE EVENT OCCURRED IN UNITED KINGDOM. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286813&pc=NKM | 2135147-2025-00492 | 2024-12-31 05:00 | Death | ABBOTT MEDICAL | 2025-01-31 05:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Incomplete Coaptation | Mitral Valve Insufficiency/ Regurgitation | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4. ONE CLIP WAS IMPLANTED, REDUCING MR TO A GRADE OF 1. ON (B)(6) 2024, THE PATIENT RETURNED TO THE HOSPITAL. IMAGING SHOWED THE IMPLANTED CLIP HAD DETACHED FROM ONE LEAFLET AND REMAINED ATTACHED TO THE OTHER LEAFLET (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA). ON (B)(6) 2025, THE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A NTW MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4. IT WAS NOTED THAT THE TISSUE WAS VERY FRIABLE AND HARD TO GRASP. THE LEAFLET FELL OFF THE CLIP DURING THE FIRST TWO GRASPS PRIOR TO/DURING CLOSING. THE NTW MITRACLIP WAS IMPLANTED, REDUCING MR TO A GRADE OF 1. IT WAS NOTED THAT ON TRANSTHORACIC ECHOCARDIOGRAM (TTE) ON (B)(6) 2024, SINGLE LEAFLET DEVICE ATTACHMENT/SLDA WAS UNNOTICED BUT PRESENT. THE PATIENT WAS DISCHARGED HOME. ON (B)(6) 2024, THE PATIENT RETURNED TO THE HOSPITAL AFTER A FALL AT HOME. THE PATIENT HAD NASAL BONE FRACTURE AND ABDOMINAL PAIN. THE PATIENT WAS ADMITTED WITH RAPID OXYGEN DESATURATION LED TO EMERGENT INTUBATION AND AGGRESSIVE THERAPY TO STABILIZE ACID-BASE STATUS. THE PATIENT HAD SEVERE METABOLIC ACIDOSIS, LACTATE, CONGESTIVE HEART FAILURE, RIGHT VENTRICULAR DYSFUNCTION, PULMONARY HYPERTENSION, MITRAL VALVE DYSFUNCTION, AND ACUTE HYPOXEMIC RESPIRATORY FAILURE. TTE ON (B)(6)2024 SHOWED THE IMPLANTED CLIP HAD DETACHED FROM ONE LEAFLET AND REMAINED ATTACHED TO THE OTHER LEAFLET (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA). THE DETACHMENT OCCURRED ON THE POSTERIOR LEAFLET. DISCUSSION WITH THE FAMILY OCCURRED AND IT WAS CHOSEN THAT THE PATIENT WOULD NOT HAVE CARDIOPULMONARY RESUSCITATION. ON (B)(6) 2025, THE PATIENT EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS FROM THE LOT. ALL AVAILABLE INFORMATION WAS INVESTIGATED AND BASED ON THE INFORMATION PROVIDED AND PER THE PHYSICIAN, THE REPORTED SINGLE LEAFLET DEVICE ATTACHMENT AND DIFFICULT OR DELAYED POSITIONING ASSOCIATED WITH LEAFLET GRASPING ARE DUE TO THIN, FRIABLE AND REDUNDANT POSTERIOR LEAFLET (PATIENT CONDITIONS). A CAUSE FOR THE REPORTED MITRAL VALVE INSUFFICIENCY/REGURGITATION, DYSPNEA, HEART FAILURE/CONGESTIVE HEART FAILURE, HYPERTENSION, HYPOXIA AND DEATH CANNOT BE DETERMINED. HOWEVER, MITRAL REGURGITATION, DEATH, DYSPNEA, HEART FAILURE AND HYPERTENSION ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. UNEXPECTED MEDICAL INTERVENTION AND HOSPITALIZATION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286832&pc=DSQ | 2916596-2025-00445 | 2023-09-26 04:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2023. NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE, AND THE PATIENT¿S EXPIRATION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT NO FURTHER INFORMATION CAN BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286876&pc=DSQ | 2916596-2025-00494 | 2025-01-09 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION WAS PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY CUSTOMER. THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286898&pc=CBK | 3013876692-2025-00006 | 2025-01-14 05:00 | Death | MAQUET CRITICAL CARE AB | 2025-01-31 05:00 | CBK | SERVO-S BASE UNIT | Connection Problem; No Flow | Cardiac Arrest; Respiratory Arrest | K123149 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTILATOR WAS NOT CONNECTED TO O2 SUPPLY WHILE ON PATIENT. THE PATIENT DIED THE DAY AFTER THE EVENT. MANUFACTURER'S REF.#: (B)(4).Event Description: MANUFACTURER'S REF.#: (B)(4). Manufacturer Narrative: THE UDI INFORMATION IS NOT AVAILABLE SINCE THE DEVICE WAS MANUFACTURED BEFORE 2015. Manufacturer Narrative: THE VENTILATOR WAS INVESTIGATED BY OUR FIELD SERVICE ENGINEER. THE VENTILATOR PASSED ALL FUNCTIONAL AND SAFETY TESTS AND WAS CLEARED FOR CLINICAL USE. EVALUATION OF THE RECEIVED LOGS REVEALED THAT THE USER STARTED THE VENTILATION AT 12:55:41 (B)(6) 2025 AND AT 12:55:48 (B)(6) 2025 A MEDIUM PRIORITY ALARM FOR LOW O2 SUPPLY PRESSURE WAS ACTIVATED (AUDIBLE). PROBABLY THE HIGH PRIORITY LOW EXPIRATORY VOLUME ALARM WAS SHOWN IN THE ALARM/MESSAGE AREA AS THIS ALARM HAS A HIGHER PRIORITY THAN THE O2 SUPPLY PRESSURE ALARM. FROM (B)(6) 2025 12:55:48 TO 12:55:54 THE USER TRIED TO ADJUST THE SET O2 CONCENTRATION. AFTER A USER ADJUSTMENT OF THE O2 CONCENTRATION, THE O2 CONCENTRATION ALARMS ARE DELAYED BY APPROXIMATELY 50 S IN ORDER TO PREVENT NUISANCE ALARM DURING STABILIZATION OF THE O2 CONCENTRATION IN THE SYSTEM. THE HIGH PRIORITY LOW O2 CONCENTRATION ALARM IS ACTIVATED (AUDIBLE AND VISIBLE) AT 12:56:40 (B)(6) 2025 AND THE USER SILENCES THE ALARMS AFTER THREE SECONDS AT 12:56:43. THE TECHNICAL LOG DOES NOT CONTAIN ANY TECHNICAL ERROR CODES TO INDICATE THAT THERE WAS A VENTILATOR MALFUNCTION. THE VENTILATOR WAS PASSED PRE-USE CHECK SUCCESSFULLY PRIOR TO VENTILATION START AT DATE OF EVENT. ACCORDING TO USER FACILITY, THE O2 SUPPLY WAS NOT CONNECTED FROM THE BEGINNING. THE VENTILATOR DETECTED THE LACK OF O2 SUPPLY PRESSURE AND ALARMED FOR IT ACCORDING TO SPECIFICATION. THERE IS NO ALLEGATION FROM THE USER THAT A VENTILATOR MALFUNCTION IN SOME WAY CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. THE CONCLUSION IS THAT THE VENTILATOR HAS BEEN WORKING ACCORDING TO SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286911&pc=BZD | 2518422-2025-101596 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION, AN ERROR 53 WAS DISCOVERED IN THE DEVICE LOGS. THE PCA WAS REPLACED. DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS RETURNED TO STOCK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21286919&pc=MKJ | 1220908-2025-00418 | 2025-01-20 05:00 | Death | ZOLL MEDICAL CORPORATION | 2025-01-31 05:00 | MKJ | R SERIES | Defibrillation/Stimulation Problem | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO MONITOR A PEDIATRIC PATIENT (AGE & GENDER UNKNOWN), THE DEVICE DISPLAYED "SYSTEM FAULT CODE 36" AND "SYSTEM FAULT CODE 37" MESSAGES. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL UNITED KINGDOM FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE ACTIVITY LOGS. HOWEVER, THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING INCLUDING STRESS TESTING USING THE CUSTOMER'S RETURNED POWER CORD AND FUNCTIONAL TESTING INCLUDING DEFIB CYCLING SHOCKS WITHOUT DUPLICATING THE REPORT. THE CUSTOMER'S MULTIFUNCTION CABLE WAS NOT RETURNED FOR EVALUATION. THE ANALOG BOARD WAS REPLACED AS A PRECAUTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287005&pc=DSQ | 2916596-2025-00740 | 2024-10-07 04:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Dyspnea; Renal Impairment; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. IT WAS LATER REPORTED THAT PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT ON (B)(6) 2024 AFTER A FALL WITH CONCERN FOR SCAPULAR FRACTURE. THE PATIENT WAS LAST SEEN IN CLINIC (B)(6) 2024 AND WAS TRANSITIONED TO HOSPICE CARE AFTER THAT VISIT. PATIENT'S SIGNIFICANT OTHER CALLED ON (B)(6) 2024 AND STATED THAT THE PATIENT PASSED IN HOSPICE CARE. IT WAS ALSO SAID THAT OUTCOME WAS NOT DEVICE RELATED, AND THE DEVICE OPERATED AS INTENDED.Event Description: IT WAS LATER REPORTED THAT THE FALL WAS NOT DEVICE OR THERAPY RELATED. THE PATIENT WAS ADMITTED TO HOSPICE CARE DUE TO WORSENING HEART FAILURE AND CHRONIC KIDNEY DISEASE WITH AN INCREASED SHORTNESS OF BREATH AND OXYGEN NEEDS. THE DECISION TO ADMITTED PATIENT TO HOSPICE CARE WAS ALSO SAID NOT BE DEVICE OR THERAPY RELATED.Event Description: PER PATIENT REQUEST, THE PATIENT WAS ADMITTED TO HOSPICE CARE DUE TO WORSENING HEART FAILURE AND CHRONIC KIDNEY DISEASE WITH AN INCREASED SHORTNESS OF BREATH AND OXYGEN NEEDS. THE DECISION TO ADMITTED PATIENT TO HOSPICE CARE WAS ALSO SAID NOT BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: SECTION B3: EVENT DATE - CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE CKD; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: A4 - PATIENT WEIGHT UPDATED. B5 NARRATIVE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTION TO B2- DEATH DATE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287032&pc=DSQ | 2916596-2025-00742 | 2025-01-15 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO FAILURE TO THRIVE. PER VENTRICULAR ASSIST DEVICE TEAM, THE PATIENT PASSED AWAY AT HOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE RELATED.Event Description: IT WAS FURTHER COMMUNICATED ON 10FEB2025 THAT THE CAUSE OF DEATH WAS DUE TO A SELF-INFLICTED GUNSHOT WOUND/SUICIDE. THE PUMP WAS STATED TO BE FUNCTIONING PROPERLY UNTIL THE PATIENT'S DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B: DATE OF DEATH CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: NO DEVICE-RELATED ISSUES WITH THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), WERE REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE-RELATED ISSUES WERE REPORTED, SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287036&pc=DSI | 9610816-2025-000084 | 2024-12-29 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-01-31 05:00 | DSI | INTELLIVUE MP70 | Device Alarm System | Insufficient Information | K151681 |  | 1 | Event Description: IT WAS REPORTED THE MONITOR DID NOT ALARM FOR ST CHANGES WHEN THE PATIENT WAS HAVING A STEMI (ST-ELEVATION MYOCARDIAL INFARCTION). WHEN NURSING NOTICED CHANGES IN ECG, A 12-LEAD ECG WAS PERFORMED TO CONFIRM THE STEMI, BUT THE PATIENT SUBSEQUENTLY PASSED AWAY. IT WAS INDICATED THE USERS NOTICED ST ELEVATION ON THE MONITOR AND, AFTER REVIEWING PATIENT DATA, DETERMINED THE PATIENT HAD BEEN HAVING ST ELEVATION FOR APPROXIMATELY 11 HOURS WITH NO ALARMS, THOUGH ST WAS TURNED ON. DURING THIS TIME, OTHER ALARMS WERE BEING GENERATED FOR LOW SPO2, LOW BLOOD PRESSURE, AND HEART RATE. Manufacturer Narrative: UPDATED NARRATIVE: A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED. THE FSE USED A SIMULATOR TO TEST THE FUNCTIONALITY OF ST AND THE DEVICE ALARMED AS EXPECTED. THE SPEAKER TESTING PASSED, AS WELL. THE CLINICAL CONFIGURATION AND CLINICAL AUDIT LOGS WERE OBTAINED BY THE FSE; HOWEVER, THE PHYSIO DATA WAS NO LONGER AVAILABLE SINCE OVER SEVEN DAYS ELAPSED. THE LOG DATA WAS REVIEWED BY THE FSE, THE CLINICAL SOLUTIONS DELIVERY CONSULTANT, A PHILIPS PRODUCT SUPPORT ENGINEER, AND PHILIPS SENIOR CLINICAL PRODUCT SPECIALIST. THE CLINICAL AUDIT LOG PROVIDES A CHRONOLOGICAL RECORD OF THE ALARMS AND ACTIONS. THE CLINICAL AUDIT SHOWED THAT THE ST ANALYSIS AND STE ANALYSIS WERE SWITCHED OFF ON (B)(6) 2024 AT 01:28:43 FOR BED C02-01. THERE IS NO INDICATION THAT THE ST ANALYSIS WERE SWITCHED BACK ON. PER THE INTELLIVUE MP20/30, MP40/50, MP60/70/80/90 INSTRUCTIONS FOR USER FOR RELEASE M WITH SOFTWARE REVISION M.0X.(B)(6) (B)(4), ¿THE INFORMATION PRINTED ON THE RECORDING STRIP INCLUDES THE PATIENT NAME AND MRN [MEDICAL RECORD NUMBER], BED NUMBER, DATE AND TIME OF RECORDING, RECORDING SPEED, AND (EXCEPT WHEN THE INFORMATION IS PRINTED ABOVE THE WAVES) RECORDING CODE. ACTIVE ALARM AND INOP MESSAGES AS WELL AS NUMERICS FOR ALL CURRENTLY MONITORED MEASUREMENTS ARE ALSO PRINTED.¿ RECORDING STRIPS WERE PROVIDED AND REVIEWED AND NO ST NUMERICS ARE DISPLAYED. NO ST ALARMS WERE DISPLAYED IN THE CLINICAL AUDIT LOG BECAUSE ST ANALYSIS AND STE ANALYSIS WERE SWITCHED OFF. THE CLINICAL AUDIT LOG ALSO SHOWS BOTH ST ALARMS AND STE ALARMS WERE SWITCHED ON THREE TIMES ON (B)(6) 2024. WHEN A MEASUREMENT IS OFF, THE MONITOR STOPS DATA ACQUISITION AND ALARMING FOR THE MEASUREMENT. HENCE, NO ST OR STE ALARMS WERE GENERATED BY THE MONITOR. THE REPORTED PROBLEM WAS NOT CONFIRMED. INFORMATION IS BEING PROVIDED TO THE CUSTOMER TO RESOLVE THE ISSUE AND THE DEVICE REMAINS IN USE. Manufacturer Narrative: THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED. THE FSE USED A SIMULATOR TO TEST THE FUNCTIONALITY OF ST AND THE DEVICE ALARMED AS EXPECTED. THE SPEAKER TESTING PASSED, AS WELL. THE CLINICAL CONFIGURATION AND CLINICAL AUDIT LOGS WERE OBTAINED BY THE FSE, HOWEVER THE PHYSIO DATA WAS NO LONGER AVAILABLE SINCE OVER SEVEN DAYS ELAPSED. THE LOG DATA WAS REVIEWED BY THE FSE AND THE CSDC AND IT WAS VERIFIED THAT THERE WAS NO ST ALARMING FOUND IN THE LOGS. THE CSDC EXPLAINED TO THE CUSTOMER THAT ST ALARMS ARE ALWAYS YELLOW ALARMS AND REQUIRE 60 SECONDS OF ST ELEVATION BEFORE ALARMING. THE CSDC FURTHER EXPLAINED TO THE CUSTOMER THAT LEAD PLACEMENT OF THE V LEAD MAY HAVE BEEN INACCURATE AND, BECAUSE OF THIS, THE ST MAY NOT HAVE BEEN AS HIGH HAS +2.0. DUE TO THE LIMITED WAVE DATA, LESS THAN 60 FULL SECONDS, FURTHER ANALYSIS IS NOT POSSIBLE AND THE ISSUE CANNOT BE CONFIRMED. THE CUSTOMER DECLINED TO QUARANTINE THE DEVICE AND FOUND THE TESTING AND EXPLANATION SATISFACTORY. THE DEVICE REMAINS IN USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287196&pc=MVK | 3008642652-2025-01205 | 2024-12-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-31 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 IN THE HOSPITAL WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:14:21 ON (B)(6) 2024. AT 22:13:15, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 70 BPM WITH CPR/MOTION ARTIFACT. AT 22:13:52, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS SINUS RHYTHM @ 70 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH CPR/MOTION ARTIFACT. AT 22:14:27, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT TIME OF TREATMENT WAS SINUS RHYTHM @ 70 BPM WITH PVCS AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM IS NOT AVAILABLE. THE DEVICE WAS SHUT DOWN AT 23:15:20 ON (B)(6) 2024. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287228&pc=DSQ | 2916596-2025-00788 | 2025-01-15 05:00 | Death | THORATEC CORPORATION | 2025-01-31 05:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cardiogenic Shock; Respiratory Failure; Heart Failure/Congestive Heart Failure; Renal Impairment; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OUTCOME, THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK SECONDARY TO RIGHT VENTRICULAR FAILURE, RESPIRATORY FAILURE, ACUTE KIDNEY INJURY. WHETHER THE OUTCOME WAS DEVICE RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE PUMP WAS GOING TO BE EXPLANTED.Event Description: IT WAS LATER REPORTED THAT THE INITIALLY REPORTED RIGHT HEART FAILURE, RESPIRATORY FAILURE, AND ACUTE KIDNEY INJURY EXISTED PRE-LVAD IMPLANT. THE PATIENT WAS SAID TO HAVE EXPERIENCED SYMPTOMS OF LETHARGY AND WAS TREATED WITH A RIGHT VENTRICULAR ASSISTS DEVICE, A TRACHEOSTOMY, AND CONTINUOUS RENAL REPLACEMENT THERAPY BEFORE THEIR DEATH. IT WAS ALSO SAID THAT DEATH WAS NOT CONSIDERED DEVICE RELATED AND THAT THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THERE WERE NO DEVICE RELATED ISSUES REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. ALTHOUGH NO DEVICE RELATED ISSUES WERE REPORTED BY THE ACCOUNT, THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. ALTHOUGH NO DEVICE RELATED ISSUES WERE REPORTED BY THE ACCOUNT, SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287507&pc=FRN | 1314492-2025-00104 | 2024-11-22 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-01-31 05:00 | FRN | SPECTRUM IQ | Complete Loss of Power | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K222048 |  | 1 | Event Description: IT WAS REPORTED FROM THE HOSPITAL¿S BIOMED DEPARTMENT, A SPECTRUM IQ INFUSION PUMP DID NOT HAVE A BATTERY CONNECTED DURING AN INFUSION AND THE PATIENT SUBSEQUENTLY DIED. DURING AN UNSPECIFIED INFUSION, THE SPECTRUM IQ PUMP WAS PLUGGED INTO THE WALL SOCKET. AT AN UNKNOWN POINT DURING THE INFUSION, THE PATIENT HAD TO BE ¿MOVED¿, AND THE PUMP WAS UNPLUGGED FROM THE WALL SOCKET. AT THIS TIME, IT WAS NOTED THERE WAS NO BATTERY PRESENT INSIDE THE SPECTRUM PUMP. THE CUSTOMER COULD NOT CLARIFY IF THE PUMP ALARMED THE ¿INSTALL BATTERY" ALARM. THIS ALARM WILL TRIGGER WHEN THE PUMP IS TURNED ON AND WILL CONTINUE TO ALARM UNTIL THE USER ACKNOWLEDGES THE ALERT. IF THE USER ACKNOWLEDGES THIS ALERT AND THERE IS STILL NO BATTERY INSTALLED; THE ONSCREEN ICON WILL FLASH CONTINUOUSLY UNTIL A BATTERY IS INSTALLED. AT SOME POINT IN TIME, THE NURSE NOTED THE PUMP WAS NOT WORKING AND ¿FOUND¿ A BATTERY. AFTER AN UNKNOWN AMOUNT OF TIME THE PATIENT DIED. THE REPORTER WAS NOT INFORMED HOW LONG AFTER THE PUMP WAS UNPLUGGED BEFORE THE PATIENT DIED. IT WAS NOT REPORTED WHAT SOLUTION WAS BEING INFUSED, IF ANY LIFE SUSTAINING TREATMENT WAS PROVIDED, THE CAUSE OF DEATH, OR IF AN AUTOPSY WAS PERFORMED. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: ADDITIONAL INFORMATION H2, H3, H6 AND H11: THE DEVICE WAS RECEIVED FOR EVALUATION. FUNCTIONAL TESTING WAS PERFORMED AND DID NOT IDENTIFY ANY ISSUES RELATED TO THE CUSTOMER REPORTED CONDITION. THE REPORTED PUMP LOST POWER AND STOPPED MID-INFUSION (PUMP DID NOT HAVE A BATTERY CONNECTED AND WAS UNPLUGGED PROBLEM FROM THE WALL SOCKET) WAS NOT REPRODUCED DURING EVALUATION. EVALUATION RAN AN INFUSION AT 33 ML/HR FOR 15 MINUTES ON THE AC POWER ADAPTER RECEIVED FROM CUSTOMER. DURING BOTH TEST INFUSIONS MANUAL AGGRAVATION OF THE POWER SUPPLIES WAS APPLIED AND THE DEVICE OPERATED AS INDENTED WITH NO SHUT DOWNS (SHUTDOWNS) OR INTERRUPTIONS OF THE TEST INFUSIONS. A SERVICE HISTORY REVIEW WAS PERFORMED AND REVEALED THAT THE DEVICE HAS NO PREVIOUS SERVICE EVENTS; THEREFORE, SERVICING DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. THE EVENT HISTORY LOG (EHL) REVIEW WAS PERFORMED AND REVEALED MULTIPLE EVENTS OF "IMPROPER SHUTDOWN!" AN "IMPROPER SHUTDOWN!" MESSAGE INDICATES THAT ALL POWER WAS LOST FROM THE PUMP. HOWEVER, THE LOG CANNOT BE USED TO DETERMINE IF THE POWER WAS MANUALLY DISCONNECTED OR SHUTDOWN WITHOUT THE INPUT OF THE USER. THE PUMP PROMPTED THE USER MULTIPLES TIMES THE BATTERY WAS NOT INSTALLED. ¿INSTALL BATTERY¿ PROMPTS WERE CONFIRMED IN THE DEVICE HISTORY. THE REPORTED CONDITION WAS NOT VERIFIED DURING DEVICE EVALUATION. BASED ON ALL AVAILABLE INFORMATION, THE DEVICE FUNCTIONED AS INTENDED. TO CONCLUDE, USER ERROR (USER DID NOT INSTALL THE BATTERY) OF THE SPECTRUM IQ INFUSION PUMP HAS POSSIBLY CAUSED OR CONTRIBUTED TO THE PATIENT'S OUTCOME. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DEVICE HAS BEEN RECEIVED, AND THE EVALUATION IS IN PROGRESS. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287516&pc=NIK | 9614453-2025-00505 | 2024-12-02 05:00 | Death | MEDTRONIC EUROPE SARL | 2025-01-31 05:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT TWELVE DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND RIGHT VENTRICULAR (RV) LEAD CHANGEOUT PROCEDURE THE PATIENT DIED. IT WAS NOTED THAT THE PATIENT DEATH WAS NOT CARDIAC OR DEVICE RELATED HOWEVER THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: 1056T LEAD; IMPLANTED: (B)(6) 2006. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287521&pc=LWS | 2649622-2025-03209 | 2024-12-02 05:00 | Death | MPRI | 2025-01-31 05:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT TWELVE DAYS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) AND RIGHT VENTRICULAR (RV) LEAD CHANGEOUT PROCEDURE THE PATIENT DIED. IT WAS NOTED THAT THE PATIENT DEATH WAS NOT CARDIAC OR DEVICE RELATED HOWEVER THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: 1056T LEAD, IMPLANTED (B)(6) 2006. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21287672&pc=CBK | 2518422-2025-101598 | 2024-11-27 05:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM A DEALER REGARDING INFORMATION THEY HAD RECEIVED FROM A FATHER OF THE PATIENT WHO HAD PASSED AWAY. THE DEALER CONTACTED THE FATHER OF THE PATIENT TO ASK THE VENTILATOR'S DATA, BUT "THE FATHER SAID THAT HE WANTED TO CONFIRM IF THERE WERE SOME MALFUNCTIONS SINCE THE PATIENT'S CONDITION WAS TEMPORARILY STABLE, AND IT SEEMED THAT HIS TONE WAS CALM." THE DEALER CONTACTED THE CLINIC FOR INFORMATION, BUT THERE WAS NO INFORMATION ON THE CAUSE OF DEATH. A SALES REPRESENTATIVE RECEIVED ADDITIONAL INFORMATION ON THIS ISSUE. ACCORDING TO THE SALE REPRESENTATIVE, THE DATE OF THE PATIENT PASSING AWAY WAS CONFIRMED. THE SALES REPRESENTATIVE STATED THAT "A VISITING NURSE ARRIVED BEFORE THE DOCTOR, AND THE MOTHER WHO WAS AT THE SCENE REPORTED NO COMPLAINTS OF DEVICE MALFUNCTION (THE PATIENT'S PARENTS DIVORCED THREE DAYS BEFORE HIS DEATH. THE COMPLAINT WAS RECEIVED FROM THE FATHER, WHO WAS NOT PRESENT AT THE TIME AND SUSPECTS THE FAMILY MIGHT HAVE TURNED OFF THE DEVICE.)". THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION FROM A DEALER REGARDING INFORMATION THEY HAD RECEIVED FROM A FATHER OF THE PATIENT WHO HAD PASSED AWAY. THE DEALER CONTACTED THE FATHER OF THE PATIENT TO ASK THE VENTILATOR'S DATA, BUT "THE FATHER SAID THAT HE WANTED TO CONFIRM IF THERE WERE SOME MALFUNCTIONS SINCE THE PATIENT'S CONDITION WAS TEMPORARILY STABLE, AND IT SEEMED THAT HIS TONE WAS CALM." THE DEALER CONTACTED THE CLINIC FOR INFORMATION, BUT THERE WAS NO INFORMATION ON THE CAUSE OF DEATH. A SALES REPRESENTATIVE RECEIVED ADDITIONAL INFORMATION ON THIS ISSUE. ACCORDING TO THE SALE REPRESENTATIVE, THE DATE OF THE PATIENT PASSING AWAY WAS CONFIRMED. THE SALES REPRESENTATIVE STATED THAT "A VISITING NURSE ARRIVED BEFORE THE DOCTOR, AND THE MOTHER WHO WAS AT THE SCENE REPORTED NO COMPLAINTS OF DEVICE MALFUNCTION (THE PATIENT'S PARENTS DIVORCED THREE DAYS BEFORE HIS DEATH. THE COMPLAINT WAS RECEIVED FROM THE FATHER, WHO WAS NOT PRESENT AT THE TIME AND SUSPECTS THE FAMILY MIGHT HAVE TURNED OFF THE DEVICE.)" THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER¿S COMPLAINT WAS NOT CONFIRMED DURING AN OPERATIONAL CHECK OF THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN EVALUATED AND OBSERVED NO ERRORS IN THE SYSTEM LOG. THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED A FUNCTIONAL TEST ON THIS DEVICE, BUT WAS NOT ABLE CONFIRM THE ISSUE ON THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN DETERMINED THE DEVICE WAS OPERATING AS NORMAL. THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED THE FINAL AND RUN-IN TESTS, ALONG WITH THE BATTERY AND VALVE CHECKS FOR THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN DETERMINED THE DEVICE WAS OPERATIONAL AND IT WAS CLEANED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21288821&pc=OZD | 1220648-2025-25946 | 2025-01-19 05:00 | Death | ABIOMED, INC. | 2025-01-31 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | No Pressure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING PURGE CASSETTE REPLACEMENT, THE DISC HOLDER ON THE AUTOMATED IMPELLA CONTROLLER (AIC) WAS INADVERTENTLY TORN OFF. THE AIC WAS REPLACED TO RESOLVE THE DAMAGE. THERE WAS NO PATIENT HARM AND SUPPORT CONTINUED WITH THE IMPLANTED PUMP.Event Description: AFTER 69.48 HOURS OF SUPPORT THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE PLACE PATIENT ON PALLIATIVE CARE AND THE PATIENT EXPIRED. THE EXACT DATE OF DEATH IS UNKNOWN. PER THE MEDICAL SAFETY OFFICER REVIEW IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9, H3, H6 AND H10. THE INVESTIGATION INTO THE REPORTED ISSUE HAS BEEN COMPLETED. CONSOLE LOGS DID NOT CONTAIN ANY DATA RELEVANT TO THIS COMPLAINT. THE CONSOLE WAS EVALUATED, AND THE REPORTED BROKEN PURGE DISC RETAINER WAS CONFIRMED. THE ROOT CAUSE OF THIS ISSUE WAS A BROKEN PURGE DISC RETAINER. BEFORE SENDING THIS CONSOLE BACK TO THE CUSTOMER, FIELD SERVICE WAS INSTRUCTED TO PERFORM THE FOLLOWING: REPLACE THE BLUE PURGE DISC RETAINER, PURGE DISC DETECT FLAG, AND PURGE FLAG SPRING, PERFORM PREVENTATIVE MAINTENANCE AND ALL FUNCTIONAL CHECKS ON THE CONSOLE AND OPTICAL SYSTEM. Manufacturer Narrative: B1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM ON MANUFACTURER DEVICE REPORT 1220648-2025-25946. B2 REVISED TO REFLECT DEATH, THE EXACT DATE OF DEATH IS UNKNOWN 1220648-2025-25946. B5 REVISED TO INCLUDE PATIENT OUTCOME AND MSO STATEMENT WHICH INADVERTENTLY WAS NOT INCLUDED ON MANUFACTURER DEVICE REPORT 1220648-2025-25946-1. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25946 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. G1 MANUFACTURING SITE NAME AND ADDRESS REVISED AS WAS INADVERTENTLY REPORTED INCORRECTLY ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25946-1. H1 TYPE OF REPORT REVISED TO REFLECT THE PATIENT OUTCOME OF DEATH ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25946. H6 REVISED TO REFLECT THE PATIENT OUTCOME OF DEATH ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25946.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21288979&pc=MNS | 2518422-2025-101599 | 2025-01-24 05:00 | Death | RESPIRONICS, INC. | 2025-01-31 05:00 | MNS | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: IT WAS REPORTED THAT THE DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. AT THIS TIME, THE DEVICE WAS RETURNED AND A FOLLOW-UP WILL BE FILED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289004&pc=NPT | 2025587-2025-00926 | 2025-01-30 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-01-31 05:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Appropriate Term/Code Not Available; Patient Device Interaction Problem | Extravasation; Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 VALVE AND THE D-EVOLUTFX-2329 DELIVERY SYSTEM, AN INJURY OCCURRED TO THE ASCENDING AORTA WHILE ATTEMPTING TO CROSS THE SINOTUBULAR JUNCTION (STJ) WITH THE DELIVERY SYSTEM. THE INJURY WAS CONFIRMED THROUGH ANGIOGRAPHY, AND THERE WAS EXTRAVASATION FROM THE ASCENDING AORTA. UNFORTUNATELY, THE PATIENT DIED ON THE TABLE.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 VALVE AND THE D-EVOLUTFX-2329 DELIVERY SYSTEM, AN INJURY OCCURRED TO THE ASCENDING AORTA WHILE ATTEMPTING TO CROSS THE SINOTUBULAR JUNCTION (STJ) WITH THE DELIVERY SYSTEM. THE INJURY WAS CONFIRMED THROUGH ANGIOGRAPHY, AND THERE WAS EXTRAVASATION FROM THE ASCENDING AORTA. UNFORTUNATELY, THE PATIENT DIED ON THE TABLE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE INJURY OCCURRED DURING ADVANCEMENT. THE VALVE REMAINED IN THE DELIVERY CATHETER SYSTEM (DCS) AND WAS NOT IMPLANTED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 VALVE AND THE D-EVOLUTFX-2329 DELIVERY SYSTEM, AN INJURY OCCURRED TO THE ASCENDING AORTA WHILE ATTEMPTING TO CROSS THE SINOTUBULAR JUNCTION (STJ) WITH THE DELIVERY SYSTEM. THE INJURY WAS CONFIRMED THROUGH ANGIOGRAPHY, AND THERE WAS EXTRAVASATION FROM THE ASCENDING AORTA. UNFORTUNATELY, THE PATIENT DIEDON THE TABLE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE INJURY OCCURRED DURING ADVANCEMENT. THE VALVE REMAINED IN THE DELIVERY CATHETER SYSTEM (DCS) AND WAS NOT IMPLANTED. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE INJURY AND DEATH WAS THE AORTIC ROOT RUPTURE FROM THE INJURY RELATED TO THE DCS DUE TO THE DIFFICULTY CROSSING THE NATIVE VALVE OVER THE NON-MEDTRONIC GUIDEWIRE. ADDITIONALLY, THE PHYSICIAN REPORTED THAT THE INJURY WAS NOT SURVIVABLE AND THE PATIENT WAS ELDERLY SO EXTREME RESUSCITATION ATTEMPTS WERE NOT PURSUED PER THE PATIENT'S WISHES. THE PATIENT'S ANATOMY HAD A SLIGHTLY HORIZONTAL AORTA AND THERE WAS CALCIUM ABOVE THE RIGHT CORONARY ARTERY THAT WAS SUSPECTED TO BE THE SITE OF THE AORTIC TEAR; HOWEVER, SEVERE CALCIFICATION WAS NOT PRESENT. Manufacturer Narrative: UPDATED DATA: B5. THIRD PARAGRAPH ADDED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVOLUTFX-29 (J162230); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289041&pc=QBJ | 3004753838-2025-028105 | 2024-11-01 04:00 | Death | DEXCOM, INC. | 2025-01-31 05:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Incorrect, Inadequate or Imprecise Result or Readings | Insufficient Information | K213919 |  | 1 | Event Description: IT WAS REPORTED BY AN ADMINISTRATOR THAT THE PATIENT PASSED AWAY WHILE USING THE DEXCOM G7 CONTINUOUS GLUCOSE MONITOR SYSTEM. ACCORDING TO THE REPORTER, PRIOR TO THE PATIENT'S DEATH, THE PATIENT HAD EXPRESSED CONCERNS TO HER FRIENDS AND FAMILY THAT THE CGM HAD BEEN PROVIDING INCORRECT READINGS. WHETHER AND HOW THESE INACCURATE READINGS WERE RELATED TO THE PATIENT'S DEATH, HOWEVER, WERE NOT SPECIFIED, NOR WERE ANY OTHER DETAILS PROVIDED. DEXCOM REPRESENTATIVES HAVE ATTEMPTED TO FOLLOW UP WITH BOTH THE REPORTER AND THE REPORTING ORGANIZATION MULTIPLE TIMES FOR FURTHER DETAILS, BUT WERE TOLD THAT IT WAS AN ONGOING INVESTIGATION, AND NO NEW INFORMATION REGARDING THE EVENT, NOR PATIENT INFORMATION WERE RECEIVED. NO DATA WAS RETURNED FOR THE ALLEGED DATE OF INCIDENT ON (B)(6) 2024. WITHOUT PERFORMANCE DATA TO INVESTIGATE, THE COMPLAINT OF DISCREPANCIES REMAINS UNCONFIRMED. THE PROBABLE CAUSE OF THE ALLEGED EVENT COULD NOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289171&pc=CBF | 2020676-2025-00007 | 2025-01-31 05:00 | Death | SECHRIST INDUSTRIES INC. | 2025-01-31 05:00 | CBF | SECHRIST INDUSTRIES INC. | Use of Device Problem; Insufficient Information | Burn(s); Insufficient Information | K140559 |  | 1 | Event Description: ON JANUARY 31, 2025, SECHRIST BECAME AWARE OF A DEATH, TO A CHILD, THAT WAS BEING TREATED IN A HYPERBARIC CHAMBER. OFFICIALS IN (B)(6), MICHIGAN CONFIRMED A CHILD DIED AFTER THE INCIDENT. Manufacturer Narrative: THIS REPORT IS BASED SOLELY ON THE AVAILABLE INFORMATION AT THIS TIME. THIS INCIDENT IS CURRENTLY UNDER INVESTIGATION BY THE (B)(6), MICHIGAN OFFICIALS AND SECHRIST HAS NOT YET BEEN OFFICIALLY NOTIFIED. AT THIS TIME, THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE ROOT CAUSE OR IF ANY DEVICE MALFUNCTION OCCURED. MANUFACTURER REFERENCE FILE NO: (B)(4). Manufacturer Narrative: THIS WILL SUPPLEMENT THE INFORMATION SUBMITTED ON JANUARY 31, 2025 (SAME DAY OF THE INCIDENT), (B)(6) 2025 REGARDING THE INCIDENT THAT INCURRED AT (B)(6) IN (B)(6), INVOLVING A SECHRIST MODEL 3300HR HYPERBARIC CHAMBER IDENTIFIED BY S/N (B)(6). ON JUNE 26, 2025, A REPRESENTATIVE OF SECHRIST ATTENDED AN INSPECTION OF THE SECHRIST MODEL 3300HR HYPERBARIC CHAMBER IDENTIFIED BY S/N (B)(6). THE INSPECTION WAS ATTENDED BY COUNSEL REPRESENTING INDIVIDUALS AGAINST WHOM CRIMINAL CHARGES HAVE BEEN FILED IN THE STATE OF MICHIGAN. THE INSPECTION WAS ATTENDED BY COUNSEL REPRESENTING THE (B)(6) FAMILY, THE (B)(6) CENTER AND THE PROPERTY OWNER / MANAGER, THE CITY OF (B)(6) (CIVIL PARTIES). THE INSPECTION WAS ATTENDED BY CONSULTANTS RETAINED BY THE PARTIES. THE INSPECTION WAS CONDUCTED PURSUANT TO A JUNE 23, 2025 ORDER OF THE HONORABLE (B)(6) IN THE CIRCUIT COURT FOR THE COUNTY OF (B)(6) (CASE NO (B)(6)) THAT LIFTED A PRIOR RESTRAINING ORDER ON THE PREMISES AND PERMITTED THE INSPECTION. REPRESENTATIVES OF THE (B)(6) FIRE DEPARTMENT HOSTED THE INSPECTION. THE CONSULTANTS RETAINED BY THE PARTIES TOOK PHOTOS OF THE INCIDENT SCENE AND SUBJECT SECHRIST MODEL 3300HR HYPERBARIC CHAMBER. THE INSPECTION INCLUDED 3D SCANS OF THE INCIDENT SCENE AND PORTIONS OF THE SECHRIST MODEL 3300HR HYPERBARIC CHAMBER. THE INSPECTION INCLUDED TESTING OF THE GROUNDING OF THE SUBJECT SECHRIST MODEL 3300HR HYPERBARIC CHAMBER AND TESTING OF THE CIRCUITS IN THE SUBJECT SECHRIST MODEL 3300HR HYPERBARIC CHAMBER. SECHRIST REPRESENTATIVE DID NOT IDENTIFY ANY EVIDENCE THAT THERE WAS A GROUNDING STRAP IN THE SUBJECT AT THE TIME OF THE INCIDENT. SECHRIST REPRESENTATIVE CONDUCTED THE TESTING OF THE CIRCUITS OF THE SUBJECT SECHRIST MODEL 3300HR HYPERBARIC CHAMBER. SECHRIST REPRESENTATIVE DID NOT IDENTIFY ANYTHING WRONG WITH ANY OF THE CIRCUITS OF THE SUBJECT SECHRIST MODEL 3300HR HYPERBARIC CHAMBER. THE SUBJECT SECHRIST MODEL 3300HR HYPERBARIC CHAMBER AND ARTIFACTS FROM THE INCIDENT INCLUDING THE STRETCHER PLATE AND TEXTILES ON THE STRETCHER PLATE ARE BEING PRESERVED BY THE PARTIES FOR FUTURE LAB TESTING PURSUANT TO A PROTOCOL TO BE AGREED UPON. MANUFACTURER REFERENCE FILE NO.: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289205&pc=MHY | 3006630150-2025-00352 | 2025-01-03 05:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-01-31 05:00 | MHY | VERCISE CARTESIA | Adverse Event Without Identified Device or Use Problem | Fever; Intracranial Hemorrhage; Pneumonia | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT EXPERIENCED AN INTER CRANIAL BRAIN BLEED SOMETIME AFTER THE IMPLANT PROCEDURE, HOWEVER IT IS UNKNOWN IF ANY MEDICAL OR SURGICAL INTERVENTION WAS PERFORMED. THE PATIENT LATER EXPERIENCED A FEVER AND PNEUMONIA UNRELATED TO THE DEVICE AND WAS HOSPITALIZED HOWEVER IT IS UNKNOWN IF ANY MEDICAL OR SURGICAL INTERVENTION WAS PERFORMED. THE PATIENT'S PNEUMONIA AND FEVER ISSUES DID NOT IMPROVE, AND PATIENT SUCCUMBED TO HIS CONDITION. IT WAS NOTED THAT THE IMPLANT PROCEDURE MAY HAVE CONTRIBUTED TO THE PATIENTS CONDITION AND THE OFFICIAL CAUSE OF DEATH WAS DUE TO THE PNEUMONIA PER THE PHYSICIAN'S ASSESSMENT HOWEVER NO INFORMATION AS TO HOW THE PROCEDURE CONTRIBUTED TO THE PATIENTS MEDICAL CONDITION WAS PROVIDED. Manufacturer Narrative: ADDITIONAL PRO CODES IN BLOCK NHL, PJS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294665&pc=LNS | MW5165548 | 2025-01-27 05:00 | Death | UNK | 2025-01-31 05:00 | LNS | EXTRACORPOREAL SHOCK WAVE LITHOSTRIPSY |  | Insufficient Information |  |  | 1 | Event Description: OR (OPERATING ROOM) DIRECTOR (B)(6) RECEIVED A PHONE CALL FROM (B)(4), REGIONAL OPERATIONS MANAGER FOR UNITED MEDICAL SYSTEMS, INDICATING THAT THEIR EQUIPMENT (EXTRACORPOREAL SHOCK WAVE LITHOTRIPSY) HAD BEEN USED DURING A LITHOTRIPSY. POST PROCEDURE, PATIENT WAS D/C(DISCHARGED) TO HOME WITH WIFE. PATIENT'S WIFE REPORTS THAT AFTER THE PROCEDURE, THE PATIENT DIED. THE PURPOSE OF THE CALL WAS TO OBTAIN SPECIFIC PATIENT INFORMATION (DATE OF PROCEDURE AND PATIENT'S NAME). PATIENT HAD A DIAGNOSIS OF RIGHT RENAL STONES. RIGHT EXTRACORPOREAL SHOCK WAVE LITHOTRIPSY WAS PERFORMED. NO COMPLICATIONS DURING PROCEDURE NOTED. PATIENT WAS D/C TO HOME WITH HIS DAUGHTER WITH WRITTEN AND VERBAL DISCHARGE INSTRUCTIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21294692&pc=QJZ | MW5165549 | 2024-11-05 05:00 | Death | FRESENIUS MEDICAL CARE/XENIOS AG | 2025-01-31 05:00 | QJZ | NOVALUNG SYSTEM FOR ECMO |  | Cardiac Arrest |  |  | 1 | Event Description: PATIENT CAME TO CATH LAB FOR A (VENTRICULAR TACHYCARDIA) VT/(VENTRICULAR ECTOPIC) VF ARREST FOR INITIATION OF ECPR (EXTRACORPOREAL CARDIOPULMONARY RESUSCITATION). THE NOVALUNG DEVICE WAS PREPPED TO HAND UP TO THE STERILE FIELD TO START ECMO. WHEN HANDED UP THE P2 PRESSURE CABLE BROKE AND WOULD NOT LIGHT UP ON THE SCREEN. ATTEMPT TO PULL CORD AND SWITCH TO ANOTHER CABLE WAS MADE HOWEVER CABLE WAS STUCK IN DEVICE. WITH IN A FEW MINUTES OF ECMO(EXTRACORPOREAL MEMBRANE OXYGENATION) SUPPORT THE P3 PRESSURE CABLE ALSO WAS NO TRANSDUCING A PRESSURE AND IT WAS FOUND TO BE IN THE DISPOSABLE DEVICE OF THE NOVALUNG BUT NOT ABLE TO READ A PRESSURE. THIS CABLE ALSO REMAINED IN THE DEVICE. ECMO SUPPORT WAS NOT STOPPED DUE TO THE URGENCY OF NEED AT THAT TIME. THE PATIENT WAS HOOKED UP TO BLOOD PRESSURE AND EKG MONITORS TO ENSURE BEST CARE PRACTICES WERE FOLLOWED. ALSO, THE NOVALUNG WAS READING A TOTAL FLOW TO PATIENT WITH RPMS. OUR BEST PRACTICE IS TO INITIATE ECMO WITH FIO2 OF 50%, NO COLOR CHANGE WAS NOTED ON ARTERIAL LIMB OF CIRCUIT UNTIL THE OXYGEN LEVEL WAS INCREASED TO MAXIMUM SUPPORT OF 100% OXYGEN. UPON ARRIVAL TO ICU THE PATIENT'S PRESSURE CABLES WERE SWITCHED TO GAIN THE PRESSURES THAT WERE PREVIOUSLY NOT ABLE TO BE ACCESSED. ONCE CHANGED THE CLINICIANS WERE ABLE TO NOTE THE NEGATIVE PRESSURE WAS -400 MM HG. NORMAL IS LESS THAN -100 IDEALLY AROUND -50. THE OTHER ASTOUNDING THING WAS THE ECMO CIRCUIT DID NOT "CHUG" TO INDICATE IT WAS INDEED DRY OR NEEDING MORE FLUID FROM THE PATIENT. ULTIMATELY, THE PATIENT DEMISED WITHIN 5 HOURS OF ARRIVAL. SENT DISPOSABLE BACK TO FRESENIUS DUE TO OXYGENATOR NOT FUNCTIONING. LABS WERE DRAWN ON THE POST OXYGENATOR SIDE AT 100% FIO2 AND THE HIGHEST YIELD WAS A PO2 OF 87 MM HG. THE PATIENTS PO2 WAS 89 MM HG INDICATING THAT THE PATIENT'S NATIVE LUNGS WERE DOING THAT WORK AS IT WAS HIGHER THAN THE DEVICE SET IN PLACE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21297553&pc=FSA | MW5165555 | 2023-07-10 04:00 | Death | MEDACURE INC. | 2025-01-31 05:00 | FSA | U FREE SPIRIT HOYER LIFT |  | Fall |  |  | 1 | Event Description: MY MOTHER BEING MOVED VIA A HOYER LIFT FROM BED TO BRODA CHAIR. HOYER LIFT BROKE, MALFUNCTIONED AND OR NOT USED PROPERLY AND MY MOM FELL TO GROUND FROM HOYER LIFT. APPEARS HOYER LIFT BROKE / CAME APART?? MOM SENT TO HOSPITAL CAME BACK TO HER COMMUNITY DIED ON (B)(6) 2023 APPROXIMATELY 2 WEEKS LATER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21298116&pc=NPT | MW5165596 | 2025-01-27 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-01-31 05:00 | NPT | EDWARDS VALVE DELIVERY SYSTEM |  | Insufficient Information |  |  | 1 | Event Description: DELIVERY SYSTEM COULD NOT BE WITHDRAWN THROUGH THE SHEATH. SHEATH LAYERS WERE SEPARATED. LOT#: 65440309.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21272194&pc=BZD | 2518422-2025-008723 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-01-30 05:00 | BZD | DREAMSTATION CPAP | Degraded | Inflammation; Pulmonary Dysfunction; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, DEATH, OTHER AND PHYNOMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 10/14/2024 AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, DEATH, OTHER AND PHYNOMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21272516&pc=PNJ | 9612164-2025-00507 | 2025-01-08 05:00 | Death | MEDTRONIC IRELAND | 2025-01-30 05:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Aortic Valve Stenosis; Heart Failure/Congestive Heart Failure | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED AFTER IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED HEART FAILURE DUE TO WORSENING OF THE SEVERE AORTIC STENOSIS AND DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21272518&pc=PNJ | 9612164-2025-00508 | 2025-01-09 05:00 | Death | MEDTRONIC IRELAND | 2025-01-30 05:00 | PNJ | MICRA¿ AV2 | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA-OPERATIVE THE PATIENT EXPERIENCED CARDIAC PERFORATION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE AND BLOOD PRESSURE DROP DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. IT WAS REPORTED THAT THE HEART WAS SMALL AND TWO IMPLANT LOCATIONS IN THE RIGHT VENTRICULAR (RV) WERE ATTEMPTED AND THE IPG SYSTEM SLIPPED. THE IPG EXHIBITED NON-CAPTURE DURING THE PROCEDURE. THE PATIENT'S CONDITION WORSENED DUE TO VENTRICULAR TAMPONADE SO THE PROCEDURE WAS DISCONTINUED. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU). PERICARDIAL PUNCTURE, DRAINAGE, RESUSCITATION, INTUBATION AND BLOOD TRANSFUSION WERE PERFORMED. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS ATTEMPTED/NOT USED. IT WAS REPORTED THAT THE PATIENT DIED. CAUSE OF DEATH WAS PROVIDED AS CARDIAC TAMPONADE. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID: [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21272540&pc=DYB | 9612164-2025-00509 | 2025-01-09 05:00 | Death | MEDTRONIC IRELAND | 2025-01-30 05:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA-OPERATIVE THE PATIENT EXPERIENCED CARDIAC PERFORATION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE AND BLOOD PRESSURE DROP DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. IT WAS REPORTED THAT THAT THE HEART WAS SMALL AND TWO IMPLANT LOCATIONS IN THE RIGHT VENTRICULAR (RV) WERE ATTEMPTED AND THE IPG SYSTEM SLIPPED. THE IPG EXHIBITED NON-CAPTURE DURING THE PROCEDURE. THE PATIENT'S CONDITION WORSENED DUE TO VENTRICULAR TAMPONADE SO THE PROCEDURE WAS DISCONTINUED. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU). PERICARDIAL PUNCTURE, DRAINAGE, RESUSCITATION, INTUBATION AND BLOOD TRANSFUSION WERE PERFORMED. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS ATTEMPTED/NOT USED. IT WAS REPORTED THAT THE PATIENT DIED. CAUSE OF DEATH WAS PROVIDED AS CARDIAC TAMPONADE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21273101&pc=BZD | 2518422-2025-101495 | 2024-12-03 05:00 | Death | RESPIRONICS, INC. | 2025-01-30 05:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: MANUFACTURER WAS CONTACTED IN REFERENCE TO A PATIENT DEATH. PATIENT'S WIFE STATED THE PATIENT PASSED AWAY (B)(6) 2023. A CLINICAL ASSESSMENT DEEMED THE DEATH AS NOT RELATED TO THE DEVICE IN THIS CASE. IT WAS REPORTED THAT SHE RETURNED THE DEVICE AND ALL HIS SUPPLIES TO HIS DME MAKING THE DEVICE NOT AVAILABLE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21273505&pc=OZD | 1220648-2025-25905 | 2025-01-05 05:00 | Death | ABIOMED, INC. | 2025-01-30 05:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Use of Device Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT AND ESCALATED TO AN IMPELLA 5.5 FOUR DAYS LATER DUE TO THE EXPECTED PROLONGATION OF SUPPORT. APPROXIMATELY AFTER 11 DAYS ON IMPELLA 5.5 SUPPORT, LEAKING FROM PURGE CASSETTE WAS OBSERVED. PER THE CLINICAL ENGINEER LIQUID WAS LEAKING FROM THE PURGE CASSETTE ON THE FOURTH DAY AFTER STARTING USE. CONSEQUENTLY, THE PURGE SET WAS REPLACED. HOWEVER, 16 DAYS AFTER START OF SUPPORT, ABNORMAL NOISE WAS HEARD AROUND THE PURGE CASSETTE INSERTION AREA OF THE AUTOMATED IMPELLA CONTROLLER (USED WITH BOTH THE IMPELLA CP AND IMPELLA 5.5), WHICH PROMPTED REPLACEMENT OF THE PURGE SET. THE ABNORMAL NOISE WAS RESOLVED AFTER REPLACING THE PURGE HOUSING. AT SOME POINT THE PATIENT DEVELOPED SEPSIS AS CASE WRAP-UP NOTED ¿CARDIAC FUNCTION RECOVERS POORLY, AND EVENTUALLY SEPSIS DISAPPEARS.¿ THE TIMEFRAME FOR THIS ADVERSE EVENT AND ITS RELATIONSHIP TO THE IMPELLA PUMP IS UNCLEAR. HOWEVER, THE PATIENT WOULD EVENTUALLY EXPIRE BASED ON REFLECTED CASE NOTE ¿EXPIRED ON SUPPORT,¿ WHICH OCCURRED ON THE 25TH DAY OF SUPPORT. REPORTEDLY THE PATIENT WENT INTO MULTIPLE ORGAN FAILURE CAUSED BY SEPSIS. ANTIBIOTICS WERE ADMINISTERED. NO DETAIL INFORMATION SURROUNDING THE PATIENT'S CONDITION, THE SEPSIS AND DEATH WERE PROVIDED, AND WITH LIMITED INFORMATION, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS A CONTRIBUTING FACTOR IN THE PATIENT¿S OUTCOME. Manufacturer Narrative: THE INVESTIGATION FOR THE CONSOLE (AIC) PURGE ISSUES HAS BEEN COMPLETED. CONSOLE LOGS WERE REVIEWED AS A PRECAUTION BUT DID NOT CONTAIN ANY DATA RELEVANT TO THE COMPLAINT. CONSOLE WAS TESTED AFTER ARRIVING IN FIELD SERVICE. THE REPORTED ABNORMAL NOISE COULD NOT BE REPRODUCED DURING TESTING. THE CONSOLE OPERATED WITHOUT ANY ISSUES AND NO NOISE WAS OBSERVED FROM THE PURGE ASSEMBLY. THE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED DUE TO THE INABILITY TO REPRODUCE. CONSOLE OPERATED NORMALLY AND NO ABNORMAL NOISES WERE OBSERVED WHILE TESTING THE CONSOLE. Manufacturer Narrative: THE AUTOMATED IMPELLA CONTROLLER HAS BEEN RECEIVED FROM THE CUSTOMER; HOWEVER, THE EVALUATION OF THE DEVICE IS NOT COMPLETE. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MDR WILL BE FILED. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5 SERIAL NUMBER (B)(6) WITH DATE OF EVENT 05-JAN-2025 AND IDENTIFIER ENDING IN (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21274735&pc=NKM | 2135147-2025-00463 | 2013-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-01-30 05:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "IMPLICATIONS OF AN OFF-HOURS SETTING IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR FOR MITRAL REGURGITATION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE PREVALENCE, CORRELATES, AND OUTCOMES OF MITRAL TEER INITIATED OFF-HOURS, I.E. BEFORE 7:30 AM, AFTER 5:30 PM, OR ON WEEKENDS/HOLIDAYS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT A NON-INFREQUENT PROCEDURE, OFF-HOURS MITRAL TEER IS PERFORMED IN HIGH-RISK CASES BUT, IN THE HANDS OF EXPERIENCED INTERVENTIONALISTS, SHOULD PROVE SAFE, FEASIBLE, AND EFFICACIOUS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS RAJ MAKKAR AT CEDARS-SINAI MEDICAL CENTER INSTITUTION WITH CORRESPONDING EMAIL RAJ.MAKKAR@CSHS.ORG] THE TIME FRAME OF THE STUDY WAS JANUARY 2013 TO DECEMBER 202. A TOTAL OF 117 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 117 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, ANEMIA, PRIOR STROKE, ATRIAL FIBRILLATION CN-275075, UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, RECURRENT MR, DEATH. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED LITERATURE ATTACHMENT: ARTICLE TITLE "IMPLICATIONS OF AN OFF-HOURS SETTING IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR FOR MITRAL REGURGITATION" INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, ANEMIA, PRIOR STROKE, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED HEART FAILURE, PROLONGED HOSPITALIZATION, RECURRENT MR, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: IMPLICATIONS OF AN OFF-HOURS SETTING IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR FOR MITRAL REGURGITATION B2 DEATH DATE IS ESTIMATED. B3 EVENT DATE IS ESTIMATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21274787&pc=NPU | 2015691-2025-00573 | 2024-12-16 05:00 | Death | EDWARDS LIFESCIENCES | 2025-01-30 05:00 | NPU | SAPIEN M3 VALVE | Device Stenosis; Central Regurgitation | Cardiac Arrest |  |  | 1 | Event Description: AS REPORTED FROM A CLINICAL STUDY, APPROXIMATELY 1 YEAR 3 MONTHS 17 DAYS POST TRANSSEPTAL TRANSCATHETER MITRAL VALVE REPLACEMENT (TMVR) PROCEDURE WITH A 29 MM SAPIEN M3 VALVE, A TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS PERFORMED REVEALING A STENOTIC BIOPROSTHETIC MITRAL VALVE (MV) WITH A MEAN GRADIENT (MG) OF 19 MMHG. THE MITRAL VALVE STENOSIS WAS NOTED TO BE MILD. THERE WAS ALSO MILD (1+) VALVULAR MITRAL REGURGITATION. SUBSEQUENTLY, ADDITIONAL INFORMATION WAS RECEIVED VIA MEDICAL RECORDS. APPROXIMATELY 1 YEAR 3 MONTHS 17 DAYS POST TMVR PROCEDURE, THE PATIENT WAS SEEN FOR CONGESTIVE HEART FAILURE (CHF) EXACERBATION AND A TTE WAS PERFORMED WHICH REVEALED THE VALVE HAD A MEAN GRADIENT OF 19 MMHG WITH MILD CENTRAL MITRAL REGURGITATION (MR) WITH A SEVERELY REDUCED EJECTION FRACTION (EF) OF 20-25% IN CHF. IT WAS NOTED THAT THE VALVE WAS STENOTIC. APPROXIMATELY TWO DAYS LATER, THE PATIENT WENT INTO PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST AND PASSED AWAY. Manufacturer Narrative: THE COMPLAINT DEVICE, SAPIEN M3 VALVE - MODEL 9880TFX29MCLUS, IS NOT SOLD OR MARKETED IN THE US; HOWEVER, IT IS DEEMED SIMILAR TO THE SAPIEN 3 TRANSCATHETER HEART VALVE - MODEL 9600TFX29, PMA NUMBER P140031. THE PMA/510(K) FIELD WILL BE BLANK. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR CORRECTION AND INCLUDES ADDITIONAL INFORMATION BASED ON INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN CORRECTED/UPDATED: H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) AND H11 (PROVIDE NARRATIVE/DATA). THE DEVICE WAS NOT RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION AS IT HAD REMAINED IMPLANTED. AS A RESULT, NO VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS COULD BE PERFORMED. IN ADDITION, THERE WAS NO IMAGERY PROVIDED FOR EVALUATION; HOWEVER, THERE WERE MEDICAL RECORDS PROVIDED FOR EVALUATION. DURING MANUFACTURING, THE DEVICE UNDERGOES MULTIPLE 100% INSPECTIONS. IN ADDITION, THE WORK ORDER UNDERWENT FUNCTIONAL PRODUCT VERIFICATION TESTING ON A SAMPLING BASIS AS A REQUIREMENT FOR LOT RELEASE. THESE INSPECTIONS AND TESTS DURING THE MANUFACTURING PROCESS SUPPORT THAT IT IS UNLIKELY THAT A NON-CONFORMANCE CONTRIBUTED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE (IFU)/TRAINING MANUALS WERE REVIEWED FOR GUIDANCE/INSTRUCTION INVOLVING THE VALVE USAGE. BASED ON THE REVIEW OF THE IFU/TRAINING MANUALS, NO DEFICIENCIES WERE IDENTIFIED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. IN THIS CASE, THE VALVE STENOSIS AND CENTRAL REGURGITATION THAT CONTRIBUTED TO A CARDIAC ARREST AND SUBSEQUENT PATIENT DEATH WERE CONFIRMED EMPIRICALLY FROM THE MEDICAL RECORDS. A REVIEW OF THE IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. AN EXISTING TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES DOCUMENTS THE ROOT CAUSE ANALYSIS ON POST-IMPLANTATION EVENTS FOR THE SAPIEN M3 VALVE. PER THE TECHNICAL SUMMARY, VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OF DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. PER REVIEW OF THE MEDICAL RECORDS, APPROXIMATELY 1 YEAR 3 MONTHS AND 17 DAYS POST TRANSCATHETER MITRAL VALVE REPLACEMENT (TMVR), A TRANSTHORACIC ECHOCARDIOGRAM (TTE) REVEALED MILD MITRAL VALVE STENOSIS WITH A MEAN GRADIENT OF 19 MMHG. THERE WAS ALSO MILD (1+) VALVULAR MITRAL REGURGITATION NOTED. AT THE PATIENT'S 30-DAY TTE FOLLOW UP, IT WAS NOTED THAT THE SAPIEN M3 VALVE WAS WELL-SEATED WITH A MEAN GRADIENT OF 4 MMHG AND TRACE VALVULAR REGURGITATION. THERE WAS NO ADDITIONAL INFORMATION AS TO WHAT MAY HAVE CAUSED THE STENOSIS AND CENTRAL REGURGITATION TO OCCUR. IN REGARD TO THE VALVE STENOSIS, THE TECHNICAL SUMMARY LISTS THE FOLLOWING AS POTENTIAL ROOT CAUSES: LEAFLET THICKENING, REDUCED LEAFLET MOTION, INADEQUATE ANTICOAGULANT THERAPY, THROMBOSIS, HYPERLIPIDEMIA, CHRONIC KIDNEY DISEASE. IN REGARD TO THE CENTRAL REGURGITATION, THE TECHNICAL SUMMARY LISTS THE FOLLOWING AS POTENTIAL ROOT CAUSES: REDUCED LEAFLET MOTION, THROMBOSIS, DIABETES, HYPERLIPIDEMIA, LEAFLET THICKENING, ENDOCARDITIS AND UNDER EXPANDED VALVE. IN THIS CASE, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME; HOWEVER, THE VALVE STENOSIS MAY HAVE BEEN DUE TO PATIENT FACTORS WHILE THE CENTRAL REGURGITATION MAY HAVE BEEN DUE TO PATIENT AND PROCEDURAL FACTORS. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275012&pc=MVK | 3008642652-2025-01143 | 2024-11-20 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-30 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise; Circuit Failure | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 21:06:57 ON (B)(6) 2024. AT 06:13:14 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 240 BPM WITH VARYING AMPLITUDES AND ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VT. VARYING AMPLITUDES AND ELECTRODE LEAD FALL OFF PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 06:13:54, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT TIME OF TREATMENT WAS VT @ 270 BPM WITH VARYING AMPLITUDES AND ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 06:29:51 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF MONITOR HAS BEEN COMPLETED. THE REPORTED PROBLEM (DID NOT PASS INCOMING TESTING) HAS BEEN CONFIRMED. UPON INVESTIGATION THE MONITOR WAS NOT PROPERLY PRODUCING A DRIVEN GROUND SIGNAL. THE CAUSE FOR THE FAILURE WAS ISOLATED TO AN OPEN R781 DRIVEN GROUND RESISTOR ON THE COMPUTER/ANALOG BOARD. THE ROOT CAUSE FOR THE OPEN RESISTOR COULD NOT BE POSITIVELY IDENTIFIED. THE OPEN R781 RESISTOR IS CONSISTENT WITH THE PATIENT RECEIVING AN EXTERNAL DEFIBRILLATION WHILE WEARING THE LIFEVEST. THE PATIENT RECEIVED AN EXTERNAL DEFIBRILLATION. THERE IS NO INDICATION THE OPEN R781 RESISTOR CAUSED OR CONTRIBUTED TO THE PATIENT'S PASSING AS THE SYSTEM WAS ABLE TO DETECT VF RHYTHM ON THE DATE OF EVENT. DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275021&pc=MVK | 3008642652-2025-01144 | 2024-12-06 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-30 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON 12/6/2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A TREATMENT FROM THE LIFEVEST. THE DEVICE WAS STARTED UP AT 21:57:33 ON (B)(6) 2024. AT 03:27:11 ON (B)(6) 2024, AN ARRHYTHMIA WAS DETECTED. ECG IS NOT AVAILABLE. AT 03:28:01, THE PATIENT RECEIVED THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT IS UNAVAILABLE. POST SHOCK RHYTHM IS UNAVAILABLE. THE ELECTRODE BELT WAS DISCONNECTED AT 03:30:44 ON (B)(6) 2024. PER REVIEW OF THE HOLTER DATA, PATIENT IS LAST SEEN IN SINUS TACHYCARDIA @ 110 BPM WITH MOTION ARTIFACT WHEN THE HOLTER DATA ENDS AT 02:27:17 ON (B)(6) 2024. REPORTING OUT OF AN ABUNDANCE OF CAUTION DUE TO THE ECG NOT BEING AVAILABLE AT THE TIME OF TREATMENT. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275648&pc=NAY | 2955842-2025-01031 | 2024-12-24 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-30 05:00 | NAY | SUREFORM | Firing Problem | Failure to Anastomose; Insufficient Information | K173721 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LOW ANTERIOR RESECTION, THE SUREFORM 60 STAPLER INSTRUMENT FAILED TO FIRE. IT WAS REPORTED THAT THE PATIENT OF THIS PROCEDURE EXPIRED ON POST-OPERATIVE DAY 8. NO INFORMATION WAS PROVIDED REGARDING THE PATIENT'S POST-OPERATIVE COURSE OR THE CAUSE OF DEATH. THE REPORT SOURCE STATED THAT THEY DO NOT THINK THE PATIENT DEATH WAS RELATED TO THE REPORTED STAPLER ISSUE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO FURTHER DETAILS HAVE BEEN RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE EVENT INVESTIGATION IS IN PROGRESS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A LOW ANTERIOR RESECTION AND LATER DIED. A FAILED STAPLER WAS REPORTED BUT HOW THE STAPLER FAILED AND HOW THIS MAY OR MAY NOT HAVE CONTRIBUTED TO THE PATIENT DEATH IS NOT REPORTED. NO DETAILS OF THE PROCEDURE, POST OPERATIVE COURSE, OR HOW THE PATIENT DIED ARE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: A REVIEW OF THE SUREFORM 60 STAPLER INSTRUMENT LOGS WAS PERFORMED. TOGETHER WITH THE RETURNED DEVICE INVESTIGATION, THE STAPLER LOGS INDICATE THAT A DRIVETRAIN FAILURE OCCURRED AT ABOUT 67% COMPLETION WITH A LOW PEAK FIRING FORCE. THE LOG DATA ALIGNS WITH THE FINDINGS OF A DISLODGED DRIVE CABLE CRIMP. IT IS LIKELY THAT THE CABLE CRIMP DISLODGED DURING THE FIRING, RESULTING IN THE DETECTED DRIVETRAIN FAILURE AND FORCE EXPIRATION OF THE DEVICE. THE INFORMATION REPORTED FROM THE CUSTOMER STATED THAT THE PATIENT EXPIRED ON POST-OPERATIVE DAY 8; HOWEVER, NO DETAILS REGARDING THE PATIENT'S POST-OPERATIVE COURSE OR THE CAUSE OF DEATH WAS PROVIDED. PER THE PREVIOUSLY SUBMITTED FOLLOW-UP REPORT, THE HOSPITAL REPORT SOURCE STATED THAT THEY DO NOT THINK THE PATIENT DEATH WAS RELATED TO THE REPORTED STAPLER ISSUE, AND PER THE MEDICAL SAFETY OFFICER REVIEW, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: UPDATED SECTIONS: D9, H3, ANNEX CODES: B, C. DEVICE EVALUATION: INTUITIVE SURGICAL, INC. (ISI) DID RECEIVE THE SUREFORM 60 STAPLER INSTRUMENT TO PERFORM FAILURE ANALYSIS (FA) AND DID NOT CONFIRM OR REPLICATE THE CUSTOMER REPORTED COMPLAINT OF "FAILED TO FIRE". INCIDENTALLY, THE INSTRUMENT COULD NOT BE TESTED AS IT WAS FOUND TO HAVE LOOSE DRIVE CABLES THAT RESULTS IN AN ENGAGEMENT FAILURE, PREVENTING THE INSTRUMENT TO ENTER INTO FOLLOWING. ADDITIONALLY, THE INSTRUMENT WAS FOUND TO HAVE LOOSE DRIVE CABLES AT THE BACK END. AFTER REMOVING THE INSTRUMENT HOUSING, THE INSTRUMENT DRIVE CABLES WERE FOUND TO BE LOOSE. NO DAMAGE ON THE CLAMPING PULLEY WAS OBSERVED. MOREOVER THE SCREWS ON THE CLAMPING PULLEY WERE TIGHTENED. THE INSTRUMENT WAS FOUND TO HAVE FAILED MECHANICAL ENGAGEMENT DURING IN-HOUSE TESTING. THE INSTRUMENT INPUTS FAILED TO ENGAGE WITH THE IN-HOUSE SYSTEM'S STERILE ADAPTER ON THREE ATTEMPTS, PREVENTING THE INSTRUMENT FROM ENTERING INTO FOLLOWING. THE ERROR LOGS FOR THE SYSTEM INSTALLS DISPLAY AN ERROR CODE, WITH PARAMETERS INDICATING THAT THE GRIP AXIS FAILED TO ENGAGE. Manufacturer Narrative: SECTION B5 ADDITIONAL INFORMATION: THE SUREFORM 60 STAPLER INSTRUMENT MISFIRED. THE PATIENT HAD AN ANASTOMOTIC LEAK POST-OPERATION WHICH WAS TREATED WITH AN ADDITIONAL SURGERY. THE SUREFORM 60 STAPLER HAS NOT BEEN RECEIVED FOR INVESTIGATION. THE STAPLER LOGS SHOW THE FIRST SUREFORM 60 STAPLER INSTRUMENT USED DURING THIS PROCEDURE FIRED A BLACK RELOAD. WHILE FIRING THE BLACK RELOAD, THE STAPLER EXPERIENCED A DRIVETRAIN FIRING FAILURE WHICH STOPPED THE FIRING PROCESS AT ABOUT 67% COMPLETION. THE INSTRUMENT BECAME FORCE EXPIRED AS DESIGNED DUE TO THIS FAILURE MAKING IT UNUSABLE AGAIN. THIS SUREFORM 60 STAPLER INSTRUMENT WAS REPLACED WITH ANOTHER TO CONTINUE THE PROCEDURE. THE SECOND SUREFORM 60 STAPLER WAS USED TO FIRE 2 BLACK RELOADS WITHOUT ISSUE. RE-REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED 8 DAYS AFTER A LOW ANTERIOR RESECTION. ALTHOUGH THE PATIENT HAD TO UNDERGO AN ADDITIONAL OPERATION DUE TO AN ANASTOMOTIC LEAK, THE SURGEON DID NOT BELIEVE THE DEATH WAS RELATED TO THIS ISSUE. THE EVENTS LEADING TO THE DEATH AND CAUSE OF DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275987&pc=DTB | 2649622-2025-03109 | 2025-01-19 05:00 | Death | MPRI | 2025-01-30 05:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Failure to Capture; Over-Sensing; Under-Sensing | Cardiac Arrest; Low Oxygen Saturation; Respiratory Failure; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND OVERSENSING DURING DETECTED VENTRICULAR ARRHYTHMIA. ADDITIONALLY, IT WAS NOTED THAT THERE WAS LACK OF CAPTURE DURING PACING/DURING VENTRICULAR HIGH RATE. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS ALSO REPORTED THAT THE PATIENT WAS BROUGHT TO THE EMERGENCY ROOM WITH SUDDEN CARDIAC ARREST. ADDITIONAL INFORMATION RECEIVED INDICATED THERE WAS A POSSIBLE SENSING ISSUE AFFECTING CAPTURE. IT WAS NOTED THAT THE NON-SUSTAINED VENTRICULAR TACHYCARDIA (VT) EPISODES CORRESPONDED WITH CARDIOPULMONARY RESUSCITATION DURING A VENTRICULAR ARRHYTHMIA. ONCE IT WAS STABILIZED, THE PATIENT WAS TREATED WITH AN INCREASED DOSAGE OF ANTIARRHYTHMIC MEDICATION. A CHANGE IN RATE WAS REQUESTED ALONG WITH AN INCREASE IN SENSITIVITY DUE TO THE UNDERSENSING. A FEW DAYS LATER, THE RATE WAS REDUCED. FOLLOW UP WITH THE CLINIC INDICATED THAT THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) FOR ACUTE RESPIRATORY FAILURE WITH HYPOXEMIA AND HEART FAILURE. THE PATIENT PASSED AWAY WHILE INPATIENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B5, B1, H1, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21276630&pc=NPT | 2135147-2025-00467 | 2022-12-01 05:00 | Death | ABBOTT MEDICAL | 2025-01-30 05:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "EARLY CLINICAL OUTCOMES OF TRANSCATHETER AORTIC VALVE IMPLANTATION USING THE NAVITOR SYSTEM", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE 30-DAY FEASIBILITY AND SAFETY OF TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) USING THE NAVITOR SYSTEM, WITH A FOCUS ON PROCEDURAL METHODS AND POSTOPERATIVE CT ASSESSMENTS. DEVICES INCLUDED IN THE STUDY WERE SOLELY NAVITOR. THE ARTICLE CONCLUDED THE NAVITOR SYSTEM FOR TAVI DEMONSTRATED FAVORABLE 30-DAY OUTCOMES, WITH HIGH TECHNICAL SUCCESS AND SAFETY RATES. HOWEVER, A NOTABLE INCIDENCE OF HALT WAS OBSERVED, HIGHLIGHTING THE NEED FOR CAREFUL FOLLOW-UP. FUTURE STUDIES WITH EXTENDED FOLLOW-UP PERIODS ARE NECESSARY TO FURTHER EVALUATE THE LONG-TERM CLINICAL EFFECTIVENESS OF THE NAVITOR SYSTEM. [THE PRIMARY AND CORRESPONDING AUTHOR WAS GAKU NAKAZAWA, DIVISION OF CARDIOLOGY, DEPARTMENT OF MEDICINE, KINDAI UNIVERSITY FACULTY OF MEDICINE, 377-2 OHNO-HIGASHI, OSAKASAYAMA, OSAKA 589-8511, JAPAN, WITH CORRESPONDING EMAIL: GNAKAZAWA@MED.KINDAI.AC.JP] THE TIME FRAME OF THE STUDY WAS FROM DECEMBER 2022 TO DECEMBER 2023. A TOTAL OF 32 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 84.5 YEARS, THE AVERAGE WEIGHT WAS 48.3KG, AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES MELLITUS, DYSLIPIDEMIA, SMOKING HISTORY, ATRIAL FIBRILLATION, PRIOR PERMANENT PACEMAKER, MYOCARDIAL INFARCTION, CORONARY STENTING, CORONARY ARTERY BYPASS GRAFT, PERIPHERAL VASCULAR DISEASE, STROKE, AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF NAVITOR WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES MELLITUS, DYSLIPIDEMIA, SMOKING HISTORY, ATRIAL FIBRILLATION, PRIOR PERMANENT PACEMAKER, MYOCARDIAL INFARCTION, CORONARY STENTING, CORONARY ARTERY BYPASS GRAFT, PERIPHERAL VASCULAR DISEASE, STROKE, AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE.. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PERMANENT PACEMAKER), UNEXPECTED MEDICAL INTERVENTION (POST-DILATATION), DISSECTION, ATRIOVENTRICULAR BLOCK, NON-UNIFORM EXPANSION, HYPO-ATTENUATED LEAFLET THICKENING; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED LITERATURE: ARTICLE TITLE: EARLY CLINICAL OUTCOMES OF TRANSCATHETER AORTIC VALVE IMPLANTATION USING THE NAVITOR SYSTEM. Manufacturer Narrative: B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT IS ESTIMATED. D4 - THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBERS WERE NOT PROVIDED. LITERATURE: ARTICLE TITLE "EARLY CLINICAL OUTCOMES OF TRANSCATHETER AORTIC VALVE IMPLANTATION USING THE NAVITOR SYSTEM". INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277000&pc=NGV | 2124215-2025-04506 | 2025-01-08 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-30 05:00 | NGV | WATCHMAN FLX PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: HEAL-LAA STUDY: IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 27MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED WITH COMPLETE DEVICE SEAL NOTED. THE PATIENT WAS DISCHARGED ON RIVAROXABAN. 184 DAYS POST INDEX PROCEDURE, THE PATIENT DIED. NO DIAGNOSTICS TEST WAS PERFORMED, AND NO ACTION WAS TAKEN. AT THE TIME OF THE EVENT, THE PATIENT WAS ON RIVAROXABAN. THE CAUSE OF DEATH REMAINS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277111&pc=MNS | 2518422-2025-009770 | 2024-03-21 04:00 | Death | RESPIRONICS, INC. | 2025-01-30 05:00 | MNS | DREAMSTATION BIPAP AUTOSV | Degraded | No Clinical Signs, Symptoms or Conditions | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA/VOMITING, NEW OR WORSENING ASTHMA, INFLAMMATORY RESPONSE, LUNG DISEASE, CHRONIC BRONCHITIS, IRRITATION, HIGH BLOOD PRESSURE AND RAPID HEARTBEAT. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277115&pc=MKJ | 3030677-2025-000334 | 2025-01-28 05:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-01-30 05:00 | MKJ | FRX DEFIBRILLATOR, UK ENGLISH, EXCHANGE | Failure to Power Up | Cardiac Arrest | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT UPON RETRIEVAL OF THE DEVICE TO BE USED ON THE PATIENT, THE DEVICE'S BATTERY WAS FLAT. THE PATIENT WAS DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277243&pc=MVK | 3008642652-2025-01145 | 2024-12-14 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-01-30 05:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 20:15:21 ON (B)(6) 2024. AT 20:15:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 20:15:21, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 23:56:46 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277634&pc=NGV | 2124215-2025-04766 | 2025-01-28 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-01-30 05:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Hypoxia; Tachycardia |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX LAA CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE CLOSURE DEVICE IMPLANTED IN THE LAA OF THE PATIENT. THERE WERE NO PATIENT COMPLICATIONS THAT WERE REPORTED TO HAVE OCCURRED DURING THE PROCEDURE. POST PROCEDURE THE PATIENT BECAME TACHYCARDIC, THEIR BLOOD PRESSURE DROPPED, AND THEIR OXYGEN SATURATION DETERIORATED. THE PATIENT WAS GIVEN CATECHOLAMINES AND THEIR OXYGEN WAS SUPPORTED. THE PATIENT STABILIZED BEFORE BEING REPOSITIONED AND DETERIORATED AGAIN. THE PATIENT WENT INTO CARDIAC ARREST, CARDIOPULMONARY RESUSCITATION WAS STARTED, AND THE PATIENT WAS INTUBATED. THE PERICARDIUM WAS REPEATEDLY EXAMINED FOR TAMPONADE AND THE ABDOMINAL CAVITY FOR FREE FLUID (BLOOD), WITH NO NEGATIVE FINDINGS. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED. THE IMPLANTING PHYSICIAN DOES NOT ASSUME THAT IT WAS DUE TO THE WATCHMAN FLX CLOSURE DEVICE OR WATCHMAN ACCESS SYSTEM BUT RATHER THAT THE PATIENT'S GENERAL CONDITION WAS THE CAUSE OF THE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277774&pc=NZU | 1018233-2025-00438 | 2024-11-21 05:00 | Death | C.R. BARD, INC. (COVINGTON) -1018233 | 2025-01-30 05:00 | NZU | PUREWICK MALE EXTERNAL CATHETER | Suction Failure | Sepsis; Urinary Tract Infection; Fungal Infection | EXEMPT |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT CONTRACTED A MALE YEAST INFECTION THAT WENT SEPTIC AFTER USING THE PUREWICK MALE EXTERNAL CATHETER. AT THE TIME PATIENT WAS MEDICALLY FRAGILE AND THEY DIED IN 2024 FROM LIVER AND KIDNEY FAILURE CAUSED PRIMARILY BY THE TOXIC ANTIFUNGALS GIVEN TO THEM TO TREAT THAT SEPTIC INFECTION. THE PRODUCT WAS SUPPOSED TO WICK AWAY URINE AND WAS SUPPOSED TO BE SUCTIONED INTO A VACUUM BOTTLE. CUSTOMER HAD A GREAT EXPERIENCE WITH THE FEMALE CATHETER AND HOPED THE MALE VERSION WOULD BE EVEN BETTER. INSTEAD, THE MALE VERSION KEPT URINE ON PATIENT SKIN OVERNIGHT AND THEY WERE AT THE EMERGENCY ROOM WITHIN 24 HRS WITH A URINARY TRACT INFECTION THAT TURNED SEPTIC. THE TESTS SHOWED CANDIDIASIS AS THE MAIN PROBLEM, AND PATIENT WAS SUBSEQUENTLY TREATED WITH STRONG ANTIFUNGAL MEDS THAT BASICALLY OVERWHELMED THEIR LIVER AND KIDNEYS. PATIENT DIED IN 2024 OF LIVER AND KIDNEY FAILURE. THIS MEDWATCH REPORT WAS RECEIVED VIA EMAIL ON 09-JAN-2025. BD HAS COMPLETED A REVIEW OF COMPLAINTS TO DETERMINE IF THIS ISSUE WAS PREVIOUSLY REPORTED; HOWEVER, THERE WERE NO COMPLAINTS IDENTIFIED RELATED TO THIS MEDWATCH REPORT. BD IS UNABLE TO OBTAIN ADDITIONAL INFORMATION AS THE INITIAL REPORTER HAS ASKED TO REMAIN CONFIDENTIAL.Event Description: IT WAS REPORTED THAT PATIENT CONTRACTED A MALE YEAST INFECTION THAT WENT SEPTIC AFTER USING THE PUREWICK MALE EXTERNAL CATHETER. AT THE TIME PATIENT WAS MEDICALLY FRAGILE AND THEY DIED IN 2024 FROM LIVER AND KIDNEY FAILURE CAUSED PRIMARILY BY THE TOXIC ANTIFUNGALS GIVEN TO THEM TO TREAT THAT SEPTIC INFECTION. THE PRODUCT WAS SUPPOSED TO WICK AWAY URINE AND WAS SUPPOSED TO BE SUCTIONED INTO A VACUUM BOTTLE. CUSTOMER HAD A GREAT EXPERIENCE WITH THE FEMALE CATHETER AND HOPED THE MALE VERSION WOULD BE EVEN BETTER. INSTEAD, THE MALE VERSION KEPT URINE ON PATIENT SKIN OVERNIGHT AND THEY WERE AT THE EMERGENCY ROOM WITHIN 24 HRS WITH A URINARY TRACT INFECTION THAT TURNED SEPTIC. THE TESTS SHOWED CANDIDIASIS AS THE MAIN PROBLEM, AND PATIENT WAS SUBSEQUENTLY TREATED WITH STRONG ANTIFUNGAL MEDS THAT BASICALLY OVERWHELMED THEIR LIVER AND KIDNEYS. PATIENT DIED IN 2024 OF LIVER AND KIDNEY FAILURE. THIS MEDWATCH REPORT WAS RECEIVED VIA EMAIL ON 09-JAN-2025. BD HAS COMPLETED A REVIEW OF COMPLAINTS TO DETERMINE IF THIS ISSUE WAS PREVIOUSLY REPORTED; HOWEVER, THERE WERE NO COMPLAINTS IDENTIFIED RELATED TO THIS MEDWATCH REPORT. BD IS UNABLE TO OBTAIN ADDITIONAL INFORMATION AS THE INITIAL REPORTER HAS ASKED TO REMAIN CONFIDENTIAL. Manufacturer Narrative: THE REPORTED EVENT WAS INCONCLUSIVE BECAUSE NO SAMPLE WAS AVAILABLE FOR EVALUATION AND FURTHER INVESTIGATION DID NOT RESULT IN ANY ADDITIONAL FINDINGS. NO ACTIONS CAN BE TAKEN AT THIS TIME SINCE A ROOT CAUSE WAS NOT IDENTIFIED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT SHOW ANY PROBLEMS OR CONDITIONS THAT WOULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE WERE FOUND ADEQUATE AND STATE THE FOLLOWING: WARNINGS: DO NOT USE ON IRRITATED SKIN. EXAMPLES INCLUDE, BUT ARE NOT LIMITED TO, RASHES, SKIN TEARS, OR BLISTERS. DO NOT COVER FRESH SURGICAL WOUNDS WITH THE DEVICE. DO NOT USE THE DEVICE WITH ANY MATERIAL OR IN ANY POSITION THAT BLOCKS THE VENTS ON THE FRONT OF THE DEVICE OR DOES NOT ALLOW FOR AIRFLOW THROUGH THE DEVICE. EXAMPLES INCLUDE, BUT ARE NOT LIMITED TO, TIGHT CLOTHING OR BRIEFS. TO AVOID POTENTIAL SKIN INJURY, NEVER PULL THE DEVICE DIRECTLY AWAY FROM THE USER. ALWAYS PEEL IN THE DIRECTION FROM HEAD TO FOOT. STOP USE IF AN ALLERGIC REACTION OCCURS. AFTER USE, THIS PRODUCT MAY BE A POTENTIAL BIOHAZARD. DISPOSE OF IN ACCORDANCE WITH APPLICABLE LOCAL, STATE, AND FEDERAL LAWS AND REGULATIONS. PRECAUTIONS: NOT RECOMMENDED FOR USERS WHO: ARE AGITATED, COMBATIVE, OR UNCOOPERATIVE AND MIGHT REMOVE THE DEVICE. HAVE FREQUENT EPISODES OF BOWEL INCONTINENCE WITHOUT A FECAL MANAGEMENT SYSTEM IN PLACE. HAVE SKIN IRRITATION OR BREAKDOWN AT THE SITE. DO NOT USE BARRIER CREAMS, LOTIONS, OR OINTMENTS ON THE PENIS OR PUBIC SKIN AROUND THE BASE OF THE PENIS WHEN USING THE DEVICE. THESE MAY IMPEDE SUCTION OR WEAKEN ADHESIVE. PROCEED WITH CAUTION IN USERS WHO HAVE HAD RECENT SURGERY OF THE EXTERNAL MALE ANATOMY. ALWAYS CHECK SKIN FOR IRRITATION OR BREAKDOWN AND CLEAN AND DRY THE SKIN PRIOR TO PLACEMENT OF A NEW DEVICE. CORRECTION: D. UDI FORMAT UPDATED WITH AVAILABLE PRODUCT INFORMATION PER GUDID. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. Manufacturer Narrative: THE INVESTIGATION IS STILL IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE FILED. UDI FORMAT UPDATED WITH AVAILABLE PRODUCT INFORMATION PER GUDID. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21278233&pc=OZD | 1220648-2025-25924 | 2025-01-15 05:00 | Death | ABIOMED, INC. | 2025-01-30 05:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT HAD IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD ACCESS SITE BLEEDING AND LIMB ISCHEMIA. THE INTERVENTION FOR THE ACCESS SITE BLEEDING IS UNKNOWN. STAFF MONITORED THE PATIENT'S LIMB ISCHEMIA AND SUPPORT CONTINUED. THE PATIENT EXPIRED AFTER 129.08 HOURS OF IMPELLA SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING AND LIMB ISCHEMIA COULD NOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. H.6 CODE 1884 WAS INADVERENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25924 AND WAS ADDED. CODE 4641 WAS REPORTED INCORRECTLY ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25924.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21278325&pc=OZD | 1220648-2025-25925 | 2025-01-14 05:00 | Death | ABIOMED, INC. | 2025-01-30 05:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE US COMPLAINANT HAD PLACED A 5.5 PUMP TO SUPPORT A PATIENT ADMITTED IN WITH MULTIPLE COMORBIDITIES WHILE AWAITING TRANSPLANT. DURING THE SUPPORT THERE WAS A CEREBROVASCULAR ACCIDENT. THE TEAM QUESTIONED CLOT INGESTION AS THE HEPARIN HAD BEEN STOPPED SYSTEMICALLY IN PREPARATION FOR THE TRANSPLANT. THE IMPELLA REMAINS ON FOR SUPPORT (WITH SODIUM BICARB IN PURGE) DESPITE THE HARM.Event Description: ADDITIONAL INFORMATION WAS RECEIVED REGARDING THE PATIENT'S OUTCOME. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO THE PATIENT'S DEATH. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE HAS NOT BEEN POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE STATE THE FOLLOWING: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. E4 SHOULD HAVE BEEN LEFT BLANK ON MANUFACTURER DEVICE REPORT 1220648-2025-25925. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED REGARDING THE PATIENTS OUTCOME. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO THE PATIENTS DEATH. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. B.2 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925 TO REFLECT THE PATIENT OUTCOME OF DEATH AND DATE OF DEATH WAS ADDED. B.5 REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925 TO ADD THE PATIENT OUTCOME OF DEATH. D.6B ADDED EXPLANT DATE AS IT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925. E.1 ADDED FAX NUMBER AS IT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925. H.1 TYPE OF REPORT WAS REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925 TO REFLECT THE PATIENT OUTCOME OF DEATH. H.6 HEALTH EFFECT - IMPACT CODE WAS REVISED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925 TO REFLECT THE PATIENT OUTCOME OF DEATH. H.10 ADDED INFORMATION THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-25925.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21278622&pc=NPT | 2015691-2025-00599 | 2024-06-22 04:00 | Death | EDWARDS LIFESCIENCES | 2025-01-30 05:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN PUERTO RICO, REGARDING A 23MM SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IMPLANT IN AORTIC POSITION, THE PATIENT TOLERATED THE IMPLANT WELL BUT CARDIOGENIC SHOCK DEVELOPED ASSOCIATED WITH ANEMIA AND RENAL FAILURE. TEE SHOWED NO MAJOR ANATOMIC DEFECTS AMENABLE TO REPAIR. THE PATIENT EXPERIENCED CLINICAL IMPROVEMENT BUT ON POD-2 DEVELOPED CARDIOGENIC SHOCK NOT RESPONSIVE TO ANY CLINICAL INTERVENTION. AFTER CAREFUL EVALUATION, IT WAS DETERMINED THAT THE PATIENT'S CONDITION COULD NOT BE IMPROVED THROUGH SURGICAL INTERVENTION AND THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS DEVICE-RELATED BLEED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (ANATOMIC DEFECTS) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT, HOWEVER DUE TO LIMITED INFORMATION PROVIDED A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21279700&pc=NPT | 2135147-2025-00472 | 2022-12-14 05:00 | Death | ABBOTT MEDICAL | 2025-01-30 05:00 | NPT | PORTICO TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Stroke/CVA; Hemorrhage/Bleeding; Unspecified Infection; Inflammation; Myocardial Infarction; Brain Injury; Heart Failure/Congestive Heart Failure; Respiratory Insufficiency; Unspecified Tissue Injury; Insufficient Information | P190023 | RWD2300134 | 80 | Event Description: IT WAS REPORTED THROUGH STS/ACC TVT REGISTRY DATA THAT PORTICO/NAVITOR DEVICES MAY BE RELATED TO 170 ADVERSE EVENTS WHICH ARE CONSIDERED DEATH INCLUDING: ACUTE MYOCARDIAL INFARCTION, RENAL, OTHER NON-CARDIOVASCULAR REASON, PULMONARY, OTHER CARDIOVASCULAR REASON, INFECTION, HEMORRHAGE, HEART FAILURE, CARDIOVASCULAR PROCEDURE, STROKE, INFLAMMATORY/IMMUNOLOGIC, SUDDEN CARDIAC DEATH, NEUROLOGICAL, UNKNOWN, CARDIOVASCULAR HEMORRHAGE, NON-CARDIOVASCULAR PROCEDURE OR SURGERY, AND TRAUMA. THE RELATIONSHIP OF THE ADVERSE EVENTS TO THE PORTICO/NAVITOR DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. STS/ACC TVT REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - RWD2300134. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 15-NOVEMBER-2021 ¿ 27-JUNE-2024. PATIENTS¿ MEAN AGE IS 83 YEARS, RANGING FROM 55 TO (B)(6) YEARS. 43% OF THE PATIENTS WERE MALE, 57% OF THE PATIENTS WERE FEMALE.Event Description: IT WAS REPORTED THROUGH STS/ACC TVT REGISTRY DATA THAT PORTICO/NAVITOR DEVICES MAY BE RELATED TO 80 ADVERSE EVENTS WHICH ARE CONSIDERED DEATH INCLUDING: HEART FAILURE, SUDDEN CARDIAC DEATH, OTHER NON-CARDIOVASCULAR REASON, UNKNOWN, INFECTION, OTHER CARDIOVASCULAR REASON, CARDIOVASCULAR PROCEDURE, PULMONARY, HEMORRHAGE, NEUROLOGICAL, STROKE, ACUTE MYOCARDIAL INFARCTION, NON-CARDIOVASCULAR PROCEDURE OR SURGERY, TRAUMA. THE RELATIONSHIP OF THE ADVERSE EVENTS TO THE PORTICO/NAVITOR DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. STS/ACC TVT REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - (B)(4). THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 14 DECEMBER 2022 ¿ 27-JUNE-2024. PATIENTS¿ MEAN AGE IS 83 YEARS, RANGING FROM 55 TO 96 YEARS. 49% OF THE PATIENTS WERE MALE, 51% OF THE PATIENTS WERE FEMALE. AS OF (B)(6) 2024, 1936 PATIENTS WERE TREATED WITH THE NAVITOR DEVICE AND 584 PATIENTS WERE TREATED WITH THE PORTICO DEVICE IN THE STS/ACC TVT REGISTRY.Event Description: Unique Complaint ID Number,Event Date,Date Entered,Manufacturer Aware Date,Brand Name,Generic Name,Model Number,Lot Number,Catalog Number,Serial Number,UDI Number,PMA / 510K Number,Date Returned to Manufacturer,Type of Reportable Event,Event Description,Manufacturer Narrative,Medical History,Patient Age,Patient Gender,Patient Weight,Date Implanted,Date Explanted,Health Effect Clinical Code,Health Effect Impact Code,Device Problem Code,Device Component Code,Investigation Type Code,Investigation Findings Code,Investigation Conclusion Code,Remedial Action Type,Latest Line Item Version
<br>40572119,6/6/2023,5/2/2025,12/17/2024,Portico,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK PORTICO,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Portico was implanted on 12/14/2022 in a 83 year old Female. On 06/06/2023, patient death was reported due to Infection. The relationship of the Death to the Portico could not be determined based on the limited data received from the registry.","An event of death due to Infection was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,83,Female,73.1,12/14/2022,NI,1930,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41273449,8/16/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 06/15/2023 in a 83 year old Male. On 08/16/2023, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,83,Male,64.28,6/15/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41798176,8/24/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 03/09/2023 in a 82 year old Female. On 08/24/2023, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,82,Female,90,3/9/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41315976,9/7/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 08/03/2023 in a 91 year old Male. On 09/07/2023, patient death was reported due to Other non-cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,91,Male,77.11,8/3/2023,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>10191801,9/14/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 08/02/2023 in a 80 year old Female. On 09/14/2023, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,80,Female,61.7,8/2/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41076823,10/4/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 06/22/2023 in a 76 year old Male. On 10/04/2023, patient death was reported due to Other non-cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,76,Male,50.1,6/22/2023,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41319948,10/11/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/11/2023 in a 88 year old Female. On 10/11/2023, patient death was reported due to Other cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,88,Female,55.85,10/11/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41516103,10/18/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/16/2023 in a 78 year old Male. On 10/18/2023, patient death was reported due to Other cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,78,Male,108,10/16/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41451097,10/18/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/18/2023 in a 55 year old Male. On 10/18/2023, patient death was reported due to Cardiovascular procedure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Cardiovascular procedure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,55,Male,85.5,10/18/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41654417,10/19/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/16/2023 in a 88 year old Male. On 10/19/2023, patient death was reported due to Sudden cardiac death. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Sudden cardiac death was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,88,Male,63,10/16/2023,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41392141,10/24/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 05/30/2023 in a 91 year old Female. On 10/24/2023, patient death was reported due to Other cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,91,Female,75.3,5/30/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41657447,10/27/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/20/2023 in a 74 year old Male. On 10/27/2023, patient death was reported due to Pulmonary. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Pulmonary was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,74,Male,85,10/20/2023,NI,4462,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41302301,10/28/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 09/05/2023 in a 86 year old Female. On 10/28/2023, patient death was reported due to Other non-cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,86,Female,62.2,9/5/2023,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41735370,10/30/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/19/2023 in a 87 year old Male. On 10/30/2023, patient death was reported due to Pulmonary. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Pulmonary was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,87,Male,74.5,10/19/2023,NI,4462,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41892498,11/7/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/31/2023 in a 86 year old Male. On 11/07/2023, patient death was reported due to Neurological. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Neurological was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,86,Male,68.1,10/31/2023,NI,2219,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41478001,11/20/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/23/2023 in a 89 year old Female. On 11/20/2023, patient death was reported due to Other non-cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,89,Female,62,10/23/2023,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41549266,11/26/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 07/05/2023 in a 83 year old Male. On 11/26/2023, patient death was reported due to Other non-cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,83,Male,71.4,7/5/2023,NI,1762,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41915471,11/28/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 11/28/2023 in a 80 year old Male. On 11/28/2023, patient death was reported due to Cardiovascular procedure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Cardiovascular procedure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,80,Male,96.4,11/28/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41527699,11/29/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 11/07/2023 in a 67 year old Female. On 11/29/2023, patient death was reported due to Other cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk assessment; therefore, the complaint is a foreseeable event. The current occurrence rates continue to remain within the documented rates in the device risk assessment. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product issue therefore, no remedial action is required at this time.",No Information,67,Female,77,11/7/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41786007,11/30/2023,5/2/2025,12/17/2024,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 11/27/2023 in a 86 year old Female. On 11/30/2023, patient death was reported due to Other non-cardiovascular reason. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Other non-cardiovascular reason was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined and there is no indication of a product quality issue with regards to manufacture, design, or labeling. A risk assessment was completed for the reported incident. All reported device and patient complications are included in the device risk as... Edited for length. See MAUDE Adverse Event Report for complete event narratives.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21280076&pc=DSQ | 2916596-2025-00649 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-01-30 05:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED "LIFETIME MANAGEMENT AFTER CONTINUOUS-FLOW LVAD IMPLANTATION AS BRIDGE TO TRANSPLANTATION" THAT HEARTMATE 3 MAY BE ASSOCIATED WITH DEATH AND TRANSPLANTATION. THE SURVIVAL RATES OF 228 PATIENTS WHO WERE IMPLANTED WITH CF-LVAD FOR BTT BETWEEN 2007 AND 2023. THERE WERE 164 MALE PATIENTS (72%) WITH A MEAN AGE OF 41.1 YEARS (11 ¿ 67). AVERAGE BODY SIZE WAS 166.8 CM IN HEIGHT, 56.3 KG IN WEIGHT AND 1.62 M2 IN BODY SURFACE AREA. MAJORITY OF ETIOLOGY FOR ADVANCED HEART FAILURE WAS IDIOPATHIC DCM IN 142 (62%), FOLLOWED BY DILATED PHASE OF HYPERTROPHIC CARDIOMYOPATHY IN 18 (8%), ISCHEMIC HEART DISEASE IN 17 (7%) AND MISCELLANEOUS. INTERMACS PROFILES AT IMPLANT WAS 1 IN 5 (2%), 2 IN 80 (35%), 3 IN 131 (57%) AND 4 IN 12 (5%). VARIOUS TYPES OF DEVICES WERE USED SUCH AS HEARTMATE II IN 72, EVAHEART IN 46, JARVIK 2000 IN 38, HEARTMATE 3 IN 30, HVAD IN 22 AND DURAHEART IN 20. FOLLOW-UP RATE WAS 100% COMPLETE. THE MEAN LVAD SUPPORT DURATION WAS 1250 DAYS. SURVIVAL AFTER LVAD IMPLANTATION WAS 94.2% FOR 1 YEAR, 91.7% FOR 2 YEARS, 88.3% FOR 3 YEARS, AND 84.1% FOR 4 YEARS. THE SURVIVAL COMBINING LVAD IMPLANTATION AND TRANSPLANTATION OR WEANING FROM THE DEVICE WAS 94.0% FOR 1 YEAR, 87.0% FOR 3 YEARS, 82.2% FOR 5 YEARS, AND 76.6% FOR 10 YEARS AFTER LVAD IMPLANTATION. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01JAN2023 AS PATIENTS WERE IMPLANTED BETWEEN 2007 AND 2023. DEPARTMENT OF CARDIOVASCULAR SURGERY, THE UNIVERSITY OF TOKYO, TOKYO, JAPAN; 2 DEPARTMENT OF CARDIOVASCULAR MEDICINE, THE UNIVERSITY OF TOKYO, TOKYO, JAPAN. ONO, M., ANDO, M., SHIMADA, S., YAMAUCHI, H., BUJO, C., AMIYA, E., & HATANO, M. (2024B). 326.2: LIFETIME MANAGEMENT AFTER CONTINUOUS-FLOW LVAD IMPLANTATION AS BRIDGE TO TRANSPLANTATION. TRANSPLANTATION, 108(9S). HTTPS://DOI.ORG/10.1097/01.TP.0001065480.21668.95 THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF 90979616). NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21280360&pc=NAY | 2955842-2025-01028 | 2023-12-27 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-01-30 05:00 | NAY | DAVINCI SI | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K081137 |  | 1 | Event Description: A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE ANALYSIS OF A TOTAL OF 282 PATIENTS THAT UNDERWENT LAPAROSCOPIC (LS) AND ROBOT-ASSISTED (RA) SURGERY IN NONGYNAECOLOGICAL PROCEDURES FROM FEBRUARY 2020 TO SEPTEMBER 2022. THE STUDY WAS CONDUCTED TO COMPARE THE OUTCOMES OF LAPAROSCOPIC (LS) AND ROBOTIC-ASSISTED (RA) SURGERY. A TOTAL OF 74 PATIENTS UNDERWENT ROBOTIC-ASSISTED SURGERIES, AND 208 PATIENTS UNDERWENT LAPAROSCOPIC SURGERIES. AMONG THE 54 PATIENTS WHO UNDERWENT ROBOTIC-ASSISTED HYSTERECTOMIES, ONE PATIENT EXPERIENCED PULMONARY EMBOLISM AFTER THE OPERATION. THE DESIGNATED AUTHOR WAS CONTACTED AND RESPONDED THAT THEY HAVE NOT FOUND A LINK BETWEEN THE DEATH OF A PATIENT FROM THROMBOEMBOLISM AND THE USE OF THE DA VINCI TECHNIQUE. THE PATIENT DIED 4 HOURS AFTER THE OPERATION DESPITE THE PREVENTION OF THROMBOEMBOLIC COMPLICATIONS IN THE FORM OF LOW MOLECULAR WEIGHT HEPARIN 12 HOURS BEFORE THE OPERATION. Manufacturer Narrative: NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. THERE IS NO INDICATION THAT ANY DA VINCI PRODUCTS CONTRIBUTED TO THE REPORTED COMPLICATION. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. A PATIENT IN THIS ARTICLE DIED FROM A PULMONARY EMBOLISM. NO OTHER INTRAOPERATIVE OR POST OPERATIVE DETAILS ARE PROVIDED. CITATION: ALIMOV V.A., GREKOV D.N., NOVIKOVA E.G., ET AL. COMPARATIVE ANALYSIS OF ROBOT-ASSISTED AND LAPAROSCOPIC OPERATIONS IN GYNECOLOGIC ONCOLOGY. TUMORS OF THE FEMALE REPRODUCTIVE SYSTEM 2024; 20(1):104¿13. DOI: HTTPS://DOI.ORG/10.17650/1994-4098-2024-20-1-104-113. SECTION A: PATIENT INFORMATION - NO SPECIFIC PATIENT DEMOGRAPHICS WERE PROVIDED FOR THIS PATIENT. THE MEDIAN AGE OF THE PATIENTS THAT UNDERWENT HYSTERECTOMIES WAS 57 YEARS. FIELD B3: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE EVENT DATE ASSOCIATED WITH THE ADVERSE EVENT, AN ALTERNATE DATE, PUBLISHED DATE HAS BEEN USED. SECTION D - SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21280395&pc=FRN | 3004548776-2025-00036 | 2024-12-30 05:00 | Death | FRESENIUS VIAL S.A.S. | 2025-01-30 05:00 | FRN | VOLUMAT AGILIA BR | Infusion or Flow Problem | Asystole |  |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED: EVENT DATE: 30/12/2024, PREMISES: UTI ADULT STERILE, HOURS: BETWEEN 11:00 AM AND 1:07 PM. "SEVERE PATIENT ON VASOACTIVE DRUG (NORADRENALINE 20 ML/H). IT WAS OBSERVED THAT THE INFUSION WAS NOT INFUSING AS PROGRAMMED IN THE BIC, AT NO TIME DID THE BIC SIGNAL INFUSION FAILURE, BUBBLES AND ALARM, WHEN IT WAS EVIDENCED BY THE ATTENDANT THAT THE PATIENT PRESENTED PCR, EVOLVING TO DEATH."INFUSION PUMP SERIAL NUMBER: (B)(6). "INFUSION PUMP TAG HGA-017/ SERIAL NUMBER: (B)(6) STOPPED INFUSING NORADRENALINE WITH THE WARNING OF BUBBLES IN THE CIRCUIT, TO THE MULTIPARAMETRIC MONITOR PATIENT WITH MAP 32MMHG, I CHECK CIRCUIT AND DO NOT IDENTIFY BUBBLES, RESTART PUMP, PATIENT ENDS UP EVOLVING TO PCR IN ASYSTOLE, RPC START FOLLOWING ACLS, BUT PATIENT EVOLVED TO DEATH." MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.Event Description: FRESENSIUS KABI RECEIVED THE FOLLOWING REQUEST FOR ADDITIONAL INFORMATION: ON JANUARY 30, 2025, WE RECEIVED MDR 3004548776-2025-00036 (IMPORTER MDR#) FOR A VOLUMAT AGILIA INFUSION PUMP (MODEL NUMBER Z019040, SERIAL NUMBER (B)(6)) STATING A PATIENT WAS ON NORADRENALINE. DURING TREATMENT, IT WAS OBSERVED THAT THE INFUSION WAS NOT INFUSING AS PROGRAMMED AND THE PATIENT WAS DETERIORATING. SPECIFICALLY, THE INFUSION PUMP STOPPED INFUSING WITH THE WARNING OF BUBBLES IN THE CIRCUIT; HOWEVER, AN INSPECTION OF THE CIRCUIT DID NOT SHOW BUBBLES. THE PATIENT EXPIRED. FRESENIUS KABI HAS ADDRESSED THE ADDITIONAL FOLLOW UP QUESTIONS AS FOLLOWS: THE PUMP TRIGGERED A FUNCTIONAL AIR ALARM, WHICH IS A NORMAL RESPONSE WHEN THE AIR SENSOR DETECTS A CERTAIN THRESHOLD OF CUMULATIVE AIR PASSING THROUGH THE LINE. SINCE THE PUMP WAS NOT RETURNED FOR EVALUATION, THE SPECIFIC VALUE FOR CUMULATIVE AIR SET IN THE PUMP'S PARAMETERS CANNOT BE DETERMINED AT THIS TIME. THE FUNCTIONAL AIR ALARM NOTIFICATION IS CLASSIFIED AS A HIGH-PRIORITY ALARM. IT TRIGGERS AN AUDIBLE SOUND, FLASHES A RED LED, AND DISPLAYS THE "AIR ALARM" MESSAGE, WHICH BLINKS ON THE SCREEN. A REVIEW OF FUNCTIONAL AIR ALARM COMPLAINTS FROM JANUARY 2023 TO APRIL 2025 WAS CONDUCTED. THERE HAVE BEEN NO ADDITIONAL INSTANCES OF PATIENT DEATH OR SERIOUS INJURY RESULTING FROM DELAYS IN RESPONDING TO A FUNCTIONAL AIR ALARM. THERE HAVE BEEN NO RECALLS FOR THE VOLUMAT AGILIA INFUSION PUMP RELATED TO REPORTS OF THE FUNCTIONAL AIR ALARM. THE INSTRUCTIONS FOR USE (IFU) CONTAIN SEVERAL REFERENCES REGARDING FUNCTIONAL AIR ALARM WARNINGS AND/OR SAFETY FEATURES, INCLUDING OPERATOR RESPONSE AND/OR PROBLEM RESOLUTION. THE INVESTIGATION AND EVALUATION CAN ONLY BE CONDUCTED THROUGH A LOG REVIEW, AS THE PUMP HAS BEEN RETAINED BY THE HOSPITAL AND REINTRODUCED FOR PATIENT USE. THE LOG REVIEW CLEARLY INDICATES THAT A FUNCTIONAL AIR ALARM OCCURRED AT 13:02:33 ON (B)(6) 2024. ON (B)(6) 2024 AT 12:14:55, THE DEVICE IS SWITCHED ON, ON BATTERY. AT 12:15:37, THE INFUSION STARTS @3 ML/H WITH A PRESSURE THRESHOLD @750 MMHG. FROM 12:15:37 TO (B)(6) 2024 AT 04:12:03, THE INFUSION TAKES PLACE WITHOUT ANY APPARENT PROBLEMS. ON (B)(6) 2024: AT 04:12:07, THE INFUSION STARTS @17 ML/H WITH A TOTAL VOLUME INFUSED OF 197.628 ML. AT 07:59:43, THE VOLUME IS RESET TO ZERO AFTER INFUSING 262.010 ML. AT 10:05:38, THE KVO MODE IS AUTOMATICALLY ACTIVATED AT THE END OF THE INFUSION WITH A TOTAL VOLUME INFUSED OF 35.618 ML AND THE FLOWRATE SWITCHED @1.000 ML/H. AT 10:18:55, THE DOWNSTREAM OCCLUSION ALARM STOPPED THE INFUSION, WITH A TOTAL VOLUME INFUSION OF 35.840 ML. AT 10:19:47, THE INFUSION RE-STARTS @17.000 ML/H. FROM (B)(6) 2024 AT 10:19:47 TO (B)(6) 2024: · AT 12:41:44, THE INFUSION TAKES PLACE WITH SEVERAL STARTS AND STOPS BY THE USER AND RESETS OF THE INFUSED VOLUME, BUT WITH NO APPARENT PROBLEMS. ON (B)(6) 2024: AT 12:41:52, THE INFUSION STARTS @20.000 ML/H WITH A TOTAL VOLUME INFUSED OF 0.696 ML. AT 13:02:33, THE AIR ALARM STOPS THE INFUSION. AT 13:23:15, THE DOOR IS OPENED (20 MINUTES AFTER THE AIR ALARM). AT 13:24:54, THE DEVICE IS SWITCHED OFF. AT 13:24:55, THE DEVICE IS SWITCHED ON. AT 13:30:02, THE INFUSION RE-STARTS @20.000 ML/H. AT 13:55:24, THE USER STOPS THE INFUSION WITH A TOTAL VOLUME INFUSED OF 8.442 ML. AT 13:55:27, THE DEVICE IS SWITCHED OFF. THE LOG REVIEW INDICATES THAT THE PUMP WAS NOT ATTENDED TO UNTIL 20 MINUTES AFTER THE ALARM WAS TRIGGERED. WITHOUT ACCESS TO THE DEVICE, DETERMINING THE ROOT CAUSE OF THE ALARM IS NOT POSSIBLE. BASED ON THE INFORMATION PROVIDED, THE CUSTOMER CONFIRMED THAT AN OPERATIONAL ERROR OCCURRED AND THAT THE PUMP HAD ALREADY BEEN RETURNED TO PATIENT USE. AS A RESULT, THE LOCAL TEAM DECIDED NOT TO REMOVE THE PUMP FOR FURTHER INVESTIGATION. IT HAS BEEN CONCLUDED THAT THE PATIENT'S DEATH RESULTED FROM A USER ERROR, SPECIFICALLY THE FAILURE TO RESPOND PROMPTLY TO THE PUMP'S FUNCTIONAL AIR ALARM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21280396&pc=FRN | 3000240707-2025-00035 | 2024-12-30 05:00 | Death | FRESENIUS VIAL S.A.S. | 2025-01-30 05:00 | FRN | VOLUMAT AGILIA BR | Infusion or Flow Problem | Asystole | K121613 |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED: EVENT DATE: 30/12/2024. PREMISES: UTI ADULT STERILE. HOURS: BETWEEN 11:00 AM AND 1:07 PM. "SEVERE PATIENT ON VASOACTIVE DRUG (NORADRENALINE 20 ML/H). IT WAS OBSERVED THAT THE INFUSION WAS NOT INFUSING AS PROGRAMMED IN THE BIC, AT NO TIME DID THE BIC SIGNAL INFUSION FAILURE, BUBBLES AND ALARM, WHEN IT WAS EVIDENCED BY THE ATTENDANT THAT THE PATIENT PRESENTED PCR, EVOLVING TO DEATH."INFUSION PUMP SERIAL NUMBER (B)(6). "INFUSION PUMP TAG HGA-017/ SERIAL NUMBER (B)(6) STOPPED INFUSING NORADRENALINE WITH THE WARNING OF BUBBLES IN THE CIRCUIT, TO THE MULTIPARAMETRIC MONITOR PATIENT WITH MAP 32MMHG, I CHECK CIRCUIT AND DO NOT IDENTIFY BUBBLES, RESTART PUMP, PATIENT ENDS UP EVOLVING TO PCR IN ASYSTOLE, RPC START FOLLOWING ACLS, BUT PATIENT EVOLVED TO DEATH" MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.Event Description: FRESENSIUS KABI RECEIVED THE FOLLOWING REQUEST FOR ADDITIONAL INFORMATION: ON JANUARY 30, 2025, WE RECEIVED MDR 3004548776-2025-00036 (IMPORTER MDR#) FOR A VOLUMAT AGILIA INFUSION PUMP (MODEL NUMBER Z019040, SERIAL NUMBER (B)(6)) STATING A PATIENT WAS ON NORADRENALINE. DURING TREATMENT, IT WAS OBSERVED THAT THE INFUSION WAS NOT INFUSING AS PROGRAMMED AND THE PATIENT WAS DETERIORATING. SPECIFICALLY, THE INFUSION PUMP STOPPED INFUSING WITH THE WARNING OF BUBBLES IN THE CIRCUIT; HOWEVER, AN INSPECTION OF THE CIRCUIT DID NOT SHOW BUBBLES. THE PATIENT EXPIRED. FRESENIUS KABI HAS ADDRESSED THE ADDITIONAL FOLLOW UP QUESTIONS AS FOLLOWS: THE PUMP TRIGGERED A FUNCTIONAL AIR ALARM, WHICH IS A NORMAL RESPONSE WHEN THE AIR SENSOR DETECTS A CERTAIN THRESHOLD OF CUMULATIVE AIR PASSING THROUGH THE LINE. SINCE THE PUMP WAS NOT RETURNED FOR EVALUATION, THE SPECIFIC VALUE FOR CUMULATIVE AIR SET IN THE PUMP'S PARAMETERS CANNOT BE DETERMINED AT THIS TIME. THE FUNCTIONAL AIR ALARM NOTIFICATION IS CLASSIFIED AS A HIGH-PRIORITY ALARM. IT TRIGGERS AN AUDIBLE SOUND, FLASHES A RED LED, AND DISPLAYS THE "AIR ALARM" MESSAGE, WHICH BLINKS ON THE SCREEN. A REVIEW OF FUNCTIONAL AIR ALARM COMPLAINTS FROM (B)(6) 2023 TO (B)(6) 2025 WAS CONDUCTED. THERE HAVE BEEN NO ADDITIONAL INSTANCES OF PATIENT DEATH OR SERIOUS INJURY RESULTING FROM DELAYS IN RESPONDING TO A FUNCTIONAL AIR ALARM. THERE HAVE BEEN NO RECALLS FOR THE VOLUMAT AGILIA INFUSION PUMP RELATED TO REPORTS OF THE FUNCTIONAL AIR ALARM. THE INSTRUCTIONS FOR USE (IFU) CONTAIN SEVERAL REFERENCES REGARDING FUNCTIONAL AIR ALARM WARNINGS AND/OR SAFETY FEATURES, INCLUDING OPERATOR RESPONSE AND/OR PROBLEM RESOLUTION. THE INVESTIGATION AND EVALUATION CAN ONLY BE CONDUCTED THROUGH A LOG REVIEW, AS THE PUMP HAS BEEN RETAINED BY THE HOSPITAL AND REINTRODUCED FOR PATIENT USE. THE LOG REVIEW CLEARLY INDICATES THAT A FUNCTIONAL AIR ALARM OCCURRED AT 13:02:33 ON (B)(6) 2024. ON (B)(6) 2024 AT 12:14:55, THE DEVICE IS SWITCHED ON, ON BATTERY. AT 12:15:37, THE INFUSION STARTS @3 ML/H WITH A PRESSURE THRESHOLD @750 MMHG. FROM 12:15:37 TO (B)(6) 2024 AT 04:12:03, THE INFUSION TAKES PLACE WITHOUT ANY APPARENT PROBLEMS. ON (B)(6) . 2024 AT 04:12:07, THE INFUSION STARTS @17 ML/H WITH A TOTAL VOLUME INFUSED OF 197.628 ML. AT 07:59:43, THE VOLUME IS RESET TO ZERO AFTER INFUSING 262.010 ML. AT 10:05:38, THE KVO MODE IS AUTOMATICALLY ACTIVATED AT THE END OF THE INFUSION WITH A TOTAL VOLUME INFUSED OF 35.618 ML AND THE FLOWRATE SWITCHED @1.000 ML/H. AT 10:18:55, THE DOWNSTREAM OCCLUSION ALARM STOPPED THE INFUSION, WITH A TOTAL VOLUME INFUSION OF 35.840 ML. AT 10:19:47, THE INFUSION RE-STARTS @17.000 ML/H. FROM (B)(6) 2024 AT 10:19:47 TO (B)(6) 2024 · AT 12:41:44, THE INFUSION TAKES PLACE WITH SEVERAL STARTS AND STOPS BY THE USER AND RESETS OF THE INFUSED VOLUME, BUT WITH NO APPARENT PROBLEMS. ON (B)(6) 2024: AT 12:41:52, THE INFUSION STARTS @20.000 ML/H WITH A TOTAL VOLUME INFUSED OF 0.696 ML. AT 13:02:33, THE AIR ALARM STOPS THE INFUSION. AT 13:23:15, THE DOOR IS OPENED (20 MINUTES AFTER THE AIR ALARM). AT 13:24:54, THE DEVICE IS SWITCHED OFF. AT 13:24:55, THE DEVICE IS SWITCHED ON. AT 13:30:02, THE INFUSION RE-STARTS @20.000 ML/H. AT 13:55:24, THE USER STOPS THE INFUSION WITH A TOTAL VOLUME INFUSED OF 8.442 ML. AT 13:55:27, THE DEVICE IS SWITCHED OFF. THE LOG REVIEW INDICATES THAT THE PUMP WAS NOT ATTENDED TO UNTIL 20 MINUTES AFTER THE ALARM WAS TRIGGERED. WITHOUT ACCESS TO THE DEVICE, DETERMINING THE ROOT CAUSE OF THE ALARM IS NOT POSSIBLE. BASED ON THE INFORMATION PROVIDED, THE CUSTOMER CONFIRMED THAT AN OPERATIONAL ERROR OCCURRED AND THAT THE PUMP HAD ALREADY BEEN RETURNED TO PATIENT USE. AS A RESULT, THE LOCAL TEAM DECIDED NOT TO REMOVE THE PUMP FOR FURTHER INVESTIGATION. IT HAS BEEN CONCLUDED THAT THE PATIENT'S DEATH RESULTED FROM A USER ERROR, SPECIFICALLY THE FAILURE TO RESPOND PROMPTLY TO THE PUMP'S FUNCTIONAL AIR ALARM. .